## Reading 1
Reference Clinical Thermometer Readout | Avg Temp. GUN Reading | Avg Temperature Difference | Regressed Temp. GUN Reading | Regressed Temperature Difference | Corrected Values | DIFF after Correction | Avg Temp GUN Reading IDEAL AVG | Avg Temperature Difference IDEAL AVG | Regressed Temp. GUN Reading IDEAL AVG | Regressed Temperature Difference IDEAL AVG | Corrected Values IDEAL AVG | DIFF after Correction | DIFF BW IDEAL and practical
Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F)
94 | 87.278 | 6.722 | 93.38 | 0.62 | 92.66 | 1.34 | 88.2733 | 5.7267 | 94.9908 | -0.99077 | 94.2708 | -0.27077 | 0.99529
94.2 | 87.797 | 6.403 | 94.22 | -0.02 | 93.5 | 0.7 | 88.1867 | 6.0133 | 94.8506 | -0.65061 | 94.1306 | 0.069387 | 0.38965
94.4 | 88.0195 | 6.3805 | 94.58 | -0.18 | 93.86 | 0.54 | 88.35 | 6.05 | 95.1149 | -0.7149 | 94.3949 | 0.0051035 | 0.33051
94.6 | 88.0566 | 6.5434 | 94.64 | -0.04 | 93.92 | 0.68 | 88.5833 | 6.0167 | 95.4925 | -0.89247 | 94.7725 | -0.17247 | 0.52674
94.8 | 88.0875 | 6.7125 | 94.69 | 0.11 | 93.97 | 0.83 | 88.52 | 6.28 | 95.39 | -0.59002 | 94.67 | 0.12998 | 0.43254
95 | 88.2975 | 6.7025 | 95.03 | -0.03 | 94.31 | 0.69 | 88.7933 | 6.2067 | 95.8323 | -0.83233 | 95.1123 | -0.11233 | 0.49576
95.2 | 88.7054 | 6.4946 | 95.69 | -0.49 | 94.97 | 0.23 | 89.0133 | 6.1867 | 96.1884 | -0.98837 | 95.4684 | -0.26837 | 0.30794
95.4 | 89.3727 | 6.0273 | 96.77 | -1.37 | 96.05 | -0.65 | 89.1767 | 6.2233 | 96.4528 | -1.0528 | 95.7328 | -0.33282 | -0.19599
95.6 | 89.6693 | 5.9307 | 97.25 | -1.65 | 96.53 | -0.93 | 89.47 | 6.13 | 96.9275 | -1.3275 | 96.2075 | -0.60749 | -0.19928
95.8 | 89.6013 | 6.1987 | 97.14 | -1.34 | 96.42 | -0.62 | 89.4467 | 6.3533 | 96.8898 | -1.0898 | 96.1698 | -0.36978 | -0.15461
96 | 89.9412 | 6.0588 | 97.69 | -1.69 | 96.97 | -0.97 | 89.56 | 6.44 | 97.0731 | -1.0731 | 96.3531 | -0.35315 | -0.38115
96.2 | 89.8485 | 6.3515 | 97.54 | -1.34 | 96.82 | -0.62 | 89.0733 | 7.1267 | 96.2855 | -0.085478 | 95.5655 | 0.63452 | -0.77517
96.4 | 89.8114 | 6.5886 | 97.48 | -1.08 | 96.76 | -0.36 | 88.9167 | 7.4833 | 96.032 | 0.36796 | 95.312 | 1.088 | -0.89469
96.6 | 90.1574 | 6.4426 | 98.04 | -1.44 | 97.32 | -0.72 | 89.4033 | 7.1967 | 96.8195 | -0.21955 | 96.0995 | 0.50045 | -0.75412
96.8 | 90.695 | 6.105 | 98.91 | -2.11 | 99.24 | -2.44 | 89.2067 | 7.5933 | 96.5014 | 0.29863 | 96.8314 | -0.031371 | -1.4883
97 | 90.7815 | 6.2185 | 99.05 | -2.05 | 99.38 | -2.38 | 89.5667 | 7.4333 | 97.084 | -0.083992 | 97.414 | -0.41399 | -1.2148
97.2 | 90.8989 | 6.3011 | 99.24 | -2.04 | 99.57 | -2.37 | 89.38 | 7.82 | 96.7818 | 0.41816 | 97.1118 | 0.088162 | -1.5189
97.4 | 91.455 | 5.945 | 100.14 | -2.74 | 100.47 | -3.07 | 89.6433 | 7.7567 | 97.208 | 0.19204 | 97.538 | -0.13796 | -1.8117
97.6 | 91.7948 | 5.8052 | 100.69 | -3.09 | 101.02 | -3.42 | 90 | 7.6 | 97.7852 | -0.18524 | 98.1152 | -0.51524 | -1.7948
97.8 | 91.8628 | 5.9372 | 100.8 | -3 | 101.13 | -3.33 | 89.74 | 8.06 | 97.3645 | 0.43554 | 97.6945 | 0.10554 | -2.1228
98 | 92.1656 | 5.8344 | 101.29 | -3.29 | 101.62 | -3.62 | 90.0367 | 7.9633 | 97.8446 | 0.15537 | 98.1746 | -0.17463 | -2.1289
98.2 | 92.9132 | 5.2868 | 102.5 | -4.3 | 102.83 | -4.63 | 90.1133 | 8.0867 | 97.9686 | 0.2314 | 98.2986 | -0.098604 | -2.7999
98.4 | 92.1718 | 6.2282 | 101.3 | -2.9 | 101.63 | -3.23 | 90.2133 | 8.1867 | 98.1304 | 0.26956 | 98.4604 | -0.060443 | -1.9585
98.6 | 92.561 | 6.039 | 101.93 | -3.33 | 102.26 | -3.66 | 90.2267 | 8.3733 | 98.1521 | 0.44787 | 98.4821 | 0.11787 | -2.3343
98.8 | 92.6043 | 6.1957 | 102 | -3.2 | 102.33 | -3.53 | 90.4867 | 8.3133 | 98.5729 | 0.22709 | 98.9029 | -0.10291 | -2.1176
99 | 92.802 | 6.198 | 102.32 | -3.32 | 102.65 | -3.65 | 90.43 | 8.57 | 98.4811 | 0.51885 | 98.8111 | 0.18885 | -2.372
99.2 | 93.2778 | 5.9222 | 103.09 | -3.89 | 103.42 | -4.22 | 90.6533 | 8.5467 | 98.8425 | 0.35747 | 99.1725 | 0.027466 | -2.6245
99.4 | 93.1357 | 6.2643 | 102.86 | -3.46 | 103.19 | -3.79 | 90.81 | 8.59 | 99.0961 | 0.30386 | 99.4261 | -0.026136 | -2.3257
99.6 | 93.3087 | 6.2913 | 103.14 | -3.54 | 103.47 | -3.87 | 90.8667 | 8.7333 | 99.1879 | 0.4121 | 99.5179 | 0.082101 | -2.442
99.8 | 93.284 | 6.516 | 103.1 | -3.3 | 103.43 | -3.63 | 91.1733 | 8.6267 | 99.6841 | 0.1159 | 100.0141 | -0.2141 | -2.1107
100 | 94.0872 | 5.9128 | 104.4 | -4.4 | 104.73 | -4.73 | 90.8867 | 9.1133 | 99.2203 | 0.77973 | 99.5503 | 0.44973 | -3.2005
100.2 | 92.8082 | 7.3918 | 102.33 | -2.13 | 102.66 | -2.46 | 91.1 | 9.1 | 99.5655 | 0.63453 | 99.8955 | 0.30453 | -1.7082
100.4 | 93.5373 | 6.8627 | 103.51 | -3.11 | 103.84 | -3.44 | 91.2933 | 9.1067 | 99.8783 | 0.5217 | 100.2083 | 0.1917 | -2.244
100.6 | 93.4323 | 7.1677 | 103.34 | -2.74 | 103.67 | -3.07 | 91.4467 | 9.1533 | 100.1266 | 0.47344 | 100.4566 | 0.14344 | -1.9856
100.8 | 93.5497 | 7.2503 | 103.53 | -2.73 | 104.51 | -3.71 | 90.9233 | 9.8767 | 99.2795 | 1.5205 | 100.2595 | 0.5405 | -2.6264
101 | 94.0749 | 6.9251 | 104.38 | -3.38 | 105.36 | -4.36 | 91.04 | 9.96 | 99.4684 | 1.5316 | 100.4484 | 0.55163 | -3.0349
101.2 | 94.2232 | 6.9768 | 104.62 | -3.42 | 105.6 | -4.4 | 91.72 | 9.48 | 100.5689 | 0.63113 | 101.5489 | -0.34887 | -2.5032
101.4 | 93.976 | 7.424 | 104.22 | -2.82 | 105.2 | -3.8 | 91.66 | 9.74 | 100.4718 | 0.92823 | 101.4518 | -0.051767 | -2.316
101.6 | 94.6619 | 6.9381 | 105.33 | -3.73 | 106.31 | -4.71 | 91.8633 | 9.7367 | 100.8008 | 0.79921 | 101.7808 | -0.18079 | -2.7986
101.8 | 95.4652 | 6.3348 | 106.63 | -4.83 | 107.61 | -5.81 | 91.86 | 9.94 | 100.7954 | 1.0046 | 101.7754 | 0.024555 | -3.6052
102 | 95.564 | 6.436 | 106.79 | -4.79 | 107.77 | -5.77 | 92.2267 | 9.7733 | 101.3889 | 0.61109 | 102.3689 | -0.36891 | -3.3373
102.2 | 95.9409 | 6.2591 | 107.4 | -5.2 | 108.38 | -6.18 | 91.95 | 10.25 | 100.9411 | 1.2589 | 101.9211 | 0.2789 | -3.9909
102.4 | 96.0089 | 6.3911 | 107.51 | -5.11 | 108.49 | -6.09 | 92.5067 | 9.8933 | 101.8421 | 0.55794 | 102.8221 | -0.42206 | -3.5022
102.6 | 95.9533 | 6.6467 | 107.42 | -4.82 | 107.18 | -4.58 | 92.8467 | 9.7533 | 102.3923 | 0.20769 | 102.1523 | 0.44769 | -3.1066
102.8 | 96.3302 | 6.4698 | 108.03 | -5.23 | 107.79 | -4.99 | 93.1267 | 9.6733 | 102.8455 | -0.04546 | 102.6055 | 0.19454 | -3.2035
103 | 96.5836 | 6.4164 | 108.44 | -5.44 | 108.2 | -5.2 | 92.7133 | 10.2867 | 102.1764 | 0.82358 | 101.9364 | 1.0636 | -3.8703
103.2 | 96.6206 | 6.5794 | 108.5 | -5.3 | 108.26 | -5.06 | 93.2233 | 9.9767 | 103.0018 | 0.1982 | 102.7618 | 0.4382 | -3.3973
103.4 | 96.6206 | 6.7794 | 108.5 | -5.1 | 108.26 | -4.86 | 93.5933 | 9.8067 | 103.6006 | -0.2006 | 103.3606 | 0.039399 | -3.0273
103.6 | 96.8245 | 6.7755 | 108.83 | -5.23 | 108.59 | -4.99 | 93.7767 | 9.8233 | 103.8974 | -0.29741 | 103.6574 | -0.057414 | -3.0478
103.8 | 97.0408 | 6.7592 | 109.18 | -5.38 | 108.94 | -5.14 | 93.96 | 9.84 | 104.1941 | -0.39406 | 103.9541 | -0.15406 | -3.0808
104 | 96.4785 | 7.5215 | 108.27 | -4.27 | 108.03 | -4.03 | 94.22 | 9.78 | 104.6148 | -0.61485 | 104.3748 | -0.37485 | -2.2585
104.2 | 98.3507 | 5.8493 | 111.3 | -7.1 | 111.06 | -6.86 | 94.43 | 9.77 | 104.9547 | -0.75471 | 104.7147 | -0.51471 | -3.9207
104.4 | 97.9368 | 6.4632 | 110.63 | -6.23 | 110.39 | -5.99 | 94.2733 | 10.1267 | 104.7011 | -0.30111 | 104.4611 | -0.061106 | -3.6635
104.6 | 98.0109 | 6.5891 | 110.75 | -6.15 | 110.51 | -5.91 | 94.5533 | 10.0467 | 105.1543 | -0.55426 | 104.9143 | -0.31426 | -3.4576
104.8 | 98.1654 | 6.6346 | 111 | -6.2 | 110.76 | -5.96 | 94.7333 | 10.0667 | 105.4456 | -0.64557 | 105.2056 | -0.40557 | -3.4321
105 | 98.4743 | 6.5257 | 111.5 | -6.5 | 111.26 | -6.26 | 94.4633 | 10.5367 | 105.0086 | -0.0086001 | 104.7686 | 0.2314 | -4.011
105.2 | 98.8142 | 6.3858 | 112.05 | -6.85 | 111.81 | -6.61 | 94.9067 | 10.2933 | 105.7262 | -0.52619 | 105.4862 | -0.28619 | -3.9075
105.4 | 98.7524 | 6.6476 | 111.95 | -6.55 | 111.71 | -6.31 | 94.9667 | 10.4333 | 105.8233 | -0.4233 | 105.5833 | -0.1833 | -3.7857
105.6 | 98.4496 | 7.1504 | 111.46 | -5.86 | 111.33 | -5.73 | 94.8633 | 10.7367 | 105.656 | -0.055956 | 105.526 | 0.074044 | -3.5863
105.8 | 98.6906 | 7.1094 | 111.85 | -6.05 | 111.72 | -5.92 | 94.9067 | 10.8933 | 105.7262 | 0.073806 | 105.5962 | 0.20381 | -3.7839
106 | 98.8574 | 7.1426 | 112.12 | -6.12 | 111.99 | -5.99 | 95.2167 | 10.7833 | 106.2279 | -0.2279 | 106.0979 | -0.097895 | -3.6407
106.2 | 99.0675 | 7.1325 | 112.46 | -6.26 | 112.33 | -6.13 | 95.34 | 10.86 | 106.4274 | -0.22744 | 106.2974 | -0.097443 | -3.7275
106.4 | 99.1787 | 7.2213 | 112.64 | -6.24 | 112.51 | -6.11 | 95.4633 | 10.9367 | 106.627 | -0.22699 | 106.497 | -0.09699 | -3.7154
106.6 | 99.463 | 7.137 | 113.1 | -6.5 | 112.97 | -6.37 | 95.4733 | 11.1267 | 106.6432 | -0.043174 | 106.5132 | 0.086826 | -3.9897
106.8 | 99.7101 | 7.0899 | 113.5 | -6.7 | 113.37 | -6.57 | 95.49 | 11.31 | 106.6702 | 0.1298 | 106.5402 | 0.2598 | -4.2201
107 | 99.5618 | 7.4382 | 113.26 | -6.26 | 113.13 | -6.13 | 95.4 | 11.6 | 106.5245 | 0.47545 | 106.3945 | 0.60545 | -4.1618
107.2 | 100.0747 | 7.1253 | 114.09 | -6.89 | 113.96 | -6.76 | 96.3267 | 10.8733 | 108.0243 | -0.82431 | 107.8943 | -0.69431 | -3.748
107.4 | 99.9326 | 7.4674 | 113.86 | -6.46 | 113.73 | -6.33 | 95.8933 | 11.5067 | 107.3229 | 0.077102 | 107.1929 | 0.2071 | -4.0393
107.6 | 100.0994 | 7.5006 | 114.13 | -6.53 | 114 | -6.4 | 96.1567 | 11.4433 | 107.7492 | -0.14918 | 107.6192 | -0.019182 | -3.9427
107.8 | 100.3033 | 7.4967 | 114.46 | -6.66 | 114.33 | -6.53 | 96.4533 | 11.3467 | 108.2292 | -0.4292 | 108.0992 | -0.2992 | -3.85
108 | 101.261 | 6.739 | 116.01 | -8.01 | 115.88 | -7.88 | 96.5033 | 11.4967 | 108.3101 | -0.31012 | 108.1801 | -0.18012 | -4.7577
ŷ = 1.61839X - 47.86986

## Reading 2
Reference Clinical Thermometer Readout | Temperature Gun Readout(2-3cm) | Temperature Difference | Lab Thermometer Readout
Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg C)
94 | 92.16 | 1.84
94.2 | 93.18 | 1.02
94.4 | 93.45 | 0.95
94.6 | 92.9 | 1.7
94.8 | 93.14 | 1.66
95 | 93.56 | 1.44
95.2 | 93.26 | 1.94
95.4 | 94.2 | 1.2
95.6 | 94.52 | 1.08
95.8 | 94.6 | 1.2
96 | 94.83 | 1.17
96.2 | 95.05 | 1.15
96.4 | 95.56 | 0.84
96.6 | 95.46 | 1.14
96.8 | 95.64 | 1.16
97 | 95.88 | 1.12
97.2 | 96.19 | 1.01
97.4 | 97.11 | 0.29
97.6 | 98.25 | -0.65
97.8 | 98.38 | -0.58
98 | 99.01 | -1.01
98.2 | 99.19 | -0.99
98.4 | 99.64 | -1.24
98.6 | 100.03 | -1.43
98.8 | 100.72 | -1.92
99 | 100.64 | -1.64
99.2 | 100.93 | -1.73
99.4 | 101.1 | -1.7
99.6 | 101.54 | -1.94
99.8 | 101.61 | -1.81
100 | 101.59 | -1.59
100.2 | 101.83 | -1.63
100.4 | 102.48 | -2.08
100.6 | 103.9 | -3.3
100.8 | 103.46 | -2.66
101 | 103.94 | -2.94
101.2 | 103.56 | -2.36
101.4 | 104.04 | -2.64
101.6 | 104.94 | -3.34
101.8 | 105.45 | -3.65
102 | 105.6 | -3.6
102.2 | 106.1 | -3.9
102.4 | 106.45 | -4.05
102.6 | 107.9 | -5.3
102.8 | 106.94 | -4.14
103 | 107.35 | -4.35
103.2 | 107.81 | -4.61
103.4 | 107.39 | -3.99
103.6 | 108.11 | -4.51
103.8 | 108.36 | -4.56
104 | 108.49 | -4.49
104.2 | 109.4 | -5.2
104.4 | 109.91 | -5.51
104.6 | 110.2 | -5.6
104.8 | 110.36 | -5.56
105 | 110.56 | -5.56
105.2 | 110.99 | -5.79
105.4 | 110.53 | -5.13
105.6 | 111.61 | -6.01
105.8 | 111.85 | -6.05
106 | 113.12 | -7.12
106.2 | 113.91 | -7.71
106.4 | 113.46 | -7.06
106.6 | 113.96 | -7.36
106.8 | 114.34 | -7.54
107 | 114.61 | -7.61
107.2 | 114.92 | -7.72
107.4 | 114.46 | -7.06
107.6 | 114.98 | -7.38
107.8 | 115.32 | -7.52
108 | 115.75 | -7.75

## NEW 1
Reference Clinical Thermometer Readout | Temperature Gun Readout(2-3cm) | AMBIENT Readout | Temperature Difference | Digital
Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F)
94 |  |  | 94 | 94.5
94.2 | 87.96 |  | 6.24
94.4 |  |  | 94.4
94.6 |  |  | 94.6
94.8 | 88.47 |  | 6.33
95 | 88.34 |  | 6.66
95.2 |  |  | 95.2
95.4 |  |  | 95.4
95.6 |  |  | 95.6
95.8 |  |  | 95.8
96 |  |  | 96
96.2 |  |  | 96.2
96.4 |  |  | 96.4
96.6 |  |  | 96.6
96.8 |  |  | 96.8
97 | 88.84 |  | 8.16
97.2 |  |  | 97.2
97.4 |  |  | 97.4
97.6 |  |  | 97.6
97.8 |  |  | 97.8
98 |  |  | 98
98.2 |  |  | 98.2
98.4 | 91.69 |  | 6.71
98.6 |  |  | 98.6
98.8 |  |  | 98.8
99 |  |  | 99
99.2 |  |  | 99.2
99.4 |  |  | 99.4
99.6 |  |  | 99.6
99.8 |  |  | 99.8
100 |  |  | 100
100.2 |  |  | 100.2
100.4 |  |  | 100.4
100.6 |  |  | 100.6
100.8 |  |  | 100.8
101 |  |  | 101
101.2 |  |  | 101.2
101.4 |  |  | 101.4
101.6 |  |  | 101.6
101.8 |  |  | 101.8
102 |  |  | 102
102.2 |  |  | 102.2
102.4 |  |  | 102.4
102.6 |  |  | 102.6
102.8 |  |  | 102.8
103 |  |  | 103
103.2 |  |  | 103.2
103.4 |  |  | 103.4
103.6 |  |  | 103.6
103.8 |  |  | 103.8
104 |  |  | 104
104.2 |  |  | 104.2
104.4 |  |  | 104.4
104.6 |  |  | 104.6
104.8 |  |  | 104.8
105 |  |  | 105
105.2 |  |  | 105.2
105.4 |  |  | 105.4
105.6 |  |  | 105.6
105.8 |  |  | 105.8
106 |  |  | 106
106.2 |  |  | 106.2
106.4 |  |  | 106.4
106.6 |  |  | 106.6
106.8 |  |  | 106.8
107 |  |  | 107
107.2 |  |  | 107.2
107.4 |  |  | 107.4
107.6 |  |  | 107.6
107.8 |  |  | 107.8
108 |  |  | 108