## Sheet1
Reference Clinical Thermometer Readout (Y var.) | SET-1 |  |  | SET-2 |  |  | SET-3 |  |  | SET-AVERAGE |  |  |  |  | Average Regression
| Reference Water Temperature (Sensor) | Calibrating contactless sensor Readout (2-3 cm) | Calibrating contactless sensor AMBIENT Readout | Reference Water Temperature (Sensor) | Calibrating contactless sensor Readout (2-3 cm) | Calibrating contactless sensor AMBIENT Readout | Reference Water Temperature (Sensor) | Calibrating contactless sensor Readout (2-3 cm) | Calibrating contactless sensor AMBIENT Readout | Reference Water Temperature (Sensor) | Calibrating contactless sensor Readout (2-3 cm) (X var.) | Calibrating contactless sensor AMBIENT Readout | DIFF between Water temp and clinical | DIFF bw clinical and sensor | Regressed values Calibrating sensor (ŷ = 1.02751X - 0.34754) | Diff between Clinical and Regressed Calibrating sensor
Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) |  | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F)
94 | 92.41 | 91.58 | 82.35 | 92.86 | 91.79 | 84.18 | 92.41 | 91.35 | 83.46 | 92.56 | 91.5733 | 83.33 | 1.44 | 2.4267 | 93.745 | 0.25502
94.4 | 92.86 | 92.57 |  | 92.97 | 92.26 |  | 93.65 | 92.95 |  | 93.16 | 92.5933 |  | 1.24 | 1.8067 | 94.793 | -0.39304
94.8 | 93.31 | 93.02 |  | 93.54 | 92.65 |  | 93.87 | 93.42 |  | 93.5733 | 93.03 |  | 1.2267 | 1.77 | 95.2417 | -0.44172
95 | 93.43 | 93.13 |  | 93.76 | 92.77 |  | 94.1 | 93.51 |  | 93.7633 | 93.1367 |  | 1.2367 | 1.8633 | 95.3513 | -0.35132
95.4 | 93.87 | 93.14 |  | 94.55 | 93.25 |  | 94.55 | 93.65 |  | 94.3233 | 93.3467 |  | 1.0767 | 2.0533 | 95.5671 | -0.16709
95.8 | 94.21 | 93.35 |  | 94.55 | 93.45 |  | 94.89 | 93.96 |  | 94.55 | 93.5867 |  | 1.25 | 2.2133 | 95.8137 | -0.013696
96 | 94.55 | 94.19 |  | 94.66 | 93.52 |  | 95 | 94.11 |  | 94.7367 | 93.94 |  | 1.2633 | 2.06 | 96.1767 | -0.17675
96.4 | 94.77 | 93.78 |  | 94.89 | 93.93 |  | 95.45 | 94.9 |  | 95.0367 | 94.2033 |  | 1.3633 | 2.1967 | 96.4473 | -0.047327
96.8 | 95.11 | 94.65 |  | 95 | 94.22 | 84.58 | 95.9 | 95.15 |  | 95.3367 | 94.6733 |  | 1.4633 | 2.1267 | 96.9303 | -0.13026
97 | 95.34 | 94.48 |  | 95.22 | 94.46 |  | 96.01 | 95 |  | 95.5233 | 94.6467 |  | 1.4767 | 2.3533 | 96.9029 | 0.097144
97.4 | 95.79 | 94.98 |  | 95.56 | 94.56 |  | 96.35 | 95.45 |  | 95.9 | 94.9967 |  | 1.5 | 2.4033 | 97.2625 | 0.13752
97.8 | 95.9 | 94.95 |  | 95.9 | 94.95 |  | 96.91 | 95.81 |  | 96.2367 | 95.2367 |  | 1.5633 | 2.5633 | 97.5091 | 0.29091
98 | 96.24 | 95.9 |  | 95.79 | 94.99 |  | 96.91 | 96.15 |  | 96.3133 | 95.68 |  | 1.6867 | 2.32 | 97.9646 | 0.035383
98.2 | 96.69 | 95.98 |  | 96.01 | 95.23 |  | 97.25 | 96.3 |  | 96.65 | 95.8367 |  | 1.55 | 2.3633 | 98.1256 | 0.074407
98.4 | 96.57 | 96.12 |  | 96.24 | 95.65 |  | 97.47 | 96.47 |  | 96.76 | 96.08 |  | 1.64 | 2.32 | 98.3756 | 0.024379
98.6 | 96.8 | 96.22 |  | 96.35 | 95.75 |  | 97.7 | 96.6 |  | 96.95 | 96.19 |  | 1.65 | 2.41 | 98.4886 | 0.11135
98.8 | 97.02 | 96.49 |  | 96.8 | 96.01 |  | 97.81 | 96.75 | 85.77 | 97.21 | 96.4167 |  | 1.59 | 2.3833 | 98.7215 | 0.078451
99 | 97.14 | 96.65 |  | 96.91 | 96.11 |  | 98.04 | 96.99 |  | 97.3633 | 96.5833 |  | 1.6367 | 2.4167 | 98.8928 | 0.1072
99.2 | 97.36 | 96.84 |  | 97.02 | 96.4 | 84.72 | 98.3 | 97.15 |  | 97.56 | 96.7967 |  | 1.64 | 2.4033 | 99.112 | 0.087997
99.4 | 97.59 | 96.93 | 86 | 97.47 | 96.75 |  | 98.49 | 97.45 |  | 97.85 | 97.0433 |  | 1.55 | 2.3567 | 99.3655 | 0.034545
99.6 | 97.81 | 97.34 |  | 97.65 | 96.96 |  | 98.64 | 97.64 | 84.51 | 98.0333 | 97.3133 |  | 1.5667 | 2.2867 | 99.6429 | -0.042883
99.8 | 98.04 | 97.61 |  | 97.92 | 97.2 |  | 98.81 | 97.78 |  | 98.2567 | 97.53 |  | 1.5433 | 2.27 | 99.8655 | -0.06551
100 | 98.26 | 97.4 |  | 98.15 | 97.45 |  | 99.05 | 97.96 |  | 98.4867 | 97.6033 |  | 1.5133 | 2.3967 | 99.9409 | 0.059139
100.2 | 98.37 | 97.7 |  | 98.37 | 97.7 | 87.06 | 99.16 | 98.1 |  | 98.6333 | 97.8333 |  | 1.5667 | 2.3667 | 100.1772 | 0.022812
100.4 | 98.6 | 97.8 |  | 98.71 | 97.78 |  | 99.27 | 98.22 |  | 98.86 | 97.9333 |  | 1.54 | 2.4667 | 100.2799 | 0.12006
100.6 | 98.82 | 98.23 |  | 98.82 | 97.83 |  | 99.35 | 98.31 |  | 98.9967 | 98.1233 |  | 1.6033 | 2.4767 | 100.4752 | 0.12483
100.8 | 98.39 | 98.45 |  | 99.27 | 98.25 |  | 99.5 | 98.35 |  | 99.0533 | 98.35 |  | 1.7467 | 2.45 | 100.7081 | 0.091932
101 | 99.16 | 98.54 |  | 99.5 | 98.6 |  | 99.72 | 98.65 |  | 99.46 | 98.5967 |  | 1.54 | 2.4033 | 100.9615 | 0.038479
101.2 | 99.39 | 98.33 |  | 99.61 | 99.02 |  | 99.95 | 98.72 |  | 99.65 | 98.69 |  | 1.55 | 2.51 | 101.0574 | 0.14258
101.4 | 99.5 | 98.85 |  | 99.72 | 99.45 |  | 100.17 | 98.85 |  | 99.7967 | 99.05 |  | 1.6033 | 2.35 | 101.4273 | -0.027325
101.6 | 99.84 | 99.22 |  | 99.95 | 99.6 |  | 100.29 | 99.03 |  | 100.0267 | 99.2833 |  | 1.5733 | 2.3167 | 101.6671 | -0.067078
101.8 | 99.95 | 99.15 |  | 100.3 | 99.32 |  | 100.51 | 99.23 |  | 100.2533 | 99.2333 |  | 1.5467 | 2.5667 | 101.6157 | 0.1843
102 | 100.17 | 99.45 |  | 100.17 | 99.5 |  | 100.62 | 99.45 |  | 100.32 | 99.4667 |  | 1.68 | 2.5333 | 101.8555 | 0.14455
102.2 | 100.4 | 99.58 |  | 100.4 | 99.69 |  | 100.74 | 99.67 |  | 100.5133 | 99.6467 |  | 1.6867 | 2.5533 | 102.0404 | 0.15959
102.4 | 100.62 | 99.81 |  | 100.51 | 99.84 |  | 101.41 | 99.9 |  | 100.8467 | 99.85 |  | 1.5533 | 2.55 | 102.2493 | 0.15067
102.6 | 101.41 | 100.45 |  | 100.96 | 100.1 | 88.14 | 101.52 | 100.3 |  | 101.2967 | 100.2833 |  | 1.3033 | 2.3167 | 102.6946 | -0.094588
102.8 | 101.4 | 100.6 |  | 101.19 | 100.26 |  | 101.64 | 100.45 |  | 101.41 | 100.4367 |  | 1.39 | 2.3633 | 102.8521 | -0.052139
103 | 101.64 | 100.9 |  | 101.3 | 100.45 |  | 101.75 | 100.53 |  | 101.5633 | 100.6267 |  | 1.4367 | 2.3733 | 103.0474 | -0.047366
103.4 | 102.09 | 101.25 |  | 101.52 | 100.57 |  | 101.89 | 100.85 |  | 101.8333 | 100.89 |  | 1.5667 | 2.51 | 103.3179 | 0.082056
103.8 | 102.54 | 101.45 |  | 102.2 | 100.85 |  | 102.42 | 101.13 |  | 102.3867 | 101.1433 |  | 1.4133 | 2.6567 | 103.5782 | 0.22175
104 | 102.76 | 101.62 |  | 102.31 | 101.29 |  | 102.6 | 101.49 |  | 102.5567 | 101.4667 |  | 1.4433 | 2.5333 | 103.9105 | 0.089525
104.4 | 103.1 | 102.65 |  | 102.54 | 101.43 | 88.5 | 103.21 | 101.97 |  | 102.95 | 102.0167 |  | 1.45 | 2.3833 | 104.4756 | -0.075605
104.8 | 103.44 | 102.72 |  | 103.21 | 101.85 |  | 103.32 | 102.2 |  | 103.3233 | 102.2567 |  | 1.4767 | 2.5433 | 104.7222 | 0.077792
105 | 103.77 | 102.84 | 87.75 | 103.32 | 102.36 |  | 103.44 | 102.45 |  | 103.51 | 102.55 |  | 1.49 | 2.45 | 105.0236 | -0.02361
105.4 | 104.22 | 103.05 |  | 103.77 | 102.74 |  | 103.89 | 102.75 |  | 103.96 | 102.8467 |  | 1.44 | 2.5533 | 105.3284 | 0.071562
105.8 | 104.45 | 103.78 |  | 104.05 | 103.15 |  | 104.34 | 102.94 |  | 104.28 | 103.29 |  | 1.52 | 2.51 | 105.784 | 0.016032
106 | 104.9 | 104.15 |  | 104.22 | 103.22 |  | 104.56 | 103.15 |  | 104.56 | 103.5067 |  | 1.44 | 2.4933 | 106.0066 | -0.0065951
106.4 | 105.12 | 104.64 |  | 104.56 | 103.75 |  | 105.01 | 103.75 |  | 104.8967 | 104.0467 |  | 1.5033 | 2.3533 | 106.5615 | -0.16145
106.8 | 105.01 | 104.68 |  | 104.67 | 103.99 |  | 105.35 | 104.15 |  | 105.01 | 104.2733 |  | 1.79 | 2.5267 | 106.7944 | 0.0056473
107 | 105.12 | 104.77 |  | 104.9 | 104.25 |  | 105.69 | 104.6 |  | 105.2367 | 104.54 |  | 1.7633 | 2.46 | 107.0684 | -0.068355
107.4 | 105.46 | 104.86 |  | 105.46 | 105.13 |  | 106.14 | 105.09 |  | 105.6867 | 105.0267 |  | 1.7133 | 2.3733 | 107.5684 | -0.16841
107.8 | 105.8 | 105.26 |  | 105.57 | 105.58 |  | 106.25 | 105.66 | 88.25 | 105.8733 | 105.5 |  | 1.9267 | 2.3 | 108.0548 | -0.25477
108 | 106.02 | 105.45 |  | 106.14 | 105.85 | 90 | 106.29 | 105.69 |  | 106.15 | 105.6633 |  | 1.85 | 2.3367 | 108.2226 | -0.22259
[11 empty] | AVERAGE OFFSET of |  | 1.524 | 2.365
ŷ = 1.02751X - 0.34754