## Sheet1
| SET-1 |  | SET-2 |  | SET-3 |  | SET-4 |  | SET-5 |  | SET-6 |  | SET-7
Reference Clinical Thermometer Readout | Calibrating contactless sensor Readout (2-3 cm) | Calibrating contactless sensor AMBIENT Readout | Calibrating contactless sensor Readout (2-3 cm) | Calibrating contactless sensor AMBIENT Readout | Calibrating contactless sensor Readout (2-3 cm) | Calibrating contactless sensor AMBIENT Readout | Calibrating contactless sensor Readout (2-3 cm) | Calibrating contactless sensor AMBIENT Readout | Calibrating contactless sensor Readout (2-3 cm) | Calibrating contactless sensor AMBIENT Readout | Calibrating contactless sensor Readout (2-3 cm) | Calibrating contactless sensor AMBIENT Readout | Calibrating contactless sensor Readout (2-3 cm) | Calibrating contactless sensor AMBIENT Readout | AVERAGE Grp1 | AVERAGE G2 | AVERAGE DIFF G1&G2 | AVERAGE AMBIENT TEMP. G2 | AVG DIFF G1 | AVG DIFF G2
Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F)
94 | 91.58 | 82.35 | 91.79 | 84.18 | 91.35 | 83.46 | 91.19 | 94.23 | 90.06 | 93.9 | 91.43 | 95.27 | 91.2 | 94.59 | 91.5733 | 90.97 | 0.60333 | 94.4975 | 2.4267 | 3.03
94.4 | 92.57 |  | 92.26 |  | 92.95 |  | 91.75 | 94.33 | 90.55 | 94.12 | 91.89 | 95.34 | 91.44 | 94.66 | 92.5933 | 91.4075 | 1.1858 | 94.6125 | 1.8067 | 2.9925
94.8 | 93.02 |  | 92.65 |  | 93.42 |  | 92.15 | 94.37 | 91.09 | 94.15 | 92.06 | 95.34 | 91.94 | 94.8 | 93.03 | 91.81 | 1.22 | 94.665 | 1.77 | 2.99
95 | 93.13 |  | 92.77 |  | 93.51 |  | 92.28 | 94.37 | 91.4 | 94.3 | 92.26 | 95.41 | 92.21 | 94.77 | 93.1367 | 92.0375 | 1.0992 | 94.7125 | 1.8633 | 2.9625
95.4 | 93.14 |  | 93.25 |  | 93.65 |  | 92.75 | 94.51 | 92.11 | 94.48 | 92.87 | 95.49 | 92.33 | 94.8 | 93.3467 | 92.515 | 0.83167 | 94.82 | 2.0533 | 2.885
95.8 | 93.35 |  | 93.45 |  | 93.96 |  | 93.15 | 94.55 | 92.51 | 94.48 | 93.14 | 95.56 | 92.69 | 94.95 | 93.5867 | 92.8725 | 0.71417 | 94.885 | 2.2133 | 2.9275
96 | 94.19 |  | 93.52 |  | 94.11 |  | 93.36 | 94.66 | 92.71 | 94.55 | 93.24 | 95.56 | 93.06 | 94.98 | 93.94 | 93.0925 | 0.8475 | 94.9375 | 2.06 | 2.9075
96.4 | 93.78 |  | 93.93 |  | 94.9 |  | 93.65 | 94.66 | 93.23 | 94.77 | 93.38 | 95.67 | 93.45 | 95.05 | 94.2033 | 93.4275 | 0.77583 | 95.0375 | 2.1967 | 2.9725
96.8 | 94.65 |  | 94.22 | 84.58 | 95.15 |  | 96.94 | 94.66 | 93.47 | 94.8 | 93.56 | 95.77 | 93.81 | 95.16 | 94.6733 | 94.445 | 0.22833 | 95.0975 | 2.1267 | 2.355
97 | 94.48 |  | 94.46 |  | 95 |  | 94.17 | 94.73 | 93.55 | 94.8 | 93.82 | 95.81 | 94.17 | 95.27 | 94.6467 | 93.9275 | 0.71917 | 95.1525 | 2.3533 | 3.0725
97.4 | 94.98 |  | 94.56 |  | 95.45 |  | 94.35 | 94.8 | 93.9 | 94.91 | 94.44 | 95.88 | 94.35 | 95.34 | 94.9967 | 94.26 | 0.73667 | 95.2325 | 2.4033 | 3.14
97.8 | 94.95 |  | 94.95 |  | 95.81 |  | 94.69 | 94.91 | 94.43 | 94.98 | 94.76 | 95.92 | 94.74 | 95.45 | 95.2367 | 94.655 | 0.58167 | 95.315 | 2.5633 | 3.145
98 | 95.9 |  | 94.99 |  | 96.15 |  | 95.07 | 94.98 | 94.86 | 94.98 | 95.02 | 95.99 | 94.95 | 95.52 | 95.68 | 94.975 | 0.705 | 95.3675 | 2.32 | 3.025
98.2 | 95.98 |  | 95.23 |  | 96.3 |  | 95.23 | 95.05 | 94.99 | 95.02 | 95.17 | 96.03 | 95.16 | 95.56 | 95.8367 | 95.1375 | 0.69917 | 95.415 | 2.3633 | 3.0625
98.4 | 96.12 |  | 95.65 |  | 96.47 |  | 95.36 | 95.05 | 95.15 | 94.98 | 95.38 | 96.21 | 95.37 | 95.59 | 96.08 | 95.315 | 0.765 | 95.4575 | 2.32 | 3.085
98.6 | 96.22 |  | 95.75 |  | 96.6 |  | 95.55 | 95.13 | 95.28 | 95.05 | 95.62 | 96.17 | 95.5 | 95.67 | 96.19 | 95.4875 | 0.7025 | 95.505 | 2.41 | 3.1125
98.8 | 96.49 |  | 96.01 |  | 96.75 | 85.77 | 95.7 | 95.16 | 95.45 | 95.13 | 95.8 | 96.21 | 95.67 | 95.7 | 96.4167 | 95.655 | 0.76167 | 95.55 | 2.3833 | 3.145
99 | 96.65 |  | 96.11 |  | 96.99 |  | 95.68 | 95.16 | 95.63 | 95.16 | 95.93 | 96.24 | 95.74 | 95.7 | 96.5833 | 95.745 | 0.83833 | 95.565 | 2.4167 | 3.255
99.2 | 96.84 |  | 96.4 | 84.72 | 97.15 |  | 95.86 | 95.16 | 95.81 | 95.16 | 96.15 | 96.24 | 95.93 | 95.77 | 96.7967 | 95.9375 | 0.85917 | 95.5825 | 2.4033 | 3.2625
99.4 | 96.93 | 86 | 96.75 |  | 97.45 |  | 96.56 | 95.2 | 95.92 | 95.27 | 96.33 | 96.35 | 96.16 | 95.81 | 97.0433 | 96.2425 | 0.80083 | 95.6575 | 2.3567 | 3.1575
99.6 | 97.34 |  | 96.96 |  | 97.64 | 84.51 | 96.84 | 95.31 | 96.06 | 95.31 | 96.48 | 96.35 | 96.26 | 95.77 | 97.3133 | 96.41 | 0.90333 | 95.685 | 2.2867 | 3.19
99.8 | 97.61 |  | 97.2 |  | 97.78 |  | 96.9 | 95.31 | 96.21 | 95.34 | 96.86 | 96.42 | 96.52 | 95.81 | 97.53 | 96.6225 | 0.9075 | 95.72 | 2.27 | 3.1775
100 | 97.4 |  | 97.45 |  | 97.96 |  | 96.86 | 95.34 | 96.49 | 95.41 | 97 | 96.42 | 96.74 | 95.85 | 97.6033 | 96.7725 | 0.83083 | 95.755 | 2.3967 | 3.2275
100.2 | 97.7 |  | 97.7 | 87.06 | 98.1 |  | 97.15 | 95.34 | 96.63 | 95.52 | 96.93 | 96.53 | 96.89 | 95.92 | 97.8333 | 96.9 | 0.93333 | 95.8275 | 2.3667 | 3.3
100.4 | 97.8 |  | 97.78 |  | 98.22 |  | 97.35 | 95.34 | 96.88 | 95.52 | 97.14 | 96.53 | 97.09 | 95.95 | 97.9333 | 97.115 | 0.81833 | 95.835 | 2.4667 | 3.285
100.6 | 98.23 |  | 97.83 |  | 98.31 |  | 97.48 | 95.41 | 97.07 | 95.59 | 97.23 | 96.53 | 97.26 | 96.03 | 98.1233 | 97.26 | 0.86333 | 95.89 | 2.4767 | 3.34
100.8 | 98.45 |  | 98.25 |  | 98.35 |  | 97.65 | 95.52 | 97.28 | 95.59 | 97.45 | 96.64 | 97.33 | 95.99 | 98.35 | 97.4275 | 0.9225 | 95.935 | 2.45 | 3.3725
101 | 98.54 |  | 98.6 |  | 98.65 |  | 97.88 | 95.41 | 97.52 | 95.56 | 97.65 | 96.67 | 97.65 | 96.1 | 98.5967 | 97.675 | 0.92167 | 95.935 | 2.4033 | 3.325
101.2 | 98.33 |  | 99.02 |  | 98.72 |  | 98.09 | 95.49 | 97.65 | 95.67 | 97.88 | 96.64 | 97.85 | 96.13 | 98.69 | 97.8675 | 0.8225 | 95.9825 | 2.51 | 3.3325
101.4 | 98.85 |  | 99.45 |  | 98.85 |  | 98.4 | 95.52 | 97.85 | 95.77 | 98.02 | 96.71 | 98.07 | 96.17 | 99.05 | 98.085 | 0.965 | 96.0425 | 2.35 | 3.315
101.6 | 99.22 |  | 99.6 |  | 99.03 |  | 98.53 | 95.42 | 98 | 95.7 | 98.18 | 96.75 | 98.16 | 96.17 | 99.2833 | 98.2175 | 1.0658 | 96.01 | 2.3167 | 3.3825
101.8 | 99.15 |  | 99.32 |  | 99.23 |  | 98.74 | 95.56 | 98.15 | 95.81 | 98.24 | 96.78 | 98.45 | 96.21 | 99.2333 | 98.395 | 0.83833 | 96.09 | 2.5667 | 3.405
102 | 99.45 |  | 99.5 |  | 99.45 |  | 99.03 | 95.7 | 98.4 | 95.85 | 98.65 | 96.78 | 98.62 | 96.21 | 99.4667 | 98.675 | 0.79167 | 96.135 | 2.5333 | 3.325
102.2 | 99.58 |  | 99.69 |  | 99.67 |  | 99.15 | 95.7 | 98.65 | 95.88 | 98.81 | 96.82 | 98.86 | 96.21 | 99.6467 | 98.8675 | 0.77917 | 96.1525 | 2.5533 | 3.3325
102.4 | 99.81 |  | 99.84 |  | 99.9 |  | 99.29 | 95.85 | 98.75 | 95.92 | 99.02 | 96.93 | 99.1 | 96.35 | 99.85 | 99.04 | 0.81 | 96.2625 | 2.55 | 3.36
102.6 | 100.45 |  | 100.1 | 88.14 | 100.3 |  | 99.45 | 95.85 | 98.92 | 96.13 | 99.19 | 96.96 | 99.24 | 96.42 | 100.2833 | 99.2 | 1.0833 | 96.34 | 2.3167 | 3.4
102.8 | 100.6 |  | 100.26 |  | 100.45 |  | 99.76 | 95.85 | 99.12 | 96.1 | 99.39 | 96.96 | 99.42 | 96.53 | 100.4367 | 99.4225 | 1.0142 | 96.36 | 2.3633 | 3.3775
103 | 100.9 |  | 100.45 |  | 100.53 |  | 99.6 | 95.81 | 99.37 | 95.99 | 99.45 | 97.07 | 99.59 | 96.57 | 100.6267 | 99.5025 | 1.1242 | 96.36 | 2.3733 | 3.4975
103.4 | 101.25 |  | 100.57 |  | 100.85 |  | 100.01 | 95.85 | 99.77 | 96.21 | 100 | 97.14 | 100.08 | 96.49 | 100.89 | 99.965 | 0.925 | 96.4225 | 2.51 | 3.435
103.8 | 101.45 |  | 100.85 |  | 101.13 |  | 100.59 | 96.03 | 100.17 | 96.28 | 100.22 | 97.11 | 100.45 | 96.71 | 101.1433 | 100.3575 | 0.78583 | 96.5325 | 2.6567 | 3.4425
104 | 101.62 |  | 101.29 |  | 101.49 |  | 100.89 | 96.1 | 100.56 | 96.31 | 100.3 | 97.18 | 100.58 | 96.78 | 101.4667 | 100.5825 | 0.88417 | 96.5925 | 2.5333 | 3.4175
104.4 | 102.65 |  | 101.43 | 88.5 | 101.97 |  | 101.46 | 96.13 | 100.69 | 96.35 | 100.62 | 97.11 | 100.85 | 96.67 | 102.0167 | 100.905 | 1.1117 | 96.565 | 2.3833 | 3.495
104.8 | 102.72 |  | 101.85 |  | 102.2 |  | 101.95 | 96.31 | 101.16 | 96.53 | 101.09 | 97.36 | 101.05 | 96.75 | 102.2567 | 101.3125 | 0.94417 | 96.7375 | 2.5433 | 3.4875
105 | 102.84 | 87.75 | 102.36 |  | 102.45 |  | 102.22 | 96.03 | 101.48 | 96.53 | 101.43 | 97.39 | 101.25 | 96.82 | 102.55 | 101.595 | 0.955 | 96.6925 | 2.45 | 3.405
105.4 | 103.05 |  | 102.74 |  | 102.75 |  | 102.47 | 96.13 | 101.88 | 96.64 | 102.27 | 97.47 | 101.56 | 96.93 | 102.8467 | 102.045 | 0.80167 | 96.7925 | 2.5533 | 3.355
105.8 | 103.78 |  | 103.15 |  | 102.94 |  | 102.9 | 96.21 | 102.25 | 96.64 | 102.65 | 97.47 | 101.81 | 96.89 | 103.29 | 102.4025 | 0.8875 | 96.8025 | 2.51 | 3.3975
106 | 104.15 |  | 103.22 |  | 103.15 |  | 103.36 | 96.31 | 102.58 | 96.64 | 102.78 | 97.61 | 102.22 | 97 | 103.5067 | 102.735 | 0.77167 | 96.89 | 2.4933 | 3.265
106.4 | 104.64 |  | 103.75 |  | 103.75 |  | 103.75 | 96.35 | 102.87 | 96.78 | 103.25 | 97.79 | 102.5 | 97.03 | 104.0467 | 103.0925 | 0.95417 | 96.9875 | 2.3533 | 3.3075
106.8 | 104.68 |  | 103.99 |  | 104.15 |  | 104.08 | 96.35 | 103.11 | 96.82 | 103.57 | 97.86 | 103.38 | 97.14 | 104.2733 | 103.535 | 0.73833 | 97.0425 | 2.5267 | 3.265
107 | 104.77 |  | 104.25 |  | 104.6 |  | 104.23 | 96.6 | 103.47 | 96.89 | 103.84 | 97.9 | 103.5 | 97.18 | 104.54 | 103.76 | 0.78 | 97.1425 | 2.46 | 3.24
107.4 | 104.86 |  | 105.13 |  | 105.09 |  | 104.51 | 96.53 | 103.82 | 97.03 | 104.16 | 97.93 | 103.88 | 97.21 | 105.0267 | 104.0925 | 0.93417 | 97.175 | 2.3733 | 3.3075
107.8 | 105.26 |  | 105.58 |  | 105.66 | 88.25 | 104.75 | 96.53 | 104.12 | 97.14 | 104.55 | 97.9 | 104.14 | 97.21 | 105.5 | 104.39 | 1.11 | 97.195 | 2.3 | 3.41
108 | 105.45 |  | 105.85 | 90 | 105.69 |  | 104.97 | 96.57 | 104.3 | 97.11 | 104.65 | 98.08 | 104.54 | 97.39 | 105.6633 | 104.615 | 1.0483 | 97.2875 | 2.3367 | 3.385
[15 empty] | 98.5747 | 97.7124 | 0.8623 |  | 2.365 | 3.2273
average shift of 0.86F when ambient temp. rises from 84 (28.8C) to 94 (34.4C) degree F, ie shif of 5.55C in temp.