## NEW 1
Reference Clinical Thermometer Readout | Reference DIGI Thermometer Readout | Reference Water Temperature (Sensor) | Calibrating contactless sensor Readout (2-3 cm) | Calibrating contactless sensor AMBIENT Readout | DIFF bw clinical and sensor | DIFF bw DIGI and sensor | DIFF BETWEEN CLINICAL AND DIGI | DIFF bwtween Water temp and clinical
Read at (deg F) |  | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F)
94 | 93.6 | 92.41 | 91.58 | 82.35 | 2.42 | 2.02 | 0.4 | 1.59
94.4 | 94.1 | 92.86 | 92.57 |  | 1.83 | 1.53 | 0.3 | 1.54
94.8 | 94.7 | 93.31 | 93.02 |  | 1.78 | 1.68 | 0.1 | 1.49
95 | 94.9 | 93.43 | 93.13 |  | 1.87 | 1.77 | 0.1 | 1.57
95.4 | 95.4 | 93.87 | 93.14 |  | 2.26 | 2.26 | 0 | 1.53
95.8 | 95.6 | 94.21 | 93.35 |  | 2.45 | 2.25 | 0.2 | 1.59
96 | 96 | 94.55 | 94.19 |  | 1.81 | 1.81 | 0 | 1.45
96.4 | 96.2 | 94.77 | 93.78 |  | 2.62 | 2.42 | 0.2 | 1.63
96.8 | 96.8 | 95.11 | 94.65 |  | 2.15 | 2.15 | 0 | 1.69
97 | 96.9 | 95.34 | 94.48 |  | 2.52 | 2.42 | 0.1 | 1.66
97.4 | 97 | 95.79 | 94.98 |  | 2.42 | 2.02 | 0.4 | 1.61
97.8 | 97.5 | 95.9 | 94.95 |  | 2.85 | 2.55 | 0.3 | 1.9
98 | 97.7 | 96.24 | 95.9 |  | 2.1 | 1.8 | 0.3 | 1.76
98.2 | 98 | 96.69 | 95.98 |  | 2.22 | 2.02 | 0.2 | 1.51
98.4 | 98 | 96.57 | 96.12 |  | 2.28 | 1.88 | 0.4 | 1.83
98.6 | 98.3 | 96.8 | 96.22 |  | 2.38 | 2.08 | 0.3 | 1.8
98.8 | 98.5 | 97.02 | 96.49 |  | 2.31 | 2.01 | 0.3 | 1.78
99 | 98.5 | 97.14 | 96.65 |  | 2.35 | 1.85 | 0.5 | 1.86
99.2 | 98.8 | 97.36 | 96.84 |  | 2.36 | 1.96 | 0.4 | 1.84
99.4 | 99 | 97.59 | 96.93 | 86 | 2.47 | 2.07 | 0.4 | 1.81
99.6 | 99.6 | 97.81 | 97.34 |  | 2.26 | 2.26 | 0 | 1.79
99.8 | 99.6 | 98.04 | 97.61 |  | 2.19 | 1.99 | 0.2 | 1.76
100 | 99.6 | 98.26 | 97.4 |  | 2.6 | 2.2 | 0.4 | 1.74
100.2 | 100 | 98.37 | 97.7 |  | 2.5 | 2.3 | 0.2 | 1.83
100.4 | 100.1 | 98.6 | 97.8 |  | 2.6 | 2.3 | 0.3 | 1.8
100.6 | 100.3 | 98.82 | 98.23 |  | 2.37 | 2.07 | 0.3 | 1.78
100.8 | 100.5 | 98.39 | 98.45 |  | 2.35 | 2.05 | 0.3 | 2.41
101 | 100.7 | 99.16 | 98.54 |  | 2.46 | 2.16 | 0.3 | 1.84
101.2 | 100.9 | 99.39 | 98.33 |  | 2.87 | 2.57 | 0.3 | 1.81
101.4 | 101.3 | 99.5 | 98.85 |  | 2.55 | 2.45 | 0.1 | 1.9
101.6 | 101.4 | 99.84 | 99.22 |  | 2.38 | 2.18 | 0.2 | 1.76
101.8 | 101.5 | 99.95 | 99.15 |  | 2.65 | 2.35 | 0.3 | 1.85
102 | 101.9 | 100.17 | 99.45 |  | 2.55 | 2.45 | 0.1 | 1.83
102.2 | 102.2 | 100.4 | 99.58 |  | 2.62 | 2.62 | 0 | 1.8
102.4 | 102.5 | 100.62 | 99.81 |  | 2.59 | 2.69 | -0.1 | 1.78
102.6 | 102.6 | 101.41 | 100.45 |  | 2.15 | 2.15 | 0 | 1.19
102.8 | 102.8 | 101.4 | 100.6 |  | 2.2 | 2.2 | 0 | 1.4
103 | 103 | 101.64 | 100.9 |  | 2.1 | 2.1 | 0 | 1.36
103.4 | 103.5 | 102.09 | 101.25 |  | 2.15 | 2.25 | -0.1 | 1.31
103.8 | 103.8 | 102.54 | 101.45 |  | 2.35 | 2.35 | 0 | 1.26
104 | 104 | 102.76 | 101.62 |  | 2.38 | 2.38 | 0 | 1.24
104.4 | 104.4 | 103.1 | 102.65 |  | 1.75 | 1.75 | 0 | 1.3
104.8 | 104.8 | 103.44 | 102.72 |  | 2.08 | 2.08 | 0 | 1.36
105 | 105 | 103.77 | 102.84 | 87.75 | 2.16 | 2.16 | 0 | 1.23
105.4 | 105.3 | 104.22 | 103.05 |  | 2.35 | 2.25 | 0.1 | 1.18
105.8 | 105.8 | 104.45 | 103.78 |  | 2.02 | 2.02 | 0 | 1.35
106 | 105.9 | 104.9 | 104.15 |  | 1.85 | 1.75 | 0.1 | 1.1
106.4 | 106.4 | 105.12 | 104.64 |  | 1.76 | 1.76 | 0 | 1.28
106.8 | 106.8 | 105.01 | 104.68 |  | 2.12 | 2.12 | 0 | 1.79
107 | 107 | 105.12 | 104.77 |  | 2.23 | 2.23 | 0 | 1.88
107.4 | 107.4 | 105.46 | 104.86 |  | 2.54 | 2.54 | 0 | 1.94
107.8 | 107.8 | 105.8 | 105.26 |  | 2.54 | 2.54 | 0 | 2
108 | 108 | 106.02 | 105.45 |  | 2.55 | 2.55 | 0 | 1.98

## RAW_HEATER
00:55:00.172 -> Water_Temp.:91.40, PWM:0
00:55:00.934 -> Water_Temp.:91.40, PWM:0
00:55:01.729 -> Water_Temp.:91.40, PWM:0
00:55:02.487 -> Water_Temp.:91.40, PWM:0
00:55:03.282 -> Water_Temp.:91.40, PWM:0
00:55:04.046 -> Water_Temp.:91.40, PWM:0
00:55:04.842 -> Water_Temp.:91.40, PWM:0
00:55:05.605 -> Water_Temp.:91.40, PWM:0
00:55:07.152 -> Water_Temp.:91.40, PWM:4
00:55:07.946 -> Water_Temp.:91.40, PWM:4
00:55:09.502 -> Water_Temp.:91.40, PWM:4
00:55:10.290 -> Water_Temp.:91.40, PWM:4
00:55:11.053 -> Water_Temp.:91.40, PWM:4
00:55:11.845 -> Water_Temp.:91.40, PWM:4
00:55:13.372 -> Water_Temp.:91.29, PWM:4
00:55:14.926 -> Water_Temp.:91.29, PWM:4
00:55:15.683 -> Water_Temp.:91.29, PWM:18
00:55:16.480 -> Water_Temp.:91.29, PWM:18
00:55:17.242 -> Water_Temp.:91.29, PWM:18
00:55:18.037 -> Water_Temp.:91.29, PWM:23
00:55:18.799 -> Water_Temp.:91.29, PWM:34
00:55:19.595 -> Water_Temp.:91.29, PWM:34
00:55:20.387 -> Water_Temp.:91.29, PWM:34
00:55:22.675 -> Water_Temp.:91.29, PWM:34
00:55:23.436 -> Water_Temp.:91.29, PWM:34
00:55:24.994 -> Water_Temp.:91.29, PWM:34
00:55:27.277 -> Water_Temp.:91.29, PWM:34
00:55:28.074 -> Water_Temp.:91.29, PWM:33
00:55:29.629 -> Water_Temp.:91.18, PWM:33
00:55:31.179 -> Water_Temp.:91.18, PWM:33
00:55:31.938 -> Water_Temp.:91.18, PWM:33
00:55:33.495 -> Water_Temp.:91.18, PWM:33
00:55:34.253 -> Water_Temp.:91.18, PWM:33
00:55:35.046 -> Water_Temp.:91.18, PWM:33
00:55:35.840 -> Water_Temp.:91.18, PWM:33
00:55:36.601 -> Water_Temp.:91.18, PWM:33
00:55:37.397 -> Water_Temp.:91.18, PWM:33
00:55:38.154 -> Water_Temp.:91.18, PWM:33
00:55:41.990 -> Water_Temp.:91.18, PWM:33
00:55:42.781 -> Water_Temp.:91.18, PWM:33
00:55:43.577 -> Water_Temp.:91.06, PWM:33
00:55:45.128 -> Water_Temp.:91.06, PWM:33
00:55:45.892 -> Water_Temp.:91.06, PWM:33
00:55:47.451 -> Water_Temp.:91.06, PWM:33
00:55:48.973 -> Water_Temp.:91.06, PWM:33
00:55:49.765 -> Water_Temp.:91.06, PWM:33
00:55:52.834 -> Water_Temp.:91.06, PWM:33
00:55:53.629 -> Water_Temp.:91.06, PWM:33
00:55:54.422 -> Water_Temp.:91.06, PWM:33
00:55:57.453 -> Water_Temp.:91.06, PWM:33
00:55:58.247 -> Water_Temp.:91.06, PWM:33
00:55:59.008 -> Water_Temp.:91.06, PWM:33
00:55:59.804 -> Water_Temp.:91.06, PWM:33
00:56:02.124 -> Water_Temp.:91.06, PWM:33
00:56:02.923 -> Water_Temp.:91.06, PWM:33
00:56:03.686 -> Water_Temp.:91.06, PWM:33
00:56:04.479 -> Water_Temp.:91.06, PWM:33
00:56:05.236 -> Water_Temp.:91.06, PWM:33
00:56:06.031 -> Water_Temp.:91.06, PWM:33
00:56:07.583 -> Water_Temp.:91.06, PWM:33
00:56:09.905 -> Water_Temp.:91.06, PWM:33
00:56:10.670 -> Water_Temp.:91.06, PWM:33
00:56:11.466 -> Water_Temp.:91.06, PWM:33
00:56:13.788 -> Water_Temp.:91.06, PWM:33
00:56:14.550 -> Water_Temp.:91.06, PWM:33
00:56:15.347 -> Water_Temp.:91.06, PWM:33
00:56:16.110 -> Water_Temp.:90.95, PWM:33
00:56:19.966 -> Water_Temp.:90.95, PWM:33
00:56:20.727 -> Water_Temp.:90.95, PWM:33
00:56:21.519 -> Water_Temp.:90.95, PWM:33
00:56:22.283 -> Water_Temp.:90.95, PWM:33
00:56:24.591 -> Water_Temp.:90.95, PWM:33
00:56:25.384 -> Water_Temp.:90.95, PWM:33
00:56:26.145 -> Water_Temp.:90.95, PWM:33
00:56:26.938 -> Water_Temp.:90.95, PWM:33
00:56:28.462 -> Water_Temp.:90.95, PWM:33
00:56:29.253 -> Water_Temp.:90.95, PWM:33
00:56:30.012 -> Water_Temp.:90.84, PWM:33
00:56:30.811 -> Water_Temp.:90.84, PWM:33
00:56:31.566 -> Water_Temp.:90.84, PWM:33
00:56:32.357 -> Water_Temp.:90.84, PWM:33
00:56:33.157 -> Water_Temp.:90.84, PWM:33
00:56:35.458 -> Water_Temp.:90.84, PWM:33
00:56:36.252 -> Water_Temp.:90.84, PWM:33
00:56:37.771 -> Water_Temp.:90.84, PWM:33
00:56:40.841 -> Water_Temp.:90.84, PWM:33
00:56:41.634 -> Water_Temp.:90.84, PWM:33
00:56:43.191 -> Water_Temp.:90.84, PWM:33
00:56:43.987 -> Water_Temp.:90.84, PWM:33
00:56:46.270 -> Water_Temp.:90.84, PWM:33
00:56:47.066 -> Water_Temp.:90.84, PWM:33
00:56:47.825 -> Water_Temp.:90.84, PWM:33
00:56:48.620 -> Water_Temp.:90.84, PWM:33
00:56:49.376 -> Water_Temp.:90.84, PWM:33
00:56:51.697 -> Water_Temp.:90.84, PWM:33
00:56:53.257 -> Water_Temp.:90.84, PWM:33
00:56:54.047 -> Water_Temp.:90.84, PWM:33
00:56:55.559 -> Water_Temp.:90.84, PWM:33
00:56:56.353 -> Water_Temp.:90.84, PWM:33
00:56:57.147 -> Water_Temp.:90.84, PWM:33
00:56:57.908 -> Water_Temp.:90.84, PWM:33
00:56:58.705 -> Water_Temp.:90.84, PWM:33
00:56:59.464 -> Water_Temp.:90.84, PWM:33
00:57:00.264 -> Water_Temp.:90.84, PWM:33
00:57:01.027 -> Water_Temp.:90.84, PWM:33
00:57:01.823 -> Water_Temp.:90.84, PWM:33
00:57:02.585 -> Water_Temp.:90.84, PWM:33
00:57:03.381 -> Water_Temp.:90.84, PWM:33
00:57:04.142 -> Water_Temp.:90.84, PWM:33
00:57:06.494 -> Water_Temp.:90.84, PWM:33
00:57:08.047 -> Water_Temp.:90.84, PWM:33
00:57:09.608 -> Water_Temp.:90.84, PWM:33
00:57:11.133 -> Water_Temp.:90.84, PWM:33
00:57:11.906 -> Water_Temp.:90.84, PWM:33
00:57:12.698 -> Water_Temp.:90.84, PWM:33
00:57:13.462 -> Water_Temp.:90.84, PWM:33
00:57:14.260 -> Water_Temp.:90.84, PWM:33
00:57:15.020 -> Water_Temp.:90.84, PWM:33
00:57:15.814 -> Water_Temp.:90.84, PWM:33
00:57:16.574 -> Water_Temp.:90.84, PWM:33
00:57:17.367 -> Water_Temp.:90.84, PWM:33
00:57:18.164 -> Water_Temp.:90.84, PWM:33
00:57:18.928 -> Water_Temp.:90.84, PWM:33
00:57:19.718 -> Water_Temp.:90.84, PWM:33
00:57:21.277 -> Water_Temp.:90.84, PWM:33
00:57:22.795 -> Water_Temp.:90.84, PWM:33
00:57:23.592 -> Water_Temp.:90.84, PWM:33
00:57:24.388 -> Water_Temp.:90.84, PWM:33
00:57:25.147 -> Water_Temp.:90.84, PWM:33
00:57:26.674 -> Water_Temp.:90.84, PWM:33
00:57:27.466 -> Water_Temp.:90.84, PWM:33
00:57:28.990 -> Water_Temp.:90.84, PWM:33
00:57:29.787 -> Water_Temp.:90.84, PWM:33
00:57:31.312 -> Water_Temp.:90.84, PWM:33
00:57:32.836 -> Water_Temp.:90.84, PWM:33
00:57:34.428 -> Water_Temp.:90.84, PWM:33
00:57:35.981 -> Water_Temp.:90.84, PWM:33
00:57:39.056 -> Water_Temp.:90.84, PWM:33
00:57:39.850 -> Water_Temp.:90.84, PWM:33
00:57:40.615 -> Water_Temp.:90.84, PWM:33
00:57:42.171 -> Water_Temp.:90.84, PWM:33
00:57:44.457 -> Water_Temp.:90.84, PWM:33
00:57:45.217 -> Water_Temp.:90.84, PWM:33
00:57:46.012 -> Water_Temp.:90.84, PWM:33
00:57:46.806 -> Water_Temp.:90.84, PWM:33
00:57:47.567 -> Water_Temp.:90.84, PWM:33
00:57:48.362 -> Water_Temp.:90.84, PWM:33
00:57:49.119 -> Water_Temp.:90.84, PWM:33
00:57:49.908 -> Water_Temp.:90.84, PWM:33
00:57:52.256 -> Water_Temp.:90.84, PWM:33
00:57:53.012 -> Water_Temp.:90.84, PWM:33
00:57:53.806 -> Water_Temp.:90.84, PWM:33
00:57:54.570 -> Water_Temp.:90.84, PWM:33
00:57:55.364 -> Water_Temp.:90.84, PWM:33
00:57:56.161 -> Water_Temp.:90.84, PWM:33
00:57:56.921 -> Water_Temp.:90.84, PWM:33
00:57:57.712 -> Water_Temp.:90.84, PWM:33
00:57:58.469 -> Water_Temp.:90.84, PWM:33
00:58:00.787 -> Water_Temp.:90.84, PWM:33
00:58:01.551 -> Water_Temp.:90.84, PWM:33
00:58:02.343 -> Water_Temp.:90.84, PWM:33
00:58:03.107 -> Water_Temp.:90.84, PWM:33
00:58:04.658 -> Water_Temp.:90.84, PWM:33
00:58:05.453 -> Water_Temp.:90.84, PWM:33
00:58:06.217 -> Water_Temp.:90.84, PWM:33
00:58:07.014 -> Water_Temp.:92.52, PWM:33
00:58:07.774 -> Water_Temp.:94.66, PWM:33
00:58:08.572 -> Water_Temp.:92.19, PWM:33
00:58:09.332 -> Water_Temp.:91.74, PWM:33
00:58:10.128 -> Water_Temp.:91.51, PWM:33
00:58:10.923 -> Water_Temp.:91.40, PWM:33
00:58:11.680 -> Water_Temp.:91.40, PWM:33
00:58:12.474 -> Water_Temp.:91.29, PWM:33
00:58:14.000 -> Water_Temp.:91.18, PWM:33
00:58:14.791 -> Water_Temp.:91.18, PWM:33
00:58:15.553 -> Water_Temp.:91.18, PWM:33
00:58:17.868 -> Water_Temp.:91.06, PWM:33
00:58:20.186 -> Water_Temp.:91.06, PWM:33
00:58:20.980 -> Water_Temp.:91.06, PWM:33
00:58:22.501 -> Water_Temp.:91.06, PWM:33
00:58:24.059 -> Water_Temp.:90.95, PWM:33
00:58:24.854 -> Water_Temp.:91.06, PWM:33
00:58:26.409 -> Water_Temp.:90.95, PWM:33
00:58:27.172 -> Water_Temp.:90.95, PWM:33
00:58:27.967 -> Water_Temp.:90.95, PWM:33
00:58:34.924 -> Water_Temp.:90.95, PWM:33
00:58:35.719 -> Water_Temp.:90.95, PWM:33
00:58:36.483 -> Water_Temp.:90.95, PWM:33
00:58:38.036 -> Water_Temp.:90.95, PWM:33
00:58:38.831 -> Water_Temp.:90.95, PWM:33
00:58:39.592 -> Water_Temp.:90.95, PWM:33
00:58:40.389 -> Water_Temp.:90.95, PWM:33
00:58:41.913 -> Water_Temp.:91.06, PWM:33
00:58:42.714 -> Water_Temp.:91.06, PWM:33
00:58:43.477 -> Water_Temp.:91.06, PWM:33
00:58:45.027 -> Water_Temp.:91.06, PWM:33
00:58:45.825 -> Water_Temp.:91.06, PWM:33
00:58:46.587 -> Water_Temp.:91.06, PWM:33
00:58:47.384 -> Water_Temp.:91.06, PWM:33
00:58:50.456 -> Water_Temp.:91.06, PWM:33
00:58:51.251 -> Water_Temp.:91.06, PWM:33
00:58:52.046 -> Water_Temp.:91.06, PWM:33
00:58:53.564 -> Water_Temp.:91.06, PWM:33
00:58:54.357 -> Water_Temp.:91.06, PWM:33
00:58:55.150 -> Water_Temp.:91.06, PWM:33
00:58:55.914 -> Water_Temp.:91.06, PWM:33
00:58:56.705 -> Water_Temp.:91.06, PWM:33
00:58:57.462 -> Water_Temp.:91.18, PWM:33
00:58:58.253 -> Water_Temp.:91.18, PWM:33
00:58:59.049 -> Water_Temp.:91.18, PWM:33
00:58:59.810 -> Water_Temp.:91.18, PWM:33
00:59:00.608 -> Water_Temp.:91.18, PWM:33
00:59:01.366 -> Water_Temp.:91.18, PWM:33
00:59:02.161 -> Water_Temp.:91.18, PWM:33
00:59:03.719 -> Water_Temp.:91.18, PWM:33
00:59:05.271 -> Water_Temp.:91.18, PWM:33
00:59:06.824 -> Water_Temp.:91.18, PWM:33
00:59:08.346 -> Water_Temp.:91.18, PWM:33
00:59:09.144 -> Water_Temp.:91.18, PWM:33
00:59:10.701 -> Water_Temp.:91.18, PWM:33
00:59:11.459 -> Water_Temp.:91.18, PWM:33
00:59:13.017 -> Water_Temp.:91.18, PWM:33
00:59:13.816 -> Water_Temp.:91.18, PWM:33
00:59:14.572 -> Water_Temp.:91.18, PWM:33
00:59:18.443 -> Water_Temp.:91.18, PWM:33
00:59:19.204 -> Water_Temp.:91.18, PWM:33
00:59:19.999 -> Water_Temp.:91.18, PWM:33
00:59:21.521 -> Water_Temp.:91.18, PWM:33
00:59:22.317 -> Water_Temp.:91.18, PWM:33
00:59:23.871 -> Water_Temp.:91.29, PWM:33
00:59:24.628 -> Water_Temp.:91.29, PWM:33
00:59:25.425 -> Water_Temp.:91.29, PWM:33
00:59:26.951 -> Water_Temp.:91.29, PWM:33
00:59:31.555 -> Water_Temp.:91.29, PWM:33
00:59:33.083 -> Water_Temp.:91.29, PWM:33
00:59:33.878 -> Water_Temp.:91.29, PWM:33
00:59:34.675 -> Water_Temp.:91.40, PWM:33
00:59:35.434 -> Water_Temp.:91.40, PWM:33
00:59:36.198 -> Water_Temp.:91.40, PWM:33
00:59:36.992 -> Water_Temp.:91.40, PWM:33
00:59:37.786 -> Water_Temp.:91.40, PWM:33
00:59:39.344 -> Water_Temp.:91.40, PWM:33
00:59:40.103 -> Water_Temp.:91.40, PWM:33
00:59:40.899 -> Water_Temp.:91.40, PWM:33
00:59:41.663 -> Water_Temp.:91.40, PWM:33
00:59:42.458 -> Water_Temp.:91.40, PWM:33
00:59:43.249 -> Water_Temp.:91.40, PWM:33
00:59:44.013 -> Water_Temp.:91.40, PWM:33
00:59:45.572 -> Water_Temp.:91.40, PWM:33
00:59:47.095 -> Water_Temp.:91.40, PWM:33
00:59:48.653 -> Water_Temp.:91.40, PWM:33
00:59:49.414 -> Water_Temp.:91.40, PWM:33
00:59:50.211 -> Water_Temp.:91.40, PWM:33
00:59:50.977 -> Water_Temp.:91.40, PWM:33
00:59:52.535 -> Water_Temp.:91.40, PWM:33
00:59:53.301 -> Water_Temp.:91.40, PWM:33
00:59:57.137 -> Water_Temp.:91.40, PWM:33
00:59:57.930 -> Water_Temp.:91.51, PWM:33
00:59:58.687 -> Water_Temp.:91.51, PWM:33
00:59:59.484 -> Water_Temp.:91.51, PWM:33
01:00:00.243 -> Water_Temp.:91.51, PWM:33
01:00:01.041 -> Water_Temp.:91.51, PWM:33
01:00:02.598 -> Water_Temp.:91.51, PWM:33
01:00:03.392 -> Water_Temp.:91.51, PWM:33
01:00:04.151 -> Water_Temp.:91.51, PWM:33
01:00:04.947 -> Water_Temp.:91.51, PWM:33
01:00:06.469 -> Water_Temp.:91.51, PWM:33
01:00:08.024 -> Water_Temp.:91.51, PWM:33
01:00:08.824 -> Water_Temp.:91.51, PWM:33
01:00:10.341 -> Water_Temp.:91.51, PWM:33
01:00:12.660 -> Water_Temp.:91.51, PWM:33
01:00:13.423 -> Water_Temp.:91.51, PWM:33
01:00:14.216 -> Water_Temp.:91.51, PWM:33
01:00:14.976 -> Water_Temp.:91.51, PWM:33
01:00:15.773 -> Water_Temp.:91.62, PWM:33
01:00:16.535 -> Water_Temp.:91.62, PWM:33
01:00:17.333 -> Water_Temp.:91.62, PWM:33
01:00:18.094 -> Water_Temp.:91.62, PWM:33
01:00:18.889 -> Water_Temp.:91.62, PWM:33
01:00:20.444 -> Water_Temp.:91.62, PWM:33
01:00:21.999 -> Water_Temp.:91.62, PWM:33
01:00:23.559 -> Water_Temp.:91.62, PWM:33
01:00:26.642 -> Water_Temp.:91.62, PWM:33
01:00:27.405 -> Water_Temp.:91.62, PWM:33
01:00:30.454 -> Water_Temp.:91.62, PWM:33
01:00:31.251 -> Water_Temp.:91.62, PWM:33
01:00:32.013 -> Water_Temp.:91.62, PWM:33
01:00:32.807 -> Water_Temp.:91.62, PWM:33
01:00:33.571 -> Water_Temp.:91.62, PWM:33
01:00:34.363 -> Water_Temp.:91.62, PWM:33
01:00:35.157 -> Water_Temp.:91.62, PWM:33
01:00:36.677 -> Water_Temp.:91.62, PWM:33
01:00:37.441 -> Water_Temp.:91.62, PWM:33
01:00:39.032 -> Water_Temp.:91.62, PWM:33
01:00:39.791 -> Water_Temp.:91.62, PWM:33
01:00:42.110 -> Water_Temp.:91.62, PWM:33
01:00:42.906 -> Water_Temp.:91.62, PWM:33
01:00:46.012 -> Water_Temp.:91.62, PWM:33
01:00:46.773 -> Water_Temp.:91.62, PWM:33
01:00:47.565 -> Water_Temp.:91.62, PWM:33
01:00:48.361 -> Water_Temp.:91.62, PWM:33
01:00:50.648 -> Water_Temp.:91.62, PWM:33
01:00:51.440 -> Water_Temp.:91.62, PWM:33
01:00:52.206 -> Water_Temp.:91.62, PWM:33
01:00:53.001 -> Water_Temp.:91.62, PWM:33
01:00:53.795 -> Water_Temp.:91.62, PWM:33
01:00:54.554 -> Water_Temp.:91.62, PWM:33
01:00:56.109 -> Water_Temp.:91.62, PWM:33
01:00:56.867 -> Water_Temp.:91.62, PWM:33
01:00:57.662 -> Water_Temp.:91.62, PWM:33
01:00:58.420 -> Water_Temp.:91.62, PWM:33
01:00:59.218 -> Water_Temp.:91.62, PWM:33
01:01:01.533 -> Water_Temp.:91.74, PWM:33
01:01:02.328 -> Water_Temp.:91.74, PWM:33
01:01:04.647 -> Water_Temp.:91.74, PWM:33
01:01:05.410 -> Water_Temp.:91.74, PWM:33
01:01:06.176 -> Water_Temp.:91.74, PWM:33
01:01:06.972 -> Water_Temp.:91.74, PWM:33
01:01:07.734 -> Water_Temp.:91.74, PWM:33
01:01:08.529 -> Water_Temp.:91.74, PWM:33
01:01:10.081 -> Water_Temp.:91.74, PWM:33
01:01:10.874 -> Water_Temp.:91.74, PWM:33
01:01:13.223 -> Water_Temp.:91.74, PWM:33
01:01:14.740 -> Water_Temp.:91.74, PWM:33
01:01:16.300 -> Water_Temp.:91.74, PWM:33
01:01:17.062 -> Water_Temp.:91.74, PWM:33
01:01:19.384 -> Water_Temp.:91.74, PWM:33
01:01:24.014 -> Water_Temp.:91.85, PWM:33
01:01:25.563 -> Water_Temp.:91.85, PWM:33
01:01:26.358 -> Water_Temp.:91.85, PWM:33
01:01:27.120 -> Water_Temp.:91.85, PWM:33
01:01:27.915 -> Water_Temp.:91.85, PWM:33
01:01:28.678 -> Water_Temp.:91.85, PWM:33
01:01:29.471 -> Water_Temp.:91.85, PWM:33
01:01:30.233 -> Water_Temp.:91.85, PWM:33
01:01:31.033 -> Water_Temp.:91.85, PWM:33
01:01:31.797 -> Water_Temp.:91.85, PWM:33
01:01:32.562 -> Water_Temp.:91.85, PWM:33
01:01:33.362 -> Water_Temp.:91.85, PWM:33
01:01:34.124 -> Water_Temp.:91.85, PWM:33
01:01:34.918 -> Water_Temp.:91.85, PWM:33
01:01:35.713 -> Water_Temp.:91.96, PWM:33
01:01:36.473 -> Water_Temp.:91.85, PWM:33
01:01:38.033 -> Water_Temp.:91.96, PWM:33
01:01:41.147 -> Water_Temp.:91.96, PWM:33
01:01:42.702 -> Water_Temp.:91.96, PWM:33
01:01:43.460 -> Water_Temp.:91.96, PWM:33
01:01:45.012 -> Water_Temp.:91.96, PWM:33
01:01:46.605 -> Water_Temp.:91.96, PWM:33
01:01:47.364 -> Water_Temp.:91.96, PWM:33
01:01:48.155 -> Water_Temp.:91.96, PWM:33
01:01:48.914 -> Water_Temp.:91.96, PWM:33
01:01:50.473 -> Water_Temp.:91.96, PWM:33
01:01:51.998 -> Water_Temp.:91.96, PWM:33
01:01:52.794 -> Water_Temp.:91.96, PWM:33
01:01:53.556 -> Water_Temp.:91.96, PWM:33
01:01:56.606 -> Water_Temp.:92.07, PWM:33
01:01:58.161 -> Water_Temp.:92.07, PWM:33
01:01:58.919 -> Water_Temp.:92.07, PWM:33
01:01:59.715 -> Water_Temp.:92.07, PWM:33
01:02:00.512 -> Water_Temp.:92.07, PWM:33
01:02:01.273 -> Water_Temp.:92.07, PWM:33
01:02:02.069 -> Water_Temp.:92.07, PWM:33
01:02:02.829 -> Water_Temp.:92.07, PWM:33
01:02:04.387 -> Water_Temp.:92.07, PWM:33
01:02:05.146 -> Water_Temp.:92.07, PWM:33
01:02:05.939 -> Water_Temp.:92.07, PWM:33
01:02:06.733 -> Water_Temp.:92.07, PWM:33
01:02:08.259 -> Water_Temp.:92.07, PWM:33
01:02:09.819 -> Water_Temp.:92.07, PWM:33
01:02:10.577 -> Water_Temp.:92.07, PWM:33
01:02:11.372 -> Water_Temp.:92.07, PWM:33
01:02:12.928 -> Water_Temp.:92.07, PWM:33
01:02:13.691 -> Water_Temp.:92.07, PWM:33
01:02:14.487 -> Water_Temp.:92.07, PWM:33
01:02:15.248 -> Water_Temp.:92.07, PWM:33
01:02:16.042 -> Water_Temp.:92.07, PWM:33
01:02:16.805 -> Water_Temp.:92.07, PWM:33
01:02:19.163 -> Water_Temp.:92.19, PWM:33
01:02:19.918 -> Water_Temp.:92.19, PWM:33
01:02:20.717 -> Water_Temp.:92.19, PWM:33
01:02:21.476 -> Water_Temp.:92.19, PWM:33
01:02:23.794 -> Water_Temp.:92.19, PWM:33
01:02:24.592 -> Water_Temp.:92.19, PWM:33
01:02:25.353 -> Water_Temp.:92.19, PWM:33
01:02:26.917 -> Water_Temp.:92.19, PWM:33
01:02:27.675 -> Water_Temp.:92.19, PWM:33
01:02:28.472 -> Water_Temp.:92.19, PWM:33
01:02:29.264 -> Water_Temp.:92.19, PWM:33
01:02:30.820 -> Water_Temp.:92.19, PWM:33
01:02:31.579 -> Water_Temp.:92.19, PWM:33
01:02:32.375 -> Water_Temp.:92.19, PWM:33
01:02:35.413 -> Water_Temp.:92.19, PWM:33
01:02:37.004 -> Water_Temp.:92.19, PWM:33
01:02:37.764 -> Water_Temp.:92.19, PWM:33
01:02:38.559 -> Water_Temp.:92.19, PWM:33
01:02:39.323 -> Water_Temp.:92.19, PWM:33
01:02:40.118 -> Water_Temp.:92.19, PWM:33
01:02:42.436 -> Water_Temp.:92.30, PWM:33
01:02:43.201 -> Water_Temp.:92.30, PWM:33
01:02:43.995 -> Water_Temp.:92.30, PWM:33
01:02:45.518 -> Water_Temp.:92.30, PWM:33
01:02:46.280 -> Water_Temp.:92.30, PWM:33
01:02:47.078 -> Water_Temp.:92.30, PWM:33
01:02:47.840 -> Water_Temp.:92.30, PWM:33
01:02:49.401 -> Water_Temp.:92.30, PWM:33
01:02:50.162 -> Water_Temp.:92.30, PWM:33
01:02:50.958 -> Water_Temp.:92.30, PWM:33
01:02:51.722 -> Water_Temp.:92.30, PWM:33
01:02:53.283 -> Water_Temp.:92.30, PWM:33
01:02:54.075 -> Water_Temp.:92.30, PWM:33
01:02:54.866 -> Water_Temp.:92.41, PWM:33
01:02:55.624 -> Water_Temp.:92.41, PWM:33
01:02:57.179 -> Water_Temp.:92.41, PWM:33
01:02:58.737 -> Water_Temp.:92.41, PWM:33
01:03:01.057 -> Water_Temp.:92.41, PWM:33
01:03:03.376 -> Water_Temp.:92.41, PWM:33
01:03:04.168 -> Water_Temp.:92.41, PWM:33
01:03:05.688 -> Water_Temp.:92.41, PWM:33
01:03:07.247 -> Water_Temp.:92.41, PWM:33
01:03:08.038 -> Water_Temp.:92.41, PWM:33
01:03:09.598 -> Water_Temp.:92.41, PWM:33
01:03:10.358 -> Water_Temp.:92.41, PWM:33
01:03:11.157 -> Water_Temp.:92.52, PWM:33
01:03:12.677 -> Water_Temp.:92.41, PWM:33
01:03:13.438 -> Water_Temp.:92.41, PWM:33
01:03:17.295 -> Water_Temp.:92.41, PWM:33
01:03:18.841 -> Water_Temp.:92.41, PWM:33
01:03:20.393 -> Water_Temp.:92.41, PWM:33
01:03:22.746 -> Water_Temp.:92.41, PWM:33
01:03:23.508 -> Water_Temp.:92.52, PWM:33
01:03:24.303 -> Water_Temp.:92.52, PWM:33
01:03:25.063 -> Water_Temp.:92.52, PWM:33
01:03:25.855 -> Water_Temp.:92.52, PWM:33
01:03:26.649 -> Water_Temp.:92.52, PWM:33
01:03:27.408 -> Water_Temp.:92.52, PWM:33
01:03:28.204 -> Water_Temp.:92.52, PWM:33
01:03:29.760 -> Water_Temp.:92.52, PWM:33
01:03:30.524 -> Water_Temp.:92.52, PWM:33
01:03:31.319 -> Water_Temp.:92.52, PWM:33
01:03:33.630 -> Water_Temp.:92.41, PWM:33
01:03:34.423 -> Water_Temp.:92.52, PWM:33
01:03:36.730 -> Water_Temp.:92.52, PWM:33
01:03:37.492 -> Water_Temp.:92.52, PWM:33
01:03:38.289 -> Water_Temp.:92.52, PWM:33
01:03:39.083 -> Water_Temp.:92.52, PWM:33
01:03:39.843 -> Water_Temp.:92.52, PWM:33
01:03:40.641 -> Water_Temp.:92.52, PWM:33
01:03:42.959 -> Water_Temp.:92.52, PWM:33
01:03:43.720 -> Water_Temp.:92.52, PWM:33
01:03:45.308 -> Water_Temp.:92.52, PWM:33
01:03:46.064 -> Water_Temp.:92.52, PWM:33
01:03:46.856 -> Water_Temp.:92.52, PWM:33
01:03:48.377 -> Water_Temp.:92.52, PWM:33
01:03:49.169 -> Water_Temp.:92.52, PWM:33
01:03:49.931 -> Water_Temp.:92.52, PWM:33
01:03:50.729 -> Water_Temp.:92.52, PWM:33
01:03:53.051 -> Water_Temp.:92.52, PWM:33
01:03:55.373 -> Water_Temp.:92.52, PWM:33
01:03:56.929 -> Water_Temp.:92.52, PWM:33
01:03:57.689 -> Water_Temp.:92.52, PWM:33
01:03:59.277 -> Water_Temp.:92.52, PWM:33
01:04:00.037 -> Water_Temp.:92.52, PWM:33
01:04:00.834 -> Water_Temp.:92.52, PWM:33
01:04:01.596 -> Water_Temp.:92.52, PWM:33
01:04:02.394 -> Water_Temp.:92.52, PWM:33
01:04:03.152 -> Water_Temp.:92.52, PWM:33
01:04:03.945 -> Water_Temp.:92.52, PWM:33
01:04:04.740 -> Water_Temp.:92.52, PWM:33
01:04:05.501 -> Water_Temp.:92.52, PWM:33
01:04:06.298 -> Water_Temp.:92.64, PWM:33
01:04:07.059 -> Water_Temp.:92.64, PWM:33
01:04:07.856 -> Water_Temp.:92.64, PWM:33
01:04:08.618 -> Water_Temp.:92.64, PWM:33
01:04:10.174 -> Water_Temp.:92.64, PWM:33
01:04:10.940 -> Water_Temp.:92.64, PWM:33
01:04:14.022 -> Water_Temp.:92.64, PWM:33
01:04:14.816 -> Water_Temp.:92.64, PWM:33
01:04:16.378 -> Water_Temp.:92.64, PWM:33
01:04:17.141 -> Water_Temp.:92.64, PWM:33
01:04:17.937 -> Water_Temp.:92.64, PWM:33
01:04:20.253 -> Water_Temp.:92.64, PWM:33
01:04:21.019 -> Water_Temp.:92.64, PWM:33
01:04:22.575 -> Water_Temp.:92.64, PWM:33
01:04:23.372 -> Water_Temp.:92.64, PWM:33
01:04:24.134 -> Water_Temp.:92.64, PWM:33
01:04:24.926 -> Water_Temp.:92.64, PWM:33
01:04:25.724 -> Water_Temp.:92.75, PWM:33
01:04:26.483 -> Water_Temp.:92.64, PWM:33
01:04:27.281 -> Water_Temp.:92.75, PWM:33
01:04:28.804 -> Water_Temp.:92.75, PWM:33
01:04:29.598 -> Water_Temp.:92.75, PWM:33
01:04:30.361 -> Water_Temp.:92.64, PWM:33
01:04:31.159 -> Water_Temp.:92.75, PWM:33
01:04:31.922 -> Water_Temp.:92.75, PWM:33
01:04:33.482 -> Water_Temp.:92.75, PWM:33
01:04:34.277 -> Water_Temp.:92.75, PWM:33
01:04:35.039 -> Water_Temp.:92.75, PWM:33
01:04:37.378 -> Water_Temp.:92.75, PWM:33
01:04:38.137 -> Water_Temp.:92.75, PWM:33
01:04:38.934 -> Water_Temp.:92.75, PWM:33
01:04:39.695 -> Water_Temp.:92.75, PWM:33
01:04:40.488 -> Water_Temp.:92.75, PWM:33
01:04:41.248 -> Water_Temp.:92.75, PWM:33
01:04:42.042 -> Water_Temp.:92.75, PWM:33
01:04:42.836 -> Water_Temp.:92.86, PWM:33
01:04:44.356 -> Water_Temp.:92.86, PWM:33
01:04:45.911 -> Water_Temp.:92.86, PWM:33
01:04:46.706 -> Water_Temp.:92.86, PWM:33
01:04:47.469 -> Water_Temp.:92.86, PWM:33
01:04:48.260 -> Water_Temp.:92.86, PWM:33
01:04:51.335 -> Water_Temp.:92.86, PWM:33
01:04:53.654 -> Water_Temp.:92.86, PWM:33
01:04:55.206 -> Water_Temp.:92.86, PWM:33
01:04:55.965 -> Water_Temp.:92.86, PWM:33
01:04:57.522 -> Water_Temp.:92.86, PWM:33
01:04:58.283 -> Water_Temp.:92.86, PWM:33
01:04:59.835 -> Water_Temp.:92.86, PWM:33
01:05:00.598 -> Water_Temp.:92.86, PWM:33
01:05:01.396 -> Water_Temp.:92.86, PWM:33
01:05:02.157 -> Water_Temp.:92.86, PWM:33
01:05:03.714 -> Water_Temp.:92.97, PWM:33
01:05:04.511 -> Water_Temp.:92.97, PWM:33
01:05:05.269 -> Water_Temp.:92.97, PWM:33
01:05:06.059 -> Water_Temp.:92.97, PWM:33
01:05:06.823 -> Water_Temp.:92.97, PWM:33
01:05:07.612 -> Water_Temp.:92.86, PWM:33
01:05:09.169 -> Water_Temp.:92.86, PWM:33
01:05:11.486 -> Water_Temp.:92.86, PWM:33
01:05:12.278 -> Water_Temp.:92.86, PWM:33
01:05:13.039 -> Water_Temp.:92.86, PWM:33
01:05:14.626 -> Water_Temp.:92.86, PWM:33
01:05:15.384 -> Water_Temp.:92.86, PWM:33
01:05:16.177 -> Water_Temp.:92.86, PWM:33
01:05:19.257 -> Water_Temp.:92.86, PWM:33
01:05:20.047 -> Water_Temp.:92.86, PWM:33
01:05:21.569 -> Water_Temp.:92.97, PWM:33
01:05:22.358 -> Water_Temp.:92.97, PWM:33
01:05:23.148 -> Water_Temp.:92.97, PWM:33
01:05:23.943 -> Water_Temp.:92.97, PWM:33
01:05:25.466 -> Water_Temp.:92.97, PWM:33
01:05:26.256 -> Water_Temp.:92.97, PWM:33
01:05:27.014 -> Water_Temp.:92.97, PWM:33
01:05:27.809 -> Water_Temp.:92.97, PWM:33
01:05:28.575 -> Water_Temp.:92.97, PWM:33
01:05:29.371 -> Water_Temp.:92.97, PWM:33
01:05:30.163 -> Water_Temp.:93.09, PWM:33
01:05:30.921 -> Water_Temp.:93.09, PWM:33
01:05:32.473 -> Water_Temp.:92.97, PWM:33
01:05:36.315 -> Water_Temp.:92.97, PWM:33
01:05:37.877 -> Water_Temp.:92.97, PWM:33
01:05:38.672 -> Water_Temp.:93.09, PWM:33
01:05:40.196 -> Water_Temp.:93.09, PWM:33
01:05:40.991 -> Water_Temp.:93.09, PWM:33
01:05:41.770 -> Water_Temp.:93.09, PWM:33
01:05:42.536 -> Water_Temp.:93.09, PWM:33
01:05:43.335 -> Water_Temp.:93.09, PWM:33
01:05:44.099 -> Water_Temp.:93.09, PWM:33
01:05:44.897 -> Water_Temp.:93.09, PWM:33
01:05:45.661 -> Water_Temp.:93.09, PWM:33
01:05:46.460 -> Water_Temp.:93.09, PWM:33
01:05:48.015 -> Water_Temp.:93.09, PWM:33
01:05:48.775 -> Water_Temp.:93.09, PWM:33
01:05:50.328 -> Water_Temp.:96.91, PWM:33
01:05:51.890 -> Water_Temp.:94.10, PWM:33
01:05:53.473 -> Water_Temp.:93.76, PWM:33
01:05:54.233 -> Water_Temp.:93.65, PWM:33
01:05:56.588 -> Water_Temp.:93.43, PWM:33
01:05:57.347 -> Water_Temp.:93.43, PWM:33
01:06:01.975 -> Water_Temp.:93.31, PWM:33
01:06:02.770 -> Water_Temp.:93.31, PWM:33
01:06:05.874 -> Water_Temp.:93.31, PWM:33
01:06:06.670 -> Water_Temp.:93.31, PWM:33
01:06:07.432 -> Water_Temp.:93.31, PWM:33
01:06:08.988 -> Water_Temp.:93.31, PWM:33
01:06:10.546 -> Water_Temp.:93.20, PWM:33
01:06:12.103 -> Water_Temp.:93.20, PWM:33
01:06:13.660 -> Water_Temp.:93.20, PWM:33
01:06:15.969 -> Water_Temp.:93.20, PWM:33
01:06:16.728 -> Water_Temp.:93.20, PWM:33
01:06:17.526 -> Water_Temp.:93.20, PWM:33
01:06:19.051 -> Water_Temp.:93.20, PWM:33
01:06:19.811 -> Water_Temp.:93.20, PWM:33
01:06:21.374 -> Water_Temp.:93.20, PWM:33
01:06:22.141 -> Water_Temp.:93.20, PWM:33
01:06:22.909 -> Water_Temp.:93.20, PWM:33
01:06:23.710 -> Water_Temp.:93.20, PWM:33
01:06:24.479 -> Water_Temp.:93.20, PWM:33
01:06:25.246 -> Water_Temp.:93.20, PWM:33
01:06:26.049 -> Water_Temp.:93.20, PWM:33
01:06:29.917 -> Water_Temp.:93.20, PWM:33
01:06:31.486 -> Water_Temp.:93.31, PWM:33
01:06:32.254 -> Water_Temp.:93.31, PWM:33
01:06:33.022 -> Water_Temp.:93.31, PWM:33
01:06:34.593 -> Water_Temp.:93.31, PWM:33
01:06:36.897 -> Water_Temp.:93.31, PWM:33
01:06:37.664 -> Water_Temp.:93.31, PWM:33
01:06:38.432 -> Water_Temp.:93.31, PWM:33
01:06:39.235 -> Water_Temp.:93.31, PWM:33
01:06:40.003 -> Water_Temp.:93.31, PWM:33
01:06:40.771 -> Water_Temp.:93.31, PWM:33
01:06:41.572 -> Water_Temp.:93.31, PWM:33
01:06:43.876 -> Water_Temp.:93.43, PWM:33
01:06:44.644 -> Water_Temp.:93.43, PWM:33
01:06:45.444 -> Water_Temp.:93.43, PWM:33
01:06:46.212 -> Water_Temp.:93.43, PWM:33
01:06:46.980 -> Water_Temp.:93.43, PWM:33
01:06:47.783 -> Water_Temp.:93.43, PWM:33
01:06:48.548 -> Water_Temp.:93.43, PWM:33
01:06:49.342 -> Water_Temp.:93.43, PWM:33
01:06:50.105 -> Water_Temp.:93.43, PWM:33
01:06:51.635 -> Water_Temp.:93.43, PWM:33
01:06:52.432 -> Water_Temp.:93.43, PWM:33
01:06:53.196 -> Water_Temp.:93.43, PWM:33
01:06:54.761 -> Water_Temp.:93.43, PWM:33
01:06:55.523 -> Water_Temp.:93.43, PWM:33
01:06:57.842 -> Water_Temp.:93.43, PWM:33
01:06:58.634 -> Water_Temp.:93.43, PWM:33
01:07:00.937 -> Water_Temp.:93.43, PWM:33
01:07:01.699 -> Water_Temp.:93.43, PWM:33
01:07:02.494 -> Water_Temp.:93.43, PWM:33
01:07:04.009 -> Water_Temp.:93.43, PWM:33
01:07:04.805 -> Water_Temp.:93.43, PWM:33
01:07:05.598 -> Water_Temp.:93.43, PWM:33
01:07:06.356 -> Water_Temp.:93.43, PWM:33
01:07:07.151 -> Water_Temp.:93.43, PWM:33
01:07:10.985 -> Water_Temp.:93.43, PWM:33
01:07:11.742 -> Water_Temp.:93.54, PWM:33
01:07:12.534 -> Water_Temp.:93.54, PWM:33
01:07:14.844 -> Water_Temp.:93.54, PWM:33
01:07:15.602 -> Water_Temp.:93.54, PWM:33
01:07:16.398 -> Water_Temp.:93.54, PWM:33
01:07:17.917 -> Water_Temp.:93.54, PWM:33
01:07:18.712 -> Water_Temp.:93.54, PWM:33
01:07:21.823 -> Water_Temp.:93.54, PWM:33
01:07:22.582 -> Water_Temp.:93.54, PWM:33
01:07:23.376 -> Water_Temp.:93.54, PWM:33
01:07:24.134 -> Water_Temp.:93.54, PWM:33
01:07:25.686 -> Water_Temp.:93.54, PWM:33
01:07:26.449 -> Water_Temp.:93.65, PWM:33
01:07:27.241 -> Water_Temp.:93.65, PWM:33
01:07:28.796 -> Water_Temp.:93.65, PWM:33
01:07:31.079 -> Water_Temp.:93.65, PWM:33
01:07:31.874 -> Water_Temp.:93.65, PWM:33
01:07:34.225 -> Water_Temp.:93.65, PWM:33
01:07:34.984 -> Water_Temp.:93.65, PWM:33
01:07:35.776 -> Water_Temp.:93.65, PWM:33
01:07:37.299 -> Water_Temp.:93.65, PWM:33
01:07:38.853 -> Water_Temp.:93.65, PWM:33
01:07:39.646 -> Water_Temp.:93.65, PWM:33
01:07:42.720 -> Water_Temp.:93.76, PWM:33
01:07:43.515 -> Water_Temp.:93.76, PWM:33
01:07:44.277 -> Water_Temp.:93.76, PWM:33
01:07:46.633 -> Water_Temp.:93.76, PWM:33
01:07:47.391 -> Water_Temp.:93.76, PWM:33
01:07:48.183 -> Water_Temp.:93.76, PWM:33
01:07:48.975 -> Water_Temp.:93.76, PWM:33
01:07:49.739 -> Water_Temp.:93.76, PWM:33
01:07:50.536 -> Water_Temp.:93.76, PWM:33
01:07:51.296 -> Water_Temp.:93.76, PWM:33
01:07:52.094 -> Water_Temp.:93.76, PWM:33
01:07:52.852 -> Water_Temp.:93.76, PWM:33
01:07:53.649 -> Water_Temp.:93.76, PWM:33
01:07:54.412 -> Water_Temp.:93.76, PWM:33
01:07:55.209 -> Water_Temp.:93.76, PWM:33
01:07:56.764 -> Water_Temp.:93.76, PWM:33
01:07:59.039 -> Water_Temp.:93.76, PWM:33
01:07:59.828 -> Water_Temp.:93.76, PWM:33
01:08:00.622 -> Water_Temp.:93.76, PWM:33
01:08:02.178 -> Water_Temp.:93.87, PWM:33
01:08:03.733 -> Water_Temp.:93.87, PWM:33
01:08:06.045 -> Water_Temp.:93.87, PWM:33
01:08:06.804 -> Water_Temp.:93.87, PWM:33
01:08:07.598 -> Water_Temp.:93.87, PWM:33
01:08:08.394 -> Water_Temp.:93.87, PWM:33
01:08:09.915 -> Water_Temp.:93.87, PWM:33
01:08:10.712 -> Water_Temp.:93.87, PWM:33
01:08:11.473 -> Water_Temp.:93.87, PWM:33
01:08:12.270 -> Water_Temp.:93.87, PWM:33
01:08:13.823 -> Water_Temp.:93.87, PWM:33
01:08:15.379 -> Water_Temp.:93.87, PWM:33
01:08:16.140 -> Water_Temp.:93.87, PWM:33
01:08:16.935 -> Water_Temp.:93.87, PWM:33
01:08:18.492 -> Water_Temp.:93.87, PWM:33
01:08:19.254 -> Water_Temp.:93.87, PWM:33
01:08:20.046 -> Water_Temp.:93.87, PWM:33
01:08:21.606 -> Water_Temp.:93.99, PWM:33
01:08:22.400 -> Water_Temp.:93.99, PWM:33
01:08:23.157 -> Water_Temp.:93.99, PWM:33
01:08:23.952 -> Water_Temp.:93.99, PWM:33
01:08:24.741 -> Water_Temp.:93.99, PWM:33
01:08:25.499 -> Water_Temp.:93.99, PWM:33
01:08:26.296 -> Water_Temp.:93.99, PWM:33
01:08:27.086 -> Water_Temp.:93.99, PWM:33
01:08:27.842 -> Water_Temp.:93.99, PWM:33
01:08:28.638 -> Water_Temp.:93.99, PWM:33
01:08:30.156 -> Water_Temp.:93.99, PWM:33
01:08:31.707 -> Water_Temp.:93.99, PWM:33
01:08:33.257 -> Water_Temp.:93.99, PWM:33
01:08:34.047 -> Water_Temp.:93.99, PWM:33
01:08:34.838 -> Water_Temp.:93.99, PWM:33
01:08:35.598 -> Water_Temp.:93.99, PWM:33
01:08:36.396 -> Water_Temp.:93.99, PWM:33
01:08:37.149 -> Water_Temp.:93.99, PWM:33
01:08:37.943 -> Water_Temp.:93.99, PWM:33
01:08:39.494 -> Water_Temp.:93.99, PWM:33
01:08:40.288 -> Water_Temp.:93.99, PWM:33
01:08:41.049 -> Water_Temp.:93.99, PWM:33
01:08:44.887 -> Water_Temp.:93.99, PWM:33
01:08:47.228 -> Water_Temp.:93.99, PWM:33
01:08:48.752 -> Water_Temp.:93.99, PWM:33
01:08:49.545 -> Water_Temp.:93.99, PWM:33
01:08:51.100 -> Water_Temp.:93.99, PWM:33
01:08:53.457 -> Water_Temp.:93.99, PWM:33
01:08:55.011 -> Water_Temp.:93.99, PWM:33
01:08:55.772 -> Water_Temp.:93.99, PWM:33
01:08:56.568 -> Water_Temp.:93.99, PWM:33
01:08:57.332 -> Water_Temp.:93.99, PWM:33
01:08:58.127 -> Water_Temp.:93.99, PWM:33
01:08:58.887 -> Water_Temp.:93.99, PWM:33
01:08:59.685 -> Water_Temp.:93.99, PWM:33
01:09:00.447 -> Water_Temp.:93.99, PWM:33
01:09:02.032 -> Water_Temp.:93.99, PWM:33
01:09:02.793 -> Water_Temp.:93.99, PWM:33
01:09:04.347 -> Water_Temp.:93.99, PWM:33
01:09:05.140 -> Water_Temp.:93.99, PWM:33
01:09:06.692 -> Water_Temp.:94.10, PWM:33
01:09:07.449 -> Water_Temp.:94.10, PWM:33
01:09:10.520 -> Water_Temp.:94.10, PWM:33
01:09:11.315 -> Water_Temp.:94.10, PWM:33
01:09:12.106 -> Water_Temp.:94.10, PWM:33
01:09:12.866 -> Water_Temp.:94.10, PWM:33
01:09:13.663 -> Water_Temp.:94.10, PWM:33
01:09:14.426 -> Water_Temp.:94.10, PWM:33
01:09:16.009 -> Water_Temp.:94.10, PWM:33
01:09:18.294 -> Water_Temp.:94.21, PWM:33
01:09:19.846 -> Water_Temp.:94.21, PWM:33
01:09:21.400 -> Water_Temp.:94.21, PWM:33
01:09:23.676 -> Water_Temp.:94.21, PWM:33
01:09:24.471 -> Water_Temp.:94.21, PWM:33
01:09:26.026 -> Water_Temp.:94.21, PWM:33
01:09:26.822 -> Water_Temp.:94.21, PWM:33
01:09:28.376 -> Water_Temp.:94.32, PWM:33
01:09:29.899 -> Water_Temp.:94.32, PWM:33
01:09:30.695 -> Water_Temp.:94.32, PWM:33
01:09:31.484 -> Water_Temp.:94.32, PWM:33
01:09:32.242 -> Water_Temp.:94.32, PWM:33
01:09:34.555 -> Water_Temp.:94.32, PWM:33
01:09:36.112 -> Water_Temp.:94.32, PWM:33
01:09:36.868 -> Water_Temp.:94.32, PWM:33
01:09:38.420 -> Water_Temp.:94.32, PWM:33
01:09:42.310 -> Water_Temp.:94.32, PWM:33
01:09:43.106 -> Water_Temp.:94.32, PWM:33
01:09:43.863 -> Water_Temp.:94.32, PWM:33
01:09:45.415 -> Water_Temp.:94.32, PWM:33
01:09:46.175 -> Water_Temp.:94.32, PWM:33
01:09:50.797 -> Water_Temp.:94.32, PWM:33
01:09:51.559 -> Water_Temp.:94.32, PWM:33
01:09:53.871 -> Water_Temp.:94.44, PWM:33
01:09:54.665 -> Water_Temp.:94.32, PWM:33
01:09:56.224 -> Water_Temp.:94.32, PWM:33
01:09:56.984 -> Water_Temp.:94.44, PWM:33
01:09:57.780 -> Water_Temp.:94.44, PWM:33
01:09:58.543 -> Water_Temp.:94.44, PWM:33
01:10:00.858 -> Water_Temp.:94.44, PWM:33
01:10:01.651 -> Water_Temp.:94.44, PWM:33
01:10:02.412 -> Water_Temp.:94.44, PWM:33
01:10:03.213 -> Water_Temp.:94.44, PWM:33
01:10:07.810 -> Water_Temp.:94.44, PWM:33
01:10:08.572 -> Water_Temp.:94.44, PWM:33
01:10:09.368 -> Water_Temp.:94.44, PWM:33
01:10:11.650 -> Water_Temp.:94.44, PWM:33
01:10:12.450 -> Water_Temp.:94.44, PWM:33
01:10:13.213 -> Water_Temp.:94.44, PWM:33
01:10:14.007 -> Water_Temp.:94.55, PWM:33
01:10:15.530 -> Water_Temp.:94.55, PWM:33
01:10:17.117 -> Water_Temp.:94.55, PWM:33
01:10:17.877 -> Water_Temp.:94.55, PWM:33
01:10:18.672 -> Water_Temp.:94.55, PWM:33
01:10:20.959 -> Water_Temp.:94.55, PWM:33
01:10:21.750 -> Water_Temp.:94.55, PWM:33
01:10:22.543 -> Water_Temp.:94.55, PWM:33
01:10:23.306 -> Water_Temp.:94.55, PWM:33
01:10:25.657 -> Water_Temp.:94.55, PWM:33
01:10:26.417 -> Water_Temp.:94.55, PWM:33
01:10:27.213 -> Water_Temp.:94.55, PWM:33
01:10:28.008 -> Water_Temp.:94.55, PWM:33
01:10:28.772 -> Water_Temp.:94.55, PWM:33
01:10:31.124 -> Water_Temp.:94.55, PWM:33
01:10:31.882 -> Water_Temp.:94.66, PWM:33
01:10:33.437 -> Water_Temp.:94.66, PWM:33
01:10:34.234 -> Water_Temp.:94.66, PWM:33
01:10:38.025 -> Water_Temp.:94.77, PWM:33
01:10:38.822 -> Water_Temp.:94.77, PWM:33
01:10:40.382 -> Water_Temp.:94.77, PWM:33
01:10:42.693 -> Water_Temp.:94.77, PWM:33
01:10:43.486 -> Water_Temp.:94.77, PWM:33
01:10:45.045 -> Water_Temp.:94.77, PWM:33
01:10:45.806 -> Water_Temp.:94.77, PWM:33
01:10:46.602 -> Water_Temp.:94.77, PWM:33
01:10:47.363 -> Water_Temp.:94.77, PWM:33
01:10:48.920 -> Water_Temp.:94.77, PWM:33
01:10:49.681 -> Water_Temp.:94.77, PWM:33
01:10:50.474 -> Water_Temp.:94.77, PWM:33
01:10:51.237 -> Water_Temp.:94.77, PWM:33
01:10:52.028 -> Water_Temp.:94.77, PWM:33
01:10:52.790 -> Water_Temp.:94.77, PWM:33
01:10:53.588 -> Water_Temp.:94.77, PWM:33
01:10:55.115 -> Water_Temp.:94.77, PWM:33
01:10:55.905 -> Water_Temp.:94.77, PWM:33
01:10:56.667 -> Water_Temp.:94.77, PWM:33
01:10:59.776 -> Water_Temp.:94.77, PWM:33
01:11:00.538 -> Water_Temp.:94.77, PWM:33
01:11:01.327 -> Water_Temp.:94.77, PWM:33
01:11:03.645 -> Water_Temp.:94.77, PWM:33
01:11:04.440 -> Water_Temp.:94.77, PWM:33
01:11:05.200 -> Water_Temp.:94.77, PWM:33
01:11:06.756 -> Water_Temp.:94.77, PWM:33
01:11:09.102 -> Water_Temp.:94.77, PWM:33
01:11:11.423 -> Water_Temp.:94.77, PWM:33
01:11:12.177 -> Water_Temp.:94.77, PWM:33
01:11:12.973 -> Water_Temp.:94.77, PWM:33
01:11:13.769 -> Water_Temp.:94.77, PWM:33
01:11:16.061 -> Water_Temp.:94.77, PWM:33
01:11:16.853 -> Water_Temp.:94.77, PWM:33
01:11:17.616 -> Water_Temp.:94.77, PWM:33
01:11:19.164 -> Water_Temp.:94.77, PWM:33
01:11:19.958 -> Water_Temp.:94.77, PWM:33
01:11:20.719 -> Water_Temp.:94.77, PWM:33
01:11:23.033 -> Water_Temp.:94.77, PWM:33
01:11:23.793 -> Water_Temp.:94.77, PWM:33
01:11:24.589 -> Water_Temp.:94.89, PWM:33
01:11:25.351 -> Water_Temp.:94.89, PWM:33
01:11:26.145 -> Water_Temp.:94.89, PWM:33
01:11:26.933 -> Water_Temp.:94.89, PWM:33
01:11:27.695 -> Water_Temp.:94.89, PWM:33
01:11:28.490 -> Water_Temp.:94.89, PWM:33
01:11:29.254 -> Water_Temp.:94.89, PWM:33
01:11:30.048 -> Water_Temp.:94.89, PWM:33
01:11:32.396 -> Water_Temp.:94.89, PWM:33
01:11:33.158 -> Water_Temp.:94.89, PWM:33
01:11:33.953 -> Water_Temp.:94.89, PWM:33
01:11:34.713 -> Water_Temp.:94.89, PWM:33
01:11:35.510 -> Water_Temp.:94.89, PWM:33
01:11:36.300 -> Water_Temp.:94.89, PWM:33
01:11:39.375 -> Water_Temp.:95.00, PWM:33
01:11:40.906 -> Water_Temp.:95.00, PWM:33
01:11:41.704 -> Water_Temp.:95.00, PWM:33
01:11:44.748 -> Water_Temp.:95.00, PWM:33
01:11:45.545 -> Water_Temp.:95.00, PWM:33
01:11:46.303 -> Water_Temp.:95.11, PWM:33
01:11:47.099 -> Water_Temp.:95.11, PWM:33
01:11:49.415 -> Water_Temp.:95.11, PWM:33
01:11:50.210 -> Water_Temp.:95.11, PWM:33
01:11:51.728 -> Water_Temp.:95.11, PWM:33
01:11:53.285 -> Water_Temp.:95.11, PWM:33
01:11:54.081 -> Water_Temp.:95.11, PWM:33
01:11:55.598 -> Water_Temp.:95.11, PWM:33
01:11:56.396 -> Water_Temp.:95.11, PWM:33
01:11:57.188 -> Water_Temp.:95.11, PWM:33
01:11:57.949 -> Water_Temp.:95.11, PWM:33
01:11:59.505 -> Water_Temp.:95.11, PWM:33
01:12:01.065 -> Water_Temp.:95.11, PWM:33
01:12:02.624 -> Water_Temp.:95.11, PWM:33
01:12:03.389 -> Water_Temp.:95.11, PWM:33
01:12:05.707 -> Water_Temp.:95.11, PWM:33
01:12:06.472 -> Water_Temp.:95.11, PWM:33
01:12:07.267 -> Water_Temp.:95.11, PWM:33
01:12:08.822 -> Water_Temp.:95.11, PWM:33
01:12:10.347 -> Water_Temp.:95.11, PWM:33
01:12:11.898 -> Water_Temp.:95.11, PWM:33
01:12:13.422 -> Water_Temp.:95.22, PWM:33
01:12:14.215 -> Water_Temp.:95.22, PWM:33
01:12:15.771 -> Water_Temp.:95.22, PWM:33
01:12:16.568 -> Water_Temp.:95.22, PWM:33
01:12:17.332 -> Water_Temp.:95.22, PWM:33
01:12:18.126 -> Water_Temp.:95.22, PWM:33
01:12:18.890 -> Water_Temp.:95.22, PWM:33
01:12:19.685 -> Water_Temp.:95.22, PWM:33
01:12:20.442 -> Water_Temp.:95.22, PWM:33
01:12:22.784 -> Water_Temp.:95.22, PWM:33
01:12:23.543 -> Water_Temp.:95.22, PWM:33
01:12:25.897 -> Water_Temp.:95.22, PWM:33
01:12:26.656 -> Water_Temp.:95.22, PWM:33
01:12:28.218 -> Water_Temp.:95.22, PWM:33
01:12:29.778 -> Water_Temp.:95.34, PWM:33
01:12:30.539 -> Water_Temp.:95.34, PWM:33
01:12:31.332 -> Water_Temp.:95.34, PWM:33
01:12:32.092 -> Water_Temp.:95.34, PWM:33
01:12:33.647 -> Water_Temp.:95.34, PWM:33
01:12:35.199 -> Water_Temp.:95.34, PWM:33
01:12:36.752 -> Water_Temp.:95.34, PWM:33
01:12:37.505 -> Water_Temp.:95.34, PWM:33
01:12:39.825 -> Water_Temp.:95.45, PWM:33
01:12:40.623 -> Water_Temp.:95.45, PWM:33
01:12:42.179 -> Water_Temp.:95.34, PWM:33
01:12:42.938 -> Water_Temp.:95.34, PWM:33
01:12:43.733 -> Water_Temp.:95.34, PWM:33
01:12:45.283 -> Water_Temp.:95.34, PWM:33
01:12:46.045 -> Water_Temp.:95.34, PWM:33
01:12:48.359 -> Water_Temp.:95.34, PWM:33
01:12:49.152 -> Water_Temp.:95.34, PWM:33
01:12:49.912 -> Water_Temp.:95.34, PWM:33
01:12:50.709 -> Water_Temp.:95.45, PWM:33
01:12:52.256 -> Water_Temp.:95.45, PWM:33
01:12:53.021 -> Water_Temp.:95.45, PWM:33
01:12:53.816 -> Water_Temp.:95.45, PWM:33
01:12:54.580 -> Water_Temp.:95.45, PWM:33
01:12:56.128 -> Water_Temp.:95.45, PWM:33
01:12:56.919 -> Water_Temp.:95.45, PWM:33
01:12:59.235 -> Water_Temp.:95.56, PWM:33
01:13:00.033 -> Water_Temp.:95.56, PWM:33
01:13:00.797 -> Water_Temp.:95.56, PWM:33
01:13:01.595 -> Water_Temp.:95.56, PWM:33
01:13:03.118 -> Water_Temp.:95.56, PWM:33
01:13:04.668 -> Water_Temp.:95.56, PWM:33
01:13:06.982 -> Water_Temp.:95.56, PWM:33
01:13:07.776 -> Water_Temp.:95.56, PWM:33
01:13:08.572 -> Water_Temp.:95.56, PWM:33
01:13:09.329 -> Water_Temp.:95.56, PWM:33
01:13:11.646 -> Water_Temp.:95.68, PWM:33
01:13:12.438 -> Water_Temp.:95.56, PWM:33
01:13:13.201 -> Water_Temp.:95.56, PWM:33
01:13:13.993 -> Water_Temp.:95.68, PWM:33
01:13:14.751 -> Water_Temp.:95.68, PWM:33
01:13:15.548 -> Water_Temp.:95.68, PWM:33
01:13:16.312 -> Water_Temp.:95.68, PWM:33
01:13:17.108 -> Water_Temp.:95.68, PWM:33
01:13:18.665 -> Water_Temp.:95.68, PWM:33
01:13:19.427 -> Water_Temp.:95.68, PWM:33
01:13:20.190 -> Water_Temp.:95.68, PWM:33
01:13:20.988 -> Water_Temp.:97.36, PWM:33
01:13:21.784 -> Water_Temp.:99.39, PWM:33
01:13:22.540 -> Water_Temp.:96.91, PWM:33
01:13:23.335 -> Water_Temp.:96.57, PWM:33
01:13:24.094 -> Water_Temp.:96.35, PWM:33
01:13:27.190 -> Water_Temp.:96.01, PWM:33
01:13:27.982 -> Water_Temp.:96.01, PWM:33
01:13:28.744 -> Water_Temp.:95.90, PWM:33
01:13:29.540 -> Water_Temp.:95.90, PWM:33
01:13:30.302 -> Water_Temp.:95.90, PWM:33
01:13:31.099 -> Water_Temp.:95.90, PWM:33
01:13:31.892 -> Water_Temp.:95.79, PWM:33
01:13:32.653 -> Water_Temp.:95.79, PWM:33
01:13:33.448 -> Water_Temp.:95.79, PWM:33
01:13:34.210 -> Water_Temp.:95.79, PWM:33
01:13:35.004 -> Water_Temp.:95.79, PWM:33
01:13:35.763 -> Water_Temp.:95.79, PWM:33
01:13:38.903 -> Water_Temp.:95.79, PWM:33
01:13:40.427 -> Water_Temp.:95.79, PWM:33
01:13:41.226 -> Water_Temp.:95.79, PWM:33
01:13:42.781 -> Water_Temp.:95.79, PWM:33
01:13:43.541 -> Water_Temp.:95.79, PWM:33
01:13:46.615 -> Water_Temp.:95.90, PWM:33
01:13:47.410 -> Water_Temp.:95.90, PWM:33
01:13:48.202 -> Water_Temp.:95.90, PWM:33
01:13:48.959 -> Water_Temp.:95.79, PWM:33
01:13:50.517 -> Water_Temp.:95.79, PWM:33
01:13:51.306 -> Water_Temp.:95.79, PWM:33
01:13:55.140 -> Water_Temp.:95.79, PWM:33
01:13:55.934 -> Water_Temp.:95.79, PWM:33
01:13:56.732 -> Water_Temp.:95.79, PWM:33
01:13:57.492 -> Water_Temp.:95.79, PWM:33
01:13:59.803 -> Water_Temp.:95.79, PWM:33
01:14:00.600 -> Water_Temp.:95.79, PWM:33
01:14:01.360 -> Water_Temp.:95.79, PWM:33
01:14:02.152 -> Water_Temp.:95.79, PWM:33
01:14:03.669 -> Water_Temp.:95.79, PWM:33
01:14:04.459 -> Water_Temp.:95.79, PWM:33
01:14:05.253 -> Water_Temp.:95.79, PWM:33
01:14:08.324 -> Water_Temp.:95.79, PWM:33
01:14:09.843 -> Water_Temp.:95.79, PWM:33
01:14:10.635 -> Water_Temp.:95.79, PWM:33
01:14:11.429 -> Water_Temp.:95.79, PWM:33
01:14:12.188 -> Water_Temp.:95.79, PWM:33
01:14:13.739 -> Water_Temp.:95.90, PWM:33
01:14:15.287 -> Water_Temp.:95.90, PWM:33
01:14:16.077 -> Water_Temp.:95.90, PWM:33
01:14:16.835 -> Water_Temp.:95.90, PWM:33
01:14:17.629 -> Water_Temp.:95.90, PWM:33
01:14:20.705 -> Water_Temp.:95.90, PWM:33
01:14:24.539 -> Water_Temp.:95.90, PWM:33
01:14:26.093 -> Water_Temp.:95.90, PWM:33
01:14:26.883 -> Water_Temp.:95.90, PWM:33
01:14:28.436 -> Water_Temp.:95.90, PWM:33
01:14:29.235 -> Water_Temp.:95.90, PWM:33
01:14:29.989 -> Water_Temp.:95.90, PWM:33
01:14:33.108 -> Water_Temp.:95.90, PWM:33
01:14:37.697 -> Water_Temp.:95.90, PWM:33
01:14:38.492 -> Water_Temp.:95.90, PWM:33
01:14:39.252 -> Water_Temp.:95.90, PWM:33
01:14:40.813 -> Water_Temp.:95.90, PWM:33
01:14:41.609 -> Water_Temp.:95.90, PWM:33
01:14:43.926 -> Water_Temp.:95.90, PWM:33
01:14:46.208 -> Water_Temp.:95.90, PWM:33
01:14:47.767 -> Water_Temp.:95.90, PWM:33
01:14:48.530 -> Water_Temp.:95.90, PWM:33
01:14:49.321 -> Water_Temp.:95.90, PWM:33
01:14:50.117 -> Water_Temp.:95.90, PWM:33
01:14:50.880 -> Water_Temp.:95.90, PWM:33
01:14:51.672 -> Water_Temp.:95.90, PWM:33
01:14:53.221 -> Water_Temp.:99.72, PWM:33
01:14:54.774 -> Water_Temp.:96.80, PWM:33
01:14:56.329 -> Water_Temp.:96.46, PWM:33
01:14:57.118 -> Water_Temp.:96.35, PWM:33
01:14:57.879 -> Water_Temp.:96.35, PWM:33
01:14:58.675 -> Water_Temp.:96.24, PWM:33
01:14:59.432 -> Water_Temp.:96.24, PWM:33
01:15:00.225 -> Water_Temp.:96.24, PWM:33
01:15:00.987 -> Water_Temp.:96.24, PWM:33
01:15:02.573 -> Water_Temp.:96.12, PWM:33
01:15:04.892 -> Water_Temp.:96.12, PWM:33
01:15:05.653 -> Water_Temp.:96.12, PWM:33
01:15:07.208 -> Water_Temp.:96.12, PWM:33
01:15:07.970 -> Water_Temp.:96.12, PWM:33
01:15:11.849 -> Water_Temp.:96.01, PWM:33
01:15:14.168 -> Water_Temp.:96.01, PWM:33
01:15:14.963 -> Water_Temp.:96.12, PWM:33
01:15:15.724 -> Water_Temp.:96.12, PWM:33
01:15:16.519 -> Water_Temp.:96.12, PWM:33
01:15:18.078 -> Water_Temp.:97.70, PWM:33
01:15:19.602 -> Water_Temp.:97.36, PWM:33
01:15:20.364 -> Water_Temp.:97.02, PWM:33
01:15:21.917 -> Water_Temp.:96.57, PWM:33
01:15:24.235 -> Water_Temp.:96.35, PWM:33
01:15:25.787 -> Water_Temp.:96.35, PWM:33
01:15:26.582 -> Water_Temp.:96.24, PWM:33
01:15:28.892 -> Water_Temp.:96.24, PWM:33
01:15:31.242 -> Water_Temp.:96.24, PWM:33
01:15:32.034 -> Water_Temp.:96.24, PWM:33
01:15:33.587 -> Water_Temp.:96.24, PWM:33
01:15:34.347 -> Water_Temp.:96.24, PWM:33
01:15:35.139 -> Water_Temp.:96.24, PWM:33
01:15:36.689 -> Water_Temp.:96.24, PWM:33
01:15:38.240 -> Water_Temp.:96.24, PWM:33
01:15:39.027 -> Water_Temp.:96.24, PWM:33
01:15:40.553 -> Water_Temp.:96.24, PWM:33
01:15:41.348 -> Water_Temp.:96.24, PWM:33
01:15:42.104 -> Water_Temp.:96.24, PWM:33
01:15:42.893 -> Water_Temp.:97.92, PWM:33
01:15:43.689 -> Water_Temp.:97.92, PWM:33
01:15:45.241 -> Water_Temp.:97.14, PWM:33
01:15:46.035 -> Water_Temp.:96.91, PWM:33
01:15:49.873 -> Water_Temp.:96.57, PWM:33
01:15:52.147 -> Water_Temp.:96.46, PWM:33
01:15:52.939 -> Water_Temp.:96.46, PWM:33
01:15:53.736 -> Water_Temp.:96.46, PWM:33
01:15:54.498 -> Water_Temp.:96.46, PWM:33
01:15:55.293 -> Water_Temp.:96.35, PWM:33
01:15:57.607 -> Water_Temp.:96.35, PWM:33
01:15:59.920 -> Water_Temp.:96.35, PWM:33
01:16:00.682 -> Water_Temp.:96.35, PWM:33
01:16:01.477 -> Water_Temp.:96.35, PWM:33
01:16:02.237 -> Water_Temp.:96.35, PWM:33
01:16:03.033 -> Water_Temp.:96.35, PWM:33
01:16:03.795 -> Water_Temp.:96.35, PWM:33
01:16:04.588 -> Water_Temp.:96.35, PWM:33
01:16:05.383 -> Water_Temp.:96.35, PWM:33
01:16:06.146 -> Water_Temp.:96.35, PWM:33
01:16:06.939 -> Water_Temp.:96.35, PWM:33
01:16:07.698 -> Water_Temp.:96.35, PWM:33
01:16:08.493 -> Water_Temp.:96.35, PWM:33
01:16:10.811 -> Water_Temp.:96.35, PWM:33
01:16:11.602 -> Water_Temp.:96.35, PWM:33
01:16:13.160 -> Water_Temp.:96.35, PWM:33
01:16:13.952 -> Water_Temp.:96.35, PWM:33
01:16:14.712 -> Water_Temp.:96.35, PWM:33
01:16:16.298 -> Water_Temp.:96.35, PWM:33
01:16:17.853 -> Water_Temp.:96.35, PWM:33
01:16:18.614 -> Water_Temp.:96.35, PWM:33
01:16:21.724 -> Water_Temp.:96.46, PWM:33
01:16:24.044 -> Water_Temp.:96.46, PWM:33
01:16:26.359 -> Water_Temp.:96.46, PWM:33
01:16:27.155 -> Water_Temp.:96.46, PWM:33
01:16:27.919 -> Water_Temp.:96.46, PWM:33
01:16:28.716 -> Water_Temp.:96.46, PWM:33
01:16:29.477 -> Water_Temp.:96.46, PWM:33
01:16:31.034 -> Water_Temp.:96.46, PWM:33
01:16:31.793 -> Water_Temp.:96.46, PWM:33
01:16:32.581 -> Water_Temp.:96.46, PWM:33
01:16:33.375 -> Water_Temp.:98.15, PWM:33
01:16:34.132 -> Water_Temp.:100.17, PWM:33
01:16:36.480 -> Water_Temp.:97.14, PWM:33
01:16:37.243 -> Water_Temp.:97.02, PWM:33
01:16:40.354 -> Water_Temp.:96.80, PWM:33
01:16:41.111 -> Water_Temp.:96.69, PWM:33
01:16:41.900 -> Water_Temp.:96.69, PWM:33
01:16:42.691 -> Water_Temp.:96.69, PWM:33
01:16:43.450 -> Water_Temp.:96.69, PWM:33
01:16:44.245 -> Water_Temp.:96.69, PWM:33
01:16:45.003 -> Water_Temp.:96.69, PWM:33
01:16:45.796 -> Water_Temp.:96.69, PWM:33
01:16:46.588 -> Water_Temp.:96.69, PWM:33
01:16:47.350 -> Water_Temp.:96.57, PWM:33
01:16:51.176 -> Water_Temp.:96.57, PWM:33
01:16:52.728 -> Water_Temp.:96.57, PWM:33
01:16:53.490 -> Water_Temp.:96.57, PWM:33
01:16:55.835 -> Water_Temp.:96.57, PWM:33
01:16:58.149 -> Water_Temp.:96.57, PWM:33
01:16:59.704 -> Water_Temp.:96.57, PWM:33
01:17:02.019 -> Water_Temp.:96.57, PWM:33
01:17:02.781 -> Water_Temp.:96.57, PWM:33
01:17:03.577 -> Water_Temp.:96.57, PWM:33
01:17:04.338 -> Water_Temp.:96.57, PWM:33
01:17:05.135 -> Water_Temp.:96.57, PWM:33
01:17:05.900 -> Water_Temp.:96.57, PWM:33
01:17:06.693 -> Water_Temp.:96.69, PWM:33
01:17:07.489 -> Water_Temp.:96.69, PWM:33
01:17:09.003 -> Water_Temp.:96.69, PWM:33
01:17:11.321 -> Water_Temp.:96.69, PWM:33
01:17:12.109 -> Water_Temp.:96.69, PWM:33
01:17:13.662 -> Water_Temp.:96.69, PWM:33
01:17:14.423 -> Water_Temp.:96.69, PWM:33
01:17:15.217 -> Water_Temp.:96.69, PWM:33
01:17:15.976 -> Water_Temp.:96.69, PWM:33
01:17:16.770 -> Water_Temp.:96.69, PWM:33
01:17:17.568 -> Water_Temp.:96.69, PWM:33
01:17:18.325 -> Water_Temp.:96.69, PWM:33
01:17:19.121 -> Water_Temp.:96.69, PWM:33
01:17:19.883 -> Water_Temp.:96.69, PWM:33
01:17:21.441 -> Water_Temp.:96.69, PWM:33
01:17:22.234 -> Water_Temp.:96.69, PWM:33
01:17:24.516 -> Water_Temp.:96.69, PWM:33
01:17:26.073 -> Water_Temp.:96.69, PWM:33
01:17:28.411 -> Water_Temp.:96.69, PWM:33
01:17:29.202 -> Water_Temp.:96.69, PWM:33
01:17:29.988 -> Water_Temp.:96.69, PWM:33
01:17:30.749 -> Water_Temp.:96.69, PWM:33
01:17:32.300 -> Water_Temp.:96.69, PWM:33
01:17:33.855 -> Water_Temp.:96.80, PWM:33
01:17:34.619 -> Water_Temp.:96.80, PWM:33
01:17:36.171 -> Water_Temp.:96.80, PWM:33
01:17:36.963 -> Water_Temp.:96.80, PWM:33
01:17:37.756 -> Water_Temp.:96.80, PWM:33
01:17:38.520 -> Water_Temp.:96.80, PWM:33
01:17:39.314 -> Water_Temp.:96.80, PWM:33
01:17:40.074 -> Water_Temp.:96.80, PWM:33
01:17:40.872 -> Water_Temp.:96.80, PWM:33
01:17:41.629 -> Water_Temp.:96.80, PWM:33
01:17:43.944 -> Water_Temp.:96.80, PWM:33
01:17:44.705 -> Water_Temp.:96.80, PWM:33
01:17:45.497 -> Water_Temp.:96.80, PWM:33
01:17:46.288 -> Water_Temp.:96.80, PWM:33
01:17:47.837 -> Water_Temp.:96.80, PWM:33
01:17:49.390 -> Water_Temp.:96.80, PWM:33
01:17:50.946 -> Water_Temp.:96.80, PWM:33
01:17:52.498 -> Water_Temp.:96.91, PWM:33
01:17:53.254 -> Water_Temp.:96.91, PWM:33
01:17:54.046 -> Water_Temp.:96.91, PWM:33
01:17:54.806 -> Water_Temp.:96.91, PWM:33
01:17:57.159 -> Water_Temp.:96.91, PWM:33
01:17:57.952 -> Water_Temp.:96.91, PWM:33
01:17:58.714 -> Water_Temp.:96.91, PWM:33
01:17:59.511 -> Water_Temp.:96.91, PWM:33
01:18:01.033 -> Water_Temp.:96.91, PWM:33
01:18:02.586 -> Water_Temp.:96.91, PWM:33
01:18:03.379 -> Water_Temp.:96.91, PWM:33
01:18:04.139 -> Water_Temp.:96.91, PWM:33
01:18:04.938 -> Water_Temp.:96.91, PWM:33
01:18:05.699 -> Water_Temp.:96.91, PWM:33
01:18:06.490 -> Water_Temp.:96.91, PWM:33
01:18:07.287 -> Water_Temp.:96.91, PWM:33
01:18:08.049 -> Water_Temp.:96.91, PWM:33
01:18:10.366 -> Water_Temp.:96.91, PWM:33
01:18:11.164 -> Water_Temp.:96.91, PWM:33
01:18:12.716 -> Water_Temp.:96.91, PWM:33
01:18:14.271 -> Water_Temp.:96.91, PWM:33
01:18:15.067 -> Water_Temp.:96.91, PWM:33
01:18:15.830 -> Water_Temp.:96.91, PWM:33
01:18:16.623 -> Water_Temp.:96.91, PWM:33
01:18:17.384 -> Water_Temp.:96.91, PWM:33
01:18:18.942 -> Water_Temp.:96.91, PWM:33
01:18:19.699 -> Water_Temp.:96.91, PWM:33
01:18:21.286 -> Water_Temp.:96.91, PWM:33
01:18:22.045 -> Water_Temp.:96.91, PWM:33
01:18:22.841 -> Water_Temp.:96.91, PWM:33
01:18:24.399 -> Water_Temp.:97.02, PWM:33
01:18:25.947 -> Water_Temp.:96.91, PWM:33
01:18:26.740 -> Water_Temp.:97.02, PWM:33
01:18:27.497 -> Water_Temp.:97.02, PWM:33
01:18:28.288 -> Water_Temp.:97.02, PWM:33
01:18:29.846 -> Water_Temp.:97.02, PWM:33
01:18:30.604 -> Water_Temp.:97.02, PWM:33
01:18:32.155 -> Water_Temp.:97.02, PWM:33
01:18:33.714 -> Water_Temp.:97.02, PWM:33
01:18:34.474 -> Water_Temp.:97.02, PWM:33
01:18:35.272 -> Water_Temp.:97.02, PWM:33
01:18:36.031 -> Water_Temp.:97.02, PWM:33
01:18:36.826 -> Water_Temp.:97.02, PWM:33
01:18:37.585 -> Water_Temp.:97.02, PWM:33
01:18:38.380 -> Water_Temp.:97.02, PWM:33
01:18:39.937 -> Water_Temp.:97.02, PWM:33
01:18:40.733 -> Water_Temp.:97.02, PWM:33
01:18:41.492 -> Water_Temp.:97.02, PWM:33
01:18:42.286 -> Water_Temp.:97.02, PWM:33
01:18:43.046 -> Water_Temp.:97.02, PWM:33
01:18:43.843 -> Water_Temp.:97.02, PWM:33
01:18:44.602 -> Water_Temp.:97.02, PWM:33
01:18:45.397 -> Water_Temp.:97.02, PWM:33
01:18:46.190 -> Water_Temp.:97.02, PWM:33
01:18:47.705 -> Water_Temp.:97.02, PWM:33
01:18:48.497 -> Water_Temp.:97.02, PWM:33
01:18:49.258 -> Water_Temp.:97.02, PWM:33
01:18:50.051 -> Water_Temp.:97.02, PWM:33
01:18:50.845 -> Water_Temp.:97.02, PWM:33
01:18:51.604 -> Water_Temp.:97.02, PWM:33
01:18:52.403 -> Water_Temp.:97.02, PWM:33
01:18:53.918 -> Water_Temp.:97.02, PWM:33
01:18:54.713 -> Water_Temp.:97.02, PWM:33
01:18:55.472 -> Water_Temp.:97.02, PWM:33
01:18:56.262 -> Water_Temp.:97.02, PWM:33
01:18:57.812 -> Water_Temp.:97.02, PWM:33
01:18:58.574 -> Water_Temp.:97.02, PWM:33
01:18:59.370 -> Water_Temp.:97.02, PWM:33
01:19:00.894 -> Water_Temp.:97.02, PWM:33
01:19:03.241 -> Water_Temp.:97.02, PWM:33
01:19:04.037 -> Water_Temp.:97.02, PWM:33
01:19:05.561 -> Water_Temp.:97.02, PWM:33
01:19:06.356 -> Water_Temp.:97.02, PWM:33
01:19:07.112 -> Water_Temp.:97.14, PWM:33
01:19:07.902 -> Water_Temp.:97.14, PWM:33
01:19:09.454 -> Water_Temp.:97.14, PWM:33
01:19:10.212 -> Water_Temp.:97.14, PWM:33
01:19:12.555 -> Water_Temp.:97.14, PWM:33
01:19:13.350 -> Water_Temp.:97.14, PWM:33
01:19:14.108 -> Water_Temp.:97.14, PWM:33
01:19:14.902 -> Water_Temp.:97.14, PWM:33
01:19:15.664 -> Water_Temp.:97.14, PWM:33
01:19:16.459 -> Water_Temp.:97.14, PWM:33
01:19:18.801 -> Water_Temp.:97.14, PWM:33
01:19:19.562 -> Water_Temp.:97.14, PWM:33
01:19:20.351 -> Water_Temp.:97.14, PWM:33
01:19:21.910 -> Water_Temp.:97.14, PWM:33
01:19:23.463 -> Water_Temp.:97.25, PWM:33
01:19:26.542 -> Water_Temp.:97.25, PWM:33
01:19:27.337 -> Water_Temp.:97.25, PWM:33
01:19:28.888 -> Water_Temp.:97.25, PWM:33
01:19:30.403 -> Water_Temp.:97.25, PWM:33
01:19:31.201 -> Water_Temp.:97.25, PWM:33
01:19:34.307 -> Water_Temp.:97.25, PWM:33
01:19:35.067 -> Water_Temp.:97.25, PWM:33
01:19:36.618 -> Water_Temp.:97.25, PWM:33
01:19:37.377 -> Water_Temp.:97.25, PWM:33
01:19:38.173 -> Water_Temp.:97.25, PWM:33
01:19:38.934 -> Water_Temp.:97.25, PWM:33
01:19:39.733 -> Water_Temp.:97.25, PWM:33
01:19:42.804 -> Water_Temp.:97.25, PWM:33
01:19:43.564 -> Water_Temp.:97.36, PWM:33
01:19:44.361 -> Water_Temp.:97.36, PWM:33
01:19:45.120 -> Water_Temp.:97.36, PWM:33
01:19:46.678 -> Water_Temp.:97.36, PWM:33
01:19:47.436 -> Water_Temp.:97.36, PWM:33
01:19:48.993 -> Water_Temp.:97.36, PWM:33
01:19:49.751 -> Water_Temp.:97.36, PWM:33
01:19:50.546 -> Water_Temp.:97.36, PWM:33
01:19:51.346 -> Water_Temp.:97.36, PWM:33
01:19:52.104 -> Water_Temp.:97.36, PWM:33
01:19:55.217 -> Water_Temp.:97.36, PWM:33
01:19:56.768 -> Water_Temp.:97.36, PWM:33
01:19:58.358 -> Water_Temp.:97.36, PWM:33
01:19:59.118 -> Water_Temp.:97.36, PWM:33
01:20:00.674 -> Water_Temp.:97.36, PWM:33
01:20:01.469 -> Water_Temp.:97.36, PWM:33
01:20:02.227 -> Water_Temp.:97.36, PWM:33
01:20:03.023 -> Water_Temp.:97.36, PWM:33
01:20:03.785 -> Water_Temp.:97.36, PWM:33
01:20:06.138 -> Water_Temp.:97.36, PWM:33
01:20:06.899 -> Water_Temp.:97.36, PWM:33
01:20:08.457 -> Water_Temp.:97.36, PWM:33
01:20:09.252 -> Water_Temp.:97.36, PWM:33
01:20:10.011 -> Water_Temp.:97.36, PWM:33
01:20:12.363 -> Water_Temp.:97.36, PWM:33
01:20:13.126 -> Water_Temp.:97.36, PWM:33
01:20:13.919 -> Water_Temp.:97.36, PWM:33
01:20:14.680 -> Water_Temp.:97.36, PWM:33
01:20:16.994 -> Water_Temp.:97.47, PWM:33
01:20:17.790 -> Water_Temp.:97.36, PWM:33
01:20:18.550 -> Water_Temp.:97.47, PWM:33
01:20:19.342 -> Water_Temp.:97.47, PWM:33
01:20:21.659 -> Water_Temp.:97.36, PWM:33
01:20:23.213 -> Water_Temp.:97.36, PWM:33
01:20:23.975 -> Water_Temp.:97.36, PWM:33
01:20:24.766 -> Water_Temp.:97.47, PWM:33
01:20:26.317 -> Water_Temp.:97.47, PWM:33
01:20:28.660 -> Water_Temp.:97.47, PWM:33
01:20:29.422 -> Water_Temp.:97.47, PWM:33
01:20:30.980 -> Water_Temp.:97.47, PWM:33
01:20:31.774 -> Water_Temp.:97.47, PWM:33
01:20:32.564 -> Water_Temp.:97.47, PWM:33
01:20:34.081 -> Water_Temp.:97.47, PWM:33
01:20:34.877 -> Water_Temp.:97.47, PWM:33
01:20:35.636 -> Water_Temp.:97.47, PWM:33
01:20:37.192 -> Water_Temp.:97.47, PWM:33
01:20:38.744 -> Water_Temp.:97.47, PWM:33
01:20:40.266 -> Water_Temp.:97.47, PWM:33
01:20:41.063 -> Water_Temp.:97.47, PWM:33
01:20:41.855 -> Water_Temp.:97.47, PWM:33
01:20:42.616 -> Water_Temp.:97.47, PWM:33
01:20:44.172 -> Water_Temp.:97.47, PWM:33
01:20:44.958 -> Water_Temp.:97.47, PWM:33
01:20:45.720 -> Water_Temp.:97.47, PWM:33
01:20:46.515 -> Water_Temp.:97.47, PWM:33
01:20:49.588 -> Water_Temp.:97.59, PWM:33
01:20:50.384 -> Water_Temp.:97.47, PWM:33
01:20:51.176 -> Water_Temp.:97.47, PWM:33
01:20:51.934 -> Water_Temp.:97.47, PWM:33
01:20:52.729 -> Water_Temp.:97.47, PWM:33
01:20:53.492 -> Water_Temp.:97.47, PWM:33
01:20:54.288 -> Water_Temp.:97.47, PWM:33
01:20:55.079 -> Water_Temp.:99.16, PWM:33
01:20:55.838 -> Water_Temp.:101.30, PWM:33
01:20:56.634 -> Water_Temp.:98.82, PWM:33
01:20:57.392 -> Water_Temp.:98.37, PWM:33
01:20:58.186 -> Water_Temp.:98.37, PWM:33
01:20:58.979 -> Water_Temp.:98.04, PWM:33
01:20:59.737 -> Water_Temp.:97.92, PWM:33
01:21:01.289 -> Water_Temp.:97.81, PWM:33
01:21:02.082 -> Water_Temp.:97.70, PWM:33
01:21:02.876 -> Water_Temp.:97.70, PWM:33
01:21:03.640 -> Water_Temp.:97.70, PWM:33
01:21:04.403 -> Water_Temp.:97.70, PWM:33
01:21:05.196 -> Water_Temp.:97.59, PWM:33
01:21:05.993 -> Water_Temp.:97.59, PWM:33
01:21:06.753 -> Water_Temp.:97.59, PWM:33
01:21:07.552 -> Water_Temp.:97.59, PWM:33
01:21:08.314 -> Water_Temp.:97.59, PWM:33
01:21:09.871 -> Water_Temp.:97.59, PWM:33
01:21:10.634 -> Water_Temp.:97.59, PWM:33
01:21:11.429 -> Water_Temp.:97.59, PWM:33
01:21:12.221 -> Water_Temp.:97.59, PWM:33
01:21:13.748 -> Water_Temp.:97.59, PWM:33
01:21:14.540 -> Water_Temp.:97.59, PWM:33
01:21:16.098 -> Water_Temp.:97.59, PWM:33
01:21:16.856 -> Water_Temp.:97.59, PWM:33
01:21:17.650 -> Water_Temp.:97.59, PWM:33
01:21:19.170 -> Water_Temp.:97.59, PWM:33
01:21:19.967 -> Water_Temp.:97.59, PWM:33
01:21:20.729 -> Water_Temp.:97.59, PWM:33
01:21:21.524 -> Water_Temp.:97.59, PWM:33
01:21:23.813 -> Water_Temp.:97.59, PWM:33
01:21:24.610 -> Water_Temp.:97.59, PWM:33
01:21:25.400 -> Water_Temp.:97.59, PWM:33
01:21:26.163 -> Water_Temp.:97.59, PWM:33
01:21:26.925 -> Water_Temp.:97.59, PWM:33
01:21:27.721 -> Water_Temp.:97.59, PWM:33
01:21:28.514 -> Water_Temp.:97.59, PWM:33
01:21:29.268 -> Water_Temp.:97.59, PWM:33
01:21:30.065 -> Water_Temp.:97.59, PWM:33
01:21:31.586 -> Water_Temp.:97.59, PWM:33
01:21:32.381 -> Water_Temp.:97.70, PWM:33
01:21:33.174 -> Water_Temp.:97.70, PWM:33
01:21:33.935 -> Water_Temp.:97.70, PWM:33
01:21:37.799 -> Water_Temp.:97.70, PWM:33
01:21:39.351 -> Water_Temp.:97.70, PWM:33
01:21:40.108 -> Water_Temp.:97.70, PWM:33
01:21:40.905 -> Water_Temp.:97.70, PWM:33
01:21:43.217 -> Water_Temp.:97.70, PWM:33
01:21:44.778 -> Water_Temp.:97.70, PWM:33
01:21:45.535 -> Water_Temp.:97.70, PWM:33
01:21:46.331 -> Water_Temp.:97.70, PWM:33
01:21:49.401 -> Water_Temp.:97.70, PWM:33
01:21:50.189 -> Water_Temp.:97.70, PWM:33
01:21:50.949 -> Water_Temp.:97.70, PWM:33
01:21:51.737 -> Water_Temp.:97.70, PWM:33
01:21:52.529 -> Water_Temp.:97.70, PWM:33
01:21:53.292 -> Water_Temp.:97.70, PWM:33
01:21:54.087 -> Water_Temp.:97.70, PWM:33
01:21:55.636 -> Water_Temp.:97.70, PWM:33
01:21:56.398 -> Water_Temp.:97.70, PWM:33
01:21:57.186 -> Water_Temp.:97.70, PWM:33
01:21:57.947 -> Water_Temp.:97.70, PWM:33
01:21:58.738 -> Water_Temp.:97.70, PWM:33
01:21:59.532 -> Water_Temp.:97.70, PWM:33
01:22:00.291 -> Water_Temp.:97.70, PWM:33
01:22:04.132 -> Water_Temp.:97.81, PWM:33
01:22:04.926 -> Water_Temp.:97.81, PWM:33
01:22:06.449 -> Water_Temp.:97.81, PWM:33
01:22:07.244 -> Water_Temp.:97.70, PWM:33
01:22:08.008 -> Water_Temp.:97.70, PWM:33
01:22:09.561 -> Water_Temp.:97.70, PWM:33
01:22:10.322 -> Water_Temp.:97.70, PWM:33
01:22:11.877 -> Water_Temp.:97.70, PWM:33
01:22:13.440 -> Water_Temp.:97.70, PWM:33
01:22:14.992 -> Water_Temp.:97.70, PWM:33
01:22:17.307 -> Water_Temp.:97.70, PWM:33
01:22:18.101 -> Water_Temp.:97.70, PWM:33
01:22:18.860 -> Water_Temp.:97.70, PWM:33
01:22:19.651 -> Water_Temp.:97.70, PWM:33
01:22:21.207 -> Water_Temp.:97.70, PWM:33
01:22:21.963 -> Water_Temp.:97.70, PWM:33
01:22:22.755 -> Water_Temp.:97.70, PWM:33
01:22:23.517 -> Water_Temp.:97.70, PWM:33
01:22:25.065 -> Water_Temp.:97.70, PWM:33
01:22:25.859 -> Water_Temp.:97.81, PWM:33
01:22:27.417 -> Water_Temp.:97.81, PWM:33
01:22:28.975 -> Water_Temp.:97.81, PWM:33
01:22:29.766 -> Water_Temp.:97.81, PWM:33
01:22:31.291 -> Water_Temp.:97.81, PWM:33
01:22:32.087 -> Water_Temp.:97.81, PWM:33
01:22:32.849 -> Water_Temp.:97.81, PWM:33
01:22:33.645 -> Water_Temp.:97.81, PWM:33
01:22:36.723 -> Water_Temp.:97.81, PWM:33
01:22:39.044 -> Water_Temp.:97.81, PWM:33
01:22:41.369 -> Water_Temp.:97.81, PWM:33
01:22:42.923 -> Water_Temp.:97.81, PWM:33
01:22:43.684 -> Water_Temp.:97.81, PWM:33
01:22:45.275 -> Water_Temp.:97.81, PWM:33
01:22:46.032 -> Water_Temp.:97.81, PWM:33
01:22:48.370 -> Water_Temp.:97.81, PWM:33
01:22:49.162 -> Water_Temp.:97.81, PWM:33
01:22:49.923 -> Water_Temp.:97.81, PWM:33
01:22:51.475 -> Water_Temp.:97.81, PWM:33
01:22:52.263 -> Water_Temp.:97.81, PWM:33
01:22:53.786 -> Water_Temp.:97.81, PWM:33
01:22:55.341 -> Water_Temp.:97.81, PWM:33
01:22:56.140 -> Water_Temp.:97.92, PWM:33
01:22:57.696 -> Water_Temp.:97.92, PWM:33
01:22:58.457 -> Water_Temp.:97.92, PWM:33
01:22:59.253 -> Water_Temp.:97.92, PWM:33
01:23:03.084 -> Water_Temp.:97.92, PWM:33
01:23:03.878 -> Water_Temp.:97.92, PWM:33
01:23:05.430 -> Water_Temp.:97.92, PWM:33
01:23:06.223 -> Water_Temp.:97.92, PWM:33
01:23:06.981 -> Water_Temp.:97.92, PWM:33
01:23:09.294 -> Water_Temp.:97.92, PWM:33
01:23:10.083 -> Water_Temp.:97.92, PWM:33
01:23:11.633 -> Water_Temp.:97.92, PWM:33
01:23:12.395 -> Water_Temp.:97.92, PWM:33
01:23:13.190 -> Water_Temp.:97.92, PWM:33
01:23:13.951 -> Water_Temp.:97.92, PWM:33
01:23:16.263 -> Water_Temp.:98.04, PWM:33
01:23:17.056 -> Water_Temp.:98.04, PWM:33
01:23:17.817 -> Water_Temp.:98.04, PWM:33
01:23:18.610 -> Water_Temp.:98.04, PWM:33
01:23:19.403 -> Water_Temp.:98.04, PWM:33
01:23:20.166 -> Water_Temp.:98.04, PWM:33
01:23:20.960 -> Water_Temp.:98.04, PWM:33
01:23:23.280 -> Water_Temp.:98.04, PWM:33
01:23:24.041 -> Water_Temp.:98.04, PWM:33
01:23:27.150 -> Water_Temp.:98.04, PWM:33
01:23:27.943 -> Water_Temp.:98.04, PWM:33
01:23:28.700 -> Water_Temp.:98.04, PWM:33
01:23:29.495 -> Water_Temp.:98.04, PWM:33
01:23:30.286 -> Water_Temp.:98.04, PWM:33
01:23:31.045 -> Water_Temp.:98.04, PWM:33
01:23:31.838 -> Water_Temp.:98.04, PWM:33
01:23:32.595 -> Water_Temp.:98.04, PWM:33
01:23:33.388 -> Water_Temp.:98.04, PWM:33
01:23:34.944 -> Water_Temp.:98.04, PWM:33
01:23:35.740 -> Water_Temp.:98.04, PWM:33
01:23:36.500 -> Water_Temp.:98.15, PWM:33
01:23:38.050 -> Water_Temp.:98.04, PWM:33
01:23:38.844 -> Water_Temp.:98.15, PWM:33
01:23:39.605 -> Water_Temp.:98.15, PWM:33
01:23:41.161 -> Water_Temp.:98.15, PWM:33
01:23:41.950 -> Water_Temp.:98.15, PWM:33
01:23:42.709 -> Water_Temp.:98.15, PWM:33
01:23:45.058 -> Water_Temp.:98.15, PWM:33
01:23:45.817 -> Water_Temp.:98.15, PWM:33
01:23:46.610 -> Water_Temp.:98.15, PWM:33
01:23:48.162 -> Water_Temp.:98.15, PWM:33
01:23:49.682 -> Water_Temp.:98.15, PWM:33
01:23:50.474 -> Water_Temp.:98.15, PWM:33
01:23:51.267 -> Water_Temp.:98.15, PWM:33
01:23:52.819 -> Water_Temp.:98.15, PWM:33
01:23:53.580 -> Water_Temp.:98.15, PWM:33
01:23:54.376 -> Water_Temp.:98.15, PWM:33
01:23:55.169 -> Water_Temp.:98.15, PWM:33
01:23:55.927 -> Water_Temp.:98.15, PWM:33
01:23:56.718 -> Water_Temp.:98.15, PWM:33
01:23:59.061 -> Water_Temp.:98.15, PWM:33
01:23:59.819 -> Water_Temp.:98.26, PWM:33
01:24:02.164 -> Water_Temp.:98.15, PWM:33
01:24:03.720 -> Water_Temp.:98.26, PWM:33
01:24:04.477 -> Water_Temp.:98.26, PWM:33
01:24:06.067 -> Water_Temp.:98.26, PWM:33
01:24:07.616 -> Water_Temp.:98.26, PWM:33
01:24:11.429 -> Water_Temp.:98.26, PWM:33
01:24:12.229 -> Water_Temp.:98.26, PWM:33
01:24:12.987 -> Water_Temp.:98.26, PWM:33
01:24:13.781 -> Water_Temp.:98.26, PWM:33
01:24:15.302 -> Water_Temp.:98.26, PWM:33
01:24:19.922 -> Water_Temp.:98.26, PWM:33
01:24:22.991 -> Water_Temp.:98.26, PWM:33
01:24:23.753 -> Water_Temp.:98.26, PWM:33
01:24:24.543 -> Water_Temp.:98.26, PWM:33
01:24:25.332 -> Water_Temp.:98.26, PWM:33
01:24:26.086 -> Water_Temp.:98.26, PWM:33
01:24:26.875 -> Water_Temp.:98.26, PWM:33
01:24:27.665 -> Water_Temp.:98.26, PWM:33
01:24:28.424 -> Water_Temp.:98.26, PWM:33
01:24:30.773 -> Water_Temp.:98.26, PWM:33
01:24:31.564 -> Water_Temp.:98.26, PWM:33
01:24:33.114 -> Water_Temp.:98.26, PWM:33
01:24:33.907 -> Water_Temp.:98.26, PWM:33
01:24:34.659 -> Water_Temp.:98.26, PWM:33
01:24:35.455 -> Water_Temp.:98.26, PWM:33
01:24:37.756 -> Water_Temp.:98.26, PWM:33
01:24:38.521 -> Water_Temp.:98.26, PWM:33
01:24:40.079 -> Water_Temp.:98.26, PWM:33
01:24:40.873 -> Water_Temp.:98.26, PWM:33
01:24:42.427 -> Water_Temp.:98.26, PWM:33
01:24:43.188 -> Water_Temp.:98.26, PWM:33
01:24:43.977 -> Water_Temp.:98.26, PWM:33
01:24:44.764 -> Water_Temp.:98.26, PWM:33
01:24:45.523 -> Water_Temp.:98.26, PWM:33
01:24:48.602 -> Water_Temp.:98.26, PWM:33
01:24:49.398 -> Water_Temp.:98.26, PWM:33
01:24:50.161 -> Water_Temp.:98.26, PWM:33
01:24:50.954 -> Water_Temp.:98.26, PWM:33
01:24:51.750 -> Water_Temp.:98.26, PWM:33
01:24:52.513 -> Water_Temp.:98.26, PWM:33
01:24:53.312 -> Water_Temp.:98.26, PWM:33
01:24:54.071 -> Water_Temp.:98.26, PWM:33
01:24:54.865 -> Water_Temp.:98.26, PWM:33
01:24:55.623 -> Water_Temp.:98.26, PWM:33
01:24:57.178 -> Water_Temp.:98.26, PWM:33
01:24:57.972 -> Water_Temp.:98.26, PWM:33
01:24:58.761 -> Water_Temp.:98.26, PWM:33
01:24:59.518 -> Water_Temp.:98.26, PWM:33
01:25:01.067 -> Water_Temp.:98.26, PWM:33
01:25:02.614 -> Water_Temp.:98.26, PWM:33
01:25:03.405 -> Water_Temp.:98.26, PWM:33
01:25:04.169 -> Water_Temp.:98.26, PWM:33
01:25:05.726 -> Water_Temp.:98.26, PWM:33
01:25:07.282 -> Water_Temp.:98.26, PWM:33
01:25:08.079 -> Water_Temp.:98.26, PWM:33
01:25:08.835 -> Water_Temp.:98.26, PWM:33
01:25:10.422 -> Water_Temp.:98.26, PWM:33
01:25:12.739 -> Water_Temp.:98.26, PWM:33
01:25:13.537 -> Water_Temp.:98.26, PWM:33
01:25:14.299 -> Water_Temp.:98.26, PWM:33
01:25:15.097 -> Water_Temp.:98.26, PWM:33
01:25:15.857 -> Water_Temp.:98.26, PWM:33
01:25:16.650 -> Water_Temp.:98.26, PWM:33
01:25:22.002 -> Water_Temp.:98.26, PWM:33
01:25:25.073 -> Water_Temp.:98.26, PWM:33
01:25:29.705 -> Water_Temp.:98.37, PWM:33
01:25:32.816 -> Water_Temp.:98.37, PWM:33
01:25:33.609 -> Water_Temp.:98.37, PWM:33
01:25:35.167 -> Water_Temp.:98.37, PWM:33
01:25:35.929 -> Water_Temp.:98.37, PWM:33
01:25:37.517 -> Water_Temp.:98.37, PWM:33
01:25:39.074 -> Water_Temp.:98.37, PWM:33
01:25:40.627 -> Water_Temp.:98.37, PWM:33
01:25:42.185 -> Water_Temp.:98.49, PWM:33
01:25:43.709 -> Water_Temp.:98.49, PWM:33
01:25:44.506 -> Water_Temp.:98.49, PWM:33
01:25:46.027 -> Water_Temp.:98.49, PWM:33
01:25:46.818 -> Water_Temp.:98.49, PWM:33
01:25:47.580 -> Water_Temp.:98.49, PWM:33
01:25:48.378 -> Water_Temp.:98.49, PWM:33
01:25:49.140 -> Water_Temp.:98.49, PWM:33
01:25:49.936 -> Water_Temp.:98.49, PWM:33
01:25:51.492 -> Water_Temp.:98.49, PWM:33
01:25:52.253 -> Water_Temp.:98.49, PWM:33
01:25:53.799 -> Water_Temp.:98.37, PWM:33
01:25:54.594 -> Water_Temp.:98.37, PWM:33
01:25:56.904 -> Water_Temp.:98.37, PWM:33
01:25:57.696 -> Water_Temp.:98.37, PWM:33
01:26:00.010 -> Water_Temp.:98.37, PWM:33
01:26:00.770 -> Water_Temp.:98.37, PWM:33
01:26:03.080 -> Water_Temp.:98.37, PWM:33
01:26:03.874 -> Water_Temp.:98.37, PWM:33
01:26:05.431 -> Water_Temp.:98.37, PWM:33
01:26:06.954 -> Water_Temp.:98.49, PWM:33
01:26:08.511 -> Water_Temp.:98.49, PWM:33
01:26:09.306 -> Water_Temp.:98.49, PWM:33
01:26:10.064 -> Water_Temp.:98.49, PWM:33
01:26:10.858 -> Water_Temp.:98.49, PWM:33
01:26:11.619 -> Water_Temp.:98.49, PWM:33
01:26:12.408 -> Water_Temp.:98.49, PWM:33
01:26:13.965 -> Water_Temp.:98.49, PWM:33
01:26:14.759 -> Water_Temp.:98.49, PWM:33
01:26:17.037 -> Water_Temp.:98.49, PWM:33
01:26:17.831 -> Water_Temp.:98.49, PWM:33
01:26:18.593 -> Water_Temp.:98.49, PWM:33
01:26:19.389 -> Water_Temp.:98.49, PWM:33
01:26:20.915 -> Water_Temp.:98.49, PWM:33
01:26:22.472 -> Water_Temp.:98.49, PWM:33
01:26:24.036 -> Water_Temp.:98.49, PWM:33
01:26:24.830 -> Water_Temp.:98.49, PWM:33
01:26:25.587 -> Water_Temp.:98.49, PWM:33
01:26:28.660 -> Water_Temp.:98.49, PWM:33
01:26:29.456 -> Water_Temp.:98.49, PWM:33
01:26:30.214 -> Water_Temp.:98.49, PWM:33
01:26:31.011 -> Water_Temp.:98.49, PWM:33
01:26:32.530 -> Water_Temp.:98.49, PWM:33
01:26:33.331 -> Water_Temp.:98.49, PWM:33
01:26:34.089 -> Water_Temp.:98.60, PWM:33
01:26:34.881 -> Water_Temp.:98.60, PWM:33
01:26:35.673 -> Water_Temp.:98.60, PWM:33
01:26:36.432 -> Water_Temp.:98.60, PWM:33
01:26:37.230 -> Water_Temp.:98.60, PWM:33
01:26:38.783 -> Water_Temp.:98.60, PWM:33
01:26:39.542 -> Water_Temp.:98.60, PWM:33
01:26:40.338 -> Water_Temp.:98.49, PWM:33
01:26:41.135 -> Water_Temp.:98.60, PWM:33
01:26:41.898 -> Water_Temp.:98.60, PWM:33
01:26:43.453 -> Water_Temp.:98.60, PWM:33
01:26:44.215 -> Water_Temp.:98.60, PWM:33
01:26:45.008 -> Water_Temp.:98.60, PWM:33
01:26:46.561 -> Water_Temp.:98.60, PWM:33
01:26:47.317 -> Water_Temp.:98.60, PWM:33
01:26:48.113 -> Water_Temp.:98.60, PWM:33
01:26:48.878 -> Water_Temp.:98.60, PWM:33
01:26:50.448 -> Water_Temp.:98.60, PWM:33
01:26:51.973 -> Water_Temp.:98.60, PWM:33
01:26:52.766 -> Water_Temp.:98.60, PWM:33
01:26:54.325 -> Water_Temp.:98.60, PWM:33
01:26:55.877 -> Water_Temp.:98.60, PWM:33
01:26:57.427 -> Water_Temp.:98.60, PWM:33
01:26:58.216 -> Water_Temp.:98.60, PWM:33
01:26:58.973 -> Water_Temp.:98.60, PWM:33
01:26:59.762 -> Water_Temp.:98.60, PWM:33
01:27:00.556 -> Water_Temp.:98.60, PWM:33
01:27:01.318 -> Water_Temp.:98.60, PWM:33
01:27:02.869 -> Water_Temp.:98.60, PWM:33
01:27:03.669 -> Water_Temp.:98.60, PWM:33
01:27:06.735 -> Water_Temp.:98.60, PWM:33
01:27:07.529 -> Water_Temp.:98.60, PWM:33
01:27:08.290 -> Water_Temp.:98.60, PWM:33
01:27:09.875 -> Water_Temp.:98.60, PWM:33
01:27:10.638 -> Water_Temp.:98.60, PWM:33
01:27:11.435 -> Water_Temp.:98.60, PWM:33
01:27:12.989 -> Water_Temp.:98.60, PWM:33
01:27:13.751 -> Water_Temp.:98.60, PWM:33
01:27:14.544 -> Water_Temp.:98.60, PWM:33
01:27:15.339 -> Water_Temp.:98.60, PWM:33
01:27:16.099 -> Water_Temp.:98.60, PWM:33
01:27:16.892 -> Water_Temp.:98.60, PWM:33
01:27:17.652 -> Water_Temp.:98.60, PWM:33
01:27:18.444 -> Water_Temp.:98.60, PWM:33
01:27:19.239 -> Water_Temp.:98.60, PWM:33
01:27:20.764 -> Water_Temp.:98.60, PWM:33
01:27:21.523 -> Water_Temp.:98.60, PWM:33
01:27:22.318 -> Water_Temp.:98.60, PWM:33
01:27:23.112 -> Water_Temp.:98.60, PWM:33
01:27:23.870 -> Water_Temp.:98.60, PWM:33
01:27:24.664 -> Water_Temp.:98.60, PWM:33
01:27:25.424 -> Water_Temp.:98.60, PWM:33
01:27:26.216 -> Water_Temp.:98.60, PWM:33
01:27:27.766 -> Water_Temp.:98.60, PWM:33
01:27:28.556 -> Water_Temp.:98.60, PWM:33
01:27:29.348 -> Water_Temp.:98.60, PWM:33
01:27:30.107 -> Water_Temp.:98.60, PWM:33
01:27:33.234 -> Water_Temp.:98.60, PWM:33
01:27:34.750 -> Water_Temp.:98.71, PWM:33
01:27:36.307 -> Water_Temp.:98.71, PWM:33
01:27:37.101 -> Water_Temp.:98.71, PWM:33
01:27:37.857 -> Water_Temp.:98.71, PWM:33
01:27:38.654 -> Water_Temp.:98.71, PWM:33
01:27:39.445 -> Water_Temp.:98.71, PWM:33
01:27:40.206 -> Water_Temp.:98.71, PWM:33
01:27:41.004 -> Water_Temp.:98.71, PWM:33
01:27:41.762 -> Water_Temp.:98.71, PWM:33
01:27:42.554 -> Water_Temp.:98.71, PWM:33
01:27:43.348 -> Water_Temp.:98.71, PWM:33
01:27:44.898 -> Water_Temp.:98.71, PWM:33
01:27:45.657 -> Water_Temp.:98.71, PWM:33
01:27:47.247 -> Water_Temp.:98.71, PWM:33
01:27:48.770 -> Water_Temp.:98.71, PWM:33
01:27:51.084 -> Water_Temp.:98.71, PWM:33
01:27:51.842 -> Water_Temp.:98.71, PWM:33
01:27:52.639 -> Water_Temp.:98.71, PWM:33
01:27:53.433 -> Water_Temp.:98.71, PWM:33
01:27:54.195 -> Water_Temp.:98.71, PWM:33
01:27:55.747 -> Water_Temp.:98.71, PWM:33
01:27:56.543 -> Water_Temp.:98.71, PWM:33
01:27:57.334 -> Water_Temp.:98.71, PWM:33
01:27:58.096 -> Water_Temp.:98.71, PWM:33
01:27:58.892 -> Water_Temp.:98.82, PWM:33
01:28:00.445 -> Water_Temp.:98.82, PWM:33
01:28:01.206 -> Water_Temp.:98.82, PWM:33
01:28:01.998 -> Water_Temp.:98.82, PWM:33
01:28:02.792 -> Water_Temp.:98.82, PWM:33
01:28:05.071 -> Water_Temp.:98.82, PWM:33
01:28:07.390 -> Water_Temp.:98.82, PWM:33
01:28:08.181 -> Water_Temp.:98.82, PWM:33
01:28:10.495 -> Water_Temp.:98.82, PWM:33
01:28:12.841 -> Water_Temp.:98.94, PWM:33
01:28:13.608 -> Water_Temp.:98.94, PWM:33
01:28:14.403 -> Water_Temp.:98.94, PWM:33
01:28:15.166 -> Water_Temp.:98.94, PWM:33
01:28:15.956 -> Water_Temp.:98.94, PWM:33
01:28:17.507 -> Water_Temp.:98.94, PWM:33
01:28:18.304 -> Water_Temp.:98.94, PWM:33
01:28:19.064 -> Water_Temp.:98.94, PWM:33
01:28:19.855 -> Water_Temp.:98.94, PWM:33
01:28:20.615 -> Water_Temp.:98.94, PWM:33
01:28:21.409 -> Water_Temp.:98.94, PWM:33
01:28:23.758 -> Water_Temp.:98.94, PWM:33
01:28:25.313 -> Water_Temp.:98.94, PWM:33
01:28:26.872 -> Water_Temp.:98.94, PWM:33
01:28:28.389 -> Water_Temp.:98.94, PWM:33
01:28:29.183 -> Water_Temp.:98.94, PWM:33
01:28:30.704 -> Water_Temp.:98.94, PWM:33
01:28:31.500 -> Water_Temp.:98.94, PWM:33
01:28:32.259 -> Water_Temp.:98.94, PWM:33
01:28:33.055 -> Water_Temp.:98.94, PWM:33
01:28:34.612 -> Water_Temp.:98.94, PWM:33
01:28:36.930 -> Water_Temp.:100.62, PWM:33
01:28:37.728 -> Water_Temp.:102.65, PWM:33
01:28:39.257 -> Water_Temp.:99.84, PWM:33
01:28:40.814 -> Water_Temp.:99.50, PWM:33
01:28:41.605 -> Water_Temp.:99.50, PWM:33
01:28:43.129 -> Water_Temp.:99.27, PWM:33
01:28:44.689 -> Water_Temp.:99.27, PWM:33
01:28:47.765 -> Water_Temp.:99.16, PWM:33
01:28:50.085 -> Water_Temp.:99.05, PWM:33
01:28:50.878 -> Water_Temp.:99.05, PWM:33
01:28:51.671 -> Water_Temp.:99.05, PWM:33
01:28:52.430 -> Water_Temp.:99.05, PWM:33
01:28:53.224 -> Water_Temp.:99.05, PWM:33
01:28:53.983 -> Water_Temp.:99.05, PWM:33
01:28:54.774 -> Water_Temp.:99.05, PWM:33
01:28:56.324 -> Water_Temp.:99.05, PWM:33
01:28:57.117 -> Water_Temp.:99.05, PWM:33
01:28:58.673 -> Water_Temp.:99.05, PWM:33
01:29:01.769 -> Water_Temp.:99.05, PWM:33
01:29:02.557 -> Water_Temp.:99.05, PWM:33
01:29:04.110 -> Water_Temp.:99.05, PWM:33
01:29:04.869 -> Water_Temp.:99.05, PWM:33
01:29:05.662 -> Water_Temp.:99.16, PWM:33
01:29:06.453 -> Water_Temp.:99.16, PWM:33
01:29:07.210 -> Water_Temp.:99.16, PWM:33
01:29:08.002 -> Water_Temp.:99.16, PWM:33
01:29:08.796 -> Water_Temp.:99.16, PWM:33
01:29:10.315 -> Water_Temp.:99.16, PWM:33
01:29:12.656 -> Water_Temp.:99.16, PWM:33
01:29:13.414 -> Water_Temp.:99.16, PWM:33
01:29:14.972 -> Water_Temp.:99.16, PWM:33
01:29:16.526 -> Water_Temp.:99.16, PWM:33
01:29:17.287 -> Water_Temp.:99.16, PWM:33
01:29:18.873 -> Water_Temp.:99.16, PWM:33
01:29:19.628 -> Water_Temp.:99.16, PWM:33
01:29:20.424 -> Water_Temp.:99.16, PWM:33
01:29:21.187 -> Water_Temp.:99.16, PWM:33
01:29:22.769 -> Water_Temp.:99.16, PWM:33
01:29:23.530 -> Water_Temp.:99.16, PWM:33
01:29:25.088 -> Water_Temp.:99.16, PWM:33
01:29:25.876 -> Water_Temp.:99.16, PWM:33
01:29:26.630 -> Water_Temp.:99.16, PWM:33
01:29:27.422 -> Water_Temp.:99.16, PWM:33
01:29:28.216 -> Water_Temp.:99.16, PWM:33
01:29:29.738 -> Water_Temp.:99.16, PWM:33
01:29:31.297 -> Water_Temp.:99.16, PWM:33
01:29:32.056 -> Water_Temp.:99.16, PWM:33
01:29:32.848 -> Water_Temp.:99.16, PWM:33
01:29:33.642 -> Water_Temp.:99.16, PWM:33
01:29:34.399 -> Water_Temp.:99.16, PWM:33
01:29:35.951 -> Water_Temp.:99.16, PWM:33
01:29:37.500 -> Water_Temp.:99.16, PWM:33
01:29:39.814 -> Water_Temp.:99.16, PWM:33
01:29:40.576 -> Water_Temp.:99.16, PWM:33
01:29:41.371 -> Water_Temp.:99.16, PWM:33
01:29:42.129 -> Water_Temp.:99.16, PWM:33
01:29:44.480 -> Water_Temp.:99.16, PWM:33
01:29:45.274 -> Water_Temp.:99.16, PWM:33
01:29:46.827 -> Water_Temp.:99.16, PWM:33
01:29:47.584 -> Water_Temp.:99.16, PWM:33
01:29:49.923 -> Water_Temp.:99.16, PWM:33
01:29:50.684 -> Water_Temp.:99.16, PWM:33
01:29:51.477 -> Water_Temp.:99.16, PWM:33
01:29:53.029 -> Water_Temp.:99.16, PWM:33
01:29:55.346 -> Water_Temp.:99.16, PWM:33
01:29:56.106 -> Water_Temp.:99.16, PWM:33
01:29:56.898 -> Water_Temp.:99.16, PWM:33
01:29:57.691 -> Water_Temp.:99.16, PWM:33
01:29:58.446 -> Water_Temp.:99.16, PWM:33
01:29:59.238 -> Water_Temp.:99.16, PWM:33
01:30:01.545 -> Water_Temp.:99.16, PWM:33
01:30:02.303 -> Water_Temp.:99.16, PWM:33
01:30:03.095 -> Water_Temp.:99.16, PWM:33
01:30:04.615 -> Water_Temp.:99.16, PWM:33
01:30:05.408 -> Water_Temp.:99.16, PWM:33
01:30:06.931 -> Water_Temp.:99.16, PWM:33
01:30:08.488 -> Water_Temp.:99.16, PWM:33
01:30:09.282 -> Water_Temp.:99.27, PWM:33
01:30:12.382 -> Water_Temp.:99.16, PWM:33
01:30:13.897 -> Water_Temp.:99.16, PWM:33
01:30:14.689 -> Water_Temp.:99.27, PWM:33
01:30:16.241 -> Water_Temp.:99.16, PWM:33
01:30:17.035 -> Water_Temp.:99.16, PWM:33
01:30:17.794 -> Water_Temp.:99.27, PWM:33
01:30:18.588 -> Water_Temp.:99.27, PWM:33
01:30:19.349 -> Water_Temp.:99.27, PWM:33
01:30:20.142 -> Water_Temp.:100.85, PWM:33
01:30:20.935 -> Water_Temp.:102.99, PWM:33
01:30:21.693 -> Water_Temp.:100.51, PWM:33
01:30:22.486 -> Water_Temp.:100.17, PWM:33
01:30:23.244 -> Water_Temp.:99.95, PWM:33
01:30:24.802 -> Water_Temp.:99.72, PWM:33
01:30:27.115 -> Water_Temp.:99.61, PWM:33
01:30:27.879 -> Water_Temp.:99.61, PWM:33
01:30:28.676 -> Water_Temp.:99.61, PWM:33
01:30:30.230 -> Water_Temp.:99.50, PWM:33
01:30:30.992 -> Water_Temp.:99.50, PWM:33
01:30:31.781 -> Water_Temp.:99.50, PWM:33
01:30:32.540 -> Water_Temp.:99.50, PWM:33
01:30:34.098 -> Water_Temp.:99.50, PWM:33
01:30:34.888 -> Water_Temp.:99.50, PWM:33
01:30:37.176 -> Water_Temp.:99.50, PWM:33
01:30:37.968 -> Water_Temp.:99.50, PWM:33
01:30:41.051 -> Water_Temp.:99.50, PWM:33
01:30:41.843 -> Water_Temp.:99.39, PWM:33
01:30:42.637 -> Water_Temp.:99.39, PWM:33
01:30:43.399 -> Water_Temp.:99.50, PWM:33
01:30:44.160 -> Water_Temp.:99.39, PWM:33
01:30:44.958 -> Water_Temp.:99.39, PWM:33
01:30:48.762 -> Water_Temp.:99.39, PWM:33
01:30:50.348 -> Water_Temp.:99.39, PWM:33
01:30:51.109 -> Water_Temp.:99.39, PWM:33
01:30:51.899 -> Water_Temp.:99.39, PWM:33
01:30:53.425 -> Water_Temp.:99.39, PWM:33
01:30:55.011 -> Water_Temp.:99.39, PWM:33
01:30:55.772 -> Water_Temp.:99.39, PWM:33
01:30:57.324 -> Water_Temp.:99.39, PWM:33
01:30:58.120 -> Water_Temp.:99.39, PWM:33
01:30:58.914 -> Water_Temp.:99.39, PWM:33
01:31:00.473 -> Water_Temp.:99.39, PWM:33
01:31:01.230 -> Water_Temp.:99.39, PWM:33
01:31:02.026 -> Water_Temp.:99.39, PWM:33
01:31:02.786 -> Water_Temp.:99.39, PWM:33
01:31:03.584 -> Water_Temp.:99.39, PWM:33
01:31:04.342 -> Water_Temp.:99.39, PWM:33
01:31:05.897 -> Water_Temp.:99.39, PWM:33
01:31:09.787 -> Water_Temp.:99.39, PWM:33
01:31:12.095 -> Water_Temp.:99.39, PWM:33
01:31:12.888 -> Water_Temp.:99.39, PWM:33
01:31:13.653 -> Water_Temp.:99.39, PWM:33
01:31:15.962 -> Water_Temp.:99.39, PWM:33
01:31:16.753 -> Water_Temp.:99.39, PWM:33
01:31:17.547 -> Water_Temp.:99.39, PWM:33
01:31:18.308 -> Water_Temp.:99.39, PWM:33
01:31:19.105 -> Water_Temp.:99.39, PWM:33
01:31:19.865 -> Water_Temp.:99.39, PWM:33
01:31:20.663 -> Water_Temp.:99.39, PWM:33
01:31:22.213 -> Water_Temp.:99.39, PWM:33
01:31:24.526 -> Water_Temp.:99.39, PWM:33
01:31:25.320 -> Water_Temp.:99.39, PWM:33
01:31:26.114 -> Water_Temp.:99.39, PWM:33
01:31:26.871 -> Water_Temp.:99.39, PWM:33
01:31:28.418 -> Water_Temp.:99.39, PWM:33
01:31:29.212 -> Water_Temp.:99.39, PWM:33
01:31:30.765 -> Water_Temp.:99.50, PWM:33
01:31:31.526 -> Water_Temp.:99.39, PWM:33
01:31:33.079 -> Water_Temp.:99.39, PWM:33
01:31:33.871 -> Water_Temp.:99.39, PWM:33
01:31:34.663 -> Water_Temp.:99.50, PWM:33
01:31:35.422 -> Water_Temp.:99.39, PWM:33
01:31:36.970 -> Water_Temp.:99.39, PWM:33
01:31:37.759 -> Water_Temp.:99.39, PWM:33
01:31:38.519 -> Water_Temp.:99.39, PWM:33
01:31:39.314 -> Water_Temp.:99.39, PWM:33
01:31:40.075 -> Water_Temp.:99.39, PWM:33
01:31:41.660 -> Water_Temp.:99.39, PWM:33
01:31:43.213 -> Water_Temp.:99.39, PWM:33
01:31:43.973 -> Water_Temp.:99.39, PWM:33
01:31:44.762 -> Water_Temp.:99.39, PWM:33
01:31:46.312 -> Water_Temp.:99.39, PWM:33
01:31:47.109 -> Water_Temp.:99.39, PWM:33
01:31:47.869 -> Water_Temp.:99.39, PWM:33
01:31:48.664 -> Water_Temp.:99.39, PWM:33
01:31:50.217 -> Water_Temp.:99.39, PWM:33
01:31:50.980 -> Water_Temp.:99.39, PWM:33
01:31:51.775 -> Water_Temp.:99.39, PWM:33
01:31:52.570 -> Water_Temp.:99.39, PWM:33
01:31:53.333 -> Water_Temp.:99.39, PWM:33
01:31:54.128 -> Water_Temp.:99.39, PWM:33
01:31:54.888 -> Water_Temp.:99.39, PWM:33
01:31:55.682 -> Water_Temp.:99.39, PWM:33
01:31:56.444 -> Water_Temp.:99.39, PWM:33
01:31:58.801 -> Water_Temp.:99.50, PWM:33
01:31:59.562 -> Water_Temp.:99.39, PWM:33
01:32:01.119 -> Water_Temp.:99.50, PWM:33
01:32:01.877 -> Water_Temp.:99.50, PWM:33
01:32:03.468 -> Water_Temp.:99.50, PWM:33
01:32:04.227 -> Water_Temp.:99.50, PWM:33
01:32:07.307 -> Water_Temp.:99.50, PWM:33
01:32:08.103 -> Water_Temp.:99.50, PWM:33
01:32:08.857 -> Water_Temp.:99.50, PWM:33
01:32:09.657 -> Water_Temp.:99.50, PWM:33
01:32:10.420 -> Water_Temp.:99.50, PWM:33
01:32:11.219 -> Water_Temp.:99.50, PWM:33
01:32:11.973 -> Water_Temp.:99.50, PWM:33
01:32:12.761 -> Water_Temp.:99.50, PWM:33
01:32:13.557 -> Water_Temp.:99.39, PWM:33
01:32:14.321 -> Water_Temp.:99.50, PWM:33
01:32:15.118 -> Water_Temp.:99.50, PWM:33
01:32:17.431 -> Water_Temp.:99.39, PWM:33
01:32:18.196 -> Water_Temp.:99.39, PWM:33
01:32:19.752 -> Water_Temp.:99.39, PWM:33
01:32:20.515 -> Water_Temp.:99.39, PWM:33
01:32:21.315 -> Water_Temp.:99.39, PWM:33
01:32:22.070 -> Water_Temp.:99.39, PWM:33
01:32:23.595 -> Water_Temp.:99.39, PWM:33
01:32:24.390 -> Water_Temp.:99.39, PWM:33
01:32:25.184 -> Water_Temp.:99.39, PWM:33
01:32:25.949 -> Water_Temp.:99.39, PWM:33
01:32:26.743 -> Water_Temp.:99.39, PWM:33
01:32:27.506 -> Water_Temp.:99.39, PWM:33
01:32:28.301 -> Water_Temp.:99.39, PWM:33
01:32:29.064 -> Water_Temp.:99.39, PWM:33
01:32:31.383 -> Water_Temp.:101.07, PWM:33
01:32:32.177 -> Water_Temp.:103.21, PWM:33
01:32:32.937 -> Water_Temp.:100.74, PWM:33
01:32:33.735 -> Water_Temp.:100.29, PWM:33
01:32:34.528 -> Water_Temp.:100.06, PWM:33
01:32:36.047 -> Water_Temp.:99.84, PWM:33
01:32:36.841 -> Water_Temp.:99.84, PWM:33
01:32:37.602 -> Water_Temp.:99.72, PWM:33
01:32:38.400 -> Water_Temp.:99.72, PWM:33
01:32:39.193 -> Water_Temp.:99.61, PWM:33
01:32:39.951 -> Water_Temp.:99.61, PWM:33
01:32:40.748 -> Water_Temp.:99.61, PWM:33
01:32:41.508 -> Water_Temp.:99.61, PWM:33
01:32:43.066 -> Water_Temp.:99.61, PWM:33
01:32:44.620 -> Water_Temp.:99.50, PWM:33
01:32:45.383 -> Water_Temp.:99.61, PWM:33
01:32:49.254 -> Water_Temp.:99.50, PWM:33
01:32:50.014 -> Water_Temp.:99.50, PWM:33
01:32:50.812 -> Water_Temp.:99.50, PWM:33
01:32:51.569 -> Water_Temp.:99.50, PWM:33
01:32:52.361 -> Water_Temp.:99.50, PWM:33
01:32:53.123 -> Water_Temp.:99.50, PWM:33
01:32:55.471 -> Water_Temp.:99.61, PWM:33
01:32:56.235 -> Water_Temp.:99.61, PWM:33
01:32:57.028 -> Water_Temp.:99.61, PWM:33
01:32:57.788 -> Water_Temp.:99.61, PWM:33
01:32:59.313 -> Water_Temp.:99.50, PWM:33
01:33:00.905 -> Water_Temp.:99.50, PWM:33
01:33:02.457 -> Water_Temp.:99.50, PWM:33
01:33:03.976 -> Water_Temp.:99.61, PWM:33
01:33:04.771 -> Water_Temp.:99.61, PWM:33
01:33:05.560 -> Water_Temp.:99.61, PWM:33
01:33:06.319 -> Water_Temp.:101.19, PWM:33
01:33:07.842 -> Water_Temp.:100.85, PWM:33
01:33:08.638 -> Water_Temp.:100.51, PWM:33
01:33:09.433 -> Water_Temp.:100.29, PWM:33
01:33:10.960 -> Water_Temp.:100.06, PWM:33
01:33:11.717 -> Water_Temp.:99.95, PWM:33
01:33:13.277 -> Water_Temp.:99.84, PWM:33
01:33:16.394 -> Water_Temp.:99.84, PWM:33
01:33:17.971 -> Water_Temp.:99.72, PWM:33
01:33:19.523 -> Water_Temp.:99.72, PWM:33
01:33:20.321 -> Water_Temp.:99.72, PWM:33
01:33:21.082 -> Water_Temp.:99.72, PWM:33
01:33:21.875 -> Water_Temp.:99.72, PWM:33
01:33:22.637 -> Water_Temp.:99.72, PWM:33
01:33:24.189 -> Water_Temp.:99.72, PWM:33
01:33:24.949 -> Water_Temp.:99.72, PWM:33
01:33:26.533 -> Water_Temp.:99.72, PWM:33
01:33:27.291 -> Water_Temp.:99.72, PWM:33
01:33:28.084 -> Water_Temp.:99.72, PWM:33
01:33:31.914 -> Water_Temp.:99.72, PWM:33
01:33:32.713 -> Water_Temp.:99.72, PWM:33
01:33:33.474 -> Water_Temp.:99.72, PWM:33
01:33:34.268 -> Water_Temp.:99.72, PWM:33
01:33:35.822 -> Water_Temp.:99.72, PWM:33
01:33:38.133 -> Water_Temp.:99.72, PWM:33
01:33:38.889 -> Water_Temp.:99.72, PWM:33
01:33:40.445 -> Water_Temp.:99.72, PWM:33
01:33:42.798 -> Water_Temp.:99.72, PWM:33
01:33:46.636 -> Water_Temp.:99.72, PWM:33
01:33:47.400 -> Water_Temp.:99.72, PWM:33
01:33:48.197 -> Water_Temp.:99.61, PWM:33
01:33:48.962 -> Water_Temp.:99.61, PWM:33
01:33:49.754 -> Water_Temp.:99.61, PWM:33
01:33:52.068 -> Water_Temp.:99.61, PWM:33
01:33:52.830 -> Water_Temp.:99.61, PWM:33
01:33:53.624 -> Water_Temp.:99.72, PWM:33
01:33:55.942 -> Water_Temp.:99.72, PWM:33
01:33:57.492 -> Water_Temp.:99.72, PWM:33
01:33:58.254 -> Water_Temp.:99.72, PWM:33
01:33:59.049 -> Water_Temp.:99.72, PWM:33
01:33:59.839 -> Water_Temp.:99.72, PWM:33
01:34:00.598 -> Water_Temp.:99.72, PWM:33
01:34:01.396 -> Water_Temp.:99.72, PWM:33
01:34:02.155 -> Water_Temp.:99.72, PWM:33
01:34:02.950 -> Water_Temp.:99.72, PWM:33
01:34:05.237 -> Water_Temp.:99.72, PWM:33
01:34:06.034 -> Water_Temp.:99.72, PWM:33
01:34:06.826 -> Water_Temp.:99.72, PWM:33
01:34:07.587 -> Water_Temp.:99.72, PWM:33
01:34:08.385 -> Water_Temp.:99.72, PWM:33
01:34:09.945 -> Water_Temp.:99.72, PWM:33
01:34:10.706 -> Water_Temp.:99.72, PWM:33
01:34:13.819 -> Water_Temp.:99.72, PWM:33
01:34:14.581 -> Water_Temp.:99.72, PWM:33
01:34:18.447 -> Water_Temp.:99.72, PWM:33
01:34:19.207 -> Water_Temp.:99.72, PWM:33
01:34:20.003 -> Water_Temp.:99.72, PWM:33
01:34:20.766 -> Water_Temp.:99.72, PWM:33
01:34:22.320 -> Water_Temp.:99.72, PWM:33
01:34:23.910 -> Water_Temp.:99.72, PWM:33
01:34:25.467 -> Water_Temp.:99.72, PWM:33
01:34:26.231 -> Water_Temp.:99.72, PWM:33
01:34:27.788 -> Water_Temp.:99.72, PWM:33
01:34:28.582 -> Water_Temp.:99.84, PWM:33
01:34:30.142 -> Water_Temp.:99.72, PWM:33
01:34:30.904 -> Water_Temp.:99.84, PWM:33
01:34:32.456 -> Water_Temp.:99.84, PWM:33
01:34:34.012 -> Water_Temp.:99.84, PWM:33
01:34:36.328 -> Water_Temp.:99.84, PWM:33
01:34:37.914 -> Water_Temp.:99.84, PWM:33
01:34:38.670 -> Water_Temp.:99.84, PWM:33
01:34:39.467 -> Water_Temp.:99.84, PWM:33
01:34:43.308 -> Water_Temp.:99.84, PWM:33
01:34:44.070 -> Water_Temp.:99.84, PWM:33
01:34:44.867 -> Water_Temp.:99.84, PWM:33
01:34:47.179 -> Water_Temp.:99.72, PWM:33
01:34:47.936 -> Water_Temp.:99.72, PWM:33
01:34:50.290 -> Water_Temp.:99.84, PWM:33
01:34:51.088 -> Water_Temp.:99.72, PWM:33
01:34:51.847 -> Water_Temp.:99.84, PWM:33
01:34:52.644 -> Water_Temp.:99.84, PWM:33
01:34:53.406 -> Water_Temp.:99.84, PWM:33
01:34:54.965 -> Water_Temp.:99.84, PWM:33
01:34:56.519 -> Water_Temp.:99.84, PWM:33
01:34:57.316 -> Water_Temp.:99.84, PWM:33
01:34:58.078 -> Water_Temp.:99.84, PWM:33
01:34:59.636 -> Water_Temp.:99.84, PWM:33
01:35:01.983 -> Water_Temp.:99.84, PWM:33
01:35:05.065 -> Water_Temp.:99.95, PWM:33
01:35:05.826 -> Water_Temp.:99.84, PWM:33
01:35:06.621 -> Water_Temp.:99.95, PWM:33
01:35:07.377 -> Water_Temp.:99.95, PWM:33
01:35:08.168 -> Water_Temp.:99.95, PWM:33
01:35:08.956 -> Water_Temp.:99.95, PWM:33
01:35:10.510 -> Water_Temp.:99.95, PWM:33
01:35:11.271 -> Water_Temp.:99.95, PWM:33
01:35:12.065 -> Water_Temp.:99.95, PWM:33
01:35:12.829 -> Water_Temp.:99.95, PWM:33
01:35:14.424 -> Water_Temp.:99.95, PWM:33
01:35:15.971 -> Water_Temp.:99.95, PWM:33
01:35:16.731 -> Water_Temp.:99.95, PWM:33
01:35:17.528 -> Water_Temp.:99.95, PWM:33
01:35:18.292 -> Water_Temp.:99.95, PWM:33
01:35:19.086 -> Water_Temp.:99.95, PWM:33
01:35:19.845 -> Water_Temp.:99.95, PWM:33
01:35:21.401 -> Water_Temp.:99.95, PWM:33
01:35:22.948 -> Water_Temp.:99.95, PWM:33
01:35:24.501 -> Water_Temp.:99.95, PWM:33
01:35:26.021 -> Water_Temp.:99.95, PWM:33
01:35:26.814 -> Water_Temp.:99.95, PWM:33
01:35:27.610 -> Water_Temp.:99.95, PWM:33
01:35:28.373 -> Water_Temp.:99.95, PWM:33
01:35:29.164 -> Water_Temp.:99.95, PWM:33
01:35:29.923 -> Water_Temp.:99.95, PWM:33
01:35:30.718 -> Water_Temp.:99.95, PWM:33
01:35:31.512 -> Water_Temp.:99.95, PWM:33
01:35:33.027 -> Water_Temp.:99.95, PWM:33
01:35:33.821 -> Water_Temp.:99.95, PWM:33
01:35:34.580 -> Water_Temp.:99.95, PWM:33
01:35:35.377 -> Water_Temp.:99.95, PWM:33
01:35:36.173 -> Water_Temp.:99.95, PWM:33
01:35:38.492 -> Water_Temp.:99.95, PWM:33
01:35:39.289 -> Water_Temp.:99.95, PWM:33
01:35:40.049 -> Water_Temp.:99.95, PWM:33
01:35:40.843 -> Water_Temp.:99.95, PWM:33
01:35:41.602 -> Water_Temp.:99.95, PWM:33
01:35:43.161 -> Water_Temp.:99.95, PWM:33
01:35:43.919 -> Water_Temp.:99.95, PWM:33
01:35:45.510 -> Water_Temp.:99.95, PWM:33
01:35:46.271 -> Water_Temp.:99.95, PWM:33
01:35:47.064 -> Water_Temp.:99.95, PWM:33
01:35:47.822 -> Water_Temp.:99.95, PWM:33
01:35:49.404 -> Water_Temp.:99.95, PWM:33
01:35:50.162 -> Water_Temp.:99.95, PWM:33
01:35:50.958 -> Water_Temp.:99.95, PWM:33
01:35:51.719 -> Water_Temp.:99.95, PWM:33
01:35:52.513 -> Water_Temp.:99.95, PWM:33
01:35:53.309 -> Water_Temp.:99.95, PWM:33
01:35:54.071 -> Water_Temp.:99.95, PWM:33
01:35:54.865 -> Water_Temp.:99.95, PWM:33
01:35:57.178 -> Water_Temp.:99.95, PWM:33
01:35:58.697 -> Water_Temp.:99.95, PWM:33
01:35:59.494 -> Water_Temp.:99.95, PWM:33
01:36:01.047 -> Water_Temp.:99.95, PWM:33
01:36:02.604 -> Water_Temp.:99.95, PWM:33
01:36:03.368 -> Water_Temp.:99.95, PWM:33
01:36:04.167 -> Water_Temp.:99.95, PWM:33
01:36:04.924 -> Water_Temp.:99.95, PWM:33
01:36:08.033 -> Water_Temp.:99.95, PWM:33
01:36:09.597 -> Water_Temp.:100.06, PWM:33
01:36:10.359 -> Water_Temp.:100.06, PWM:33
01:36:11.158 -> Water_Temp.:100.06, PWM:33
01:36:12.677 -> Water_Temp.:100.06, PWM:33
01:36:13.468 -> Water_Temp.:100.06, PWM:33
01:36:14.264 -> Water_Temp.:100.06, PWM:33
01:36:15.786 -> Water_Temp.:100.06, PWM:33
01:36:17.374 -> Water_Temp.:100.06, PWM:33
01:36:18.134 -> Water_Temp.:100.06, PWM:33
01:36:18.928 -> Water_Temp.:100.06, PWM:33
01:36:21.242 -> Water_Temp.:100.06, PWM:33
01:36:22.797 -> Water_Temp.:100.06, PWM:33
01:36:23.594 -> Water_Temp.:100.06, PWM:33
01:36:24.356 -> Water_Temp.:100.06, PWM:33
01:36:25.906 -> Water_Temp.:100.06, PWM:33
01:36:27.457 -> Water_Temp.:100.06, PWM:33
01:36:30.589 -> Water_Temp.:100.06, PWM:33
01:36:32.107 -> Water_Temp.:100.06, PWM:33
01:36:32.905 -> Water_Temp.:100.06, PWM:33
01:36:33.662 -> Water_Temp.:100.06, PWM:33
01:36:35.219 -> Water_Temp.:99.95, PWM:33
01:36:39.050 -> Water_Temp.:99.95, PWM:33
01:36:39.843 -> Water_Temp.:99.95, PWM:33
01:36:40.636 -> Water_Temp.:99.95, PWM:33
01:36:41.402 -> Water_Temp.:99.95, PWM:33
01:36:42.196 -> Water_Temp.:99.95, PWM:33
01:36:43.718 -> Water_Temp.:99.95, PWM:33
01:36:45.270 -> Water_Temp.:99.95, PWM:33
01:36:46.031 -> Water_Temp.:99.95, PWM:33
01:36:49.142 -> Water_Temp.:99.95, PWM:33
01:36:49.939 -> Water_Temp.:99.95, PWM:33
01:36:50.697 -> Water_Temp.:99.95, PWM:33
01:36:51.491 -> Water_Temp.:99.95, PWM:33
01:36:52.251 -> Water_Temp.:99.95, PWM:33
01:36:53.814 -> Water_Temp.:99.95, PWM:33
01:36:55.366 -> Water_Temp.:100.06, PWM:33
01:36:56.159 -> Water_Temp.:99.95, PWM:33
01:36:56.952 -> Water_Temp.:99.95, PWM:33
01:36:59.996 -> Water_Temp.:99.95, PWM:33
01:37:00.754 -> Water_Temp.:99.95, PWM:33
01:37:02.310 -> Water_Temp.:99.95, PWM:33
01:37:03.899 -> Water_Temp.:100.06, PWM:33
01:37:04.658 -> Water_Temp.:100.06, PWM:33
01:37:05.453 -> Water_Temp.:100.06, PWM:33
01:37:06.214 -> Water_Temp.:100.06, PWM:33
01:37:07.005 -> Water_Temp.:100.06, PWM:33
01:37:08.564 -> Water_Temp.:100.06, PWM:33
01:37:09.329 -> Water_Temp.:100.06, PWM:33
01:37:10.122 -> Water_Temp.:100.06, PWM:33
01:37:10.885 -> Water_Temp.:100.06, PWM:33
01:37:11.678 -> Water_Temp.:100.06, PWM:33
01:37:13.999 -> Water_Temp.:100.06, PWM:33
01:37:14.759 -> Water_Temp.:100.06, PWM:33
01:37:15.559 -> Water_Temp.:100.06, PWM:33
01:37:16.322 -> Water_Temp.:100.06, PWM:33
01:37:17.119 -> Water_Temp.:100.06, PWM:33
01:37:17.879 -> Water_Temp.:100.06, PWM:33
01:37:19.439 -> Water_Temp.:99.95, PWM:33
01:37:20.994 -> Water_Temp.:99.95, PWM:33
01:37:21.754 -> Water_Temp.:99.95, PWM:33
01:37:22.552 -> Water_Temp.:99.95, PWM:33
01:37:25.634 -> Water_Temp.:99.95, PWM:33
01:37:26.397 -> Water_Temp.:99.95, PWM:33
01:37:27.193 -> Water_Temp.:99.95, PWM:33
01:37:27.989 -> Water_Temp.:99.95, PWM:33
01:37:29.539 -> Water_Temp.:99.84, PWM:33
01:37:30.300 -> Water_Temp.:99.84, PWM:33
01:37:31.885 -> Water_Temp.:99.84, PWM:33
01:37:32.646 -> Water_Temp.:99.84, PWM:33
01:37:33.443 -> Water_Temp.:99.84, PWM:33
01:37:35.002 -> Water_Temp.:99.84, PWM:33
01:37:37.286 -> Water_Temp.:99.84, PWM:33
01:37:38.837 -> Water_Temp.:99.84, PWM:33
01:37:40.393 -> Water_Temp.:99.84, PWM:33
01:37:41.910 -> Water_Temp.:99.84, PWM:33
01:37:42.707 -> Water_Temp.:99.95, PWM:33
01:37:44.261 -> Water_Temp.:99.95, PWM:33
01:37:45.056 -> Water_Temp.:99.95, PWM:33
01:37:45.817 -> Water_Temp.:99.95, PWM:33
01:37:51.250 -> Water_Temp.:99.95, PWM:33
01:37:52.044 -> Water_Temp.:99.95, PWM:33
01:37:52.805 -> Water_Temp.:99.95, PWM:33
01:37:54.363 -> Water_Temp.:99.95, PWM:33
01:37:55.159 -> Water_Temp.:99.95, PWM:33
01:37:56.682 -> Water_Temp.:99.95, PWM:33
01:37:59.034 -> Water_Temp.:99.95, PWM:33
01:37:59.793 -> Water_Temp.:99.95, PWM:33
01:38:01.347 -> Water_Temp.:99.95, PWM:33
01:38:02.141 -> Water_Temp.:99.95, PWM:33
01:38:02.899 -> Water_Temp.:99.95, PWM:33
01:38:05.215 -> Water_Temp.:99.95, PWM:33
01:38:08.323 -> Water_Temp.:99.95, PWM:33
01:38:09.116 -> Water_Temp.:99.95, PWM:33
01:38:10.671 -> Water_Temp.:99.95, PWM:33
01:38:12.990 -> Water_Temp.:100.06, PWM:33
01:38:13.789 -> Water_Temp.:100.06, PWM:33
01:38:14.551 -> Water_Temp.:100.06, PWM:33
01:38:16.073 -> Water_Temp.:100.06, PWM:33
01:38:17.630 -> Water_Temp.:100.06, PWM:33
01:38:18.426 -> Water_Temp.:100.06, PWM:33
01:38:19.986 -> Water_Temp.:100.06, PWM:33
01:38:20.780 -> Water_Temp.:100.06, PWM:33
01:38:21.540 -> Water_Temp.:100.17, PWM:33
01:38:22.334 -> Water_Temp.:100.17, PWM:33
01:38:23.096 -> Water_Temp.:100.06, PWM:33
01:38:23.889 -> Water_Temp.:100.06, PWM:33
01:38:25.447 -> Water_Temp.:100.17, PWM:33
01:38:27.005 -> Water_Temp.:100.17, PWM:33
01:38:27.765 -> Water_Temp.:100.17, PWM:33
01:38:28.530 -> Water_Temp.:100.17, PWM:33
01:38:30.083 -> Water_Temp.:100.17, PWM:33
01:38:30.877 -> Water_Temp.:100.17, PWM:33
01:38:33.195 -> Water_Temp.:100.17, PWM:33
01:38:33.989 -> Water_Temp.:100.17, PWM:33
01:38:36.276 -> Water_Temp.:100.17, PWM:33
01:38:37.073 -> Water_Temp.:100.17, PWM:33
01:38:38.631 -> Water_Temp.:100.17, PWM:33
01:38:39.395 -> Water_Temp.:100.17, PWM:33
01:38:40.194 -> Water_Temp.:100.17, PWM:33
01:38:40.989 -> Water_Temp.:100.06, PWM:33
01:38:41.753 -> Water_Temp.:100.06, PWM:33
01:38:43.308 -> Water_Temp.:100.06, PWM:33
01:38:44.102 -> Water_Temp.:100.06, PWM:33
01:38:44.865 -> Water_Temp.:100.06, PWM:33
01:38:45.661 -> Water_Temp.:100.06, PWM:33
01:38:47.217 -> Water_Temp.:100.06, PWM:33
01:38:49.533 -> Water_Temp.:100.06, PWM:33
01:38:50.328 -> Water_Temp.:100.06, PWM:33
01:38:51.884 -> Water_Temp.:100.06, PWM:33
01:38:52.682 -> Water_Temp.:100.06, PWM:33
01:38:53.441 -> Water_Temp.:100.06, PWM:33
01:38:54.235 -> Water_Temp.:100.06, PWM:33
01:38:54.991 -> Water_Temp.:100.06, PWM:33
01:38:56.546 -> Water_Temp.:100.06, PWM:33
01:38:57.344 -> Water_Temp.:100.06, PWM:33
01:38:58.107 -> Water_Temp.:100.06, PWM:33
01:38:59.665 -> Water_Temp.:100.06, PWM:33
01:39:00.460 -> Water_Temp.:100.06, PWM:33
01:39:02.747 -> Water_Temp.:100.06, PWM:33
01:39:03.546 -> Water_Temp.:100.06, PWM:33
01:39:05.834 -> Water_Temp.:100.06, PWM:33
01:39:06.626 -> Water_Temp.:100.06, PWM:33
01:39:08.942 -> Water_Temp.:100.06, PWM:33
01:39:10.501 -> Water_Temp.:100.06, PWM:33
01:39:11.265 -> Water_Temp.:100.06, PWM:33
01:39:12.062 -> Water_Temp.:100.06, PWM:33
01:39:12.825 -> Water_Temp.:100.06, PWM:33
01:39:13.621 -> Water_Temp.:100.06, PWM:33
01:39:15.173 -> Water_Temp.:100.06, PWM:33
01:39:16.725 -> Water_Temp.:100.17, PWM:33
01:39:18.257 -> Water_Temp.:100.06, PWM:33
01:39:19.050 -> Water_Temp.:100.17, PWM:33
01:39:19.810 -> Water_Temp.:100.17, PWM:33
01:39:22.921 -> Water_Temp.:100.17, PWM:33
01:39:26.005 -> Water_Temp.:100.06, PWM:33
01:39:26.801 -> Water_Temp.:100.06, PWM:33
01:39:27.599 -> Water_Temp.:100.06, PWM:33
01:39:28.359 -> Water_Temp.:100.06, PWM:33
01:39:29.155 -> Water_Temp.:100.06, PWM:33
01:39:30.678 -> Water_Temp.:100.06, PWM:33
01:39:31.469 -> Water_Temp.:100.06, PWM:33
01:39:33.027 -> Water_Temp.:100.06, PWM:33
01:39:35.309 -> Water_Temp.:100.17, PWM:33
01:39:36.827 -> Water_Temp.:100.17, PWM:33
01:39:37.623 -> Water_Temp.:100.17, PWM:33
01:39:39.177 -> Water_Temp.:100.17, PWM:33
01:39:39.973 -> Water_Temp.:100.17, PWM:33
01:39:41.497 -> Water_Temp.:100.17, PWM:37
01:39:42.294 -> Water_Temp.:100.17, PWM:37
01:39:43.057 -> Water_Temp.:100.17, PWM:37
01:39:43.851 -> Water_Temp.:100.17, PWM:37
01:39:44.613 -> Water_Temp.:100.17, PWM:37
01:39:46.166 -> Water_Temp.:100.17, PWM:41
01:39:49.315 -> Water_Temp.:100.17, PWM:41
01:39:50.075 -> Water_Temp.:100.17, PWM:41
01:39:50.871 -> Water_Temp.:100.17, PWM:41
01:39:52.390 -> Water_Temp.:100.17, PWM:41
01:39:53.185 -> Water_Temp.:100.17, PWM:41
01:39:53.943 -> Water_Temp.:100.17, PWM:41
01:39:54.737 -> Water_Temp.:100.17, PWM:41
01:39:55.534 -> Water_Temp.:100.17, PWM:41
01:39:57.089 -> Water_Temp.:100.17, PWM:41
01:39:57.849 -> Water_Temp.:100.17, PWM:41
01:39:59.407 -> Water_Temp.:100.17, PWM:41
01:40:00.203 -> Water_Temp.:100.17, PWM:41
01:40:00.967 -> Water_Temp.:100.17, PWM:41
01:40:03.317 -> Water_Temp.:100.17, PWM:41
01:40:04.866 -> Water_Temp.:101.75, PWM:41
01:40:06.390 -> Water_Temp.:101.41, PWM:41
01:40:07.182 -> Water_Temp.:101.07, PWM:41
01:40:08.740 -> Water_Temp.:100.74, PWM:41
01:40:09.504 -> Water_Temp.:100.62, PWM:41
01:40:13.386 -> Water_Temp.:100.40, PWM:41
01:40:15.702 -> Water_Temp.:100.29, PWM:38
01:40:16.457 -> Water_Temp.:100.29, PWM:37
01:40:17.250 -> Water_Temp.:100.29, PWM:37
01:40:18.800 -> Water_Temp.:100.29, PWM:36
01:40:19.560 -> Water_Temp.:100.29, PWM:36
01:40:24.190 -> Water_Temp.:100.17, PWM:36
01:40:24.989 -> Water_Temp.:100.17, PWM:37
01:40:25.747 -> Water_Temp.:100.17, PWM:36
01:40:26.543 -> Water_Temp.:100.17, PWM:36
01:40:27.334 -> Water_Temp.:100.17, PWM:36
01:40:29.642 -> Water_Temp.:100.17, PWM:36
01:40:30.440 -> Water_Temp.:100.17, PWM:36
01:40:31.205 -> Water_Temp.:100.17, PWM:36
01:40:31.997 -> Water_Temp.:100.17, PWM:36
01:40:32.758 -> Water_Temp.:100.29, PWM:36
01:40:33.556 -> Water_Temp.:100.29, PWM:36
01:40:34.312 -> Water_Temp.:100.29, PWM:36
01:40:35.861 -> Water_Temp.:100.29, PWM:36
01:40:36.621 -> Water_Temp.:100.29, PWM:36
01:40:37.417 -> Water_Temp.:100.29, PWM:36
01:40:38.181 -> Water_Temp.:100.29, PWM:36
01:40:38.973 -> Water_Temp.:100.29, PWM:36
01:40:39.767 -> Water_Temp.:100.29, PWM:36
01:40:40.533 -> Water_Temp.:100.29, PWM:36
01:40:42.088 -> Water_Temp.:100.29, PWM:36
01:40:42.853 -> Water_Temp.:100.29, PWM:36
01:40:43.647 -> Water_Temp.:100.29, PWM:36
01:40:44.408 -> Water_Temp.:100.29, PWM:36
01:40:45.972 -> Water_Temp.:100.29, PWM:36
01:40:47.525 -> Water_Temp.:100.29, PWM:36
01:40:48.317 -> Water_Temp.:100.29, PWM:36
01:40:51.399 -> Water_Temp.:100.29, PWM:36
01:40:52.194 -> Water_Temp.:100.29, PWM:36
01:40:52.957 -> Water_Temp.:100.29, PWM:36
01:40:53.755 -> Water_Temp.:100.29, PWM:36
01:40:55.313 -> Water_Temp.:100.29, PWM:36
01:40:56.073 -> Water_Temp.:100.29, PWM:36
01:40:56.867 -> Water_Temp.:100.29, PWM:36
01:40:57.625 -> Water_Temp.:100.29, PWM:36
01:40:59.187 -> Water_Temp.:100.40, PWM:36
01:40:59.947 -> Water_Temp.:100.40, PWM:36
01:41:01.504 -> Water_Temp.:100.40, PWM:36
01:41:02.298 -> Water_Temp.:100.40, PWM:36
01:41:03.091 -> Water_Temp.:100.40, PWM:36
01:41:03.849 -> Water_Temp.:100.40, PWM:36
01:41:04.648 -> Water_Temp.:100.40, PWM:36
01:41:06.168 -> Water_Temp.:100.40, PWM:36
01:41:06.960 -> Water_Temp.:100.40, PWM:36
01:41:07.725 -> Water_Temp.:100.40, PWM:36
01:41:09.316 -> Water_Temp.:100.40, PWM:36
01:41:10.075 -> Water_Temp.:100.40, PWM:36
01:41:10.872 -> Water_Temp.:100.40, PWM:36
01:41:13.192 -> Water_Temp.:100.40, PWM:36
01:41:13.984 -> Water_Temp.:100.40, PWM:36
01:41:14.743 -> Water_Temp.:100.40, PWM:36
01:41:16.292 -> Water_Temp.:100.40, PWM:37
01:41:17.084 -> Water_Temp.:100.40, PWM:37
01:41:17.843 -> Water_Temp.:100.40, PWM:36
01:41:18.637 -> Water_Temp.:100.40, PWM:36
01:41:19.398 -> Water_Temp.:100.40, PWM:36
01:41:20.192 -> Water_Temp.:100.40, PWM:36
01:41:20.987 -> Water_Temp.:100.29, PWM:36
01:41:22.509 -> Water_Temp.:100.29, PWM:36
01:41:23.308 -> Water_Temp.:100.29, PWM:36
01:41:24.069 -> Water_Temp.:100.40, PWM:36
01:41:24.864 -> Water_Temp.:100.29, PWM:36
01:41:26.420 -> Water_Temp.:100.40, PWM:36
01:41:27.177 -> Water_Temp.:100.40, PWM:36
01:41:27.969 -> Water_Temp.:100.40, PWM:36
01:41:28.763 -> Water_Temp.:100.40, PWM:36
01:41:29.523 -> Water_Temp.:100.40, PWM:36
01:41:31.835 -> Water_Temp.:100.40, PWM:36
01:41:32.595 -> Water_Temp.:100.40, PWM:36
01:41:33.391 -> Water_Temp.:100.40, PWM:36
01:41:34.941 -> Water_Temp.:100.40, PWM:36
01:41:36.499 -> Water_Temp.:100.40, PWM:36
01:41:37.287 -> Water_Temp.:100.40, PWM:36
01:41:38.082 -> Water_Temp.:100.40, PWM:36
01:41:38.845 -> Water_Temp.:100.40, PWM:36
01:41:39.634 -> Water_Temp.:100.40, PWM:36
01:41:41.158 -> Water_Temp.:100.40, PWM:36
01:41:42.717 -> Water_Temp.:100.51, PWM:37
01:41:43.515 -> Water_Temp.:100.51, PWM:36
01:41:44.275 -> Water_Temp.:100.51, PWM:36
01:41:45.831 -> Water_Temp.:100.51, PWM:36
01:41:46.585 -> Water_Temp.:100.51, PWM:36
01:41:47.382 -> Water_Temp.:100.51, PWM:36
01:41:48.940 -> Water_Temp.:100.51, PWM:36
01:41:50.460 -> Water_Temp.:100.51, PWM:36
01:41:51.256 -> Water_Temp.:100.51, PWM:36
01:41:52.017 -> Water_Temp.:100.51, PWM:36
01:41:53.573 -> Water_Temp.:100.51, PWM:36
01:41:55.134 -> Water_Temp.:100.51, PWM:36
01:41:56.690 -> Water_Temp.:100.51, PWM:36
01:41:57.487 -> Water_Temp.:100.51, PWM:36
01:41:58.252 -> Water_Temp.:100.51, PWM:36
01:41:59.049 -> Water_Temp.:100.51, PWM:36
01:41:59.811 -> Water_Temp.:100.51, PWM:36
01:42:00.605 -> Water_Temp.:100.51, PWM:36
01:42:01.369 -> Water_Temp.:100.51, PWM:36
01:42:03.690 -> Water_Temp.:100.51, PWM:36
01:42:05.247 -> Water_Temp.:100.62, PWM:36
01:42:06.039 -> Water_Temp.:100.62, PWM:36
01:42:08.388 -> Water_Temp.:100.62, PWM:36
01:42:10.700 -> Water_Temp.:100.62, PWM:36
01:42:11.494 -> Water_Temp.:100.62, PWM:36
01:42:12.286 -> Water_Temp.:100.62, PWM:36
01:42:14.570 -> Water_Temp.:100.62, PWM:36
01:42:15.364 -> Water_Temp.:100.62, PWM:36
01:42:16.123 -> Water_Temp.:100.62, PWM:36
01:42:16.915 -> Water_Temp.:100.62, PWM:36
01:42:17.680 -> Water_Temp.:100.62, PWM:36
01:42:18.479 -> Water_Temp.:100.62, PWM:36
01:42:19.242 -> Water_Temp.:100.62, PWM:36
01:42:20.768 -> Water_Temp.:100.62, PWM:36
01:42:21.567 -> Water_Temp.:100.62, PWM:36
01:42:22.329 -> Water_Temp.:100.62, PWM:36
01:42:23.125 -> Water_Temp.:100.62, PWM:36
01:42:23.923 -> Water_Temp.:100.62, PWM:36
01:42:25.450 -> Water_Temp.:100.62, PWM:36
01:42:26.207 -> Water_Temp.:100.62, PWM:36
01:42:27.757 -> Water_Temp.:100.62, PWM:36
01:42:29.315 -> Water_Temp.:100.62, PWM:36
01:42:30.870 -> Water_Temp.:100.62, PWM:36
01:42:32.458 -> Water_Temp.:100.62, PWM:36
01:42:34.017 -> Water_Temp.:100.62, PWM:36
01:42:36.298 -> Water_Temp.:100.62, PWM:36
01:42:37.856 -> Water_Temp.:100.62, PWM:36
01:42:38.619 -> Water_Temp.:100.62, PWM:36
01:42:40.178 -> Water_Temp.:100.62, PWM:36
01:42:41.733 -> Water_Temp.:100.62, PWM:36
01:42:43.258 -> Water_Temp.:100.62, PWM:36
01:42:44.050 -> Water_Temp.:100.62, PWM:36
01:42:44.842 -> Water_Temp.:100.62, PWM:36
01:42:45.604 -> Water_Temp.:100.62, PWM:36
01:42:46.399 -> Water_Temp.:100.62, PWM:36
01:42:47.160 -> Water_Temp.:100.62, PWM:36
01:42:47.952 -> Water_Temp.:100.74, PWM:36
01:42:51.035 -> Water_Temp.:100.74, PWM:36
01:42:51.794 -> Water_Temp.:100.74, PWM:36
01:42:52.586 -> Water_Temp.:100.74, PWM:36
01:42:57.184 -> Water_Temp.:100.74, PWM:36
01:42:57.981 -> Water_Temp.:100.74, PWM:36
01:42:58.744 -> Water_Temp.:100.74, PWM:36
01:43:00.299 -> Water_Temp.:100.74, PWM:36
01:43:01.062 -> Water_Temp.:100.74, PWM:36
01:43:01.857 -> Water_Temp.:100.74, PWM:36
01:43:02.619 -> Water_Temp.:100.74, PWM:36
01:43:04.180 -> Water_Temp.:100.74, PWM:36
01:43:04.945 -> Water_Temp.:100.74, PWM:36
01:43:05.741 -> Water_Temp.:100.74, PWM:36
01:43:06.505 -> Water_Temp.:100.74, PWM:36
01:43:07.268 -> Water_Temp.:100.74, PWM:36
01:43:08.064 -> Water_Temp.:100.74, PWM:36
01:43:11.910 -> Water_Temp.:100.74, PWM:36
01:43:12.708 -> Water_Temp.:100.74, PWM:36
01:43:13.467 -> Water_Temp.:100.74, PWM:36
01:43:14.262 -> Water_Temp.:100.74, PWM:36
01:43:15.021 -> Water_Temp.:100.74, PWM:36
01:43:15.814 -> Water_Temp.:100.74, PWM:36
01:43:17.368 -> Water_Temp.:100.74, PWM:36
01:43:18.919 -> Water_Temp.:100.74, PWM:36
01:43:19.681 -> Water_Temp.:100.74, PWM:36
01:43:20.481 -> Water_Temp.:100.74, PWM:36
01:43:21.245 -> Water_Temp.:100.74, PWM:36
01:43:22.794 -> Water_Temp.:100.74, PWM:36
01:43:24.347 -> Water_Temp.:100.74, PWM:36
01:43:25.108 -> Water_Temp.:100.74, PWM:36
01:43:25.905 -> Water_Temp.:100.74, PWM:36
01:43:26.667 -> Water_Temp.:100.74, PWM:36
01:43:27.463 -> Water_Temp.:100.74, PWM:36
01:43:28.256 -> Water_Temp.:100.74, PWM:36
01:43:29.020 -> Water_Temp.:100.74, PWM:36
01:43:32.824 -> Water_Temp.:100.74, PWM:36
01:43:33.620 -> Water_Temp.:100.74, PWM:36
01:43:35.175 -> Water_Temp.:100.74, PWM:36
01:43:37.494 -> Water_Temp.:100.74, PWM:36
01:43:39.054 -> Water_Temp.:100.74, PWM:36
01:43:39.850 -> Water_Temp.:100.74, PWM:36
01:43:42.955 -> Water_Temp.:100.74, PWM:36
01:43:43.751 -> Water_Temp.:100.74, PWM:36
01:43:45.301 -> Water_Temp.:100.74, PWM:36
01:43:46.059 -> Water_Temp.:100.74, PWM:36
01:43:46.851 -> Water_Temp.:100.74, PWM:36
01:43:48.406 -> Water_Temp.:100.74, PWM:36
01:43:49.199 -> Water_Temp.:100.74, PWM:36
01:43:49.958 -> Water_Temp.:100.74, PWM:36
01:43:50.755 -> Water_Temp.:100.74, PWM:36
01:43:51.517 -> Water_Temp.:100.74, PWM:36
01:43:52.314 -> Water_Temp.:100.74, PWM:36
01:43:53.072 -> Water_Temp.:100.74, PWM:36
01:43:54.630 -> Water_Temp.:100.74, PWM:36
01:43:55.427 -> Water_Temp.:100.74, PWM:36
01:43:56.188 -> Water_Temp.:100.74, PWM:36
01:43:56.984 -> Water_Temp.:100.74, PWM:36
01:43:57.744 -> Water_Temp.:100.74, PWM:36
01:43:58.542 -> Water_Temp.:100.74, PWM:36
01:43:59.306 -> Water_Temp.:100.74, PWM:36
01:44:00.869 -> Water_Temp.:100.74, PWM:36
01:44:02.426 -> Water_Temp.:100.74, PWM:36
01:44:03.183 -> Water_Temp.:100.74, PWM:36
01:44:03.974 -> Water_Temp.:100.74, PWM:36
01:44:04.770 -> Water_Temp.:100.74, PWM:36
01:44:05.531 -> Water_Temp.:100.74, PWM:36
01:44:07.088 -> Water_Temp.:100.74, PWM:36
01:44:07.849 -> Water_Temp.:100.85, PWM:36
01:44:08.644 -> Water_Temp.:100.74, PWM:36
01:44:09.404 -> Water_Temp.:100.74, PWM:36
01:44:11.755 -> Water_Temp.:100.85, PWM:36
01:44:12.513 -> Water_Temp.:100.85, PWM:36
01:44:13.311 -> Water_Temp.:100.85, PWM:36
01:44:14.071 -> Water_Temp.:100.85, PWM:36
01:44:14.868 -> Water_Temp.:100.85, PWM:36
01:44:15.631 -> Water_Temp.:100.85, PWM:36
01:44:16.425 -> Water_Temp.:100.85, PWM:36
01:44:17.216 -> Water_Temp.:100.85, PWM:36
01:44:17.976 -> Water_Temp.:100.85, PWM:36
01:44:18.771 -> Water_Temp.:100.85, PWM:36
01:44:19.535 -> Water_Temp.:100.85, PWM:36
01:44:21.092 -> Water_Temp.:100.85, PWM:36
01:44:21.849 -> Water_Temp.:100.85, PWM:36
01:44:22.647 -> Water_Temp.:100.85, PWM:36
01:44:23.407 -> Water_Temp.:100.85, PWM:36
01:44:24.203 -> Water_Temp.:100.85, PWM:36
01:44:25.762 -> Water_Temp.:100.85, PWM:36
01:44:26.528 -> Water_Temp.:100.85, PWM:36
01:44:28.084 -> Water_Temp.:100.85, PWM:36
01:44:29.673 -> Water_Temp.:100.85, PWM:36
01:44:31.193 -> Water_Temp.:100.85, PWM:36
01:44:31.984 -> Water_Temp.:100.85, PWM:36
01:44:32.778 -> Water_Temp.:100.85, PWM:36
01:44:33.539 -> Water_Temp.:100.85, PWM:36
01:44:35.093 -> Water_Temp.:100.96, PWM:37
01:44:37.414 -> Water_Temp.:100.96, PWM:36
01:44:38.171 -> Water_Temp.:100.96, PWM:36
01:44:38.967 -> Water_Temp.:100.96, PWM:36
01:44:40.487 -> Water_Temp.:100.96, PWM:36
01:44:42.045 -> Water_Temp.:100.96, PWM:36
01:44:42.842 -> Water_Temp.:100.96, PWM:36
01:44:43.605 -> Water_Temp.:100.96, PWM:36
01:44:44.402 -> Water_Temp.:100.96, PWM:36
01:44:45.957 -> Water_Temp.:100.96, PWM:36
01:44:46.717 -> Water_Temp.:100.96, PWM:36
01:44:47.512 -> Water_Temp.:100.96, PWM:36
01:44:48.306 -> Water_Temp.:100.96, PWM:36
01:44:49.067 -> Water_Temp.:100.96, PWM:36
01:44:49.865 -> Water_Temp.:100.96, PWM:36
01:44:50.625 -> Water_Temp.:100.96, PWM:36
01:44:51.417 -> Water_Temp.:100.96, PWM:36
01:44:52.179 -> Water_Temp.:100.96, PWM:36
01:44:53.728 -> Water_Temp.:100.96, PWM:36
01:44:54.526 -> Water_Temp.:100.96, PWM:36
01:44:56.077 -> Water_Temp.:100.96, PWM:36
01:44:56.870 -> Water_Temp.:100.96, PWM:36
01:44:58.393 -> Water_Temp.:100.96, PWM:36
01:44:59.155 -> Water_Temp.:100.96, PWM:36
01:44:59.948 -> Water_Temp.:100.96, PWM:36
01:45:00.739 -> Water_Temp.:100.96, PWM:36
01:45:01.502 -> Water_Temp.:100.96, PWM:36
01:45:02.298 -> Water_Temp.:100.96, PWM:36
01:45:03.853 -> Water_Temp.:100.96, PWM:36
01:45:04.614 -> Water_Temp.:100.96, PWM:36
01:45:05.407 -> Water_Temp.:100.96, PWM:36
01:45:07.696 -> Water_Temp.:100.96, PWM:36
01:45:08.488 -> Water_Temp.:100.96, PWM:36
01:45:09.249 -> Water_Temp.:100.96, PWM:36
01:45:10.044 -> Water_Temp.:100.96, PWM:36
01:45:10.807 -> Water_Temp.:101.07, PWM:36
01:45:12.363 -> Water_Temp.:101.07, PWM:36
01:45:13.914 -> Water_Temp.:101.07, PWM:36
01:45:14.676 -> Water_Temp.:101.07, PWM:36
01:45:15.472 -> Water_Temp.:101.07, PWM:36
01:45:17.028 -> Water_Temp.:101.07, PWM:36
01:45:17.789 -> Water_Temp.:101.07, PWM:36
01:45:18.581 -> Water_Temp.:101.07, PWM:36
01:45:19.343 -> Water_Temp.:101.07, PWM:36
01:45:20.140 -> Water_Temp.:101.07, PWM:36
01:45:20.902 -> Water_Temp.:101.07, PWM:36
01:45:22.459 -> Water_Temp.:101.07, PWM:36
01:45:24.015 -> Water_Temp.:101.07, PWM:36
01:45:24.776 -> Water_Temp.:101.07, PWM:36
01:45:25.572 -> Water_Temp.:101.19, PWM:36
01:45:26.333 -> Water_Temp.:101.19, PWM:36
01:45:27.126 -> Water_Temp.:101.19, PWM:36
01:45:28.680 -> Water_Temp.:101.19, PWM:36
01:45:29.475 -> Water_Temp.:101.19, PWM:36
01:45:30.239 -> Water_Temp.:101.19, PWM:36
01:45:31.036 -> Water_Temp.:101.19, PWM:36
01:45:31.796 -> Water_Temp.:101.19, PWM:36
01:45:33.349 -> Water_Temp.:101.19, PWM:36
01:45:34.143 -> Water_Temp.:101.19, PWM:36
01:45:34.900 -> Water_Temp.:101.19, PWM:36
01:45:35.697 -> Water_Temp.:101.19, PWM:36
01:45:37.256 -> Water_Temp.:101.19, PWM:36
01:45:38.015 -> Water_Temp.:101.19, PWM:36
01:45:39.574 -> Water_Temp.:101.19, PWM:36
01:45:40.367 -> Water_Temp.:101.19, PWM:36
01:45:41.133 -> Water_Temp.:101.19, PWM:39
01:45:44.243 -> Water_Temp.:101.19, PWM:39
01:45:45.041 -> Water_Temp.:102.87, PWM:40
01:45:45.800 -> Water_Temp.:104.90, PWM:39
01:45:47.365 -> Water_Temp.:102.09, PWM:40
01:45:48.128 -> Water_Temp.:101.86, PWM:40
01:45:51.242 -> Water_Temp.:101.52, PWM:39
01:45:52.830 -> Water_Temp.:101.52, PWM:40
01:45:53.590 -> Water_Temp.:101.41, PWM:39
01:45:55.143 -> Water_Temp.:101.41, PWM:39
01:45:55.939 -> Water_Temp.:101.41, PWM:40
01:45:56.701 -> Water_Temp.:101.41, PWM:39
01:45:57.494 -> Water_Temp.:101.41, PWM:40
01:45:58.250 -> Water_Temp.:101.41, PWM:40
01:45:59.042 -> Water_Temp.:101.41, PWM:39
01:46:00.597 -> Water_Temp.:101.41, PWM:39
01:46:03.676 -> Water_Temp.:101.30, PWM:39
01:46:04.470 -> Water_Temp.:101.30, PWM:40
01:46:05.229 -> Water_Temp.:101.30, PWM:40
01:46:06.027 -> Water_Temp.:101.30, PWM:39
01:46:06.820 -> Water_Temp.:101.30, PWM:39
01:46:07.581 -> Water_Temp.:101.30, PWM:39
01:46:08.377 -> Water_Temp.:101.30, PWM:39
01:46:09.897 -> Water_Temp.:101.30, PWM:39
01:46:12.251 -> Water_Temp.:101.30, PWM:40
01:46:14.534 -> Water_Temp.:101.30, PWM:39
01:46:15.329 -> Water_Temp.:101.30, PWM:40
01:46:16.123 -> Water_Temp.:101.30, PWM:39
01:46:16.882 -> Water_Temp.:101.30, PWM:39
01:46:17.677 -> Water_Temp.:101.30, PWM:40
01:46:18.436 -> Water_Temp.:101.30, PWM:39
01:46:20.022 -> Water_Temp.:101.30, PWM:39
01:46:20.777 -> Water_Temp.:101.30, PWM:39
01:46:22.360 -> Water_Temp.:101.30, PWM:40
01:46:23.882 -> Water_Temp.:101.30, PWM:40
01:46:26.175 -> Water_Temp.:101.30, PWM:39
01:46:26.969 -> Water_Temp.:101.30, PWM:39
01:46:27.729 -> Water_Temp.:101.30, PWM:40
01:46:30.812 -> Water_Temp.:101.30, PWM:39
01:46:31.608 -> Water_Temp.:101.30, PWM:39
01:46:32.368 -> Water_Temp.:101.30, PWM:39
01:46:33.924 -> Water_Temp.:101.41, PWM:39
01:46:34.718 -> Water_Temp.:101.41, PWM:40
01:46:37.824 -> Water_Temp.:101.41, PWM:39
01:46:38.618 -> Water_Temp.:101.41, PWM:39
01:46:39.379 -> Water_Temp.:101.41, PWM:40
01:46:40.932 -> Water_Temp.:101.41, PWM:39
01:46:41.725 -> Water_Temp.:101.41, PWM:39
01:46:44.042 -> Water_Temp.:101.41, PWM:39
01:46:46.387 -> Water_Temp.:101.41, PWM:40
01:46:47.182 -> Water_Temp.:101.41, PWM:39
01:46:47.941 -> Water_Temp.:101.41, PWM:39
01:46:48.739 -> Water_Temp.:101.41, PWM:39
01:46:50.296 -> Water_Temp.:101.41, PWM:39
01:46:51.058 -> Water_Temp.:101.41, PWM:40
01:46:51.854 -> Water_Temp.:101.41, PWM:39
01:46:52.616 -> Water_Temp.:101.41, PWM:40
01:46:53.412 -> Water_Temp.:101.41, PWM:39
01:46:54.206 -> Water_Temp.:101.41, PWM:40
01:46:54.967 -> Water_Temp.:101.41, PWM:39
01:46:56.517 -> Water_Temp.:101.41, PWM:39
01:46:57.279 -> Water_Temp.:101.41, PWM:39
01:46:58.076 -> Water_Temp.:101.41, PWM:40
01:46:58.840 -> Water_Temp.:101.41, PWM:39
01:47:00.396 -> Water_Temp.:101.52, PWM:39
01:47:01.193 -> Water_Temp.:101.52, PWM:39
01:47:01.952 -> Water_Temp.:101.52, PWM:39
01:47:02.750 -> Water_Temp.:101.52, PWM:39
01:47:03.509 -> Water_Temp.:101.52, PWM:40
01:47:04.302 -> Water_Temp.:101.41, PWM:39
01:47:05.096 -> Water_Temp.:101.52, PWM:39
01:47:06.639 -> Water_Temp.:101.52, PWM:40
01:47:07.397 -> Water_Temp.:101.52, PWM:40
01:47:08.954 -> Water_Temp.:101.52, PWM:40
01:47:09.752 -> Water_Temp.:101.52, PWM:40
01:47:11.310 -> Water_Temp.:101.52, PWM:40
01:47:12.830 -> Water_Temp.:101.52, PWM:39
01:47:13.623 -> Water_Temp.:101.52, PWM:40
01:47:14.383 -> Water_Temp.:101.52, PWM:39
01:47:15.177 -> Water_Temp.:101.52, PWM:40
01:47:15.970 -> Water_Temp.:101.52, PWM:39
01:47:16.729 -> Water_Temp.:101.52, PWM:39
01:47:17.517 -> Water_Temp.:101.52, PWM:39
01:47:18.307 -> Water_Temp.:101.52, PWM:40
01:47:19.071 -> Water_Temp.:101.52, PWM:39
01:47:19.868 -> Water_Temp.:101.52, PWM:39
01:47:21.421 -> Water_Temp.:101.52, PWM:40
01:47:22.211 -> Water_Temp.:101.52, PWM:40
01:47:22.969 -> Water_Temp.:101.52, PWM:39
01:47:24.528 -> Water_Temp.:101.52, PWM:40
01:47:26.847 -> Water_Temp.:101.52, PWM:40
01:47:27.645 -> Water_Temp.:101.52, PWM:39
01:47:28.408 -> Water_Temp.:101.52, PWM:39
01:47:29.203 -> Water_Temp.:101.52, PWM:39
01:47:29.995 -> Water_Temp.:101.52, PWM:39
01:47:30.758 -> Water_Temp.:101.52, PWM:39
01:47:31.554 -> Water_Temp.:101.52, PWM:39
01:47:32.318 -> Water_Temp.:101.52, PWM:40
01:47:33.870 -> Water_Temp.:101.52, PWM:39
01:47:35.390 -> Water_Temp.:101.52, PWM:40
01:47:36.185 -> Water_Temp.:101.52, PWM:40
01:47:36.974 -> Water_Temp.:101.52, PWM:40
01:47:38.493 -> Water_Temp.:101.52, PWM:40
01:47:42.369 -> Water_Temp.:101.64, PWM:40
01:47:43.164 -> Water_Temp.:101.64, PWM:40
01:47:43.924 -> Water_Temp.:101.64, PWM:40
01:47:44.722 -> Water_Temp.:101.64, PWM:39
01:47:45.482 -> Water_Temp.:101.64, PWM:40
01:47:47.037 -> Water_Temp.:101.64, PWM:39
01:47:50.144 -> Water_Temp.:101.64, PWM:39
01:47:51.670 -> Water_Temp.:101.64, PWM:40
01:47:52.467 -> Water_Temp.:101.64, PWM:39
01:47:53.231 -> Water_Temp.:101.64, PWM:39
01:47:54.022 -> Water_Temp.:101.64, PWM:39
01:47:54.780 -> Water_Temp.:101.64, PWM:39
01:47:55.571 -> Water_Temp.:101.64, PWM:39
01:47:56.364 -> Water_Temp.:101.64, PWM:40
01:47:57.119 -> Water_Temp.:101.64, PWM:40
01:47:57.910 -> Water_Temp.:101.64, PWM:39
01:48:00.217 -> Water_Temp.:101.64, PWM:39
01:48:01.015 -> Water_Temp.:101.64, PWM:39
01:48:03.359 -> Water_Temp.:101.64, PWM:39
01:48:04.118 -> Water_Temp.:101.64, PWM:39
01:48:04.908 -> Water_Temp.:101.64, PWM:39
01:48:05.700 -> Water_Temp.:101.75, PWM:40
01:48:06.457 -> Water_Temp.:101.75, PWM:39
01:48:07.251 -> Water_Temp.:101.75, PWM:40
01:48:08.806 -> Water_Temp.:101.75, PWM:39
01:48:09.566 -> Water_Temp.:101.75, PWM:40
01:48:10.360 -> Water_Temp.:101.75, PWM:39
01:48:11.883 -> Water_Temp.:101.75, PWM:39
01:48:14.229 -> Water_Temp.:101.75, PWM:40
01:48:15.780 -> Water_Temp.:101.75, PWM:39
01:48:16.576 -> Water_Temp.:101.64, PWM:39
01:48:17.337 -> Water_Temp.:101.64, PWM:40
01:48:18.133 -> Water_Temp.:101.64, PWM:39
01:48:18.893 -> Water_Temp.:101.64, PWM:39
01:48:19.684 -> Water_Temp.:101.64, PWM:39
01:48:22.727 -> Water_Temp.:101.64, PWM:39
01:48:24.283 -> Water_Temp.:101.64, PWM:39
01:48:25.078 -> Water_Temp.:101.64, PWM:40
01:48:25.841 -> Water_Temp.:101.64, PWM:40
01:48:26.631 -> Water_Temp.:101.64, PWM:39
01:48:27.395 -> Water_Temp.:101.64, PWM:39
01:48:28.192 -> Water_Temp.:101.64, PWM:40
01:48:28.985 -> Water_Temp.:101.64, PWM:40
01:48:29.750 -> Water_Temp.:101.64, PWM:39
01:48:30.547 -> Water_Temp.:101.64, PWM:39
01:48:31.310 -> Water_Temp.:101.64, PWM:40
01:48:32.102 -> Water_Temp.:101.64, PWM:39
01:48:33.627 -> Water_Temp.:101.75, PWM:39
01:48:34.419 -> Water_Temp.:101.75, PWM:40
01:48:35.174 -> Water_Temp.:101.75, PWM:39
01:48:35.963 -> Water_Temp.:101.75, PWM:39
01:48:36.761 -> Water_Temp.:101.75, PWM:39
01:48:37.522 -> Water_Temp.:101.75, PWM:39
01:48:38.317 -> Water_Temp.:101.75, PWM:39
01:48:39.071 -> Water_Temp.:101.75, PWM:39
01:48:39.864 -> Water_Temp.:101.75, PWM:40
01:48:42.908 -> Water_Temp.:101.75, PWM:40
01:48:45.257 -> Water_Temp.:103.44, PWM:39
01:48:46.018 -> Water_Temp.:105.57, PWM:39
01:48:46.808 -> Water_Temp.:103.10, PWM:39
01:48:47.604 -> Water_Temp.:102.65, PWM:40
01:48:48.364 -> Water_Temp.:102.54, PWM:39
01:48:49.162 -> Water_Temp.:102.31, PWM:39
01:48:51.446 -> Water_Temp.:102.09, PWM:40
01:48:53.007 -> Water_Temp.:102.09, PWM:40
01:48:53.768 -> Water_Temp.:101.97, PWM:40
01:48:54.563 -> Water_Temp.:101.97, PWM:39
01:48:55.357 -> Water_Temp.:101.97, PWM:40
01:48:56.114 -> Water_Temp.:101.97, PWM:40
01:48:56.908 -> Water_Temp.:101.97, PWM:39
01:48:58.431 -> Water_Temp.:101.97, PWM:39
01:48:59.227 -> Water_Temp.:101.86, PWM:39
01:49:01.576 -> Water_Temp.:101.86, PWM:39
01:49:02.336 -> Water_Temp.:101.86, PWM:39
01:49:03.131 -> Water_Temp.:101.86, PWM:39
01:49:04.686 -> Water_Temp.:101.86, PWM:39
01:49:06.239 -> Water_Temp.:101.97, PWM:40
01:49:07.033 -> Water_Temp.:101.97, PWM:39
01:49:08.589 -> Water_Temp.:101.97, PWM:39
01:49:09.350 -> Water_Temp.:101.97, PWM:40
01:49:10.142 -> Water_Temp.:101.97, PWM:40
01:49:11.664 -> Water_Temp.:101.97, PWM:39
01:49:12.460 -> Water_Temp.:101.97, PWM:40
01:49:13.254 -> Water_Temp.:101.97, PWM:39
01:49:14.810 -> Water_Temp.:101.97, PWM:39
01:49:15.570 -> Water_Temp.:101.97, PWM:39
01:49:16.363 -> Water_Temp.:101.97, PWM:40
01:49:17.121 -> Water_Temp.:101.97, PWM:39
01:49:18.708 -> Water_Temp.:101.97, PWM:39
01:49:19.467 -> Water_Temp.:101.86, PWM:39
01:49:20.261 -> Water_Temp.:101.86, PWM:39
01:49:22.578 -> Water_Temp.:101.86, PWM:40
01:49:24.134 -> Water_Temp.:101.86, PWM:40
01:49:24.927 -> Water_Temp.:101.86, PWM:40
01:49:26.451 -> Water_Temp.:101.86, PWM:40
01:49:27.248 -> Water_Temp.:101.97, PWM:40
01:49:29.563 -> Water_Temp.:101.97, PWM:39
01:49:31.120 -> Water_Temp.:101.86, PWM:39
01:49:31.880 -> Water_Temp.:101.86, PWM:39
01:49:32.671 -> Water_Temp.:101.86, PWM:40
01:49:34.193 -> Water_Temp.:101.97, PWM:39
01:49:34.989 -> Water_Temp.:101.97, PWM:39
01:49:35.749 -> Water_Temp.:101.97, PWM:39
01:49:37.304 -> Water_Temp.:101.97, PWM:39
01:49:38.100 -> Water_Temp.:101.97, PWM:39
01:49:39.664 -> Water_Temp.:101.97, PWM:39
01:49:40.424 -> Water_Temp.:101.97, PWM:39
01:49:41.979 -> Water_Temp.:101.97, PWM:39
01:49:43.541 -> Water_Temp.:101.97, PWM:39
01:49:44.337 -> Water_Temp.:101.97, PWM:40
01:49:45.102 -> Water_Temp.:101.97, PWM:39
01:49:45.898 -> Water_Temp.:101.97, PWM:39
01:49:46.656 -> Water_Temp.:101.97, PWM:39
01:49:47.453 -> Water_Temp.:101.97, PWM:39
01:49:48.216 -> Water_Temp.:101.97, PWM:39
01:49:49.016 -> Water_Temp.:101.97, PWM:39
01:49:49.780 -> Water_Temp.:101.97, PWM:40
01:49:50.546 -> Water_Temp.:101.97, PWM:40
01:49:51.341 -> Water_Temp.:101.97, PWM:39
01:49:52.135 -> Water_Temp.:102.09, PWM:39
01:49:54.421 -> Water_Temp.:102.09, PWM:39
01:49:57.538 -> Water_Temp.:102.09, PWM:39
01:49:58.303 -> Water_Temp.:102.09, PWM:39
01:49:59.097 -> Water_Temp.:102.09, PWM:39
01:50:00.613 -> Water_Temp.:102.09, PWM:39
01:50:02.960 -> Water_Temp.:102.09, PWM:40
01:50:04.514 -> Water_Temp.:102.09, PWM:39
01:50:05.301 -> Water_Temp.:102.09, PWM:39
01:50:06.095 -> Water_Temp.:102.09, PWM:40
01:50:06.851 -> Water_Temp.:102.09, PWM:39
01:50:07.641 -> Water_Temp.:102.09, PWM:39
01:50:08.402 -> Water_Temp.:102.09, PWM:39
01:50:10.752 -> Water_Temp.:102.09, PWM:39
01:50:11.509 -> Water_Temp.:102.09, PWM:40
01:50:12.303 -> Water_Temp.:102.09, PWM:40
01:50:13.061 -> Water_Temp.:102.09, PWM:39
01:50:13.855 -> Water_Temp.:102.09, PWM:40
01:50:15.408 -> Water_Temp.:102.09, PWM:39
01:50:16.961 -> Water_Temp.:102.09, PWM:40
01:50:20.069 -> Water_Temp.:102.09, PWM:40
01:50:20.860 -> Water_Temp.:102.09, PWM:39
01:50:22.408 -> Water_Temp.:102.09, PWM:39
01:50:23.200 -> Water_Temp.:102.09, PWM:39
01:50:23.960 -> Water_Temp.:102.09, PWM:40
01:50:24.755 -> Water_Temp.:102.09, PWM:40
01:50:25.516 -> Water_Temp.:102.09, PWM:39
01:50:26.303 -> Water_Temp.:102.09, PWM:39
01:50:27.097 -> Water_Temp.:102.09, PWM:39
01:50:27.856 -> Water_Temp.:102.09, PWM:39
01:50:28.648 -> Water_Temp.:102.09, PWM:40
01:50:29.439 -> Water_Temp.:102.09, PWM:39
01:50:30.198 -> Water_Temp.:102.09, PWM:39
01:50:30.995 -> Water_Temp.:102.09, PWM:39
01:50:32.550 -> Water_Temp.:102.09, PWM:40
01:50:33.310 -> Water_Temp.:102.09, PWM:39
01:50:34.101 -> Water_Temp.:102.09, PWM:39
01:50:34.860 -> Water_Temp.:102.09, PWM:39
01:50:35.658 -> Water_Temp.:102.09, PWM:39
01:50:37.967 -> Water_Temp.:102.09, PWM:39
01:50:38.729 -> Water_Temp.:102.09, PWM:39
01:50:39.528 -> Water_Temp.:102.09, PWM:40
01:50:40.323 -> Water_Temp.:102.09, PWM:40
01:50:41.876 -> Water_Temp.:102.09, PWM:39
01:50:42.640 -> Water_Temp.:102.09, PWM:40
01:50:44.199 -> Water_Temp.:102.09, PWM:40
01:50:44.996 -> Water_Temp.:102.09, PWM:39
01:50:45.757 -> Water_Temp.:102.09, PWM:39
01:50:46.549 -> Water_Temp.:102.09, PWM:40
01:50:48.102 -> Water_Temp.:102.09, PWM:39
01:50:48.861 -> Water_Temp.:102.09, PWM:39
01:50:49.655 -> Water_Temp.:102.09, PWM:39
01:50:50.414 -> Water_Temp.:102.09, PWM:39
01:50:51.208 -> Water_Temp.:102.20, PWM:39
01:50:52.003 -> Water_Temp.:102.09, PWM:39
01:50:53.529 -> Water_Temp.:102.20, PWM:40
01:50:54.319 -> Water_Temp.:102.20, PWM:40
01:50:55.084 -> Water_Temp.:102.20, PWM:40
01:50:55.880 -> Water_Temp.:102.20, PWM:39
01:50:57.435 -> Water_Temp.:102.20, PWM:40
01:50:58.199 -> Water_Temp.:102.20, PWM:39
01:50:58.993 -> Water_Temp.:102.20, PWM:40
01:50:59.753 -> Water_Temp.:102.20, PWM:39
01:51:00.543 -> Water_Temp.:102.20, PWM:39
01:51:02.097 -> Water_Temp.:102.20, PWM:40
01:51:02.860 -> Water_Temp.:102.20, PWM:39
01:51:03.653 -> Water_Temp.:102.20, PWM:39
01:51:04.415 -> Water_Temp.:102.20, PWM:40
01:51:05.973 -> Water_Temp.:102.20, PWM:39
01:51:06.769 -> Water_Temp.:102.20, PWM:39
01:51:07.531 -> Water_Temp.:102.20, PWM:40
01:51:08.330 -> Water_Temp.:102.20, PWM:39
01:51:11.434 -> Water_Temp.:102.31, PWM:39
01:51:13.756 -> Water_Temp.:102.31, PWM:39
01:51:14.553 -> Water_Temp.:102.31, PWM:40
01:51:16.840 -> Water_Temp.:102.31, PWM:39
01:51:19.195 -> Water_Temp.:102.31, PWM:39
01:51:19.956 -> Water_Temp.:102.31, PWM:39
01:51:20.756 -> Water_Temp.:102.31, PWM:39
01:51:21.517 -> Water_Temp.:102.31, PWM:40
01:51:23.111 -> Water_Temp.:102.31, PWM:40
01:51:24.637 -> Water_Temp.:102.31, PWM:40
01:51:26.950 -> Water_Temp.:102.31, PWM:39
01:51:27.747 -> Water_Temp.:102.42, PWM:39
01:51:28.509 -> Water_Temp.:102.31, PWM:40
01:51:30.827 -> Water_Temp.:102.31, PWM:40
01:51:31.622 -> Water_Temp.:102.31, PWM:40
01:51:32.380 -> Water_Temp.:102.31, PWM:40
01:51:33.177 -> Water_Temp.:102.31, PWM:39
01:51:34.731 -> Water_Temp.:102.31, PWM:39
01:51:36.288 -> Water_Temp.:102.31, PWM:39
01:51:37.050 -> Water_Temp.:102.31, PWM:40
01:51:37.843 -> Water_Temp.:104.00, PWM:40
01:51:39.394 -> Water_Temp.:103.66, PWM:40
01:51:40.189 -> Water_Temp.:103.21, PWM:39
01:51:40.953 -> Water_Temp.:102.99, PWM:40
01:51:41.745 -> Water_Temp.:102.87, PWM:39
01:51:42.511 -> Water_Temp.:102.87, PWM:40
01:51:46.387 -> Water_Temp.:102.54, PWM:39
01:51:47.149 -> Water_Temp.:102.54, PWM:39
01:51:47.944 -> Water_Temp.:102.54, PWM:39
01:51:48.703 -> Water_Temp.:102.54, PWM:40
01:51:49.498 -> Water_Temp.:102.42, PWM:39
01:51:50.290 -> Water_Temp.:102.42, PWM:40
01:51:51.805 -> Water_Temp.:102.42, PWM:39
01:51:52.601 -> Water_Temp.:102.42, PWM:39
01:51:53.364 -> Water_Temp.:102.42, PWM:40
01:51:54.157 -> Water_Temp.:102.42, PWM:40
01:51:54.920 -> Water_Temp.:102.42, PWM:39
01:51:55.717 -> Water_Temp.:102.42, PWM:39
01:51:57.270 -> Water_Temp.:102.42, PWM:40
01:51:58.031 -> Water_Temp.:102.42, PWM:39
01:51:58.826 -> Water_Temp.:102.42, PWM:39
01:51:59.620 -> Water_Temp.:102.42, PWM:39
01:52:00.383 -> Water_Temp.:102.42, PWM:40
01:52:01.942 -> Water_Temp.:102.42, PWM:39
01:52:02.704 -> Water_Temp.:102.42, PWM:39
01:52:03.497 -> Water_Temp.:102.42, PWM:39
01:52:05.846 -> Water_Temp.:102.42, PWM:39
01:52:06.605 -> Water_Temp.:102.42, PWM:40
01:52:07.398 -> Water_Temp.:102.42, PWM:40
01:52:08.952 -> Water_Temp.:102.42, PWM:39
01:52:09.715 -> Water_Temp.:102.42, PWM:39
01:52:10.510 -> Water_Temp.:102.42, PWM:39
01:52:12.833 -> Water_Temp.:102.42, PWM:40
01:52:14.392 -> Water_Temp.:102.42, PWM:39
01:52:16.745 -> Water_Temp.:102.42, PWM:39
01:52:18.305 -> Water_Temp.:102.42, PWM:39
01:52:19.070 -> Water_Temp.:102.54, PWM:40
01:52:20.629 -> Water_Temp.:102.54, PWM:40
01:52:21.427 -> Water_Temp.:102.54, PWM:40
01:52:22.182 -> Water_Temp.:102.54, PWM:39
01:52:22.976 -> Water_Temp.:102.54, PWM:40
01:52:24.497 -> Water_Temp.:102.54, PWM:40
01:52:26.850 -> Water_Temp.:102.54, PWM:40
01:52:27.605 -> Water_Temp.:102.54, PWM:39
01:52:28.402 -> Water_Temp.:102.54, PWM:39
01:52:29.922 -> Water_Temp.:102.54, PWM:39
01:52:31.480 -> Water_Temp.:102.54, PWM:39
01:52:32.273 -> Water_Temp.:102.54, PWM:39
01:52:33.035 -> Water_Temp.:102.54, PWM:39
01:52:33.831 -> Water_Temp.:102.54, PWM:39
01:52:34.590 -> Water_Temp.:102.54, PWM:39
01:52:36.144 -> Water_Temp.:102.54, PWM:39
01:52:37.698 -> Water_Temp.:102.54, PWM:39
01:52:39.253 -> Water_Temp.:102.54, PWM:39
01:52:40.814 -> Water_Temp.:102.54, PWM:39
01:52:41.573 -> Water_Temp.:102.54, PWM:40
01:52:42.366 -> Water_Temp.:102.54, PWM:39
01:52:43.126 -> Water_Temp.:102.54, PWM:40
01:52:44.686 -> Water_Temp.:102.54, PWM:39
01:52:45.482 -> Water_Temp.:102.54, PWM:39
01:52:46.242 -> Water_Temp.:102.54, PWM:39
01:52:47.792 -> Water_Temp.:102.54, PWM:39
01:52:50.136 -> Water_Temp.:102.54, PWM:39
01:52:51.693 -> Water_Temp.:102.54, PWM:39
01:52:53.244 -> Water_Temp.:102.54, PWM:40
01:52:54.799 -> Water_Temp.:102.54, PWM:39
01:52:57.140 -> Water_Temp.:102.54, PWM:39
01:52:58.663 -> Water_Temp.:102.54, PWM:40
01:53:01.000 -> Water_Temp.:102.54, PWM:40
01:53:02.557 -> Water_Temp.:102.54, PWM:39
01:53:04.112 -> Water_Temp.:102.54, PWM:39
01:53:05.667 -> Water_Temp.:102.54, PWM:39
01:53:06.427 -> Water_Temp.:102.54, PWM:39
01:53:07.218 -> Water_Temp.:102.54, PWM:39
01:53:08.013 -> Water_Temp.:102.54, PWM:39
01:53:08.773 -> Water_Temp.:102.54, PWM:39
01:53:11.089 -> Water_Temp.:102.54, PWM:40
01:53:12.641 -> Water_Temp.:102.54, PWM:40
01:53:13.435 -> Water_Temp.:102.54, PWM:39
01:53:14.197 -> Water_Temp.:102.54, PWM:39
01:53:14.993 -> Water_Temp.:102.54, PWM:40
01:53:15.753 -> Water_Temp.:102.54, PWM:40
01:53:16.549 -> Water_Temp.:102.54, PWM:40
01:53:17.344 -> Water_Temp.:102.54, PWM:40
01:53:18.103 -> Water_Temp.:102.54, PWM:39
01:53:18.903 -> Water_Temp.:102.54, PWM:39
01:53:20.460 -> Water_Temp.:102.54, PWM:39
01:53:21.978 -> Water_Temp.:102.54, PWM:40
01:53:24.331 -> Water_Temp.:102.54, PWM:40
01:53:25.126 -> Water_Temp.:102.54, PWM:40
01:53:26.683 -> Water_Temp.:102.54, PWM:39
01:53:27.443 -> Water_Temp.:104.22, PWM:40
01:53:28.239 -> Water_Temp.:106.36, PWM:39
01:53:29.795 -> Water_Temp.:103.44, PWM:39
01:53:30.556 -> Water_Temp.:103.32, PWM:39
01:53:31.349 -> Water_Temp.:103.32, PWM:39
01:53:32.900 -> Water_Temp.:102.99, PWM:39
01:53:33.694 -> Water_Temp.:102.99, PWM:39
01:53:34.455 -> Water_Temp.:102.87, PWM:39
01:53:35.250 -> Water_Temp.:102.87, PWM:40
01:53:39.098 -> Water_Temp.:102.76, PWM:39
01:53:41.418 -> Water_Temp.:102.76, PWM:40
01:53:43.736 -> Water_Temp.:102.76, PWM:40
01:53:44.497 -> Water_Temp.:102.65, PWM:39
01:53:45.293 -> Water_Temp.:102.65, PWM:40
01:53:46.054 -> Water_Temp.:102.65, PWM:40
01:53:46.851 -> Water_Temp.:102.65, PWM:39
01:53:48.367 -> Water_Temp.:102.65, PWM:39
01:53:49.164 -> Water_Temp.:102.65, PWM:39
01:53:49.927 -> Water_Temp.:102.76, PWM:39
01:53:50.725 -> Water_Temp.:102.76, PWM:39
01:53:51.482 -> Water_Temp.:102.76, PWM:40
01:53:53.072 -> Water_Temp.:102.76, PWM:39
01:53:53.835 -> Water_Temp.:102.76, PWM:39
01:53:54.631 -> Water_Temp.:102.76, PWM:40
01:53:55.393 -> Water_Temp.:102.76, PWM:39
01:53:56.947 -> Water_Temp.:102.76, PWM:40
01:53:57.704 -> Water_Temp.:102.76, PWM:39
01:53:58.497 -> Water_Temp.:102.76, PWM:40
01:53:59.291 -> Water_Temp.:102.76, PWM:39
01:54:00.846 -> Water_Temp.:102.76, PWM:39
01:54:01.604 -> Water_Temp.:102.76, PWM:40
01:54:02.399 -> Water_Temp.:102.76, PWM:39
01:54:04.716 -> Water_Temp.:102.76, PWM:39
01:54:07.028 -> Water_Temp.:102.76, PWM:40
01:54:09.309 -> Water_Temp.:102.76, PWM:39
01:54:10.101 -> Water_Temp.:102.76, PWM:40
01:54:10.860 -> Water_Temp.:102.65, PWM:40
01:54:11.654 -> Water_Temp.:102.65, PWM:40
01:54:13.205 -> Water_Temp.:102.65, PWM:39
01:54:14.001 -> Water_Temp.:102.65, PWM:39
01:54:14.759 -> Water_Temp.:102.65, PWM:39
01:54:16.311 -> Water_Temp.:102.65, PWM:39
01:54:17.071 -> Water_Temp.:102.65, PWM:40
01:54:17.867 -> Water_Temp.:102.65, PWM:40
01:54:18.661 -> Water_Temp.:102.65, PWM:39
01:54:19.423 -> Water_Temp.:102.65, PWM:39
01:54:20.219 -> Water_Temp.:102.76, PWM:39
01:54:20.977 -> Water_Temp.:102.76, PWM:39
01:54:21.771 -> Water_Temp.:102.65, PWM:39
01:54:22.531 -> Water_Temp.:102.65, PWM:39
01:54:23.326 -> Water_Temp.:104.34, PWM:39
01:54:24.120 -> Water_Temp.:106.47, PWM:40
01:54:24.876 -> Water_Temp.:104.00, PWM:39
01:54:25.670 -> Water_Temp.:103.55, PWM:40
01:54:26.430 -> Water_Temp.:103.44, PWM:39
01:54:27.227 -> Water_Temp.:103.21, PWM:39
01:54:27.988 -> Water_Temp.:103.10, PWM:39
01:54:29.545 -> Water_Temp.:102.99, PWM:39
01:54:30.342 -> Water_Temp.:102.99, PWM:40
01:54:31.101 -> Water_Temp.:102.99, PWM:40
01:54:31.896 -> Water_Temp.:102.87, PWM:39
01:54:32.656 -> Water_Temp.:102.87, PWM:40
01:54:33.455 -> Water_Temp.:102.87, PWM:40
01:54:34.216 -> Water_Temp.:102.87, PWM:40
01:54:36.568 -> Water_Temp.:102.87, PWM:39
01:54:37.325 -> Water_Temp.:102.87, PWM:39
01:54:38.121 -> Water_Temp.:102.87, PWM:39
01:54:38.882 -> Water_Temp.:102.87, PWM:39
01:54:42.747 -> Water_Temp.:102.87, PWM:39
01:54:44.300 -> Water_Temp.:102.87, PWM:39
01:54:45.058 -> Water_Temp.:102.87, PWM:39
01:54:45.853 -> Water_Temp.:102.76, PWM:39
01:54:47.371 -> Water_Temp.:102.76, PWM:40
01:54:48.959 -> Water_Temp.:102.76, PWM:39
01:54:49.719 -> Water_Temp.:102.76, PWM:39
01:54:50.513 -> Water_Temp.:102.87, PWM:40
01:54:52.793 -> Water_Temp.:102.87, PWM:39
01:54:53.587 -> Water_Temp.:102.87, PWM:40
01:54:54.379 -> Water_Temp.:102.87, PWM:39
01:54:55.138 -> Water_Temp.:102.87, PWM:39
01:54:55.933 -> Water_Temp.:102.87, PWM:39
01:54:56.692 -> Water_Temp.:102.87, PWM:39
01:54:57.484 -> Water_Temp.:102.76, PWM:39
01:54:59.034 -> Water_Temp.:102.87, PWM:39
01:54:59.794 -> Water_Temp.:102.87, PWM:40
01:55:00.587 -> Water_Temp.:102.87, PWM:39
01:55:02.139 -> Water_Temp.:102.87, PWM:39
01:55:02.934 -> Water_Temp.:102.87, PWM:39
01:55:03.694 -> Water_Temp.:102.87, PWM:39
01:55:05.280 -> Water_Temp.:102.87, PWM:39
01:55:06.037 -> Water_Temp.:102.87, PWM:40
01:55:06.831 -> Water_Temp.:102.87, PWM:39
01:55:07.626 -> Water_Temp.:102.87, PWM:39
01:55:08.389 -> Water_Temp.:102.87, PWM:39
01:55:09.946 -> Water_Temp.:102.87, PWM:39
01:55:11.500 -> Water_Temp.:102.87, PWM:39
01:55:13.053 -> Water_Temp.:102.99, PWM:39
01:55:13.816 -> Water_Temp.:102.99, PWM:39
01:55:14.611 -> Water_Temp.:102.99, PWM:39
01:55:16.923 -> Water_Temp.:102.99, PWM:39
01:55:18.479 -> Water_Temp.:102.99, PWM:39
01:55:20.036 -> Water_Temp.:102.99, PWM:39
01:55:20.834 -> Water_Temp.:102.99, PWM:39
01:55:21.623 -> Water_Temp.:102.99, PWM:40
01:55:22.388 -> Water_Temp.:102.99, PWM:40
01:55:23.186 -> Water_Temp.:102.99, PWM:40
01:55:26.259 -> Water_Temp.:102.99, PWM:39
01:55:27.019 -> Water_Temp.:102.99, PWM:39
01:55:27.815 -> Water_Temp.:102.99, PWM:40
01:55:29.369 -> Water_Temp.:102.99, PWM:39
01:55:31.680 -> Water_Temp.:102.99, PWM:39
01:55:32.445 -> Water_Temp.:102.99, PWM:39
01:55:33.969 -> Water_Temp.:102.99, PWM:40
01:55:34.769 -> Water_Temp.:102.99, PWM:40
01:55:35.528 -> Water_Temp.:102.99, PWM:39
01:55:36.327 -> Water_Temp.:102.99, PWM:40
01:55:37.091 -> Water_Temp.:102.99, PWM:40
01:55:38.645 -> Water_Temp.:102.99, PWM:39
01:55:40.234 -> Water_Temp.:102.99, PWM:40
01:55:41.753 -> Water_Temp.:102.99, PWM:39
01:55:42.549 -> Water_Temp.:102.99, PWM:39
01:55:43.311 -> Water_Temp.:102.99, PWM:40
01:55:45.622 -> Water_Temp.:102.99, PWM:39
01:55:46.423 -> Water_Temp.:102.99, PWM:40
01:55:47.978 -> Water_Temp.:102.99, PWM:39
01:55:48.775 -> Water_Temp.:102.99, PWM:40
01:55:51.888 -> Water_Temp.:102.99, PWM:40
01:55:52.648 -> Water_Temp.:102.99, PWM:40
01:55:54.204 -> Water_Temp.:103.10, PWM:39
01:55:54.966 -> Water_Temp.:103.10, PWM:40
01:55:55.761 -> Water_Temp.:103.10, PWM:39
01:55:56.553 -> Water_Temp.:103.10, PWM:40
01:55:57.314 -> Water_Temp.:103.10, PWM:40
01:56:00.420 -> Water_Temp.:103.10, PWM:40
01:56:01.979 -> Water_Temp.:103.10, PWM:39
01:56:02.742 -> Water_Temp.:103.10, PWM:39
01:56:03.538 -> Water_Temp.:103.21, PWM:39
01:56:04.299 -> Water_Temp.:103.21, PWM:40
01:56:05.099 -> Water_Temp.:103.21, PWM:40
01:56:05.861 -> Water_Temp.:103.21, PWM:39
01:56:06.654 -> Water_Temp.:103.21, PWM:39
01:56:07.416 -> Water_Temp.:103.21, PWM:40
01:56:09.733 -> Water_Temp.:103.21, PWM:40
01:56:10.527 -> Water_Temp.:103.21, PWM:40
01:56:11.282 -> Water_Temp.:103.21, PWM:39
01:56:12.072 -> Water_Temp.:103.21, PWM:40
01:56:13.631 -> Water_Temp.:103.21, PWM:39
01:56:14.427 -> Water_Temp.:103.21, PWM:39
01:56:15.185 -> Water_Temp.:103.21, PWM:39
01:56:17.502 -> Water_Temp.:103.21, PWM:39
01:56:19.055 -> Water_Temp.:103.21, PWM:39
01:56:19.814 -> Water_Temp.:103.21, PWM:39
01:56:20.611 -> Water_Temp.:103.21, PWM:40
01:56:21.405 -> Water_Temp.:103.21, PWM:40
01:56:22.163 -> Water_Temp.:103.21, PWM:39
01:56:24.517 -> Water_Temp.:103.10, PWM:39
01:56:26.040 -> Water_Temp.:103.10, PWM:39
01:56:26.834 -> Water_Temp.:103.10, PWM:39
01:56:27.626 -> Water_Temp.:103.10, PWM:39
01:56:28.381 -> Water_Temp.:103.10, PWM:39
01:56:29.175 -> Water_Temp.:103.10, PWM:39
01:56:29.936 -> Water_Temp.:103.10, PWM:40
01:56:30.729 -> Water_Temp.:103.10, PWM:39
01:56:31.524 -> Water_Temp.:103.10, PWM:39
01:56:33.082 -> Water_Temp.:103.10, PWM:39
01:56:33.840 -> Water_Temp.:103.10, PWM:39
01:56:34.633 -> Water_Temp.:103.10, PWM:39
01:56:36.185 -> Water_Temp.:103.10, PWM:40
01:56:36.976 -> Water_Temp.:103.10, PWM:40
01:56:37.734 -> Water_Temp.:103.10, PWM:39
01:56:38.529 -> Water_Temp.:103.10, PWM:40
01:56:40.084 -> Water_Temp.:103.10, PWM:39
01:56:40.847 -> Water_Temp.:103.10, PWM:40
01:56:42.398 -> Water_Temp.:103.10, PWM:39
01:56:43.160 -> Water_Temp.:103.21, PWM:39
01:56:43.960 -> Water_Temp.:103.21, PWM:39
01:56:44.752 -> Water_Temp.:103.21, PWM:40
01:56:45.511 -> Water_Temp.:103.21, PWM:39
01:56:46.302 -> Water_Temp.:103.21, PWM:40
01:56:47.095 -> Water_Temp.:103.21, PWM:39
01:56:47.854 -> Water_Temp.:103.21, PWM:39
01:56:48.644 -> Water_Temp.:103.21, PWM:39
01:56:50.951 -> Water_Temp.:103.21, PWM:40
01:56:51.740 -> Water_Temp.:103.21, PWM:40
01:56:52.504 -> Water_Temp.:103.21, PWM:39
01:56:53.297 -> Water_Temp.:103.21, PWM:39
01:56:54.092 -> Water_Temp.:103.21, PWM:39
01:56:57.159 -> Water_Temp.:103.21, PWM:40
01:56:57.948 -> Water_Temp.:103.21, PWM:40
01:56:58.708 -> Water_Temp.:103.21, PWM:39
01:56:59.495 -> Water_Temp.:103.21, PWM:39
01:57:01.051 -> Water_Temp.:103.32, PWM:40
01:57:02.610 -> Water_Temp.:103.32, PWM:39
01:57:03.411 -> Water_Temp.:103.32, PWM:39
01:57:04.171 -> Water_Temp.:103.32, PWM:39
01:57:04.967 -> Water_Temp.:103.32, PWM:39
01:57:05.727 -> Water_Temp.:103.32, PWM:40
01:57:07.308 -> Water_Temp.:103.32, PWM:40
01:57:08.067 -> Water_Temp.:103.32, PWM:40
01:57:08.861 -> Water_Temp.:103.32, PWM:40
01:57:10.380 -> Water_Temp.:103.32, PWM:39
01:57:11.973 -> Water_Temp.:103.32, PWM:40
01:57:12.734 -> Water_Temp.:103.32, PWM:40
01:57:14.286 -> Water_Temp.:103.32, PWM:40
01:57:15.078 -> Water_Temp.:103.32, PWM:39
01:57:15.838 -> Water_Temp.:103.32, PWM:40
01:57:16.635 -> Water_Temp.:103.32, PWM:40
01:57:17.429 -> Water_Temp.:103.32, PWM:40
01:57:18.988 -> Water_Temp.:103.32, PWM:40
01:57:19.748 -> Water_Temp.:103.32, PWM:40
01:57:20.541 -> Water_Temp.:103.32, PWM:39
01:57:21.301 -> Water_Temp.:103.32, PWM:39
01:57:22.851 -> Water_Temp.:103.32, PWM:40
01:57:23.649 -> Water_Temp.:103.32, PWM:40
01:57:24.408 -> Water_Temp.:103.32, PWM:39
01:57:25.205 -> Water_Temp.:103.32, PWM:39
01:57:25.967 -> Water_Temp.:103.32, PWM:39
01:57:26.761 -> Water_Temp.:103.32, PWM:39
01:57:27.554 -> Water_Temp.:103.32, PWM:39
01:57:29.870 -> Water_Temp.:103.32, PWM:40
01:57:30.628 -> Water_Temp.:103.32, PWM:39
01:57:31.427 -> Water_Temp.:103.32, PWM:39
01:57:32.189 -> Water_Temp.:103.32, PWM:40
01:57:32.983 -> Water_Temp.:103.32, PWM:40
01:57:33.775 -> Water_Temp.:103.32, PWM:39
01:57:35.293 -> Water_Temp.:103.32, PWM:39
01:57:36.848 -> Water_Temp.:103.32, PWM:40
01:57:38.404 -> Water_Temp.:103.32, PWM:39
01:57:39.166 -> Water_Temp.:103.32, PWM:39
01:57:39.961 -> Water_Temp.:103.44, PWM:39
01:57:42.270 -> Water_Temp.:103.44, PWM:39
01:57:43.032 -> Water_Temp.:103.44, PWM:39
01:57:43.795 -> Water_Temp.:103.44, PWM:39
01:57:44.588 -> Water_Temp.:103.44, PWM:40
01:57:45.383 -> Water_Temp.:103.44, PWM:39
01:57:46.143 -> Water_Temp.:103.44, PWM:39
01:57:46.938 -> Water_Temp.:103.44, PWM:40
01:57:47.700 -> Water_Temp.:103.44, PWM:39
01:57:48.495 -> Water_Temp.:103.44, PWM:39
01:57:49.257 -> Water_Temp.:103.44, PWM:40
01:57:51.603 -> Water_Temp.:103.44, PWM:40
01:57:52.362 -> Water_Temp.:103.44, PWM:39
01:57:53.158 -> Water_Temp.:103.44, PWM:39
01:57:54.680 -> Water_Temp.:103.44, PWM:40
01:57:56.235 -> Water_Temp.:103.44, PWM:39
01:57:57.792 -> Water_Temp.:103.44, PWM:39
01:57:59.347 -> Water_Temp.:103.44, PWM:39
01:58:00.110 -> Water_Temp.:103.44, PWM:40
01:58:00.909 -> Water_Temp.:103.44, PWM:40
01:58:02.462 -> Water_Temp.:103.44, PWM:39
01:58:03.228 -> Water_Temp.:103.44, PWM:39
01:58:04.023 -> Water_Temp.:103.44, PWM:40
01:58:04.816 -> Water_Temp.:103.44, PWM:39
01:58:05.578 -> Water_Temp.:103.44, PWM:39
01:58:06.375 -> Water_Temp.:103.44, PWM:39
01:58:07.134 -> Water_Temp.:103.44, PWM:39
01:58:09.452 -> Water_Temp.:103.44, PWM:40
01:58:10.252 -> Water_Temp.:103.44, PWM:40
01:58:11.017 -> Water_Temp.:103.44, PWM:39
01:58:11.812 -> Water_Temp.:103.44, PWM:40
01:58:12.574 -> Water_Temp.:103.44, PWM:39
01:58:13.368 -> Water_Temp.:103.44, PWM:40
01:58:14.928 -> Water_Temp.:103.44, PWM:39
01:58:15.686 -> Water_Temp.:103.44, PWM:39
01:58:18.038 -> Water_Temp.:103.44, PWM:39
01:58:19.599 -> Water_Temp.:103.44, PWM:39
01:58:20.359 -> Water_Temp.:103.44, PWM:39
01:58:21.153 -> Water_Temp.:103.44, PWM:39
01:58:21.911 -> Water_Temp.:103.44, PWM:40
01:58:24.265 -> Water_Temp.:103.44, PWM:39
01:58:25.785 -> Water_Temp.:103.44, PWM:40
01:58:27.370 -> Water_Temp.:103.44, PWM:39
01:58:30.444 -> Water_Temp.:103.44, PWM:39
01:58:31.240 -> Water_Temp.:103.44, PWM:40
01:58:31.998 -> Water_Temp.:103.44, PWM:39
01:58:32.795 -> Water_Temp.:103.44, PWM:39
01:58:34.348 -> Water_Temp.:103.44, PWM:39
01:58:35.902 -> Water_Temp.:103.55, PWM:39
01:58:36.658 -> Water_Temp.:103.55, PWM:39
01:58:37.448 -> Water_Temp.:103.55, PWM:39
01:58:38.244 -> Water_Temp.:103.55, PWM:39
01:58:39.796 -> Water_Temp.:103.66, PWM:39
01:58:40.558 -> Water_Temp.:103.55, PWM:40
01:58:41.353 -> Water_Temp.:103.66, PWM:39
01:58:42.902 -> Water_Temp.:103.66, PWM:39
01:58:45.220 -> Water_Temp.:103.55, PWM:39
01:58:48.258 -> Water_Temp.:103.55, PWM:40
01:58:49.051 -> Water_Temp.:103.55, PWM:39
01:58:49.810 -> Water_Temp.:103.55, PWM:39
01:58:50.603 -> Water_Temp.:103.55, PWM:39
01:58:52.159 -> Water_Temp.:103.55, PWM:39
01:58:52.953 -> Water_Temp.:103.55, PWM:39
01:58:54.509 -> Water_Temp.:103.55, PWM:40
01:58:55.300 -> Water_Temp.:103.55, PWM:40
01:58:56.060 -> Water_Temp.:103.44, PWM:40
01:58:57.607 -> Water_Temp.:103.55, PWM:39
01:58:58.400 -> Water_Temp.:103.44, PWM:39
01:58:59.959 -> Water_Temp.:103.55, PWM:40
01:59:01.514 -> Water_Temp.:103.55, PWM:39
01:59:02.303 -> Water_Temp.:103.55, PWM:40
01:59:03.063 -> Water_Temp.:103.55, PWM:39
01:59:03.859 -> Water_Temp.:103.55, PWM:40
01:59:04.620 -> Water_Temp.:103.55, PWM:40
01:59:05.416 -> Water_Temp.:103.55, PWM:39
01:59:06.975 -> Water_Temp.:103.55, PWM:40
01:59:07.735 -> Water_Temp.:103.55, PWM:39
01:59:08.531 -> Water_Temp.:103.55, PWM:40
01:59:09.292 -> Water_Temp.:103.55, PWM:40
01:59:10.847 -> Water_Temp.:103.66, PWM:39
01:59:11.642 -> Water_Temp.:103.66, PWM:40
01:59:12.403 -> Water_Temp.:103.66, PWM:39
01:59:14.760 -> Water_Temp.:103.66, PWM:40
01:59:17.078 -> Water_Temp.:103.66, PWM:39
01:59:17.866 -> Water_Temp.:103.66, PWM:40
01:59:18.661 -> Water_Temp.:103.66, PWM:39
01:59:19.421 -> Water_Temp.:103.66, PWM:39
01:59:20.977 -> Water_Temp.:103.66, PWM:40
01:59:22.528 -> Water_Temp.:103.66, PWM:39
01:59:23.326 -> Water_Temp.:103.66, PWM:39
01:59:24.087 -> Water_Temp.:103.77, PWM:39
01:59:25.643 -> Water_Temp.:103.77, PWM:39
01:59:26.443 -> Water_Temp.:103.77, PWM:39
01:59:27.203 -> Water_Temp.:103.77, PWM:39
01:59:28.001 -> Water_Temp.:103.77, PWM:39
01:59:28.766 -> Water_Temp.:103.77, PWM:39
01:59:30.322 -> Water_Temp.:103.77, PWM:39
01:59:31.872 -> Water_Temp.:103.77, PWM:39
01:59:32.665 -> Water_Temp.:103.77, PWM:39
01:59:33.426 -> Water_Temp.:103.77, PWM:40
01:59:34.219 -> Water_Temp.:103.77, PWM:39
01:59:34.982 -> Water_Temp.:103.77, PWM:39
01:59:35.778 -> Water_Temp.:103.77, PWM:39
01:59:36.570 -> Water_Temp.:103.77, PWM:39
01:59:38.091 -> Water_Temp.:103.77, PWM:39
01:59:38.890 -> Water_Temp.:103.77, PWM:40
01:59:40.445 -> Water_Temp.:103.77, PWM:39
01:59:41.207 -> Water_Temp.:103.77, PWM:39
01:59:42.000 -> Water_Temp.:103.77, PWM:40
01:59:42.792 -> Water_Temp.:103.77, PWM:39
01:59:43.553 -> Water_Temp.:103.77, PWM:40
01:59:45.106 -> Water_Temp.:103.77, PWM:40
01:59:46.661 -> Water_Temp.:103.77, PWM:39
01:59:48.245 -> Water_Temp.:103.77, PWM:39
01:59:49.001 -> Water_Temp.:103.77, PWM:39
01:59:50.555 -> Water_Temp.:103.77, PWM:40
01:59:52.105 -> Water_Temp.:103.77, PWM:39
01:59:52.900 -> Water_Temp.:103.77, PWM:39
01:59:53.663 -> Water_Temp.:103.77, PWM:40
01:59:54.457 -> Water_Temp.:103.89, PWM:39
01:59:55.213 -> Water_Temp.:103.89, PWM:40
01:59:56.010 -> Water_Temp.:103.89, PWM:40
01:59:57.569 -> Water_Temp.:103.89, PWM:39
01:59:59.880 -> Water_Temp.:103.89, PWM:40
02:00:00.639 -> Water_Temp.:103.89, PWM:39
02:00:01.436 -> Water_Temp.:103.89, PWM:40
02:00:02.196 -> Water_Temp.:103.89, PWM:39
02:00:03.753 -> Water_Temp.:103.89, PWM:39
02:00:04.547 -> Water_Temp.:103.89, PWM:39
02:00:05.342 -> Water_Temp.:103.89, PWM:39
02:00:06.894 -> Water_Temp.:103.89, PWM:39
02:00:07.654 -> Water_Temp.:103.89, PWM:39
02:00:08.448 -> Water_Temp.:103.89, PWM:40
02:00:09.210 -> Water_Temp.:103.89, PWM:39
02:00:10.006 -> Water_Temp.:103.89, PWM:40
02:00:10.800 -> Water_Temp.:103.89, PWM:40
02:00:12.350 -> Water_Temp.:103.89, PWM:40
02:00:13.113 -> Water_Temp.:103.77, PWM:39
02:00:13.911 -> Water_Temp.:103.77, PWM:40
02:00:16.222 -> Water_Temp.:103.77, PWM:39
02:00:17.015 -> Water_Temp.:103.77, PWM:39
02:00:17.771 -> Water_Temp.:103.89, PWM:39
02:00:19.328 -> Water_Temp.:103.89, PWM:39
02:00:21.675 -> Water_Temp.:103.89, PWM:40
02:00:24.788 -> Water_Temp.:103.77, PWM:39
02:00:25.583 -> Water_Temp.:103.77, PWM:39
02:00:26.375 -> Water_Temp.:103.77, PWM:39
02:00:27.139 -> Water_Temp.:103.77, PWM:39
02:00:27.934 -> Water_Temp.:103.77, PWM:40
02:00:28.696 -> Water_Temp.:103.77, PWM:39
02:00:32.564 -> Water_Temp.:103.66, PWM:39
02:00:33.325 -> Water_Temp.:103.66, PWM:40
02:00:34.881 -> Water_Temp.:103.66, PWM:39
02:00:35.674 -> Water_Temp.:103.66, PWM:39
02:00:36.433 -> Water_Temp.:103.66, PWM:39
02:00:37.226 -> Water_Temp.:103.66, PWM:39
02:00:37.986 -> Water_Temp.:103.66, PWM:40
02:00:39.569 -> Water_Temp.:103.66, PWM:39
02:00:41.880 -> Water_Temp.:103.66, PWM:39
02:00:42.676 -> Water_Temp.:103.66, PWM:39
02:00:43.440 -> Water_Temp.:103.66, PWM:39
02:00:46.517 -> Water_Temp.:103.77, PWM:40
02:00:47.314 -> Water_Temp.:103.77, PWM:39
02:00:48.071 -> Water_Temp.:103.77, PWM:39
02:00:48.865 -> Water_Temp.:103.77, PWM:39
02:00:49.661 -> Water_Temp.:103.89, PWM:40
02:00:50.421 -> Water_Temp.:103.89, PWM:39
02:00:51.220 -> Water_Temp.:103.89, PWM:39
02:00:52.745 -> Water_Temp.:103.89, PWM:40
02:00:53.543 -> Water_Temp.:103.89, PWM:39
02:00:55.069 -> Water_Temp.:104.00, PWM:39
02:00:55.865 -> Water_Temp.:104.00, PWM:39
02:00:56.658 -> Water_Temp.:104.00, PWM:40
02:00:57.416 -> Water_Temp.:104.00, PWM:39
02:00:58.209 -> Water_Temp.:104.00, PWM:39
02:01:02.050 -> Water_Temp.:104.00, PWM:40
02:01:02.846 -> Water_Temp.:104.00, PWM:39
02:01:03.610 -> Water_Temp.:104.00, PWM:40
02:01:04.407 -> Water_Temp.:104.00, PWM:40
02:01:05.167 -> Water_Temp.:104.00, PWM:40
02:01:05.965 -> Water_Temp.:104.00, PWM:39
02:01:06.725 -> Water_Temp.:104.00, PWM:40
02:01:09.036 -> Water_Temp.:103.89, PWM:40
02:01:09.824 -> Water_Temp.:103.89, PWM:39
02:01:11.377 -> Water_Temp.:103.89, PWM:40
02:01:12.174 -> Water_Temp.:104.00, PWM:39
02:01:13.726 -> Water_Temp.:104.00, PWM:40
02:01:15.280 -> Water_Temp.:103.89, PWM:39
02:01:16.041 -> Water_Temp.:104.00, PWM:39
02:01:16.836 -> Water_Temp.:104.00, PWM:39
02:01:18.356 -> Water_Temp.:104.00, PWM:39
02:01:19.150 -> Water_Temp.:104.00, PWM:39
02:01:19.911 -> Water_Temp.:104.00, PWM:40
02:01:21.464 -> Water_Temp.:104.00, PWM:39
02:01:22.255 -> Water_Temp.:104.00, PWM:39
02:01:23.815 -> Water_Temp.:104.00, PWM:39
02:01:24.573 -> Water_Temp.:104.00, PWM:40
02:01:26.129 -> Water_Temp.:103.89, PWM:40
02:01:26.889 -> Water_Temp.:103.89, PWM:39
02:01:27.685 -> Water_Temp.:103.89, PWM:40
02:01:28.444 -> Water_Temp.:103.89, PWM:39
02:01:29.242 -> Water_Temp.:103.89, PWM:39
02:01:30.004 -> Water_Temp.:103.89, PWM:39
02:01:31.564 -> Water_Temp.:103.89, PWM:39
02:01:33.122 -> Water_Temp.:103.89, PWM:40
02:01:33.914 -> Water_Temp.:103.89, PWM:39
02:01:34.679 -> Water_Temp.:103.89, PWM:39
02:01:35.441 -> Water_Temp.:104.00, PWM:40
02:01:36.236 -> Water_Temp.:104.00, PWM:39
02:01:37.031 -> Water_Temp.:104.00, PWM:39
02:01:39.345 -> Water_Temp.:104.00, PWM:39
02:01:41.664 -> Water_Temp.:104.00, PWM:39
02:01:43.220 -> Water_Temp.:104.00, PWM:39
02:01:44.016 -> Water_Temp.:104.00, PWM:40
02:01:44.777 -> Water_Temp.:104.00, PWM:39
02:01:45.569 -> Water_Temp.:104.00, PWM:39
02:01:46.362 -> Water_Temp.:104.00, PWM:39
02:01:47.119 -> Water_Temp.:104.00, PWM:39
02:01:48.674 -> Water_Temp.:104.00, PWM:39
02:01:49.469 -> Water_Temp.:104.00, PWM:39
02:01:50.232 -> Water_Temp.:104.00, PWM:39
02:01:51.027 -> Water_Temp.:104.00, PWM:39
02:01:52.579 -> Water_Temp.:104.00, PWM:39
02:01:54.101 -> Water_Temp.:104.00, PWM:39
02:01:54.895 -> Water_Temp.:104.00, PWM:40
02:01:55.693 -> Water_Temp.:104.00, PWM:40
02:01:56.454 -> Water_Temp.:104.00, PWM:39
02:01:58.007 -> Water_Temp.:104.00, PWM:40
02:02:00.324 -> Water_Temp.:104.11, PWM:39
02:02:01.116 -> Water_Temp.:104.11, PWM:39
02:02:01.910 -> Water_Temp.:104.11, PWM:40
02:02:02.673 -> Water_Temp.:104.11, PWM:40
02:02:04.237 -> Water_Temp.:104.11, PWM:39
02:02:05.032 -> Water_Temp.:104.11, PWM:40
02:02:05.790 -> Water_Temp.:104.11, PWM:39
02:02:06.585 -> Water_Temp.:104.11, PWM:39
02:02:08.110 -> Water_Temp.:104.11, PWM:39
02:02:09.664 -> Water_Temp.:104.11, PWM:40
02:02:10.459 -> Water_Temp.:104.11, PWM:40
02:02:12.018 -> Water_Temp.:104.11, PWM:40
02:02:12.778 -> Water_Temp.:104.00, PWM:39
02:02:13.573 -> Water_Temp.:104.00, PWM:39
02:02:15.914 -> Water_Temp.:104.00, PWM:39
02:02:17.465 -> Water_Temp.:104.00, PWM:40
02:02:18.226 -> Water_Temp.:104.00, PWM:40
02:02:19.782 -> Water_Temp.:104.00, PWM:40
02:02:20.575 -> Water_Temp.:104.00, PWM:40
02:02:21.334 -> Water_Temp.:104.00, PWM:39
02:02:22.133 -> Water_Temp.:104.00, PWM:40
02:02:22.898 -> Water_Temp.:104.00, PWM:40
02:02:23.693 -> Water_Temp.:104.00, PWM:39
02:02:24.452 -> Water_Temp.:104.00, PWM:40
02:02:26.007 -> Water_Temp.:104.00, PWM:40
02:02:26.800 -> Water_Temp.:104.00, PWM:39
02:02:27.558 -> Water_Temp.:104.00, PWM:39
02:02:28.356 -> Water_Temp.:104.00, PWM:40
02:02:31.397 -> Water_Temp.:104.00, PWM:40
02:02:32.190 -> Water_Temp.:104.00, PWM:39
02:02:32.984 -> Water_Temp.:104.00, PWM:39
02:02:33.746 -> Water_Temp.:104.00, PWM:39
02:02:34.538 -> Water_Temp.:104.00, PWM:39
02:02:36.854 -> Water_Temp.:104.00, PWM:40
02:02:37.646 -> Water_Temp.:104.11, PWM:40
02:02:38.436 -> Water_Temp.:104.11, PWM:39
02:02:39.195 -> Water_Temp.:104.11, PWM:39
02:02:39.989 -> Water_Temp.:104.11, PWM:39
02:02:40.748 -> Water_Temp.:104.11, PWM:39
02:02:43.092 -> Water_Temp.:104.11, PWM:39
02:02:43.881 -> Water_Temp.:104.11, PWM:39
02:02:44.639 -> Water_Temp.:104.11, PWM:39
02:02:45.434 -> Water_Temp.:104.11, PWM:39
02:02:46.226 -> Water_Temp.:104.11, PWM:39
02:02:48.539 -> Water_Temp.:104.11, PWM:39
02:02:49.334 -> Water_Temp.:104.22, PWM:40
02:02:50.093 -> Water_Temp.:104.22, PWM:39
02:02:50.887 -> Water_Temp.:104.22, PWM:39
02:02:51.678 -> Water_Temp.:104.22, PWM:39
02:02:52.438 -> Water_Temp.:104.22, PWM:39
02:02:53.996 -> Water_Temp.:104.22, PWM:40
02:02:54.791 -> Water_Temp.:104.22, PWM:39
02:02:55.551 -> Water_Temp.:104.22, PWM:40
02:02:57.106 -> Water_Temp.:104.34, PWM:39
02:02:57.867 -> Water_Temp.:104.34, PWM:40
02:02:58.659 -> Water_Temp.:104.34, PWM:39
02:02:59.452 -> Water_Temp.:104.34, PWM:39
02:03:00.210 -> Water_Temp.:104.34, PWM:40
02:03:01.765 -> Water_Temp.:104.34, PWM:40
02:03:02.560 -> Water_Temp.:104.34, PWM:39
02:03:04.911 -> Water_Temp.:104.22, PWM:40
02:03:05.670 -> Water_Temp.:105.91, PWM:40
02:03:07.226 -> Water_Temp.:105.57, PWM:39
02:03:08.024 -> Water_Temp.:105.12, PWM:39
02:03:09.583 -> Water_Temp.:104.79, PWM:40
02:03:11.141 -> Water_Temp.:104.56, PWM:40
02:03:11.898 -> Water_Temp.:104.56, PWM:40
02:03:15.702 -> Water_Temp.:104.34, PWM:39
02:03:16.496 -> Water_Temp.:104.34, PWM:40
02:03:17.254 -> Water_Temp.:104.34, PWM:39
02:03:18.049 -> Water_Temp.:104.34, PWM:40
02:03:18.843 -> Water_Temp.:104.34, PWM:39
02:03:19.602 -> Water_Temp.:104.22, PWM:39
02:03:20.397 -> Water_Temp.:104.22, PWM:39
02:03:21.158 -> Water_Temp.:104.22, PWM:39
02:03:22.707 -> Water_Temp.:104.22, PWM:39
02:03:23.502 -> Water_Temp.:104.22, PWM:39
02:03:25.054 -> Water_Temp.:104.22, PWM:40
02:03:25.847 -> Water_Temp.:104.22, PWM:39
02:03:27.399 -> Water_Temp.:104.22, PWM:39
02:03:28.189 -> Water_Temp.:104.22, PWM:39
02:03:28.948 -> Water_Temp.:104.22, PWM:39
02:03:29.747 -> Water_Temp.:104.22, PWM:40
02:03:30.510 -> Water_Temp.:104.22, PWM:40
02:03:34.348 -> Water_Temp.:104.22, PWM:39
02:03:35.141 -> Water_Temp.:104.22, PWM:40
02:03:35.937 -> Water_Temp.:104.22, PWM:39
02:03:36.691 -> Water_Temp.:104.22, PWM:40
02:03:38.240 -> Water_Temp.:104.22, PWM:40
02:03:39.032 -> Water_Temp.:104.22, PWM:40
02:03:41.317 -> Water_Temp.:104.22, PWM:40
02:03:42.108 -> Water_Temp.:104.22, PWM:40
02:03:42.903 -> Water_Temp.:104.22, PWM:39
02:03:45.220 -> Water_Temp.:104.22, PWM:39
02:03:46.010 -> Water_Temp.:104.34, PWM:39
02:03:46.768 -> Water_Temp.:104.22, PWM:40
02:03:49.115 -> Water_Temp.:104.22, PWM:40
02:03:50.698 -> Water_Temp.:104.22, PWM:39
02:03:51.458 -> Water_Temp.:104.22, PWM:39
02:03:52.252 -> Water_Temp.:104.22, PWM:39
02:03:53.011 -> Water_Temp.:104.22, PWM:39
02:03:53.809 -> Water_Temp.:104.22, PWM:39
02:03:58.411 -> Water_Temp.:104.22, PWM:40
02:03:59.203 -> Water_Temp.:104.22, PWM:40
02:03:59.961 -> Water_Temp.:104.22, PWM:39
02:04:00.755 -> Water_Temp.:104.22, PWM:39
02:04:01.514 -> Water_Temp.:104.22, PWM:40
02:04:02.307 -> Water_Temp.:104.22, PWM:40
02:04:03.103 -> Water_Temp.:104.22, PWM:39
02:04:03.867 -> Water_Temp.:104.22, PWM:40
02:04:04.663 -> Water_Temp.:104.34, PWM:39
02:04:05.424 -> Water_Temp.:104.34, PWM:40
02:04:07.746 -> Water_Temp.:104.34, PWM:40
02:04:10.853 -> Water_Temp.:104.34, PWM:39
02:04:11.609 -> Water_Temp.:104.22, PWM:39
02:04:12.408 -> Water_Temp.:104.22, PWM:39
02:04:13.168 -> Water_Temp.:104.22, PWM:40
02:04:15.488 -> Water_Temp.:104.22, PWM:39
02:04:16.289 -> Water_Temp.:104.22, PWM:39
02:04:17.052 -> Water_Temp.:104.22, PWM:39
02:04:17.844 -> Water_Temp.:104.22, PWM:40
02:04:19.369 -> Water_Temp.:104.22, PWM:40
02:04:20.133 -> Water_Temp.:104.22, PWM:40
02:04:20.928 -> Water_Temp.:104.22, PWM:40
02:04:21.721 -> Water_Temp.:104.22, PWM:39
02:04:22.486 -> Water_Temp.:104.22, PWM:39
02:04:24.811 -> Water_Temp.:104.22, PWM:40
02:04:25.605 -> Water_Temp.:104.22, PWM:39
02:04:27.924 -> Water_Temp.:104.22, PWM:39
02:04:28.721 -> Water_Temp.:104.22, PWM:39
02:04:29.483 -> Water_Temp.:104.34, PWM:39
02:04:30.279 -> Water_Temp.:104.22, PWM:39
02:04:31.839 -> Water_Temp.:104.34, PWM:39
02:04:32.597 -> Water_Temp.:104.34, PWM:40
02:04:33.395 -> Water_Temp.:104.34, PWM:39
02:04:34.187 -> Water_Temp.:104.34, PWM:40
02:04:35.705 -> Water_Temp.:104.34, PWM:39
02:04:36.494 -> Water_Temp.:104.34, PWM:40
02:04:37.286 -> Water_Temp.:104.34, PWM:39
02:04:38.044 -> Water_Temp.:104.34, PWM:40
02:04:41.879 -> Water_Temp.:104.45, PWM:39
02:04:42.673 -> Water_Temp.:104.56, PWM:40
02:04:43.434 -> Water_Temp.:104.56, PWM:39
02:04:44.952 -> Water_Temp.:104.56, PWM:40
02:04:45.748 -> Water_Temp.:104.56, PWM:40
02:04:46.541 -> Water_Temp.:104.56, PWM:40
02:04:50.378 -> Water_Temp.:104.45, PWM:40
02:04:51.900 -> Water_Temp.:104.45, PWM:40
02:04:52.698 -> Water_Temp.:104.45, PWM:39
02:04:53.461 -> Water_Temp.:104.45, PWM:39
02:04:54.257 -> Water_Temp.:104.45, PWM:39
02:04:55.805 -> Water_Temp.:104.34, PWM:39
02:04:56.601 -> Water_Temp.:104.34, PWM:39
02:04:58.880 -> Water_Temp.:104.34, PWM:39
02:05:00.437 -> Water_Temp.:104.34, PWM:39
02:05:01.202 -> Water_Temp.:104.34, PWM:39
02:05:01.995 -> Water_Temp.:104.34, PWM:39
02:05:04.313 -> Water_Temp.:104.34, PWM:39
02:05:05.106 -> Water_Temp.:104.34, PWM:40
02:05:05.866 -> Water_Temp.:104.34, PWM:39
02:05:06.663 -> Water_Temp.:104.34, PWM:40
02:05:08.183 -> Water_Temp.:104.34, PWM:39
02:05:08.943 -> Water_Temp.:104.34, PWM:39
02:05:11.260 -> Water_Temp.:104.34, PWM:39
02:05:13.574 -> Water_Temp.:104.34, PWM:39
02:05:14.335 -> Water_Temp.:104.45, PWM:39
02:05:15.130 -> Water_Temp.:104.45, PWM:40
02:05:15.893 -> Water_Temp.:104.45, PWM:39
02:05:16.684 -> Water_Temp.:104.45, PWM:39
02:05:17.481 -> Water_Temp.:104.45, PWM:39
02:05:18.244 -> Water_Temp.:104.45, PWM:39
02:05:19.040 -> Water_Temp.:104.45, PWM:39
02:05:19.797 -> Water_Temp.:104.45, PWM:40
02:05:20.595 -> Water_Temp.:104.45, PWM:39
02:05:21.359 -> Water_Temp.:104.45, PWM:39
02:05:22.154 -> Water_Temp.:104.45, PWM:39
02:05:22.915 -> Water_Temp.:104.45, PWM:39
02:05:26.026 -> Water_Temp.:104.45, PWM:39
02:05:26.819 -> Water_Temp.:104.45, PWM:40
02:05:27.584 -> Water_Temp.:104.45, PWM:39
02:05:28.379 -> Water_Temp.:106.14, PWM:39
02:05:29.143 -> Water_Temp.:108.16, PWM:39
02:05:29.940 -> Water_Temp.:105.80, PWM:39
02:05:30.702 -> Water_Temp.:105.35, PWM:39
02:05:32.259 -> Water_Temp.:105.01, PWM:39
02:05:33.818 -> Water_Temp.:104.79, PWM:40
02:05:35.381 -> Water_Temp.:104.79, PWM:39
02:05:36.140 -> Water_Temp.:104.67, PWM:39
02:05:37.692 -> Water_Temp.:104.67, PWM:39
02:05:38.453 -> Water_Temp.:104.56, PWM:40
02:05:39.248 -> Water_Temp.:104.56, PWM:40
02:05:40.806 -> Water_Temp.:104.56, PWM:40
02:05:41.564 -> Water_Temp.:104.56, PWM:39
02:05:43.117 -> Water_Temp.:104.56, PWM:39
02:05:43.908 -> Water_Temp.:104.56, PWM:39
02:05:44.669 -> Water_Temp.:104.56, PWM:39
02:05:45.466 -> Water_Temp.:104.56, PWM:40
02:05:47.023 -> Water_Temp.:104.56, PWM:39
02:05:47.819 -> Water_Temp.:104.56, PWM:40
02:05:49.374 -> Water_Temp.:104.56, PWM:40
02:05:50.132 -> Water_Temp.:104.56, PWM:40
02:05:50.929 -> Water_Temp.:104.67, PWM:39
02:05:51.688 -> Water_Temp.:104.67, PWM:39
02:05:52.482 -> Water_Temp.:104.67, PWM:39
02:05:53.273 -> Water_Temp.:104.67, PWM:39
02:05:54.037 -> Water_Temp.:104.56, PWM:39
02:05:57.136 -> Water_Temp.:104.56, PWM:39
02:05:58.661 -> Water_Temp.:104.56, PWM:39
02:05:59.459 -> Water_Temp.:104.56, PWM:39
02:06:00.224 -> Water_Temp.:104.56, PWM:39
02:06:01.016 -> Water_Temp.:104.56, PWM:39
02:06:02.576 -> Water_Temp.:104.45, PWM:40
02:06:04.897 -> Water_Temp.:104.45, PWM:39
02:06:05.697 -> Water_Temp.:104.45, PWM:39
02:06:07.255 -> Water_Temp.:104.45, PWM:40
02:06:08.773 -> Water_Temp.:104.45, PWM:39
02:06:10.328 -> Water_Temp.:104.45, PWM:40
02:06:11.127 -> Water_Temp.:104.45, PWM:39
02:06:11.882 -> Water_Temp.:104.45, PWM:39
02:06:13.435 -> Water_Temp.:104.45, PWM:39
02:06:14.196 -> Water_Temp.:104.45, PWM:39
02:06:14.993 -> Water_Temp.:104.45, PWM:39
02:06:15.787 -> Water_Temp.:104.45, PWM:40
02:06:17.339 -> Water_Temp.:104.45, PWM:39
02:06:18.891 -> Water_Temp.:104.56, PWM:40
02:06:19.683 -> Water_Temp.:104.45, PWM:39
02:06:20.448 -> Water_Temp.:104.45, PWM:40
02:06:21.967 -> Water_Temp.:104.56, PWM:40
02:06:24.290 -> Water_Temp.:104.56, PWM:40
02:06:27.398 -> Water_Temp.:108.27, PWM:40
02:06:28.986 -> Water_Temp.:105.46, PWM:39
02:06:30.534 -> Water_Temp.:105.12, PWM:40
02:06:32.853 -> Water_Temp.:104.90, PWM:39
02:06:34.409 -> Water_Temp.:104.79, PWM:39
02:06:35.165 -> Water_Temp.:104.79, PWM:39
02:06:35.960 -> Water_Temp.:104.79, PWM:39
02:06:36.720 -> Water_Temp.:104.79, PWM:40
02:06:39.029 -> Water_Temp.:104.79, PWM:39
02:06:39.821 -> Water_Temp.:104.79, PWM:39
02:06:40.612 -> Water_Temp.:104.79, PWM:40
02:06:41.372 -> Water_Temp.:104.79, PWM:40
02:06:42.166 -> Water_Temp.:104.79, PWM:39
02:06:42.960 -> Water_Temp.:104.79, PWM:39
02:06:44.516 -> Water_Temp.:106.36, PWM:39
02:06:45.274 -> Water_Temp.:108.50, PWM:39
02:06:46.068 -> Water_Temp.:106.02, PWM:39
02:06:46.828 -> Water_Temp.:105.57, PWM:39
02:06:48.384 -> Water_Temp.:105.24, PWM:39
02:06:49.148 -> Water_Temp.:105.12, PWM:39
02:06:49.942 -> Water_Temp.:105.01, PWM:40
02:06:51.498 -> Water_Temp.:104.90, PWM:39
02:06:53.019 -> Water_Temp.:104.79, PWM:39
02:06:55.372 -> Water_Temp.:104.67, PWM:39
02:06:56.132 -> Water_Temp.:104.67, PWM:39
02:06:56.926 -> Water_Temp.:104.67, PWM:40
02:06:57.685 -> Water_Temp.:104.56, PWM:40
02:06:58.478 -> Water_Temp.:104.56, PWM:39
02:06:59.275 -> Water_Temp.:104.56, PWM:40
02:07:00.036 -> Water_Temp.:104.56, PWM:39
02:07:02.383 -> Water_Temp.:104.56, PWM:40
02:07:03.143 -> Water_Temp.:104.56, PWM:40
02:07:03.938 -> Water_Temp.:104.56, PWM:39
02:07:05.492 -> Water_Temp.:104.45, PWM:39
02:07:06.284 -> Water_Temp.:104.45, PWM:39
02:07:07.043 -> Water_Temp.:104.45, PWM:40
02:07:07.834 -> Water_Temp.:104.56, PWM:39
02:07:08.628 -> Water_Temp.:104.45, PWM:39
02:07:10.154 -> Water_Temp.:104.56, PWM:39
02:07:10.915 -> Water_Temp.:104.56, PWM:39
02:07:11.707 -> Water_Temp.:104.56, PWM:39
02:07:13.233 -> Water_Temp.:104.56, PWM:40
02:07:14.030 -> Water_Temp.:104.56, PWM:40
02:07:14.792 -> Water_Temp.:104.56, PWM:39
02:07:15.588 -> Water_Temp.:104.56, PWM:39
02:07:16.347 -> Water_Temp.:104.56, PWM:39
02:07:17.140 -> Water_Temp.:104.56, PWM:39
02:07:18.662 -> Water_Temp.:104.56, PWM:39
02:07:19.456 -> Water_Temp.:104.56, PWM:39
02:07:20.217 -> Water_Temp.:104.56, PWM:40
02:07:22.534 -> Water_Temp.:104.56, PWM:39
02:07:23.331 -> Water_Temp.:104.56, PWM:40
02:07:24.089 -> Water_Temp.:104.67, PWM:40
02:07:24.882 -> Water_Temp.:104.67, PWM:39
02:07:25.644 -> Water_Temp.:104.67, PWM:39
02:07:26.440 -> Water_Temp.:104.67, PWM:39
02:07:30.277 -> Water_Temp.:104.79, PWM:40
02:07:31.072 -> Water_Temp.:104.79, PWM:40
02:07:33.392 -> Water_Temp.:104.79, PWM:40
02:07:34.947 -> Water_Temp.:104.90, PWM:40
02:07:36.505 -> Water_Temp.:104.90, PWM:40
02:07:37.301 -> Water_Temp.:105.01, PWM:40
02:07:38.061 -> Water_Temp.:105.01, PWM:39
02:07:41.143 -> Water_Temp.:105.01, PWM:40
02:07:41.938 -> Water_Temp.:104.90, PWM:39
02:07:43.460 -> Water_Temp.:104.90, PWM:39
02:07:44.253 -> Water_Temp.:104.90, PWM:39
02:07:45.774 -> Water_Temp.:104.90, PWM:40
02:07:47.329 -> Water_Temp.:104.90, PWM:39
02:07:48.127 -> Water_Temp.:104.90, PWM:40
02:07:48.889 -> Water_Temp.:104.90, PWM:40
02:07:50.449 -> Water_Temp.:104.90, PWM:39
02:07:52.003 -> Water_Temp.:104.79, PWM:40
02:07:52.799 -> Water_Temp.:104.79, PWM:39
02:07:53.561 -> Water_Temp.:104.79, PWM:40
02:07:54.357 -> Water_Temp.:104.79, PWM:39
02:07:55.113 -> Water_Temp.:104.79, PWM:39
02:07:55.909 -> Water_Temp.:104.79, PWM:39
02:07:59.747 -> Water_Temp.:104.79, PWM:40
02:08:00.507 -> Water_Temp.:104.79, PWM:39
02:08:01.307 -> Water_Temp.:104.79, PWM:40
02:08:03.620 -> Water_Temp.:104.79, PWM:39
02:08:04.416 -> Water_Temp.:104.79, PWM:40
02:08:05.174 -> Water_Temp.:104.79, PWM:40
02:08:07.493 -> Water_Temp.:104.79, PWM:39
02:08:08.284 -> Water_Temp.:104.79, PWM:39
02:08:09.077 -> Water_Temp.:104.79, PWM:40
02:08:09.839 -> Water_Temp.:104.79, PWM:39
02:08:12.950 -> Water_Temp.:104.90, PWM:40
02:08:13.747 -> Water_Temp.:104.90, PWM:39
02:08:16.065 -> Water_Temp.:104.90, PWM:40
02:08:16.860 -> Water_Temp.:104.90, PWM:39
02:08:17.621 -> Water_Temp.:104.90, PWM:39
02:08:18.410 -> Water_Temp.:104.90, PWM:39
02:08:20.764 -> Water_Temp.:105.01, PWM:40
02:08:21.523 -> Water_Temp.:105.01, PWM:39
02:08:22.320 -> Water_Temp.:105.01, PWM:39
02:08:24.633 -> Water_Temp.:105.01, PWM:39
02:08:25.394 -> Water_Temp.:105.01, PWM:40
02:08:26.186 -> Water_Temp.:105.01, PWM:39
02:08:26.949 -> Water_Temp.:105.01, PWM:40
02:08:28.501 -> Water_Temp.:104.90, PWM:39
02:08:29.262 -> Water_Temp.:105.01, PWM:40
02:08:30.058 -> Water_Temp.:104.90, PWM:39
02:08:30.819 -> Water_Temp.:104.90, PWM:39
02:08:31.613 -> Water_Temp.:104.90, PWM:40
02:08:32.408 -> Water_Temp.:105.01, PWM:39
02:08:33.163 -> Water_Temp.:105.01, PWM:39
02:08:34.727 -> Water_Temp.:105.01, PWM:39
02:08:35.517 -> Water_Temp.:104.90, PWM:40
02:08:36.281 -> Water_Temp.:104.90, PWM:39
02:08:37.073 -> Water_Temp.:104.90, PWM:39
02:08:39.358 -> Water_Temp.:104.90, PWM:40
02:08:40.945 -> Water_Temp.:105.01, PWM:39
02:08:42.501 -> Water_Temp.:104.90, PWM:40
02:08:43.261 -> Water_Temp.:104.90, PWM:39
02:08:44.057 -> Water_Temp.:104.90, PWM:39
02:08:45.613 -> Water_Temp.:104.90, PWM:40
02:08:46.376 -> Water_Temp.:104.90, PWM:40
02:08:47.173 -> Water_Temp.:104.90, PWM:39
02:08:47.934 -> Water_Temp.:104.90, PWM:40
02:08:48.732 -> Water_Temp.:104.90, PWM:40
02:08:49.491 -> Water_Temp.:104.90, PWM:39
02:08:51.812 -> Water_Temp.:105.01, PWM:39
02:08:52.605 -> Water_Temp.:105.01, PWM:40
02:08:54.163 -> Water_Temp.:105.01, PWM:39
02:08:55.679 -> Water_Temp.:104.90, PWM:39
02:08:57.266 -> Water_Temp.:104.90, PWM:40
02:08:58.023 -> Water_Temp.:104.90, PWM:40
02:08:58.816 -> Water_Temp.:104.90, PWM:39
02:08:59.607 -> Water_Temp.:104.90, PWM:39
02:09:00.364 -> Water_Temp.:104.90, PWM:39
02:09:01.915 -> Water_Temp.:104.90, PWM:39
02:09:02.712 -> Water_Temp.:104.90, PWM:39
02:09:03.474 -> Water_Temp.:104.90, PWM:40
02:09:05.825 -> Water_Temp.:104.90, PWM:39
02:09:06.584 -> Water_Temp.:104.90, PWM:39
02:09:08.137 -> Water_Temp.:104.90, PWM:40
02:09:08.933 -> Water_Temp.:105.01, PWM:39
02:09:10.491 -> Water_Temp.:105.01, PWM:39
02:09:11.254 -> Water_Temp.:105.01, PWM:40
02:09:13.603 -> Water_Temp.:106.59, PWM:40
02:09:14.365 -> Water_Temp.:108.72, PWM:40
02:09:15.955 -> Water_Temp.:105.91, PWM:40
02:09:18.272 -> Water_Temp.:105.35, PWM:39
02:09:19.067 -> Water_Temp.:105.35, PWM:40
02:09:19.826 -> Water_Temp.:105.24, PWM:39
02:09:22.110 -> Water_Temp.:105.12, PWM:40
02:09:22.902 -> Water_Temp.:105.12, PWM:39
02:09:23.666 -> Water_Temp.:105.12, PWM:39
02:09:24.460 -> Water_Temp.:105.12, PWM:39
02:09:25.259 -> Water_Temp.:105.12, PWM:39
02:09:26.021 -> Water_Temp.:105.12, PWM:39
02:09:26.814 -> Water_Temp.:105.12, PWM:40
02:09:29.124 -> Water_Temp.:105.12, PWM:39
02:09:29.884 -> Water_Temp.:105.12, PWM:40
02:09:30.680 -> Water_Temp.:105.01, PWM:39
02:09:31.440 -> Water_Temp.:105.01, PWM:39
02:09:32.236 -> Water_Temp.:105.01, PWM:40
02:09:32.997 -> Water_Temp.:105.01, PWM:39
02:09:33.792 -> Water_Temp.:105.12, PWM:39
02:09:34.590 -> Water_Temp.:105.12, PWM:39
02:09:35.354 -> Water_Temp.:105.12, PWM:40
02:09:36.880 -> Water_Temp.:105.12, PWM:40
02:09:37.674 -> Water_Temp.:105.12, PWM:40
02:09:38.470 -> Water_Temp.:105.12, PWM:39
02:09:39.231 -> Water_Temp.:105.12, PWM:40
02:09:40.026 -> Water_Temp.:105.12, PWM:40
02:09:40.789 -> Water_Temp.:105.12, PWM:40
02:09:41.590 -> Water_Temp.:105.12, PWM:40
02:09:46.188 -> Water_Temp.:105.12, PWM:40
02:09:48.510 -> Water_Temp.:105.12, PWM:39
02:09:49.272 -> Water_Temp.:105.12, PWM:39
02:09:50.823 -> Water_Temp.:105.12, PWM:39
02:09:51.616 -> Water_Temp.:105.01, PWM:39
02:09:53.166 -> Water_Temp.:105.12, PWM:39
02:09:53.928 -> Water_Temp.:105.12, PWM:40
02:09:54.723 -> Water_Temp.:105.12, PWM:40
02:09:55.485 -> Water_Temp.:105.12, PWM:39
02:09:56.282 -> Water_Temp.:105.12, PWM:39
02:09:57.835 -> Water_Temp.:105.01, PWM:39
02:09:58.595 -> Water_Temp.:105.01, PWM:39
02:10:00.155 -> Water_Temp.:105.01, PWM:40
02:10:00.952 -> Water_Temp.:105.01, PWM:39
02:10:01.710 -> Water_Temp.:105.01, PWM:39
02:10:02.509 -> Water_Temp.:105.01, PWM:39
02:10:03.269 -> Water_Temp.:105.01, PWM:39
02:10:04.826 -> Water_Temp.:105.01, PWM:40
02:10:06.382 -> Water_Temp.:104.90, PWM:40
02:10:07.934 -> Water_Temp.:104.90, PWM:40
02:10:08.728 -> Water_Temp.:104.90, PWM:39
02:10:09.490 -> Water_Temp.:104.90, PWM:39
02:10:11.045 -> Water_Temp.:104.90, PWM:39
02:10:14.911 -> Water_Temp.:104.90, PWM:39
02:10:15.672 -> Water_Temp.:104.90, PWM:40
02:10:16.469 -> Water_Temp.:104.90, PWM:39
02:10:17.228 -> Water_Temp.:104.90, PWM:39
02:10:18.022 -> Water_Temp.:104.90, PWM:39
02:10:19.577 -> Water_Temp.:104.90, PWM:39
02:10:20.376 -> Water_Temp.:104.90, PWM:40
02:10:21.140 -> Water_Temp.:104.90, PWM:39
02:10:21.933 -> Water_Temp.:104.90, PWM:39
02:10:23.452 -> Water_Temp.:104.90, PWM:39
02:10:24.248 -> Water_Temp.:104.90, PWM:40
02:10:25.010 -> Water_Temp.:104.90, PWM:39
02:10:25.804 -> Water_Temp.:104.90, PWM:39
02:10:27.329 -> Water_Temp.:105.01, PWM:40
02:10:29.647 -> Water_Temp.:105.01, PWM:40
02:10:30.444 -> Water_Temp.:105.01, PWM:39
02:10:32.001 -> Water_Temp.:105.01, PWM:40
02:10:32.761 -> Water_Temp.:105.01, PWM:40
02:10:33.554 -> Water_Temp.:105.01, PWM:39
02:10:35.114 -> Water_Temp.:105.01, PWM:40
02:10:36.633 -> Water_Temp.:105.01, PWM:39
02:10:37.430 -> Water_Temp.:105.01, PWM:39
02:10:38.222 -> Water_Temp.:105.01, PWM:39
02:10:38.977 -> Water_Temp.:105.01, PWM:39
02:10:39.769 -> Water_Temp.:105.01, PWM:39
02:10:43.607 -> Water_Temp.:105.01, PWM:39
02:10:44.405 -> Water_Temp.:105.01, PWM:40
02:10:45.960 -> Water_Temp.:105.01, PWM:39
02:10:46.720 -> Water_Temp.:105.01, PWM:39
02:10:48.280 -> Water_Temp.:105.01, PWM:39
02:10:49.076 -> Water_Temp.:105.01, PWM:39
02:10:52.186 -> Water_Temp.:105.12, PWM:39
02:10:52.948 -> Water_Temp.:105.01, PWM:40
02:10:53.745 -> Water_Temp.:105.12, PWM:39
02:10:54.506 -> Water_Temp.:105.01, PWM:39
02:10:56.856 -> Water_Temp.:105.12, PWM:40
02:10:58.412 -> Water_Temp.:105.12, PWM:39
02:10:59.170 -> Water_Temp.:105.12, PWM:40
02:10:59.966 -> Water_Temp.:105.12, PWM:40
02:11:00.721 -> Water_Temp.:105.12, PWM:40
02:11:01.518 -> Water_Temp.:105.12, PWM:40
02:11:03.072 -> Water_Temp.:105.12, PWM:39
02:11:03.871 -> Water_Temp.:105.12, PWM:39
02:11:04.633 -> Water_Temp.:105.12, PWM:39
02:11:05.426 -> Water_Temp.:105.12, PWM:39
02:11:06.186 -> Water_Temp.:105.12, PWM:39
02:11:07.741 -> Water_Temp.:105.12, PWM:40
02:11:08.534 -> Water_Temp.:105.12, PWM:39
02:11:09.323 -> Water_Temp.:105.12, PWM:39
02:11:10.081 -> Water_Temp.:105.12, PWM:39
02:11:11.637 -> Water_Temp.:105.12, PWM:40
02:11:13.956 -> Water_Temp.:105.01, PWM:40
02:11:16.317 -> Water_Temp.:105.01, PWM:39
02:11:17.071 -> Water_Temp.:105.01, PWM:39
02:11:18.627 -> Water_Temp.:105.01, PWM:39
02:11:19.423 -> Water_Temp.:105.01, PWM:39
02:11:20.217 -> Water_Temp.:105.01, PWM:39
02:11:20.982 -> Water_Temp.:105.01, PWM:40
02:11:21.776 -> Water_Temp.:105.01, PWM:39
02:11:22.535 -> Water_Temp.:105.01, PWM:40
02:11:24.081 -> Water_Temp.:105.12, PWM:40
02:11:24.878 -> Water_Temp.:105.12, PWM:40
02:11:27.191 -> Water_Temp.:105.12, PWM:40
02:11:28.745 -> Water_Temp.:105.12, PWM:39
02:11:30.303 -> Water_Temp.:105.12, PWM:40
02:11:31.103 -> Water_Temp.:105.12, PWM:39
02:11:31.895 -> Water_Temp.:105.12, PWM:39
02:11:32.657 -> Water_Temp.:105.12, PWM:40
02:11:33.423 -> Water_Temp.:105.12, PWM:39
02:11:34.219 -> Water_Temp.:105.12, PWM:39
02:11:35.739 -> Water_Temp.:105.12, PWM:40
02:11:36.530 -> Water_Temp.:105.12, PWM:40
02:11:37.321 -> Water_Temp.:105.12, PWM:40
02:11:38.079 -> Water_Temp.:105.12, PWM:39
02:11:39.636 -> Water_Temp.:105.12, PWM:39
02:11:40.430 -> Water_Temp.:105.12, PWM:40
02:11:41.225 -> Water_Temp.:105.12, PWM:40
02:11:44.291 -> Water_Temp.:105.12, PWM:39
02:11:45.054 -> Water_Temp.:105.12, PWM:39
02:11:45.847 -> Water_Temp.:105.12, PWM:40
02:11:46.636 -> Water_Temp.:105.12, PWM:39
02:11:47.398 -> Water_Temp.:105.12, PWM:40
02:11:48.193 -> Water_Temp.:105.24, PWM:40
02:11:48.984 -> Water_Temp.:105.24, PWM:40
02:11:49.746 -> Water_Temp.:105.24, PWM:39
02:11:52.822 -> Water_Temp.:105.24, PWM:39
02:11:56.658 -> Water_Temp.:105.12, PWM:39
02:11:58.212 -> Water_Temp.:105.12, PWM:39
02:11:59.008 -> Water_Temp.:105.12, PWM:39
02:11:59.802 -> Water_Temp.:105.12, PWM:40
02:12:01.357 -> Water_Temp.:105.12, PWM:39
02:12:02.116 -> Water_Temp.:105.12, PWM:39
02:12:02.913 -> Water_Temp.:105.12, PWM:40
02:12:03.671 -> Water_Temp.:106.81, PWM:39
02:12:04.467 -> Water_Temp.:108.84, PWM:40
02:12:06.025 -> Water_Temp.:106.36, PWM:40
02:12:06.818 -> Water_Temp.:105.80, PWM:40
02:12:07.579 -> Water_Temp.:105.69, PWM:39
02:12:08.375 -> Water_Temp.:105.57, PWM:39
02:12:09.938 -> Water_Temp.:105.46, PWM:39
02:12:10.698 -> Water_Temp.:105.46, PWM:40
02:12:11.495 -> Water_Temp.:105.46, PWM:40
02:12:13.780 -> Water_Temp.:105.46, PWM:40
02:12:15.334 -> Water_Temp.:105.46, PWM:40
02:12:16.094 -> Water_Temp.:105.46, PWM:40
02:12:16.889 -> Water_Temp.:105.35, PWM:40
02:12:17.651 -> Water_Temp.:105.35, PWM:39
02:12:18.447 -> Water_Temp.:105.35, PWM:40
02:12:20.004 -> Water_Temp.:105.35, PWM:40
02:12:20.801 -> Water_Temp.:105.35, PWM:39
02:12:22.359 -> Water_Temp.:105.35, PWM:39
02:12:23.879 -> Water_Temp.:105.35, PWM:40
02:12:24.679 -> Water_Temp.:105.46, PWM:40
02:12:25.442 -> Water_Temp.:105.46, PWM:39
02:12:26.235 -> Water_Temp.:105.46, PWM:39
02:12:27.793 -> Water_Temp.:105.35, PWM:39
02:12:28.554 -> Water_Temp.:105.35, PWM:39
02:12:29.351 -> Water_Temp.:105.35, PWM:40
02:12:30.112 -> Water_Temp.:105.35, PWM:39
02:12:31.672 -> Water_Temp.:105.35, PWM:39
02:12:33.225 -> Water_Temp.:105.35, PWM:39
02:12:34.018 -> Water_Temp.:105.24, PWM:40
02:12:36.340 -> Water_Temp.:105.24, PWM:39
02:12:37.134 -> Water_Temp.:105.24, PWM:40
02:12:37.894 -> Water_Temp.:105.24, PWM:39
02:12:38.687 -> Water_Temp.:105.24, PWM:39
02:12:39.450 -> Water_Temp.:105.24, PWM:40
02:12:40.247 -> Water_Temp.:105.24, PWM:39
02:12:41.042 -> Water_Temp.:105.12, PWM:39
02:12:41.802 -> Water_Temp.:105.12, PWM:39
02:12:44.152 -> Water_Temp.:105.12, PWM:39
02:12:47.224 -> Water_Temp.:105.12, PWM:39
02:12:48.779 -> Water_Temp.:105.12, PWM:40
02:12:49.572 -> Water_Temp.:105.12, PWM:40
02:12:50.336 -> Water_Temp.:105.12, PWM:39
02:12:51.893 -> Water_Temp.:105.12, PWM:39
02:12:53.454 -> Water_Temp.:105.12, PWM:39
02:12:54.248 -> Water_Temp.:105.12, PWM:39
02:12:56.569 -> Water_Temp.:105.12, PWM:40
02:12:57.332 -> Water_Temp.:105.12, PWM:39
02:12:59.642 -> Water_Temp.:105.12, PWM:39
02:13:00.438 -> Water_Temp.:105.12, PWM:40
02:13:01.201 -> Water_Temp.:105.12, PWM:39
02:13:01.999 -> Water_Temp.:105.12, PWM:40
02:13:02.760 -> Water_Temp.:105.12, PWM:39
02:13:05.868 -> Water_Temp.:105.12, PWM:40
02:13:06.629 -> Water_Temp.:105.12, PWM:39
02:13:07.424 -> Water_Temp.:105.12, PWM:39
02:13:08.219 -> Water_Temp.:105.12, PWM:39
02:13:09.734 -> Water_Temp.:105.12, PWM:40
02:13:10.531 -> Water_Temp.:105.12, PWM:40
02:13:11.293 -> Water_Temp.:105.12, PWM:39
02:13:12.088 -> Water_Temp.:105.12, PWM:39
02:13:12.847 -> Water_Temp.:105.12, PWM:40
02:13:13.645 -> Water_Temp.:105.12, PWM:39
02:13:15.166 -> Water_Temp.:105.12, PWM:40
02:13:15.960 -> Water_Temp.:105.12, PWM:39
02:13:16.753 -> Water_Temp.:105.12, PWM:40
02:13:17.513 -> Water_Temp.:105.12, PWM:40
02:13:19.869 -> Water_Temp.:105.01, PWM:39
02:13:20.631 -> Water_Temp.:105.01, PWM:39
02:13:21.427 -> Water_Temp.:105.01, PWM:39
02:13:22.184 -> Water_Temp.:105.01, PWM:40
02:13:22.980 -> Water_Temp.:105.01, PWM:40
02:13:24.536 -> Water_Temp.:105.01, PWM:40
02:13:26.093 -> Water_Temp.:105.01, PWM:39
02:13:27.648 -> Water_Temp.:105.01, PWM:39
02:13:28.409 -> Water_Temp.:105.01, PWM:39
02:13:30.768 -> Water_Temp.:105.01, PWM:39
02:13:31.525 -> Water_Temp.:105.01, PWM:39
02:13:32.317 -> Water_Temp.:105.01, PWM:40
02:13:33.075 -> Water_Temp.:105.01, PWM:39
02:13:36.914 -> Water_Temp.:105.01, PWM:39
02:13:37.677 -> Water_Temp.:105.01, PWM:40
02:13:38.473 -> Water_Temp.:105.01, PWM:40
02:13:40.025 -> Water_Temp.:105.01, PWM:39
02:13:40.789 -> Water_Temp.:105.01, PWM:39
02:13:42.344 -> Water_Temp.:105.01, PWM:40
02:13:43.138 -> Water_Temp.:105.01, PWM:39
02:13:45.458 -> Water_Temp.:105.01, PWM:40
02:13:46.222 -> Water_Temp.:105.01, PWM:40
02:13:47.016 -> Water_Temp.:105.01, PWM:40
02:13:48.575 -> Water_Temp.:105.01, PWM:39
02:13:50.131 -> Water_Temp.:105.01, PWM:39
02:13:50.895 -> Water_Temp.:105.01, PWM:39
02:13:51.689 -> Water_Temp.:105.01, PWM:40
02:13:52.449 -> Water_Temp.:105.01, PWM:40
02:13:53.246 -> Water_Temp.:105.01, PWM:40
02:13:56.332 -> Water_Temp.:105.01, PWM:39
02:13:57.128 -> Water_Temp.:105.01, PWM:39
02:13:57.891 -> Water_Temp.:105.01, PWM:40
02:13:59.453 -> Water_Temp.:105.01, PWM:39
02:14:00.213 -> Water_Temp.:105.01, PWM:40
02:14:01.011 -> Water_Temp.:105.01, PWM:40
02:14:01.770 -> Water_Temp.:105.01, PWM:40
02:14:03.324 -> Water_Temp.:105.01, PWM:39
02:14:04.118 -> Water_Temp.:105.01, PWM:40
02:14:04.913 -> Water_Temp.:105.01, PWM:39
02:14:05.672 -> Water_Temp.:105.01, PWM:39
02:14:06.466 -> Water_Temp.:105.01, PWM:39
02:14:07.229 -> Water_Temp.:105.01, PWM:40
02:14:08.026 -> Water_Temp.:105.01, PWM:40
02:14:11.072 -> Water_Temp.:105.01, PWM:40
02:14:11.864 -> Water_Temp.:105.01, PWM:39
02:14:13.422 -> Water_Temp.:105.01, PWM:39
02:14:14.216 -> Water_Temp.:105.01, PWM:40
02:14:14.975 -> Water_Temp.:105.01, PWM:40
02:14:17.290 -> Water_Temp.:105.01, PWM:39
02:14:18.086 -> Water_Temp.:105.01, PWM:40
02:14:18.845 -> Water_Temp.:105.01, PWM:39
02:14:19.642 -> Water_Temp.:105.01, PWM:39
02:14:21.993 -> Water_Temp.:105.01, PWM:39
02:14:22.754 -> Water_Temp.:105.01, PWM:39
02:14:23.552 -> Water_Temp.:105.01, PWM:39
02:14:24.314 -> Water_Temp.:105.01, PWM:39
02:14:25.865 -> Water_Temp.:105.01, PWM:40
02:14:28.178 -> Water_Temp.:105.01, PWM:39
02:14:28.974 -> Water_Temp.:105.01, PWM:39
02:14:29.736 -> Water_Temp.:105.01, PWM:39
02:14:30.532 -> Water_Temp.:105.01, PWM:39
02:14:33.577 -> Water_Temp.:104.90, PWM:39
02:14:34.338 -> Water_Temp.:105.01, PWM:40
02:14:35.130 -> Water_Temp.:104.90, PWM:40
02:14:35.889 -> Water_Temp.:104.90, PWM:39
02:14:37.479 -> Water_Temp.:104.90, PWM:39
02:14:38.240 -> Water_Temp.:104.90, PWM:39
02:14:40.549 -> Water_Temp.:104.90, PWM:39
02:14:41.345 -> Water_Temp.:104.90, PWM:39
02:14:42.142 -> Water_Temp.:104.90, PWM:39
02:14:42.897 -> Water_Temp.:104.90, PWM:40
02:14:43.692 -> Water_Temp.:104.90, PWM:39
02:14:44.457 -> Water_Temp.:104.90, PWM:40
02:14:45.252 -> Water_Temp.:104.90, PWM:40
02:14:46.009 -> Water_Temp.:104.90, PWM:39
02:14:46.801 -> Water_Temp.:104.90, PWM:40
02:14:47.599 -> Water_Temp.:104.90, PWM:39
02:14:48.361 -> Water_Temp.:104.90, PWM:40
02:14:49.158 -> Water_Temp.:104.90, PWM:39
02:14:49.917 -> Water_Temp.:104.90, PWM:39
02:14:50.714 -> Water_Temp.:104.90, PWM:39
02:14:51.475 -> Water_Temp.:104.90, PWM:39
02:14:52.270 -> Water_Temp.:104.90, PWM:39
02:14:53.830 -> Water_Temp.:104.90, PWM:39
02:14:54.592 -> Water_Temp.:104.90, PWM:40
02:14:55.388 -> Water_Temp.:104.90, PWM:39
02:14:56.151 -> Water_Temp.:104.90, PWM:40
02:14:56.949 -> Water_Temp.:104.90, PWM:39
02:14:57.712 -> Water_Temp.:104.90, PWM:40
02:14:59.267 -> Water_Temp.:104.90, PWM:40
02:15:00.061 -> Water_Temp.:104.90, PWM:40
02:15:00.860 -> Water_Temp.:104.90, PWM:39
02:15:02.374 -> Water_Temp.:104.90, PWM:40
02:15:03.168 -> Water_Temp.:104.90, PWM:40
02:15:04.725 -> Water_Temp.:104.90, PWM:40
02:15:06.281 -> Water_Temp.:104.90, PWM:40
02:15:07.075 -> Water_Temp.:104.90, PWM:40
02:15:07.831 -> Water_Temp.:104.90, PWM:40
02:15:09.419 -> Water_Temp.:104.90, PWM:39
02:15:10.943 -> Water_Temp.:104.90, PWM:40
02:15:11.740 -> Water_Temp.:104.90, PWM:40
02:15:14.055 -> Water_Temp.:104.90, PWM:39
02:15:14.850 -> Water_Temp.:104.90, PWM:39
02:15:15.642 -> Water_Temp.:104.90, PWM:40
02:15:16.402 -> Water_Temp.:104.90, PWM:39
02:15:17.197 -> Water_Temp.:104.90, PWM:39
02:15:17.957 -> Water_Temp.:104.90, PWM:39
02:15:21.036 -> Water_Temp.:104.90, PWM:39
02:15:21.830 -> Water_Temp.:104.90, PWM:39
02:15:22.590 -> Water_Temp.:104.79, PWM:39
02:15:23.384 -> Water_Temp.:104.90, PWM:39
02:15:24.141 -> Water_Temp.:104.90, PWM:39
02:15:24.935 -> Water_Temp.:104.90, PWM:40
02:15:25.700 -> Water_Temp.:104.90, PWM:40
02:15:27.261 -> Water_Temp.:104.90, PWM:39
02:15:28.023 -> Water_Temp.:104.90, PWM:39
02:15:28.823 -> Water_Temp.:104.90, PWM:40
02:15:30.345 -> Water_Temp.:104.90, PWM:39
02:15:31.141 -> Water_Temp.:104.90, PWM:39
02:15:31.928 -> Water_Temp.:104.90, PWM:39
02:15:34.976 -> Water_Temp.:105.01, PWM:40
02:15:35.769 -> Water_Temp.:105.01, PWM:39
02:15:36.529 -> Water_Temp.:106.59, PWM:40
02:15:37.327 -> Water_Temp.:108.72, PWM:39
02:15:38.086 -> Water_Temp.:106.25, PWM:40
02:15:38.879 -> Water_Temp.:106.25, PWM:39
02:15:40.438 -> Water_Temp.:105.46, PWM:39
02:15:41.203 -> Water_Temp.:105.35, PWM:40
02:15:41.993 -> Water_Temp.:105.35, PWM:39
02:15:42.788 -> Water_Temp.:105.24, PWM:39
02:15:43.546 -> Water_Temp.:105.24, PWM:40
02:15:44.340 -> Water_Temp.:105.12, PWM:40
02:15:45.864 -> Water_Temp.:105.12, PWM:39
02:15:46.654 -> Water_Temp.:105.12, PWM:39
02:15:47.448 -> Water_Temp.:105.12, PWM:39
02:15:48.212 -> Water_Temp.:105.01, PWM:40
02:15:49.007 -> Water_Temp.:105.01, PWM:39
02:15:50.563 -> Water_Temp.:105.01, PWM:39
02:15:51.323 -> Water_Temp.:105.01, PWM:40
02:15:53.642 -> Water_Temp.:105.01, PWM:40
02:15:54.433 -> Water_Temp.:105.01, PWM:40
02:15:55.993 -> Water_Temp.:104.90, PWM:39
02:15:56.789 -> Water_Temp.:104.90, PWM:39
02:15:57.551 -> Water_Temp.:104.90, PWM:39
02:16:00.635 -> Water_Temp.:104.90, PWM:39
02:16:01.432 -> Water_Temp.:104.90, PWM:40
02:16:02.188 -> Water_Temp.:104.90, PWM:39
02:16:02.985 -> Water_Temp.:104.90, PWM:39
02:16:04.504 -> Water_Temp.:104.90, PWM:39
02:16:05.300 -> Water_Temp.:104.90, PWM:39
02:16:06.088 -> Water_Temp.:104.90, PWM:40
02:16:06.848 -> Water_Temp.:104.90, PWM:39
02:16:07.645 -> Water_Temp.:104.90, PWM:39
02:16:08.411 -> Water_Temp.:104.90, PWM:40
02:16:09.206 -> Water_Temp.:104.90, PWM:39
02:16:10.761 -> Water_Temp.:105.01, PWM:39
02:16:11.555 -> Water_Temp.:105.01, PWM:39
02:16:13.875 -> Water_Temp.:105.01, PWM:40
02:16:14.634 -> Water_Temp.:104.90, PWM:39
02:16:16.187 -> Water_Temp.:104.90, PWM:39
02:16:16.982 -> Water_Temp.:104.90, PWM:40
02:16:18.539 -> Water_Temp.:104.90, PWM:39
02:16:20.098 -> Water_Temp.:104.90, PWM:39
02:16:20.861 -> Water_Temp.:104.90, PWM:39
02:16:21.659 -> Water_Temp.:104.90, PWM:40
02:16:22.422 -> Water_Temp.:104.90, PWM:39
02:16:23.216 -> Water_Temp.:104.90, PWM:39
02:16:23.977 -> Water_Temp.:104.90, PWM:39
02:16:24.771 -> Water_Temp.:104.90, PWM:39
02:16:26.322 -> Water_Temp.:104.79, PWM:40
02:16:27.115 -> Water_Temp.:104.79, PWM:39
02:16:29.439 -> Water_Temp.:104.90, PWM:39
02:16:30.200 -> Water_Temp.:104.90, PWM:39
02:16:31.750 -> Water_Temp.:104.90, PWM:39
02:16:33.309 -> Water_Temp.:104.90, PWM:40
02:16:34.101 -> Water_Temp.:104.90, PWM:39
02:16:34.898 -> Water_Temp.:104.90, PWM:39
02:16:35.661 -> Water_Temp.:104.90, PWM:39
02:16:36.456 -> Water_Temp.:104.90, PWM:39
02:16:37.214 -> Water_Temp.:104.90, PWM:40
02:16:38.010 -> Water_Temp.:104.90, PWM:40
02:16:39.531 -> Water_Temp.:104.90, PWM:40
02:16:40.326 -> Water_Temp.:104.90, PWM:39
02:16:41.884 -> Water_Temp.:104.90, PWM:39
02:16:42.647 -> Water_Temp.:104.90, PWM:39
02:16:44.202 -> Water_Temp.:104.90, PWM:39
02:16:46.548 -> Water_Temp.:105.01, PWM:40
02:16:47.339 -> Water_Temp.:105.01, PWM:39
02:16:48.102 -> Water_Temp.:105.01, PWM:40
02:16:48.891 -> Water_Temp.:105.01, PWM:40
02:16:49.684 -> Water_Temp.:105.01, PWM:40
02:16:50.444 -> Water_Temp.:105.01, PWM:39
02:16:51.243 -> Water_Temp.:105.01, PWM:39
02:16:52.762 -> Water_Temp.:105.01, PWM:39
02:16:53.524 -> Water_Temp.:105.01, PWM:39
02:16:55.871 -> Water_Temp.:105.01, PWM:40
02:16:56.661 -> Water_Temp.:105.01, PWM:40
02:16:57.452 -> Water_Temp.:105.01, PWM:40
02:16:58.211 -> Water_Temp.:105.01, PWM:40
02:17:00.513 -> Water_Temp.:105.01, PWM:39
02:17:01.278 -> Water_Temp.:105.01, PWM:39
02:17:02.869 -> Water_Temp.:105.01, PWM:40
02:17:03.634 -> Water_Temp.:105.01, PWM:39
02:17:04.425 -> Water_Temp.:105.01, PWM:39
02:17:05.184 -> Water_Temp.:105.01, PWM:39
02:17:06.739 -> Water_Temp.:105.01, PWM:39
02:17:07.536 -> Water_Temp.:105.01, PWM:39
02:17:09.816 -> Water_Temp.:105.01, PWM:39
02:17:10.610 -> Water_Temp.:105.01, PWM:39
02:17:11.406 -> Water_Temp.:105.01, PWM:40
02:17:12.160 -> Water_Temp.:105.01, PWM:39
02:17:13.722 -> Water_Temp.:104.90, PWM:39
02:17:14.485 -> Water_Temp.:105.01, PWM:40
02:17:15.282 -> Water_Temp.:105.01, PWM:39
02:17:16.043 -> Water_Temp.:104.90, PWM:40
02:17:20.710 -> Water_Temp.:105.01, PWM:39
02:17:21.506 -> Water_Temp.:104.90, PWM:39
02:17:22.268 -> Water_Temp.:104.90, PWM:40
02:17:23.823 -> Water_Temp.:104.90, PWM:40
02:17:24.585 -> Water_Temp.:104.90, PWM:39
02:17:25.382 -> Water_Temp.:104.90, PWM:39
02:17:26.936 -> Water_Temp.:104.90, PWM:39
02:17:27.698 -> Water_Temp.:104.90, PWM:39
02:17:29.285 -> Water_Temp.:104.90, PWM:40
02:17:30.047 -> Water_Temp.:104.90, PWM:39
02:17:30.845 -> Water_Temp.:104.90, PWM:40
02:17:31.606 -> Water_Temp.:104.90, PWM:39
02:17:32.404 -> Water_Temp.:104.90, PWM:39
02:17:34.719 -> Water_Temp.:104.90, PWM:39
02:17:36.278 -> Water_Temp.:104.90, PWM:43
02:17:37.037 -> Water_Temp.:104.90, PWM:43
02:17:37.832 -> Water_Temp.:104.90, PWM:43
02:17:38.597 -> Water_Temp.:104.90, PWM:43
02:17:40.155 -> Water_Temp.:104.90, PWM:43
02:17:41.680 -> Water_Temp.:104.90, PWM:43
02:17:42.473 -> Water_Temp.:104.90, PWM:43
02:17:43.236 -> Water_Temp.:104.90, PWM:43
02:17:47.108 -> Water_Temp.:104.90, PWM:43
02:17:50.156 -> Water_Temp.:104.90, PWM:43
02:17:50.917 -> Water_Temp.:104.90, PWM:43
02:17:51.713 -> Water_Temp.:104.90, PWM:43
02:17:52.510 -> Water_Temp.:104.90, PWM:43
02:17:54.033 -> Water_Temp.:104.90, PWM:43
02:17:54.794 -> Water_Temp.:104.90, PWM:43
02:17:55.590 -> Water_Temp.:104.90, PWM:43
02:17:56.384 -> Water_Temp.:104.90, PWM:43
02:17:57.143 -> Water_Temp.:104.90, PWM:43
02:17:58.700 -> Water_Temp.:104.90, PWM:43
02:17:59.497 -> Water_Temp.:104.90, PWM:43
02:18:00.258 -> Water_Temp.:104.90, PWM:43
02:18:01.814 -> Water_Temp.:104.90, PWM:43
02:18:02.606 -> Water_Temp.:104.90, PWM:43
02:18:04.130 -> Water_Temp.:104.90, PWM:43
02:18:05.724 -> Water_Temp.:104.90, PWM:43
02:18:06.483 -> Water_Temp.:104.90, PWM:43
02:18:08.044 -> Water_Temp.:104.90, PWM:43
02:18:08.806 -> Water_Temp.:104.90, PWM:43
02:18:09.603 -> Water_Temp.:104.90, PWM:43
02:18:11.159 -> Water_Temp.:104.90, PWM:43
02:18:11.918 -> Water_Temp.:104.90, PWM:43
02:18:12.715 -> Water_Temp.:104.90, PWM:43
02:18:13.477 -> Water_Temp.:104.90, PWM:43
02:18:15.034 -> Water_Temp.:104.90, PWM:43
02:18:16.548 -> Water_Temp.:104.90, PWM:43
02:18:17.342 -> Water_Temp.:105.01, PWM:43
02:18:18.101 -> Water_Temp.:104.90, PWM:43
02:18:18.891 -> Water_Temp.:104.90, PWM:43
02:18:19.684 -> Water_Temp.:105.01, PWM:43
02:18:20.445 -> Water_Temp.:105.01, PWM:43
02:18:22.757 -> Water_Temp.:105.01, PWM:43
02:18:23.548 -> Water_Temp.:105.01, PWM:43
02:18:25.859 -> Water_Temp.:105.01, PWM:43
02:18:26.654 -> Water_Temp.:105.01, PWM:43
02:18:28.207 -> Water_Temp.:105.01, PWM:43
02:18:28.963 -> Water_Temp.:105.01, PWM:43
02:18:29.756 -> Water_Temp.:105.01, PWM:43
02:18:30.553 -> Water_Temp.:105.01, PWM:43
02:18:31.316 -> Water_Temp.:105.01, PWM:43
02:18:32.111 -> Water_Temp.:105.01, PWM:43
02:18:32.873 -> Water_Temp.:105.01, PWM:43
02:18:35.183 -> Water_Temp.:105.01, PWM:43
02:18:37.534 -> Water_Temp.:105.01, PWM:43
02:18:38.293 -> Water_Temp.:105.01, PWM:43
02:18:39.084 -> Water_Temp.:105.01, PWM:43
02:18:39.879 -> Water_Temp.:105.01, PWM:43
02:18:40.639 -> Water_Temp.:105.01, PWM:43
02:18:41.435 -> Water_Temp.:105.01, PWM:43
02:18:42.195 -> Water_Temp.:106.59, PWM:43
02:18:42.990 -> Water_Temp.:108.72, PWM:43
02:18:44.552 -> Water_Temp.:105.80, PWM:43
02:18:45.309 -> Water_Temp.:105.69, PWM:43
02:18:46.103 -> Water_Temp.:105.46, PWM:43
02:18:48.414 -> Water_Temp.:105.24, PWM:43
02:18:49.999 -> Water_Temp.:105.12, PWM:43
02:18:51.523 -> Water_Temp.:105.12, PWM:43
02:18:52.315 -> Water_Temp.:105.01, PWM:43
02:18:53.074 -> Water_Temp.:105.01, PWM:43
02:18:54.634 -> Water_Temp.:105.01, PWM:43
02:18:55.428 -> Water_Temp.:105.01, PWM:43
02:18:59.268 -> Water_Temp.:104.90, PWM:43
02:19:00.031 -> Water_Temp.:104.90, PWM:43
02:19:00.825 -> Water_Temp.:104.90, PWM:43
02:19:01.618 -> Water_Temp.:104.90, PWM:43
02:19:02.377 -> Water_Temp.:104.90, PWM:43
02:19:03.170 -> Water_Temp.:104.90, PWM:43
02:19:03.929 -> Water_Temp.:104.90, PWM:43
02:19:04.724 -> Water_Temp.:105.01, PWM:43
02:19:05.521 -> Water_Temp.:105.01, PWM:43
02:19:07.040 -> Water_Temp.:105.01, PWM:43
02:19:07.834 -> Water_Temp.:105.01, PWM:43
02:19:08.593 -> Water_Temp.:105.01, PWM:43
02:19:09.394 -> Water_Temp.:105.01, PWM:43
02:19:10.919 -> Water_Temp.:105.01, PWM:43
02:19:11.711 -> Water_Temp.:105.01, PWM:43
02:19:12.472 -> Water_Temp.:105.01, PWM:43
02:19:14.027 -> Water_Temp.:105.01, PWM:43
02:19:14.785 -> Water_Temp.:105.01, PWM:43
02:19:15.579 -> Water_Temp.:105.01, PWM:43
02:19:16.338 -> Water_Temp.:105.01, PWM:43
02:19:17.134 -> Water_Temp.:105.01, PWM:43
02:19:18.689 -> Water_Temp.:105.01, PWM:43
02:19:19.446 -> Water_Temp.:105.01, PWM:43
02:19:20.240 -> Water_Temp.:105.01, PWM:43
02:19:21.003 -> Water_Temp.:105.01, PWM:43
02:19:21.794 -> Water_Temp.:105.01, PWM:43
02:19:22.585 -> Water_Temp.:105.01, PWM:43
02:19:24.107 -> Water_Temp.:105.01, PWM:43
02:19:24.900 -> Water_Temp.:105.01, PWM:43
02:19:25.693 -> Water_Temp.:105.01, PWM:43
02:19:28.770 -> Water_Temp.:105.01, PWM:43
02:19:29.564 -> Water_Temp.:105.01, PWM:43
02:19:31.123 -> Water_Temp.:105.01, PWM:43
02:19:31.879 -> Water_Temp.:105.01, PWM:43
02:19:32.677 -> Water_Temp.:105.01, PWM:43
02:19:37.312 -> Water_Temp.:105.01, PWM:43
02:19:38.869 -> Water_Temp.:105.01, PWM:43
02:19:39.666 -> Water_Temp.:105.01, PWM:43
02:19:40.459 -> Water_Temp.:105.01, PWM:43
02:19:41.982 -> Water_Temp.:105.01, PWM:43
02:19:42.746 -> Water_Temp.:105.01, PWM:43
02:19:43.542 -> Water_Temp.:105.01, PWM:43
02:19:45.064 -> Water_Temp.:105.01, PWM:43
02:19:45.858 -> Water_Temp.:105.01, PWM:43
02:19:46.617 -> Water_Temp.:105.01, PWM:43
02:19:48.201 -> Water_Temp.:105.01, PWM:43
02:19:49.731 -> Water_Temp.:105.01, PWM:43
02:19:50.528 -> Water_Temp.:105.01, PWM:43
02:19:51.284 -> Water_Temp.:105.01, PWM:43
02:19:52.079 -> Water_Temp.:105.01, PWM:43
02:19:52.841 -> Water_Temp.:105.01, PWM:43
02:19:53.640 -> Water_Temp.:105.01, PWM:43
02:19:54.405 -> Water_Temp.:104.90, PWM:43
02:19:55.200 -> Water_Temp.:105.01, PWM:43
02:19:55.993 -> Water_Temp.:104.90, PWM:43
02:19:56.750 -> Water_Temp.:104.90, PWM:43
02:19:57.546 -> Water_Temp.:104.90, PWM:43
02:19:58.303 -> Water_Temp.:105.01, PWM:43
02:19:59.095 -> Water_Temp.:104.90, PWM:43
02:19:59.892 -> Water_Temp.:104.90, PWM:43
02:20:00.654 -> Water_Temp.:104.90, PWM:43
02:20:01.449 -> Water_Temp.:104.90, PWM:43
02:20:02.209 -> Water_Temp.:104.90, PWM:43
02:20:03.002 -> Water_Temp.:104.90, PWM:43
02:20:03.763 -> Water_Temp.:104.90, PWM:43
02:20:04.555 -> Water_Temp.:104.90, PWM:43
02:20:06.899 -> Water_Temp.:104.90, PWM:43
02:20:07.658 -> Water_Temp.:104.90, PWM:43
02:20:08.452 -> Water_Temp.:104.90, PWM:43
02:20:09.247 -> Water_Temp.:104.90, PWM:43
02:20:10.003 -> Water_Temp.:104.90, PWM:43
02:20:11.559 -> Water_Temp.:105.01, PWM:43
02:20:13.073 -> Water_Temp.:105.01, PWM:43
02:20:13.868 -> Water_Temp.:105.01, PWM:43
02:20:14.630 -> Water_Temp.:105.01, PWM:43
02:20:15.423 -> Water_Temp.:105.01, PWM:43
02:20:16.215 -> Water_Temp.:105.01, PWM:43
02:20:17.728 -> Water_Temp.:105.01, PWM:43
02:20:18.523 -> Water_Temp.:105.01, PWM:43
02:20:20.075 -> Water_Temp.:105.12, PWM:43
02:20:20.869 -> Water_Temp.:105.12, PWM:43
02:20:21.630 -> Water_Temp.:105.12, PWM:43
02:20:22.416 -> Water_Temp.:105.12, PWM:43
02:20:24.760 -> Water_Temp.:105.12, PWM:43
02:20:25.553 -> Water_Temp.:105.12, PWM:43
02:20:27.107 -> Water_Temp.:105.12, PWM:43
02:20:27.867 -> Water_Temp.:105.12, PWM:43
02:20:29.428 -> Water_Temp.:105.12, PWM:43
02:20:30.980 -> Water_Temp.:105.12, PWM:43
02:20:31.737 -> Water_Temp.:105.12, PWM:43
02:20:33.288 -> Water_Temp.:105.24, PWM:43
02:20:34.080 -> Water_Temp.:105.12, PWM:43
02:20:35.625 -> Water_Temp.:105.12, PWM:43
02:20:36.381 -> Water_Temp.:105.12, PWM:43
02:20:37.176 -> Water_Temp.:105.12, PWM:43
02:20:37.936 -> Water_Temp.:105.24, PWM:43
02:20:39.526 -> Water_Temp.:105.12, PWM:43
02:20:40.289 -> Water_Temp.:105.12, PWM:43
02:20:41.084 -> Water_Temp.:105.12, PWM:43
02:20:41.842 -> Water_Temp.:105.12, PWM:43
02:20:42.635 -> Water_Temp.:105.12, PWM:43
02:20:43.399 -> Water_Temp.:105.12, PWM:43
02:20:44.950 -> Water_Temp.:105.01, PWM:43
02:20:45.743 -> Water_Temp.:105.01, PWM:43
02:20:47.292 -> Water_Temp.:105.01, PWM:43
02:20:48.087 -> Water_Temp.:105.01, PWM:43
02:20:48.847 -> Water_Temp.:105.01, PWM:43
02:20:49.639 -> Water_Temp.:105.01, PWM:43
02:20:51.195 -> Water_Temp.:105.01, PWM:43
02:20:51.986 -> Water_Temp.:105.01, PWM:43
02:20:54.269 -> Water_Temp.:105.01, PWM:43
02:20:55.060 -> Water_Temp.:105.01, PWM:43
02:20:56.616 -> Water_Temp.:105.01, PWM:43
02:20:57.404 -> Water_Temp.:105.01, PWM:43
02:20:58.199 -> Water_Temp.:105.01, PWM:43
02:21:01.271 -> Water_Temp.:105.12, PWM:43
02:21:02.034 -> Water_Temp.:105.01, PWM:43
02:21:03.589 -> Water_Temp.:105.12, PWM:43
02:21:05.142 -> Water_Temp.:105.12, PWM:43
02:21:05.939 -> Water_Temp.:105.12, PWM:43
02:21:07.495 -> Water_Temp.:105.12, PWM:43
02:21:08.285 -> Water_Temp.:105.01, PWM:47
02:21:09.043 -> Water_Temp.:105.12, PWM:47
02:21:09.834 -> Water_Temp.:105.12, PWM:47
02:21:11.387 -> Water_Temp.:105.01, PWM:47
02:21:12.147 -> Water_Temp.:105.12, PWM:47
02:21:14.505 -> Water_Temp.:105.12, PWM:47
02:21:15.267 -> Water_Temp.:105.01, PWM:47
02:21:16.823 -> Water_Temp.:105.01, PWM:47
02:21:18.380 -> Water_Temp.:105.01, PWM:47
02:21:20.728 -> Water_Temp.:105.01, PWM:47
02:21:22.288 -> Water_Temp.:105.12, PWM:47
02:21:23.049 -> Water_Temp.:105.12, PWM:47
02:21:23.846 -> Water_Temp.:105.12, PWM:47
02:21:24.607 -> Water_Temp.:105.12, PWM:47
02:21:26.164 -> Water_Temp.:105.12, PWM:47
02:21:26.956 -> Water_Temp.:105.12, PWM:47
02:21:27.749 -> Water_Temp.:105.12, PWM:47
02:21:29.304 -> Water_Temp.:105.12, PWM:47
02:21:30.815 -> Water_Temp.:105.12, PWM:47
02:21:33.128 -> Water_Temp.:105.01, PWM:47
02:21:35.446 -> Water_Temp.:105.12, PWM:47
02:21:36.242 -> Water_Temp.:105.01, PWM:47
02:21:37.003 -> Water_Temp.:105.01, PWM:47
02:21:37.798 -> Water_Temp.:105.01, PWM:47
02:21:40.879 -> Water_Temp.:105.12, PWM:47
02:21:43.193 -> Water_Temp.:105.12, PWM:47
02:21:43.991 -> Water_Temp.:105.12, PWM:47
02:21:44.786 -> Water_Temp.:105.12, PWM:47
02:21:45.546 -> Water_Temp.:105.12, PWM:47
02:21:46.341 -> Water_Temp.:105.12, PWM:47
02:21:47.100 -> Water_Temp.:105.12, PWM:47
02:21:48.686 -> Water_Temp.:105.12, PWM:47
02:21:49.446 -> Water_Temp.:105.12, PWM:47
02:21:50.240 -> Water_Temp.:105.12, PWM:47
02:21:52.559 -> Water_Temp.:105.12, PWM:47
02:21:53.322 -> Water_Temp.:105.12, PWM:47
02:21:54.115 -> Water_Temp.:105.12, PWM:47
02:21:54.874 -> Water_Temp.:105.12, PWM:47
02:21:55.668 -> Water_Temp.:105.12, PWM:47
02:21:56.461 -> Water_Temp.:105.12, PWM:47
02:21:57.219 -> Water_Temp.:105.12, PWM:47
02:21:58.775 -> Water_Temp.:105.12, PWM:47
02:22:00.333 -> Water_Temp.:105.12, PWM:47
02:22:01.093 -> Water_Temp.:105.12, PWM:47
02:22:01.884 -> Water_Temp.:105.12, PWM:47
02:22:02.680 -> Water_Temp.:105.12, PWM:47
02:22:03.439 -> Water_Temp.:105.12, PWM:47
02:22:04.993 -> Water_Temp.:105.12, PWM:47
02:22:05.789 -> Water_Temp.:105.12, PWM:47
02:22:06.579 -> Water_Temp.:105.12, PWM:47
02:22:08.897 -> Water_Temp.:105.12, PWM:47
02:22:10.446 -> Water_Temp.:105.12, PWM:47
02:22:11.241 -> Water_Temp.:105.12, PWM:47
02:22:12.000 -> Water_Temp.:105.12, PWM:47
02:22:12.795 -> Water_Temp.:105.12, PWM:47
02:22:15.147 -> Water_Temp.:105.24, PWM:47
02:22:15.906 -> Water_Temp.:105.24, PWM:47
02:22:16.700 -> Water_Temp.:105.24, PWM:47
02:22:17.459 -> Water_Temp.:105.24, PWM:47
02:22:18.254 -> Water_Temp.:105.24, PWM:47
02:22:19.047 -> Water_Temp.:105.24, PWM:47
02:22:20.564 -> Water_Temp.:105.24, PWM:47
02:22:21.364 -> Water_Temp.:105.24, PWM:47
02:22:22.120 -> Water_Temp.:105.24, PWM:47
02:22:23.678 -> Water_Temp.:105.24, PWM:47
02:22:24.475 -> Water_Temp.:105.24, PWM:47
02:22:26.035 -> Water_Temp.:105.24, PWM:47
02:22:27.587 -> Water_Temp.:105.24, PWM:47
02:22:28.342 -> Water_Temp.:105.24, PWM:47
02:22:29.901 -> Water_Temp.:105.24, PWM:47
02:22:31.459 -> Water_Temp.:105.24, PWM:47
02:22:32.249 -> Water_Temp.:105.24, PWM:47
02:22:33.047 -> Water_Temp.:105.24, PWM:47
02:22:33.805 -> Water_Temp.:105.24, PWM:47
02:22:34.598 -> Water_Temp.:105.24, PWM:47
02:22:36.156 -> Water_Temp.:105.24, PWM:47
02:22:36.945 -> Water_Temp.:105.24, PWM:47
02:22:37.708 -> Water_Temp.:105.24, PWM:47
02:22:38.505 -> Water_Temp.:105.24, PWM:47
02:22:40.025 -> Water_Temp.:105.24, PWM:47
02:22:40.819 -> Water_Temp.:105.24, PWM:47
02:22:41.605 -> Water_Temp.:105.24, PWM:47
02:22:42.364 -> Water_Temp.:105.24, PWM:47
02:22:43.953 -> Water_Temp.:105.24, PWM:47
02:22:44.717 -> Water_Temp.:105.24, PWM:47
02:22:46.271 -> Water_Temp.:105.35, PWM:47
02:22:47.828 -> Water_Temp.:105.35, PWM:47
02:22:48.588 -> Water_Temp.:105.35, PWM:47
02:22:49.384 -> Water_Temp.:105.35, PWM:47
02:22:50.941 -> Water_Temp.:105.35, PWM:47
02:22:52.492 -> Water_Temp.:105.35, PWM:47
02:22:53.252 -> Water_Temp.:105.35, PWM:47
02:22:54.048 -> Water_Temp.:105.35, PWM:47
02:22:55.575 -> Water_Temp.:105.35, PWM:47
02:22:56.371 -> Water_Temp.:105.35, PWM:47
02:23:00.206 -> Water_Temp.:105.35, PWM:47
02:23:01.002 -> Water_Temp.:105.35, PWM:47
02:23:01.762 -> Water_Temp.:105.35, PWM:47
02:23:02.555 -> Water_Temp.:105.46, PWM:47
02:23:04.076 -> Water_Temp.:105.46, PWM:47
02:23:05.601 -> Water_Temp.:105.46, PWM:47
02:23:06.393 -> Water_Temp.:105.46, PWM:47
02:23:07.190 -> Water_Temp.:105.46, PWM:47
02:23:07.951 -> Water_Temp.:105.46, PWM:47
02:23:09.498 -> Water_Temp.:105.46, PWM:47
02:23:10.257 -> Water_Temp.:105.46, PWM:47
02:23:11.054 -> Water_Temp.:105.46, PWM:47
02:23:11.846 -> Water_Temp.:105.46, PWM:47
02:23:12.608 -> Water_Temp.:105.46, PWM:47
02:23:13.405 -> Water_Temp.:105.46, PWM:47
02:23:14.163 -> Water_Temp.:105.46, PWM:47
02:23:14.956 -> Water_Temp.:105.46, PWM:47
02:23:15.715 -> Water_Temp.:105.46, PWM:47
02:23:18.031 -> Water_Temp.:105.46, PWM:47
02:23:18.827 -> Water_Temp.:105.46, PWM:47
02:23:19.590 -> Water_Temp.:105.46, PWM:47
02:23:21.897 -> Water_Temp.:105.46, PWM:47
02:23:22.691 -> Water_Temp.:105.46, PWM:47
02:23:24.245 -> Water_Temp.:105.46, PWM:47
02:23:25.040 -> Water_Temp.:105.46, PWM:47
02:23:27.353 -> Water_Temp.:105.46, PWM:47
02:23:28.112 -> Water_Temp.:105.46, PWM:47
02:23:28.904 -> Water_Temp.:105.46, PWM:47
02:23:30.452 -> Water_Temp.:105.46, PWM:47
02:23:31.217 -> Water_Temp.:105.46, PWM:47
02:23:32.015 -> Water_Temp.:105.46, PWM:47
02:23:33.567 -> Water_Temp.:105.46, PWM:47
02:23:35.119 -> Water_Temp.:105.46, PWM:47
02:23:35.876 -> Water_Temp.:105.46, PWM:47
02:23:36.671 -> Water_Temp.:105.46, PWM:47
02:23:37.466 -> Water_Temp.:105.46, PWM:47
02:23:38.227 -> Water_Temp.:105.46, PWM:47
02:23:39.023 -> Water_Temp.:105.46, PWM:47
02:23:40.544 -> Water_Temp.:105.57, PWM:47
02:23:42.858 -> Water_Temp.:105.46, PWM:47
02:23:43.650 -> Water_Temp.:105.57, PWM:47
02:23:46.731 -> Water_Temp.:105.57, PWM:47
02:23:48.288 -> Water_Temp.:105.57, PWM:47
02:23:49.081 -> Water_Temp.:105.57, PWM:47
02:23:49.874 -> Water_Temp.:105.57, PWM:47
02:23:50.633 -> Water_Temp.:105.57, PWM:47
02:23:51.429 -> Water_Temp.:105.57, PWM:47
02:23:52.189 -> Water_Temp.:105.57, PWM:47
02:23:53.749 -> Water_Temp.:105.57, PWM:47
02:23:54.544 -> Water_Temp.:105.57, PWM:47
02:23:56.063 -> Water_Temp.:105.57, PWM:47
02:23:57.654 -> Water_Temp.:105.57, PWM:47
02:23:58.414 -> Water_Temp.:105.57, PWM:47
02:23:59.209 -> Water_Temp.:105.57, PWM:47
02:23:59.969 -> Water_Temp.:105.57, PWM:47
02:24:03.073 -> Water_Temp.:105.57, PWM:47
02:24:03.833 -> Water_Temp.:105.57, PWM:47
02:24:05.415 -> Water_Temp.:105.57, PWM:47
02:24:06.177 -> Water_Temp.:105.57, PWM:47
02:24:06.967 -> Water_Temp.:105.57, PWM:47
02:24:07.729 -> Water_Temp.:105.57, PWM:47
02:24:08.521 -> Water_Temp.:105.57, PWM:47
02:24:09.312 -> Water_Temp.:105.57, PWM:47
02:24:10.830 -> Water_Temp.:105.57, PWM:47
02:24:11.622 -> Water_Temp.:105.57, PWM:47
02:24:12.417 -> Water_Temp.:105.57, PWM:47
02:24:13.183 -> Water_Temp.:105.57, PWM:47
02:24:13.949 -> Water_Temp.:105.57, PWM:47
02:24:14.745 -> Water_Temp.:105.57, PWM:47
02:24:15.506 -> Water_Temp.:105.57, PWM:47
02:24:16.303 -> Water_Temp.:105.57, PWM:47
02:24:17.859 -> Water_Temp.:105.57, PWM:47
02:24:18.621 -> Water_Temp.:105.57, PWM:47
02:24:19.411 -> Water_Temp.:105.57, PWM:47
02:24:20.168 -> Water_Temp.:105.57, PWM:47
02:24:21.724 -> Water_Temp.:105.57, PWM:47
02:24:23.245 -> Water_Temp.:105.57, PWM:47
02:24:24.036 -> Water_Temp.:105.57, PWM:47
02:24:26.309 -> Water_Temp.:105.57, PWM:47
02:24:27.099 -> Water_Temp.:105.57, PWM:47
02:24:27.864 -> Water_Temp.:105.57, PWM:47
02:24:29.457 -> Water_Temp.:105.69, PWM:47
02:24:30.217 -> Water_Temp.:105.69, PWM:47
02:24:31.773 -> Water_Temp.:105.69, PWM:47
02:24:32.565 -> Water_Temp.:105.69, PWM:47
02:24:33.328 -> Water_Temp.:105.69, PWM:47
02:24:34.126 -> Water_Temp.:105.69, PWM:47
02:24:34.890 -> Water_Temp.:105.69, PWM:47
02:24:35.680 -> Water_Temp.:105.69, PWM:47
02:24:38.715 -> Water_Temp.:105.69, PWM:47
02:24:40.265 -> Water_Temp.:105.69, PWM:47
02:24:42.569 -> Water_Temp.:105.69, PWM:47
02:24:44.109 -> Water_Temp.:105.69, PWM:47
02:24:44.902 -> Water_Temp.:105.69, PWM:47
02:24:46.448 -> Water_Temp.:105.80, PWM:47
02:24:47.236 -> Water_Temp.:105.80, PWM:47
02:24:49.547 -> Water_Temp.:105.80, PWM:47
02:24:51.102 -> Water_Temp.:105.80, PWM:47
02:24:51.864 -> Water_Temp.:105.80, PWM:47
02:24:52.661 -> Water_Temp.:105.80, PWM:47
02:24:53.421 -> Water_Temp.:105.80, PWM:47
02:24:54.216 -> Water_Temp.:105.80, PWM:47
02:24:54.978 -> Water_Temp.:105.80, PWM:47
02:24:58.085 -> Water_Temp.:105.80, PWM:47
02:25:00.401 -> Water_Temp.:105.80, PWM:47
02:25:01.160 -> Water_Temp.:105.80, PWM:47
02:25:02.717 -> Water_Temp.:105.80, PWM:47
02:25:03.515 -> Water_Temp.:105.80, PWM:47
02:25:04.274 -> Water_Temp.:105.80, PWM:47
02:25:05.068 -> Water_Temp.:105.80, PWM:47
02:25:05.863 -> Water_Temp.:105.80, PWM:47
02:25:06.623 -> Water_Temp.:105.80, PWM:47
02:25:08.183 -> Water_Temp.:105.80, PWM:47
02:25:08.943 -> Water_Temp.:105.80, PWM:47
02:25:09.741 -> Water_Temp.:105.80, PWM:47
02:25:11.298 -> Water_Temp.:105.80, PWM:47
02:25:12.090 -> Water_Temp.:105.80, PWM:47
02:25:12.855 -> Water_Temp.:105.80, PWM:47
02:25:14.420 -> Water_Temp.:105.80, PWM:47
02:25:15.978 -> Water_Temp.:105.80, PWM:47
02:25:16.737 -> Water_Temp.:105.91, PWM:47
02:25:17.530 -> Water_Temp.:105.91, PWM:47
02:25:18.321 -> Water_Temp.:105.91, PWM:47
02:25:19.079 -> Water_Temp.:105.91, PWM:47
02:25:20.633 -> Water_Temp.:105.91, PWM:47
02:25:21.396 -> Water_Temp.:105.91, PWM:47
02:25:22.981 -> Water_Temp.:105.91, PWM:47
02:25:23.742 -> Water_Temp.:105.91, PWM:47
02:25:24.527 -> Water_Temp.:105.91, PWM:47
02:25:25.325 -> Water_Temp.:105.91, PWM:47
02:25:26.081 -> Water_Temp.:105.91, PWM:47
02:25:27.638 -> Water_Temp.:105.91, PWM:47
02:25:28.402 -> Water_Temp.:105.91, PWM:47
02:25:29.198 -> Water_Temp.:105.91, PWM:47
02:25:29.958 -> Water_Temp.:105.91, PWM:47
02:25:31.541 -> Water_Temp.:105.91, PWM:47
02:25:32.297 -> Water_Temp.:105.91, PWM:47
02:25:33.094 -> Water_Temp.:105.91, PWM:47
02:25:34.645 -> Water_Temp.:105.91, PWM:47
02:25:35.440 -> Water_Temp.:105.91, PWM:47
02:25:36.200 -> Water_Temp.:105.91, PWM:47
02:25:38.542 -> Water_Temp.:105.91, PWM:47
02:25:39.302 -> Water_Temp.:105.91, PWM:47
02:25:40.094 -> Water_Temp.:105.91, PWM:47
02:25:40.856 -> Water_Temp.:105.91, PWM:47
02:25:42.405 -> Water_Temp.:105.91, PWM:47
02:25:43.197 -> Water_Temp.:105.91, PWM:47
02:25:44.719 -> Water_Temp.:105.91, PWM:47
02:25:45.479 -> Water_Temp.:105.91, PWM:47
02:25:46.274 -> Water_Temp.:105.91, PWM:47
02:25:47.071 -> Water_Temp.:105.91, PWM:47
02:25:49.388 -> Water_Temp.:106.02, PWM:47
02:25:50.186 -> Water_Temp.:105.91, PWM:47
02:25:50.951 -> Water_Temp.:105.91, PWM:47
02:25:51.745 -> Water_Temp.:106.02, PWM:47
02:25:52.505 -> Water_Temp.:106.02, PWM:47
02:25:54.083 -> Water_Temp.:106.02, PWM:47
02:25:54.842 -> Water_Temp.:106.02, PWM:47
02:25:55.635 -> Water_Temp.:106.02, PWM:47
02:25:56.429 -> Water_Temp.:106.02, PWM:47
02:25:57.186 -> Water_Temp.:106.02, PWM:47
02:25:57.981 -> Water_Temp.:106.02, PWM:47
02:25:58.744 -> Water_Temp.:106.02, PWM:47
02:26:00.301 -> Water_Temp.:106.02, PWM:47
02:26:01.859 -> Water_Temp.:106.02, PWM:47
02:26:02.653 -> Water_Temp.:106.02, PWM:47
02:26:04.937 -> Water_Temp.:106.02, PWM:47
02:26:05.728 -> Water_Temp.:106.02, PWM:47
02:26:06.521 -> Water_Temp.:106.02, PWM:47
02:26:07.284 -> Water_Temp.:106.02, PWM:47
02:26:09.597 -> Water_Temp.:106.02, PWM:47
02:26:10.359 -> Water_Temp.:106.02, PWM:47
02:26:11.153 -> Water_Temp.:106.02, PWM:47
02:26:11.915 -> Water_Temp.:106.02, PWM:47
02:26:13.473 -> Water_Temp.:106.02, PWM:47
02:26:14.266 -> Water_Temp.:106.14, PWM:47
02:26:15.029 -> Water_Temp.:106.02, PWM:47
02:26:15.825 -> Water_Temp.:106.02, PWM:47
02:26:17.380 -> Water_Temp.:106.02, PWM:47
02:26:18.934 -> Water_Temp.:106.02, PWM:47
02:26:19.727 -> Water_Temp.:106.02, PWM:47
02:26:20.523 -> Water_Temp.:106.02, PWM:47
02:26:21.288 -> Water_Temp.:106.02, PWM:47
02:26:22.046 -> Water_Temp.:106.02, PWM:47
02:26:22.843 -> Water_Temp.:106.02, PWM:47
02:26:23.637 -> Water_Temp.:106.02, PWM:47
02:26:24.397 -> Water_Temp.:106.02, PWM:47
02:26:25.194 -> Water_Temp.:106.14, PWM:47
02:26:25.958 -> Water_Temp.:106.02, PWM:47
02:26:28.271 -> Water_Temp.:106.02, PWM:47
02:26:29.827 -> Water_Temp.:106.14, PWM:47
02:26:30.622 -> Water_Temp.:106.14, PWM:47
02:26:32.176 -> Water_Temp.:106.14, PWM:47
02:26:33.734 -> Water_Temp.:106.14, PWM:47
02:26:35.256 -> Water_Temp.:106.14, PWM:47
02:26:36.051 -> Water_Temp.:106.14, PWM:47
02:26:36.811 -> Water_Temp.:106.14, PWM:47
02:26:39.162 -> Water_Temp.:106.14, PWM:47
02:26:39.927 -> Water_Temp.:106.14, PWM:47
02:26:41.487 -> Water_Temp.:106.14, PWM:47
02:26:42.250 -> Water_Temp.:106.14, PWM:47
02:26:43.806 -> Water_Temp.:106.14, PWM:47
02:26:45.366 -> Water_Temp.:106.14, PWM:47
02:26:46.164 -> Water_Temp.:106.14, PWM:47
02:26:46.925 -> Water_Temp.:106.14, PWM:47
02:26:48.508 -> Water_Temp.:106.14, PWM:47
02:26:49.266 -> Water_Temp.:106.14, PWM:47
02:26:50.056 -> Water_Temp.:106.14, PWM:47
02:26:51.612 -> Water_Temp.:106.14, PWM:47
02:26:53.166 -> Water_Temp.:106.14, PWM:47
02:26:54.685 -> Water_Temp.:106.14, PWM:47
02:26:56.248 -> Water_Temp.:106.14, PWM:47
02:27:00.907 -> Water_Temp.:106.14, PWM:47
02:27:03.224 -> Water_Temp.:106.14, PWM:47
02:27:03.985 -> Water_Temp.:106.14, PWM:47
02:27:04.781 -> Water_Temp.:106.14, PWM:47
02:27:05.542 -> Water_Temp.:106.14, PWM:47
02:27:07.127 -> Water_Temp.:106.14, PWM:47
02:27:07.886 -> Water_Temp.:106.25, PWM:47
02:27:08.683 -> Water_Temp.:106.25, PWM:47
02:27:10.201 -> Water_Temp.:106.25, PWM:47
02:27:10.999 -> Water_Temp.:106.25, PWM:47
02:27:11.761 -> Water_Temp.:106.25, PWM:47
02:27:12.553 -> Water_Temp.:106.25, PWM:47
02:27:13.315 -> Water_Temp.:106.25, PWM:47
02:27:14.112 -> Water_Temp.:106.25, PWM:47
02:27:14.873 -> Water_Temp.:106.25, PWM:47
02:27:16.422 -> Water_Temp.:106.25, PWM:47
02:27:17.214 -> Water_Temp.:106.25, PWM:47
02:27:18.004 -> Water_Temp.:106.25, PWM:47
02:27:18.761 -> Water_Temp.:106.25, PWM:47
02:27:19.552 -> Water_Temp.:106.14, PWM:47
02:27:20.343 -> Water_Temp.:106.14, PWM:47
02:27:21.102 -> Water_Temp.:106.25, PWM:47
02:27:21.895 -> Water_Temp.:106.14, PWM:47
02:27:22.655 -> Water_Temp.:106.25, PWM:47
02:27:24.212 -> Water_Temp.:106.25, PWM:47
02:27:25.006 -> Water_Temp.:106.25, PWM:47
02:27:25.767 -> Water_Temp.:106.25, PWM:47
02:27:26.557 -> Water_Temp.:106.25, PWM:47
02:27:28.109 -> Water_Temp.:106.25, PWM:47
02:27:28.905 -> Water_Temp.:106.25, PWM:0
02:27:29.664 -> Water_Temp.:106.25, PWM:0
02:27:30.456 -> Water_Temp.:106.25, PWM:0
02:27:31.218 -> Water_Temp.:106.25, PWM:0
02:27:32.008 -> Water_Temp.:106.25, PWM:0
02:27:32.802 -> Water_Temp.:106.25, PWM:0
02:27:33.561 -> Water_Temp.:106.25, PWM:0
02:27:34.355 -> Water_Temp.:106.25, PWM:0
02:27:35.147 -> Water_Temp.:106.25, PWM:0
02:27:35.905 -> Water_Temp.:106.25, PWM:0
02:27:36.701 -> Water_Temp.:106.25, PWM:0
02:27:37.464 -> Water_Temp.:106.25, PWM:0
02:27:38.252 -> Water_Temp.:106.25, PWM:0
02:27:39.045 -> Water_Temp.:106.25, PWM:0
02:27:39.803 -> Water_Temp.:106.25, PWM:0
02:27:40.599 -> Water_Temp.:106.25, PWM:0
02:27:41.360 -> Water_Temp.:106.25, PWM:0
02:27:42.150 -> Water_Temp.:106.25, PWM:0
02:27:42.943 -> Water_Temp.:106.25, PWM:0
02:27:43.705 -> Water_Temp.:106.25, PWM:0
02:27:44.500 -> Water_Temp.:106.36, PWM:0
02:27:45.261 -> Water_Temp.:106.25, PWM:0
02:27:46.056 -> Water_Temp.:106.36, PWM:0
02:27:46.814 -> Water_Temp.:106.36, PWM:0
02:27:47.607 -> Water_Temp.:106.36, PWM:0
02:27:48.400 -> Water_Temp.:106.36, PWM:0
02:27:49.161 -> Water_Temp.:106.36, PWM:0
02:27:49.951 -> Water_Temp.:106.36, PWM:0
02:27:50.710 -> Water_Temp.:106.36, PWM:0
02:27:51.503 -> Water_Temp.:106.36, PWM:0
02:27:52.301 -> Water_Temp.:106.36, PWM:0
02:27:53.059 -> Water_Temp.:106.36, PWM:0
02:27:53.850 -> Water_Temp.:106.36, PWM:0
02:27:54.614 -> Water_Temp.:106.36, PWM:0
02:27:55.404 -> Water_Temp.:106.36, PWM:0
02:27:56.200 -> Water_Temp.:106.36, PWM:0
02:27:56.962 -> Water_Temp.:106.36, PWM:0
02:27:57.760 -> Water_Temp.:106.36, PWM:0
02:27:58.518 -> Water_Temp.:106.36, PWM:0
02:27:59.313 -> Water_Temp.:106.36, PWM:0
02:28:00.103 -> Water_Temp.:106.36, PWM:0
02:28:00.864 -> Water_Temp.:106.36, PWM:0
02:28:01.653 -> Water_Temp.:106.36, PWM:0
02:28:02.415 -> Water_Temp.:106.36, PWM:0
02:28:03.209 -> Water_Temp.:106.36, PWM:0
02:28:03.972 -> Water_Temp.:106.36, PWM:0
02:28:04.767 -> Water_Temp.:106.36, PWM:0
02:28:05.561 -> Water_Temp.:106.36, PWM:0
02:28:06.323 -> Water_Temp.:106.36, PWM:0
02:28:07.115 -> Water_Temp.:106.36, PWM:0
02:28:07.874 -> Water_Temp.:106.36, PWM:0
02:28:08.669 -> Water_Temp.:106.36, PWM:0
02:28:09.429 -> Water_Temp.:106.36, PWM:0
02:28:10.226 -> Water_Temp.:106.36, PWM:0
02:28:11.020 -> Water_Temp.:106.36, PWM:0
02:28:11.775 -> Water_Temp.:106.36, PWM:0
02:28:12.567 -> Water_Temp.:106.36, PWM:0
02:28:13.359 -> Water_Temp.:106.36, PWM:0
02:28:14.108 -> Water_Temp.:106.36, PWM:0
02:28:14.902 -> Water_Temp.:106.36, PWM:0
02:28:15.690 -> Water_Temp.:106.36, PWM:0
02:28:16.451 -> Water_Temp.:106.25, PWM:0
02:28:17.244 -> Water_Temp.:106.25, PWM:0
02:28:18.032 -> Water_Temp.:106.25, PWM:0
02:28:18.791 -> Water_Temp.:106.25, PWM:0
02:28:19.586 -> Water_Temp.:106.25, PWM:0
02:28:20.347 -> Water_Temp.:106.25, PWM:0
02:28:21.143 -> Water_Temp.:106.25, PWM:0
02:28:21.933 -> Water_Temp.:106.25, PWM:0
02:28:22.693 -> Water_Temp.:106.25, PWM:0
02:28:23.484 -> Water_Temp.:106.25, PWM:0
02:28:24.277 -> Water_Temp.:106.25, PWM:0
02:28:25.037 -> Water_Temp.:106.25, PWM:0
02:28:25.836 -> Water_Temp.:106.25, PWM:0
02:28:26.597 -> Water_Temp.:106.25, PWM:0
02:28:27.386 -> Water_Temp.:106.25, PWM:0
02:28:28.178 -> Water_Temp.:106.25, PWM:0
02:28:28.938 -> Water_Temp.:106.25, PWM:0
02:28:29.725 -> Water_Temp.:106.25, PWM:0
02:28:30.515 -> Water_Temp.:106.25, PWM:0
02:28:31.270 -> Water_Temp.:106.25, PWM:0
02:28:32.056 -> Water_Temp.:106.25, PWM:0

## RAW_CAL_MLX
00:55:00.969 -> Average Cal sensor Temp.: 90.40 Ambient Temp.: 82.67 F
00:55:02.522 -> Average Cal sensor Temp.: 90.49 Ambient Temp.: 82.63 F
00:55:04.081 -> Average Cal sensor Temp.: 90.34 Ambient Temp.: 82.71 F
00:55:05.640 -> Average Cal sensor Temp.: 90.48 Ambient Temp.: 82.67 F
00:55:07.187 -> Average Cal sensor Temp.: 90.61 Ambient Temp.: 82.74 F
00:55:08.740 -> Average Cal sensor Temp.: 90.44 Ambient Temp.: 82.74 F
00:55:10.325 -> Average Cal sensor Temp.: 90.35 Ambient Temp.: 82.85 F
00:55:11.880 -> Average Cal sensor Temp.: 90.38 Ambient Temp.: 82.74 F
00:55:13.406 -> Average Cal sensor Temp.: 90.48 Ambient Temp.: 82.78 F
00:55:14.994 -> Average Cal sensor Temp.: 90.33 Ambient Temp.: 82.78 F
00:55:16.549 -> Average Cal sensor Temp.: 89.99 Ambient Temp.: 82.85 F
00:55:18.106 -> Average Cal sensor Temp.: 90.00 Ambient Temp.: 82.92 F
00:55:19.664 -> Average Cal sensor Temp.: 89.97 Ambient Temp.: 82.85 F
00:55:21.218 -> Average Cal sensor Temp.: 90.17 Ambient Temp.: 82.89 F
00:55:22.778 -> Average Cal sensor Temp.: 90.04 Ambient Temp.: 82.78 F
00:55:24.337 -> Average Cal sensor Temp.: 89.82 Ambient Temp.: 82.85 F
00:55:25.891 -> Average Cal sensor Temp.: 89.96 Ambient Temp.: 82.89 F
00:55:27.449 -> Average Cal sensor Temp.: 90.06 Ambient Temp.: 82.89 F
00:55:29.004 -> Average Cal sensor Temp.: 90.16 Ambient Temp.: 82.89 F
00:55:30.562 -> Average Cal sensor Temp.: 90.22 Ambient Temp.: 82.89 F
00:55:32.111 -> Average Cal sensor Temp.: 90.07 Ambient Temp.: 82.89 F
00:55:33.668 -> Average Cal sensor Temp.: 90.16 Ambient Temp.: 82.89 F
00:55:35.219 -> Average Cal sensor Temp.: 90.20 Ambient Temp.: 82.85 F
00:55:36.775 -> Average Cal sensor Temp.: 90.05 Ambient Temp.: 82.85 F
00:55:38.327 -> Average Cal sensor Temp.: 89.88 Ambient Temp.: 82.85 F
00:55:39.909 -> Average Cal sensor Temp.: 89.98 Ambient Temp.: 82.78 F
00:55:41.465 -> Average Cal sensor Temp.: 89.77 Ambient Temp.: 82.89 F
00:55:43.022 -> Average Cal sensor Temp.: 89.76 Ambient Temp.: 82.78 F
00:55:44.578 -> Average Cal sensor Temp.: 89.87 Ambient Temp.: 82.78 F
00:55:46.134 -> Average Cal sensor Temp.: 89.55 Ambient Temp.: 82.85 F
00:55:47.693 -> Average Cal sensor Temp.: 89.51 Ambient Temp.: 82.74 F
00:55:49.251 -> Average Cal sensor Temp.: 89.91 Ambient Temp.: 82.74 F
00:55:50.803 -> Average Cal sensor Temp.: 89.95 Ambient Temp.: 82.71 F
00:55:52.352 -> Average Cal sensor Temp.: 90.47 Ambient Temp.: 82.71 F
00:55:53.907 -> Average Cal sensor Temp.: 90.41 Ambient Temp.: 82.74 F
00:55:55.459 -> Average Cal sensor Temp.: 90.49 Ambient Temp.: 82.71 F
00:55:57.009 -> Average Cal sensor Temp.: 90.27 Ambient Temp.: 82.71 F
00:55:58.563 -> Average Cal sensor Temp.: 90.15 Ambient Temp.: 82.67 F
00:56:00.117 -> Average Cal sensor Temp.: 90.54 Ambient Temp.: 82.63 F
00:56:01.674 -> Average Cal sensor Temp.: 90.04 Ambient Temp.: 82.60 F
00:56:03.237 -> Average Cal sensor Temp.: 90.18 Ambient Temp.: 82.60 F
00:56:04.788 -> Average Cal sensor Temp.: 90.21 Ambient Temp.: 82.63 F
00:56:06.342 -> Average Cal sensor Temp.: 90.25 Ambient Temp.: 82.63 F
00:56:07.896 -> Average Cal sensor Temp.: 90.54 Ambient Temp.: 82.67 F
00:56:09.456 -> Average Cal sensor Temp.: 91.02 Ambient Temp.: 82.63 F
00:56:11.017 -> Average Cal sensor Temp.: 90.34 Ambient Temp.: 82.63 F
00:56:12.566 -> Average Cal sensor Temp.: 90.41 Ambient Temp.: 82.67 F
00:56:14.123 -> Average Cal sensor Temp.: 90.34 Ambient Temp.: 82.67 F
00:56:15.680 -> Average Cal sensor Temp.: 90.18 Ambient Temp.: 82.63 F
00:56:17.262 -> Average Cal sensor Temp.: 89.99 Ambient Temp.: 82.60 F
00:56:18.809 -> Average Cal sensor Temp.: 89.92 Ambient Temp.: 82.63 F
00:56:20.360 -> Average Cal sensor Temp.: 89.81 Ambient Temp.: 82.56 F
00:56:21.916 -> Average Cal sensor Temp.: 89.76 Ambient Temp.: 82.60 F
00:56:23.471 -> Average Cal sensor Temp.: 89.75 Ambient Temp.: 82.56 F
00:56:25.025 -> Average Cal sensor Temp.: 90.12 Ambient Temp.: 82.53 F
00:56:26.583 -> Average Cal sensor Temp.: 89.73 Ambient Temp.: 82.49 F
00:56:28.135 -> Average Cal sensor Temp.: 89.57 Ambient Temp.: 82.49 F
00:56:29.685 -> Average Cal sensor Temp.: 90.15 Ambient Temp.: 82.45 F
00:56:31.272 -> Average Cal sensor Temp.: 90.17 Ambient Temp.: 82.45 F
00:56:32.820 -> Average Cal sensor Temp.: 89.65 Ambient Temp.: 82.45 F
00:56:34.375 -> Average Cal sensor Temp.: 89.86 Ambient Temp.: 82.38 F
00:56:35.923 -> Average Cal sensor Temp.: 90.03 Ambient Temp.: 82.38 F
00:56:37.483 -> Average Cal sensor Temp.: 90.03 Ambient Temp.: 82.31 F
00:56:39.039 -> Average Cal sensor Temp.: 89.42 Ambient Temp.: 82.31 F
00:56:40.600 -> Average Cal sensor Temp.: 89.70 Ambient Temp.: 82.27 F
00:56:42.151 -> Average Cal sensor Temp.: 90.47 Ambient Temp.: 82.24 F
00:56:43.709 -> Average Cal sensor Temp.: 90.48 Ambient Temp.: 82.24 F
00:56:45.268 -> Average Cal sensor Temp.: 90.43 Ambient Temp.: 82.20 F
00:56:46.824 -> Average Cal sensor Temp.: 90.77 Ambient Temp.: 82.20 F
00:56:48.378 -> Average Cal sensor Temp.: 91.13 Ambient Temp.: 82.27 F
00:56:49.928 -> Average Cal sensor Temp.: 91.24 Ambient Temp.: 82.24 F
00:56:51.489 -> Average Cal sensor Temp.: 91.11 Ambient Temp.: 82.24 F
00:56:53.049 -> Average Cal sensor Temp.: 90.62 Ambient Temp.: 82.20 F
00:56:54.594 -> Average Cal sensor Temp.: 90.57 Ambient Temp.: 82.24 F
00:56:56.146 -> Average Cal sensor Temp.: 90.40 Ambient Temp.: 82.20 F
00:56:57.701 -> Average Cal sensor Temp.: 90.72 Ambient Temp.: 82.20 F
00:56:59.259 -> Average Cal sensor Temp.: 90.68 Ambient Temp.: 82.20 F
00:57:00.820 -> Average Cal sensor Temp.: 90.35 Ambient Temp.: 82.20 F
00:57:02.377 -> Average Cal sensor Temp.: 90.44 Ambient Temp.: 82.24 F
00:57:03.935 -> Average Cal sensor Temp.: 90.50 Ambient Temp.: 82.24 F
00:57:05.487 -> Average Cal sensor Temp.: 90.32 Ambient Temp.: 82.27 F
00:57:07.047 -> Average Cal sensor Temp.: 90.50 Ambient Temp.: 82.24 F
00:57:08.603 -> Average Cal sensor Temp.: 90.49 Ambient Temp.: 82.24 F
00:57:10.160 -> Average Cal sensor Temp.: 90.25 Ambient Temp.: 82.24 F
00:57:11.714 -> Average Cal sensor Temp.: 90.28 Ambient Temp.: 82.24 F
00:57:13.270 -> Average Cal sensor Temp.: 90.70 Ambient Temp.: 82.27 F
00:57:14.820 -> Average Cal sensor Temp.: 90.76 Ambient Temp.: 82.27 F
00:57:16.377 -> Average Cal sensor Temp.: 90.78 Ambient Temp.: 82.35 F
00:57:17.939 -> Average Cal sensor Temp.: 90.55 Ambient Temp.: 82.27 F
00:57:19.499 -> Average Cal sensor Temp.: 90.65 Ambient Temp.: 82.35 F
00:57:21.056 -> Average Cal sensor Temp.: 90.82 Ambient Temp.: 82.38 F
00:57:22.607 -> Average Cal sensor Temp.: 90.60 Ambient Temp.: 82.35 F
00:57:24.164 -> Average Cal sensor Temp.: 90.58 Ambient Temp.: 82.35 F
00:57:25.721 -> Average Cal sensor Temp.: 90.64 Ambient Temp.: 82.45 F
00:57:27.282 -> Average Cal sensor Temp.: 90.56 Ambient Temp.: 82.45 F
00:57:28.839 -> Average Cal sensor Temp.: 90.67 Ambient Temp.: 82.45 F
00:57:30.398 -> Average Cal sensor Temp.: 90.28 Ambient Temp.: 82.45 F
00:57:31.956 -> Average Cal sensor Temp.: 90.55 Ambient Temp.: 82.38 F
00:57:33.514 -> Average Cal sensor Temp.: 90.34 Ambient Temp.: 82.45 F
00:57:35.073 -> Average Cal sensor Temp.: 90.28 Ambient Temp.: 82.45 F
00:57:36.627 -> Average Cal sensor Temp.: 90.51 Ambient Temp.: 82.49 F
00:57:38.187 -> Average Cal sensor Temp.: 90.30 Ambient Temp.: 82.45 F
00:57:39.740 -> Average Cal sensor Temp.: 90.44 Ambient Temp.: 82.49 F
00:57:41.301 -> Average Cal sensor Temp.: 90.16 Ambient Temp.: 82.49 F
00:57:42.860 -> Average Cal sensor Temp.: 90.49 Ambient Temp.: 82.45 F
00:57:44.413 -> Average Cal sensor Temp.: 90.25 Ambient Temp.: 82.53 F
00:57:45.968 -> Average Cal sensor Temp.: 90.17 Ambient Temp.: 82.45 F
00:57:47.524 -> Average Cal sensor Temp.: 90.37 Ambient Temp.: 82.49 F
00:57:49.080 -> Average Cal sensor Temp.: 90.13 Ambient Temp.: 82.56 F
00:57:50.636 -> Average Cal sensor Temp.: 90.05 Ambient Temp.: 82.49 F
00:57:52.187 -> Average Cal sensor Temp.: 90.58 Ambient Temp.: 82.53 F
00:57:53.737 -> Average Cal sensor Temp.: 90.22 Ambient Temp.: 82.53 F
00:57:55.294 -> Average Cal sensor Temp.: 90.23 Ambient Temp.: 82.53 F
00:57:56.851 -> Average Cal sensor Temp.: 89.91 Ambient Temp.: 82.49 F
00:57:58.402 -> Average Cal sensor Temp.: 89.82 Ambient Temp.: 82.49 F
00:57:59.957 -> Average Cal sensor Temp.: 90.25 Ambient Temp.: 82.45 F
00:58:01.516 -> Average Cal sensor Temp.: 90.32 Ambient Temp.: 82.45 F
00:58:03.072 -> Average Cal sensor Temp.: 90.35 Ambient Temp.: 82.49 F
00:58:04.625 -> Average Cal sensor Temp.: 90.47 Ambient Temp.: 82.49 F
00:58:06.182 -> Average Cal sensor Temp.: 90.54 Ambient Temp.: 82.49 F
00:58:07.734 -> Average Cal sensor Temp.: 90.72 Ambient Temp.: 82.53 F
00:58:09.317 -> Average Cal sensor Temp.: 90.69 Ambient Temp.: 82.53 F
00:58:10.845 -> Average Cal sensor Temp.: 90.72 Ambient Temp.: 82.60 F
00:58:12.404 -> Average Cal sensor Temp.: 90.46 Ambient Temp.: 82.56 F
00:58:13.958 -> Average Cal sensor Temp.: 90.53 Ambient Temp.: 82.63 F
00:58:15.518 -> Average Cal sensor Temp.: 90.63 Ambient Temp.: 82.60 F
00:58:17.075 -> Average Cal sensor Temp.: 90.64 Ambient Temp.: 82.63 F
00:58:18.630 -> Average Cal sensor Temp.: 90.73 Ambient Temp.: 82.63 F
00:58:20.188 -> Average Cal sensor Temp.: 90.89 Ambient Temp.: 82.63 F
00:58:21.747 -> Average Cal sensor Temp.: 91.07 Ambient Temp.: 82.74 F
00:58:23.304 -> Average Cal sensor Temp.: 90.60 Ambient Temp.: 82.67 F
00:58:24.856 -> Average Cal sensor Temp.: 90.78 Ambient Temp.: 82.67 F
00:58:26.444 -> Average Cal sensor Temp.: 90.54 Ambient Temp.: 82.74 F
00:58:28.002 -> Average Cal sensor Temp.: 90.78 Ambient Temp.: 82.78 F
00:58:29.526 -> Average Cal sensor Temp.: 90.63 Ambient Temp.: 82.78 F
00:58:31.084 -> Average Cal sensor Temp.: 90.94 Ambient Temp.: 82.74 F
00:58:32.640 -> Average Cal sensor Temp.: 90.59 Ambient Temp.: 82.78 F
00:58:34.201 -> Average Cal sensor Temp.: 90.85 Ambient Temp.: 82.85 F
00:58:35.754 -> Average Cal sensor Temp.: 90.80 Ambient Temp.: 82.78 F
00:58:37.314 -> Average Cal sensor Temp.: 90.89 Ambient Temp.: 82.89 F
00:58:38.901 -> Average Cal sensor Temp.: 90.81 Ambient Temp.: 82.85 F
00:58:40.459 -> Average Cal sensor Temp.: 90.46 Ambient Temp.: 82.92 F
00:58:42.016 -> Average Cal sensor Temp.: 90.66 Ambient Temp.: 82.89 F
00:58:43.572 -> Average Cal sensor Temp.: 90.89 Ambient Temp.: 82.89 F
00:58:45.126 -> Average Cal sensor Temp.: 91.15 Ambient Temp.: 82.92 F
00:58:46.684 -> Average Cal sensor Temp.: 91.11 Ambient Temp.: 82.92 F
00:58:48.240 -> Average Cal sensor Temp.: 91.18 Ambient Temp.: 82.92 F
00:58:49.764 -> Average Cal sensor Temp.: 90.92 Ambient Temp.: 82.96 F
00:58:51.321 -> Average Cal sensor Temp.: 90.76 Ambient Temp.: 82.92 F
00:58:52.905 -> Average Cal sensor Temp.: 90.92 Ambient Temp.: 82.92 F
00:58:54.457 -> Average Cal sensor Temp.: 91.16 Ambient Temp.: 82.96 F
00:58:56.009 -> Average Cal sensor Temp.: 91.08 Ambient Temp.: 83.03 F
00:58:57.566 -> Average Cal sensor Temp.: 91.26 Ambient Temp.: 82.99 F
00:58:59.119 -> Average Cal sensor Temp.: 91.13 Ambient Temp.: 82.96 F
00:59:00.677 -> Average Cal sensor Temp.: 90.97 Ambient Temp.: 82.99 F
00:59:02.230 -> Average Cal sensor Temp.: 91.04 Ambient Temp.: 83.03 F
00:59:03.787 -> Average Cal sensor Temp.: 91.18 Ambient Temp.: 83.07 F
00:59:05.339 -> Average Cal sensor Temp.: 91.28 Ambient Temp.: 83.07 F
00:59:06.892 -> Average Cal sensor Temp.: 91.44 Ambient Temp.: 83.07 F
00:59:08.450 -> Average Cal sensor Temp.: 91.88 Ambient Temp.: 83.14 F
00:59:10.008 -> Average Cal sensor Temp.: 91.32 Ambient Temp.: 83.14 F
00:59:11.563 -> Average Cal sensor Temp.: 91.31 Ambient Temp.: 83.25 F
00:59:13.123 -> Average Cal sensor Temp.: 91.03 Ambient Temp.: 83.25 F
00:59:14.679 -> Average Cal sensor Temp.: 91.00 Ambient Temp.: 83.28 F
00:59:16.237 -> Average Cal sensor Temp.: 91.03 Ambient Temp.: 83.28 F
00:59:17.792 -> Average Cal sensor Temp.: 91.08 Ambient Temp.: 83.32 F
00:59:19.350 -> Average Cal sensor Temp.: 91.06 Ambient Temp.: 83.32 F
00:59:20.902 -> Average Cal sensor Temp.: 91.06 Ambient Temp.: 83.35 F
00:59:22.457 -> Average Cal sensor Temp.: 91.13 Ambient Temp.: 83.35 F
00:59:24.013 -> Average Cal sensor Temp.: 91.08 Ambient Temp.: 83.35 F
00:59:25.565 -> Average Cal sensor Temp.: 91.12 Ambient Temp.: 83.43 F
00:59:27.125 -> Average Cal sensor Temp.: 91.21 Ambient Temp.: 83.50 F
00:59:28.680 -> Average Cal sensor Temp.: 90.98 Ambient Temp.: 83.53 F
00:59:30.239 -> Average Cal sensor Temp.: 90.52 Ambient Temp.: 83.50 F
00:59:31.764 -> Average Cal sensor Temp.: 90.75 Ambient Temp.: 83.53 F
00:59:33.326 -> Average Cal sensor Temp.: 91.04 Ambient Temp.: 83.53 F
00:59:34.882 -> Average Cal sensor Temp.: 90.91 Ambient Temp.: 83.64 F
00:59:36.441 -> Average Cal sensor Temp.: 91.17 Ambient Temp.: 83.57 F
00:59:37.995 -> Average Cal sensor Temp.: 90.73 Ambient Temp.: 83.57 F
00:59:39.548 -> Average Cal sensor Temp.: 90.95 Ambient Temp.: 83.64 F
00:59:41.107 -> Average Cal sensor Temp.: 90.95 Ambient Temp.: 83.64 F
00:59:42.664 -> Average Cal sensor Temp.: 90.84 Ambient Temp.: 83.71 F
00:59:44.220 -> Average Cal sensor Temp.: 90.86 Ambient Temp.: 83.68 F
00:59:45.779 -> Average Cal sensor Temp.: 91.08 Ambient Temp.: 83.68 F
00:59:47.339 -> Average Cal sensor Temp.: 91.30 Ambient Temp.: 83.71 F
00:59:48.895 -> Average Cal sensor Temp.: 91.29 Ambient Temp.: 83.79 F
00:59:50.455 -> Average Cal sensor Temp.: 91.13 Ambient Temp.: 83.79 F
00:59:52.014 -> Average Cal sensor Temp.: 90.83 Ambient Temp.: 83.79 F
00:59:53.569 -> Average Cal sensor Temp.: 90.77 Ambient Temp.: 83.79 F
00:59:55.124 -> Average Cal sensor Temp.: 91.22 Ambient Temp.: 83.82 F
00:59:56.681 -> Average Cal sensor Temp.: 91.07 Ambient Temp.: 83.82 F
00:59:58.235 -> Average Cal sensor Temp.: 90.93 Ambient Temp.: 83.86 F
00:59:59.786 -> Average Cal sensor Temp.: 90.91 Ambient Temp.: 83.86 F
01:00:01.340 -> Average Cal sensor Temp.: 91.10 Ambient Temp.: 83.86 F
01:00:02.895 -> Average Cal sensor Temp.: 91.33 Ambient Temp.: 83.86 F
01:00:04.450 -> Average Cal sensor Temp.: 91.35 Ambient Temp.: 83.86 F
01:00:06.006 -> Average Cal sensor Temp.: 91.05 Ambient Temp.: 83.89 F
01:00:07.561 -> Average Cal sensor Temp.: 91.04 Ambient Temp.: 83.93 F
01:00:09.115 -> Average Cal sensor Temp.: 91.32 Ambient Temp.: 83.89 F
01:00:10.673 -> Average Cal sensor Temp.: 91.57 Ambient Temp.: 83.97 F
01:00:12.231 -> Average Cal sensor Temp.: 91.13 Ambient Temp.: 83.97 F
01:00:13.790 -> Average Cal sensor Temp.: 91.04 Ambient Temp.: 83.97 F
01:00:15.337 -> Average Cal sensor Temp.: 91.24 Ambient Temp.: 84.00 F
01:00:16.884 -> Average Cal sensor Temp.: 91.21 Ambient Temp.: 84.00 F
01:00:18.440 -> Average Cal sensor Temp.: 91.02 Ambient Temp.: 84.00 F
01:00:19.996 -> Average Cal sensor Temp.: 90.84 Ambient Temp.: 84.11 F
01:00:21.549 -> Average Cal sensor Temp.: 91.11 Ambient Temp.: 84.07 F
01:00:23.112 -> Average Cal sensor Temp.: 90.99 Ambient Temp.: 84.15 F
01:00:24.691 -> Average Cal sensor Temp.: 90.83 Ambient Temp.: 84.15 F
01:00:26.220 -> Average Cal sensor Temp.: 91.09 Ambient Temp.: 84.11 F
01:00:27.774 -> Average Cal sensor Temp.: 91.01 Ambient Temp.: 84.18 F
01:00:29.332 -> Average Cal sensor Temp.: 91.18 Ambient Temp.: 84.15 F
01:00:30.890 -> Average Cal sensor Temp.: 91.34 Ambient Temp.: 84.18 F
01:00:32.443 -> Average Cal sensor Temp.: 91.51 Ambient Temp.: 84.18 F
01:00:34.003 -> Average Cal sensor Temp.: 91.31 Ambient Temp.: 84.18 F
01:00:35.556 -> Average Cal sensor Temp.: 91.36 Ambient Temp.: 84.25 F
01:00:37.110 -> Average Cal sensor Temp.: 91.24 Ambient Temp.: 84.25 F
01:00:38.663 -> Average Cal sensor Temp.: 91.23 Ambient Temp.: 84.29 F
01:00:40.220 -> Average Cal sensor Temp.: 91.20 Ambient Temp.: 84.25 F
01:00:41.772 -> Average Cal sensor Temp.: 91.09 Ambient Temp.: 84.29 F
01:00:43.332 -> Average Cal sensor Temp.: 91.13 Ambient Temp.: 84.33 F
01:00:44.888 -> Average Cal sensor Temp.: 91.16 Ambient Temp.: 84.25 F
01:00:46.448 -> Average Cal sensor Temp.: 91.04 Ambient Temp.: 84.29 F
01:00:47.999 -> Average Cal sensor Temp.: 91.39 Ambient Temp.: 84.33 F
01:00:49.556 -> Average Cal sensor Temp.: 91.46 Ambient Temp.: 84.33 F
01:00:51.116 -> Average Cal sensor Temp.: 91.62 Ambient Temp.: 84.36 F
01:00:52.673 -> Average Cal sensor Temp.: 91.32 Ambient Temp.: 84.36 F
01:00:54.229 -> Average Cal sensor Temp.: 91.49 Ambient Temp.: 84.36 F
01:00:55.777 -> Average Cal sensor Temp.: 91.49 Ambient Temp.: 84.51 F
01:00:57.335 -> Average Cal sensor Temp.: 91.35 Ambient Temp.: 84.47 F
01:00:58.889 -> Average Cal sensor Temp.: 91.49 Ambient Temp.: 84.40 F
01:01:00.444 -> Average Cal sensor Temp.: 91.41 Ambient Temp.: 84.51 F
01:01:01.998 -> Average Cal sensor Temp.: 91.74 Ambient Temp.: 84.54 F
01:01:03.548 -> Average Cal sensor Temp.: 91.81 Ambient Temp.: 84.54 F
01:01:05.105 -> Average Cal sensor Temp.: 91.53 Ambient Temp.: 84.54 F
01:01:06.660 -> Average Cal sensor Temp.: 91.62 Ambient Temp.: 84.54 F
01:01:08.217 -> Average Cal sensor Temp.: 91.31 Ambient Temp.: 84.58 F
01:01:09.775 -> Average Cal sensor Temp.: 91.30 Ambient Temp.: 84.58 F
01:01:11.333 -> Average Cal sensor Temp.: 91.43 Ambient Temp.: 84.65 F
01:01:12.893 -> Average Cal sensor Temp.: 90.98 Ambient Temp.: 84.58 F
01:01:14.450 -> Average Cal sensor Temp.: 91.42 Ambient Temp.: 84.61 F
01:01:16.008 -> Average Cal sensor Temp.: 91.54 Ambient Temp.: 84.61 F
01:01:17.567 -> Average Cal sensor Temp.: 91.29 Ambient Temp.: 84.69 F
01:01:19.123 -> Average Cal sensor Temp.: 90.97 Ambient Temp.: 84.65 F
01:01:20.680 -> Average Cal sensor Temp.: 90.89 Ambient Temp.: 84.61 F
01:01:22.233 -> Average Cal sensor Temp.: 91.18 Ambient Temp.: 84.65 F
01:01:23.789 -> Average Cal sensor Temp.: 91.02 Ambient Temp.: 84.69 F
01:01:25.336 -> Average Cal sensor Temp.: 91.18 Ambient Temp.: 84.69 F
01:01:26.890 -> Average Cal sensor Temp.: 91.17 Ambient Temp.: 84.69 F
01:01:28.451 -> Average Cal sensor Temp.: 91.01 Ambient Temp.: 84.61 F
01:01:30.008 -> Average Cal sensor Temp.: 90.98 Ambient Temp.: 84.61 F
01:01:31.567 -> Average Cal sensor Temp.: 90.86 Ambient Temp.: 84.69 F
01:01:33.129 -> Average Cal sensor Temp.: 91.18 Ambient Temp.: 84.65 F
01:01:34.684 -> Average Cal sensor Temp.: 91.06 Ambient Temp.: 84.61 F
01:01:36.239 -> Average Cal sensor Temp.: 90.94 Ambient Temp.: 84.58 F
01:01:37.796 -> Average Cal sensor Temp.: 90.98 Ambient Temp.: 84.54 F
01:01:39.347 -> Average Cal sensor Temp.: 91.19 Ambient Temp.: 84.58 F
01:01:40.904 -> Average Cal sensor Temp.: 90.80 Ambient Temp.: 84.61 F
01:01:42.460 -> Average Cal sensor Temp.: 91.18 Ambient Temp.: 84.61 F
01:01:44.010 -> Average Cal sensor Temp.: 91.08 Ambient Temp.: 84.65 F
01:01:45.568 -> Average Cal sensor Temp.: 91.24 Ambient Temp.: 84.61 F
01:01:47.121 -> Average Cal sensor Temp.: 91.20 Ambient Temp.: 84.61 F
01:01:48.673 -> Average Cal sensor Temp.: 91.22 Ambient Temp.: 84.61 F
01:01:50.234 -> Average Cal sensor Temp.: 91.58 Ambient Temp.: 84.58 F
01:01:51.790 -> Average Cal sensor Temp.: 91.60 Ambient Temp.: 84.58 F
01:01:53.350 -> Average Cal sensor Temp.: 91.17 Ambient Temp.: 84.51 F
01:01:54.907 -> Average Cal sensor Temp.: 91.47 Ambient Temp.: 84.51 F
01:01:56.468 -> Average Cal sensor Temp.: 91.46 Ambient Temp.: 84.51 F
01:01:58.023 -> Average Cal sensor Temp.: 91.78 Ambient Temp.: 84.54 F
01:01:59.575 -> Average Cal sensor Temp.: 91.62 Ambient Temp.: 84.51 F
01:02:01.135 -> Average Cal sensor Temp.: 91.72 Ambient Temp.: 84.51 F
01:02:02.691 -> Average Cal sensor Temp.: 91.42 Ambient Temp.: 84.51 F
01:02:04.248 -> Average Cal sensor Temp.: 91.17 Ambient Temp.: 84.47 F
01:02:05.801 -> Average Cal sensor Temp.: 91.37 Ambient Temp.: 84.47 F
01:02:07.358 -> Average Cal sensor Temp.: 91.31 Ambient Temp.: 84.51 F
01:02:08.883 -> Average Cal sensor Temp.: 91.66 Ambient Temp.: 84.40 F
01:02:10.474 -> Average Cal sensor Temp.: 91.84 Ambient Temp.: 84.51 F
01:02:11.998 -> Average Cal sensor Temp.: 91.80 Ambient Temp.: 84.47 F
01:02:13.553 -> Average Cal sensor Temp.: 91.56 Ambient Temp.: 84.47 F
01:02:15.109 -> Average Cal sensor Temp.: 91.79 Ambient Temp.: 84.51 F
01:02:16.666 -> Average Cal sensor Temp.: 91.96 Ambient Temp.: 84.51 F
01:02:18.225 -> Average Cal sensor Temp.: 91.53 Ambient Temp.: 84.51 F
01:02:19.785 -> Average Cal sensor Temp.: 91.44 Ambient Temp.: 84.47 F
01:02:21.339 -> Average Cal sensor Temp.: 91.61 Ambient Temp.: 84.51 F
01:02:22.892 -> Average Cal sensor Temp.: 91.40 Ambient Temp.: 84.47 F
01:02:24.454 -> Average Cal sensor Temp.: 91.75 Ambient Temp.: 84.51 F
01:02:26.013 -> Average Cal sensor Temp.: 91.70 Ambient Temp.: 84.47 F
01:02:27.572 -> Average Cal sensor Temp.: 91.94 Ambient Temp.: 84.51 F
01:02:29.126 -> Average Cal sensor Temp.: 92.10 Ambient Temp.: 84.47 F
01:02:30.681 -> Average Cal sensor Temp.: 91.96 Ambient Temp.: 84.51 F
01:02:32.233 -> Average Cal sensor Temp.: 91.45 Ambient Temp.: 84.47 F
01:02:33.789 -> Average Cal sensor Temp.: 91.47 Ambient Temp.: 84.51 F
01:02:35.344 -> Average Cal sensor Temp.: 91.44 Ambient Temp.: 84.36 F
01:02:36.900 -> Average Cal sensor Temp.: 91.53 Ambient Temp.: 84.36 F
01:02:38.456 -> Average Cal sensor Temp.: 91.67 Ambient Temp.: 84.47 F
01:02:40.015 -> Average Cal sensor Temp.: 91.84 Ambient Temp.: 84.40 F
01:02:41.571 -> Average Cal sensor Temp.: 91.83 Ambient Temp.: 84.47 F
01:02:43.131 -> Average Cal sensor Temp.: 91.90 Ambient Temp.: 84.47 F
01:02:44.686 -> Average Cal sensor Temp.: 91.32 Ambient Temp.: 84.47 F
01:02:46.246 -> Average Cal sensor Temp.: 91.54 Ambient Temp.: 84.36 F
01:02:47.801 -> Average Cal sensor Temp.: 91.42 Ambient Temp.: 84.36 F
01:02:49.326 -> Average Cal sensor Temp.: 91.59 Ambient Temp.: 84.36 F
01:02:50.884 -> Average Cal sensor Temp.: 91.44 Ambient Temp.: 84.33 F
01:02:52.443 -> Average Cal sensor Temp.: 91.86 Ambient Temp.: 84.25 F
01:02:53.998 -> Average Cal sensor Temp.: 91.87 Ambient Temp.: 84.36 F
01:02:55.554 -> Average Cal sensor Temp.: 91.87 Ambient Temp.: 84.33 F
01:02:57.111 -> Average Cal sensor Temp.: 91.68 Ambient Temp.: 84.33 F
01:02:58.669 -> Average Cal sensor Temp.: 91.50 Ambient Temp.: 84.36 F
01:03:00.223 -> Average Cal sensor Temp.: 91.76 Ambient Temp.: 84.33 F
01:03:01.786 -> Average Cal sensor Temp.: 91.56 Ambient Temp.: 84.25 F
01:03:03.342 -> Average Cal sensor Temp.: 91.65 Ambient Temp.: 84.29 F
01:03:04.893 -> Average Cal sensor Temp.: 91.68 Ambient Temp.: 84.25 F
01:03:06.448 -> Average Cal sensor Temp.: 92.02 Ambient Temp.: 84.25 F
01:03:08.003 -> Average Cal sensor Temp.: 92.01 Ambient Temp.: 84.22 F
01:03:09.564 -> Average Cal sensor Temp.: 92.01 Ambient Temp.: 84.25 F
01:03:11.122 -> Average Cal sensor Temp.: 91.76 Ambient Temp.: 84.22 F
01:03:12.677 -> Average Cal sensor Temp.: 92.02 Ambient Temp.: 84.22 F
01:03:14.230 -> Average Cal sensor Temp.: 92.04 Ambient Temp.: 84.22 F
01:03:15.779 -> Average Cal sensor Temp.: 91.58 Ambient Temp.: 84.22 F
01:03:17.329 -> Average Cal sensor Temp.: 91.87 Ambient Temp.: 84.18 F
01:03:18.909 -> Average Cal sensor Temp.: 91.59 Ambient Temp.: 84.18 F
01:03:20.462 -> Average Cal sensor Temp.: 91.66 Ambient Temp.: 84.15 F
01:03:22.019 -> Average Cal sensor Temp.: 91.80 Ambient Temp.: 84.11 F
01:03:23.577 -> Average Cal sensor Temp.: 92.05 Ambient Temp.: 84.11 F
01:03:25.130 -> Average Cal sensor Temp.: 91.71 Ambient Temp.: 84.11 F
01:03:26.683 -> Average Cal sensor Temp.: 91.76 Ambient Temp.: 84.15 F
01:03:28.239 -> Average Cal sensor Temp.: 91.72 Ambient Temp.: 84.11 F
01:03:29.788 -> Average Cal sensor Temp.: 91.76 Ambient Temp.: 84.15 F
01:03:31.338 -> Average Cal sensor Temp.: 91.72 Ambient Temp.: 84.11 F
01:03:32.891 -> Average Cal sensor Temp.: 91.88 Ambient Temp.: 84.15 F
01:03:34.442 -> Average Cal sensor Temp.: 91.70 Ambient Temp.: 84.15 F
01:03:35.995 -> Average Cal sensor Temp.: 91.65 Ambient Temp.: 84.18 F
01:03:37.547 -> Average Cal sensor Temp.: 91.43 Ambient Temp.: 84.11 F
01:03:39.104 -> Average Cal sensor Temp.: 91.63 Ambient Temp.: 84.15 F
01:03:40.662 -> Average Cal sensor Temp.: 91.38 Ambient Temp.: 84.11 F
01:03:42.220 -> Average Cal sensor Temp.: 91.85 Ambient Temp.: 84.11 F
01:03:43.776 -> Average Cal sensor Temp.: 92.12 Ambient Temp.: 84.15 F
01:03:45.333 -> Average Cal sensor Temp.: 92.02 Ambient Temp.: 84.11 F
01:03:46.890 -> Average Cal sensor Temp.: 92.12 Ambient Temp.: 84.15 F
01:03:48.447 -> Average Cal sensor Temp.: 91.99 Ambient Temp.: 84.15 F
01:03:50.000 -> Average Cal sensor Temp.: 92.15 Ambient Temp.: 84.15 F
01:03:51.560 -> Average Cal sensor Temp.: 92.13 Ambient Temp.: 84.18 F
01:03:53.117 -> Average Cal sensor Temp.: 92.19 Ambient Temp.: 84.18 F
01:03:54.668 -> Average Cal sensor Temp.: 92.11 Ambient Temp.: 84.15 F
01:03:56.227 -> Average Cal sensor Temp.: 91.84 Ambient Temp.: 84.15 F
01:03:57.785 -> Average Cal sensor Temp.: 92.10 Ambient Temp.: 84.15 F
01:03:59.311 -> Average Cal sensor Temp.: 92.32 Ambient Temp.: 84.15 F
01:04:00.870 -> Average Cal sensor Temp.: 92.34 Ambient Temp.: 84.22 F
01:04:02.428 -> Average Cal sensor Temp.: 92.18 Ambient Temp.: 84.15 F
01:04:03.980 -> Average Cal sensor Temp.: 92.14 Ambient Temp.: 84.18 F
01:04:05.536 -> Average Cal sensor Temp.: 91.99 Ambient Temp.: 84.18 F
01:04:07.094 -> Average Cal sensor Temp.: 91.89 Ambient Temp.: 84.15 F
01:04:08.649 -> Average Cal sensor Temp.: 91.98 Ambient Temp.: 84.15 F
01:04:10.209 -> Average Cal sensor Temp.: 92.07 Ambient Temp.: 84.15 F
01:04:11.771 -> Average Cal sensor Temp.: 92.08 Ambient Temp.: 84.18 F
01:04:13.329 -> Average Cal sensor Temp.: 92.36 Ambient Temp.: 84.18 F
01:04:14.886 -> Average Cal sensor Temp.: 92.43 Ambient Temp.: 84.18 F
01:04:16.448 -> Average Cal sensor Temp.: 92.40 Ambient Temp.: 84.22 F
01:04:18.006 -> Average Cal sensor Temp.: 92.54 Ambient Temp.: 84.18 F
01:04:19.563 -> Average Cal sensor Temp.: 92.42 Ambient Temp.: 84.25 F
01:04:21.122 -> Average Cal sensor Temp.: 92.19 Ambient Temp.: 84.22 F
01:04:22.643 -> Average Cal sensor Temp.: 92.39 Ambient Temp.: 84.22 F
01:04:24.203 -> Average Cal sensor Temp.: 92.20 Ambient Temp.: 84.18 F
01:04:25.757 -> Average Cal sensor Temp.: 92.55 Ambient Temp.: 84.25 F
01:04:27.315 -> Average Cal sensor Temp.: 92.49 Ambient Temp.: 84.29 F
01:04:28.871 -> Average Cal sensor Temp.: 92.44 Ambient Temp.: 84.29 F
01:04:30.431 -> Average Cal sensor Temp.: 92.44 Ambient Temp.: 84.29 F
01:04:31.992 -> Average Cal sensor Temp.: 92.57 Ambient Temp.: 84.33 F
01:04:33.551 -> Average Cal sensor Temp.: 92.39 Ambient Temp.: 84.29 F
01:04:35.108 -> Average Cal sensor Temp.: 92.30 Ambient Temp.: 84.29 F
01:04:36.658 -> Average Cal sensor Temp.: 92.52 Ambient Temp.: 84.36 F
01:04:38.207 -> Average Cal sensor Temp.: 92.57 Ambient Temp.: 84.33 F
01:04:39.762 -> Average Cal sensor Temp.: 92.23 Ambient Temp.: 84.29 F
01:04:41.317 -> Average Cal sensor Temp.: 92.36 Ambient Temp.: 84.36 F
01:04:42.906 -> Average Cal sensor Temp.: 92.48 Ambient Temp.: 84.33 F
01:04:44.459 -> Average Cal sensor Temp.: 92.19 Ambient Temp.: 84.36 F
01:04:46.016 -> Average Cal sensor Temp.: 92.25 Ambient Temp.: 84.36 F
01:04:47.570 -> Average Cal sensor Temp.: 92.52 Ambient Temp.: 84.33 F
01:04:49.119 -> Average Cal sensor Temp.: 92.36 Ambient Temp.: 84.36 F
01:04:50.680 -> Average Cal sensor Temp.: 92.11 Ambient Temp.: 84.40 F
01:04:52.233 -> Average Cal sensor Temp.: 92.47 Ambient Temp.: 84.40 F
01:04:53.792 -> Average Cal sensor Temp.: 92.22 Ambient Temp.: 84.40 F
01:04:55.344 -> Average Cal sensor Temp.: 92.52 Ambient Temp.: 84.51 F
01:04:56.899 -> Average Cal sensor Temp.: 92.47 Ambient Temp.: 84.47 F
01:04:58.455 -> Average Cal sensor Temp.: 92.20 Ambient Temp.: 84.47 F
01:05:00.009 -> Average Cal sensor Temp.: 92.48 Ambient Temp.: 84.47 F
01:05:01.569 -> Average Cal sensor Temp.: 92.54 Ambient Temp.: 84.51 F
01:05:03.128 -> Average Cal sensor Temp.: 92.53 Ambient Temp.: 84.54 F
01:05:04.681 -> Average Cal sensor Temp.: 92.59 Ambient Temp.: 84.54 F
01:05:06.234 -> Average Cal sensor Temp.: 92.31 Ambient Temp.: 84.58 F
01:05:07.784 -> Average Cal sensor Temp.: 92.28 Ambient Temp.: 84.58 F
01:05:09.343 -> Average Cal sensor Temp.: 92.57 Ambient Temp.: 84.58 F
01:05:10.898 -> Average Cal sensor Temp.: 92.61 Ambient Temp.: 84.58 F
01:05:12.449 -> Average Cal sensor Temp.: 92.53 Ambient Temp.: 84.65 F
01:05:14.007 -> Average Cal sensor Temp.: 92.77 Ambient Temp.: 84.65 F
01:05:15.557 -> Average Cal sensor Temp.: 92.75 Ambient Temp.: 84.61 F
01:05:17.112 -> Average Cal sensor Temp.: 92.93 Ambient Temp.: 84.69 F
01:05:18.670 -> Average Cal sensor Temp.: 92.89 Ambient Temp.: 84.65 F
01:05:20.220 -> Average Cal sensor Temp.: 92.84 Ambient Temp.: 84.72 F
01:05:21.773 -> Average Cal sensor Temp.: 92.42 Ambient Temp.: 84.69 F
01:05:23.360 -> Average Cal sensor Temp.: 92.70 Ambient Temp.: 84.72 F
01:05:24.911 -> Average Cal sensor Temp.: 92.46 Ambient Temp.: 84.72 F
01:05:26.461 -> Average Cal sensor Temp.: 92.64 Ambient Temp.: 84.76 F
01:05:28.018 -> Average Cal sensor Temp.: 92.46 Ambient Temp.: 84.72 F
01:05:29.577 -> Average Cal sensor Temp.: 92.82 Ambient Temp.: 84.79 F
01:05:31.129 -> Average Cal sensor Temp.: 92.41 Ambient Temp.: 84.76 F
01:05:32.682 -> Average Cal sensor Temp.: 92.19 Ambient Temp.: 84.79 F
01:05:34.238 -> Average Cal sensor Temp.: 92.49 Ambient Temp.: 84.87 F
01:05:35.796 -> Average Cal sensor Temp.: 92.44 Ambient Temp.: 84.87 F
01:05:37.354 -> Average Cal sensor Temp.: 92.40 Ambient Temp.: 84.87 F
01:05:38.915 -> Average Cal sensor Temp.: 92.63 Ambient Temp.: 84.90 F
01:05:40.474 -> Average Cal sensor Temp.: 92.55 Ambient Temp.: 84.94 F
01:05:42.015 -> Average Cal sensor Temp.: 92.62 Ambient Temp.: 84.94 F
01:05:43.578 -> Average Cal sensor Temp.: 92.62 Ambient Temp.: 84.97 F
01:05:45.141 -> Average Cal sensor Temp.: 92.72 Ambient Temp.: 84.97 F
01:05:46.707 -> Average Cal sensor Temp.: 92.67 Ambient Temp.: 84.97 F
01:05:48.263 -> Average Cal sensor Temp.: 92.82 Ambient Temp.: 85.01 F
01:05:49.815 -> Average Cal sensor Temp.: 92.83 Ambient Temp.: 84.94 F
01:05:51.375 -> Average Cal sensor Temp.: 92.65 Ambient Temp.: 85.01 F
01:05:52.934 -> Average Cal sensor Temp.: 92.70 Ambient Temp.: 85.01 F
01:05:54.485 -> Average Cal sensor Temp.: 92.69 Ambient Temp.: 84.97 F
01:05:56.041 -> Average Cal sensor Temp.: 92.71 Ambient Temp.: 84.97 F
01:05:57.599 -> Average Cal sensor Temp.: 92.81 Ambient Temp.: 85.05 F
01:05:59.154 -> Average Cal sensor Temp.: 92.86 Ambient Temp.: 85.01 F
01:06:00.710 -> Average Cal sensor Temp.: 92.62 Ambient Temp.: 85.05 F
01:06:02.268 -> Average Cal sensor Temp.: 92.79 Ambient Temp.: 85.08 F
01:06:03.822 -> Average Cal sensor Temp.: 92.85 Ambient Temp.: 85.12 F
01:06:05.380 -> Average Cal sensor Temp.: 92.80 Ambient Temp.: 85.12 F
01:06:06.937 -> Average Cal sensor Temp.: 92.84 Ambient Temp.: 85.12 F
01:06:08.495 -> Average Cal sensor Temp.: 92.80 Ambient Temp.: 85.12 F
01:06:10.019 -> Average Cal sensor Temp.: 92.71 Ambient Temp.: 85.15 F
01:06:11.581 -> Average Cal sensor Temp.: 92.82 Ambient Temp.: 85.19 F
01:06:13.136 -> Average Cal sensor Temp.: 92.83 Ambient Temp.: 85.26 F
01:06:14.695 -> Average Cal sensor Temp.: 92.47 Ambient Temp.: 85.26 F
01:06:16.251 -> Average Cal sensor Temp.: 92.54 Ambient Temp.: 85.26 F
01:06:17.802 -> Average Cal sensor Temp.: 92.73 Ambient Temp.: 85.30 F
01:06:19.360 -> Average Cal sensor Temp.: 92.99 Ambient Temp.: 85.30 F
01:06:20.923 -> Average Cal sensor Temp.: 92.49 Ambient Temp.: 85.30 F
01:06:22.490 -> Average Cal sensor Temp.: 92.78 Ambient Temp.: 85.30 F
01:06:24.024 -> Average Cal sensor Temp.: 92.70 Ambient Temp.: 85.30 F
01:06:25.591 -> Average Cal sensor Temp.: 93.01 Ambient Temp.: 85.30 F
01:06:27.162 -> Average Cal sensor Temp.: 92.94 Ambient Temp.: 85.33 F
01:06:28.697 -> Average Cal sensor Temp.: 93.08 Ambient Temp.: 85.33 F
01:06:30.266 -> Average Cal sensor Temp.: 93.06 Ambient Temp.: 85.33 F
01:06:31.800 -> Average Cal sensor Temp.: 93.07 Ambient Temp.: 85.33 F
01:06:33.371 -> Average Cal sensor Temp.: 93.06 Ambient Temp.: 85.41 F
01:06:34.942 -> Average Cal sensor Temp.: 93.07 Ambient Temp.: 85.37 F
01:06:36.478 -> Average Cal sensor Temp.: 92.85 Ambient Temp.: 85.37 F
01:06:38.048 -> Average Cal sensor Temp.: 93.02 Ambient Temp.: 85.37 F
01:06:39.584 -> Average Cal sensor Temp.: 93.07 Ambient Temp.: 85.44 F
01:06:41.153 -> Average Cal sensor Temp.: 92.70 Ambient Temp.: 85.41 F
01:06:42.689 -> Average Cal sensor Temp.: 92.73 Ambient Temp.: 85.41 F
01:06:44.260 -> Average Cal sensor Temp.: 92.74 Ambient Temp.: 85.37 F
01:06:45.828 -> Average Cal sensor Temp.: 92.74 Ambient Temp.: 85.41 F
01:06:47.364 -> Average Cal sensor Temp.: 93.06 Ambient Temp.: 85.48 F
01:06:48.929 -> Average Cal sensor Temp.: 92.99 Ambient Temp.: 85.44 F
01:06:50.487 -> Average Cal sensor Temp.: 92.97 Ambient Temp.: 85.48 F
01:06:52.049 -> Average Cal sensor Temp.: 92.63 Ambient Temp.: 85.48 F
01:06:53.579 -> Average Cal sensor Temp.: 92.80 Ambient Temp.: 85.41 F
01:06:55.143 -> Average Cal sensor Temp.: 92.76 Ambient Temp.: 85.41 F
01:06:56.704 -> Average Cal sensor Temp.: 92.84 Ambient Temp.: 85.48 F
01:06:58.251 -> Average Cal sensor Temp.: 92.57 Ambient Temp.: 85.51 F
01:06:59.802 -> Average Cal sensor Temp.: 93.07 Ambient Temp.: 85.51 F
01:07:01.389 -> Average Cal sensor Temp.: 92.80 Ambient Temp.: 85.48 F
01:07:02.940 -> Average Cal sensor Temp.: 92.93 Ambient Temp.: 85.48 F
01:07:04.493 -> Average Cal sensor Temp.: 92.55 Ambient Temp.: 85.44 F
01:07:06.047 -> Average Cal sensor Temp.: 92.55 Ambient Temp.: 85.44 F
01:07:07.601 -> Average Cal sensor Temp.: 92.99 Ambient Temp.: 85.44 F
01:07:09.156 -> Average Cal sensor Temp.: 93.11 Ambient Temp.: 85.48 F
01:07:10.710 -> Average Cal sensor Temp.: 92.97 Ambient Temp.: 85.41 F
01:07:12.258 -> Average Cal sensor Temp.: 92.59 Ambient Temp.: 85.30 F
01:07:13.808 -> Average Cal sensor Temp.: 92.82 Ambient Temp.: 85.37 F
01:07:15.361 -> Average Cal sensor Temp.: 92.49 Ambient Temp.: 85.37 F
01:07:16.916 -> Average Cal sensor Temp.: 93.09 Ambient Temp.: 85.41 F
01:07:18.505 -> Average Cal sensor Temp.: 93.05 Ambient Temp.: 85.37 F
01:07:20.060 -> Average Cal sensor Temp.: 93.07 Ambient Temp.: 85.33 F
01:07:21.583 -> Average Cal sensor Temp.: 93.37 Ambient Temp.: 85.37 F
01:07:23.169 -> Average Cal sensor Temp.: 93.10 Ambient Temp.: 85.37 F
01:07:24.721 -> Average Cal sensor Temp.: 93.00 Ambient Temp.: 85.37 F
01:07:26.276 -> Average Cal sensor Temp.: 92.79 Ambient Temp.: 85.41 F
01:07:27.829 -> Average Cal sensor Temp.: 92.89 Ambient Temp.: 85.37 F
01:07:29.381 -> Average Cal sensor Temp.: 92.89 Ambient Temp.: 85.41 F
01:07:30.940 -> Average Cal sensor Temp.: 92.73 Ambient Temp.: 85.37 F
01:07:32.496 -> Average Cal sensor Temp.: 92.80 Ambient Temp.: 85.37 F
01:07:34.053 -> Average Cal sensor Temp.: 92.80 Ambient Temp.: 85.30 F
01:07:35.606 -> Average Cal sensor Temp.: 93.13 Ambient Temp.: 85.41 F
01:07:37.162 -> Average Cal sensor Temp.: 93.11 Ambient Temp.: 85.33 F
01:07:38.715 -> Average Cal sensor Temp.: 93.19 Ambient Temp.: 85.33 F
01:07:40.273 -> Average Cal sensor Temp.: 93.16 Ambient Temp.: 85.41 F
01:07:41.824 -> Average Cal sensor Temp.: 93.20 Ambient Temp.: 85.44 F
01:07:43.377 -> Average Cal sensor Temp.: 93.24 Ambient Temp.: 85.44 F
01:07:44.936 -> Average Cal sensor Temp.: 92.93 Ambient Temp.: 85.44 F
01:07:46.496 -> Average Cal sensor Temp.: 92.99 Ambient Temp.: 85.44 F
01:07:48.046 -> Average Cal sensor Temp.: 93.05 Ambient Temp.: 85.44 F
01:07:49.601 -> Average Cal sensor Temp.: 92.98 Ambient Temp.: 85.44 F
01:07:51.159 -> Average Cal sensor Temp.: 93.08 Ambient Temp.: 85.44 F
01:07:52.713 -> Average Cal sensor Temp.: 93.33 Ambient Temp.: 85.44 F
01:07:54.273 -> Average Cal sensor Temp.: 93.02 Ambient Temp.: 85.41 F
01:07:55.833 -> Average Cal sensor Temp.: 93.38 Ambient Temp.: 85.41 F
01:07:57.381 -> Average Cal sensor Temp.: 93.33 Ambient Temp.: 85.44 F
01:07:58.935 -> Average Cal sensor Temp.: 93.08 Ambient Temp.: 85.41 F
01:08:00.484 -> Average Cal sensor Temp.: 93.35 Ambient Temp.: 85.41 F
01:08:02.040 -> Average Cal sensor Temp.: 93.49 Ambient Temp.: 85.44 F
01:08:03.595 -> Average Cal sensor Temp.: 93.43 Ambient Temp.: 85.37 F
01:08:05.150 -> Average Cal sensor Temp.: 93.49 Ambient Temp.: 85.44 F
01:08:06.700 -> Average Cal sensor Temp.: 93.56 Ambient Temp.: 85.48 F
01:08:08.255 -> Average Cal sensor Temp.: 93.47 Ambient Temp.: 85.48 F
01:08:09.810 -> Average Cal sensor Temp.: 93.60 Ambient Temp.: 85.48 F
01:08:11.370 -> Average Cal sensor Temp.: 93.33 Ambient Temp.: 85.48 F
01:08:12.954 -> Average Cal sensor Temp.: 93.15 Ambient Temp.: 85.51 F
01:08:14.476 -> Average Cal sensor Temp.: 93.24 Ambient Temp.: 85.48 F
01:08:16.035 -> Average Cal sensor Temp.: 93.11 Ambient Temp.: 85.48 F
01:08:17.594 -> Average Cal sensor Temp.: 93.16 Ambient Temp.: 85.48 F
01:08:19.150 -> Average Cal sensor Temp.: 92.95 Ambient Temp.: 85.44 F
01:08:20.706 -> Average Cal sensor Temp.: 93.20 Ambient Temp.: 85.44 F
01:08:22.260 -> Average Cal sensor Temp.: 92.69 Ambient Temp.: 85.41 F
01:08:23.813 -> Average Cal sensor Temp.: 92.97 Ambient Temp.: 85.41 F
01:08:25.397 -> Average Cal sensor Temp.: 92.93 Ambient Temp.: 85.33 F
01:08:26.948 -> Average Cal sensor Temp.: 93.11 Ambient Temp.: 85.41 F
01:08:28.500 -> Average Cal sensor Temp.: 93.26 Ambient Temp.: 85.41 F
01:08:30.055 -> Average Cal sensor Temp.: 93.20 Ambient Temp.: 85.41 F
01:08:31.603 -> Average Cal sensor Temp.: 93.00 Ambient Temp.: 85.37 F
01:08:33.154 -> Average Cal sensor Temp.: 93.15 Ambient Temp.: 85.37 F
01:08:34.700 -> Average Cal sensor Temp.: 93.14 Ambient Temp.: 85.41 F
01:08:36.257 -> Average Cal sensor Temp.: 93.14 Ambient Temp.: 85.41 F
01:08:37.839 -> Average Cal sensor Temp.: 93.18 Ambient Temp.: 85.44 F
01:08:39.391 -> Average Cal sensor Temp.: 93.14 Ambient Temp.: 85.41 F
01:08:40.945 -> Average Cal sensor Temp.: 92.84 Ambient Temp.: 85.30 F
01:08:42.504 -> Average Cal sensor Temp.: 93.29 Ambient Temp.: 85.37 F
01:08:44.057 -> Average Cal sensor Temp.: 93.26 Ambient Temp.: 85.30 F
01:08:45.611 -> Average Cal sensor Temp.: 93.01 Ambient Temp.: 85.33 F
01:08:47.160 -> Average Cal sensor Temp.: 93.33 Ambient Temp.: 85.26 F
01:08:48.715 -> Average Cal sensor Temp.: 92.85 Ambient Temp.: 85.30 F
01:08:50.270 -> Average Cal sensor Temp.: 93.37 Ambient Temp.: 85.30 F
01:08:51.827 -> Average Cal sensor Temp.: 93.38 Ambient Temp.: 85.26 F
01:08:53.387 -> Average Cal sensor Temp.: 92.96 Ambient Temp.: 85.26 F
01:08:54.942 -> Average Cal sensor Temp.: 93.14 Ambient Temp.: 85.26 F
01:08:56.500 -> Average Cal sensor Temp.: 93.24 Ambient Temp.: 85.15 F
01:08:58.056 -> Average Cal sensor Temp.: 93.02 Ambient Temp.: 85.19 F
01:08:59.616 -> Average Cal sensor Temp.: 93.14 Ambient Temp.: 85.12 F
01:09:01.141 -> Average Cal sensor Temp.: 93.23 Ambient Temp.: 85.12 F
01:09:02.699 -> Average Cal sensor Temp.: 93.29 Ambient Temp.: 85.08 F
01:09:04.249 -> Average Cal sensor Temp.: 93.39 Ambient Temp.: 85.12 F
01:09:05.832 -> Average Cal sensor Temp.: 93.41 Ambient Temp.: 85.08 F
01:09:07.381 -> Average Cal sensor Temp.: 93.12 Ambient Temp.: 85.08 F
01:09:08.931 -> Average Cal sensor Temp.: 93.44 Ambient Temp.: 85.05 F
01:09:10.486 -> Average Cal sensor Temp.: 93.70 Ambient Temp.: 85.12 F
01:09:12.037 -> Average Cal sensor Temp.: 93.48 Ambient Temp.: 85.05 F
01:09:13.595 -> Average Cal sensor Temp.: 93.07 Ambient Temp.: 85.01 F
01:09:15.148 -> Average Cal sensor Temp.: 93.54 Ambient Temp.: 85.05 F
01:09:16.698 -> Average Cal sensor Temp.: 93.70 Ambient Temp.: 85.05 F
01:09:18.259 -> Average Cal sensor Temp.: 93.61 Ambient Temp.: 85.12 F
01:09:19.811 -> Average Cal sensor Temp.: 93.37 Ambient Temp.: 85.08 F
01:09:21.365 -> Average Cal sensor Temp.: 93.46 Ambient Temp.: 85.08 F
01:09:22.949 -> Average Cal sensor Temp.: 93.46 Ambient Temp.: 85.08 F
01:09:24.506 -> Average Cal sensor Temp.: 93.52 Ambient Temp.: 85.05 F
01:09:26.061 -> Average Cal sensor Temp.: 93.63 Ambient Temp.: 85.05 F
01:09:27.611 -> Average Cal sensor Temp.: 93.56 Ambient Temp.: 85.05 F
01:09:29.173 -> Average Cal sensor Temp.: 93.49 Ambient Temp.: 85.08 F
01:09:30.729 -> Average Cal sensor Temp.: 93.23 Ambient Temp.: 85.05 F
01:09:32.276 -> Average Cal sensor Temp.: 93.34 Ambient Temp.: 85.05 F
01:09:33.828 -> Average Cal sensor Temp.: 93.39 Ambient Temp.: 85.08 F
01:09:35.386 -> Average Cal sensor Temp.: 93.43 Ambient Temp.: 85.05 F
01:09:36.932 -> Average Cal sensor Temp.: 93.22 Ambient Temp.: 85.01 F
01:09:38.483 -> Average Cal sensor Temp.: 93.14 Ambient Temp.: 84.97 F
01:09:40.039 -> Average Cal sensor Temp.: 93.35 Ambient Temp.: 84.94 F
01:09:41.588 -> Average Cal sensor Temp.: 93.50 Ambient Temp.: 85.01 F
01:09:43.157 -> Average Cal sensor Temp.: 93.65 Ambient Temp.: 84.97 F
01:09:44.700 -> Average Cal sensor Temp.: 93.65 Ambient Temp.: 85.01 F
01:09:46.254 -> Average Cal sensor Temp.: 93.67 Ambient Temp.: 84.97 F
01:09:47.840 -> Average Cal sensor Temp.: 93.11 Ambient Temp.: 84.97 F
01:09:49.397 -> Average Cal sensor Temp.: 93.08 Ambient Temp.: 84.90 F
01:09:50.951 -> Average Cal sensor Temp.: 93.14 Ambient Temp.: 84.90 F
01:09:52.507 -> Average Cal sensor Temp.: 93.28 Ambient Temp.: 84.87 F
01:09:54.059 -> Average Cal sensor Temp.: 93.57 Ambient Temp.: 84.94 F
01:09:55.611 -> Average Cal sensor Temp.: 93.43 Ambient Temp.: 84.90 F
01:09:57.165 -> Average Cal sensor Temp.: 93.66 Ambient Temp.: 84.90 F
01:09:58.722 -> Average Cal sensor Temp.: 93.56 Ambient Temp.: 84.90 F
01:10:00.282 -> Average Cal sensor Temp.: 93.17 Ambient Temp.: 84.90 F
01:10:01.837 -> Average Cal sensor Temp.: 93.32 Ambient Temp.: 84.94 F
01:10:03.396 -> Average Cal sensor Temp.: 93.95 Ambient Temp.: 84.90 F
01:10:04.942 -> Average Cal sensor Temp.: 94.13 Ambient Temp.: 84.90 F
01:10:06.494 -> Average Cal sensor Temp.: 94.62 Ambient Temp.: 84.94 F
01:10:08.053 -> Average Cal sensor Temp.: 94.15 Ambient Temp.: 84.94 F
01:10:09.610 -> Average Cal sensor Temp.: 94.50 Ambient Temp.: 85.01 F
01:10:11.168 -> Average Cal sensor Temp.: 94.34 Ambient Temp.: 84.97 F
01:10:12.727 -> Average Cal sensor Temp.: 94.32 Ambient Temp.: 85.01 F
01:10:14.283 -> Average Cal sensor Temp.: 94.33 Ambient Temp.: 85.05 F
01:10:15.841 -> Average Cal sensor Temp.: 94.17 Ambient Temp.: 85.01 F
01:10:17.395 -> Average Cal sensor Temp.: 94.31 Ambient Temp.: 85.08 F
01:10:18.950 -> Average Cal sensor Temp.: 94.08 Ambient Temp.: 85.12 F
01:10:20.508 -> Average Cal sensor Temp.: 94.13 Ambient Temp.: 85.08 F
01:10:22.060 -> Average Cal sensor Temp.: 93.96 Ambient Temp.: 85.08 F
01:10:23.615 -> Average Cal sensor Temp.: 94.17 Ambient Temp.: 85.08 F
01:10:25.172 -> Average Cal sensor Temp.: 94.22 Ambient Temp.: 85.12 F
01:10:26.725 -> Average Cal sensor Temp.: 94.19 Ambient Temp.: 85.15 F
01:10:28.285 -> Average Cal sensor Temp.: 93.64 Ambient Temp.: 85.12 F
01:10:29.842 -> Average Cal sensor Temp.: 93.73 Ambient Temp.: 85.26 F
01:10:31.400 -> Average Cal sensor Temp.: 93.78 Ambient Temp.: 85.12 F
01:10:32.953 -> Average Cal sensor Temp.: 94.00 Ambient Temp.: 85.15 F
01:10:34.510 -> Average Cal sensor Temp.: 93.78 Ambient Temp.: 85.19 F
01:10:36.063 -> Average Cal sensor Temp.: 94.05 Ambient Temp.: 85.15 F
01:10:37.614 -> Average Cal sensor Temp.: 93.94 Ambient Temp.: 85.15 F
01:10:39.161 -> Average Cal sensor Temp.: 94.02 Ambient Temp.: 85.30 F
01:10:40.721 -> Average Cal sensor Temp.: 93.97 Ambient Temp.: 85.26 F
01:10:42.271 -> Average Cal sensor Temp.: 93.72 Ambient Temp.: 85.30 F
01:10:43.852 -> Average Cal sensor Temp.: 93.75 Ambient Temp.: 85.30 F
01:10:45.405 -> Average Cal sensor Temp.: 93.74 Ambient Temp.: 85.30 F
01:10:46.961 -> Average Cal sensor Temp.: 93.97 Ambient Temp.: 85.26 F
01:10:48.516 -> Average Cal sensor Temp.: 93.78 Ambient Temp.: 85.30 F
01:10:50.075 -> Average Cal sensor Temp.: 93.95 Ambient Temp.: 85.33 F
01:10:51.628 -> Average Cal sensor Temp.: 94.03 Ambient Temp.: 85.37 F
01:10:53.182 -> Average Cal sensor Temp.: 94.01 Ambient Temp.: 85.33 F
01:10:54.734 -> Average Cal sensor Temp.: 94.07 Ambient Temp.: 85.37 F
01:10:56.288 -> Average Cal sensor Temp.: 93.80 Ambient Temp.: 85.37 F
01:10:57.840 -> Average Cal sensor Temp.: 93.83 Ambient Temp.: 85.33 F
01:10:59.399 -> Average Cal sensor Temp.: 93.91 Ambient Temp.: 85.33 F
01:11:00.951 -> Average Cal sensor Temp.: 93.82 Ambient Temp.: 85.37 F
01:11:02.503 -> Average Cal sensor Temp.: 93.63 Ambient Temp.: 85.41 F
01:11:04.060 -> Average Cal sensor Temp.: 93.78 Ambient Temp.: 85.37 F
01:11:05.617 -> Average Cal sensor Temp.: 93.99 Ambient Temp.: 85.41 F
01:11:07.169 -> Average Cal sensor Temp.: 94.01 Ambient Temp.: 85.37 F
01:11:08.723 -> Average Cal sensor Temp.: 93.96 Ambient Temp.: 85.41 F
01:11:10.279 -> Average Cal sensor Temp.: 93.98 Ambient Temp.: 85.37 F
01:11:11.834 -> Average Cal sensor Temp.: 94.01 Ambient Temp.: 85.41 F
01:11:13.390 -> Average Cal sensor Temp.: 94.34 Ambient Temp.: 85.44 F
01:11:14.948 -> Average Cal sensor Temp.: 94.32 Ambient Temp.: 85.41 F
01:11:16.511 -> Average Cal sensor Temp.: 94.49 Ambient Temp.: 85.41 F
01:11:18.065 -> Average Cal sensor Temp.: 94.06 Ambient Temp.: 85.41 F
01:11:19.615 -> Average Cal sensor Temp.: 94.20 Ambient Temp.: 85.44 F
01:11:21.167 -> Average Cal sensor Temp.: 94.30 Ambient Temp.: 85.48 F
01:11:22.721 -> Average Cal sensor Temp.: 94.37 Ambient Temp.: 85.48 F
01:11:24.310 -> Average Cal sensor Temp.: 94.20 Ambient Temp.: 85.55 F
01:11:25.833 -> Average Cal sensor Temp.: 94.31 Ambient Temp.: 85.59 F
01:11:27.420 -> Average Cal sensor Temp.: 94.24 Ambient Temp.: 85.55 F
01:11:28.977 -> Average Cal sensor Temp.: 94.24 Ambient Temp.: 85.66 F
01:11:30.500 -> Average Cal sensor Temp.: 94.39 Ambient Temp.: 85.55 F
01:11:32.089 -> Average Cal sensor Temp.: 94.43 Ambient Temp.: 85.59 F
01:11:33.643 -> Average Cal sensor Temp.: 94.37 Ambient Temp.: 85.59 F
01:11:35.198 -> Average Cal sensor Temp.: 94.36 Ambient Temp.: 85.66 F
01:11:36.748 -> Average Cal sensor Temp.: 94.73 Ambient Temp.: 85.69 F
01:11:38.301 -> Average Cal sensor Temp.: 94.63 Ambient Temp.: 85.73 F
01:11:39.864 -> Average Cal sensor Temp.: 94.63 Ambient Temp.: 85.73 F
01:11:41.423 -> Average Cal sensor Temp.: 94.51 Ambient Temp.: 85.77 F
01:11:42.974 -> Average Cal sensor Temp.: 94.32 Ambient Temp.: 85.80 F
01:11:44.536 -> Average Cal sensor Temp.: 94.26 Ambient Temp.: 85.80 F
01:11:46.092 -> Average Cal sensor Temp.: 94.20 Ambient Temp.: 85.84 F
01:11:47.642 -> Average Cal sensor Temp.: 94.29 Ambient Temp.: 85.80 F
01:11:49.196 -> Average Cal sensor Temp.: 94.30 Ambient Temp.: 85.87 F
01:11:50.749 -> Average Cal sensor Temp.: 94.35 Ambient Temp.: 85.84 F
01:11:52.306 -> Average Cal sensor Temp.: 94.67 Ambient Temp.: 85.87 F
01:11:53.861 -> Average Cal sensor Temp.: 94.57 Ambient Temp.: 85.91 F
01:11:55.416 -> Average Cal sensor Temp.: 94.43 Ambient Temp.: 85.95 F
01:11:56.975 -> Average Cal sensor Temp.: 94.62 Ambient Temp.: 85.95 F
01:11:58.531 -> Average Cal sensor Temp.: 94.77 Ambient Temp.: 86.05 F
01:12:00.092 -> Average Cal sensor Temp.: 94.69 Ambient Temp.: 85.98 F
01:12:01.653 -> Average Cal sensor Temp.: 94.37 Ambient Temp.: 86.05 F
01:12:03.215 -> Average Cal sensor Temp.: 94.62 Ambient Temp.: 85.98 F
01:12:04.771 -> Average Cal sensor Temp.: 94.46 Ambient Temp.: 86.05 F
01:12:06.300 -> Average Cal sensor Temp.: 94.63 Ambient Temp.: 86.09 F
01:12:07.889 -> Average Cal sensor Temp.: 94.60 Ambient Temp.: 86.13 F
01:12:09.447 -> Average Cal sensor Temp.: 94.73 Ambient Temp.: 86.13 F
01:12:10.968 -> Average Cal sensor Temp.: 94.68 Ambient Temp.: 86.13 F
01:12:12.555 -> Average Cal sensor Temp.: 94.29 Ambient Temp.: 86.13 F
01:12:14.112 -> Average Cal sensor Temp.: 94.19 Ambient Temp.: 86.09 F
01:12:15.667 -> Average Cal sensor Temp.: 94.68 Ambient Temp.: 86.16 F
01:12:17.227 -> Average Cal sensor Temp.: 94.64 Ambient Temp.: 86.16 F
01:12:18.785 -> Average Cal sensor Temp.: 94.61 Ambient Temp.: 86.16 F
01:12:20.338 -> Average Cal sensor Temp.: 94.79 Ambient Temp.: 86.16 F
01:12:21.886 -> Average Cal sensor Temp.: 94.92 Ambient Temp.: 86.13 F
01:12:23.439 -> Average Cal sensor Temp.: 94.73 Ambient Temp.: 86.20 F
01:12:24.998 -> Average Cal sensor Temp.: 94.64 Ambient Temp.: 86.20 F
01:12:26.553 -> Average Cal sensor Temp.: 94.59 Ambient Temp.: 86.16 F
01:12:28.114 -> Average Cal sensor Temp.: 94.36 Ambient Temp.: 86.20 F
01:12:29.674 -> Average Cal sensor Temp.: 94.29 Ambient Temp.: 86.16 F
01:12:31.229 -> Average Cal sensor Temp.: 94.10 Ambient Temp.: 86.13 F
01:12:32.780 -> Average Cal sensor Temp.: 94.44 Ambient Temp.: 86.16 F
01:12:34.336 -> Average Cal sensor Temp.: 94.84 Ambient Temp.: 86.16 F
01:12:35.887 -> Average Cal sensor Temp.: 94.34 Ambient Temp.: 86.16 F
01:12:37.440 -> Average Cal sensor Temp.: 94.64 Ambient Temp.: 86.16 F
01:12:38.994 -> Average Cal sensor Temp.: 94.65 Ambient Temp.: 86.20 F
01:12:40.553 -> Average Cal sensor Temp.: 94.50 Ambient Temp.: 86.13 F
01:12:42.111 -> Average Cal sensor Temp.: 94.49 Ambient Temp.: 86.16 F
01:12:43.698 -> Average Cal sensor Temp.: 94.44 Ambient Temp.: 86.16 F
01:12:45.249 -> Average Cal sensor Temp.: 94.26 Ambient Temp.: 86.16 F
01:12:46.802 -> Average Cal sensor Temp.: 94.44 Ambient Temp.: 86.16 F
01:12:48.355 -> Average Cal sensor Temp.: 94.82 Ambient Temp.: 86.20 F
01:12:49.907 -> Average Cal sensor Temp.: 94.83 Ambient Temp.: 86.20 F
01:12:51.465 -> Average Cal sensor Temp.: 94.73 Ambient Temp.: 86.23 F
01:12:53.021 -> Average Cal sensor Temp.: 94.48 Ambient Temp.: 86.16 F
01:12:54.580 -> Average Cal sensor Temp.: 94.39 Ambient Temp.: 86.20 F
01:12:56.139 -> Average Cal sensor Temp.: 94.38 Ambient Temp.: 86.16 F
01:12:57.690 -> Average Cal sensor Temp.: 93.99 Ambient Temp.: 86.13 F
01:12:59.247 -> Average Cal sensor Temp.: 94.19 Ambient Temp.: 86.13 F
01:13:00.805 -> Average Cal sensor Temp.: 94.44 Ambient Temp.: 86.05 F
01:13:02.359 -> Average Cal sensor Temp.: 94.17 Ambient Temp.: 86.09 F
01:13:03.910 -> Average Cal sensor Temp.: 94.44 Ambient Temp.: 85.98 F
01:13:05.466 -> Average Cal sensor Temp.: 94.47 Ambient Temp.: 85.98 F
01:13:07.051 -> Average Cal sensor Temp.: 94.42 Ambient Temp.: 85.95 F
01:13:08.606 -> Average Cal sensor Temp.: 94.38 Ambient Temp.: 86.05 F
01:13:10.157 -> Average Cal sensor Temp.: 94.38 Ambient Temp.: 86.05 F
01:13:11.714 -> Average Cal sensor Temp.: 94.15 Ambient Temp.: 85.95 F
01:13:13.270 -> Average Cal sensor Temp.: 94.25 Ambient Temp.: 85.95 F
01:13:14.820 -> Average Cal sensor Temp.: 94.77 Ambient Temp.: 85.91 F
01:13:16.380 -> Average Cal sensor Temp.: 94.59 Ambient Temp.: 86.05 F
01:13:17.937 -> Average Cal sensor Temp.: 94.46 Ambient Temp.: 85.98 F
01:13:19.497 -> Average Cal sensor Temp.: 94.95 Ambient Temp.: 85.95 F
01:13:21.058 -> Average Cal sensor Temp.: 94.77 Ambient Temp.: 85.95 F
01:13:22.606 -> Average Cal sensor Temp.: 94.81 Ambient Temp.: 85.98 F
01:13:24.158 -> Average Cal sensor Temp.: 94.84 Ambient Temp.: 85.98 F
01:13:25.710 -> Average Cal sensor Temp.: 94.68 Ambient Temp.: 85.98 F
01:13:27.258 -> Average Cal sensor Temp.: 94.69 Ambient Temp.: 85.98 F
01:13:28.848 -> Average Cal sensor Temp.: 94.57 Ambient Temp.: 85.98 F
01:13:30.372 -> Average Cal sensor Temp.: 94.84 Ambient Temp.: 85.98 F
01:13:31.961 -> Average Cal sensor Temp.: 94.77 Ambient Temp.: 85.95 F
01:13:33.516 -> Average Cal sensor Temp.: 95.12 Ambient Temp.: 85.98 F
01:13:35.073 -> Average Cal sensor Temp.: 94.99 Ambient Temp.: 85.95 F
01:13:36.627 -> Average Cal sensor Temp.: 95.04 Ambient Temp.: 86.05 F
01:13:38.176 -> Average Cal sensor Temp.: 94.98 Ambient Temp.: 85.98 F
01:13:39.733 -> Average Cal sensor Temp.: 95.01 Ambient Temp.: 85.95 F
01:13:41.296 -> Average Cal sensor Temp.: 94.70 Ambient Temp.: 86.05 F
01:13:42.824 -> Average Cal sensor Temp.: 95.10 Ambient Temp.: 86.05 F
01:13:44.382 -> Average Cal sensor Temp.: 95.22 Ambient Temp.: 86.05 F
01:13:45.935 -> Average Cal sensor Temp.: 94.84 Ambient Temp.: 86.05 F
01:13:47.523 -> Average Cal sensor Temp.: 94.37 Ambient Temp.: 85.95 F
01:13:49.070 -> Average Cal sensor Temp.: 94.32 Ambient Temp.: 85.95 F
01:13:50.628 -> Average Cal sensor Temp.: 94.48 Ambient Temp.: 85.98 F
01:13:52.178 -> Average Cal sensor Temp.: 94.85 Ambient Temp.: 85.98 F
01:13:53.734 -> Average Cal sensor Temp.: 94.55 Ambient Temp.: 85.98 F
01:13:55.281 -> Average Cal sensor Temp.: 94.23 Ambient Temp.: 85.98 F
01:13:56.838 -> Average Cal sensor Temp.: 94.59 Ambient Temp.: 85.95 F
01:13:58.391 -> Average Cal sensor Temp.: 94.58 Ambient Temp.: 85.95 F
01:13:59.976 -> Average Cal sensor Temp.: 94.99 Ambient Temp.: 85.87 F
01:14:01.532 -> Average Cal sensor Temp.: 94.85 Ambient Temp.: 85.91 F
01:14:03.083 -> Average Cal sensor Temp.: 94.75 Ambient Temp.: 85.87 F
01:14:04.632 -> Average Cal sensor Temp.: 94.17 Ambient Temp.: 85.87 F
01:14:06.185 -> Average Cal sensor Temp.: 94.02 Ambient Temp.: 85.87 F
01:14:07.738 -> Average Cal sensor Temp.: 94.57 Ambient Temp.: 85.80 F
01:14:09.291 -> Average Cal sensor Temp.: 94.87 Ambient Temp.: 85.80 F
01:14:10.842 -> Average Cal sensor Temp.: 94.83 Ambient Temp.: 85.80 F
01:14:12.430 -> Average Cal sensor Temp.: 94.93 Ambient Temp.: 85.80 F
01:14:13.989 -> Average Cal sensor Temp.: 95.02 Ambient Temp.: 85.80 F
01:14:15.535 -> Average Cal sensor Temp.: 94.79 Ambient Temp.: 85.77 F
01:14:17.082 -> Average Cal sensor Temp.: 94.68 Ambient Temp.: 85.77 F
01:14:18.634 -> Average Cal sensor Temp.: 94.59 Ambient Temp.: 85.73 F
01:14:20.192 -> Average Cal sensor Temp.: 94.41 Ambient Temp.: 85.73 F
01:14:21.744 -> Average Cal sensor Temp.: 94.19 Ambient Temp.: 85.66 F
01:14:23.293 -> Average Cal sensor Temp.: 94.34 Ambient Temp.: 85.59 F
01:14:24.871 -> Average Cal sensor Temp.: 94.46 Ambient Temp.: 85.66 F
01:14:26.418 -> Average Cal sensor Temp.: 94.10 Ambient Temp.: 85.55 F
01:14:27.964 -> Average Cal sensor Temp.: 94.39 Ambient Temp.: 85.55 F
01:14:29.518 -> Average Cal sensor Temp.: 94.42 Ambient Temp.: 85.51 F
01:14:31.073 -> Average Cal sensor Temp.: 94.46 Ambient Temp.: 85.51 F
01:14:32.630 -> Average Cal sensor Temp.: 94.44 Ambient Temp.: 85.51 F
01:14:34.219 -> Average Cal sensor Temp.: 94.68 Ambient Temp.: 85.51 F
01:14:35.763 -> Average Cal sensor Temp.: 94.81 Ambient Temp.: 85.44 F
01:14:37.320 -> Average Cal sensor Temp.: 95.04 Ambient Temp.: 85.44 F
01:14:38.880 -> Average Cal sensor Temp.: 95.00 Ambient Temp.: 85.44 F
01:14:40.436 -> Average Cal sensor Temp.: 94.86 Ambient Temp.: 85.41 F
01:14:41.997 -> Average Cal sensor Temp.: 95.06 Ambient Temp.: 85.44 F
01:14:43.556 -> Average Cal sensor Temp.: 95.03 Ambient Temp.: 85.37 F
01:14:45.112 -> Average Cal sensor Temp.: 95.10 Ambient Temp.: 85.41 F
01:14:46.664 -> Average Cal sensor Temp.: 95.13 Ambient Temp.: 85.37 F
01:14:48.216 -> Average Cal sensor Temp.: 95.35 Ambient Temp.: 85.41 F
01:14:49.771 -> Average Cal sensor Temp.: 94.84 Ambient Temp.: 85.37 F
01:14:51.329 -> Average Cal sensor Temp.: 94.24 Ambient Temp.: 85.33 F
01:14:52.876 -> Average Cal sensor Temp.: 94.64 Ambient Temp.: 85.37 F
01:14:54.426 -> Average Cal sensor Temp.: 94.95 Ambient Temp.: 85.37 F
01:14:55.984 -> Average Cal sensor Temp.: 94.93 Ambient Temp.: 85.41 F
01:14:57.535 -> Average Cal sensor Temp.: 95.26 Ambient Temp.: 85.37 F
01:14:59.086 -> Average Cal sensor Temp.: 95.21 Ambient Temp.: 85.37 F
01:15:00.640 -> Average Cal sensor Temp.: 94.89 Ambient Temp.: 85.37 F
01:15:02.192 -> Average Cal sensor Temp.: 94.65 Ambient Temp.: 85.37 F
01:15:03.752 -> Average Cal sensor Temp.: 94.59 Ambient Temp.: 85.37 F
01:15:05.308 -> Average Cal sensor Temp.: 94.77 Ambient Temp.: 85.33 F
01:15:06.864 -> Average Cal sensor Temp.: 94.83 Ambient Temp.: 85.33 F
01:15:08.423 -> Average Cal sensor Temp.: 94.71 Ambient Temp.: 85.33 F
01:15:09.981 -> Average Cal sensor Temp.: 94.67 Ambient Temp.: 85.30 F
01:15:11.532 -> Average Cal sensor Temp.: 95.04 Ambient Temp.: 85.30 F
01:15:13.086 -> Average Cal sensor Temp.: 94.83 Ambient Temp.: 85.19 F
01:15:14.645 -> Average Cal sensor Temp.: 94.83 Ambient Temp.: 85.19 F
01:15:16.226 -> Average Cal sensor Temp.: 94.65 Ambient Temp.: 85.19 F
01:15:17.780 -> Average Cal sensor Temp.: 94.82 Ambient Temp.: 85.19 F
01:15:19.306 -> Average Cal sensor Temp.: 95.00 Ambient Temp.: 85.19 F
01:15:20.865 -> Average Cal sensor Temp.: 95.13 Ambient Temp.: 85.12 F
01:15:22.425 -> Average Cal sensor Temp.: 95.46 Ambient Temp.: 85.19 F
01:15:23.980 -> Average Cal sensor Temp.: 95.73 Ambient Temp.: 85.12 F
01:15:25.533 -> Average Cal sensor Temp.: 95.27 Ambient Temp.: 85.19 F
01:15:27.113 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 85.15 F
01:15:28.639 -> Average Cal sensor Temp.: 95.49 Ambient Temp.: 85.19 F
01:15:30.224 -> Average Cal sensor Temp.: 95.66 Ambient Temp.: 85.19 F
01:15:31.774 -> Average Cal sensor Temp.: 95.72 Ambient Temp.: 85.19 F
01:15:33.324 -> Average Cal sensor Temp.: 95.46 Ambient Temp.: 85.26 F
01:15:34.882 -> Average Cal sensor Temp.: 95.49 Ambient Temp.: 85.26 F
01:15:36.435 -> Average Cal sensor Temp.: 95.66 Ambient Temp.: 85.26 F
01:15:37.984 -> Average Cal sensor Temp.: 95.90 Ambient Temp.: 85.30 F
01:15:39.540 -> Average Cal sensor Temp.: 95.89 Ambient Temp.: 85.30 F
01:15:41.099 -> Average Cal sensor Temp.: 95.49 Ambient Temp.: 85.30 F
01:15:42.645 -> Average Cal sensor Temp.: 95.62 Ambient Temp.: 85.33 F
01:15:44.203 -> Average Cal sensor Temp.: 95.80 Ambient Temp.: 85.41 F
01:15:45.784 -> Average Cal sensor Temp.: 95.87 Ambient Temp.: 85.37 F
01:15:47.339 -> Average Cal sensor Temp.: 96.29 Ambient Temp.: 85.44 F
01:15:48.891 -> Average Cal sensor Temp.: 95.47 Ambient Temp.: 85.41 F
01:15:50.446 -> Average Cal sensor Temp.: 95.26 Ambient Temp.: 85.44 F
01:15:51.995 -> Average Cal sensor Temp.: 95.50 Ambient Temp.: 85.44 F
01:15:53.552 -> Average Cal sensor Temp.: 95.73 Ambient Temp.: 85.48 F
01:15:55.110 -> Average Cal sensor Temp.: 95.48 Ambient Temp.: 85.48 F
01:15:56.663 -> Average Cal sensor Temp.: 95.29 Ambient Temp.: 85.48 F
01:15:58.215 -> Average Cal sensor Temp.: 95.55 Ambient Temp.: 85.51 F
01:15:59.772 -> Average Cal sensor Temp.: 95.73 Ambient Temp.: 85.55 F
01:16:01.322 -> Average Cal sensor Temp.: 96.11 Ambient Temp.: 85.59 F
01:16:02.880 -> Average Cal sensor Temp.: 95.61 Ambient Temp.: 85.55 F
01:16:04.437 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 85.66 F
01:16:05.993 -> Average Cal sensor Temp.: 95.60 Ambient Temp.: 85.66 F
01:16:07.549 -> Average Cal sensor Temp.: 95.92 Ambient Temp.: 85.66 F
01:16:09.104 -> Average Cal sensor Temp.: 95.85 Ambient Temp.: 85.66 F
01:16:10.660 -> Average Cal sensor Temp.: 95.72 Ambient Temp.: 85.69 F
01:16:12.210 -> Average Cal sensor Temp.: 95.36 Ambient Temp.: 85.69 F
01:16:13.759 -> Average Cal sensor Temp.: 95.59 Ambient Temp.: 85.73 F
01:16:15.316 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 85.66 F
01:16:16.869 -> Average Cal sensor Temp.: 96.03 Ambient Temp.: 85.69 F
01:16:18.426 -> Average Cal sensor Temp.: 96.11 Ambient Temp.: 85.77 F
01:16:19.986 -> Average Cal sensor Temp.: 96.10 Ambient Temp.: 85.77 F
01:16:21.539 -> Average Cal sensor Temp.: 96.51 Ambient Temp.: 85.80 F
01:16:23.095 -> Average Cal sensor Temp.: 95.90 Ambient Temp.: 85.80 F
01:16:24.651 -> Average Cal sensor Temp.: 95.41 Ambient Temp.: 85.80 F
01:16:26.208 -> Average Cal sensor Temp.: 95.54 Ambient Temp.: 85.73 F
01:16:27.761 -> Average Cal sensor Temp.: 95.95 Ambient Temp.: 85.80 F
01:16:29.342 -> Average Cal sensor Temp.: 95.96 Ambient Temp.: 85.84 F
01:16:30.896 -> Average Cal sensor Temp.: 95.98 Ambient Temp.: 85.84 F
01:16:32.445 -> Average Cal sensor Temp.: 95.82 Ambient Temp.: 85.84 F
01:16:33.997 -> Average Cal sensor Temp.: 95.95 Ambient Temp.: 85.87 F
01:16:35.556 -> Average Cal sensor Temp.: 95.96 Ambient Temp.: 85.91 F
01:16:37.114 -> Average Cal sensor Temp.: 95.74 Ambient Temp.: 85.91 F
01:16:38.671 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 85.91 F
01:16:40.225 -> Average Cal sensor Temp.: 95.88 Ambient Temp.: 85.91 F
01:16:41.776 -> Average Cal sensor Temp.: 96.04 Ambient Temp.: 85.95 F
01:16:43.332 -> Average Cal sensor Temp.: 96.04 Ambient Temp.: 85.95 F
01:16:44.886 -> Average Cal sensor Temp.: 95.92 Ambient Temp.: 85.95 F
01:16:46.437 -> Average Cal sensor Temp.: 95.98 Ambient Temp.: 86.05 F
01:16:48.024 -> Average Cal sensor Temp.: 95.97 Ambient Temp.: 85.98 F
01:16:49.578 -> Average Cal sensor Temp.: 96.13 Ambient Temp.: 86.09 F
01:16:51.128 -> Average Cal sensor Temp.: 96.03 Ambient Temp.: 86.05 F
01:16:52.684 -> Average Cal sensor Temp.: 95.50 Ambient Temp.: 85.98 F
01:16:54.242 -> Average Cal sensor Temp.: 95.66 Ambient Temp.: 86.05 F
01:16:55.797 -> Average Cal sensor Temp.: 96.04 Ambient Temp.: 86.09 F
01:16:57.352 -> Average Cal sensor Temp.: 95.78 Ambient Temp.: 86.09 F
01:16:58.909 -> Average Cal sensor Temp.: 95.91 Ambient Temp.: 85.98 F
01:17:00.463 -> Average Cal sensor Temp.: 96.07 Ambient Temp.: 86.13 F
01:17:02.015 -> Average Cal sensor Temp.: 96.12 Ambient Temp.: 86.09 F
01:17:03.579 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 86.16 F
01:17:05.135 -> Average Cal sensor Temp.: 96.09 Ambient Temp.: 86.13 F
01:17:06.693 -> Average Cal sensor Temp.: 96.13 Ambient Temp.: 86.16 F
01:17:08.249 -> Average Cal sensor Temp.: 96.12 Ambient Temp.: 86.13 F
01:17:09.799 -> Average Cal sensor Temp.: 95.99 Ambient Temp.: 86.13 F
01:17:11.350 -> Average Cal sensor Temp.: 96.14 Ambient Temp.: 86.16 F
01:17:12.902 -> Average Cal sensor Temp.: 96.05 Ambient Temp.: 86.16 F
01:17:14.458 -> Average Cal sensor Temp.: 96.25 Ambient Temp.: 86.16 F
01:17:16.045 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 86.20 F
01:17:17.568 -> Average Cal sensor Temp.: 95.98 Ambient Temp.: 86.16 F
01:17:19.155 -> Average Cal sensor Temp.: 96.13 Ambient Temp.: 86.16 F
01:17:20.680 -> Average Cal sensor Temp.: 95.96 Ambient Temp.: 86.23 F
01:17:22.267 -> Average Cal sensor Temp.: 96.14 Ambient Temp.: 86.20 F
01:17:23.822 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 86.23 F
01:17:25.380 -> Average Cal sensor Temp.: 96.50 Ambient Temp.: 86.23 F
01:17:26.932 -> Average Cal sensor Temp.: 96.28 Ambient Temp.: 86.23 F
01:17:28.480 -> Average Cal sensor Temp.: 96.19 Ambient Temp.: 86.23 F
01:17:30.021 -> Average Cal sensor Temp.: 95.95 Ambient Temp.: 86.27 F
01:17:31.608 -> Average Cal sensor Temp.: 96.21 Ambient Temp.: 86.23 F
01:17:33.163 -> Average Cal sensor Temp.: 96.27 Ambient Temp.: 86.27 F
01:17:34.687 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 86.20 F
01:17:36.275 -> Average Cal sensor Temp.: 96.01 Ambient Temp.: 86.27 F
01:17:37.831 -> Average Cal sensor Temp.: 96.09 Ambient Temp.: 86.27 F
01:17:39.390 -> Average Cal sensor Temp.: 96.38 Ambient Temp.: 86.31 F
01:17:40.937 -> Average Cal sensor Temp.: 97.14 Ambient Temp.: 86.31 F
01:17:42.485 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 86.27 F
01:17:44.046 -> Average Cal sensor Temp.: 96.71 Ambient Temp.: 86.34 F
01:17:45.602 -> Average Cal sensor Temp.: 96.22 Ambient Temp.: 86.34 F
01:17:47.147 -> Average Cal sensor Temp.: 96.07 Ambient Temp.: 86.38 F
01:17:48.701 -> Average Cal sensor Temp.: 96.04 Ambient Temp.: 86.38 F
01:17:50.287 -> Average Cal sensor Temp.: 96.16 Ambient Temp.: 86.38 F
01:17:51.843 -> Average Cal sensor Temp.: 96.37 Ambient Temp.: 86.38 F
01:17:53.392 -> Average Cal sensor Temp.: 96.47 Ambient Temp.: 86.49 F
01:17:54.944 -> Average Cal sensor Temp.: 96.16 Ambient Temp.: 86.34 F
01:17:56.502 -> Average Cal sensor Temp.: 96.58 Ambient Temp.: 86.38 F
01:17:58.055 -> Average Cal sensor Temp.: 96.33 Ambient Temp.: 86.38 F
01:17:59.615 -> Average Cal sensor Temp.: 96.20 Ambient Temp.: 86.45 F
01:18:01.171 -> Average Cal sensor Temp.: 96.38 Ambient Temp.: 86.38 F
01:18:02.725 -> Average Cal sensor Temp.: 96.35 Ambient Temp.: 86.45 F
01:18:04.279 -> Average Cal sensor Temp.: 96.18 Ambient Temp.: 86.45 F
01:18:05.838 -> Average Cal sensor Temp.: 96.25 Ambient Temp.: 86.45 F
01:18:07.388 -> Average Cal sensor Temp.: 96.26 Ambient Temp.: 86.38 F
01:18:08.972 -> Average Cal sensor Temp.: 96.34 Ambient Temp.: 86.45 F
01:18:10.525 -> Average Cal sensor Temp.: 96.51 Ambient Temp.: 86.45 F
01:18:12.069 -> Average Cal sensor Temp.: 96.57 Ambient Temp.: 86.49 F
01:18:13.627 -> Average Cal sensor Temp.: 96.05 Ambient Temp.: 86.38 F
01:18:15.182 -> Average Cal sensor Temp.: 95.65 Ambient Temp.: 86.45 F
01:18:16.737 -> Average Cal sensor Temp.: 95.59 Ambient Temp.: 86.38 F
01:18:18.301 -> Average Cal sensor Temp.: 95.78 Ambient Temp.: 86.38 F
01:18:19.851 -> Average Cal sensor Temp.: 95.79 Ambient Temp.: 86.31 F
01:18:21.409 -> Average Cal sensor Temp.: 95.91 Ambient Temp.: 86.34 F
01:18:22.962 -> Average Cal sensor Temp.: 96.22 Ambient Temp.: 86.38 F
01:18:24.524 -> Average Cal sensor Temp.: 95.95 Ambient Temp.: 86.27 F
01:18:26.083 -> Average Cal sensor Temp.: 96.03 Ambient Temp.: 86.31 F
01:18:27.635 -> Average Cal sensor Temp.: 96.48 Ambient Temp.: 86.38 F
01:18:29.186 -> Average Cal sensor Temp.: 96.02 Ambient Temp.: 86.31 F
01:18:30.744 -> Average Cal sensor Temp.: 95.57 Ambient Temp.: 86.31 F
01:18:32.297 -> Average Cal sensor Temp.: 95.76 Ambient Temp.: 86.31 F
01:18:33.853 -> Average Cal sensor Temp.: 95.92 Ambient Temp.: 86.31 F
01:18:35.411 -> Average Cal sensor Temp.: 96.01 Ambient Temp.: 86.23 F
01:18:36.965 -> Average Cal sensor Temp.: 95.89 Ambient Temp.: 86.31 F
01:18:38.518 -> Average Cal sensor Temp.: 95.96 Ambient Temp.: 86.27 F
01:18:40.073 -> Average Cal sensor Temp.: 95.95 Ambient Temp.: 86.27 F
01:18:41.627 -> Average Cal sensor Temp.: 96.09 Ambient Temp.: 86.27 F
01:18:43.183 -> Average Cal sensor Temp.: 96.15 Ambient Temp.: 86.31 F
01:18:44.739 -> Average Cal sensor Temp.: 96.39 Ambient Temp.: 86.23 F
01:18:46.292 -> Average Cal sensor Temp.: 96.31 Ambient Temp.: 86.27 F
01:18:47.877 -> Average Cal sensor Temp.: 96.49 Ambient Temp.: 86.31 F
01:18:49.430 -> Average Cal sensor Temp.: 96.72 Ambient Temp.: 86.31 F
01:18:50.984 -> Average Cal sensor Temp.: 96.37 Ambient Temp.: 86.27 F
01:18:52.536 -> Average Cal sensor Temp.: 96.22 Ambient Temp.: 86.23 F
01:18:54.084 -> Average Cal sensor Temp.: 96.33 Ambient Temp.: 86.27 F
01:18:55.640 -> Average Cal sensor Temp.: 96.53 Ambient Temp.: 86.27 F
01:18:57.189 -> Average Cal sensor Temp.: 96.35 Ambient Temp.: 86.23 F
01:18:58.746 -> Average Cal sensor Temp.: 96.42 Ambient Temp.: 86.31 F
01:19:00.304 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 86.31 F
01:19:01.860 -> Average Cal sensor Temp.: 96.04 Ambient Temp.: 86.27 F
01:19:03.415 -> Average Cal sensor Temp.: 96.41 Ambient Temp.: 86.23 F
01:19:04.972 -> Average Cal sensor Temp.: 96.23 Ambient Temp.: 86.23 F
01:19:06.525 -> Average Cal sensor Temp.: 96.69 Ambient Temp.: 86.20 F
01:19:08.107 -> Average Cal sensor Temp.: 96.38 Ambient Temp.: 86.23 F
01:19:09.661 -> Average Cal sensor Temp.: 96.19 Ambient Temp.: 86.23 F
01:19:11.212 -> Average Cal sensor Temp.: 96.24 Ambient Temp.: 86.23 F
01:19:12.762 -> Average Cal sensor Temp.: 96.43 Ambient Temp.: 86.20 F
01:19:14.315 -> Average Cal sensor Temp.: 96.70 Ambient Temp.: 86.23 F
01:19:15.870 -> Average Cal sensor Temp.: 96.24 Ambient Temp.: 86.20 F
01:19:17.427 -> Average Cal sensor Temp.: 96.56 Ambient Temp.: 86.20 F
01:19:19.008 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 86.23 F
01:19:20.559 -> Average Cal sensor Temp.: 96.65 Ambient Temp.: 86.20 F
01:19:22.117 -> Average Cal sensor Temp.: 96.04 Ambient Temp.: 86.20 F
01:19:23.675 -> Average Cal sensor Temp.: 95.88 Ambient Temp.: 86.16 F
01:19:25.227 -> Average Cal sensor Temp.: 96.15 Ambient Temp.: 86.13 F
01:19:26.784 -> Average Cal sensor Temp.: 96.54 Ambient Temp.: 86.16 F
01:19:28.339 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 86.09 F
01:19:29.888 -> Average Cal sensor Temp.: 96.10 Ambient Temp.: 86.09 F
01:19:31.464 -> Average Cal sensor Temp.: 96.01 Ambient Temp.: 86.05 F
01:19:33.015 -> Average Cal sensor Temp.: 96.01 Ambient Temp.: 85.98 F
01:19:34.569 -> Average Cal sensor Temp.: 96.07 Ambient Temp.: 85.95 F
01:19:36.125 -> Average Cal sensor Temp.: 96.53 Ambient Temp.: 85.95 F
01:19:37.676 -> Average Cal sensor Temp.: 96.48 Ambient Temp.: 85.95 F
01:19:39.237 -> Average Cal sensor Temp.: 96.62 Ambient Temp.: 85.95 F
01:19:40.795 -> Average Cal sensor Temp.: 96.53 Ambient Temp.: 85.95 F
01:19:42.351 -> Average Cal sensor Temp.: 96.51 Ambient Temp.: 85.91 F
01:19:43.908 -> Average Cal sensor Temp.: 96.62 Ambient Temp.: 85.95 F
01:19:45.463 -> Average Cal sensor Temp.: 96.61 Ambient Temp.: 85.95 F
01:19:47.020 -> Average Cal sensor Temp.: 96.56 Ambient Temp.: 85.91 F
01:19:48.577 -> Average Cal sensor Temp.: 96.55 Ambient Temp.: 85.95 F
01:19:50.129 -> Average Cal sensor Temp.: 96.55 Ambient Temp.: 85.95 F
01:19:51.689 -> Average Cal sensor Temp.: 96.48 Ambient Temp.: 85.87 F
01:19:53.245 -> Average Cal sensor Temp.: 96.45 Ambient Temp.: 85.87 F
01:19:54.799 -> Average Cal sensor Temp.: 96.62 Ambient Temp.: 85.91 F
01:19:56.354 -> Average Cal sensor Temp.: 96.27 Ambient Temp.: 85.91 F
01:19:57.910 -> Average Cal sensor Temp.: 96.78 Ambient Temp.: 85.95 F
01:19:59.466 -> Average Cal sensor Temp.: 96.25 Ambient Temp.: 85.87 F
01:20:01.021 -> Average Cal sensor Temp.: 96.10 Ambient Temp.: 85.91 F
01:20:02.574 -> Average Cal sensor Temp.: 96.53 Ambient Temp.: 85.91 F
01:20:04.130 -> Average Cal sensor Temp.: 96.27 Ambient Temp.: 85.91 F
01:20:05.691 -> Average Cal sensor Temp.: 96.11 Ambient Temp.: 85.95 F
01:20:07.248 -> Average Cal sensor Temp.: 96.38 Ambient Temp.: 85.87 F
01:20:08.803 -> Average Cal sensor Temp.: 96.12 Ambient Temp.: 85.87 F
01:20:10.356 -> Average Cal sensor Temp.: 96.02 Ambient Temp.: 85.80 F
01:20:11.914 -> Average Cal sensor Temp.: 96.23 Ambient Temp.: 85.84 F
01:20:13.471 -> Average Cal sensor Temp.: 96.02 Ambient Temp.: 85.80 F
01:20:15.025 -> Average Cal sensor Temp.: 96.12 Ambient Temp.: 85.84 F
01:20:16.582 -> Average Cal sensor Temp.: 96.76 Ambient Temp.: 85.77 F
01:20:18.137 -> Average Cal sensor Temp.: 96.84 Ambient Temp.: 85.80 F
01:20:19.689 -> Average Cal sensor Temp.: 96.89 Ambient Temp.: 85.87 F
01:20:21.246 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 85.77 F
01:20:22.797 -> Average Cal sensor Temp.: 96.38 Ambient Temp.: 85.80 F
01:20:24.353 -> Average Cal sensor Temp.: 96.58 Ambient Temp.: 85.80 F
01:20:25.935 -> Average Cal sensor Temp.: 96.48 Ambient Temp.: 85.84 F
01:20:27.487 -> Average Cal sensor Temp.: 96.56 Ambient Temp.: 85.84 F
01:20:29.041 -> Average Cal sensor Temp.: 96.61 Ambient Temp.: 85.80 F
01:20:30.601 -> Average Cal sensor Temp.: 96.81 Ambient Temp.: 85.80 F
01:20:32.150 -> Average Cal sensor Temp.: 96.46 Ambient Temp.: 85.80 F
01:20:33.704 -> Average Cal sensor Temp.: 96.44 Ambient Temp.: 85.77 F
01:20:35.255 -> Average Cal sensor Temp.: 96.73 Ambient Temp.: 85.77 F
01:20:36.815 -> Average Cal sensor Temp.: 96.66 Ambient Temp.: 85.87 F
01:20:38.366 -> Average Cal sensor Temp.: 96.61 Ambient Temp.: 85.80 F
01:20:39.923 -> Average Cal sensor Temp.: 96.23 Ambient Temp.: 85.84 F
01:20:41.476 -> Average Cal sensor Temp.: 96.63 Ambient Temp.: 85.84 F
01:20:43.066 -> Average Cal sensor Temp.: 96.62 Ambient Temp.: 85.77 F
01:20:44.617 -> Average Cal sensor Temp.: 97.04 Ambient Temp.: 85.80 F
01:20:46.169 -> Average Cal sensor Temp.: 96.89 Ambient Temp.: 85.84 F
01:20:47.725 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 85.87 F
01:20:49.278 -> Average Cal sensor Temp.: 96.52 Ambient Temp.: 85.87 F
01:20:50.833 -> Average Cal sensor Temp.: 96.55 Ambient Temp.: 85.87 F
01:20:52.384 -> Average Cal sensor Temp.: 96.80 Ambient Temp.: 85.87 F
01:20:53.939 -> Average Cal sensor Temp.: 96.37 Ambient Temp.: 85.87 F
01:20:55.490 -> Average Cal sensor Temp.: 97.15 Ambient Temp.: 85.91 F
01:20:57.048 -> Average Cal sensor Temp.: 96.72 Ambient Temp.: 85.91 F
01:20:58.601 -> Average Cal sensor Temp.: 96.73 Ambient Temp.: 85.95 F
01:21:00.152 -> Average Cal sensor Temp.: 97.07 Ambient Temp.: 85.91 F
01:21:01.709 -> Average Cal sensor Temp.: 96.41 Ambient Temp.: 85.98 F
01:21:03.273 -> Average Cal sensor Temp.: 96.60 Ambient Temp.: 85.98 F
01:21:04.827 -> Average Cal sensor Temp.: 97.04 Ambient Temp.: 86.05 F
01:21:06.379 -> Average Cal sensor Temp.: 96.90 Ambient Temp.: 85.98 F
01:21:07.937 -> Average Cal sensor Temp.: 96.66 Ambient Temp.: 85.95 F
01:21:09.494 -> Average Cal sensor Temp.: 96.88 Ambient Temp.: 85.98 F
01:21:11.084 -> Average Cal sensor Temp.: 96.47 Ambient Temp.: 85.95 F
01:21:12.636 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 86.05 F
01:21:14.195 -> Average Cal sensor Temp.: 96.69 Ambient Temp.: 85.98 F
01:21:15.753 -> Average Cal sensor Temp.: 96.64 Ambient Temp.: 86.05 F
01:21:17.305 -> Average Cal sensor Temp.: 96.79 Ambient Temp.: 86.05 F
01:21:18.858 -> Average Cal sensor Temp.: 96.53 Ambient Temp.: 85.98 F
01:21:20.416 -> Average Cal sensor Temp.: 96.74 Ambient Temp.: 85.98 F
01:21:21.974 -> Average Cal sensor Temp.: 97.29 Ambient Temp.: 86.09 F
01:21:23.536 -> Average Cal sensor Temp.: 96.10 Ambient Temp.: 85.98 F
01:21:25.090 -> Average Cal sensor Temp.: 96.34 Ambient Temp.: 85.95 F
01:21:26.649 -> Average Cal sensor Temp.: 96.23 Ambient Temp.: 85.95 F
01:21:28.204 -> Average Cal sensor Temp.: 96.16 Ambient Temp.: 85.91 F
01:21:29.754 -> Average Cal sensor Temp.: 96.46 Ambient Temp.: 85.87 F
01:21:31.311 -> Average Cal sensor Temp.: 96.64 Ambient Temp.: 85.87 F
01:21:32.862 -> Average Cal sensor Temp.: 96.39 Ambient Temp.: 85.95 F
01:21:34.418 -> Average Cal sensor Temp.: 96.54 Ambient Temp.: 85.95 F
01:21:35.970 -> Average Cal sensor Temp.: 96.84 Ambient Temp.: 85.91 F
01:21:37.521 -> Average Cal sensor Temp.: 97.11 Ambient Temp.: 85.87 F
01:21:39.074 -> Average Cal sensor Temp.: 96.82 Ambient Temp.: 85.95 F
01:21:40.661 -> Average Cal sensor Temp.: 96.99 Ambient Temp.: 85.91 F
01:21:42.216 -> Average Cal sensor Temp.: 96.93 Ambient Temp.: 85.87 F
01:21:43.773 -> Average Cal sensor Temp.: 97.04 Ambient Temp.: 85.91 F
01:21:45.330 -> Average Cal sensor Temp.: 97.30 Ambient Temp.: 85.87 F
01:21:46.870 -> Average Cal sensor Temp.: 97.25 Ambient Temp.: 85.91 F
01:21:48.426 -> Average Cal sensor Temp.: 96.93 Ambient Temp.: 85.91 F
01:21:49.974 -> Average Cal sensor Temp.: 96.43 Ambient Temp.: 85.95 F
01:21:51.557 -> Average Cal sensor Temp.: 96.48 Ambient Temp.: 85.87 F
01:21:53.107 -> Average Cal sensor Temp.: 96.78 Ambient Temp.: 85.91 F
01:21:54.661 -> Average Cal sensor Temp.: 96.74 Ambient Temp.: 85.87 F
01:21:56.214 -> Average Cal sensor Temp.: 96.66 Ambient Temp.: 85.87 F
01:21:57.767 -> Average Cal sensor Temp.: 96.64 Ambient Temp.: 85.95 F
01:21:59.327 -> Average Cal sensor Temp.: 97.02 Ambient Temp.: 85.91 F
01:22:00.878 -> Average Cal sensor Temp.: 96.51 Ambient Temp.: 85.91 F
01:22:02.433 -> Average Cal sensor Temp.: 96.71 Ambient Temp.: 85.95 F
01:22:03.995 -> Average Cal sensor Temp.: 97.29 Ambient Temp.: 85.95 F
01:22:05.549 -> Average Cal sensor Temp.: 97.19 Ambient Temp.: 85.95 F
01:22:07.106 -> Average Cal sensor Temp.: 97.42 Ambient Temp.: 86.05 F
01:22:08.666 -> Average Cal sensor Temp.: 97.45 Ambient Temp.: 86.05 F
01:22:10.221 -> Average Cal sensor Temp.: 96.72 Ambient Temp.: 85.98 F
01:22:11.775 -> Average Cal sensor Temp.: 96.80 Ambient Temp.: 86.05 F
01:22:13.334 -> Average Cal sensor Temp.: 96.84 Ambient Temp.: 85.98 F
01:22:14.885 -> Average Cal sensor Temp.: 97.07 Ambient Temp.: 86.05 F
01:22:16.442 -> Average Cal sensor Temp.: 96.99 Ambient Temp.: 86.13 F
01:22:17.988 -> Average Cal sensor Temp.: 96.80 Ambient Temp.: 86.09 F
01:22:19.577 -> Average Cal sensor Temp.: 96.88 Ambient Temp.: 86.16 F
01:22:21.131 -> Average Cal sensor Temp.: 96.71 Ambient Temp.: 86.16 F
01:22:22.684 -> Average Cal sensor Temp.: 96.70 Ambient Temp.: 86.13 F
01:22:24.241 -> Average Cal sensor Temp.: 97.16 Ambient Temp.: 86.13 F
01:22:25.796 -> Average Cal sensor Temp.: 97.35 Ambient Temp.: 86.16 F
01:22:27.350 -> Average Cal sensor Temp.: 97.15 Ambient Temp.: 86.20 F
01:22:28.909 -> Average Cal sensor Temp.: 97.21 Ambient Temp.: 86.13 F
01:22:30.464 -> Average Cal sensor Temp.: 97.14 Ambient Temp.: 86.16 F
01:22:32.020 -> Average Cal sensor Temp.: 97.31 Ambient Temp.: 86.13 F
01:22:33.576 -> Average Cal sensor Temp.: 96.89 Ambient Temp.: 86.16 F
01:22:35.132 -> Average Cal sensor Temp.: 97.06 Ambient Temp.: 86.13 F
01:22:36.688 -> Average Cal sensor Temp.: 97.20 Ambient Temp.: 86.23 F
01:22:38.247 -> Average Cal sensor Temp.: 97.25 Ambient Temp.: 86.23 F
01:22:39.772 -> Average Cal sensor Temp.: 97.83 Ambient Temp.: 86.34 F
01:22:41.335 -> Average Cal sensor Temp.: 97.24 Ambient Temp.: 86.27 F
01:22:42.889 -> Average Cal sensor Temp.: 97.31 Ambient Temp.: 86.27 F
01:22:44.445 -> Average Cal sensor Temp.: 97.30 Ambient Temp.: 86.34 F
01:22:45.998 -> Average Cal sensor Temp.: 97.20 Ambient Temp.: 86.34 F
01:22:47.576 -> Average Cal sensor Temp.: 97.26 Ambient Temp.: 86.34 F
01:22:49.128 -> Average Cal sensor Temp.: 97.34 Ambient Temp.: 86.34 F
01:22:50.684 -> Average Cal sensor Temp.: 97.27 Ambient Temp.: 86.45 F
01:22:52.237 -> Average Cal sensor Temp.: 97.49 Ambient Temp.: 86.49 F
01:22:53.795 -> Average Cal sensor Temp.: 97.14 Ambient Temp.: 86.49 F
01:22:55.354 -> Average Cal sensor Temp.: 97.25 Ambient Temp.: 86.45 F
01:22:56.912 -> Average Cal sensor Temp.: 97.36 Ambient Temp.: 86.52 F
01:22:58.462 -> Average Cal sensor Temp.: 96.93 Ambient Temp.: 86.52 F
01:23:00.016 -> Average Cal sensor Temp.: 97.54 Ambient Temp.: 86.52 F
01:23:01.572 -> Average Cal sensor Temp.: 97.67 Ambient Temp.: 86.56 F
01:23:03.119 -> Average Cal sensor Temp.: 97.56 Ambient Temp.: 86.56 F
01:23:04.675 -> Average Cal sensor Temp.: 97.58 Ambient Temp.: 86.59 F
01:23:06.231 -> Average Cal sensor Temp.: 97.58 Ambient Temp.: 86.67 F
01:23:07.782 -> Average Cal sensor Temp.: 97.36 Ambient Temp.: 86.67 F
01:23:09.338 -> Average Cal sensor Temp.: 97.09 Ambient Temp.: 86.63 F
01:23:10.925 -> Average Cal sensor Temp.: 97.07 Ambient Temp.: 86.70 F
01:23:12.476 -> Average Cal sensor Temp.: 97.21 Ambient Temp.: 86.67 F
01:23:14.035 -> Average Cal sensor Temp.: 97.10 Ambient Temp.: 86.70 F
01:23:15.587 -> Average Cal sensor Temp.: 97.43 Ambient Temp.: 86.70 F
01:23:17.147 -> Average Cal sensor Temp.: 97.40 Ambient Temp.: 86.74 F
01:23:18.701 -> Average Cal sensor Temp.: 97.70 Ambient Temp.: 86.70 F
01:23:20.227 -> Average Cal sensor Temp.: 97.46 Ambient Temp.: 86.88 F
01:23:21.817 -> Average Cal sensor Temp.: 97.81 Ambient Temp.: 86.81 F
01:23:23.341 -> Average Cal sensor Temp.: 97.31 Ambient Temp.: 86.77 F
01:23:24.924 -> Average Cal sensor Temp.: 97.39 Ambient Temp.: 86.77 F
01:23:26.479 -> Average Cal sensor Temp.: 97.69 Ambient Temp.: 86.81 F
01:23:28.024 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 86.88 F
01:23:29.579 -> Average Cal sensor Temp.: 97.89 Ambient Temp.: 86.88 F
01:23:31.137 -> Average Cal sensor Temp.: 97.78 Ambient Temp.: 86.88 F
01:23:32.687 -> Average Cal sensor Temp.: 97.43 Ambient Temp.: 86.92 F
01:23:34.239 -> Average Cal sensor Temp.: 97.40 Ambient Temp.: 86.92 F
01:23:35.822 -> Average Cal sensor Temp.: 97.52 Ambient Temp.: 86.92 F
01:23:37.381 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 86.92 F
01:23:38.930 -> Average Cal sensor Temp.: 97.70 Ambient Temp.: 86.92 F
01:23:40.481 -> Average Cal sensor Temp.: 97.53 Ambient Temp.: 86.92 F
01:23:42.034 -> Average Cal sensor Temp.: 97.50 Ambient Temp.: 86.95 F
01:23:43.588 -> Average Cal sensor Temp.: 97.47 Ambient Temp.: 86.95 F
01:23:45.139 -> Average Cal sensor Temp.: 97.14 Ambient Temp.: 86.99 F
01:23:46.692 -> Average Cal sensor Temp.: 97.38 Ambient Temp.: 87.03 F
01:23:48.252 -> Average Cal sensor Temp.: 97.74 Ambient Temp.: 87.03 F
01:23:49.803 -> Average Cal sensor Temp.: 97.60 Ambient Temp.: 87.03 F
01:23:51.359 -> Average Cal sensor Temp.: 97.35 Ambient Temp.: 86.99 F
01:23:52.943 -> Average Cal sensor Temp.: 97.02 Ambient Temp.: 87.06 F
01:23:54.497 -> Average Cal sensor Temp.: 97.44 Ambient Temp.: 86.99 F
01:23:56.045 -> Average Cal sensor Temp.: 97.36 Ambient Temp.: 86.99 F
01:23:57.593 -> Average Cal sensor Temp.: 96.90 Ambient Temp.: 86.99 F
01:23:59.146 -> Average Cal sensor Temp.: 97.12 Ambient Temp.: 86.99 F
01:24:00.698 -> Average Cal sensor Temp.: 97.28 Ambient Temp.: 86.99 F
01:24:02.255 -> Average Cal sensor Temp.: 97.40 Ambient Temp.: 87.03 F
01:24:03.845 -> Average Cal sensor Temp.: 96.94 Ambient Temp.: 86.99 F
01:24:05.401 -> Average Cal sensor Temp.: 96.71 Ambient Temp.: 86.95 F
01:24:06.954 -> Average Cal sensor Temp.: 97.09 Ambient Temp.: 86.95 F
01:24:08.505 -> Average Cal sensor Temp.: 97.04 Ambient Temp.: 86.95 F
01:24:10.059 -> Average Cal sensor Temp.: 97.50 Ambient Temp.: 86.95 F
01:24:11.621 -> Average Cal sensor Temp.: 96.90 Ambient Temp.: 86.92 F
01:24:13.176 -> Average Cal sensor Temp.: 96.80 Ambient Temp.: 86.92 F
01:24:14.736 -> Average Cal sensor Temp.: 97.11 Ambient Temp.: 86.99 F
01:24:16.292 -> Average Cal sensor Temp.: 96.93 Ambient Temp.: 86.88 F
01:24:17.851 -> Average Cal sensor Temp.: 97.57 Ambient Temp.: 86.81 F
01:24:19.403 -> Average Cal sensor Temp.: 97.16 Ambient Temp.: 86.88 F
01:24:20.959 -> Average Cal sensor Temp.: 97.33 Ambient Temp.: 86.77 F
01:24:22.503 -> Average Cal sensor Temp.: 97.37 Ambient Temp.: 86.81 F
01:24:24.062 -> Average Cal sensor Temp.: 96.97 Ambient Temp.: 86.77 F
01:24:25.606 -> Average Cal sensor Temp.: 97.14 Ambient Temp.: 86.77 F
01:24:27.189 -> Average Cal sensor Temp.: 97.24 Ambient Temp.: 86.74 F
01:24:28.738 -> Average Cal sensor Temp.: 97.21 Ambient Temp.: 86.74 F
01:24:30.289 -> Average Cal sensor Temp.: 97.33 Ambient Temp.: 86.77 F
01:24:31.839 -> Average Cal sensor Temp.: 97.31 Ambient Temp.: 86.74 F
01:24:33.395 -> Average Cal sensor Temp.: 97.33 Ambient Temp.: 86.70 F
01:24:34.947 -> Average Cal sensor Temp.: 97.42 Ambient Temp.: 86.70 F
01:24:36.501 -> Average Cal sensor Temp.: 97.52 Ambient Temp.: 86.70 F
01:24:38.052 -> Average Cal sensor Temp.: 97.64 Ambient Temp.: 86.74 F
01:24:39.611 -> Average Cal sensor Temp.: 97.34 Ambient Temp.: 86.70 F
01:24:41.169 -> Average Cal sensor Temp.: 97.43 Ambient Temp.: 86.74 F
01:24:42.724 -> Average Cal sensor Temp.: 97.30 Ambient Temp.: 86.74 F
01:24:44.275 -> Average Cal sensor Temp.: 97.45 Ambient Temp.: 86.77 F
01:24:45.827 -> Average Cal sensor Temp.: 97.65 Ambient Temp.: 86.81 F
01:24:47.411 -> Average Cal sensor Temp.: 97.34 Ambient Temp.: 86.81 F
01:24:48.969 -> Average Cal sensor Temp.: 97.33 Ambient Temp.: 86.74 F
01:24:50.527 -> Average Cal sensor Temp.: 97.36 Ambient Temp.: 86.77 F
01:24:52.082 -> Average Cal sensor Temp.: 97.51 Ambient Temp.: 86.77 F
01:24:53.607 -> Average Cal sensor Temp.: 97.43 Ambient Temp.: 86.81 F
01:24:55.167 -> Average Cal sensor Temp.: 97.59 Ambient Temp.: 86.77 F
01:24:56.743 -> Average Cal sensor Temp.: 97.86 Ambient Temp.: 86.81 F
01:24:58.304 -> Average Cal sensor Temp.: 97.46 Ambient Temp.: 86.88 F
01:24:59.857 -> Average Cal sensor Temp.: 97.11 Ambient Temp.: 86.81 F
01:25:01.410 -> Average Cal sensor Temp.: 97.13 Ambient Temp.: 86.88 F
01:25:02.965 -> Average Cal sensor Temp.: 97.29 Ambient Temp.: 86.81 F
01:25:04.521 -> Average Cal sensor Temp.: 97.25 Ambient Temp.: 86.88 F
01:25:06.079 -> Average Cal sensor Temp.: 97.81 Ambient Temp.: 86.81 F
01:25:07.636 -> Average Cal sensor Temp.: 97.54 Ambient Temp.: 86.81 F
01:25:09.188 -> Average Cal sensor Temp.: 97.41 Ambient Temp.: 86.81 F
01:25:10.742 -> Average Cal sensor Temp.: 97.21 Ambient Temp.: 86.88 F
01:25:12.297 -> Average Cal sensor Temp.: 97.41 Ambient Temp.: 86.77 F
01:25:13.853 -> Average Cal sensor Temp.: 97.30 Ambient Temp.: 86.74 F
01:25:15.408 -> Average Cal sensor Temp.: 97.27 Ambient Temp.: 86.77 F
01:25:16.962 -> Average Cal sensor Temp.: 97.36 Ambient Temp.: 86.88 F
01:25:18.509 -> Average Cal sensor Temp.: 96.93 Ambient Temp.: 86.88 F
01:25:20.059 -> Average Cal sensor Temp.: 96.80 Ambient Temp.: 86.77 F
01:25:21.650 -> Average Cal sensor Temp.: 97.26 Ambient Temp.: 86.74 F
01:25:23.201 -> Average Cal sensor Temp.: 97.35 Ambient Temp.: 86.77 F
01:25:24.753 -> Average Cal sensor Temp.: 97.63 Ambient Temp.: 86.74 F
01:25:26.311 -> Average Cal sensor Temp.: 97.49 Ambient Temp.: 86.74 F
01:25:27.860 -> Average Cal sensor Temp.: 97.32 Ambient Temp.: 86.70 F
01:25:29.418 -> Average Cal sensor Temp.: 97.38 Ambient Temp.: 86.74 F
01:25:30.979 -> Average Cal sensor Temp.: 98.00 Ambient Temp.: 86.74 F
01:25:32.533 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 86.74 F
01:25:34.091 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 86.77 F
01:25:35.652 -> Average Cal sensor Temp.: 97.38 Ambient Temp.: 86.70 F
01:25:37.203 -> Average Cal sensor Temp.: 97.34 Ambient Temp.: 86.77 F
01:25:38.764 -> Average Cal sensor Temp.: 97.41 Ambient Temp.: 86.74 F
01:25:40.314 -> Average Cal sensor Temp.: 97.29 Ambient Temp.: 86.74 F
01:25:41.875 -> Average Cal sensor Temp.: 97.11 Ambient Temp.: 86.74 F
01:25:43.430 -> Average Cal sensor Temp.: 97.34 Ambient Temp.: 86.74 F
01:25:44.989 -> Average Cal sensor Temp.: 97.45 Ambient Temp.: 86.77 F
01:25:46.542 -> Average Cal sensor Temp.: 97.29 Ambient Temp.: 86.74 F
01:25:48.100 -> Average Cal sensor Temp.: 97.42 Ambient Temp.: 86.70 F
01:25:49.659 -> Average Cal sensor Temp.: 97.43 Ambient Temp.: 86.67 F
01:25:51.210 -> Average Cal sensor Temp.: 97.47 Ambient Temp.: 86.70 F
01:25:52.764 -> Average Cal sensor Temp.: 97.49 Ambient Temp.: 86.67 F
01:25:54.318 -> Average Cal sensor Temp.: 97.93 Ambient Temp.: 86.67 F
01:25:55.874 -> Average Cal sensor Temp.: 97.71 Ambient Temp.: 86.70 F
01:25:57.421 -> Average Cal sensor Temp.: 96.84 Ambient Temp.: 86.63 F
01:25:58.976 -> Average Cal sensor Temp.: 97.36 Ambient Temp.: 86.63 F
01:26:00.528 -> Average Cal sensor Temp.: 97.56 Ambient Temp.: 86.67 F
01:26:02.079 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 86.63 F
01:26:03.632 -> Average Cal sensor Temp.: 97.33 Ambient Temp.: 86.63 F
01:26:05.188 -> Average Cal sensor Temp.: 97.53 Ambient Temp.: 86.63 F
01:26:06.775 -> Average Cal sensor Temp.: 97.76 Ambient Temp.: 86.63 F
01:26:08.329 -> Average Cal sensor Temp.: 97.75 Ambient Temp.: 86.63 F
01:26:09.878 -> Average Cal sensor Temp.: 98.04 Ambient Temp.: 86.59 F
01:26:11.436 -> Average Cal sensor Temp.: 97.65 Ambient Temp.: 86.56 F
01:26:12.993 -> Average Cal sensor Temp.: 97.96 Ambient Temp.: 86.59 F
01:26:14.549 -> Average Cal sensor Temp.: 97.30 Ambient Temp.: 86.59 F
01:26:16.102 -> Average Cal sensor Temp.: 97.20 Ambient Temp.: 86.52 F
01:26:17.655 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 86.59 F
01:26:19.208 -> Average Cal sensor Temp.: 97.74 Ambient Temp.: 86.56 F
01:26:20.764 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 86.59 F
01:26:22.314 -> Average Cal sensor Temp.: 97.68 Ambient Temp.: 86.59 F
01:26:23.899 -> Average Cal sensor Temp.: 97.72 Ambient Temp.: 86.52 F
01:26:25.423 -> Average Cal sensor Temp.: 97.63 Ambient Temp.: 86.56 F
01:26:26.978 -> Average Cal sensor Temp.: 97.46 Ambient Temp.: 86.56 F
01:26:28.541 -> Average Cal sensor Temp.: 97.13 Ambient Temp.: 86.59 F
01:26:30.098 -> Average Cal sensor Temp.: 97.22 Ambient Temp.: 86.56 F
01:26:31.651 -> Average Cal sensor Temp.: 97.10 Ambient Temp.: 86.52 F
01:26:33.203 -> Average Cal sensor Temp.: 97.81 Ambient Temp.: 86.56 F
01:26:34.761 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 86.49 F
01:26:36.317 -> Average Cal sensor Temp.: 97.63 Ambient Temp.: 86.49 F
01:26:37.903 -> Average Cal sensor Temp.: 98.08 Ambient Temp.: 86.49 F
01:26:39.461 -> Average Cal sensor Temp.: 98.28 Ambient Temp.: 86.52 F
01:26:41.017 -> Average Cal sensor Temp.: 98.16 Ambient Temp.: 86.59 F
01:26:42.570 -> Average Cal sensor Temp.: 97.89 Ambient Temp.: 86.56 F
01:26:44.127 -> Average Cal sensor Temp.: 97.62 Ambient Temp.: 86.56 F
01:26:45.679 -> Average Cal sensor Temp.: 97.47 Ambient Temp.: 86.52 F
01:26:47.233 -> Average Cal sensor Temp.: 97.75 Ambient Temp.: 86.56 F
01:26:48.789 -> Average Cal sensor Temp.: 97.78 Ambient Temp.: 86.52 F
01:26:50.351 -> Average Cal sensor Temp.: 97.08 Ambient Temp.: 86.49 F
01:26:51.916 -> Average Cal sensor Temp.: 97.04 Ambient Temp.: 86.49 F
01:26:53.471 -> Average Cal sensor Temp.: 97.25 Ambient Temp.: 86.49 F
01:26:55.023 -> Average Cal sensor Temp.: 96.90 Ambient Temp.: 86.49 F
01:26:56.570 -> Average Cal sensor Temp.: 96.89 Ambient Temp.: 86.38 F
01:26:58.124 -> Average Cal sensor Temp.: 97.05 Ambient Temp.: 86.38 F
01:26:59.673 -> Average Cal sensor Temp.: 97.46 Ambient Temp.: 86.38 F
01:27:01.231 -> Average Cal sensor Temp.: 97.63 Ambient Temp.: 86.38 F
01:27:02.809 -> Average Cal sensor Temp.: 97.39 Ambient Temp.: 86.38 F
01:27:04.351 -> Average Cal sensor Temp.: 97.86 Ambient Temp.: 86.45 F
01:27:05.906 -> Average Cal sensor Temp.: 97.96 Ambient Temp.: 86.34 F
01:27:07.460 -> Average Cal sensor Temp.: 97.97 Ambient Temp.: 86.45 F
01:27:09.018 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 86.38 F
01:27:10.573 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 86.49 F
01:27:12.131 -> Average Cal sensor Temp.: 97.96 Ambient Temp.: 86.52 F
01:27:13.686 -> Average Cal sensor Temp.: 97.83 Ambient Temp.: 86.49 F
01:27:15.272 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 86.52 F
01:27:16.823 -> Average Cal sensor Temp.: 97.85 Ambient Temp.: 86.52 F
01:27:18.375 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 86.52 F
01:27:19.930 -> Average Cal sensor Temp.: 97.73 Ambient Temp.: 86.56 F
01:27:21.489 -> Average Cal sensor Temp.: 97.77 Ambient Temp.: 86.56 F
01:27:23.043 -> Average Cal sensor Temp.: 98.10 Ambient Temp.: 86.56 F
01:27:24.596 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 86.52 F
01:27:26.147 -> Average Cal sensor Temp.: 98.04 Ambient Temp.: 86.56 F
01:27:27.698 -> Average Cal sensor Temp.: 97.95 Ambient Temp.: 86.59 F
01:27:29.280 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 86.63 F
01:27:30.838 -> Average Cal sensor Temp.: 97.86 Ambient Temp.: 86.63 F
01:27:32.366 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 86.67 F
01:27:33.923 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 86.67 F
01:27:35.478 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 86.67 F
01:27:37.031 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 86.70 F
01:27:38.620 -> Average Cal sensor Temp.: 98.05 Ambient Temp.: 86.70 F
01:27:40.171 -> Average Cal sensor Temp.: 98.16 Ambient Temp.: 86.74 F
01:27:41.728 -> Average Cal sensor Temp.: 98.24 Ambient Temp.: 86.77 F
01:27:43.278 -> Average Cal sensor Temp.: 98.04 Ambient Temp.: 86.70 F
01:27:44.829 -> Average Cal sensor Temp.: 98.35 Ambient Temp.: 86.74 F
01:27:46.383 -> Average Cal sensor Temp.: 97.84 Ambient Temp.: 86.74 F
01:27:47.939 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 86.77 F
01:27:49.497 -> Average Cal sensor Temp.: 98.17 Ambient Temp.: 86.74 F
01:27:51.050 -> Average Cal sensor Temp.: 98.28 Ambient Temp.: 86.77 F
01:27:52.604 -> Average Cal sensor Temp.: 98.30 Ambient Temp.: 86.77 F
01:27:54.160 -> Average Cal sensor Temp.: 98.12 Ambient Temp.: 86.88 F
01:27:55.747 -> Average Cal sensor Temp.: 97.97 Ambient Temp.: 86.81 F
01:27:57.300 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 86.81 F
01:27:58.857 -> Average Cal sensor Temp.: 98.10 Ambient Temp.: 86.81 F
01:28:00.410 -> Average Cal sensor Temp.: 98.22 Ambient Temp.: 86.81 F
01:28:01.964 -> Average Cal sensor Temp.: 98.16 Ambient Temp.: 86.81 F
01:28:03.515 -> Average Cal sensor Temp.: 98.24 Ambient Temp.: 86.88 F
01:28:05.071 -> Average Cal sensor Temp.: 98.44 Ambient Temp.: 86.81 F
01:28:06.629 -> Average Cal sensor Temp.: 98.42 Ambient Temp.: 86.88 F
01:28:08.186 -> Average Cal sensor Temp.: 98.00 Ambient Temp.: 86.88 F
01:28:09.738 -> Average Cal sensor Temp.: 98.00 Ambient Temp.: 86.88 F
01:28:11.296 -> Average Cal sensor Temp.: 98.30 Ambient Temp.: 86.88 F
01:28:12.851 -> Average Cal sensor Temp.: 98.25 Ambient Temp.: 86.88 F
01:28:14.407 -> Average Cal sensor Temp.: 98.38 Ambient Temp.: 86.92 F
01:28:15.964 -> Average Cal sensor Temp.: 98.22 Ambient Temp.: 86.95 F
01:28:17.520 -> Average Cal sensor Temp.: 98.19 Ambient Temp.: 86.92 F
01:28:19.077 -> Average Cal sensor Temp.: 98.33 Ambient Temp.: 86.95 F
01:28:20.666 -> Average Cal sensor Temp.: 97.78 Ambient Temp.: 86.99 F
01:28:22.217 -> Average Cal sensor Temp.: 98.04 Ambient Temp.: 86.99 F
01:28:23.771 -> Average Cal sensor Temp.: 98.09 Ambient Temp.: 86.99 F
01:28:25.328 -> Average Cal sensor Temp.: 97.97 Ambient Temp.: 87.03 F
01:28:26.882 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 86.99 F
01:28:28.431 -> Average Cal sensor Temp.: 97.70 Ambient Temp.: 87.03 F
01:28:29.988 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 87.03 F
01:28:31.546 -> Average Cal sensor Temp.: 98.32 Ambient Temp.: 87.03 F
01:28:33.100 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 87.06 F
01:28:34.656 -> Average Cal sensor Temp.: 97.98 Ambient Temp.: 87.03 F
01:28:36.210 -> Average Cal sensor Temp.: 98.13 Ambient Temp.: 87.10 F
01:28:37.763 -> Average Cal sensor Temp.: 97.98 Ambient Temp.: 87.06 F
01:28:39.322 -> Average Cal sensor Temp.: 98.41 Ambient Temp.: 87.06 F
01:28:40.878 -> Average Cal sensor Temp.: 98.53 Ambient Temp.: 87.10 F
01:28:42.444 -> Average Cal sensor Temp.: 98.52 Ambient Temp.: 87.13 F
01:28:44.005 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 87.10 F
01:28:45.567 -> Average Cal sensor Temp.: 98.40 Ambient Temp.: 87.10 F
01:28:47.124 -> Average Cal sensor Temp.: 98.29 Ambient Temp.: 87.06 F
01:28:48.652 -> Average Cal sensor Temp.: 98.31 Ambient Temp.: 87.17 F
01:28:50.206 -> Average Cal sensor Temp.: 98.54 Ambient Temp.: 87.17 F
01:28:51.765 -> Average Cal sensor Temp.: 98.70 Ambient Temp.: 87.13 F
01:28:53.323 -> Average Cal sensor Temp.: 98.51 Ambient Temp.: 87.17 F
01:28:54.872 -> Average Cal sensor Temp.: 98.82 Ambient Temp.: 87.21 F
01:28:56.461 -> Average Cal sensor Temp.: 98.39 Ambient Temp.: 87.21 F
01:28:57.983 -> Average Cal sensor Temp.: 98.46 Ambient Temp.: 87.21 F
01:28:59.573 -> Average Cal sensor Temp.: 98.57 Ambient Temp.: 87.28 F
01:29:01.120 -> Average Cal sensor Temp.: 98.36 Ambient Temp.: 87.21 F
01:29:02.662 -> Average Cal sensor Temp.: 98.46 Ambient Temp.: 87.31 F
01:29:04.215 -> Average Cal sensor Temp.: 97.89 Ambient Temp.: 87.28 F
01:29:05.767 -> Average Cal sensor Temp.: 98.44 Ambient Temp.: 87.28 F
01:29:07.346 -> Average Cal sensor Temp.: 98.67 Ambient Temp.: 87.31 F
01:29:08.896 -> Average Cal sensor Temp.: 98.64 Ambient Temp.: 87.28 F
01:29:10.454 -> Average Cal sensor Temp.: 97.80 Ambient Temp.: 87.31 F
01:29:12.002 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 87.31 F
01:29:13.553 -> Average Cal sensor Temp.: 98.46 Ambient Temp.: 87.28 F
01:29:15.112 -> Average Cal sensor Temp.: 98.77 Ambient Temp.: 87.31 F
01:29:16.666 -> Average Cal sensor Temp.: 98.40 Ambient Temp.: 87.31 F
01:29:18.222 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 87.31 F
01:29:19.802 -> Average Cal sensor Temp.: 98.46 Ambient Temp.: 87.31 F
01:29:21.359 -> Average Cal sensor Temp.: 98.25 Ambient Temp.: 87.35 F
01:29:22.911 -> Average Cal sensor Temp.: 98.03 Ambient Temp.: 87.31 F
01:29:24.462 -> Average Cal sensor Temp.: 97.99 Ambient Temp.: 87.28 F
01:29:26.015 -> Average Cal sensor Temp.: 98.36 Ambient Temp.: 87.28 F
01:29:27.559 -> Average Cal sensor Temp.: 98.05 Ambient Temp.: 87.28 F
01:29:29.115 -> Average Cal sensor Temp.: 98.38 Ambient Temp.: 87.17 F
01:29:30.675 -> Average Cal sensor Temp.: 98.72 Ambient Temp.: 87.21 F
01:29:32.225 -> Average Cal sensor Temp.: 97.85 Ambient Temp.: 87.17 F
01:29:33.780 -> Average Cal sensor Temp.: 98.19 Ambient Temp.: 87.21 F
01:29:35.365 -> Average Cal sensor Temp.: 98.29 Ambient Temp.: 87.17 F
01:29:36.904 -> Average Cal sensor Temp.: 98.16 Ambient Temp.: 87.17 F
01:29:38.458 -> Average Cal sensor Temp.: 97.96 Ambient Temp.: 87.13 F
01:29:40.017 -> Average Cal sensor Temp.: 98.32 Ambient Temp.: 87.13 F
01:29:41.573 -> Average Cal sensor Temp.: 98.45 Ambient Temp.: 87.17 F
01:29:43.130 -> Average Cal sensor Temp.: 98.49 Ambient Temp.: 87.17 F
01:29:44.687 -> Average Cal sensor Temp.: 98.53 Ambient Temp.: 87.17 F
01:29:46.242 -> Average Cal sensor Temp.: 98.50 Ambient Temp.: 87.21 F
01:29:47.793 -> Average Cal sensor Temp.: 97.53 Ambient Temp.: 87.17 F
01:29:49.338 -> Average Cal sensor Temp.: 97.63 Ambient Temp.: 87.17 F
01:29:50.893 -> Average Cal sensor Temp.: 98.14 Ambient Temp.: 87.17 F
01:29:52.443 -> Average Cal sensor Temp.: 98.39 Ambient Temp.: 87.13 F
01:29:54.029 -> Average Cal sensor Temp.: 98.84 Ambient Temp.: 87.17 F
01:29:55.552 -> Average Cal sensor Temp.: 98.26 Ambient Temp.: 87.17 F
01:29:57.142 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 87.13 F
01:29:58.694 -> Average Cal sensor Temp.: 98.29 Ambient Temp.: 87.10 F
01:30:00.244 -> Average Cal sensor Temp.: 98.37 Ambient Temp.: 87.13 F
01:30:01.795 -> Average Cal sensor Temp.: 98.43 Ambient Temp.: 87.13 F
01:30:03.349 -> Average Cal sensor Temp.: 98.60 Ambient Temp.: 87.13 F
01:30:04.897 -> Average Cal sensor Temp.: 98.47 Ambient Temp.: 87.10 F
01:30:06.455 -> Average Cal sensor Temp.: 97.75 Ambient Temp.: 87.13 F
01:30:08.005 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 87.10 F
01:30:09.558 -> Average Cal sensor Temp.: 98.39 Ambient Temp.: 87.10 F
01:30:11.142 -> Average Cal sensor Temp.: 98.38 Ambient Temp.: 87.13 F
01:30:12.690 -> Average Cal sensor Temp.: 98.45 Ambient Temp.: 87.17 F
01:30:14.243 -> Average Cal sensor Temp.: 98.09 Ambient Temp.: 87.13 F
01:30:15.782 -> Average Cal sensor Temp.: 97.74 Ambient Temp.: 87.10 F
01:30:17.335 -> Average Cal sensor Temp.: 98.05 Ambient Temp.: 87.10 F
01:30:18.898 -> Average Cal sensor Temp.: 98.31 Ambient Temp.: 87.10 F
01:30:20.454 -> Average Cal sensor Temp.: 98.56 Ambient Temp.: 87.10 F
01:30:22.005 -> Average Cal sensor Temp.: 98.21 Ambient Temp.: 87.06 F
01:30:23.558 -> Average Cal sensor Temp.: 98.12 Ambient Temp.: 87.10 F
01:30:25.111 -> Average Cal sensor Temp.: 98.51 Ambient Temp.: 87.10 F
01:30:26.693 -> Average Cal sensor Temp.: 98.13 Ambient Temp.: 87.10 F
01:30:28.254 -> Average Cal sensor Temp.: 97.83 Ambient Temp.: 87.10 F
01:30:29.778 -> Average Cal sensor Temp.: 98.37 Ambient Temp.: 87.03 F
01:30:31.335 -> Average Cal sensor Temp.: 98.56 Ambient Temp.: 87.10 F
01:30:32.919 -> Average Cal sensor Temp.: 98.17 Ambient Temp.: 87.03 F
01:30:34.474 -> Average Cal sensor Temp.: 98.25 Ambient Temp.: 87.10 F
01:30:36.030 -> Average Cal sensor Temp.: 97.95 Ambient Temp.: 87.10 F
01:30:37.583 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 87.03 F
01:30:39.136 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 87.06 F
01:30:40.694 -> Average Cal sensor Temp.: 98.19 Ambient Temp.: 86.99 F
01:30:42.246 -> Average Cal sensor Temp.: 98.23 Ambient Temp.: 87.06 F
01:30:43.804 -> Average Cal sensor Temp.: 97.93 Ambient Temp.: 86.99 F
01:30:45.364 -> Average Cal sensor Temp.: 97.99 Ambient Temp.: 86.95 F
01:30:46.914 -> Average Cal sensor Temp.: 98.37 Ambient Temp.: 86.95 F
01:30:48.479 -> Average Cal sensor Temp.: 98.30 Ambient Temp.: 86.99 F
01:30:50.036 -> Average Cal sensor Temp.: 97.99 Ambient Temp.: 86.95 F
01:30:51.588 -> Average Cal sensor Temp.: 98.58 Ambient Temp.: 86.95 F
01:30:53.147 -> Average Cal sensor Temp.: 98.35 Ambient Temp.: 86.99 F
01:30:54.706 -> Average Cal sensor Temp.: 98.10 Ambient Temp.: 86.95 F
01:30:56.262 -> Average Cal sensor Temp.: 98.12 Ambient Temp.: 86.92 F
01:30:57.817 -> Average Cal sensor Temp.: 98.27 Ambient Temp.: 86.95 F
01:30:59.376 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 86.92 F
01:31:00.935 -> Average Cal sensor Temp.: 98.46 Ambient Temp.: 86.92 F
01:31:02.492 -> Average Cal sensor Temp.: 98.48 Ambient Temp.: 86.88 F
01:31:04.043 -> Average Cal sensor Temp.: 98.66 Ambient Temp.: 86.92 F
01:31:05.596 -> Average Cal sensor Temp.: 98.45 Ambient Temp.: 86.92 F
01:31:07.174 -> Average Cal sensor Temp.: 98.38 Ambient Temp.: 86.92 F
01:31:08.730 -> Average Cal sensor Temp.: 98.75 Ambient Temp.: 86.95 F
01:31:10.280 -> Average Cal sensor Temp.: 97.85 Ambient Temp.: 86.92 F
01:31:11.838 -> Average Cal sensor Temp.: 98.05 Ambient Temp.: 86.92 F
01:31:13.393 -> Average Cal sensor Temp.: 98.47 Ambient Temp.: 86.88 F
01:31:14.937 -> Average Cal sensor Temp.: 97.93 Ambient Temp.: 86.88 F
01:31:16.520 -> Average Cal sensor Temp.: 97.97 Ambient Temp.: 86.95 F
01:31:18.068 -> Average Cal sensor Temp.: 98.04 Ambient Temp.: 86.88 F
01:31:19.623 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 86.92 F
01:31:21.182 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 86.88 F
01:31:22.728 -> Average Cal sensor Temp.: 98.18 Ambient Temp.: 86.88 F
01:31:24.283 -> Average Cal sensor Temp.: 98.33 Ambient Temp.: 86.81 F
01:31:25.838 -> Average Cal sensor Temp.: 98.31 Ambient Temp.: 86.88 F
01:31:27.422 -> Average Cal sensor Temp.: 98.22 Ambient Temp.: 86.81 F
01:31:28.966 -> Average Cal sensor Temp.: 98.19 Ambient Temp.: 86.88 F
01:31:30.521 -> Average Cal sensor Temp.: 98.59 Ambient Temp.: 86.81 F
01:31:32.078 -> Average Cal sensor Temp.: 98.77 Ambient Temp.: 86.88 F
01:31:33.630 -> Average Cal sensor Temp.: 97.93 Ambient Temp.: 86.81 F
01:31:35.179 -> Average Cal sensor Temp.: 97.51 Ambient Temp.: 86.77 F
01:31:36.764 -> Average Cal sensor Temp.: 98.12 Ambient Temp.: 86.77 F
01:31:38.312 -> Average Cal sensor Temp.: 98.33 Ambient Temp.: 86.77 F
01:31:39.872 -> Average Cal sensor Temp.: 98.18 Ambient Temp.: 86.70 F
01:31:41.423 -> Average Cal sensor Temp.: 98.45 Ambient Temp.: 86.70 F
01:31:42.972 -> Average Cal sensor Temp.: 98.26 Ambient Temp.: 86.70 F
01:31:44.523 -> Average Cal sensor Temp.: 98.41 Ambient Temp.: 86.70 F
01:31:46.077 -> Average Cal sensor Temp.: 98.49 Ambient Temp.: 86.67 F
01:31:47.665 -> Average Cal sensor Temp.: 98.59 Ambient Temp.: 86.67 F
01:31:49.215 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 86.63 F
01:31:50.772 -> Average Cal sensor Temp.: 98.39 Ambient Temp.: 86.67 F
01:31:52.328 -> Average Cal sensor Temp.: 98.19 Ambient Temp.: 86.67 F
01:31:53.887 -> Average Cal sensor Temp.: 97.74 Ambient Temp.: 86.63 F
01:31:55.443 -> Average Cal sensor Temp.: 98.03 Ambient Temp.: 86.56 F
01:31:57.002 -> Average Cal sensor Temp.: 98.09 Ambient Temp.: 86.59 F
01:31:58.560 -> Average Cal sensor Temp.: 98.27 Ambient Temp.: 86.56 F
01:32:00.116 -> Average Cal sensor Temp.: 98.08 Ambient Temp.: 86.56 F
01:32:01.668 -> Average Cal sensor Temp.: 98.09 Ambient Temp.: 86.56 F
01:32:03.224 -> Average Cal sensor Temp.: 97.65 Ambient Temp.: 86.52 F
01:32:04.779 -> Average Cal sensor Temp.: 97.85 Ambient Temp.: 86.56 F
01:32:06.336 -> Average Cal sensor Temp.: 97.37 Ambient Temp.: 86.45 F
01:32:07.895 -> Average Cal sensor Temp.: 98.08 Ambient Temp.: 86.45 F
01:32:09.448 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 86.45 F
01:32:11.012 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 86.38 F
01:32:12.551 -> Average Cal sensor Temp.: 98.29 Ambient Temp.: 86.34 F
01:32:14.108 -> Average Cal sensor Temp.: 98.54 Ambient Temp.: 86.45 F
01:32:15.663 -> Average Cal sensor Temp.: 98.31 Ambient Temp.: 86.45 F
01:32:17.217 -> Average Cal sensor Temp.: 98.41 Ambient Temp.: 86.38 F
01:32:18.779 -> Average Cal sensor Temp.: 98.43 Ambient Temp.: 86.45 F
01:32:20.338 -> Average Cal sensor Temp.: 98.30 Ambient Temp.: 86.38 F
01:32:21.893 -> Average Cal sensor Temp.: 98.39 Ambient Temp.: 86.34 F
01:32:23.450 -> Average Cal sensor Temp.: 98.68 Ambient Temp.: 86.38 F
01:32:25.006 -> Average Cal sensor Temp.: 99.32 Ambient Temp.: 86.38 F
01:32:26.561 -> Average Cal sensor Temp.: 98.69 Ambient Temp.: 86.45 F
01:32:28.119 -> Average Cal sensor Temp.: 98.34 Ambient Temp.: 86.45 F
01:32:29.676 -> Average Cal sensor Temp.: 98.29 Ambient Temp.: 86.45 F
01:32:31.237 -> Average Cal sensor Temp.: 98.55 Ambient Temp.: 86.34 F
01:32:32.791 -> Average Cal sensor Temp.: 98.83 Ambient Temp.: 86.52 F
01:32:34.352 -> Average Cal sensor Temp.: 98.65 Ambient Temp.: 86.38 F
01:32:35.908 -> Average Cal sensor Temp.: 98.64 Ambient Temp.: 86.49 F
01:32:37.463 -> Average Cal sensor Temp.: 98.30 Ambient Temp.: 86.45 F
01:32:39.020 -> Average Cal sensor Temp.: 98.49 Ambient Temp.: 86.38 F
01:32:40.608 -> Average Cal sensor Temp.: 98.67 Ambient Temp.: 86.45 F
01:32:42.166 -> Average Cal sensor Temp.: 98.95 Ambient Temp.: 86.38 F
01:32:43.724 -> Average Cal sensor Temp.: 98.88 Ambient Temp.: 86.52 F
01:32:45.278 -> Average Cal sensor Temp.: 98.70 Ambient Temp.: 86.49 F
01:32:46.828 -> Average Cal sensor Temp.: 98.87 Ambient Temp.: 86.52 F
01:32:48.384 -> Average Cal sensor Temp.: 98.96 Ambient Temp.: 86.52 F
01:32:49.946 -> Average Cal sensor Temp.: 99.08 Ambient Temp.: 86.52 F
01:32:51.500 -> Average Cal sensor Temp.: 98.83 Ambient Temp.: 86.52 F
01:32:53.054 -> Average Cal sensor Temp.: 98.74 Ambient Temp.: 86.49 F
01:32:54.612 -> Average Cal sensor Temp.: 99.14 Ambient Temp.: 86.56 F
01:32:56.166 -> Average Cal sensor Temp.: 98.98 Ambient Temp.: 86.63 F
01:32:57.719 -> Average Cal sensor Temp.: 99.02 Ambient Temp.: 86.59 F
01:32:59.279 -> Average Cal sensor Temp.: 98.95 Ambient Temp.: 86.63 F
01:33:00.837 -> Average Cal sensor Temp.: 99.13 Ambient Temp.: 86.67 F
01:33:02.394 -> Average Cal sensor Temp.: 98.54 Ambient Temp.: 86.67 F
01:33:03.945 -> Average Cal sensor Temp.: 98.94 Ambient Temp.: 86.70 F
01:33:05.500 -> Average Cal sensor Temp.: 98.48 Ambient Temp.: 86.70 F
01:33:07.089 -> Average Cal sensor Temp.: 98.41 Ambient Temp.: 86.67 F
01:33:08.643 -> Average Cal sensor Temp.: 98.58 Ambient Temp.: 86.70 F
01:33:10.200 -> Average Cal sensor Temp.: 98.51 Ambient Temp.: 86.67 F
01:33:11.751 -> Average Cal sensor Temp.: 98.38 Ambient Temp.: 86.63 F
01:33:13.312 -> Average Cal sensor Temp.: 98.54 Ambient Temp.: 86.70 F
01:33:14.870 -> Average Cal sensor Temp.: 98.78 Ambient Temp.: 86.67 F
01:33:16.429 -> Average Cal sensor Temp.: 99.25 Ambient Temp.: 86.67 F
01:33:17.971 -> Average Cal sensor Temp.: 99.27 Ambient Temp.: 86.74 F
01:33:19.523 -> Average Cal sensor Temp.: 99.11 Ambient Temp.: 86.70 F
01:33:21.082 -> Average Cal sensor Temp.: 98.76 Ambient Temp.: 86.74 F
01:33:22.630 -> Average Cal sensor Temp.: 98.28 Ambient Temp.: 86.77 F
01:33:24.189 -> Average Cal sensor Temp.: 98.66 Ambient Temp.: 86.70 F
01:33:25.745 -> Average Cal sensor Temp.: 98.87 Ambient Temp.: 86.70 F
01:33:27.326 -> Average Cal sensor Temp.: 98.95 Ambient Temp.: 86.70 F
01:33:28.877 -> Average Cal sensor Temp.: 98.97 Ambient Temp.: 86.74 F
01:33:30.426 -> Average Cal sensor Temp.: 98.73 Ambient Temp.: 86.74 F
01:33:31.984 -> Average Cal sensor Temp.: 98.83 Ambient Temp.: 86.74 F
01:33:33.544 -> Average Cal sensor Temp.: 99.11 Ambient Temp.: 86.81 F
01:33:35.098 -> Average Cal sensor Temp.: 98.60 Ambient Temp.: 86.77 F
01:33:36.646 -> Average Cal sensor Temp.: 98.31 Ambient Temp.: 86.74 F
01:33:38.200 -> Average Cal sensor Temp.: 98.51 Ambient Temp.: 86.77 F
01:33:39.750 -> Average Cal sensor Temp.: 98.97 Ambient Temp.: 86.81 F
01:33:41.313 -> Average Cal sensor Temp.: 98.65 Ambient Temp.: 86.77 F
01:33:42.867 -> Average Cal sensor Temp.: 98.99 Ambient Temp.: 86.81 F
01:33:44.424 -> Average Cal sensor Temp.: 98.62 Ambient Temp.: 86.77 F
01:33:46.013 -> Average Cal sensor Temp.: 98.89 Ambient Temp.: 86.77 F
01:33:47.569 -> Average Cal sensor Temp.: 99.10 Ambient Temp.: 86.77 F
01:33:49.123 -> Average Cal sensor Temp.: 99.10 Ambient Temp.: 86.81 F
01:33:50.682 -> Average Cal sensor Temp.: 98.81 Ambient Temp.: 86.77 F
01:33:52.238 -> Average Cal sensor Temp.: 99.34 Ambient Temp.: 86.81 F
01:33:53.791 -> Average Cal sensor Temp.: 98.95 Ambient Temp.: 86.74 F
01:33:55.348 -> Average Cal sensor Temp.: 99.23 Ambient Temp.: 86.81 F
01:33:56.896 -> Average Cal sensor Temp.: 99.15 Ambient Temp.: 86.81 F
01:33:58.454 -> Average Cal sensor Temp.: 99.16 Ambient Temp.: 86.88 F
01:34:00.006 -> Average Cal sensor Temp.: 98.90 Ambient Temp.: 86.88 F
01:34:01.556 -> Average Cal sensor Temp.: 99.24 Ambient Temp.: 86.92 F
01:34:03.114 -> Average Cal sensor Temp.: 98.97 Ambient Temp.: 86.88 F
01:34:04.672 -> Average Cal sensor Temp.: 99.28 Ambient Temp.: 86.88 F
01:34:06.229 -> Average Cal sensor Temp.: 99.35 Ambient Temp.: 86.95 F
01:34:07.785 -> Average Cal sensor Temp.: 99.35 Ambient Temp.: 86.92 F
01:34:09.343 -> Average Cal sensor Temp.: 99.35 Ambient Temp.: 86.95 F
01:34:10.899 -> Average Cal sensor Temp.: 98.98 Ambient Temp.: 86.99 F
01:34:12.457 -> Average Cal sensor Temp.: 99.13 Ambient Temp.: 86.99 F
01:34:14.014 -> Average Cal sensor Temp.: 98.82 Ambient Temp.: 86.95 F
01:34:15.576 -> Average Cal sensor Temp.: 98.86 Ambient Temp.: 87.03 F
01:34:17.134 -> Average Cal sensor Temp.: 98.72 Ambient Temp.: 86.92 F
01:34:18.686 -> Average Cal sensor Temp.: 99.41 Ambient Temp.: 86.95 F
01:34:20.245 -> Average Cal sensor Temp.: 99.19 Ambient Temp.: 86.95 F
01:34:21.798 -> Average Cal sensor Temp.: 99.36 Ambient Temp.: 86.95 F
01:34:23.355 -> Average Cal sensor Temp.: 99.41 Ambient Temp.: 87.06 F
01:34:24.912 -> Average Cal sensor Temp.: 98.97 Ambient Temp.: 87.06 F
01:34:26.472 -> Average Cal sensor Temp.: 98.97 Ambient Temp.: 86.99 F
01:34:28.031 -> Average Cal sensor Temp.: 99.20 Ambient Temp.: 87.03 F
01:34:29.588 -> Average Cal sensor Temp.: 99.82 Ambient Temp.: 87.10 F
01:34:31.146 -> Average Cal sensor Temp.: 99.07 Ambient Temp.: 87.03 F
01:34:32.733 -> Average Cal sensor Temp.: 98.92 Ambient Temp.: 87.06 F
01:34:34.291 -> Average Cal sensor Temp.: 99.22 Ambient Temp.: 87.03 F
01:34:35.842 -> Average Cal sensor Temp.: 99.23 Ambient Temp.: 87.03 F
01:34:37.392 -> Average Cal sensor Temp.: 99.46 Ambient Temp.: 87.10 F
01:34:38.942 -> Average Cal sensor Temp.: 99.75 Ambient Temp.: 87.10 F
01:34:40.498 -> Average Cal sensor Temp.: 99.49 Ambient Temp.: 87.10 F
01:34:42.053 -> Average Cal sensor Temp.: 99.11 Ambient Temp.: 87.10 F
01:34:43.609 -> Average Cal sensor Temp.: 99.24 Ambient Temp.: 87.13 F
01:34:45.168 -> Average Cal sensor Temp.: 99.43 Ambient Temp.: 87.21 F
01:34:46.718 -> Average Cal sensor Temp.: 99.36 Ambient Temp.: 87.17 F
01:34:48.275 -> Average Cal sensor Temp.: 98.87 Ambient Temp.: 87.21 F
01:34:49.828 -> Average Cal sensor Temp.: 98.88 Ambient Temp.: 87.17 F
01:34:51.387 -> Average Cal sensor Temp.: 99.37 Ambient Temp.: 87.21 F
01:34:52.944 -> Average Cal sensor Temp.: 99.62 Ambient Temp.: 87.21 F
01:34:54.501 -> Average Cal sensor Temp.: 99.39 Ambient Temp.: 87.21 F
01:34:56.057 -> Average Cal sensor Temp.: 99.50 Ambient Temp.: 87.31 F
01:34:57.617 -> Average Cal sensor Temp.: 99.56 Ambient Temp.: 87.31 F
01:34:59.180 -> Average Cal sensor Temp.: 99.54 Ambient Temp.: 87.35 F
01:35:00.741 -> Average Cal sensor Temp.: 99.09 Ambient Temp.: 87.31 F
01:35:02.297 -> Average Cal sensor Temp.: 98.90 Ambient Temp.: 87.28 F
01:35:03.854 -> Average Cal sensor Temp.: 99.31 Ambient Temp.: 87.39 F
01:35:05.411 -> Average Cal sensor Temp.: 99.05 Ambient Temp.: 87.35 F
01:35:06.965 -> Average Cal sensor Temp.: 98.79 Ambient Temp.: 87.39 F
01:35:08.546 -> Average Cal sensor Temp.: 99.10 Ambient Temp.: 87.42 F
01:35:10.102 -> Average Cal sensor Temp.: 99.22 Ambient Temp.: 87.42 F
01:35:11.626 -> Average Cal sensor Temp.: 99.26 Ambient Temp.: 87.35 F
01:35:13.214 -> Average Cal sensor Temp.: 99.19 Ambient Temp.: 87.46 F
01:35:14.767 -> Average Cal sensor Temp.: 99.17 Ambient Temp.: 87.46 F
01:35:16.316 -> Average Cal sensor Temp.: 99.17 Ambient Temp.: 87.46 F
01:35:17.875 -> Average Cal sensor Temp.: 99.10 Ambient Temp.: 87.49 F
01:35:19.431 -> Average Cal sensor Temp.: 99.39 Ambient Temp.: 87.46 F
01:35:20.986 -> Average Cal sensor Temp.: 99.28 Ambient Temp.: 87.46 F
01:35:22.537 -> Average Cal sensor Temp.: 99.12 Ambient Temp.: 87.46 F
01:35:24.092 -> Average Cal sensor Temp.: 99.20 Ambient Temp.: 87.49 F
01:35:25.648 -> Average Cal sensor Temp.: 99.36 Ambient Temp.: 87.49 F
01:35:27.202 -> Average Cal sensor Temp.: 99.24 Ambient Temp.: 87.53 F
01:35:28.757 -> Average Cal sensor Temp.: 98.52 Ambient Temp.: 87.46 F
01:35:30.314 -> Average Cal sensor Temp.: 98.17 Ambient Temp.: 87.46 F
01:35:31.869 -> Average Cal sensor Temp.: 98.40 Ambient Temp.: 87.46 F
01:35:33.427 -> Average Cal sensor Temp.: 98.80 Ambient Temp.: 87.46 F
01:35:34.983 -> Average Cal sensor Temp.: 98.75 Ambient Temp.: 87.39 F
01:35:36.537 -> Average Cal sensor Temp.: 98.52 Ambient Temp.: 87.39 F
01:35:38.092 -> Average Cal sensor Temp.: 99.09 Ambient Temp.: 87.42 F
01:35:39.650 -> Average Cal sensor Temp.: 99.14 Ambient Temp.: 87.46 F
01:35:41.211 -> Average Cal sensor Temp.: 99.12 Ambient Temp.: 87.42 F
01:35:42.761 -> Average Cal sensor Temp.: 98.92 Ambient Temp.: 87.46 F
01:35:44.316 -> Average Cal sensor Temp.: 99.00 Ambient Temp.: 87.49 F
01:35:45.907 -> Average Cal sensor Temp.: 99.18 Ambient Temp.: 87.46 F
01:35:47.460 -> Average Cal sensor Temp.: 98.93 Ambient Temp.: 87.49 F
01:35:49.018 -> Average Cal sensor Temp.: 98.89 Ambient Temp.: 87.46 F
01:35:50.568 -> Average Cal sensor Temp.: 99.15 Ambient Temp.: 87.49 F
01:35:52.124 -> Average Cal sensor Temp.: 98.99 Ambient Temp.: 87.46 F
01:35:53.682 -> Average Cal sensor Temp.: 99.34 Ambient Temp.: 87.49 F
01:35:55.241 -> Average Cal sensor Temp.: 99.27 Ambient Temp.: 87.46 F
01:35:56.794 -> Average Cal sensor Temp.: 98.54 Ambient Temp.: 87.42 F
01:35:58.346 -> Average Cal sensor Temp.: 99.19 Ambient Temp.: 87.42 F
01:35:59.902 -> Average Cal sensor Temp.: 99.47 Ambient Temp.: 87.46 F
01:36:01.460 -> Average Cal sensor Temp.: 99.27 Ambient Temp.: 87.46 F
01:36:03.011 -> Average Cal sensor Temp.: 99.13 Ambient Temp.: 87.49 F
01:36:04.565 -> Average Cal sensor Temp.: 100.00 Ambient Temp.: 87.53 F
01:36:06.154 -> Average Cal sensor Temp.: 98.89 Ambient Temp.: 87.53 F
01:36:07.710 -> Average Cal sensor Temp.: 99.40 Ambient Temp.: 87.49 F
01:36:09.267 -> Average Cal sensor Temp.: 99.14 Ambient Temp.: 87.46 F
01:36:10.827 -> Average Cal sensor Temp.: 98.89 Ambient Temp.: 87.53 F
01:36:12.384 -> Average Cal sensor Temp.: 98.68 Ambient Temp.: 87.46 F
01:36:13.939 -> Average Cal sensor Temp.: 98.35 Ambient Temp.: 87.46 F
01:36:15.495 -> Average Cal sensor Temp.: 98.65 Ambient Temp.: 87.46 F
01:36:17.054 -> Average Cal sensor Temp.: 98.83 Ambient Temp.: 87.49 F
01:36:18.603 -> Average Cal sensor Temp.: 99.11 Ambient Temp.: 87.46 F
01:36:20.159 -> Average Cal sensor Temp.: 98.75 Ambient Temp.: 87.42 F
01:36:21.710 -> Average Cal sensor Temp.: 98.90 Ambient Temp.: 87.46 F
01:36:23.263 -> Average Cal sensor Temp.: 99.14 Ambient Temp.: 87.39 F
01:36:24.821 -> Average Cal sensor Temp.: 99.27 Ambient Temp.: 87.39 F
01:36:26.375 -> Average Cal sensor Temp.: 99.32 Ambient Temp.: 87.46 F
01:36:27.932 -> Average Cal sensor Temp.: 98.57 Ambient Temp.: 87.42 F
01:36:29.483 -> Average Cal sensor Temp.: 98.78 Ambient Temp.: 87.39 F
01:36:31.039 -> Average Cal sensor Temp.: 99.12 Ambient Temp.: 87.39 F
01:36:32.593 -> Average Cal sensor Temp.: 99.20 Ambient Temp.: 87.42 F
01:36:34.147 -> Average Cal sensor Temp.: 98.66 Ambient Temp.: 87.39 F
01:36:35.704 -> Average Cal sensor Temp.: 98.68 Ambient Temp.: 87.39 F
01:36:37.287 -> Average Cal sensor Temp.: 98.38 Ambient Temp.: 87.31 F
01:36:38.841 -> Average Cal sensor Temp.: 98.57 Ambient Temp.: 87.31 F
01:36:40.395 -> Average Cal sensor Temp.: 98.89 Ambient Temp.: 87.21 F
01:36:41.954 -> Average Cal sensor Temp.: 98.71 Ambient Temp.: 87.28 F
01:36:43.512 -> Average Cal sensor Temp.: 98.78 Ambient Temp.: 87.21 F
01:36:45.063 -> Average Cal sensor Temp.: 99.05 Ambient Temp.: 87.28 F
01:36:46.616 -> Average Cal sensor Temp.: 99.03 Ambient Temp.: 87.17 F
01:36:48.174 -> Average Cal sensor Temp.: 99.19 Ambient Temp.: 87.28 F
01:36:49.729 -> Average Cal sensor Temp.: 99.32 Ambient Temp.: 87.17 F
01:36:51.286 -> Average Cal sensor Temp.: 99.30 Ambient Temp.: 87.17 F
01:36:52.839 -> Average Cal sensor Temp.: 98.89 Ambient Temp.: 87.21 F
01:36:54.400 -> Average Cal sensor Temp.: 98.87 Ambient Temp.: 87.21 F
01:36:55.947 -> Average Cal sensor Temp.: 98.90 Ambient Temp.: 87.13 F
01:36:57.517 -> Average Cal sensor Temp.: 99.01 Ambient Temp.: 87.17 F
01:36:59.072 -> Average Cal sensor Temp.: 99.04 Ambient Temp.: 87.21 F
01:37:00.626 -> Average Cal sensor Temp.: 98.83 Ambient Temp.: 87.21 F
01:37:02.181 -> Average Cal sensor Temp.: 98.27 Ambient Temp.: 87.13 F
01:37:03.739 -> Average Cal sensor Temp.: 98.63 Ambient Temp.: 87.10 F
01:37:05.293 -> Average Cal sensor Temp.: 98.89 Ambient Temp.: 87.13 F
01:37:06.845 -> Average Cal sensor Temp.: 98.62 Ambient Temp.: 87.10 F
01:37:08.403 -> Average Cal sensor Temp.: 98.90 Ambient Temp.: 87.10 F
01:37:09.964 -> Average Cal sensor Temp.: 98.96 Ambient Temp.: 87.10 F
01:37:11.514 -> Average Cal sensor Temp.: 98.82 Ambient Temp.: 87.10 F
01:37:13.066 -> Average Cal sensor Temp.: 98.73 Ambient Temp.: 87.06 F
01:37:14.625 -> Average Cal sensor Temp.: 98.62 Ambient Temp.: 87.06 F
01:37:16.182 -> Average Cal sensor Temp.: 98.79 Ambient Temp.: 87.10 F
01:37:17.739 -> Average Cal sensor Temp.: 98.99 Ambient Temp.: 86.99 F
01:37:19.301 -> Average Cal sensor Temp.: 99.06 Ambient Temp.: 87.06 F
01:37:20.855 -> Average Cal sensor Temp.: 98.45 Ambient Temp.: 87.06 F
01:37:22.413 -> Average Cal sensor Temp.: 98.90 Ambient Temp.: 87.03 F
01:37:23.971 -> Average Cal sensor Temp.: 98.83 Ambient Temp.: 87.03 F
01:37:25.531 -> Average Cal sensor Temp.: 98.51 Ambient Temp.: 86.99 F
01:37:27.089 -> Average Cal sensor Temp.: 98.65 Ambient Temp.: 86.92 F
01:37:28.641 -> Average Cal sensor Temp.: 98.61 Ambient Temp.: 86.95 F
01:37:30.196 -> Average Cal sensor Temp.: 98.78 Ambient Temp.: 86.99 F
01:37:31.745 -> Average Cal sensor Temp.: 98.67 Ambient Temp.: 86.95 F
01:37:33.305 -> Average Cal sensor Temp.: 98.73 Ambient Temp.: 86.92 F
01:37:34.863 -> Average Cal sensor Temp.: 99.13 Ambient Temp.: 86.92 F
01:37:36.421 -> Average Cal sensor Temp.: 98.69 Ambient Temp.: 86.88 F
01:37:37.978 -> Average Cal sensor Temp.: 99.05 Ambient Temp.: 86.81 F
01:37:39.527 -> Average Cal sensor Temp.: 99.14 Ambient Temp.: 86.88 F
01:37:41.085 -> Average Cal sensor Temp.: 99.06 Ambient Temp.: 86.81 F
01:37:42.676 -> Average Cal sensor Temp.: 99.33 Ambient Temp.: 86.92 F
01:37:44.226 -> Average Cal sensor Temp.: 99.07 Ambient Temp.: 86.92 F
01:37:45.782 -> Average Cal sensor Temp.: 99.03 Ambient Temp.: 86.88 F
01:37:47.341 -> Average Cal sensor Temp.: 99.47 Ambient Temp.: 86.88 F
01:37:48.899 -> Average Cal sensor Temp.: 99.60 Ambient Temp.: 86.88 F
01:37:50.455 -> Average Cal sensor Temp.: 99.14 Ambient Temp.: 86.88 F
01:37:52.009 -> Average Cal sensor Temp.: 99.04 Ambient Temp.: 86.92 F
01:37:53.564 -> Average Cal sensor Temp.: 98.91 Ambient Temp.: 86.88 F
01:37:55.125 -> Average Cal sensor Temp.: 98.96 Ambient Temp.: 86.88 F
01:37:56.682 -> Average Cal sensor Temp.: 98.98 Ambient Temp.: 86.88 F
01:37:58.239 -> Average Cal sensor Temp.: 98.80 Ambient Temp.: 86.95 F
01:37:59.793 -> Average Cal sensor Temp.: 99.09 Ambient Temp.: 86.92 F
01:38:01.347 -> Average Cal sensor Temp.: 98.97 Ambient Temp.: 86.92 F
01:38:02.934 -> Average Cal sensor Temp.: 98.92 Ambient Temp.: 86.88 F
01:38:04.491 -> Average Cal sensor Temp.: 98.73 Ambient Temp.: 86.88 F
01:38:06.047 -> Average Cal sensor Temp.: 98.64 Ambient Temp.: 86.92 F
01:38:07.598 -> Average Cal sensor Temp.: 98.34 Ambient Temp.: 86.81 F
01:38:09.151 -> Average Cal sensor Temp.: 98.08 Ambient Temp.: 86.81 F
01:38:10.706 -> Average Cal sensor Temp.: 98.67 Ambient Temp.: 86.81 F
01:38:12.265 -> Average Cal sensor Temp.: 99.34 Ambient Temp.: 86.81 F
01:38:13.824 -> Average Cal sensor Temp.: 98.63 Ambient Temp.: 86.77 F
01:38:15.384 -> Average Cal sensor Temp.: 98.70 Ambient Temp.: 86.77 F
01:38:16.934 -> Average Cal sensor Temp.: 99.16 Ambient Temp.: 86.74 F
01:38:18.496 -> Average Cal sensor Temp.: 99.26 Ambient Temp.: 86.77 F
01:38:20.056 -> Average Cal sensor Temp.: 99.43 Ambient Temp.: 86.77 F
01:38:21.608 -> Average Cal sensor Temp.: 98.92 Ambient Temp.: 86.74 F
01:38:23.165 -> Average Cal sensor Temp.: 99.15 Ambient Temp.: 86.74 F
01:38:24.721 -> Average Cal sensor Temp.: 99.24 Ambient Temp.: 86.77 F
01:38:26.280 -> Average Cal sensor Temp.: 99.46 Ambient Temp.: 86.77 F
01:38:27.835 -> Average Cal sensor Temp.: 99.47 Ambient Temp.: 86.74 F
01:38:29.391 -> Average Cal sensor Temp.: 99.46 Ambient Temp.: 86.74 F
01:38:30.946 -> Average Cal sensor Temp.: 98.79 Ambient Temp.: 86.74 F
01:38:32.504 -> Average Cal sensor Temp.: 99.03 Ambient Temp.: 86.67 F
01:38:34.058 -> Average Cal sensor Temp.: 99.60 Ambient Temp.: 86.70 F
01:38:35.617 -> Average Cal sensor Temp.: 99.59 Ambient Temp.: 86.70 F
01:38:37.177 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 86.74 F
01:38:38.769 -> Average Cal sensor Temp.: 99.56 Ambient Temp.: 86.74 F
01:38:40.332 -> Average Cal sensor Temp.: 99.14 Ambient Temp.: 86.70 F
01:38:41.856 -> Average Cal sensor Temp.: 98.88 Ambient Temp.: 86.70 F
01:38:43.446 -> Average Cal sensor Temp.: 99.25 Ambient Temp.: 86.67 F
01:38:45.004 -> Average Cal sensor Temp.: 98.96 Ambient Temp.: 86.67 F
01:38:46.560 -> Average Cal sensor Temp.: 99.27 Ambient Temp.: 86.70 F
01:38:48.118 -> Average Cal sensor Temp.: 99.23 Ambient Temp.: 86.74 F
01:38:49.671 -> Average Cal sensor Temp.: 99.26 Ambient Temp.: 86.77 F
01:38:51.228 -> Average Cal sensor Temp.: 99.34 Ambient Temp.: 86.74 F
01:38:52.786 -> Average Cal sensor Temp.: 99.26 Ambient Temp.: 86.70 F
01:38:54.346 -> Average Cal sensor Temp.: 99.37 Ambient Temp.: 86.70 F
01:38:55.900 -> Average Cal sensor Temp.: 99.33 Ambient Temp.: 86.74 F
01:38:57.427 -> Average Cal sensor Temp.: 99.78 Ambient Temp.: 86.74 F
01:38:59.013 -> Average Cal sensor Temp.: 99.14 Ambient Temp.: 86.70 F
01:39:00.569 -> Average Cal sensor Temp.: 99.12 Ambient Temp.: 86.70 F
01:39:02.123 -> Average Cal sensor Temp.: 99.28 Ambient Temp.: 86.70 F
01:39:03.685 -> Average Cal sensor Temp.: 99.74 Ambient Temp.: 86.74 F
01:39:05.245 -> Average Cal sensor Temp.: 99.17 Ambient Temp.: 86.77 F
01:39:06.798 -> Average Cal sensor Temp.: 99.04 Ambient Temp.: 86.77 F
01:39:08.356 -> Average Cal sensor Temp.: 98.60 Ambient Temp.: 86.70 F
01:39:09.910 -> Average Cal sensor Temp.: 99.15 Ambient Temp.: 86.70 F
01:39:11.459 -> Average Cal sensor Temp.: 98.97 Ambient Temp.: 86.70 F
01:39:13.013 -> Average Cal sensor Temp.: 99.33 Ambient Temp.: 86.70 F
01:39:14.572 -> Average Cal sensor Temp.: 99.28 Ambient Temp.: 86.70 F
01:39:16.119 -> Average Cal sensor Temp.: 99.55 Ambient Temp.: 86.77 F
01:39:17.707 -> Average Cal sensor Temp.: 99.44 Ambient Temp.: 86.77 F
01:39:19.270 -> Average Cal sensor Temp.: 99.03 Ambient Temp.: 86.77 F
01:39:20.797 -> Average Cal sensor Temp.: 99.33 Ambient Temp.: 86.81 F
01:39:22.352 -> Average Cal sensor Temp.: 99.47 Ambient Temp.: 86.77 F
01:39:23.914 -> Average Cal sensor Temp.: 99.43 Ambient Temp.: 86.88 F
01:39:25.470 -> Average Cal sensor Temp.: 99.43 Ambient Temp.: 86.88 F
01:39:27.025 -> Average Cal sensor Temp.: 99.24 Ambient Temp.: 86.92 F
01:39:28.582 -> Average Cal sensor Temp.: 99.53 Ambient Temp.: 86.92 F
01:39:30.172 -> Average Cal sensor Temp.: 99.38 Ambient Temp.: 86.92 F
01:39:31.734 -> Average Cal sensor Temp.: 99.27 Ambient Temp.: 86.95 F
01:39:33.283 -> Average Cal sensor Temp.: 99.46 Ambient Temp.: 86.99 F
01:39:34.839 -> Average Cal sensor Temp.: 99.32 Ambient Temp.: 86.99 F
01:39:36.400 -> Average Cal sensor Temp.: 99.52 Ambient Temp.: 87.03 F
01:39:37.956 -> Average Cal sensor Temp.: 99.53 Ambient Temp.: 87.03 F
01:39:39.511 -> Average Cal sensor Temp.: 98.89 Ambient Temp.: 86.99 F
01:39:41.075 -> Average Cal sensor Temp.: 99.16 Ambient Temp.: 87.03 F
01:39:42.633 -> Average Cal sensor Temp.: 99.70 Ambient Temp.: 87.03 F
01:39:44.194 -> Average Cal sensor Temp.: 99.16 Ambient Temp.: 87.06 F
01:39:45.752 -> Average Cal sensor Temp.: 99.29 Ambient Temp.: 87.06 F
01:39:47.305 -> Average Cal sensor Temp.: 99.50 Ambient Temp.: 87.06 F
01:39:48.865 -> Average Cal sensor Temp.: 99.46 Ambient Temp.: 87.10 F
01:39:50.421 -> Average Cal sensor Temp.: 100.37 Ambient Temp.: 87.13 F
01:39:51.980 -> Average Cal sensor Temp.: 99.38 Ambient Temp.: 87.13 F
01:39:53.530 -> Average Cal sensor Temp.: 99.55 Ambient Temp.: 87.13 F
01:39:55.085 -> Average Cal sensor Temp.: 99.28 Ambient Temp.: 87.13 F
01:39:56.639 -> Average Cal sensor Temp.: 99.76 Ambient Temp.: 87.17 F
01:39:58.231 -> Average Cal sensor Temp.: 99.38 Ambient Temp.: 87.13 F
01:39:59.755 -> Average Cal sensor Temp.: 99.51 Ambient Temp.: 87.21 F
01:40:01.315 -> Average Cal sensor Temp.: 99.22 Ambient Temp.: 87.17 F
01:40:02.868 -> Average Cal sensor Temp.: 99.70 Ambient Temp.: 87.28 F
01:40:04.452 -> Average Cal sensor Temp.: 100.00 Ambient Temp.: 87.21 F
01:40:06.009 -> Average Cal sensor Temp.: 99.69 Ambient Temp.: 87.28 F
01:40:07.564 -> Average Cal sensor Temp.: 99.81 Ambient Temp.: 87.31 F
01:40:09.123 -> Average Cal sensor Temp.: 99.76 Ambient Temp.: 87.35 F
01:40:10.684 -> Average Cal sensor Temp.: 99.65 Ambient Temp.: 87.35 F
01:40:12.275 -> Average Cal sensor Temp.: 99.72 Ambient Temp.: 87.42 F
01:40:13.832 -> Average Cal sensor Temp.: 99.80 Ambient Temp.: 87.42 F
01:40:15.392 -> Average Cal sensor Temp.: 99.83 Ambient Temp.: 87.42 F
01:40:16.939 -> Average Cal sensor Temp.: 99.77 Ambient Temp.: 87.49 F
01:40:18.490 -> Average Cal sensor Temp.: 99.80 Ambient Temp.: 87.53 F
01:40:20.079 -> Average Cal sensor Temp.: 99.18 Ambient Temp.: 87.53 F
01:40:21.637 -> Average Cal sensor Temp.: 99.29 Ambient Temp.: 87.49 F
01:40:23.192 -> Average Cal sensor Temp.: 99.57 Ambient Temp.: 87.49 F
01:40:24.744 -> Average Cal sensor Temp.: 99.77 Ambient Temp.: 87.57 F
01:40:26.299 -> Average Cal sensor Temp.: 99.41 Ambient Temp.: 87.49 F
01:40:27.849 -> Average Cal sensor Temp.: 99.65 Ambient Temp.: 87.57 F
01:40:29.434 -> Average Cal sensor Temp.: 99.86 Ambient Temp.: 87.60 F
01:40:30.996 -> Average Cal sensor Temp.: 99.46 Ambient Temp.: 87.57 F
01:40:32.551 -> Average Cal sensor Temp.: 99.62 Ambient Temp.: 87.57 F
01:40:34.106 -> Average Cal sensor Temp.: 99.60 Ambient Temp.: 87.57 F
01:40:35.653 -> Average Cal sensor Temp.: 99.69 Ambient Temp.: 87.67 F
01:40:37.245 -> Average Cal sensor Temp.: 99.79 Ambient Temp.: 87.60 F
01:40:38.799 -> Average Cal sensor Temp.: 99.61 Ambient Temp.: 87.67 F
01:40:40.359 -> Average Cal sensor Temp.: 99.86 Ambient Temp.: 87.60 F
01:40:41.915 -> Average Cal sensor Temp.: 100.10 Ambient Temp.: 87.57 F
01:40:43.475 -> Average Cal sensor Temp.: 100.02 Ambient Temp.: 87.67 F
01:40:45.034 -> Average Cal sensor Temp.: 99.91 Ambient Temp.: 87.60 F
01:40:46.595 -> Average Cal sensor Temp.: 99.94 Ambient Temp.: 87.60 F
01:40:48.181 -> Average Cal sensor Temp.: 99.95 Ambient Temp.: 87.60 F
01:40:49.732 -> Average Cal sensor Temp.: 99.91 Ambient Temp.: 87.67 F
01:40:51.295 -> Average Cal sensor Temp.: 100.04 Ambient Temp.: 87.60 F
01:40:52.853 -> Average Cal sensor Temp.: 99.66 Ambient Temp.: 87.60 F
01:40:54.414 -> Average Cal sensor Temp.: 99.65 Ambient Temp.: 87.67 F
01:40:56.005 -> Average Cal sensor Temp.: 99.84 Ambient Temp.: 87.71 F
01:40:57.557 -> Average Cal sensor Temp.: 99.94 Ambient Temp.: 87.75 F
01:40:59.117 -> Average Cal sensor Temp.: 99.50 Ambient Temp.: 87.71 F
01:41:00.672 -> Average Cal sensor Temp.: 99.64 Ambient Temp.: 87.75 F
01:41:02.229 -> Average Cal sensor Temp.: 99.39 Ambient Temp.: 87.71 F
01:41:03.777 -> Average Cal sensor Temp.: 99.23 Ambient Temp.: 87.60 F
01:41:05.332 -> Average Cal sensor Temp.: 98.81 Ambient Temp.: 87.67 F
01:41:06.912 -> Average Cal sensor Temp.: 99.11 Ambient Temp.: 87.60 F
01:41:08.471 -> Average Cal sensor Temp.: 99.30 Ambient Temp.: 87.60 F
01:41:10.026 -> Average Cal sensor Temp.: 99.44 Ambient Temp.: 87.67 F
01:41:11.578 -> Average Cal sensor Temp.: 99.06 Ambient Temp.: 87.57 F
01:41:13.124 -> Average Cal sensor Temp.: 98.67 Ambient Temp.: 87.57 F
01:41:14.674 -> Average Cal sensor Temp.: 98.58 Ambient Temp.: 87.57 F
01:41:16.255 -> Average Cal sensor Temp.: 98.65 Ambient Temp.: 87.49 F
01:41:17.808 -> Average Cal sensor Temp.: 99.15 Ambient Temp.: 87.49 F
01:41:19.363 -> Average Cal sensor Temp.: 99.62 Ambient Temp.: 87.49 F
01:41:20.916 -> Average Cal sensor Temp.: 100.03 Ambient Temp.: 87.53 F
01:41:22.474 -> Average Cal sensor Temp.: 99.58 Ambient Temp.: 87.49 F
01:41:24.026 -> Average Cal sensor Temp.: 98.54 Ambient Temp.: 87.46 F
01:41:25.585 -> Average Cal sensor Temp.: 98.90 Ambient Temp.: 87.42 F
01:41:27.138 -> Average Cal sensor Temp.: 98.93 Ambient Temp.: 87.39 F
01:41:28.728 -> Average Cal sensor Temp.: 99.00 Ambient Temp.: 87.42 F
01:41:30.280 -> Average Cal sensor Temp.: 99.42 Ambient Temp.: 87.42 F
01:41:31.835 -> Average Cal sensor Temp.: 99.37 Ambient Temp.: 87.42 F
01:41:33.391 -> Average Cal sensor Temp.: 99.51 Ambient Temp.: 87.42 F
01:41:34.941 -> Average Cal sensor Temp.: 99.61 Ambient Temp.: 87.39 F
01:41:36.499 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 87.39 F
01:41:38.082 -> Average Cal sensor Temp.: 99.90 Ambient Temp.: 87.42 F
01:41:39.634 -> Average Cal sensor Temp.: 99.41 Ambient Temp.: 87.46 F
01:41:41.193 -> Average Cal sensor Temp.: 98.95 Ambient Temp.: 87.39 F
01:41:42.752 -> Average Cal sensor Temp.: 99.14 Ambient Temp.: 87.42 F
01:41:44.310 -> Average Cal sensor Temp.: 99.46 Ambient Temp.: 87.42 F
01:41:45.865 -> Average Cal sensor Temp.: 99.92 Ambient Temp.: 87.35 F
01:41:46.449 -> Ambient Temp.: 87.35 F
01:41:48.005 -> Average Cal sensor Temp.: 94.77 Ambient Temp.: 87.42 F
01:41:49.561 -> Average Cal sensor Temp.: 94.62 Ambient Temp.: 87.42 F
01:41:51.118 -> Average Cal sensor Temp.: 100.25 Ambient Temp.: 87.42 F
01:41:52.672 -> Average Cal sensor Temp.: 99.91 Ambient Temp.: 87.39 F
01:41:54.233 -> Average Cal sensor Temp.: 99.47 Ambient Temp.: 87.39 F
01:41:55.791 -> Average Cal sensor Temp.: 99.81 Ambient Temp.: 87.39 F
01:41:57.348 -> Average Cal sensor Temp.: 100.03 Ambient Temp.: 87.42 F
01:41:58.911 -> Average Cal sensor Temp.: 99.81 Ambient Temp.: 87.39 F
01:42:00.467 -> Average Cal sensor Temp.: 99.91 Ambient Temp.: 87.39 F
01:42:02.027 -> Average Cal sensor Temp.: 99.68 Ambient Temp.: 87.39 F
01:42:03.620 -> Average Cal sensor Temp.: 99.86 Ambient Temp.: 87.39 F
01:42:05.178 -> Average Cal sensor Temp.: 100.04 Ambient Temp.: 87.42 F
01:42:06.731 -> Average Cal sensor Temp.: 99.19 Ambient Temp.: 87.39 F
01:42:08.288 -> Average Cal sensor Temp.: 99.31 Ambient Temp.: 87.35 F
01:42:09.843 -> Average Cal sensor Temp.: 99.50 Ambient Temp.: 87.39 F
01:42:11.399 -> Average Cal sensor Temp.: 99.81 Ambient Temp.: 87.42 F
01:42:12.982 -> Average Cal sensor Temp.: 99.88 Ambient Temp.: 87.39 F
01:42:14.540 -> Average Cal sensor Temp.: 99.45 Ambient Temp.: 87.35 F
01:42:16.094 -> Average Cal sensor Temp.: 99.09 Ambient Temp.: 87.39 F
01:42:17.646 -> Average Cal sensor Temp.: 98.93 Ambient Temp.: 87.28 F
01:42:19.208 -> Average Cal sensor Temp.: 99.73 Ambient Temp.: 87.31 F
01:42:20.768 -> Average Cal sensor Temp.: 99.60 Ambient Temp.: 87.31 F
01:42:22.329 -> Average Cal sensor Temp.: 99.23 Ambient Temp.: 87.28 F
01:42:23.888 -> Average Cal sensor Temp.: 99.48 Ambient Temp.: 87.28 F
01:42:25.450 -> Average Cal sensor Temp.: 99.46 Ambient Temp.: 87.28 F
01:42:26.998 -> Average Cal sensor Temp.: 99.73 Ambient Temp.: 87.28 F
01:42:28.547 -> Average Cal sensor Temp.: 99.50 Ambient Temp.: 87.31 F
01:42:30.107 -> Average Cal sensor Temp.: 99.60 Ambient Temp.: 87.21 F
01:42:31.663 -> Average Cal sensor Temp.: 99.86 Ambient Temp.: 87.21 F
01:42:33.256 -> Average Cal sensor Temp.: 99.58 Ambient Temp.: 87.21 F
01:42:34.806 -> Average Cal sensor Temp.: 99.77 Ambient Temp.: 87.21 F
01:42:36.368 -> Average Cal sensor Temp.: 99.78 Ambient Temp.: 87.28 F
01:42:37.926 -> Average Cal sensor Temp.: 99.92 Ambient Temp.: 87.28 F
01:42:39.488 -> Average Cal sensor Temp.: 99.60 Ambient Temp.: 87.21 F
01:42:41.046 -> Average Cal sensor Temp.: 100.09 Ambient Temp.: 87.28 F
01:42:42.597 -> Average Cal sensor Temp.: 99.34 Ambient Temp.: 87.28 F
01:42:44.153 -> Average Cal sensor Temp.: 99.48 Ambient Temp.: 87.21 F
01:42:45.741 -> Average Cal sensor Temp.: 99.46 Ambient Temp.: 87.17 F
01:42:47.298 -> Average Cal sensor Temp.: 99.03 Ambient Temp.: 87.17 F
01:42:48.852 -> Average Cal sensor Temp.: 99.60 Ambient Temp.: 87.13 F
01:42:50.412 -> Average Cal sensor Temp.: 99.42 Ambient Temp.: 87.17 F
01:42:51.967 -> Average Cal sensor Temp.: 99.54 Ambient Temp.: 87.17 F
01:42:53.521 -> Average Cal sensor Temp.: 98.82 Ambient Temp.: 87.10 F
01:42:55.080 -> Average Cal sensor Temp.: 99.56 Ambient Temp.: 87.10 F
01:42:56.632 -> Average Cal sensor Temp.: 99.11 Ambient Temp.: 87.10 F
01:42:58.189 -> Average Cal sensor Temp.: 99.07 Ambient Temp.: 87.06 F
01:42:59.743 -> Average Cal sensor Temp.: 99.47 Ambient Temp.: 87.10 F
01:43:01.305 -> Average Cal sensor Temp.: 99.07 Ambient Temp.: 87.06 F
01:43:02.861 -> Average Cal sensor Temp.: 98.95 Ambient Temp.: 86.95 F
01:43:04.424 -> Average Cal sensor Temp.: 99.32 Ambient Temp.: 86.95 F
01:43:05.985 -> Average Cal sensor Temp.: 99.35 Ambient Temp.: 86.95 F
01:43:07.545 -> Average Cal sensor Temp.: 99.53 Ambient Temp.: 86.92 F
01:43:09.104 -> Average Cal sensor Temp.: 100.10 Ambient Temp.: 86.99 F
01:43:10.666 -> Average Cal sensor Temp.: 100.01 Ambient Temp.: 86.92 F
01:43:12.223 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 86.88 F
01:43:13.778 -> Average Cal sensor Temp.: 99.95 Ambient Temp.: 86.92 F
01:43:15.366 -> Average Cal sensor Temp.: 100.09 Ambient Temp.: 86.92 F
01:43:16.912 -> Average Cal sensor Temp.: 100.14 Ambient Temp.: 86.88 F
01:43:18.471 -> Average Cal sensor Temp.: 99.85 Ambient Temp.: 86.88 F
01:43:20.063 -> Average Cal sensor Temp.: 99.72 Ambient Temp.: 86.95 F
01:43:21.620 -> Average Cal sensor Temp.: 99.68 Ambient Temp.: 86.92 F
01:43:23.175 -> Average Cal sensor Temp.: 99.79 Ambient Temp.: 86.88 F
01:43:24.728 -> Average Cal sensor Temp.: 99.84 Ambient Temp.: 86.92 F
01:43:26.287 -> Average Cal sensor Temp.: 100.05 Ambient Temp.: 86.95 F
01:43:27.876 -> Average Cal sensor Temp.: 99.77 Ambient Temp.: 86.95 F
01:43:29.434 -> Average Cal sensor Temp.: 100.11 Ambient Temp.: 86.95 F
01:43:30.993 -> Average Cal sensor Temp.: 99.89 Ambient Temp.: 86.92 F
01:43:32.546 -> Average Cal sensor Temp.: 99.88 Ambient Temp.: 86.99 F
01:43:34.138 -> Average Cal sensor Temp.: 100.10 Ambient Temp.: 86.95 F
01:43:35.693 -> Average Cal sensor Temp.: 99.37 Ambient Temp.: 86.95 F
01:43:37.251 -> Average Cal sensor Temp.: 99.25 Ambient Temp.: 86.99 F
01:43:38.812 -> Average Cal sensor Temp.: 99.93 Ambient Temp.: 86.95 F
01:43:40.368 -> Average Cal sensor Temp.: 99.65 Ambient Temp.: 86.95 F
01:43:41.919 -> Average Cal sensor Temp.: 99.54 Ambient Temp.: 86.95 F
01:43:43.476 -> Average Cal sensor Temp.: 100.10 Ambient Temp.: 86.92 F
01:43:45.060 -> Average Cal sensor Temp.: 99.92 Ambient Temp.: 86.92 F
01:43:46.607 -> Average Cal sensor Temp.: 99.30 Ambient Temp.: 86.81 F
01:43:48.165 -> Average Cal sensor Temp.: 99.38 Ambient Temp.: 86.92 F
01:43:49.716 -> Average Cal sensor Temp.: 99.64 Ambient Temp.: 86.92 F
01:43:51.309 -> Average Cal sensor Temp.: 99.50 Ambient Temp.: 86.92 F
01:43:52.866 -> Average Cal sensor Temp.: 99.73 Ambient Temp.: 86.92 F
01:43:54.422 -> Average Cal sensor Temp.: 99.72 Ambient Temp.: 86.92 F
01:43:55.975 -> Average Cal sensor Temp.: 99.45 Ambient Temp.: 86.88 F
01:43:57.536 -> Average Cal sensor Temp.: 99.88 Ambient Temp.: 86.92 F
01:43:59.090 -> Average Cal sensor Temp.: 99.47 Ambient Temp.: 86.95 F
01:44:00.677 -> Average Cal sensor Temp.: 99.32 Ambient Temp.: 86.81 F
01:44:02.202 -> Average Cal sensor Temp.: 99.73 Ambient Temp.: 86.88 F
01:44:03.764 -> Average Cal sensor Temp.: 99.94 Ambient Temp.: 86.81 F
01:44:05.357 -> Average Cal sensor Temp.: 99.88 Ambient Temp.: 86.81 F
01:44:06.909 -> Average Cal sensor Temp.: 99.59 Ambient Temp.: 86.81 F
01:44:08.471 -> Average Cal sensor Temp.: 100.09 Ambient Temp.: 86.88 F
01:44:10.026 -> Average Cal sensor Temp.: 99.87 Ambient Temp.: 86.92 F
01:44:11.616 -> Average Cal sensor Temp.: 100.26 Ambient Temp.: 86.95 F
01:44:13.172 -> Average Cal sensor Temp.: 100.00 Ambient Temp.: 86.95 F
01:44:14.729 -> Average Cal sensor Temp.: 100.28 Ambient Temp.: 86.95 F
01:44:16.287 -> Average Cal sensor Temp.: 100.18 Ambient Temp.: 86.95 F
01:44:17.872 -> Average Cal sensor Temp.: 99.96 Ambient Temp.: 86.99 F
01:44:19.431 -> Average Cal sensor Temp.: 99.86 Ambient Temp.: 86.95 F
01:44:20.988 -> Average Cal sensor Temp.: 100.05 Ambient Temp.: 86.99 F
01:44:22.545 -> Average Cal sensor Temp.: 100.18 Ambient Temp.: 87.03 F
01:44:24.133 -> Average Cal sensor Temp.: 100.07 Ambient Temp.: 87.03 F
01:44:25.693 -> Average Cal sensor Temp.: 100.04 Ambient Temp.: 87.10 F
01:44:27.254 -> Average Cal sensor Temp.: 100.04 Ambient Temp.: 87.03 F
01:44:28.808 -> Average Cal sensor Temp.: 99.65 Ambient Temp.: 87.06 F
01:44:30.396 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 87.03 F
01:44:31.951 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 87.03 F
01:44:33.505 -> Average Cal sensor Temp.: 100.06 Ambient Temp.: 87.03 F
01:44:35.093 -> Average Cal sensor Temp.: 100.31 Ambient Temp.: 87.03 F
01:44:36.650 -> Average Cal sensor Temp.: 100.76 Ambient Temp.: 87.03 F
01:44:38.206 -> Average Cal sensor Temp.: 100.21 Ambient Temp.: 87.03 F
01:44:39.760 -> Average Cal sensor Temp.: 100.04 Ambient Temp.: 86.99 F
01:44:41.354 -> Average Cal sensor Temp.: 100.80 Ambient Temp.: 87.10 F
01:44:42.911 -> Average Cal sensor Temp.: 100.35 Ambient Temp.: 87.13 F
01:44:44.472 -> Average Cal sensor Temp.: 100.44 Ambient Temp.: 87.13 F
01:44:46.026 -> Average Cal sensor Temp.: 100.07 Ambient Temp.: 87.10 F
01:44:47.607 -> Average Cal sensor Temp.: 100.15 Ambient Temp.: 87.06 F
01:44:49.155 -> Average Cal sensor Temp.: 100.16 Ambient Temp.: 87.13 F
01:44:50.712 -> Average Cal sensor Temp.: 100.30 Ambient Temp.: 87.10 F
01:44:52.302 -> Average Cal sensor Temp.: 99.79 Ambient Temp.: 87.13 F
01:44:53.858 -> Average Cal sensor Temp.: 99.90 Ambient Temp.: 87.13 F
01:44:55.420 -> Average Cal sensor Temp.: 99.99 Ambient Temp.: 87.17 F
01:44:56.974 -> Average Cal sensor Temp.: 100.06 Ambient Temp.: 87.10 F
01:44:58.531 -> Average Cal sensor Temp.: 100.26 Ambient Temp.: 87.03 F
01:45:00.120 -> Average Cal sensor Temp.: 100.28 Ambient Temp.: 87.10 F
01:45:01.674 -> Average Cal sensor Temp.: 100.77 Ambient Temp.: 87.10 F
01:45:03.232 -> Average Cal sensor Temp.: 100.50 Ambient Temp.: 87.10 F
01:45:04.789 -> Average Cal sensor Temp.: 100.22 Ambient Temp.: 87.10 F
01:45:06.377 -> Average Cal sensor Temp.: 100.26 Ambient Temp.: 87.13 F
01:45:07.930 -> Average Cal sensor Temp.: 100.62 Ambient Temp.: 87.13 F
01:45:09.489 -> Average Cal sensor Temp.: 100.51 Ambient Temp.: 87.17 F
01:45:11.049 -> Average Cal sensor Temp.: 100.55 Ambient Temp.: 87.17 F
01:45:12.606 -> Average Cal sensor Temp.: 100.59 Ambient Temp.: 87.13 F
01:45:14.193 -> Average Cal sensor Temp.: 100.71 Ambient Temp.: 87.10 F
01:45:15.749 -> Average Cal sensor Temp.: 100.57 Ambient Temp.: 87.17 F
01:45:17.305 -> Average Cal sensor Temp.: 100.50 Ambient Temp.: 87.13 F
01:45:18.857 -> Average Cal sensor Temp.: 100.48 Ambient Temp.: 87.13 F
01:45:20.452 -> Average Cal sensor Temp.: 100.72 Ambient Temp.: 87.17 F
01:45:22.007 -> Average Cal sensor Temp.: 100.50 Ambient Temp.: 87.21 F
01:45:23.567 -> Average Cal sensor Temp.: 100.37 Ambient Temp.: 87.17 F
01:45:25.121 -> Average Cal sensor Temp.: 100.40 Ambient Temp.: 87.21 F
01:45:26.710 -> Average Cal sensor Temp.: 100.13 Ambient Temp.: 87.28 F
01:45:28.262 -> Average Cal sensor Temp.: 100.41 Ambient Temp.: 87.21 F
01:45:29.816 -> Average Cal sensor Temp.: 100.26 Ambient Temp.: 87.17 F
01:45:31.380 -> Average Cal sensor Temp.: 100.52 Ambient Temp.: 87.21 F
01:45:32.968 -> Average Cal sensor Temp.: 100.80 Ambient Temp.: 87.28 F
01:45:34.520 -> Average Cal sensor Temp.: 100.50 Ambient Temp.: 87.28 F
01:45:36.078 -> Average Cal sensor Temp.: 100.32 Ambient Temp.: 87.17 F
01:45:37.635 -> Average Cal sensor Temp.: 100.46 Ambient Temp.: 87.17 F
01:45:39.225 -> Average Cal sensor Temp.: 100.26 Ambient Temp.: 87.28 F
01:45:40.785 -> Average Cal sensor Temp.: 100.89 Ambient Temp.: 87.28 F
01:45:42.339 -> Average Cal sensor Temp.: 100.48 Ambient Temp.: 87.21 F
01:45:43.896 -> Average Cal sensor Temp.: 100.88 Ambient Temp.: 87.28 F
01:45:45.489 -> Average Cal sensor Temp.: 101.17 Ambient Temp.: 87.31 F
01:45:47.051 -> Average Cal sensor Temp.: 100.65 Ambient Temp.: 87.31 F
01:45:48.613 -> Average Cal sensor Temp.: 100.71 Ambient Temp.: 87.31 F
01:45:50.171 -> Average Cal sensor Temp.: 101.31 Ambient Temp.: 87.35 F
01:45:51.724 -> Average Cal sensor Temp.: 100.71 Ambient Temp.: 87.31 F
01:45:53.279 -> Average Cal sensor Temp.: 100.66 Ambient Temp.: 87.31 F
01:45:54.869 -> Average Cal sensor Temp.: 100.76 Ambient Temp.: 87.31 F
01:45:56.423 -> Average Cal sensor Temp.: 100.84 Ambient Temp.: 87.31 F
01:45:57.974 -> Average Cal sensor Temp.: 100.51 Ambient Temp.: 87.28 F
01:45:59.560 -> Average Cal sensor Temp.: 100.65 Ambient Temp.: 87.31 F
01:46:01.118 -> Average Cal sensor Temp.: 100.50 Ambient Temp.: 87.35 F
01:46:02.670 -> Average Cal sensor Temp.: 100.84 Ambient Temp.: 87.35 F
01:46:04.227 -> Average Cal sensor Temp.: 100.66 Ambient Temp.: 87.35 F
01:46:05.818 -> Average Cal sensor Temp.: 100.94 Ambient Temp.: 87.35 F
01:46:07.373 -> Average Cal sensor Temp.: 100.78 Ambient Temp.: 87.42 F
01:46:08.930 -> Average Cal sensor Temp.: 101.17 Ambient Temp.: 87.39 F
01:46:10.519 -> Average Cal sensor Temp.: 100.90 Ambient Temp.: 87.39 F
01:46:12.077 -> Average Cal sensor Temp.: 100.78 Ambient Temp.: 87.46 F
01:46:13.636 -> Average Cal sensor Temp.: 100.33 Ambient Temp.: 87.42 F
01:46:15.191 -> Average Cal sensor Temp.: 100.75 Ambient Temp.: 87.42 F
01:46:16.778 -> Average Cal sensor Temp.: 100.71 Ambient Temp.: 87.46 F
01:46:18.333 -> Average Cal sensor Temp.: 100.65 Ambient Temp.: 87.49 F
01:46:19.887 -> Average Cal sensor Temp.: 100.50 Ambient Temp.: 87.42 F
01:46:21.460 -> Average Cal sensor Temp.: 101.03 Ambient Temp.: 87.46 F
01:46:23.017 -> Average Cal sensor Temp.: 100.69 Ambient Temp.: 87.53 F
01:46:24.578 -> Average Cal sensor Temp.: 100.59 Ambient Temp.: 87.49 F
01:46:26.141 -> Average Cal sensor Temp.: 100.67 Ambient Temp.: 87.53 F
01:46:27.729 -> Average Cal sensor Temp.: 100.63 Ambient Temp.: 87.49 F
01:46:29.291 -> Average Cal sensor Temp.: 100.64 Ambient Temp.: 87.53 F
01:46:30.846 -> Average Cal sensor Temp.: 100.24 Ambient Temp.: 87.46 F
01:46:32.403 -> Average Cal sensor Temp.: 100.65 Ambient Temp.: 87.46 F
01:46:33.994 -> Average Cal sensor Temp.: 100.45 Ambient Temp.: 87.42 F
01:46:35.545 -> Average Cal sensor Temp.: 100.25 Ambient Temp.: 87.42 F
01:46:37.101 -> Average Cal sensor Temp.: 100.83 Ambient Temp.: 87.39 F
01:46:38.688 -> Average Cal sensor Temp.: 100.74 Ambient Temp.: 87.42 F
01:46:40.241 -> Average Cal sensor Temp.: 100.67 Ambient Temp.: 87.46 F
01:46:41.793 -> Average Cal sensor Temp.: 100.31 Ambient Temp.: 87.42 F
01:46:43.385 -> Average Cal sensor Temp.: 100.30 Ambient Temp.: 87.39 F
01:46:44.939 -> Average Cal sensor Temp.: 100.14 Ambient Temp.: 87.39 F
01:46:46.491 -> Average Cal sensor Temp.: 99.88 Ambient Temp.: 87.39 F
01:46:48.078 -> Average Cal sensor Temp.: 100.22 Ambient Temp.: 87.35 F
01:46:49.640 -> Average Cal sensor Temp.: 100.69 Ambient Temp.: 87.35 F
01:46:51.196 -> Average Cal sensor Temp.: 100.73 Ambient Temp.: 87.35 F
01:46:52.752 -> Average Cal sensor Temp.: 100.68 Ambient Temp.: 87.39 F
01:46:54.342 -> Average Cal sensor Temp.: 100.51 Ambient Temp.: 87.35 F
01:46:55.898 -> Average Cal sensor Temp.: 100.61 Ambient Temp.: 87.35 F
01:46:57.451 -> Average Cal sensor Temp.: 100.57 Ambient Temp.: 87.35 F
01:46:59.013 -> Average Cal sensor Temp.: 100.51 Ambient Temp.: 87.39 F
01:47:00.569 -> Average Cal sensor Temp.: 100.82 Ambient Temp.: 87.39 F
01:47:02.159 -> Average Cal sensor Temp.: 100.13 Ambient Temp.: 87.35 F
01:47:03.716 -> Average Cal sensor Temp.: 100.31 Ambient Temp.: 87.31 F
01:47:05.267 -> Average Cal sensor Temp.: 100.12 Ambient Temp.: 87.28 F
01:47:06.854 -> Average Cal sensor Temp.: 100.05 Ambient Temp.: 87.28 F
01:47:08.403 -> Average Cal sensor Temp.: 99.85 Ambient Temp.: 87.21 F
01:47:09.958 -> Average Cal sensor Temp.: 100.35 Ambient Temp.: 87.28 F
01:47:11.541 -> Average Cal sensor Temp.: 100.68 Ambient Temp.: 87.28 F
01:47:13.098 -> Average Cal sensor Temp.: 100.71 Ambient Temp.: 87.28 F
01:47:14.649 -> Average Cal sensor Temp.: 100.50 Ambient Temp.: 87.31 F
01:47:16.237 -> Average Cal sensor Temp.: 100.61 Ambient Temp.: 87.28 F
01:47:17.787 -> Average Cal sensor Temp.: 100.67 Ambient Temp.: 87.28 F
01:47:19.351 -> Average Cal sensor Temp.: 100.66 Ambient Temp.: 87.28 F
01:47:20.939 -> Average Cal sensor Temp.: 100.93 Ambient Temp.: 87.31 F
01:47:22.488 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 87.31 F
01:47:24.042 -> Average Cal sensor Temp.: 100.51 Ambient Temp.: 87.31 F
01:47:25.602 -> Average Cal sensor Temp.: 100.00 Ambient Temp.: 87.28 F
01:47:27.194 -> Average Cal sensor Temp.: 99.86 Ambient Temp.: 87.28 F
01:47:28.753 -> Average Cal sensor Temp.: 100.15 Ambient Temp.: 87.21 F
01:47:30.306 -> Average Cal sensor Temp.: 100.46 Ambient Temp.: 87.21 F
01:47:31.866 -> Average Cal sensor Temp.: 100.71 Ambient Temp.: 87.21 F
01:47:33.421 -> Average Cal sensor Temp.: 100.78 Ambient Temp.: 87.28 F
01:47:35.010 -> Average Cal sensor Temp.: 100.71 Ambient Temp.: 87.28 F
01:47:36.557 -> Average Cal sensor Temp.: 100.75 Ambient Temp.: 87.21 F
01:47:38.113 -> Average Cal sensor Temp.: 100.63 Ambient Temp.: 87.28 F
01:47:39.702 -> Average Cal sensor Temp.: 100.55 Ambient Temp.: 87.21 F
01:47:41.265 -> Average Cal sensor Temp.: 100.45 Ambient Temp.: 87.21 F
01:47:42.822 -> Average Cal sensor Temp.: 100.64 Ambient Temp.: 87.31 F
01:47:44.381 -> Average Cal sensor Temp.: 100.58 Ambient Temp.: 87.28 F
01:47:45.944 -> Average Cal sensor Temp.: 100.78 Ambient Temp.: 87.31 F
01:47:47.500 -> Average Cal sensor Temp.: 100.77 Ambient Temp.: 87.31 F
01:47:49.061 -> Average Cal sensor Temp.: 100.65 Ambient Temp.: 87.35 F
01:47:50.651 -> Average Cal sensor Temp.: 100.61 Ambient Temp.: 87.39 F
01:47:52.201 -> Average Cal sensor Temp.: 100.27 Ambient Temp.: 87.39 F
01:47:53.757 -> Average Cal sensor Temp.: 100.63 Ambient Temp.: 87.42 F
01:47:55.344 -> Average Cal sensor Temp.: 100.50 Ambient Temp.: 87.39 F
01:47:56.902 -> Average Cal sensor Temp.: 100.13 Ambient Temp.: 87.35 F
01:47:58.452 -> Average Cal sensor Temp.: 100.27 Ambient Temp.: 87.35 F
01:48:00.033 -> Average Cal sensor Temp.: 100.70 Ambient Temp.: 87.35 F
01:48:01.590 -> Average Cal sensor Temp.: 100.92 Ambient Temp.: 87.39 F
01:48:03.175 -> Average Cal sensor Temp.: 100.89 Ambient Temp.: 87.39 F
01:48:04.727 -> Average Cal sensor Temp.: 101.00 Ambient Temp.: 87.39 F
01:48:06.308 -> Average Cal sensor Temp.: 101.01 Ambient Temp.: 87.39 F
01:48:07.863 -> Average Cal sensor Temp.: 101.16 Ambient Temp.: 87.39 F
01:48:09.412 -> Average Cal sensor Temp.: 100.92 Ambient Temp.: 87.42 F
01:48:10.998 -> Average Cal sensor Temp.: 101.00 Ambient Temp.: 87.39 F
01:48:12.546 -> Average Cal sensor Temp.: 100.82 Ambient Temp.: 87.39 F
01:48:14.124 -> Average Cal sensor Temp.: 100.74 Ambient Temp.: 87.39 F
01:48:15.679 -> Average Cal sensor Temp.: 100.32 Ambient Temp.: 87.42 F
01:48:17.234 -> Average Cal sensor Temp.: 100.59 Ambient Temp.: 87.39 F
01:48:18.824 -> Average Cal sensor Temp.: 101.25 Ambient Temp.: 87.39 F
01:48:20.375 -> Average Cal sensor Temp.: 100.91 Ambient Temp.: 87.39 F
01:48:21.930 -> Average Cal sensor Temp.: 100.17 Ambient Temp.: 87.42 F
01:48:23.487 -> Average Cal sensor Temp.: 100.36 Ambient Temp.: 87.39 F
01:48:25.075 -> Average Cal sensor Temp.: 100.01 Ambient Temp.: 87.35 F
01:48:26.631 -> Average Cal sensor Temp.: 99.99 Ambient Temp.: 87.31 F
01:48:28.192 -> Average Cal sensor Temp.: 99.98 Ambient Temp.: 87.31 F
01:48:29.750 -> Average Cal sensor Temp.: 100.65 Ambient Temp.: 87.31 F
01:48:31.310 -> Average Cal sensor Temp.: 100.71 Ambient Temp.: 87.31 F
01:48:32.865 -> Average Cal sensor Temp.: 100.85 Ambient Temp.: 87.31 F
01:48:34.453 -> Average Cal sensor Temp.: 100.80 Ambient Temp.: 87.28 F
01:48:36.000 -> Average Cal sensor Temp.: 100.74 Ambient Temp.: 87.31 F
01:48:37.591 -> Average Cal sensor Temp.: 99.34 Ambient Temp.: 87.28 F
01:48:39.141 -> Average Cal sensor Temp.: 99.81 Ambient Temp.: 87.21 F
01:48:40.697 -> Average Cal sensor Temp.: 100.15 Ambient Temp.: 87.21 F
01:48:42.250 -> Average Cal sensor Temp.: 100.66 Ambient Temp.: 87.21 F
01:48:43.806 -> Average Cal sensor Temp.: 101.02 Ambient Temp.: 87.21 F
01:48:45.396 -> Average Cal sensor Temp.: 101.06 Ambient Temp.: 87.21 F
01:48:46.948 -> Average Cal sensor Temp.: 100.54 Ambient Temp.: 87.21 F
01:48:48.503 -> Average Cal sensor Temp.: 100.45 Ambient Temp.: 87.21 F
01:48:50.093 -> Average Cal sensor Temp.: 100.55 Ambient Temp.: 87.21 F
01:48:51.655 -> Average Cal sensor Temp.: 100.98 Ambient Temp.: 87.17 F
01:48:53.215 -> Average Cal sensor Temp.: 101.04 Ambient Temp.: 87.17 F
01:48:54.771 -> Average Cal sensor Temp.: 101.59 Ambient Temp.: 87.21 F
01:48:56.325 -> Average Cal sensor Temp.: 100.58 Ambient Temp.: 87.17 F
01:48:57.881 -> Average Cal sensor Temp.: 101.24 Ambient Temp.: 87.21 F
01:48:59.473 -> Average Cal sensor Temp.: 100.80 Ambient Temp.: 87.21 F
01:49:01.033 -> Average Cal sensor Temp.: 100.13 Ambient Temp.: 87.17 F
01:49:02.581 -> Average Cal sensor Temp.: 99.98 Ambient Temp.: 87.13 F
01:49:04.132 -> Average Cal sensor Temp.: 99.95 Ambient Temp.: 87.06 F
01:49:05.720 -> Average Cal sensor Temp.: 100.14 Ambient Temp.: 87.06 F
01:49:07.274 -> Average Cal sensor Temp.: 100.20 Ambient Temp.: 87.06 F
01:49:08.829 -> Average Cal sensor Temp.: 100.02 Ambient Temp.: 86.95 F
01:49:10.417 -> Average Cal sensor Temp.: 100.64 Ambient Temp.: 87.03 F
01:49:11.974 -> Average Cal sensor Temp.: 100.68 Ambient Temp.: 86.95 F
01:49:13.531 -> Average Cal sensor Temp.: 101.33 Ambient Temp.: 86.99 F
01:49:15.121 -> Average Cal sensor Temp.: 100.72 Ambient Temp.: 86.99 F
01:49:16.674 -> Average Cal sensor Temp.: 100.69 Ambient Temp.: 86.99 F
01:49:18.226 -> Average Cal sensor Temp.: 100.90 Ambient Temp.: 86.95 F
01:49:19.813 -> Average Cal sensor Temp.: 101.03 Ambient Temp.: 86.92 F
01:49:21.368 -> Average Cal sensor Temp.: 101.52 Ambient Temp.: 86.99 F
01:49:22.924 -> Average Cal sensor Temp.: 101.16 Ambient Temp.: 86.99 F
01:49:24.514 -> Average Cal sensor Temp.: 101.60 Ambient Temp.: 87.03 F
01:49:26.070 -> Average Cal sensor Temp.: 101.19 Ambient Temp.: 86.99 F
01:49:27.621 -> Average Cal sensor Temp.: 101.22 Ambient Temp.: 87.03 F
01:49:29.202 -> Average Cal sensor Temp.: 101.13 Ambient Temp.: 86.99 F
01:49:30.761 -> Average Cal sensor Temp.: 100.84 Ambient Temp.: 86.99 F
01:49:32.312 -> Average Cal sensor Temp.: 100.64 Ambient Temp.: 87.03 F
01:49:33.904 -> Average Cal sensor Temp.: 100.71 Ambient Temp.: 87.03 F
01:49:35.459 -> Average Cal sensor Temp.: 101.01 Ambient Temp.: 86.99 F
01:49:37.016 -> Average Cal sensor Temp.: 101.09 Ambient Temp.: 87.06 F
01:49:38.574 -> Average Cal sensor Temp.: 101.32 Ambient Temp.: 87.03 F
01:49:40.132 -> Average Cal sensor Temp.: 101.35 Ambient Temp.: 87.06 F
01:49:41.722 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 87.13 F
01:49:43.281 -> Average Cal sensor Temp.: 101.48 Ambient Temp.: 87.10 F
01:49:44.838 -> Average Cal sensor Temp.: 101.39 Ambient Temp.: 87.17 F
01:49:46.396 -> Average Cal sensor Temp.: 100.87 Ambient Temp.: 87.17 F
01:49:47.949 -> Average Cal sensor Temp.: 100.74 Ambient Temp.: 87.17 F
01:49:49.515 -> Average Cal sensor Temp.: 100.87 Ambient Temp.: 87.13 F
01:49:51.079 -> Average Cal sensor Temp.: 101.01 Ambient Temp.: 87.13 F
01:49:52.656 -> Average Cal sensor Temp.: 100.93 Ambient Temp.: 87.17 F
01:49:54.212 -> Average Cal sensor Temp.: 101.17 Ambient Temp.: 87.13 F
01:49:55.775 -> Average Cal sensor Temp.: 101.21 Ambient Temp.: 87.13 F
01:49:57.330 -> Average Cal sensor Temp.: 101.32 Ambient Temp.: 87.21 F
01:49:58.890 -> Average Cal sensor Temp.: 101.34 Ambient Temp.: 87.17 F
01:50:00.479 -> Average Cal sensor Temp.: 100.96 Ambient Temp.: 87.13 F
01:50:02.031 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 87.28 F
01:50:03.583 -> Average Cal sensor Temp.: 100.73 Ambient Temp.: 87.17 F
01:50:05.163 -> Average Cal sensor Temp.: 100.71 Ambient Temp.: 87.21 F
01:50:06.713 -> Average Cal sensor Temp.: 100.59 Ambient Temp.: 87.21 F
01:50:08.285 -> Average Cal sensor Temp.: 100.58 Ambient Temp.: 87.17 F
01:50:09.833 -> Average Cal sensor Temp.: 100.35 Ambient Temp.: 87.21 F
01:50:11.422 -> Average Cal sensor Temp.: 100.21 Ambient Temp.: 87.17 F
01:50:12.977 -> Average Cal sensor Temp.: 100.20 Ambient Temp.: 87.21 F
01:50:14.531 -> Average Cal sensor Temp.: 100.06 Ambient Temp.: 87.10 F
01:50:16.087 -> Average Cal sensor Temp.: 100.61 Ambient Temp.: 87.06 F
01:50:17.671 -> Average Cal sensor Temp.: 100.60 Ambient Temp.: 87.10 F
01:50:19.227 -> Average Cal sensor Temp.: 101.34 Ambient Temp.: 87.10 F
01:50:20.810 -> Average Cal sensor Temp.: 101.40 Ambient Temp.: 87.06 F
01:50:22.362 -> Average Cal sensor Temp.: 101.60 Ambient Temp.: 87.06 F
01:50:23.915 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 87.10 F
01:50:25.504 -> Average Cal sensor Temp.: 100.84 Ambient Temp.: 87.03 F
01:50:27.051 -> Average Cal sensor Temp.: 101.26 Ambient Temp.: 87.06 F
01:50:28.605 -> Average Cal sensor Temp.: 101.32 Ambient Temp.: 86.99 F
01:50:30.193 -> Average Cal sensor Temp.: 100.97 Ambient Temp.: 87.06 F
01:50:31.749 -> Average Cal sensor Temp.: 101.57 Ambient Temp.: 86.95 F
01:50:33.303 -> Average Cal sensor Temp.: 101.32 Ambient Temp.: 87.03 F
01:50:34.859 -> Average Cal sensor Temp.: 101.56 Ambient Temp.: 87.03 F
01:50:36.449 -> Average Cal sensor Temp.: 101.56 Ambient Temp.: 87.03 F
01:50:37.991 -> Average Cal sensor Temp.: 101.31 Ambient Temp.: 87.06 F
01:50:39.551 -> Average Cal sensor Temp.: 101.61 Ambient Temp.: 87.03 F
01:50:41.144 -> Average Cal sensor Temp.: 101.78 Ambient Temp.: 87.06 F
01:50:42.702 -> Average Cal sensor Temp.: 101.79 Ambient Temp.: 87.10 F
01:50:44.257 -> Average Cal sensor Temp.: 101.44 Ambient Temp.: 87.13 F
01:50:45.816 -> Average Cal sensor Temp.: 101.31 Ambient Temp.: 87.06 F
01:50:47.371 -> Average Cal sensor Temp.: 101.58 Ambient Temp.: 87.10 F
01:50:48.958 -> Average Cal sensor Temp.: 102.05 Ambient Temp.: 87.17 F
01:50:50.511 -> Average Cal sensor Temp.: 101.99 Ambient Temp.: 87.17 F
01:50:52.067 -> Average Cal sensor Temp.: 101.84 Ambient Temp.: 87.21 F
01:50:53.625 -> Average Cal sensor Temp.: 101.93 Ambient Temp.: 87.28 F
01:50:55.181 -> Average Cal sensor Temp.: 101.68 Ambient Temp.: 87.21 F
01:50:56.772 -> Average Cal sensor Temp.: 101.57 Ambient Temp.: 87.31 F
01:50:58.334 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 87.28 F
01:50:59.889 -> Average Cal sensor Temp.: 101.42 Ambient Temp.: 87.28 F
01:51:01.444 -> Average Cal sensor Temp.: 101.37 Ambient Temp.: 87.28 F
01:51:02.999 -> Average Cal sensor Temp.: 101.64 Ambient Temp.: 87.31 F
01:51:04.590 -> Average Cal sensor Temp.: 101.87 Ambient Temp.: 87.35 F
01:51:06.146 -> Average Cal sensor Temp.: 101.58 Ambient Temp.: 87.35 F
01:51:07.704 -> Average Cal sensor Temp.: 101.58 Ambient Temp.: 87.39 F
01:51:09.259 -> Average Cal sensor Temp.: 101.54 Ambient Temp.: 87.39 F
01:51:10.815 -> Average Cal sensor Temp.: 101.54 Ambient Temp.: 87.35 F
01:51:12.405 -> Average Cal sensor Temp.: 101.58 Ambient Temp.: 87.42 F
01:51:13.964 -> Average Cal sensor Temp.: 101.85 Ambient Temp.: 87.39 F
01:51:15.522 -> Average Cal sensor Temp.: 101.25 Ambient Temp.: 87.46 F
01:51:17.081 -> Average Cal sensor Temp.: 101.41 Ambient Temp.: 87.42 F
01:51:18.642 -> Average Cal sensor Temp.: 101.35 Ambient Temp.: 87.35 F
01:51:20.200 -> Average Cal sensor Temp.: 101.80 Ambient Temp.: 87.39 F
01:51:21.756 -> Average Cal sensor Temp.: 101.49 Ambient Temp.: 87.42 F
01:51:23.353 -> Average Cal sensor Temp.: 101.50 Ambient Temp.: 87.46 F
01:51:24.907 -> Average Cal sensor Temp.: 101.97 Ambient Temp.: 87.46 F
01:51:26.457 -> Average Cal sensor Temp.: 101.58 Ambient Temp.: 87.53 F
01:51:28.015 -> Average Cal sensor Temp.: 101.51 Ambient Temp.: 87.49 F
01:51:29.574 -> Average Cal sensor Temp.: 101.18 Ambient Temp.: 87.53 F
01:51:31.169 -> Average Cal sensor Temp.: 101.11 Ambient Temp.: 87.53 F
01:51:32.727 -> Average Cal sensor Temp.: 101.62 Ambient Temp.: 87.57 F
01:51:34.283 -> Average Cal sensor Temp.: 101.62 Ambient Temp.: 87.60 F
01:51:35.840 -> Average Cal sensor Temp.: 101.69 Ambient Temp.: 87.60 F
01:51:37.401 -> Average Cal sensor Temp.: 101.56 Ambient Temp.: 87.57 F
01:51:38.959 -> Average Cal sensor Temp.: 101.14 Ambient Temp.: 87.60 F
01:51:40.544 -> Average Cal sensor Temp.: 101.25 Ambient Temp.: 87.57 F
01:51:42.102 -> Average Cal sensor Temp.: 101.77 Ambient Temp.: 87.57 F
01:51:43.659 -> Average Cal sensor Temp.: 101.78 Ambient Temp.: 87.67 F
01:51:45.212 -> Average Cal sensor Temp.: 101.80 Ambient Temp.: 87.67 F
01:51:46.766 -> Average Cal sensor Temp.: 102.26 Ambient Temp.: 87.67 F
01:51:48.354 -> Average Cal sensor Temp.: 101.92 Ambient Temp.: 87.75 F
01:51:49.908 -> Average Cal sensor Temp.: 102.06 Ambient Temp.: 87.71 F
01:51:51.461 -> Average Cal sensor Temp.: 101.48 Ambient Temp.: 87.82 F
01:51:53.049 -> Average Cal sensor Temp.: 101.45 Ambient Temp.: 87.75 F
01:51:54.599 -> Average Cal sensor Temp.: 101.39 Ambient Temp.: 87.82 F
01:51:56.154 -> Average Cal sensor Temp.: 101.17 Ambient Temp.: 87.78 F
01:51:57.741 -> Average Cal sensor Temp.: 101.44 Ambient Temp.: 87.75 F
01:51:59.293 -> Average Cal sensor Temp.: 101.64 Ambient Temp.: 87.85 F
01:52:00.851 -> Average Cal sensor Temp.: 101.66 Ambient Temp.: 87.75 F
01:52:02.409 -> Average Cal sensor Temp.: 101.84 Ambient Temp.: 87.75 F
01:52:03.966 -> Average Cal sensor Temp.: 101.53 Ambient Temp.: 87.85 F
01:52:05.560 -> Average Cal sensor Temp.: 101.54 Ambient Temp.: 87.75 F
01:52:07.115 -> Average Cal sensor Temp.: 101.83 Ambient Temp.: 87.82 F
01:52:08.672 -> Average Cal sensor Temp.: 101.61 Ambient Temp.: 87.78 F
01:52:10.228 -> Average Cal sensor Temp.: 101.84 Ambient Temp.: 87.82 F
01:52:11.783 -> Average Cal sensor Temp.: 101.69 Ambient Temp.: 87.78 F
01:52:13.374 -> Average Cal sensor Temp.: 101.12 Ambient Temp.: 87.82 F
01:52:14.931 -> Average Cal sensor Temp.: 101.66 Ambient Temp.: 87.82 F
01:52:16.483 -> Average Cal sensor Temp.: 101.64 Ambient Temp.: 87.85 F
01:52:18.046 -> Average Cal sensor Temp.: 101.75 Ambient Temp.: 87.85 F
01:52:19.603 -> Average Cal sensor Temp.: 101.55 Ambient Temp.: 87.89 F
01:52:21.192 -> Average Cal sensor Temp.: 101.37 Ambient Temp.: 87.89 F
01:52:22.749 -> Average Cal sensor Temp.: 101.59 Ambient Temp.: 87.85 F
01:52:24.304 -> Average Cal sensor Temp.: 101.94 Ambient Temp.: 87.82 F
01:52:25.856 -> Average Cal sensor Temp.: 101.42 Ambient Temp.: 87.85 F
01:52:27.447 -> Average Cal sensor Temp.: 101.16 Ambient Temp.: 87.93 F
01:52:29.005 -> Average Cal sensor Temp.: 101.33 Ambient Temp.: 87.89 F
01:52:30.565 -> Average Cal sensor Temp.: 101.21 Ambient Temp.: 87.82 F
01:52:32.120 -> Average Cal sensor Temp.: 101.79 Ambient Temp.: 87.82 F
01:52:33.680 -> Average Cal sensor Temp.: 101.52 Ambient Temp.: 87.89 F
01:52:35.235 -> Average Cal sensor Temp.: 101.47 Ambient Temp.: 87.82 F
01:52:36.827 -> Average Cal sensor Temp.: 101.79 Ambient Temp.: 87.82 F
01:52:38.380 -> Average Cal sensor Temp.: 101.38 Ambient Temp.: 87.82 F
01:52:39.937 -> Average Cal sensor Temp.: 101.66 Ambient Temp.: 87.75 F
01:52:41.493 -> Average Cal sensor Temp.: 101.52 Ambient Temp.: 87.82 F
01:52:43.081 -> Average Cal sensor Temp.: 101.43 Ambient Temp.: 87.85 F
01:52:44.638 -> Average Cal sensor Temp.: 101.44 Ambient Temp.: 87.82 F
01:52:46.189 -> Average Cal sensor Temp.: 101.61 Ambient Temp.: 87.82 F
01:52:47.777 -> Average Cal sensor Temp.: 101.21 Ambient Temp.: 87.85 F
01:52:49.325 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 87.82 F
01:52:50.879 -> Average Cal sensor Temp.: 101.59 Ambient Temp.: 87.82 F
01:52:52.466 -> Average Cal sensor Temp.: 101.61 Ambient Temp.: 87.82 F
01:52:54.018 -> Average Cal sensor Temp.: 100.94 Ambient Temp.: 87.82 F
01:52:55.573 -> Average Cal sensor Temp.: 101.27 Ambient Temp.: 87.82 F
01:52:57.158 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 87.78 F
01:52:58.715 -> Average Cal sensor Temp.: 101.16 Ambient Temp.: 87.82 F
01:53:00.267 -> Average Cal sensor Temp.: 101.51 Ambient Temp.: 87.78 F
01:53:01.822 -> Average Cal sensor Temp.: 101.22 Ambient Temp.: 87.75 F
01:53:03.381 -> Average Cal sensor Temp.: 101.22 Ambient Temp.: 87.78 F
01:53:04.967 -> Average Cal sensor Temp.: 100.50 Ambient Temp.: 87.71 F
01:53:06.514 -> Average Cal sensor Temp.: 101.24 Ambient Temp.: 87.71 F
01:53:08.069 -> Average Cal sensor Temp.: 101.59 Ambient Temp.: 87.67 F
01:53:09.661 -> Average Cal sensor Temp.: 101.48 Ambient Temp.: 87.71 F
01:53:11.216 -> Average Cal sensor Temp.: 101.52 Ambient Temp.: 87.67 F
01:53:12.766 -> Average Cal sensor Temp.: 101.10 Ambient Temp.: 87.67 F
01:53:14.324 -> Average Cal sensor Temp.: 101.64 Ambient Temp.: 87.71 F
01:53:15.914 -> Average Cal sensor Temp.: 100.95 Ambient Temp.: 87.67 F
01:53:17.464 -> Average Cal sensor Temp.: 101.05 Ambient Temp.: 87.60 F
01:53:19.015 -> Average Cal sensor Temp.: 101.09 Ambient Temp.: 87.57 F
01:53:20.604 -> Average Cal sensor Temp.: 101.60 Ambient Temp.: 87.57 F
01:53:22.156 -> Average Cal sensor Temp.: 101.51 Ambient Temp.: 87.60 F
01:53:23.714 -> Average Cal sensor Temp.: 101.64 Ambient Temp.: 87.57 F
01:53:25.272 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 87.60 F
01:53:26.831 -> Average Cal sensor Temp.: 101.26 Ambient Temp.: 87.67 F
01:53:28.421 -> Average Cal sensor Temp.: 100.96 Ambient Temp.: 87.57 F
01:53:29.981 -> Average Cal sensor Temp.: 101.18 Ambient Temp.: 87.53 F
01:53:31.534 -> Average Cal sensor Temp.: 101.09 Ambient Temp.: 87.57 F
01:53:33.086 -> Average Cal sensor Temp.: 100.79 Ambient Temp.: 87.53 F
01:53:34.678 -> Average Cal sensor Temp.: 101.63 Ambient Temp.: 87.49 F
01:53:36.230 -> Average Cal sensor Temp.: 101.34 Ambient Temp.: 87.53 F
01:53:37.791 -> Average Cal sensor Temp.: 101.48 Ambient Temp.: 87.49 F
01:53:39.353 -> Average Cal sensor Temp.: 101.26 Ambient Temp.: 87.46 F
01:53:40.905 -> Average Cal sensor Temp.: 101.48 Ambient Temp.: 87.53 F
01:53:42.489 -> Average Cal sensor Temp.: 101.94 Ambient Temp.: 87.46 F
01:53:44.047 -> Average Cal sensor Temp.: 102.00 Ambient Temp.: 87.53 F
01:53:45.606 -> Average Cal sensor Temp.: 101.95 Ambient Temp.: 87.53 F
01:53:47.160 -> Average Cal sensor Temp.: 101.69 Ambient Temp.: 87.53 F
01:53:48.749 -> Average Cal sensor Temp.: 101.71 Ambient Temp.: 87.49 F
01:53:50.309 -> Average Cal sensor Temp.: 101.77 Ambient Temp.: 87.60 F
01:53:51.861 -> Average Cal sensor Temp.: 101.93 Ambient Temp.: 87.57 F
01:53:53.421 -> Average Cal sensor Temp.: 101.84 Ambient Temp.: 87.57 F
01:53:54.976 -> Average Cal sensor Temp.: 101.69 Ambient Temp.: 87.60 F
01:53:56.568 -> Average Cal sensor Temp.: 101.83 Ambient Temp.: 87.57 F
01:53:58.119 -> Average Cal sensor Temp.: 101.27 Ambient Temp.: 87.60 F
01:53:59.676 -> Average Cal sensor Temp.: 101.09 Ambient Temp.: 87.53 F
01:54:01.264 -> Average Cal sensor Temp.: 101.50 Ambient Temp.: 87.53 F
01:54:02.813 -> Average Cal sensor Temp.: 101.71 Ambient Temp.: 87.49 F
01:54:04.373 -> Average Cal sensor Temp.: 102.10 Ambient Temp.: 87.57 F
01:54:05.928 -> Average Cal sensor Temp.: 101.92 Ambient Temp.: 87.53 F
01:54:07.482 -> Average Cal sensor Temp.: 101.70 Ambient Temp.: 87.57 F
01:54:09.071 -> Average Cal sensor Temp.: 101.33 Ambient Temp.: 87.57 F
01:54:10.623 -> Average Cal sensor Temp.: 101.19 Ambient Temp.: 87.57 F
01:54:12.181 -> Average Cal sensor Temp.: 101.62 Ambient Temp.: 87.57 F
01:54:13.736 -> Average Cal sensor Temp.: 101.86 Ambient Temp.: 87.57 F
01:54:15.321 -> Average Cal sensor Temp.: 102.26 Ambient Temp.: 87.60 F
01:54:16.874 -> Average Cal sensor Temp.: 101.47 Ambient Temp.: 87.57 F
01:54:18.431 -> Average Cal sensor Temp.: 101.31 Ambient Temp.: 87.53 F
01:54:19.990 -> Average Cal sensor Temp.: 101.26 Ambient Temp.: 87.53 F
01:54:21.548 -> Average Cal sensor Temp.: 101.63 Ambient Temp.: 87.57 F
01:54:23.136 -> Average Cal sensor Temp.: 101.32 Ambient Temp.: 87.57 F
01:54:24.692 -> Average Cal sensor Temp.: 101.37 Ambient Temp.: 87.46 F
01:54:26.251 -> Average Cal sensor Temp.: 101.24 Ambient Temp.: 87.46 F
01:54:27.808 -> Average Cal sensor Temp.: 101.61 Ambient Temp.: 87.49 F
01:54:29.370 -> Average Cal sensor Temp.: 101.85 Ambient Temp.: 87.46 F
01:54:30.929 -> Average Cal sensor Temp.: 101.49 Ambient Temp.: 87.49 F
01:54:32.517 -> Average Cal sensor Temp.: 101.55 Ambient Temp.: 87.46 F
01:54:34.078 -> Average Cal sensor Temp.: 101.02 Ambient Temp.: 87.39 F
01:54:35.636 -> Average Cal sensor Temp.: 101.67 Ambient Temp.: 87.39 F
01:54:37.192 -> Average Cal sensor Temp.: 101.52 Ambient Temp.: 87.42 F
01:54:38.751 -> Average Cal sensor Temp.: 101.40 Ambient Temp.: 87.42 F
01:54:40.330 -> Average Cal sensor Temp.: 101.68 Ambient Temp.: 87.42 F
01:54:41.885 -> Average Cal sensor Temp.: 101.81 Ambient Temp.: 87.35 F
01:54:43.437 -> Average Cal sensor Temp.: 101.73 Ambient Temp.: 87.39 F
01:54:45.024 -> Average Cal sensor Temp.: 101.89 Ambient Temp.: 87.42 F
01:54:46.580 -> Average Cal sensor Temp.: 101.40 Ambient Temp.: 87.35 F
01:54:48.133 -> Average Cal sensor Temp.: 101.77 Ambient Temp.: 87.39 F
01:54:49.719 -> Average Cal sensor Temp.: 101.53 Ambient Temp.: 87.35 F
01:54:51.274 -> Average Cal sensor Temp.: 101.15 Ambient Temp.: 87.31 F
01:54:52.828 -> Average Cal sensor Temp.: 101.48 Ambient Temp.: 87.21 F
01:54:54.382 -> Average Cal sensor Temp.: 101.48 Ambient Temp.: 87.28 F
01:54:55.968 -> Average Cal sensor Temp.: 101.46 Ambient Temp.: 87.28 F
01:54:57.518 -> Average Cal sensor Temp.: 101.41 Ambient Temp.: 87.31 F
01:54:59.069 -> Average Cal sensor Temp.: 101.74 Ambient Temp.: 87.28 F
01:55:00.648 -> Average Cal sensor Temp.: 102.02 Ambient Temp.: 87.31 F
01:55:02.199 -> Average Cal sensor Temp.: 101.66 Ambient Temp.: 87.35 F
01:55:03.787 -> Average Cal sensor Temp.: 101.79 Ambient Temp.: 87.35 F
01:55:05.334 -> Average Cal sensor Temp.: 101.99 Ambient Temp.: 87.39 F
01:55:06.887 -> Average Cal sensor Temp.: 101.97 Ambient Temp.: 87.35 F
01:55:08.479 -> Average Cal sensor Temp.: 101.75 Ambient Temp.: 87.46 F
01:55:10.035 -> Average Cal sensor Temp.: 101.76 Ambient Temp.: 87.49 F
01:55:11.589 -> Average Cal sensor Temp.: 101.66 Ambient Temp.: 87.49 F
01:55:13.145 -> Average Cal sensor Temp.: 101.80 Ambient Temp.: 87.46 F
01:55:14.734 -> Average Cal sensor Temp.: 102.41 Ambient Temp.: 87.49 F
01:55:16.290 -> Average Cal sensor Temp.: 102.56 Ambient Temp.: 87.49 F
01:55:17.845 -> Average Cal sensor Temp.: 102.35 Ambient Temp.: 87.57 F
01:55:19.397 -> Average Cal sensor Temp.: 102.31 Ambient Temp.: 87.53 F
01:55:20.987 -> Average Cal sensor Temp.: 102.18 Ambient Temp.: 87.53 F
01:55:22.542 -> Average Cal sensor Temp.: 101.93 Ambient Temp.: 87.57 F
01:55:24.100 -> Average Cal sensor Temp.: 102.26 Ambient Temp.: 87.60 F
01:55:25.652 -> Average Cal sensor Temp.: 101.75 Ambient Temp.: 87.60 F
01:55:27.211 -> Average Cal sensor Temp.: 101.60 Ambient Temp.: 87.60 F
01:55:28.801 -> Average Cal sensor Temp.: 101.37 Ambient Temp.: 87.60 F
01:55:30.356 -> Average Cal sensor Temp.: 102.14 Ambient Temp.: 87.57 F
01:55:31.916 -> Average Cal sensor Temp.: 100.94 Ambient Temp.: 87.57 F
01:55:33.472 -> Average Cal sensor Temp.: 101.53 Ambient Temp.: 87.57 F
01:55:35.034 -> Average Cal sensor Temp.: 102.43 Ambient Temp.: 87.57 F
01:55:36.621 -> Average Cal sensor Temp.: 102.27 Ambient Temp.: 87.71 F
01:55:38.180 -> Average Cal sensor Temp.: 102.63 Ambient Temp.: 87.60 F
01:55:39.737 -> Average Cal sensor Temp.: 102.16 Ambient Temp.: 87.60 F
01:55:41.296 -> Average Cal sensor Temp.: 102.01 Ambient Temp.: 87.67 F
01:55:42.856 -> Average Cal sensor Temp.: 102.38 Ambient Temp.: 87.60 F
01:55:44.413 -> Average Cal sensor Temp.: 102.19 Ambient Temp.: 87.67 F
01:55:46.005 -> Average Cal sensor Temp.: 102.01 Ambient Temp.: 87.57 F
01:55:47.563 -> Average Cal sensor Temp.: 101.70 Ambient Temp.: 87.60 F
01:55:49.123 -> Average Cal sensor Temp.: 101.80 Ambient Temp.: 87.67 F
01:55:50.684 -> Average Cal sensor Temp.: 101.53 Ambient Temp.: 87.57 F
01:55:52.236 -> Average Cal sensor Temp.: 101.96 Ambient Temp.: 87.57 F
01:55:53.795 -> Average Cal sensor Temp.: 101.54 Ambient Temp.: 87.49 F
01:55:55.380 -> Average Cal sensor Temp.: 101.77 Ambient Temp.: 87.49 F
01:55:56.934 -> Average Cal sensor Temp.: 102.17 Ambient Temp.: 87.53 F
01:55:58.490 -> Average Cal sensor Temp.: 102.48 Ambient Temp.: 87.53 F
01:56:00.076 -> Average Cal sensor Temp.: 102.59 Ambient Temp.: 87.53 F
01:56:01.632 -> Average Cal sensor Temp.: 102.20 Ambient Temp.: 87.49 F
01:56:03.192 -> Average Cal sensor Temp.: 102.09 Ambient Temp.: 87.53 F
01:56:04.751 -> Average Cal sensor Temp.: 102.42 Ambient Temp.: 87.53 F
01:56:06.311 -> Average Cal sensor Temp.: 102.67 Ambient Temp.: 87.53 F
01:56:07.866 -> Average Cal sensor Temp.: 102.31 Ambient Temp.: 87.60 F
01:56:09.454 -> Average Cal sensor Temp.: 102.73 Ambient Temp.: 87.57 F
01:56:11.007 -> Average Cal sensor Temp.: 102.65 Ambient Temp.: 87.57 F
01:56:12.556 -> Average Cal sensor Temp.: 102.51 Ambient Temp.: 87.60 F
01:56:14.151 -> Average Cal sensor Temp.: 102.59 Ambient Temp.: 87.53 F
01:56:15.707 -> Average Cal sensor Temp.: 101.92 Ambient Temp.: 87.57 F
01:56:17.261 -> Average Cal sensor Temp.: 102.24 Ambient Temp.: 87.67 F
01:56:18.813 -> Average Cal sensor Temp.: 102.55 Ambient Temp.: 87.60 F
01:56:20.402 -> Average Cal sensor Temp.: 102.51 Ambient Temp.: 87.60 F
01:56:21.956 -> Average Cal sensor Temp.: 102.47 Ambient Temp.: 87.57 F
01:56:23.513 -> Average Cal sensor Temp.: 102.29 Ambient Temp.: 87.60 F
01:56:25.069 -> Average Cal sensor Temp.: 102.08 Ambient Temp.: 87.57 F
01:56:26.661 -> Average Cal sensor Temp.: 101.91 Ambient Temp.: 87.60 F
01:56:28.209 -> Average Cal sensor Temp.: 101.68 Ambient Temp.: 87.60 F
01:56:29.763 -> Average Cal sensor Temp.: 102.34 Ambient Temp.: 87.57 F
01:56:31.353 -> Average Cal sensor Temp.: 102.26 Ambient Temp.: 87.57 F
01:56:32.911 -> Average Cal sensor Temp.: 102.69 Ambient Temp.: 87.60 F
01:56:34.460 -> Average Cal sensor Temp.: 102.11 Ambient Temp.: 87.67 F
01:56:36.014 -> Average Cal sensor Temp.: 102.29 Ambient Temp.: 87.57 F
01:56:37.596 -> Average Cal sensor Temp.: 102.44 Ambient Temp.: 87.57 F
01:56:39.150 -> Average Cal sensor Temp.: 102.35 Ambient Temp.: 87.53 F
01:56:40.708 -> Average Cal sensor Temp.: 102.23 Ambient Temp.: 87.53 F
01:56:42.294 -> Average Cal sensor Temp.: 102.65 Ambient Temp.: 87.57 F
01:56:43.854 -> Average Cal sensor Temp.: 103.00 Ambient Temp.: 87.53 F
01:56:45.408 -> Average Cal sensor Temp.: 102.60 Ambient Temp.: 87.53 F
01:56:46.991 -> Average Cal sensor Temp.: 102.25 Ambient Temp.: 87.57 F
01:56:48.542 -> Average Cal sensor Temp.: 102.26 Ambient Temp.: 87.57 F
01:56:50.086 -> Average Cal sensor Temp.: 102.02 Ambient Temp.: 87.53 F
01:56:51.672 -> Average Cal sensor Temp.: 102.59 Ambient Temp.: 87.57 F
01:56:53.228 -> Average Cal sensor Temp.: 102.19 Ambient Temp.: 87.53 F
01:56:54.780 -> Average Cal sensor Temp.: 102.53 Ambient Temp.: 87.53 F
01:56:56.366 -> Average Cal sensor Temp.: 102.35 Ambient Temp.: 87.53 F
01:56:57.913 -> Average Cal sensor Temp.: 102.65 Ambient Temp.: 87.53 F
01:56:59.495 -> Average Cal sensor Temp.: 102.40 Ambient Temp.: 87.57 F
01:57:01.051 -> Average Cal sensor Temp.: 102.53 Ambient Temp.: 87.57 F
01:57:02.610 -> Average Cal sensor Temp.: 102.31 Ambient Temp.: 87.57 F
01:57:04.171 -> Average Cal sensor Temp.: 102.70 Ambient Temp.: 87.57 F
01:57:05.725 -> Average Cal sensor Temp.: 102.87 Ambient Temp.: 87.60 F
01:57:07.308 -> Average Cal sensor Temp.: 102.41 Ambient Temp.: 87.60 F
01:57:08.861 -> Average Cal sensor Temp.: 102.08 Ambient Temp.: 87.60 F
01:57:10.415 -> Average Cal sensor Temp.: 102.76 Ambient Temp.: 87.67 F
01:57:11.973 -> Average Cal sensor Temp.: 103.13 Ambient Temp.: 87.60 F
01:57:13.563 -> Average Cal sensor Temp.: 103.09 Ambient Temp.: 87.60 F
01:57:15.116 -> Average Cal sensor Temp.: 102.66 Ambient Temp.: 87.71 F
01:57:16.670 -> Average Cal sensor Temp.: 102.71 Ambient Temp.: 87.71 F
01:57:18.258 -> Average Cal sensor Temp.: 102.65 Ambient Temp.: 87.71 F
01:57:19.817 -> Average Cal sensor Temp.: 102.46 Ambient Temp.: 87.67 F
01:57:21.371 -> Average Cal sensor Temp.: 102.80 Ambient Temp.: 87.78 F
01:57:22.920 -> Average Cal sensor Temp.: 103.00 Ambient Temp.: 87.75 F
01:57:24.512 -> Average Cal sensor Temp.: 103.09 Ambient Temp.: 87.82 F
01:57:26.071 -> Average Cal sensor Temp.: 102.97 Ambient Temp.: 87.78 F
01:57:27.622 -> Average Cal sensor Temp.: 102.48 Ambient Temp.: 87.82 F
01:57:29.212 -> Average Cal sensor Temp.: 102.89 Ambient Temp.: 87.85 F
01:57:30.772 -> Average Cal sensor Temp.: 102.06 Ambient Temp.: 87.82 F
01:57:32.325 -> Average Cal sensor Temp.: 102.06 Ambient Temp.: 87.78 F
01:57:33.886 -> Average Cal sensor Temp.: 102.74 Ambient Temp.: 87.82 F
01:57:35.441 -> Average Cal sensor Temp.: 103.14 Ambient Temp.: 87.93 F
01:57:37.031 -> Average Cal sensor Temp.: 102.60 Ambient Temp.: 87.85 F
01:57:38.588 -> Average Cal sensor Temp.: 102.69 Ambient Temp.: 87.89 F
01:57:40.142 -> Average Cal sensor Temp.: 102.17 Ambient Temp.: 87.89 F
01:57:41.726 -> Average Cal sensor Temp.: 102.69 Ambient Temp.: 87.89 F
01:57:43.283 -> Average Cal sensor Temp.: 102.47 Ambient Temp.: 87.93 F
01:57:44.837 -> Average Cal sensor Temp.: 102.56 Ambient Temp.: 87.96 F
01:57:46.396 -> Average Cal sensor Temp.: 102.58 Ambient Temp.: 87.93 F
01:57:47.986 -> Average Cal sensor Temp.: 102.23 Ambient Temp.: 87.89 F
01:57:49.544 -> Average Cal sensor Temp.: 102.73 Ambient Temp.: 87.89 F
01:57:51.102 -> Average Cal sensor Temp.: 102.87 Ambient Temp.: 87.93 F
01:57:52.652 -> Average Cal sensor Temp.: 103.35 Ambient Temp.: 87.93 F
01:57:54.245 -> Average Cal sensor Temp.: 102.53 Ambient Temp.: 87.93 F
01:57:55.800 -> Average Cal sensor Temp.: 102.80 Ambient Temp.: 87.96 F
01:57:57.355 -> Average Cal sensor Temp.: 102.72 Ambient Temp.: 87.93 F
01:57:58.909 -> Average Cal sensor Temp.: 102.59 Ambient Temp.: 88.00 F
01:58:00.501 -> Average Cal sensor Temp.: 102.61 Ambient Temp.: 87.93 F
01:58:02.055 -> Average Cal sensor Temp.: 102.07 Ambient Temp.: 88.00 F
01:58:03.612 -> Average Cal sensor Temp.: 102.05 Ambient Temp.: 87.93 F
01:58:05.174 -> Average Cal sensor Temp.: 102.15 Ambient Temp.: 88.00 F
01:58:06.730 -> Average Cal sensor Temp.: 101.93 Ambient Temp.: 87.89 F
01:58:08.315 -> Average Cal sensor Temp.: 102.31 Ambient Temp.: 88.07 F
01:58:09.871 -> Average Cal sensor Temp.: 102.26 Ambient Temp.: 87.96 F
01:58:11.432 -> Average Cal sensor Temp.: 102.28 Ambient Temp.: 87.96 F
01:58:12.989 -> Average Cal sensor Temp.: 102.24 Ambient Temp.: 87.96 F
01:58:14.582 -> Average Cal sensor Temp.: 102.58 Ambient Temp.: 87.89 F
01:58:16.136 -> Average Cal sensor Temp.: 102.42 Ambient Temp.: 87.96 F
01:58:17.688 -> Average Cal sensor Temp.: 102.64 Ambient Temp.: 87.93 F
01:58:19.279 -> Average Cal sensor Temp.: 102.64 Ambient Temp.: 87.96 F
01:58:20.838 -> Average Cal sensor Temp.: 102.12 Ambient Temp.: 87.96 F
01:58:22.388 -> Average Cal sensor Temp.: 102.03 Ambient Temp.: 87.96 F
01:58:23.943 -> Average Cal sensor Temp.: 102.12 Ambient Temp.: 87.96 F
01:58:25.536 -> Average Cal sensor Temp.: 102.01 Ambient Temp.: 87.96 F
01:58:27.089 -> Average Cal sensor Temp.: 102.93 Ambient Temp.: 88.07 F
01:58:28.639 -> Average Cal sensor Temp.: 102.42 Ambient Temp.: 88.00 F
01:58:30.226 -> Average Cal sensor Temp.: 102.86 Ambient Temp.: 88.00 F
01:58:31.784 -> Average Cal sensor Temp.: 102.32 Ambient Temp.: 88.07 F
01:58:33.346 -> Average Cal sensor Temp.: 102.33 Ambient Temp.: 87.96 F
01:58:34.900 -> Average Cal sensor Temp.: 102.41 Ambient Temp.: 87.96 F
01:58:36.488 -> Average Cal sensor Temp.: 102.81 Ambient Temp.: 87.96 F
01:58:38.041 -> Average Cal sensor Temp.: 102.50 Ambient Temp.: 88.07 F
01:58:39.595 -> Average Cal sensor Temp.: 102.56 Ambient Temp.: 88.00 F
01:58:41.183 -> Average Cal sensor Temp.: 101.72 Ambient Temp.: 88.00 F
01:58:42.732 -> Average Cal sensor Temp.: 101.85 Ambient Temp.: 87.96 F
01:58:44.285 -> Average Cal sensor Temp.: 102.27 Ambient Temp.: 88.00 F
01:58:45.868 -> Average Cal sensor Temp.: 101.83 Ambient Temp.: 88.00 F
01:58:47.418 -> Average Cal sensor Temp.: 101.69 Ambient Temp.: 88.07 F
01:58:48.974 -> Average Cal sensor Temp.: 101.60 Ambient Temp.: 87.93 F
01:58:50.558 -> Average Cal sensor Temp.: 102.51 Ambient Temp.: 88.00 F
01:58:52.107 -> Average Cal sensor Temp.: 102.64 Ambient Temp.: 87.93 F
01:58:53.692 -> Average Cal sensor Temp.: 102.35 Ambient Temp.: 87.96 F
01:58:55.244 -> Average Cal sensor Temp.: 102.63 Ambient Temp.: 88.00 F
01:58:56.831 -> Average Cal sensor Temp.: 101.97 Ambient Temp.: 87.96 F
01:58:58.381 -> Average Cal sensor Temp.: 102.06 Ambient Temp.: 87.93 F
01:58:59.933 -> Average Cal sensor Temp.: 101.80 Ambient Temp.: 87.89 F
01:59:01.517 -> Average Cal sensor Temp.: 102.28 Ambient Temp.: 87.85 F
01:59:03.063 -> Average Cal sensor Temp.: 101.92 Ambient Temp.: 87.82 F
01:59:04.620 -> Average Cal sensor Temp.: 102.34 Ambient Temp.: 87.85 F
01:59:06.215 -> Average Cal sensor Temp.: 102.21 Ambient Temp.: 87.71 F
01:59:07.770 -> Average Cal sensor Temp.: 102.70 Ambient Temp.: 87.78 F
01:59:09.327 -> Average Cal sensor Temp.: 102.20 Ambient Temp.: 87.78 F
01:59:10.881 -> Average Cal sensor Temp.: 102.13 Ambient Temp.: 87.71 F
01:59:12.438 -> Average Cal sensor Temp.: 102.11 Ambient Temp.: 87.71 F
01:59:14.031 -> Average Cal sensor Temp.: 102.33 Ambient Temp.: 87.75 F
01:59:15.588 -> Average Cal sensor Temp.: 102.56 Ambient Temp.: 87.78 F
01:59:17.140 -> Average Cal sensor Temp.: 102.06 Ambient Temp.: 87.71 F
01:59:18.696 -> Average Cal sensor Temp.: 102.40 Ambient Temp.: 87.75 F
01:59:20.283 -> Average Cal sensor Temp.: 102.11 Ambient Temp.: 87.75 F
01:59:21.836 -> Average Cal sensor Temp.: 102.45 Ambient Temp.: 87.71 F
01:59:23.396 -> Average Cal sensor Temp.: 101.85 Ambient Temp.: 87.67 F
01:59:24.953 -> Average Cal sensor Temp.: 102.00 Ambient Temp.: 87.60 F
01:59:26.511 -> Average Cal sensor Temp.: 102.74 Ambient Temp.: 87.60 F
01:59:28.070 -> Average Cal sensor Temp.: 102.13 Ambient Temp.: 87.67 F
01:59:29.663 -> Average Cal sensor Temp.: 102.33 Ambient Temp.: 87.67 F
01:59:31.219 -> Average Cal sensor Temp.: 102.65 Ambient Temp.: 87.71 F
01:59:32.768 -> Average Cal sensor Temp.: 102.87 Ambient Temp.: 87.67 F
01:59:34.327 -> Average Cal sensor Temp.: 101.96 Ambient Temp.: 87.67 F
01:59:35.915 -> Average Cal sensor Temp.: 102.12 Ambient Temp.: 87.67 F
01:59:37.469 -> Average Cal sensor Temp.: 102.24 Ambient Temp.: 87.60 F
01:59:39.027 -> Average Cal sensor Temp.: 102.16 Ambient Temp.: 87.67 F
01:59:40.579 -> Average Cal sensor Temp.: 102.80 Ambient Temp.: 87.60 F
01:59:42.155 -> Average Cal sensor Temp.: 102.39 Ambient Temp.: 87.60 F
01:59:43.743 -> Average Cal sensor Temp.: 101.84 Ambient Temp.: 87.60 F
01:59:45.304 -> Average Cal sensor Temp.: 101.84 Ambient Temp.: 87.53 F
01:59:46.860 -> Average Cal sensor Temp.: 102.81 Ambient Temp.: 87.60 F
01:59:48.410 -> Average Cal sensor Temp.: 102.80 Ambient Temp.: 87.57 F
01:59:49.993 -> Average Cal sensor Temp.: 102.63 Ambient Temp.: 87.57 F
01:59:51.548 -> Average Cal sensor Temp.: 102.47 Ambient Temp.: 87.60 F
01:59:53.126 -> Average Cal sensor Temp.: 101.84 Ambient Temp.: 87.49 F
01:59:54.669 -> Average Cal sensor Temp.: 101.76 Ambient Temp.: 87.46 F
01:59:56.226 -> Average Cal sensor Temp.: 102.28 Ambient Temp.: 87.46 F
01:59:57.821 -> Average Cal sensor Temp.: 102.05 Ambient Temp.: 87.42 F
01:59:59.373 -> Average Cal sensor Temp.: 101.75 Ambient Temp.: 87.35 F
02:00:00.929 -> Average Cal sensor Temp.: 102.29 Ambient Temp.: 87.39 F
02:00:02.488 -> Average Cal sensor Temp.: 102.52 Ambient Temp.: 87.39 F
02:00:04.047 -> Average Cal sensor Temp.: 101.85 Ambient Temp.: 87.31 F
02:00:05.605 -> Average Cal sensor Temp.: 101.72 Ambient Temp.: 87.28 F
02:00:07.190 -> Average Cal sensor Temp.: 101.78 Ambient Temp.: 87.28 F
02:00:08.746 -> Average Cal sensor Temp.: 101.49 Ambient Temp.: 87.13 F
02:00:10.303 -> Average Cal sensor Temp.: 102.12 Ambient Temp.: 87.17 F
02:00:11.858 -> Average Cal sensor Temp.: 101.87 Ambient Temp.: 87.13 F
02:00:13.419 -> Average Cal sensor Temp.: 101.97 Ambient Temp.: 87.06 F
02:00:15.010 -> Average Cal sensor Temp.: 102.42 Ambient Temp.: 87.06 F
02:00:16.563 -> Average Cal sensor Temp.: 102.25 Ambient Temp.: 87.03 F
02:00:18.114 -> Average Cal sensor Temp.: 102.21 Ambient Temp.: 87.03 F
02:00:19.675 -> Average Cal sensor Temp.: 101.94 Ambient Temp.: 87.03 F
02:00:21.261 -> Average Cal sensor Temp.: 102.44 Ambient Temp.: 86.99 F
02:00:22.815 -> Average Cal sensor Temp.: 102.54 Ambient Temp.: 87.03 F
02:00:24.374 -> Average Cal sensor Temp.: 102.32 Ambient Temp.: 87.06 F
02:00:25.963 -> Average Cal sensor Temp.: 102.42 Ambient Temp.: 87.06 F
02:00:27.516 -> Average Cal sensor Temp.: 102.77 Ambient Temp.: 87.03 F
02:00:29.062 -> Average Cal sensor Temp.: 102.58 Ambient Temp.: 87.06 F
02:00:30.659 -> Average Cal sensor Temp.: 102.51 Ambient Temp.: 87.06 F
02:00:32.215 -> Average Cal sensor Temp.: 102.42 Ambient Temp.: 87.10 F
02:00:33.773 -> Average Cal sensor Temp.: 102.46 Ambient Temp.: 87.10 F
02:00:35.326 -> Average Cal sensor Temp.: 102.60 Ambient Temp.: 87.06 F
02:00:36.914 -> Average Cal sensor Temp.: 102.54 Ambient Temp.: 87.13 F
02:00:38.469 -> Average Cal sensor Temp.: 102.24 Ambient Temp.: 87.13 F
02:00:40.018 -> Average Cal sensor Temp.: 102.42 Ambient Temp.: 87.10 F
02:00:41.605 -> Average Cal sensor Temp.: 103.27 Ambient Temp.: 87.13 F
02:00:43.162 -> Average Cal sensor Temp.: 102.75 Ambient Temp.: 87.13 F
02:00:44.722 -> Average Cal sensor Temp.: 102.39 Ambient Temp.: 87.17 F
02:00:46.278 -> Average Cal sensor Temp.: 103.24 Ambient Temp.: 87.21 F
02:00:47.828 -> Average Cal sensor Temp.: 103.22 Ambient Temp.: 87.28 F
02:00:49.418 -> Average Cal sensor Temp.: 103.33 Ambient Temp.: 87.17 F
02:00:50.976 -> Average Cal sensor Temp.: 102.96 Ambient Temp.: 87.28 F
02:00:52.537 -> Average Cal sensor Temp.: 103.20 Ambient Temp.: 87.28 F
02:00:54.089 -> Average Cal sensor Temp.: 103.80 Ambient Temp.: 87.35 F
02:00:55.648 -> Average Cal sensor Temp.: 103.03 Ambient Temp.: 87.39 F
02:00:57.237 -> Average Cal sensor Temp.: 102.87 Ambient Temp.: 87.46 F
02:00:58.798 -> Average Cal sensor Temp.: 102.85 Ambient Temp.: 87.42 F
02:01:00.350 -> Average Cal sensor Temp.: 102.68 Ambient Temp.: 87.42 F
02:01:01.903 -> Average Cal sensor Temp.: 102.77 Ambient Temp.: 87.39 F
02:01:03.492 -> Average Cal sensor Temp.: 102.83 Ambient Temp.: 87.42 F
02:01:05.051 -> Average Cal sensor Temp.: 102.65 Ambient Temp.: 87.42 F
02:01:06.607 -> Average Cal sensor Temp.: 103.10 Ambient Temp.: 87.46 F
02:01:08.186 -> Average Cal sensor Temp.: 102.93 Ambient Temp.: 87.42 F
02:01:09.737 -> Average Cal sensor Temp.: 103.27 Ambient Temp.: 87.42 F
02:01:11.291 -> Average Cal sensor Temp.: 103.03 Ambient Temp.: 87.42 F
02:01:12.875 -> Average Cal sensor Temp.: 103.05 Ambient Temp.: 87.49 F
02:01:14.432 -> Average Cal sensor Temp.: 102.79 Ambient Temp.: 87.49 F
02:01:15.989 -> Average Cal sensor Temp.: 103.59 Ambient Temp.: 87.49 F
02:01:17.543 -> Average Cal sensor Temp.: 104.02 Ambient Temp.: 87.53 F
02:01:19.132 -> Average Cal sensor Temp.: 102.86 Ambient Temp.: 87.57 F
02:01:20.685 -> Average Cal sensor Temp.: 101.89 Ambient Temp.: 87.53 F
02:01:22.243 -> Average Cal sensor Temp.: 102.40 Ambient Temp.: 87.53 F
02:01:23.803 -> Average Cal sensor Temp.: 102.83 Ambient Temp.: 87.57 F
02:01:25.395 -> Average Cal sensor Temp.: 103.01 Ambient Temp.: 87.53 F
02:01:26.945 -> Average Cal sensor Temp.: 102.87 Ambient Temp.: 87.53 F
02:01:28.504 -> Average Cal sensor Temp.: 102.33 Ambient Temp.: 87.53 F
02:01:30.054 -> Average Cal sensor Temp.: 102.49 Ambient Temp.: 87.53 F
02:01:31.646 -> Average Cal sensor Temp.: 102.67 Ambient Temp.: 87.49 F
02:01:33.200 -> Average Cal sensor Temp.: 102.90 Ambient Temp.: 87.53 F
02:01:34.760 -> Average Cal sensor Temp.: 102.83 Ambient Temp.: 87.57 F
02:01:36.317 -> Average Cal sensor Temp.: 102.91 Ambient Temp.: 87.53 F
02:01:37.902 -> Average Cal sensor Temp.: 103.36 Ambient Temp.: 87.57 F
02:01:39.452 -> Average Cal sensor Temp.: 103.25 Ambient Temp.: 87.57 F
02:01:41.006 -> Average Cal sensor Temp.: 103.55 Ambient Temp.: 87.57 F
02:01:42.564 -> Average Cal sensor Temp.: 103.49 Ambient Temp.: 87.60 F
02:01:44.154 -> Average Cal sensor Temp.: 103.69 Ambient Temp.: 87.57 F
02:01:45.707 -> Average Cal sensor Temp.: 103.15 Ambient Temp.: 87.67 F
02:01:47.258 -> Average Cal sensor Temp.: 103.14 Ambient Temp.: 87.60 F
02:01:48.848 -> Average Cal sensor Temp.: 103.77 Ambient Temp.: 87.67 F
02:01:50.405 -> Average Cal sensor Temp.: 103.23 Ambient Temp.: 87.67 F
02:01:51.959 -> Average Cal sensor Temp.: 102.94 Ambient Temp.: 87.71 F
02:01:53.515 -> Average Cal sensor Temp.: 103.15 Ambient Temp.: 87.71 F
02:01:55.068 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 87.75 F
02:01:56.661 -> Average Cal sensor Temp.: 103.19 Ambient Temp.: 87.75 F
02:01:58.215 -> Average Cal sensor Temp.: 103.19 Ambient Temp.: 87.75 F
02:01:59.770 -> Average Cal sensor Temp.: 102.75 Ambient Temp.: 87.75 F
02:02:01.321 -> Average Cal sensor Temp.: 103.13 Ambient Temp.: 87.75 F
02:02:02.912 -> Average Cal sensor Temp.: 103.30 Ambient Temp.: 87.82 F
02:02:04.465 -> Average Cal sensor Temp.: 103.01 Ambient Temp.: 87.78 F
02:02:06.019 -> Average Cal sensor Temp.: 103.01 Ambient Temp.: 87.78 F
02:02:07.574 -> Average Cal sensor Temp.: 103.00 Ambient Temp.: 87.85 F
02:02:09.164 -> Average Cal sensor Temp.: 102.66 Ambient Temp.: 87.75 F
02:02:10.718 -> Average Cal sensor Temp.: 102.83 Ambient Temp.: 87.82 F
02:02:12.267 -> Average Cal sensor Temp.: 102.96 Ambient Temp.: 87.78 F
02:02:13.851 -> Average Cal sensor Temp.: 103.03 Ambient Temp.: 87.78 F
02:02:15.395 -> Average Cal sensor Temp.: 103.43 Ambient Temp.: 87.78 F
02:02:16.983 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 87.82 F
02:02:18.537 -> Average Cal sensor Temp.: 103.33 Ambient Temp.: 87.82 F
02:02:20.089 -> Average Cal sensor Temp.: 103.10 Ambient Temp.: 87.82 F
02:02:21.681 -> Average Cal sensor Temp.: 102.97 Ambient Temp.: 87.85 F
02:02:23.243 -> Average Cal sensor Temp.: 102.92 Ambient Temp.: 87.85 F
02:02:24.798 -> Average Cal sensor Temp.: 103.06 Ambient Temp.: 87.85 F
02:02:26.352 -> Average Cal sensor Temp.: 103.53 Ambient Temp.: 87.89 F
02:02:27.939 -> Average Cal sensor Temp.: 103.59 Ambient Temp.: 87.89 F
02:02:29.497 -> Average Cal sensor Temp.: 103.21 Ambient Temp.: 87.93 F
02:02:31.046 -> Average Cal sensor Temp.: 103.03 Ambient Temp.: 87.93 F
02:02:32.634 -> Average Cal sensor Temp.: 103.10 Ambient Temp.: 87.85 F
02:02:34.188 -> Average Cal sensor Temp.: 103.27 Ambient Temp.: 87.89 F
02:02:35.741 -> Average Cal sensor Temp.: 103.09 Ambient Temp.: 87.89 F
02:02:37.329 -> Average Cal sensor Temp.: 103.51 Ambient Temp.: 87.93 F
02:02:38.882 -> Average Cal sensor Temp.: 103.91 Ambient Temp.: 87.96 F
02:02:40.436 -> Average Cal sensor Temp.: 103.33 Ambient Temp.: 87.96 F
02:02:41.990 -> Average Cal sensor Temp.: 103.68 Ambient Temp.: 87.96 F
02:02:43.575 -> Average Cal sensor Temp.: 103.16 Ambient Temp.: 87.96 F
02:02:45.133 -> Average Cal sensor Temp.: 103.32 Ambient Temp.: 88.00 F
02:02:46.687 -> Average Cal sensor Temp.: 103.65 Ambient Temp.: 88.11 F
02:02:48.275 -> Average Cal sensor Temp.: 102.82 Ambient Temp.: 88.07 F
02:02:49.827 -> Average Cal sensor Temp.: 102.92 Ambient Temp.: 88.00 F
02:02:51.379 -> Average Cal sensor Temp.: 103.97 Ambient Temp.: 88.07 F
02:02:52.961 -> Average Cal sensor Temp.: 102.73 Ambient Temp.: 88.07 F
02:02:54.515 -> Average Cal sensor Temp.: 102.91 Ambient Temp.: 88.00 F
02:02:56.102 -> Average Cal sensor Temp.: 103.38 Ambient Temp.: 88.00 F
02:02:57.659 -> Average Cal sensor Temp.: 103.22 Ambient Temp.: 88.00 F
02:02:59.211 -> Average Cal sensor Temp.: 103.54 Ambient Temp.: 88.11 F
02:03:00.796 -> Average Cal sensor Temp.: 103.59 Ambient Temp.: 88.11 F
02:03:02.346 -> Average Cal sensor Temp.: 103.23 Ambient Temp.: 88.07 F
02:03:03.904 -> Average Cal sensor Temp.: 103.57 Ambient Temp.: 88.11 F
02:03:05.460 -> Average Cal sensor Temp.: 103.08 Ambient Temp.: 88.11 F
02:03:07.045 -> Average Cal sensor Temp.: 103.14 Ambient Temp.: 88.11 F
02:03:08.602 -> Average Cal sensor Temp.: 103.25 Ambient Temp.: 88.11 F
02:03:10.160 -> Average Cal sensor Temp.: 103.22 Ambient Temp.: 88.11 F
02:03:11.716 -> Average Cal sensor Temp.: 103.46 Ambient Temp.: 88.11 F
02:03:13.276 -> Average Cal sensor Temp.: 103.55 Ambient Temp.: 88.11 F
02:03:14.862 -> Average Cal sensor Temp.: 103.59 Ambient Temp.: 88.11 F
02:03:16.420 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 88.11 F
02:03:17.971 -> Average Cal sensor Temp.: 103.67 Ambient Temp.: 88.21 F
02:03:19.533 -> Average Cal sensor Temp.: 103.03 Ambient Temp.: 88.21 F
02:03:21.088 -> Average Cal sensor Temp.: 102.39 Ambient Temp.: 88.14 F
02:03:22.673 -> Average Cal sensor Temp.: 103.30 Ambient Temp.: 88.18 F
02:03:24.225 -> Average Cal sensor Temp.: 103.43 Ambient Temp.: 88.18 F
02:03:25.812 -> Average Cal sensor Temp.: 103.48 Ambient Temp.: 88.25 F
02:03:27.364 -> Average Cal sensor Temp.: 102.99 Ambient Temp.: 88.18 F
02:03:28.913 -> Average Cal sensor Temp.: 103.03 Ambient Temp.: 88.21 F
02:03:30.508 -> Average Cal sensor Temp.: 102.67 Ambient Temp.: 88.18 F
02:03:32.062 -> Average Cal sensor Temp.: 102.60 Ambient Temp.: 88.21 F
02:03:33.614 -> Average Cal sensor Temp.: 102.58 Ambient Temp.: 88.18 F
02:03:35.169 -> Average Cal sensor Temp.: 102.56 Ambient Temp.: 88.18 F
02:03:36.752 -> Average Cal sensor Temp.: 102.22 Ambient Temp.: 88.14 F
02:03:38.303 -> Average Cal sensor Temp.: 101.92 Ambient Temp.: 88.11 F
02:03:39.888 -> Average Cal sensor Temp.: 102.12 Ambient Temp.: 88.07 F
02:03:41.440 -> Average Cal sensor Temp.: 102.65 Ambient Temp.: 88.00 F
02:03:43.021 -> Average Cal sensor Temp.: 102.77 Ambient Temp.: 88.00 F
02:03:44.576 -> Average Cal sensor Temp.: 102.83 Ambient Temp.: 87.93 F
02:03:46.128 -> Average Cal sensor Temp.: 102.95 Ambient Temp.: 88.00 F
02:03:47.710 -> Average Cal sensor Temp.: 103.04 Ambient Temp.: 87.96 F
02:03:49.264 -> Average Cal sensor Temp.: 103.15 Ambient Temp.: 88.00 F
02:03:50.848 -> Average Cal sensor Temp.: 102.62 Ambient Temp.: 88.00 F
02:03:52.402 -> Average Cal sensor Temp.: 103.10 Ambient Temp.: 88.00 F
02:03:53.962 -> Average Cal sensor Temp.: 103.48 Ambient Temp.: 88.00 F
02:03:55.515 -> Average Cal sensor Temp.: 103.21 Ambient Temp.: 88.00 F
02:03:57.105 -> Average Cal sensor Temp.: 103.31 Ambient Temp.: 88.07 F
02:03:58.666 -> Average Cal sensor Temp.: 103.70 Ambient Temp.: 88.11 F
02:04:00.211 -> Average Cal sensor Temp.: 103.62 Ambient Temp.: 88.14 F
02:04:01.800 -> Average Cal sensor Temp.: 103.24 Ambient Temp.: 88.14 F
02:04:03.353 -> Average Cal sensor Temp.: 103.70 Ambient Temp.: 88.18 F
02:04:04.911 -> Average Cal sensor Temp.: 103.16 Ambient Temp.: 88.18 F
02:04:06.467 -> Average Cal sensor Temp.: 102.82 Ambient Temp.: 88.18 F
02:04:08.028 -> Average Cal sensor Temp.: 102.99 Ambient Temp.: 88.21 F
02:04:09.618 -> Average Cal sensor Temp.: 103.41 Ambient Temp.: 88.21 F
02:04:11.169 -> Average Cal sensor Temp.: 102.06 Ambient Temp.: 88.18 F
02:04:12.722 -> Average Cal sensor Temp.: 102.27 Ambient Temp.: 88.14 F
02:04:14.313 -> Average Cal sensor Temp.: 102.84 Ambient Temp.: 88.21 F
02:04:15.874 -> Average Cal sensor Temp.: 103.14 Ambient Temp.: 88.18 F
02:04:17.431 -> Average Cal sensor Temp.: 103.20 Ambient Temp.: 88.25 F
02:04:18.988 -> Average Cal sensor Temp.: 102.79 Ambient Temp.: 88.25 F
02:04:20.547 -> Average Cal sensor Temp.: 103.15 Ambient Temp.: 88.25 F
02:04:22.104 -> Average Cal sensor Temp.: 102.47 Ambient Temp.: 88.21 F
02:04:23.697 -> Average Cal sensor Temp.: 102.38 Ambient Temp.: 88.21 F
02:04:25.261 -> Average Cal sensor Temp.: 102.68 Ambient Temp.: 88.25 F
02:04:26.814 -> Average Cal sensor Temp.: 102.72 Ambient Temp.: 88.18 F
02:04:28.376 -> Average Cal sensor Temp.: 102.50 Ambient Temp.: 88.21 F
02:04:29.933 -> Average Cal sensor Temp.: 102.97 Ambient Temp.: 88.18 F
02:04:31.493 -> Average Cal sensor Temp.: 103.41 Ambient Temp.: 88.25 F
02:04:33.084 -> Average Cal sensor Temp.: 103.46 Ambient Temp.: 88.25 F
02:04:34.642 -> Average Cal sensor Temp.: 103.44 Ambient Temp.: 88.29 F
02:04:36.198 -> Average Cal sensor Temp.: 102.89 Ambient Temp.: 88.29 F
02:04:37.752 -> Average Cal sensor Temp.: 102.78 Ambient Temp.: 88.32 F
02:04:39.342 -> Average Cal sensor Temp.: 102.82 Ambient Temp.: 88.25 F
02:04:40.897 -> Average Cal sensor Temp.: 102.82 Ambient Temp.: 88.25 F
02:04:42.456 -> Average Cal sensor Temp.: 103.12 Ambient Temp.: 88.32 F
02:04:44.007 -> Average Cal sensor Temp.: 103.06 Ambient Temp.: 88.25 F
02:04:45.589 -> Average Cal sensor Temp.: 103.08 Ambient Temp.: 88.25 F
02:04:47.137 -> Average Cal sensor Temp.: 102.61 Ambient Temp.: 88.29 F
02:04:48.726 -> Average Cal sensor Temp.: 102.45 Ambient Temp.: 88.32 F
02:04:50.281 -> Average Cal sensor Temp.: 102.75 Ambient Temp.: 88.29 F
02:04:51.839 -> Average Cal sensor Temp.: 102.76 Ambient Temp.: 88.29 F
02:04:53.395 -> Average Cal sensor Temp.: 103.68 Ambient Temp.: 88.25 F
02:04:54.983 -> Average Cal sensor Temp.: 103.21 Ambient Temp.: 88.29 F
02:04:56.537 -> Average Cal sensor Temp.: 104.50 Ambient Temp.: 88.29 F
02:04:58.093 -> Average Cal sensor Temp.: 104.30 Ambient Temp.: 88.36 F
02:04:59.682 -> Average Cal sensor Temp.: 103.46 Ambient Temp.: 88.39 F
02:05:01.235 -> Average Cal sensor Temp.: 103.08 Ambient Temp.: 88.36 F
02:05:02.791 -> Average Cal sensor Temp.: 102.90 Ambient Temp.: 88.39 F
02:05:04.383 -> Average Cal sensor Temp.: 103.65 Ambient Temp.: 88.39 F
02:05:05.936 -> Average Cal sensor Temp.: 102.96 Ambient Temp.: 88.39 F
02:05:07.492 -> Average Cal sensor Temp.: 102.92 Ambient Temp.: 88.39 F
02:05:09.082 -> Average Cal sensor Temp.: 102.69 Ambient Temp.: 88.39 F
02:05:10.637 -> Average Cal sensor Temp.: 102.33 Ambient Temp.: 88.36 F
02:05:12.192 -> Average Cal sensor Temp.: 102.69 Ambient Temp.: 88.36 F
02:05:13.746 -> Average Cal sensor Temp.: 102.88 Ambient Temp.: 88.29 F
02:05:15.338 -> Average Cal sensor Temp.: 103.40 Ambient Temp.: 88.36 F
02:05:16.892 -> Average Cal sensor Temp.: 102.97 Ambient Temp.: 88.32 F
02:05:18.450 -> Average Cal sensor Temp.: 103.32 Ambient Temp.: 88.29 F
02:05:20.009 -> Average Cal sensor Temp.: 103.63 Ambient Temp.: 88.32 F
02:05:21.604 -> Average Cal sensor Temp.: 103.23 Ambient Temp.: 88.29 F
02:05:23.160 -> Average Cal sensor Temp.: 103.01 Ambient Temp.: 88.29 F
02:05:24.716 -> Average Cal sensor Temp.: 102.53 Ambient Temp.: 88.21 F
02:05:26.270 -> Average Cal sensor Temp.: 102.17 Ambient Temp.: 88.25 F
02:05:27.827 -> Average Cal sensor Temp.: 101.99 Ambient Temp.: 88.14 F
02:05:29.420 -> Average Cal sensor Temp.: 102.30 Ambient Temp.: 88.11 F
02:05:30.979 -> Average Cal sensor Temp.: 103.01 Ambient Temp.: 88.11 F
02:05:32.535 -> Average Cal sensor Temp.: 103.02 Ambient Temp.: 88.14 F
02:05:34.095 -> Average Cal sensor Temp.: 103.23 Ambient Temp.: 88.07 F
02:05:35.670 -> Average Cal sensor Temp.: 103.14 Ambient Temp.: 88.11 F
02:05:37.215 -> Average Cal sensor Temp.: 103.18 Ambient Temp.: 88.07 F
02:05:38.805 -> Average Cal sensor Temp.: 103.66 Ambient Temp.: 88.07 F
02:05:40.360 -> Average Cal sensor Temp.: 103.64 Ambient Temp.: 88.11 F
02:05:41.917 -> Average Cal sensor Temp.: 103.65 Ambient Temp.: 88.07 F
02:05:43.474 -> Average Cal sensor Temp.: 103.41 Ambient Temp.: 88.11 F
02:05:45.030 -> Average Cal sensor Temp.: 103.44 Ambient Temp.: 88.14 F
02:05:46.619 -> Average Cal sensor Temp.: 103.60 Ambient Temp.: 88.11 F
02:05:48.173 -> Average Cal sensor Temp.: 103.19 Ambient Temp.: 88.11 F
02:05:49.729 -> Average Cal sensor Temp.: 103.36 Ambient Temp.: 88.14 F
02:05:51.318 -> Average Cal sensor Temp.: 103.33 Ambient Temp.: 88.14 F
02:05:52.865 -> Average Cal sensor Temp.: 103.11 Ambient Temp.: 88.14 F
02:05:54.417 -> Average Cal sensor Temp.: 102.92 Ambient Temp.: 88.14 F
02:05:56.011 -> Average Cal sensor Temp.: 103.28 Ambient Temp.: 88.14 F
02:05:57.564 -> Average Cal sensor Temp.: 103.19 Ambient Temp.: 88.14 F
02:05:59.124 -> Average Cal sensor Temp.: 103.05 Ambient Temp.: 88.14 F
02:06:00.676 -> Average Cal sensor Temp.: 102.99 Ambient Temp.: 88.00 F
02:06:02.269 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 88.14 F
02:06:03.828 -> Average Cal sensor Temp.: 103.52 Ambient Temp.: 88.11 F
02:06:05.384 -> Average Cal sensor Temp.: 103.60 Ambient Temp.: 88.07 F
02:06:06.943 -> Average Cal sensor Temp.: 104.12 Ambient Temp.: 88.00 F
02:06:08.497 -> Average Cal sensor Temp.: 103.28 Ambient Temp.: 88.07 F
02:06:10.081 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 88.07 F
02:06:11.632 -> Average Cal sensor Temp.: 103.58 Ambient Temp.: 88.07 F
02:06:13.218 -> Average Cal sensor Temp.: 103.83 Ambient Temp.: 88.07 F
02:06:14.772 -> Average Cal sensor Temp.: 103.21 Ambient Temp.: 88.11 F
02:06:16.325 -> Average Cal sensor Temp.: 102.82 Ambient Temp.: 88.00 F
02:06:17.907 -> Average Cal sensor Temp.: 102.76 Ambient Temp.: 88.00 F
02:06:19.461 -> Average Cal sensor Temp.: 103.79 Ambient Temp.: 88.00 F
02:06:21.018 -> Average Cal sensor Temp.: 103.32 Ambient Temp.: 88.07 F
02:06:22.578 -> Average Cal sensor Temp.: 103.38 Ambient Temp.: 88.00 F
02:06:24.169 -> Average Cal sensor Temp.: 103.75 Ambient Temp.: 88.07 F
02:06:25.721 -> Average Cal sensor Temp.: 104.08 Ambient Temp.: 88.00 F
02:06:27.272 -> Average Cal sensor Temp.: 103.95 Ambient Temp.: 87.96 F
02:06:28.863 -> Average Cal sensor Temp.: 103.71 Ambient Temp.: 88.00 F
02:06:30.412 -> Average Cal sensor Temp.: 103.53 Ambient Temp.: 88.07 F
02:06:31.967 -> Average Cal sensor Temp.: 103.99 Ambient Temp.: 88.07 F
02:06:33.556 -> Average Cal sensor Temp.: 103.96 Ambient Temp.: 88.11 F
02:06:35.114 -> Average Cal sensor Temp.: 104.02 Ambient Temp.: 88.00 F
02:06:36.667 -> Average Cal sensor Temp.: 104.62 Ambient Temp.: 88.07 F
02:06:38.254 -> Average Cal sensor Temp.: 103.72 Ambient Temp.: 88.14 F
02:06:39.805 -> Average Cal sensor Temp.: 103.55 Ambient Temp.: 88.11 F
02:06:41.354 -> Average Cal sensor Temp.: 103.68 Ambient Temp.: 88.14 F
02:06:42.942 -> Average Cal sensor Temp.: 102.89 Ambient Temp.: 88.07 F
02:06:44.500 -> Average Cal sensor Temp.: 102.60 Ambient Temp.: 88.07 F
02:06:46.083 -> Average Cal sensor Temp.: 102.47 Ambient Temp.: 87.96 F
02:06:47.636 -> Average Cal sensor Temp.: 103.10 Ambient Temp.: 88.00 F
02:06:49.192 -> Average Cal sensor Temp.: 103.77 Ambient Temp.: 88.07 F
02:06:50.776 -> Average Cal sensor Temp.: 103.68 Ambient Temp.: 88.07 F
02:06:52.326 -> Average Cal sensor Temp.: 104.17 Ambient Temp.: 88.07 F
02:06:53.885 -> Average Cal sensor Temp.: 103.98 Ambient Temp.: 88.14 F
02:06:55.442 -> Average Cal sensor Temp.: 102.57 Ambient Temp.: 88.11 F
02:06:57.029 -> Average Cal sensor Temp.: 103.09 Ambient Temp.: 88.11 F
02:06:58.581 -> Average Cal sensor Temp.: 103.97 Ambient Temp.: 88.11 F
02:07:00.137 -> Average Cal sensor Temp.: 103.56 Ambient Temp.: 88.11 F
02:07:01.692 -> Average Cal sensor Temp.: 103.60 Ambient Temp.: 88.07 F
02:07:03.283 -> Average Cal sensor Temp.: 103.92 Ambient Temp.: 88.14 F
02:07:04.839 -> Average Cal sensor Temp.: 103.67 Ambient Temp.: 88.07 F
02:07:06.393 -> Average Cal sensor Temp.: 104.35 Ambient Temp.: 88.14 F
02:07:07.978 -> Average Cal sensor Temp.: 104.29 Ambient Temp.: 88.18 F
02:07:09.534 -> Average Cal sensor Temp.: 104.30 Ambient Temp.: 88.21 F
02:07:11.091 -> Average Cal sensor Temp.: 103.59 Ambient Temp.: 88.21 F
02:07:12.647 -> Average Cal sensor Temp.: 103.67 Ambient Temp.: 88.29 F
02:07:14.244 -> Average Cal sensor Temp.: 103.99 Ambient Temp.: 88.29 F
02:07:15.802 -> Average Cal sensor Temp.: 103.72 Ambient Temp.: 88.29 F
02:07:17.361 -> Average Cal sensor Temp.: 103.49 Ambient Temp.: 88.29 F
02:07:18.919 -> Average Cal sensor Temp.: 103.60 Ambient Temp.: 88.36 F
02:07:20.479 -> Average Cal sensor Temp.: 103.90 Ambient Temp.: 88.39 F
02:07:22.071 -> Average Cal sensor Temp.: 103.29 Ambient Temp.: 88.47 F
02:07:23.631 -> Average Cal sensor Temp.: 103.59 Ambient Temp.: 88.39 F
02:07:25.180 -> Average Cal sensor Temp.: 103.74 Ambient Temp.: 88.47 F
02:07:26.733 -> Average Cal sensor Temp.: 104.04 Ambient Temp.: 88.39 F
02:07:28.325 -> Average Cal sensor Temp.: 103.58 Ambient Temp.: 88.50 F
02:07:29.882 -> Average Cal sensor Temp.: 103.37 Ambient Temp.: 88.50 F
02:07:31.438 -> Average Cal sensor Temp.: 103.77 Ambient Temp.: 88.54 F
02:07:33.022 -> Average Cal sensor Temp.: 103.96 Ambient Temp.: 88.54 F
02:07:34.583 -> Average Cal sensor Temp.: 104.09 Ambient Temp.: 88.54 F
02:07:36.144 -> Average Cal sensor Temp.: 104.16 Ambient Temp.: 88.50 F
02:07:37.696 -> Average Cal sensor Temp.: 103.95 Ambient Temp.: 88.61 F
02:07:39.254 -> Average Cal sensor Temp.: 103.90 Ambient Temp.: 88.65 F
02:07:40.843 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 88.65 F
02:07:42.402 -> Average Cal sensor Temp.: 103.79 Ambient Temp.: 88.68 F
02:07:43.965 -> Average Cal sensor Temp.: 103.79 Ambient Temp.: 88.61 F
02:07:45.527 -> Average Cal sensor Temp.: 103.67 Ambient Temp.: 88.75 F
02:07:47.083 -> Average Cal sensor Temp.: 103.43 Ambient Temp.: 88.72 F
02:07:48.674 -> Average Cal sensor Temp.: 103.32 Ambient Temp.: 88.75 F
02:07:50.232 -> Average Cal sensor Temp.: 103.69 Ambient Temp.: 88.75 F
02:07:51.790 -> Average Cal sensor Temp.: 103.54 Ambient Temp.: 88.79 F
02:07:53.345 -> Average Cal sensor Temp.: 103.41 Ambient Temp.: 88.86 F
02:07:54.900 -> Average Cal sensor Temp.: 103.18 Ambient Temp.: 88.86 F
02:07:56.493 -> Average Cal sensor Temp.: 103.43 Ambient Temp.: 88.86 F
02:07:58.049 -> Average Cal sensor Temp.: 103.79 Ambient Temp.: 88.90 F
02:07:59.607 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 88.86 F
02:08:01.167 -> Average Cal sensor Temp.: 103.58 Ambient Temp.: 88.86 F
02:08:02.755 -> Average Cal sensor Temp.: 103.38 Ambient Temp.: 88.97 F
02:08:04.312 -> Average Cal sensor Temp.: 103.40 Ambient Temp.: 88.93 F
02:08:05.866 -> Average Cal sensor Temp.: 104.36 Ambient Temp.: 88.90 F
02:08:07.424 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 88.97 F
02:08:09.005 -> Average Cal sensor Temp.: 103.45 Ambient Temp.: 88.97 F
02:08:10.564 -> Average Cal sensor Temp.: 103.64 Ambient Temp.: 89.01 F
02:08:12.118 -> Average Cal sensor Temp.: 103.88 Ambient Temp.: 89.01 F
02:08:13.713 -> Average Cal sensor Temp.: 103.95 Ambient Temp.: 89.01 F
02:08:15.269 -> Average Cal sensor Temp.: 104.31 Ambient Temp.: 89.08 F
02:08:16.826 -> Average Cal sensor Temp.: 104.18 Ambient Temp.: 89.08 F
02:08:18.410 -> Average Cal sensor Temp.: 103.82 Ambient Temp.: 89.08 F
02:08:19.968 -> Average Cal sensor Temp.: 103.82 Ambient Temp.: 89.08 F
02:08:21.525 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 89.11 F
02:08:23.080 -> Average Cal sensor Temp.: 103.72 Ambient Temp.: 89.11 F
02:08:24.668 -> Average Cal sensor Temp.: 103.70 Ambient Temp.: 89.11 F
02:08:26.221 -> Average Cal sensor Temp.: 103.75 Ambient Temp.: 89.11 F
02:08:27.776 -> Average Cal sensor Temp.: 104.05 Ambient Temp.: 89.11 F
02:08:29.332 -> Average Cal sensor Temp.: 104.14 Ambient Temp.: 89.19 F
02:08:30.922 -> Average Cal sensor Temp.: 104.13 Ambient Temp.: 89.19 F
02:08:32.477 -> Average Cal sensor Temp.: 104.13 Ambient Temp.: 89.19 F
02:08:34.030 -> Average Cal sensor Temp.: 103.87 Ambient Temp.: 89.19 F
02:08:35.586 -> Average Cal sensor Temp.: 103.32 Ambient Temp.: 89.26 F
02:08:37.176 -> Average Cal sensor Temp.: 103.55 Ambient Temp.: 89.19 F
02:08:38.732 -> Average Cal sensor Temp.: 103.67 Ambient Temp.: 89.19 F
02:08:40.290 -> Average Cal sensor Temp.: 103.71 Ambient Temp.: 89.19 F
02:08:41.843 -> Average Cal sensor Temp.: 103.67 Ambient Temp.: 89.26 F
02:08:43.432 -> Average Cal sensor Temp.: 103.78 Ambient Temp.: 89.26 F
02:08:44.990 -> Average Cal sensor Temp.: 103.72 Ambient Temp.: 89.26 F
02:08:46.549 -> Average Cal sensor Temp.: 103.82 Ambient Temp.: 89.26 F
02:08:48.109 -> Average Cal sensor Temp.: 103.80 Ambient Temp.: 89.29 F
02:08:49.698 -> Average Cal sensor Temp.: 104.28 Ambient Temp.: 89.29 F
02:08:51.259 -> Average Cal sensor Temp.: 103.84 Ambient Temp.: 89.29 F
02:08:52.812 -> Average Cal sensor Temp.: 103.84 Ambient Temp.: 89.26 F
02:08:54.370 -> Average Cal sensor Temp.: 104.12 Ambient Temp.: 89.29 F
02:08:55.955 -> Average Cal sensor Temp.: 104.35 Ambient Temp.: 89.29 F
02:08:57.506 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 89.29 F
02:08:59.058 -> Average Cal sensor Temp.: 103.77 Ambient Temp.: 89.26 F
02:09:00.641 -> Average Cal sensor Temp.: 103.74 Ambient Temp.: 89.29 F
02:09:02.194 -> Average Cal sensor Temp.: 103.93 Ambient Temp.: 89.26 F
02:09:03.782 -> Average Cal sensor Temp.: 104.27 Ambient Temp.: 89.26 F
02:09:05.338 -> Average Cal sensor Temp.: 104.32 Ambient Temp.: 89.19 F
02:09:06.889 -> Average Cal sensor Temp.: 104.46 Ambient Temp.: 89.29 F
02:09:08.476 -> Average Cal sensor Temp.: 104.74 Ambient Temp.: 89.29 F
02:09:10.030 -> Average Cal sensor Temp.: 103.82 Ambient Temp.: 89.29 F
02:09:11.586 -> Average Cal sensor Temp.: 103.77 Ambient Temp.: 89.29 F
02:09:13.144 -> Average Cal sensor Temp.: 103.62 Ambient Temp.: 89.37 F
02:09:14.739 -> Average Cal sensor Temp.: 103.32 Ambient Temp.: 89.26 F
02:09:16.290 -> Average Cal sensor Temp.: 103.41 Ambient Temp.: 89.33 F
02:09:17.845 -> Average Cal sensor Temp.: 103.44 Ambient Temp.: 89.33 F
02:09:19.402 -> Average Cal sensor Temp.: 103.55 Ambient Temp.: 89.26 F
02:09:20.996 -> Average Cal sensor Temp.: 103.72 Ambient Temp.: 89.33 F
02:09:22.550 -> Average Cal sensor Temp.: 103.51 Ambient Temp.: 89.26 F
02:09:24.107 -> Average Cal sensor Temp.: 103.83 Ambient Temp.: 89.19 F
02:09:25.667 -> Average Cal sensor Temp.: 103.80 Ambient Temp.: 89.26 F
02:09:27.221 -> Average Cal sensor Temp.: 104.05 Ambient Temp.: 89.33 F
02:09:28.811 -> Average Cal sensor Temp.: 104.22 Ambient Temp.: 89.29 F
02:09:30.370 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 89.26 F
02:09:31.924 -> Average Cal sensor Temp.: 103.40 Ambient Temp.: 89.19 F
02:09:33.482 -> Average Cal sensor Temp.: 103.77 Ambient Temp.: 89.26 F
02:09:35.077 -> Average Cal sensor Temp.: 103.76 Ambient Temp.: 89.29 F
02:09:36.632 -> Average Cal sensor Temp.: 103.71 Ambient Temp.: 89.26 F
02:09:38.188 -> Average Cal sensor Temp.: 103.62 Ambient Temp.: 89.26 F
02:09:39.742 -> Average Cal sensor Temp.: 103.76 Ambient Temp.: 89.19 F
02:09:41.303 -> Average Cal sensor Temp.: 103.83 Ambient Temp.: 89.15 F
02:09:42.895 -> Average Cal sensor Temp.: 103.72 Ambient Temp.: 89.11 F
02:09:44.456 -> Average Cal sensor Temp.: 103.64 Ambient Temp.: 89.19 F
02:09:46.017 -> Average Cal sensor Temp.: 103.74 Ambient Temp.: 89.15 F
02:09:47.575 -> Average Cal sensor Temp.: 104.17 Ambient Temp.: 89.19 F
02:09:49.134 -> Average Cal sensor Temp.: 103.64 Ambient Temp.: 89.15 F
02:09:50.688 -> Average Cal sensor Temp.: 103.40 Ambient Temp.: 89.15 F
02:09:52.279 -> Average Cal sensor Temp.: 103.13 Ambient Temp.: 89.08 F
02:09:53.840 -> Average Cal sensor Temp.: 102.96 Ambient Temp.: 89.04 F
02:09:55.398 -> Average Cal sensor Temp.: 103.47 Ambient Temp.: 89.01 F
02:09:56.955 -> Average Cal sensor Temp.: 102.91 Ambient Temp.: 89.01 F
02:09:58.518 -> Average Cal sensor Temp.: 103.17 Ambient Temp.: 89.01 F
02:10:00.071 -> Average Cal sensor Temp.: 103.47 Ambient Temp.: 89.01 F
02:10:01.662 -> Average Cal sensor Temp.: 103.64 Ambient Temp.: 88.93 F
02:10:03.221 -> Average Cal sensor Temp.: 103.65 Ambient Temp.: 88.93 F
02:10:04.780 -> Average Cal sensor Temp.: 103.96 Ambient Temp.: 88.97 F
02:10:06.335 -> Average Cal sensor Temp.: 103.93 Ambient Temp.: 88.93 F
02:10:07.920 -> Average Cal sensor Temp.: 103.04 Ambient Temp.: 88.93 F
02:10:09.476 -> Average Cal sensor Temp.: 103.23 Ambient Temp.: 88.93 F
02:10:11.030 -> Average Cal sensor Temp.: 103.57 Ambient Temp.: 88.90 F
02:10:12.619 -> Average Cal sensor Temp.: 104.03 Ambient Temp.: 88.93 F
02:10:14.182 -> Average Cal sensor Temp.: 103.16 Ambient Temp.: 88.86 F
02:10:15.741 -> Average Cal sensor Temp.: 103.52 Ambient Temp.: 88.86 F
02:10:17.297 -> Average Cal sensor Temp.: 103.60 Ambient Temp.: 88.79 F
02:10:18.885 -> Average Cal sensor Temp.: 103.97 Ambient Temp.: 88.86 F
02:10:20.444 -> Average Cal sensor Temp.: 103.96 Ambient Temp.: 88.75 F
02:10:22.002 -> Average Cal sensor Temp.: 104.03 Ambient Temp.: 88.79 F
02:10:23.555 -> Average Cal sensor Temp.: 103.97 Ambient Temp.: 88.79 F
02:10:25.148 -> Average Cal sensor Temp.: 103.64 Ambient Temp.: 88.79 F
02:10:26.704 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 88.75 F
02:10:28.259 -> Average Cal sensor Temp.: 103.67 Ambient Temp.: 88.72 F
02:10:29.812 -> Average Cal sensor Temp.: 103.17 Ambient Temp.: 88.72 F
02:10:31.403 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 88.75 F
02:10:32.959 -> Average Cal sensor Temp.: 102.97 Ambient Temp.: 88.68 F
02:10:34.515 -> Average Cal sensor Temp.: 103.59 Ambient Temp.: 88.68 F
02:10:36.073 -> Average Cal sensor Temp.: 103.55 Ambient Temp.: 88.65 F
02:10:37.629 -> Average Cal sensor Temp.: 103.02 Ambient Temp.: 88.65 F
02:10:39.218 -> Average Cal sensor Temp.: 103.64 Ambient Temp.: 88.65 F
02:10:40.772 -> Average Cal sensor Temp.: 103.55 Ambient Temp.: 88.61 F
02:10:42.328 -> Average Cal sensor Temp.: 103.10 Ambient Temp.: 88.61 F
02:10:43.919 -> Average Cal sensor Temp.: 103.32 Ambient Temp.: 88.57 F
02:10:45.477 -> Average Cal sensor Temp.: 103.69 Ambient Temp.: 88.50 F
02:10:47.033 -> Average Cal sensor Temp.: 103.60 Ambient Temp.: 88.54 F
02:10:48.591 -> Average Cal sensor Temp.: 104.01 Ambient Temp.: 88.50 F
02:10:50.180 -> Average Cal sensor Temp.: 104.42 Ambient Temp.: 88.54 F
02:10:51.733 -> Average Cal sensor Temp.: 104.15 Ambient Temp.: 88.50 F
02:10:53.295 -> Average Cal sensor Temp.: 104.09 Ambient Temp.: 88.50 F
02:10:54.849 -> Average Cal sensor Temp.: 103.67 Ambient Temp.: 88.54 F
02:10:56.439 -> Average Cal sensor Temp.: 103.95 Ambient Temp.: 88.50 F
02:10:57.996 -> Average Cal sensor Temp.: 104.20 Ambient Temp.: 88.54 F
02:10:59.550 -> Average Cal sensor Temp.: 103.19 Ambient Temp.: 88.50 F
02:11:01.102 -> Average Cal sensor Temp.: 103.58 Ambient Temp.: 88.47 F
02:11:02.692 -> Average Cal sensor Temp.: 103.39 Ambient Temp.: 88.50 F
02:11:04.254 -> Average Cal sensor Temp.: 103.69 Ambient Temp.: 88.47 F
02:11:05.806 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 88.36 F
02:11:07.395 -> Average Cal sensor Temp.: 103.87 Ambient Temp.: 88.36 F
02:11:08.946 -> Average Cal sensor Temp.: 103.72 Ambient Temp.: 88.36 F
02:11:10.497 -> Average Cal sensor Temp.: 104.16 Ambient Temp.: 88.32 F
02:11:12.086 -> Average Cal sensor Temp.: 104.02 Ambient Temp.: 88.39 F
02:11:13.645 -> Average Cal sensor Temp.: 104.10 Ambient Temp.: 88.39 F
02:11:15.207 -> Average Cal sensor Temp.: 103.68 Ambient Temp.: 88.32 F
02:11:16.762 -> Average Cal sensor Temp.: 103.45 Ambient Temp.: 88.36 F
02:11:18.345 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 88.36 F
02:11:19.903 -> Average Cal sensor Temp.: 103.46 Ambient Temp.: 88.32 F
02:11:21.464 -> Average Cal sensor Temp.: 103.83 Ambient Temp.: 88.32 F
02:11:23.015 -> Average Cal sensor Temp.: 103.93 Ambient Temp.: 88.36 F
02:11:24.601 -> Average Cal sensor Temp.: 103.96 Ambient Temp.: 88.36 F
02:11:26.154 -> Average Cal sensor Temp.: 103.82 Ambient Temp.: 88.39 F
02:11:27.740 -> Average Cal sensor Temp.: 104.23 Ambient Temp.: 88.36 F
02:11:29.299 -> Average Cal sensor Temp.: 104.29 Ambient Temp.: 88.36 F
02:11:30.863 -> Average Cal sensor Temp.: 104.47 Ambient Temp.: 88.47 F
02:11:32.417 -> Average Cal sensor Temp.: 104.23 Ambient Temp.: 88.39 F
02:11:33.978 -> Average Cal sensor Temp.: 104.07 Ambient Temp.: 88.47 F
02:11:35.566 -> Average Cal sensor Temp.: 103.95 Ambient Temp.: 88.47 F
02:11:37.113 -> Average Cal sensor Temp.: 104.30 Ambient Temp.: 88.50 F
02:11:38.666 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 88.54 F
02:11:40.259 -> Average Cal sensor Temp.: 104.07 Ambient Temp.: 88.47 F
02:11:41.810 -> Average Cal sensor Temp.: 104.22 Ambient Temp.: 88.50 F
02:11:43.359 -> Average Cal sensor Temp.: 104.02 Ambient Temp.: 88.47 F
02:11:44.950 -> Average Cal sensor Temp.: 104.41 Ambient Temp.: 88.57 F
02:11:46.500 -> Average Cal sensor Temp.: 104.64 Ambient Temp.: 88.54 F
02:11:48.089 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 88.54 F
02:11:49.644 -> Average Cal sensor Temp.: 103.15 Ambient Temp.: 88.54 F
02:11:51.198 -> Average Cal sensor Temp.: 103.51 Ambient Temp.: 88.47 F
02:11:52.753 -> Average Cal sensor Temp.: 103.62 Ambient Temp.: 88.50 F
02:11:54.315 -> Average Cal sensor Temp.: 103.55 Ambient Temp.: 88.39 F
02:11:55.902 -> Average Cal sensor Temp.: 103.75 Ambient Temp.: 88.50 F
02:11:57.452 -> Average Cal sensor Temp.: 104.31 Ambient Temp.: 88.50 F
02:11:59.008 -> Average Cal sensor Temp.: 104.48 Ambient Temp.: 88.50 F
02:12:00.598 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 88.50 F
02:12:02.151 -> Average Cal sensor Temp.: 103.78 Ambient Temp.: 88.50 F
02:12:03.709 -> Average Cal sensor Temp.: 103.82 Ambient Temp.: 88.47 F
02:12:05.300 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 88.47 F
02:12:06.848 -> Average Cal sensor Temp.: 103.80 Ambient Temp.: 88.36 F
02:12:08.410 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 88.32 F
02:12:09.972 -> Average Cal sensor Temp.: 104.09 Ambient Temp.: 88.32 F
02:12:11.530 -> Average Cal sensor Temp.: 104.06 Ambient Temp.: 88.36 F
02:12:13.119 -> Average Cal sensor Temp.: 104.20 Ambient Temp.: 88.36 F
02:12:14.679 -> Average Cal sensor Temp.: 103.93 Ambient Temp.: 88.32 F
02:12:16.233 -> Average Cal sensor Temp.: 103.53 Ambient Temp.: 88.29 F
02:12:17.823 -> Average Cal sensor Temp.: 103.59 Ambient Temp.: 88.21 F
02:12:19.379 -> Average Cal sensor Temp.: 103.37 Ambient Temp.: 88.18 F
02:12:20.940 -> Average Cal sensor Temp.: 104.30 Ambient Temp.: 88.14 F
02:12:22.499 -> Average Cal sensor Temp.: 104.62 Ambient Temp.: 88.21 F
02:12:24.052 -> Average Cal sensor Temp.: 104.23 Ambient Temp.: 88.18 F
02:12:25.614 -> Average Cal sensor Temp.: 103.97 Ambient Temp.: 88.18 F
02:12:27.204 -> Average Cal sensor Temp.: 103.96 Ambient Temp.: 88.14 F
02:12:28.760 -> Average Cal sensor Temp.: 104.32 Ambient Temp.: 88.14 F
02:12:30.317 -> Average Cal sensor Temp.: 104.72 Ambient Temp.: 88.18 F
02:12:31.875 -> Average Cal sensor Temp.: 104.10 Ambient Temp.: 88.11 F
02:12:33.468 -> Average Cal sensor Temp.: 103.65 Ambient Temp.: 88.11 F
02:12:35.014 -> Average Cal sensor Temp.: 104.27 Ambient Temp.: 88.11 F
02:12:36.566 -> Average Cal sensor Temp.: 104.37 Ambient Temp.: 88.11 F
02:12:38.155 -> Average Cal sensor Temp.: 104.95 Ambient Temp.: 88.11 F
02:12:39.709 -> Average Cal sensor Temp.: 105.02 Ambient Temp.: 88.11 F
02:12:41.266 -> Average Cal sensor Temp.: 104.99 Ambient Temp.: 88.18 F
02:12:42.851 -> Average Cal sensor Temp.: 103.88 Ambient Temp.: 88.18 F
02:12:44.412 -> Average Cal sensor Temp.: 103.15 Ambient Temp.: 88.21 F
02:12:45.963 -> Average Cal sensor Temp.: 104.41 Ambient Temp.: 88.18 F
02:12:47.516 -> Average Cal sensor Temp.: 104.84 Ambient Temp.: 88.18 F
02:12:49.107 -> Average Cal sensor Temp.: 104.85 Ambient Temp.: 88.25 F
02:12:50.664 -> Average Cal sensor Temp.: 104.87 Ambient Temp.: 88.25 F
02:12:52.220 -> Average Cal sensor Temp.: 104.61 Ambient Temp.: 88.29 F
02:12:53.784 -> Average Cal sensor Temp.: 104.90 Ambient Temp.: 88.32 F
02:12:55.341 -> Average Cal sensor Temp.: 104.70 Ambient Temp.: 88.29 F
02:12:56.929 -> Average Cal sensor Temp.: 104.53 Ambient Temp.: 88.36 F
02:12:58.489 -> Average Cal sensor Temp.: 104.36 Ambient Temp.: 88.36 F
02:13:00.036 -> Average Cal sensor Temp.: 104.34 Ambient Temp.: 88.39 F
02:13:01.631 -> Average Cal sensor Temp.: 104.56 Ambient Temp.: 88.39 F
02:13:03.192 -> Average Cal sensor Temp.: 104.62 Ambient Temp.: 88.47 F
02:13:04.747 -> Average Cal sensor Temp.: 104.31 Ambient Temp.: 88.47 F
02:13:06.304 -> Average Cal sensor Temp.: 104.13 Ambient Temp.: 88.50 F
02:13:07.887 -> Average Cal sensor Temp.: 104.18 Ambient Temp.: 88.54 F
02:13:09.441 -> Average Cal sensor Temp.: 103.97 Ambient Temp.: 88.50 F
02:13:11.000 -> Average Cal sensor Temp.: 104.26 Ambient Temp.: 88.57 F
02:13:12.588 -> Average Cal sensor Temp.: 103.69 Ambient Temp.: 88.61 F
02:13:14.141 -> Average Cal sensor Temp.: 103.95 Ambient Temp.: 88.61 F
02:13:15.699 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 88.65 F
02:13:17.253 -> Average Cal sensor Temp.: 103.95 Ambient Temp.: 88.61 F
02:13:18.845 -> Average Cal sensor Temp.: 103.82 Ambient Temp.: 88.65 F
02:13:20.395 -> Average Cal sensor Temp.: 104.26 Ambient Temp.: 88.68 F
02:13:21.951 -> Average Cal sensor Temp.: 104.06 Ambient Temp.: 88.68 F
02:13:23.541 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 88.72 F
02:13:25.095 -> Average Cal sensor Temp.: 103.64 Ambient Temp.: 88.68 F
02:13:26.651 -> Average Cal sensor Temp.: 103.71 Ambient Temp.: 88.72 F
02:13:28.209 -> Average Cal sensor Temp.: 103.84 Ambient Temp.: 88.79 F
02:13:29.795 -> Average Cal sensor Temp.: 103.82 Ambient Temp.: 88.79 F
02:13:31.353 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 88.79 F
02:13:32.937 -> Average Cal sensor Temp.: 103.58 Ambient Temp.: 88.75 F
02:13:34.498 -> Average Cal sensor Temp.: 103.81 Ambient Temp.: 88.86 F
02:13:36.052 -> Average Cal sensor Temp.: 103.82 Ambient Temp.: 88.75 F
02:13:37.607 -> Average Cal sensor Temp.: 104.12 Ambient Temp.: 88.86 F
02:13:39.194 -> Average Cal sensor Temp.: 104.29 Ambient Temp.: 88.86 F
02:13:40.754 -> Average Cal sensor Temp.: 103.99 Ambient Temp.: 88.90 F
02:13:42.310 -> Average Cal sensor Temp.: 103.98 Ambient Temp.: 88.90 F
02:13:43.868 -> Average Cal sensor Temp.: 104.28 Ambient Temp.: 88.86 F
02:13:45.458 -> Average Cal sensor Temp.: 104.40 Ambient Temp.: 88.93 F
02:13:47.006 -> Average Cal sensor Temp.: 104.35 Ambient Temp.: 88.93 F
02:13:48.564 -> Average Cal sensor Temp.: 104.25 Ambient Temp.: 88.97 F
02:13:50.158 -> Average Cal sensor Temp.: 103.91 Ambient Temp.: 88.97 F
02:13:51.714 -> Average Cal sensor Temp.: 103.57 Ambient Temp.: 89.01 F
02:13:53.276 -> Average Cal sensor Temp.: 103.67 Ambient Temp.: 88.97 F
02:13:54.838 -> Average Cal sensor Temp.: 104.02 Ambient Temp.: 88.97 F
02:13:56.395 -> Average Cal sensor Temp.: 104.60 Ambient Temp.: 89.01 F
02:13:57.984 -> Average Cal sensor Temp.: 104.25 Ambient Temp.: 89.04 F
02:13:59.543 -> Average Cal sensor Temp.: 104.30 Ambient Temp.: 89.11 F
02:14:01.099 -> Average Cal sensor Temp.: 103.75 Ambient Temp.: 89.11 F
02:14:02.650 -> Average Cal sensor Temp.: 103.69 Ambient Temp.: 89.04 F
02:14:04.240 -> Average Cal sensor Temp.: 103.82 Ambient Temp.: 89.19 F
02:14:05.800 -> Average Cal sensor Temp.: 104.04 Ambient Temp.: 89.11 F
02:14:07.357 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 89.11 F
02:14:08.917 -> Average Cal sensor Temp.: 104.21 Ambient Temp.: 89.11 F
02:14:10.472 -> Average Cal sensor Temp.: 103.91 Ambient Temp.: 89.15 F
02:14:12.033 -> Average Cal sensor Temp.: 103.48 Ambient Temp.: 89.11 F
02:14:13.594 -> Average Cal sensor Temp.: 103.51 Ambient Temp.: 89.11 F
02:14:15.180 -> Average Cal sensor Temp.: 103.87 Ambient Temp.: 89.08 F
02:14:16.737 -> Average Cal sensor Temp.: 104.09 Ambient Temp.: 89.08 F
02:14:18.294 -> Average Cal sensor Temp.: 104.32 Ambient Temp.: 89.11 F
02:14:19.883 -> Average Cal sensor Temp.: 103.59 Ambient Temp.: 89.08 F
02:14:21.443 -> Average Cal sensor Temp.: 103.62 Ambient Temp.: 89.11 F
02:14:22.994 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 89.04 F
02:14:24.555 -> Average Cal sensor Temp.: 103.65 Ambient Temp.: 89.04 F
02:14:26.144 -> Average Cal sensor Temp.: 103.61 Ambient Temp.: 89.01 F
02:14:27.696 -> Average Cal sensor Temp.: 103.95 Ambient Temp.: 89.04 F
02:14:29.251 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 89.04 F
02:14:30.812 -> Average Cal sensor Temp.: 103.67 Ambient Temp.: 89.01 F
02:14:32.404 -> Average Cal sensor Temp.: 103.81 Ambient Temp.: 89.04 F
02:14:33.956 -> Average Cal sensor Temp.: 103.71 Ambient Temp.: 89.01 F
02:14:35.512 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 89.04 F
02:14:37.097 -> Average Cal sensor Temp.: 103.64 Ambient Temp.: 89.08 F
02:14:38.656 -> Average Cal sensor Temp.: 103.55 Ambient Temp.: 89.08 F
02:14:40.206 -> Average Cal sensor Temp.: 103.49 Ambient Temp.: 89.01 F
02:14:41.793 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 89.01 F
02:14:43.349 -> Average Cal sensor Temp.: 103.57 Ambient Temp.: 89.04 F
02:14:44.905 -> Average Cal sensor Temp.: 103.34 Ambient Temp.: 89.04 F
02:14:46.464 -> Average Cal sensor Temp.: 103.76 Ambient Temp.: 89.08 F
02:14:48.019 -> Average Cal sensor Temp.: 103.68 Ambient Temp.: 89.08 F
02:14:49.606 -> Average Cal sensor Temp.: 103.54 Ambient Temp.: 89.04 F
02:14:51.165 -> Average Cal sensor Temp.: 104.17 Ambient Temp.: 89.08 F
02:14:52.718 -> Average Cal sensor Temp.: 103.84 Ambient Temp.: 89.08 F
02:14:54.315 -> Average Cal sensor Temp.: 103.59 Ambient Temp.: 89.08 F
02:14:55.871 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 89.08 F
02:14:57.435 -> Average Cal sensor Temp.: 103.55 Ambient Temp.: 89.19 F
02:14:58.990 -> Average Cal sensor Temp.: 103.39 Ambient Temp.: 89.11 F
02:15:00.551 -> Average Cal sensor Temp.: 103.61 Ambient Temp.: 89.11 F
02:15:02.102 -> Average Cal sensor Temp.: 103.48 Ambient Temp.: 89.11 F
02:15:03.693 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 89.04 F
02:15:05.248 -> Average Cal sensor Temp.: 104.08 Ambient Temp.: 89.11 F
02:15:06.806 -> Average Cal sensor Temp.: 103.67 Ambient Temp.: 89.08 F
02:15:08.394 -> Average Cal sensor Temp.: 103.51 Ambient Temp.: 89.08 F
02:15:09.953 -> Average Cal sensor Temp.: 103.52 Ambient Temp.: 89.08 F
02:15:11.510 -> Average Cal sensor Temp.: 103.63 Ambient Temp.: 89.08 F
02:15:13.067 -> Average Cal sensor Temp.: 104.09 Ambient Temp.: 89.15 F
02:15:14.624 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 89.08 F
02:15:16.211 -> Average Cal sensor Temp.: 103.96 Ambient Temp.: 89.11 F
02:15:17.762 -> Average Cal sensor Temp.: 104.12 Ambient Temp.: 89.11 F
02:15:19.317 -> Average Cal sensor Temp.: 104.27 Ambient Temp.: 89.11 F
02:15:20.911 -> Average Cal sensor Temp.: 103.59 Ambient Temp.: 89.11 F
02:15:22.469 -> Average Cal sensor Temp.: 103.51 Ambient Temp.: 89.08 F
02:15:24.022 -> Average Cal sensor Temp.: 104.70 Ambient Temp.: 89.11 F
02:15:25.583 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 89.08 F
02:15:27.171 -> Average Cal sensor Temp.: 104.26 Ambient Temp.: 89.08 F
02:15:28.730 -> Average Cal sensor Temp.: 104.22 Ambient Temp.: 89.15 F
02:15:30.287 -> Average Cal sensor Temp.: 104.26 Ambient Temp.: 89.15 F
02:15:31.839 -> Average Cal sensor Temp.: 104.06 Ambient Temp.: 89.15 F
02:15:33.398 -> Average Cal sensor Temp.: 104.27 Ambient Temp.: 89.19 F
02:15:34.987 -> Average Cal sensor Temp.: 103.93 Ambient Temp.: 89.19 F
02:15:36.543 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 89.26 F
02:15:38.098 -> Average Cal sensor Temp.: 104.05 Ambient Temp.: 89.26 F
02:15:39.688 -> Average Cal sensor Temp.: 102.82 Ambient Temp.: 89.15 F
02:15:41.240 -> Average Cal sensor Temp.: 103.06 Ambient Temp.: 89.15 F
02:15:42.796 -> Average Cal sensor Temp.: 103.01 Ambient Temp.: 89.08 F
02:15:44.351 -> Average Cal sensor Temp.: 103.28 Ambient Temp.: 89.08 F
02:15:45.941 -> Average Cal sensor Temp.: 103.37 Ambient Temp.: 89.08 F
02:15:47.500 -> Average Cal sensor Temp.: 103.81 Ambient Temp.: 89.11 F
02:15:49.052 -> Average Cal sensor Temp.: 103.96 Ambient Temp.: 89.08 F
02:15:50.615 -> Average Cal sensor Temp.: 103.93 Ambient Temp.: 89.08 F
02:15:52.206 -> Average Cal sensor Temp.: 103.89 Ambient Temp.: 89.04 F
02:15:53.765 -> Average Cal sensor Temp.: 103.25 Ambient Temp.: 89.01 F
02:15:55.320 -> Average Cal sensor Temp.: 103.26 Ambient Temp.: 89.01 F
02:15:56.874 -> Average Cal sensor Temp.: 103.31 Ambient Temp.: 88.97 F
02:15:58.465 -> Average Cal sensor Temp.: 103.46 Ambient Temp.: 88.97 F
02:16:00.017 -> Average Cal sensor Temp.: 102.89 Ambient Temp.: 88.90 F
02:16:01.580 -> Average Cal sensor Temp.: 103.04 Ambient Temp.: 88.90 F
02:16:03.135 -> Average Cal sensor Temp.: 103.07 Ambient Temp.: 88.86 F
02:16:04.722 -> Average Cal sensor Temp.: 103.56 Ambient Temp.: 88.79 F
02:16:06.276 -> Average Cal sensor Temp.: 103.56 Ambient Temp.: 88.75 F
02:16:07.859 -> Average Cal sensor Temp.: 103.80 Ambient Temp.: 88.79 F
02:16:09.416 -> Average Cal sensor Temp.: 104.05 Ambient Temp.: 88.75 F
02:16:10.974 -> Average Cal sensor Temp.: 104.41 Ambient Temp.: 88.79 F
02:16:12.532 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 88.75 F
02:16:14.118 -> Average Cal sensor Temp.: 103.96 Ambient Temp.: 88.72 F
02:16:15.673 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 88.72 F
02:16:17.230 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 88.65 F
02:16:18.786 -> Average Cal sensor Temp.: 104.14 Ambient Temp.: 88.68 F
02:16:20.383 -> Average Cal sensor Temp.: 104.13 Ambient Temp.: 88.68 F
02:16:21.945 -> Average Cal sensor Temp.: 103.77 Ambient Temp.: 88.61 F
02:16:23.503 -> Average Cal sensor Temp.: 103.93 Ambient Temp.: 88.61 F
02:16:25.061 -> Average Cal sensor Temp.: 103.91 Ambient Temp.: 88.54 F
02:16:26.621 -> Average Cal sensor Temp.: 103.84 Ambient Temp.: 88.47 F
02:16:28.182 -> Average Cal sensor Temp.: 103.82 Ambient Temp.: 88.47 F
02:16:29.738 -> Average Cal sensor Temp.: 104.07 Ambient Temp.: 88.39 F
02:16:31.334 -> Average Cal sensor Temp.: 104.49 Ambient Temp.: 88.39 F
02:16:32.885 -> Average Cal sensor Temp.: 103.64 Ambient Temp.: 88.47 F
02:16:34.441 -> Average Cal sensor Temp.: 103.68 Ambient Temp.: 88.39 F
02:16:36.028 -> Average Cal sensor Temp.: 103.43 Ambient Temp.: 88.39 F
02:16:37.583 -> Average Cal sensor Temp.: 103.59 Ambient Temp.: 88.32 F
02:16:39.137 -> Average Cal sensor Temp.: 103.59 Ambient Temp.: 88.36 F
02:16:40.727 -> Average Cal sensor Temp.: 103.66 Ambient Temp.: 88.29 F
02:16:42.281 -> Average Cal sensor Temp.: 103.64 Ambient Temp.: 88.29 F
02:16:43.835 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 88.29 F
02:16:45.418 -> Average Cal sensor Temp.: 103.85 Ambient Temp.: 88.29 F
02:16:46.969 -> Average Cal sensor Temp.: 103.63 Ambient Temp.: 88.29 F
02:16:48.522 -> Average Cal sensor Temp.: 103.98 Ambient Temp.: 88.25 F
02:16:50.108 -> Average Cal sensor Temp.: 104.20 Ambient Temp.: 88.21 F
02:16:51.664 -> Average Cal sensor Temp.: 104.08 Ambient Temp.: 88.21 F
02:16:53.250 -> Average Cal sensor Temp.: 104.23 Ambient Temp.: 88.25 F
02:16:54.803 -> Average Cal sensor Temp.: 104.29 Ambient Temp.: 88.25 F
02:16:56.353 -> Average Cal sensor Temp.: 103.80 Ambient Temp.: 88.18 F
02:16:57.940 -> Average Cal sensor Temp.: 103.70 Ambient Temp.: 88.14 F
02:16:59.491 -> Average Cal sensor Temp.: 103.58 Ambient Temp.: 88.18 F
02:17:01.048 -> Average Cal sensor Temp.: 103.97 Ambient Temp.: 88.11 F
02:17:02.636 -> Average Cal sensor Temp.: 104.03 Ambient Temp.: 88.18 F
02:17:04.194 -> Average Cal sensor Temp.: 104.31 Ambient Temp.: 88.14 F
02:17:05.750 -> Average Cal sensor Temp.: 104.32 Ambient Temp.: 88.14 F
02:17:07.310 -> Average Cal sensor Temp.: 104.45 Ambient Temp.: 88.18 F
02:17:08.890 -> Average Cal sensor Temp.: 103.93 Ambient Temp.: 88.18 F
02:17:10.440 -> Average Cal sensor Temp.: 104.31 Ambient Temp.: 88.14 F
02:17:12.027 -> Average Cal sensor Temp.: 103.91 Ambient Temp.: 88.21 F
02:17:13.584 -> Average Cal sensor Temp.: 104.05 Ambient Temp.: 88.25 F
02:17:15.142 -> Average Cal sensor Temp.: 104.97 Ambient Temp.: 88.25 F
02:17:16.701 -> Average Cal sensor Temp.: 104.57 Ambient Temp.: 88.32 F
02:17:18.255 -> Average Cal sensor Temp.: 104.03 Ambient Temp.: 88.29 F
02:17:19.845 -> Average Cal sensor Temp.: 103.82 Ambient Temp.: 88.32 F
02:17:21.403 -> Average Cal sensor Temp.: 104.14 Ambient Temp.: 88.29 F
02:17:22.958 -> Average Cal sensor Temp.: 104.47 Ambient Temp.: 88.32 F
02:17:24.515 -> Average Cal sensor Temp.: 104.42 Ambient Temp.: 88.36 F
02:17:26.109 -> Average Cal sensor Temp.: 103.97 Ambient Temp.: 88.36 F
02:17:27.665 -> Average Cal sensor Temp.: 103.94 Ambient Temp.: 88.39 F
02:17:29.215 -> Average Cal sensor Temp.: 103.92 Ambient Temp.: 88.36 F
02:17:30.776 -> Average Cal sensor Temp.: 104.14 Ambient Temp.: 88.36 F
02:17:32.367 -> Average Cal sensor Temp.: 104.02 Ambient Temp.: 88.47 F
02:17:33.926 -> Average Cal sensor Temp.: 103.85 Ambient Temp.: 88.39 F
02:17:35.480 -> Average Cal sensor Temp.: 104.17 Ambient Temp.: 88.39 F
02:17:37.037 -> Average Cal sensor Temp.: 104.57 Ambient Temp.: 88.50 F
02:17:38.597 -> Average Cal sensor Temp.: 104.39 Ambient Temp.: 88.54 F
02:17:40.190 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 88.47 F
02:17:41.749 -> Average Cal sensor Temp.: 103.58 Ambient Temp.: 88.50 F
02:17:43.306 -> Average Cal sensor Temp.: 103.67 Ambient Temp.: 88.39 F
02:17:44.860 -> Average Cal sensor Temp.: 104.16 Ambient Temp.: 88.36 F
02:17:46.453 -> Average Cal sensor Temp.: 103.89 Ambient Temp.: 88.39 F
02:17:48.010 -> Average Cal sensor Temp.: 104.42 Ambient Temp.: 88.39 F
02:17:49.569 -> Average Cal sensor Temp.: 104.25 Ambient Temp.: 88.39 F
02:17:51.124 -> Average Cal sensor Temp.: 103.93 Ambient Temp.: 88.39 F
02:17:52.684 -> Average Cal sensor Temp.: 103.55 Ambient Temp.: 88.47 F
02:17:54.275 -> Average Cal sensor Temp.: 103.75 Ambient Temp.: 88.36 F
02:17:55.831 -> Average Cal sensor Temp.: 103.40 Ambient Temp.: 88.36 F
02:17:57.384 -> Average Cal sensor Temp.: 103.74 Ambient Temp.: 88.36 F
02:17:58.944 -> Average Cal sensor Temp.: 103.76 Ambient Temp.: 88.47 F
02:18:00.533 -> Average Cal sensor Temp.: 103.66 Ambient Temp.: 88.32 F
02:18:02.090 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 88.29 F
02:18:03.647 -> Average Cal sensor Temp.: 103.92 Ambient Temp.: 88.29 F
02:18:05.204 -> Average Cal sensor Temp.: 103.96 Ambient Temp.: 88.29 F
02:18:06.794 -> Average Cal sensor Temp.: 103.61 Ambient Temp.: 88.32 F
02:18:08.357 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 88.25 F
02:18:09.913 -> Average Cal sensor Temp.: 103.96 Ambient Temp.: 88.29 F
02:18:11.470 -> Average Cal sensor Temp.: 103.70 Ambient Temp.: 88.29 F
02:18:13.026 -> Average Cal sensor Temp.: 103.91 Ambient Temp.: 88.29 F
02:18:14.617 -> Average Cal sensor Temp.: 104.07 Ambient Temp.: 88.25 F
02:18:16.167 -> Average Cal sensor Temp.: 104.29 Ambient Temp.: 88.25 F
02:18:17.753 -> Average Cal sensor Temp.: 104.10 Ambient Temp.: 88.25 F
02:18:19.307 -> Average Cal sensor Temp.: 104.46 Ambient Temp.: 88.29 F
02:18:20.862 -> Average Cal sensor Temp.: 104.21 Ambient Temp.: 88.29 F
02:18:22.446 -> Average Cal sensor Temp.: 104.19 Ambient Temp.: 88.25 F
02:18:23.997 -> Average Cal sensor Temp.: 104.59 Ambient Temp.: 88.32 F
02:18:25.549 -> Average Cal sensor Temp.: 103.68 Ambient Temp.: 88.25 F
02:18:27.135 -> Average Cal sensor Temp.: 104.33 Ambient Temp.: 88.32 F
02:18:28.690 -> Average Cal sensor Temp.: 104.05 Ambient Temp.: 88.32 F
02:18:30.242 -> Average Cal sensor Temp.: 104.13 Ambient Temp.: 88.32 F
02:18:31.834 -> Average Cal sensor Temp.: 104.29 Ambient Temp.: 88.32 F
02:18:33.393 -> Average Cal sensor Temp.: 104.31 Ambient Temp.: 88.29 F
02:18:34.942 -> Average Cal sensor Temp.: 104.18 Ambient Temp.: 88.29 F
02:18:36.531 -> Average Cal sensor Temp.: 104.14 Ambient Temp.: 88.32 F
02:18:38.084 -> Average Cal sensor Temp.: 103.87 Ambient Temp.: 88.29 F
02:18:39.638 -> Average Cal sensor Temp.: 103.69 Ambient Temp.: 88.29 F
02:18:41.194 -> Average Cal sensor Temp.: 103.85 Ambient Temp.: 88.29 F
02:18:42.783 -> Average Cal sensor Temp.: 103.79 Ambient Temp.: 88.32 F
02:18:44.344 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 88.29 F
02:18:45.893 -> Average Cal sensor Temp.: 103.98 Ambient Temp.: 88.29 F
02:18:47.484 -> Average Cal sensor Temp.: 103.69 Ambient Temp.: 88.32 F
02:18:49.032 -> Average Cal sensor Temp.: 103.82 Ambient Temp.: 88.29 F
02:18:50.588 -> Average Cal sensor Temp.: 104.04 Ambient Temp.: 88.21 F
02:18:52.176 -> Average Cal sensor Temp.: 104.46 Ambient Temp.: 88.25 F
02:18:53.735 -> Average Cal sensor Temp.: 104.35 Ambient Temp.: 88.25 F
02:18:55.284 -> Average Cal sensor Temp.: 103.57 Ambient Temp.: 88.21 F
02:18:56.854 -> Average Cal sensor Temp.: 104.05 Ambient Temp.: 88.21 F
02:18:58.439 -> Average Cal sensor Temp.: 104.31 Ambient Temp.: 88.25 F
02:18:59.996 -> Average Cal sensor Temp.: 104.74 Ambient Temp.: 88.25 F
02:19:01.549 -> Average Cal sensor Temp.: 104.25 Ambient Temp.: 88.21 F
02:19:03.103 -> Average Cal sensor Temp.: 104.43 Ambient Temp.: 88.29 F
02:19:04.690 -> Average Cal sensor Temp.: 104.05 Ambient Temp.: 88.29 F
02:19:06.245 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 88.25 F
02:19:07.799 -> Average Cal sensor Temp.: 104.36 Ambient Temp.: 88.32 F
02:19:09.391 -> Average Cal sensor Temp.: 103.97 Ambient Temp.: 88.36 F
02:19:10.942 -> Average Cal sensor Temp.: 104.54 Ambient Temp.: 88.32 F
02:19:12.494 -> Average Cal sensor Temp.: 104.39 Ambient Temp.: 88.32 F
02:19:14.082 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 88.29 F
02:19:15.637 -> Average Cal sensor Temp.: 103.98 Ambient Temp.: 88.32 F
02:19:17.217 -> Average Cal sensor Temp.: 103.84 Ambient Temp.: 88.32 F
02:19:18.774 -> Average Cal sensor Temp.: 104.05 Ambient Temp.: 88.29 F
02:19:20.331 -> Average Cal sensor Temp.: 103.99 Ambient Temp.: 88.32 F
02:19:21.916 -> Average Cal sensor Temp.: 103.43 Ambient Temp.: 88.25 F
02:19:23.476 -> Average Cal sensor Temp.: 103.89 Ambient Temp.: 88.29 F
02:19:25.020 -> Average Cal sensor Temp.: 104.24 Ambient Temp.: 88.18 F
02:19:26.615 -> Average Cal sensor Temp.: 104.44 Ambient Temp.: 88.21 F
02:19:28.174 -> Average Cal sensor Temp.: 103.99 Ambient Temp.: 88.25 F
02:19:29.728 -> Average Cal sensor Temp.: 104.76 Ambient Temp.: 88.25 F
02:19:31.287 -> Average Cal sensor Temp.: 104.24 Ambient Temp.: 88.25 F
02:19:32.842 -> Average Cal sensor Temp.: 104.25 Ambient Temp.: 88.21 F
02:19:34.439 -> Average Cal sensor Temp.: 103.95 Ambient Temp.: 88.21 F
02:19:35.999 -> Average Cal sensor Temp.: 104.10 Ambient Temp.: 88.21 F
02:19:37.555 -> Average Cal sensor Temp.: 104.36 Ambient Temp.: 88.21 F
02:19:39.111 -> Average Cal sensor Temp.: 104.47 Ambient Temp.: 88.21 F
02:19:40.668 -> Average Cal sensor Temp.: 104.33 Ambient Temp.: 88.25 F
02:19:42.259 -> Average Cal sensor Temp.: 104.28 Ambient Temp.: 88.21 F
02:19:43.818 -> Average Cal sensor Temp.: 104.56 Ambient Temp.: 88.25 F
02:19:45.371 -> Average Cal sensor Temp.: 104.34 Ambient Temp.: 88.21 F
02:19:46.923 -> Average Cal sensor Temp.: 104.31 Ambient Temp.: 88.29 F
02:19:48.513 -> Average Cal sensor Temp.: 104.19 Ambient Temp.: 88.32 F
02:19:50.078 -> Average Cal sensor Temp.: 104.09 Ambient Temp.: 88.29 F
02:19:51.630 -> Average Cal sensor Temp.: 104.67 Ambient Temp.: 88.32 F
02:19:53.188 -> Average Cal sensor Temp.: 103.91 Ambient Temp.: 88.36 F
02:19:54.778 -> Average Cal sensor Temp.: 104.37 Ambient Temp.: 88.29 F
02:19:56.330 -> Average Cal sensor Temp.: 104.52 Ambient Temp.: 88.32 F
02:19:57.909 -> Average Cal sensor Temp.: 104.15 Ambient Temp.: 88.25 F
02:19:59.465 -> Average Cal sensor Temp.: 104.87 Ambient Temp.: 88.29 F
02:20:01.022 -> Average Cal sensor Temp.: 104.29 Ambient Temp.: 88.32 F
02:20:02.570 -> Average Cal sensor Temp.: 103.92 Ambient Temp.: 88.32 F
02:20:04.161 -> Average Cal sensor Temp.: 104.13 Ambient Temp.: 88.36 F
02:20:05.713 -> Average Cal sensor Temp.: 104.58 Ambient Temp.: 88.32 F
02:20:07.265 -> Average Cal sensor Temp.: 103.98 Ambient Temp.: 88.29 F
02:20:08.854 -> Average Cal sensor Temp.: 104.32 Ambient Temp.: 88.29 F
02:20:10.405 -> Average Cal sensor Temp.: 104.82 Ambient Temp.: 88.32 F
02:20:11.963 -> Average Cal sensor Temp.: 104.47 Ambient Temp.: 88.32 F
02:20:13.548 -> Average Cal sensor Temp.: 104.30 Ambient Temp.: 88.36 F
02:20:15.099 -> Average Cal sensor Temp.: 104.67 Ambient Temp.: 88.39 F
02:20:16.685 -> Average Cal sensor Temp.: 104.07 Ambient Temp.: 88.39 F
02:20:18.239 -> Average Cal sensor Temp.: 103.97 Ambient Temp.: 88.36 F
02:20:19.788 -> Average Cal sensor Temp.: 104.19 Ambient Temp.: 88.36 F
02:20:21.379 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 88.32 F
02:20:22.925 -> Average Cal sensor Temp.: 103.92 Ambient Temp.: 88.32 F
02:20:24.509 -> Average Cal sensor Temp.: 104.58 Ambient Temp.: 88.32 F
02:20:26.064 -> Average Cal sensor Temp.: 104.46 Ambient Temp.: 88.36 F
02:20:27.615 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 88.32 F
02:20:29.208 -> Average Cal sensor Temp.: 103.51 Ambient Temp.: 88.36 F
02:20:30.760 -> Average Cal sensor Temp.: 103.12 Ambient Temp.: 88.29 F
02:20:32.303 -> Average Cal sensor Temp.: 103.61 Ambient Temp.: 88.29 F
02:20:33.895 -> Average Cal sensor Temp.: 103.74 Ambient Temp.: 88.25 F
02:20:35.440 -> Average Cal sensor Temp.: 103.92 Ambient Temp.: 88.25 F
02:20:37.029 -> Average Cal sensor Temp.: 103.88 Ambient Temp.: 88.29 F
02:20:38.587 -> Average Cal sensor Temp.: 104.07 Ambient Temp.: 88.29 F
02:20:40.145 -> Average Cal sensor Temp.: 103.14 Ambient Temp.: 88.25 F
02:20:41.699 -> Average Cal sensor Temp.: 103.37 Ambient Temp.: 88.21 F
02:20:43.288 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 88.29 F
02:20:44.845 -> Average Cal sensor Temp.: 103.68 Ambient Temp.: 88.25 F
02:20:46.395 -> Average Cal sensor Temp.: 103.45 Ambient Temp.: 88.18 F
02:20:47.983 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 88.18 F
02:20:49.538 -> Average Cal sensor Temp.: 103.63 Ambient Temp.: 88.21 F
02:20:51.092 -> Average Cal sensor Temp.: 103.53 Ambient Temp.: 88.18 F
02:20:52.645 -> Average Cal sensor Temp.: 103.90 Ambient Temp.: 88.18 F
02:20:54.235 -> Average Cal sensor Temp.: 104.10 Ambient Temp.: 88.21 F
02:20:55.792 -> Average Cal sensor Temp.: 104.10 Ambient Temp.: 88.21 F
02:20:57.339 -> Average Cal sensor Temp.: 104.04 Ambient Temp.: 88.21 F
02:20:58.922 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 88.21 F
02:21:00.475 -> Average Cal sensor Temp.: 103.92 Ambient Temp.: 88.21 F
02:21:02.034 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 88.21 F
02:21:03.621 -> Average Cal sensor Temp.: 104.05 Ambient Temp.: 88.18 F
02:21:05.177 -> Average Cal sensor Temp.: 104.62 Ambient Temp.: 88.25 F
02:21:06.731 -> Average Cal sensor Temp.: 103.82 Ambient Temp.: 88.14 F
02:21:08.320 -> Average Cal sensor Temp.: 103.59 Ambient Temp.: 88.14 F
02:21:09.869 -> Average Cal sensor Temp.: 103.68 Ambient Temp.: 88.14 F
02:21:11.422 -> Average Cal sensor Temp.: 103.87 Ambient Temp.: 88.07 F
02:21:13.015 -> Average Cal sensor Temp.: 104.06 Ambient Temp.: 88.11 F
02:21:14.575 -> Average Cal sensor Temp.: 104.18 Ambient Temp.: 88.11 F
02:21:16.134 -> Average Cal sensor Temp.: 104.02 Ambient Temp.: 88.14 F
02:21:17.687 -> Average Cal sensor Temp.: 103.93 Ambient Temp.: 88.07 F
02:21:19.242 -> Average Cal sensor Temp.: 103.53 Ambient Temp.: 87.96 F
02:21:20.833 -> Average Cal sensor Temp.: 103.53 Ambient Temp.: 87.96 F
02:21:22.392 -> Average Cal sensor Temp.: 104.10 Ambient Temp.: 88.00 F
02:21:23.948 -> Average Cal sensor Temp.: 104.02 Ambient Temp.: 88.00 F
02:21:25.503 -> Average Cal sensor Temp.: 103.87 Ambient Temp.: 87.96 F
02:21:27.091 -> Average Cal sensor Temp.: 104.14 Ambient Temp.: 87.93 F
02:21:28.645 -> Average Cal sensor Temp.: 103.64 Ambient Temp.: 87.96 F
02:21:30.196 -> Average Cal sensor Temp.: 103.74 Ambient Temp.: 87.93 F
02:21:31.780 -> Average Cal sensor Temp.: 103.98 Ambient Temp.: 87.89 F
02:21:33.337 -> Average Cal sensor Temp.: 104.43 Ambient Temp.: 87.85 F
02:21:34.894 -> Average Cal sensor Temp.: 103.80 Ambient Temp.: 87.78 F
02:21:36.483 -> Average Cal sensor Temp.: 103.59 Ambient Temp.: 87.78 F
02:21:38.042 -> Average Cal sensor Temp.: 104.03 Ambient Temp.: 87.82 F
02:21:39.597 -> Average Cal sensor Temp.: 104.35 Ambient Temp.: 87.82 F
02:21:41.153 -> Average Cal sensor Temp.: 104.09 Ambient Temp.: 87.82 F
02:21:42.738 -> Average Cal sensor Temp.: 103.91 Ambient Temp.: 87.78 F
02:21:44.288 -> Average Cal sensor Temp.: 104.08 Ambient Temp.: 87.78 F
02:21:45.848 -> Average Cal sensor Temp.: 104.77 Ambient Temp.: 87.78 F
02:21:47.438 -> Average Cal sensor Temp.: 104.33 Ambient Temp.: 87.78 F
02:21:48.994 -> Average Cal sensor Temp.: 104.94 Ambient Temp.: 87.78 F
02:21:50.547 -> Average Cal sensor Temp.: 103.88 Ambient Temp.: 87.82 F
02:21:52.133 -> Average Cal sensor Temp.: 103.66 Ambient Temp.: 87.78 F
02:21:53.692 -> Average Cal sensor Temp.: 103.77 Ambient Temp.: 87.82 F
02:21:55.251 -> Average Cal sensor Temp.: 103.84 Ambient Temp.: 87.75 F
02:21:56.807 -> Average Cal sensor Temp.: 103.98 Ambient Temp.: 87.71 F
02:21:58.363 -> Average Cal sensor Temp.: 104.02 Ambient Temp.: 87.71 F
02:21:59.922 -> Average Cal sensor Temp.: 103.83 Ambient Temp.: 87.71 F
02:22:01.510 -> Average Cal sensor Temp.: 103.67 Ambient Temp.: 87.67 F
02:22:03.069 -> Average Cal sensor Temp.: 104.06 Ambient Temp.: 87.71 F
02:22:04.629 -> Average Cal sensor Temp.: 104.18 Ambient Temp.: 87.67 F
02:22:06.189 -> Average Cal sensor Temp.: 104.23 Ambient Temp.: 87.67 F
02:22:07.775 -> Average Cal sensor Temp.: 104.13 Ambient Temp.: 87.53 F
02:22:09.332 -> Average Cal sensor Temp.: 103.98 Ambient Temp.: 87.57 F
02:22:10.885 -> Average Cal sensor Temp.: 104.07 Ambient Temp.: 87.60 F
02:22:12.435 -> Average Cal sensor Temp.: 104.29 Ambient Temp.: 87.53 F
02:22:14.023 -> Average Cal sensor Temp.: 104.03 Ambient Temp.: 87.57 F
02:22:15.580 -> Average Cal sensor Temp.: 104.23 Ambient Temp.: 87.49 F
02:22:17.134 -> Average Cal sensor Temp.: 103.68 Ambient Temp.: 87.49 F
02:22:18.725 -> Average Cal sensor Temp.: 104.56 Ambient Temp.: 87.49 F
02:22:20.273 -> Average Cal sensor Temp.: 104.17 Ambient Temp.: 87.53 F
02:22:21.830 -> Average Cal sensor Temp.: 104.31 Ambient Temp.: 87.42 F
02:22:23.393 -> Average Cal sensor Temp.: 104.61 Ambient Temp.: 87.49 F
02:22:24.979 -> Average Cal sensor Temp.: 104.92 Ambient Temp.: 87.46 F
02:22:26.532 -> Average Cal sensor Temp.: 104.96 Ambient Temp.: 87.42 F
02:22:28.085 -> Average Cal sensor Temp.: 104.78 Ambient Temp.: 87.42 F
02:22:29.674 -> Average Cal sensor Temp.: 104.92 Ambient Temp.: 87.42 F
02:22:31.229 -> Average Cal sensor Temp.: 104.84 Ambient Temp.: 87.46 F
02:22:32.782 -> Average Cal sensor Temp.: 104.14 Ambient Temp.: 87.46 F
02:22:34.373 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 87.46 F
02:22:35.932 -> Average Cal sensor Temp.: 103.54 Ambient Temp.: 87.42 F
02:22:37.484 -> Average Cal sensor Temp.: 103.97 Ambient Temp.: 87.46 F
02:22:39.038 -> Average Cal sensor Temp.: 104.36 Ambient Temp.: 87.46 F
02:22:40.629 -> Average Cal sensor Temp.: 104.32 Ambient Temp.: 87.46 F
02:22:42.179 -> Average Cal sensor Temp.: 104.70 Ambient Temp.: 87.46 F
02:22:43.762 -> Average Cal sensor Temp.: 104.63 Ambient Temp.: 87.39 F
02:22:45.313 -> Average Cal sensor Temp.: 104.45 Ambient Temp.: 87.49 F
02:22:46.863 -> Average Cal sensor Temp.: 104.78 Ambient Temp.: 87.42 F
02:22:48.452 -> Average Cal sensor Temp.: 105.06 Ambient Temp.: 87.49 F
02:22:50.007 -> Average Cal sensor Temp.: 104.81 Ambient Temp.: 87.53 F
02:22:51.559 -> Average Cal sensor Temp.: 104.13 Ambient Temp.: 87.49 F
02:22:53.148 -> Average Cal sensor Temp.: 104.23 Ambient Temp.: 87.46 F
02:22:54.710 -> Average Cal sensor Temp.: 104.90 Ambient Temp.: 87.53 F
02:22:56.266 -> Average Cal sensor Temp.: 104.63 Ambient Temp.: 87.57 F
02:22:57.822 -> Average Cal sensor Temp.: 104.45 Ambient Temp.: 87.49 F
02:22:59.375 -> Average Cal sensor Temp.: 104.49 Ambient Temp.: 87.57 F
02:23:00.967 -> Average Cal sensor Temp.: 104.90 Ambient Temp.: 87.49 F
02:23:02.522 -> Average Cal sensor Temp.: 104.81 Ambient Temp.: 87.53 F
02:23:04.077 -> Average Cal sensor Temp.: 104.40 Ambient Temp.: 87.57 F
02:23:05.635 -> Average Cal sensor Temp.: 104.91 Ambient Temp.: 87.60 F
02:23:07.225 -> Average Cal sensor Temp.: 104.75 Ambient Temp.: 87.60 F
02:23:08.776 -> Average Cal sensor Temp.: 104.12 Ambient Temp.: 87.57 F
02:23:10.333 -> Average Cal sensor Temp.: 104.26 Ambient Temp.: 87.60 F
02:23:11.921 -> Average Cal sensor Temp.: 104.04 Ambient Temp.: 87.71 F
02:23:13.478 -> Average Cal sensor Temp.: 103.98 Ambient Temp.: 87.60 F
02:23:15.033 -> Average Cal sensor Temp.: 104.36 Ambient Temp.: 87.71 F
02:23:16.586 -> Average Cal sensor Temp.: 104.27 Ambient Temp.: 87.67 F
02:23:18.175 -> Average Cal sensor Temp.: 104.95 Ambient Temp.: 87.71 F
02:23:19.734 -> Average Cal sensor Temp.: 105.22 Ambient Temp.: 87.71 F
02:23:21.289 -> Average Cal sensor Temp.: 104.73 Ambient Temp.: 87.75 F
02:23:22.844 -> Average Cal sensor Temp.: 105.21 Ambient Temp.: 87.78 F
02:23:24.434 -> Average Cal sensor Temp.: 104.92 Ambient Temp.: 87.78 F
02:23:25.989 -> Average Cal sensor Temp.: 104.72 Ambient Temp.: 87.78 F
02:23:27.543 -> Average Cal sensor Temp.: 104.52 Ambient Temp.: 87.78 F
02:23:29.128 -> Average Cal sensor Temp.: 104.66 Ambient Temp.: 87.78 F
02:23:30.686 -> Average Cal sensor Temp.: 104.10 Ambient Temp.: 87.78 F
02:23:32.241 -> Average Cal sensor Temp.: 104.10 Ambient Temp.: 87.78 F
02:23:33.795 -> Average Cal sensor Temp.: 103.92 Ambient Temp.: 87.75 F
02:23:35.381 -> Average Cal sensor Temp.: 104.85 Ambient Temp.: 87.78 F
02:23:36.933 -> Average Cal sensor Temp.: 104.30 Ambient Temp.: 87.71 F
02:23:38.490 -> Average Cal sensor Temp.: 104.62 Ambient Temp.: 87.75 F
02:23:40.068 -> Average Cal sensor Temp.: 104.68 Ambient Temp.: 87.71 F
02:23:41.617 -> Average Cal sensor Temp.: 104.35 Ambient Temp.: 87.75 F
02:23:43.201 -> Average Cal sensor Temp.: 104.78 Ambient Temp.: 87.67 F
02:23:44.760 -> Average Cal sensor Temp.: 105.04 Ambient Temp.: 87.67 F
02:23:46.316 -> Average Cal sensor Temp.: 105.19 Ambient Temp.: 87.75 F
02:23:47.872 -> Average Cal sensor Temp.: 104.85 Ambient Temp.: 87.71 F
02:23:49.462 -> Average Cal sensor Temp.: 104.85 Ambient Temp.: 87.67 F
02:23:51.015 -> Average Cal sensor Temp.: 105.17 Ambient Temp.: 87.67 F
02:23:52.572 -> Average Cal sensor Temp.: 105.03 Ambient Temp.: 87.71 F
02:23:54.130 -> Average Cal sensor Temp.: 104.83 Ambient Temp.: 87.71 F
02:23:55.715 -> Average Cal sensor Temp.: 105.20 Ambient Temp.: 87.67 F
02:23:57.273 -> Average Cal sensor Temp.: 104.72 Ambient Temp.: 87.67 F
02:23:58.828 -> Average Cal sensor Temp.: 105.06 Ambient Temp.: 87.67 F
02:24:00.380 -> Average Cal sensor Temp.: 105.31 Ambient Temp.: 87.67 F
02:24:01.969 -> Average Cal sensor Temp.: 104.84 Ambient Temp.: 87.60 F
02:24:03.523 -> Average Cal sensor Temp.: 104.98 Ambient Temp.: 87.67 F
02:24:05.070 -> Average Cal sensor Temp.: 105.04 Ambient Temp.: 87.57 F
02:24:06.659 -> Average Cal sensor Temp.: 105.46 Ambient Temp.: 87.60 F
02:24:08.211 -> Average Cal sensor Temp.: 105.01 Ambient Temp.: 87.60 F
02:24:09.763 -> Average Cal sensor Temp.: 104.96 Ambient Temp.: 87.60 F
02:24:11.349 -> Average Cal sensor Temp.: 105.63 Ambient Temp.: 87.71 F
02:24:12.903 -> Average Cal sensor Temp.: 105.29 Ambient Temp.: 87.71 F
02:24:14.470 -> Average Cal sensor Temp.: 105.03 Ambient Temp.: 87.71 F
02:24:16.025 -> Average Cal sensor Temp.: 104.49 Ambient Temp.: 87.71 F
02:24:17.615 -> Average Cal sensor Temp.: 104.76 Ambient Temp.: 87.67 F
02:24:19.169 -> Average Cal sensor Temp.: 105.08 Ambient Temp.: 87.71 F
02:24:20.725 -> Average Cal sensor Temp.: 105.35 Ambient Temp.: 87.75 F
02:24:22.277 -> Average Cal sensor Temp.: 104.58 Ambient Temp.: 87.75 F
02:24:23.865 -> Average Cal sensor Temp.: 104.72 Ambient Temp.: 87.71 F
02:24:25.416 -> Average Cal sensor Temp.: 104.58 Ambient Temp.: 87.71 F
02:24:26.963 -> Average Cal sensor Temp.: 104.88 Ambient Temp.: 87.75 F
02:24:28.556 -> Average Cal sensor Temp.: 105.14 Ambient Temp.: 87.71 F
02:24:30.114 -> Average Cal sensor Temp.: 104.93 Ambient Temp.: 87.67 F
02:24:31.670 -> Average Cal sensor Temp.: 105.30 Ambient Temp.: 87.71 F
02:24:33.223 -> Average Cal sensor Temp.: 105.18 Ambient Temp.: 87.71 F
02:24:34.786 -> Average Cal sensor Temp.: 105.05 Ambient Temp.: 87.78 F
02:24:36.372 -> Average Cal sensor Temp.: 105.06 Ambient Temp.: 87.71 F
02:24:37.921 -> Average Cal sensor Temp.: 104.90 Ambient Temp.: 87.75 F
02:24:39.475 -> Average Cal sensor Temp.: 104.62 Ambient Temp.: 87.75 F
02:24:41.056 -> Average Cal sensor Temp.: 104.79 Ambient Temp.: 87.75 F
02:24:42.603 -> Average Cal sensor Temp.: 105.61 Ambient Temp.: 87.78 F
02:24:44.176 -> Average Cal sensor Temp.: 105.13 Ambient Temp.: 87.75 F
02:24:45.761 -> Average Cal sensor Temp.: 105.02 Ambient Temp.: 87.75 F
02:24:47.306 -> Average Cal sensor Temp.: 104.72 Ambient Temp.: 87.78 F
02:24:48.856 -> Average Cal sensor Temp.: 104.74 Ambient Temp.: 87.71 F
02:24:50.445 -> Average Cal sensor Temp.: 104.85 Ambient Temp.: 87.71 F
02:24:52.002 -> Average Cal sensor Temp.: 104.79 Ambient Temp.: 87.71 F
02:24:53.560 -> Average Cal sensor Temp.: 104.74 Ambient Temp.: 87.71 F
02:24:55.116 -> Average Cal sensor Temp.: 104.92 Ambient Temp.: 87.67 F
02:24:56.706 -> Average Cal sensor Temp.: 104.91 Ambient Temp.: 87.57 F
02:24:58.256 -> Average Cal sensor Temp.: 105.09 Ambient Temp.: 87.75 F
02:24:59.816 -> Average Cal sensor Temp.: 105.35 Ambient Temp.: 87.67 F
02:25:01.366 -> Average Cal sensor Temp.: 105.12 Ambient Temp.: 87.71 F
02:25:02.959 -> Average Cal sensor Temp.: 105.56 Ambient Temp.: 87.71 F
02:25:04.517 -> Average Cal sensor Temp.: 105.62 Ambient Temp.: 87.71 F
02:25:06.069 -> Average Cal sensor Temp.: 105.95 Ambient Temp.: 87.75 F
02:25:07.630 -> Average Cal sensor Temp.: 105.32 Ambient Temp.: 87.78 F
02:25:09.187 -> Average Cal sensor Temp.: 105.43 Ambient Temp.: 87.75 F
02:25:10.747 -> Average Cal sensor Temp.: 105.05 Ambient Temp.: 87.71 F
02:25:12.334 -> Average Cal sensor Temp.: 105.50 Ambient Temp.: 87.78 F
02:25:13.893 -> Average Cal sensor Temp.: 105.61 Ambient Temp.: 87.82 F
02:25:15.447 -> Average Cal sensor Temp.: 105.67 Ambient Temp.: 87.85 F
02:25:17.005 -> Average Cal sensor Temp.: 105.13 Ambient Temp.: 87.85 F
02:25:18.593 -> Average Cal sensor Temp.: 104.99 Ambient Temp.: 87.93 F
02:25:20.147 -> Average Cal sensor Temp.: 105.11 Ambient Temp.: 87.85 F
02:25:21.707 -> Average Cal sensor Temp.: 105.37 Ambient Temp.: 87.89 F
02:25:23.265 -> Average Cal sensor Temp.: 105.75 Ambient Temp.: 87.89 F
02:25:24.820 -> Average Cal sensor Temp.: 105.40 Ambient Temp.: 87.96 F
02:25:26.403 -> Average Cal sensor Temp.: 105.79 Ambient Temp.: 87.96 F
02:25:27.957 -> Average Cal sensor Temp.: 105.90 Ambient Temp.: 87.96 F
02:25:29.545 -> Average Cal sensor Temp.: 105.08 Ambient Temp.: 87.93 F
02:25:31.094 -> Average Cal sensor Temp.: 105.52 Ambient Temp.: 87.96 F
02:25:32.644 -> Average Cal sensor Temp.: 105.70 Ambient Temp.: 88.14 F
02:25:34.230 -> Average Cal sensor Temp.: 105.48 Ambient Temp.: 88.07 F
02:25:35.787 -> Average Cal sensor Temp.: 105.66 Ambient Temp.: 88.11 F
02:25:37.336 -> Average Cal sensor Temp.: 105.17 Ambient Temp.: 88.07 F
02:25:38.922 -> Average Cal sensor Temp.: 105.60 Ambient Temp.: 88.18 F
02:25:40.474 -> Average Cal sensor Temp.: 105.41 Ambient Temp.: 88.18 F
02:25:42.027 -> Average Cal sensor Temp.: 105.39 Ambient Temp.: 88.18 F
02:25:43.610 -> Average Cal sensor Temp.: 105.56 Ambient Temp.: 88.21 F
02:25:45.168 -> Average Cal sensor Temp.: 105.31 Ambient Temp.: 88.21 F
02:25:46.722 -> Average Cal sensor Temp.: 104.80 Ambient Temp.: 88.21 F
02:25:48.312 -> Average Cal sensor Temp.: 105.32 Ambient Temp.: 88.21 F
02:25:49.868 -> Average Cal sensor Temp.: 105.22 Ambient Temp.: 88.25 F
02:25:51.430 -> Average Cal sensor Temp.: 104.66 Ambient Temp.: 88.25 F
02:25:52.982 -> Average Cal sensor Temp.: 104.72 Ambient Temp.: 88.29 F
02:25:54.565 -> Average Cal sensor Temp.: 105.31 Ambient Temp.: 88.29 F
02:25:56.120 -> Average Cal sensor Temp.: 106.26 Ambient Temp.: 88.36 F
02:25:57.700 -> Average Cal sensor Temp.: 105.51 Ambient Temp.: 88.39 F
02:25:59.254 -> Average Cal sensor Temp.: 104.17 Ambient Temp.: 88.32 F
02:26:00.812 -> Average Cal sensor Temp.: 104.79 Ambient Temp.: 88.32 F
02:26:02.367 -> Average Cal sensor Temp.: 104.87 Ambient Temp.: 88.36 F
02:26:03.960 -> Average Cal sensor Temp.: 105.36 Ambient Temp.: 88.39 F
02:26:05.510 -> Average Cal sensor Temp.: 104.93 Ambient Temp.: 88.36 F
02:26:07.060 -> Average Cal sensor Temp.: 105.02 Ambient Temp.: 88.36 F
02:26:08.651 -> Average Cal sensor Temp.: 105.14 Ambient Temp.: 88.36 F
02:26:10.205 -> Average Cal sensor Temp.: 105.09 Ambient Temp.: 88.36 F
02:26:11.760 -> Average Cal sensor Temp.: 104.98 Ambient Temp.: 88.36 F
02:26:13.321 -> Average Cal sensor Temp.: 105.26 Ambient Temp.: 88.36 F
02:26:14.909 -> Average Cal sensor Temp.: 105.24 Ambient Temp.: 88.36 F
02:26:16.462 -> Average Cal sensor Temp.: 105.20 Ambient Temp.: 88.39 F
02:26:18.019 -> Average Cal sensor Temp.: 105.11 Ambient Temp.: 88.39 F
02:26:19.607 -> Average Cal sensor Temp.: 105.18 Ambient Temp.: 88.47 F
02:26:21.164 -> Average Cal sensor Temp.: 105.25 Ambient Temp.: 88.47 F
02:26:22.719 -> Average Cal sensor Temp.: 104.54 Ambient Temp.: 88.39 F
02:26:24.305 -> Average Cal sensor Temp.: 104.53 Ambient Temp.: 88.39 F
02:26:25.859 -> Average Cal sensor Temp.: 104.36 Ambient Temp.: 88.36 F
02:26:27.411 -> Average Cal sensor Temp.: 104.83 Ambient Temp.: 88.36 F
02:26:28.995 -> Average Cal sensor Temp.: 104.94 Ambient Temp.: 88.36 F
02:26:30.554 -> Average Cal sensor Temp.: 104.83 Ambient Temp.: 88.29 F
02:26:32.107 -> Average Cal sensor Temp.: 104.83 Ambient Temp.: 88.32 F
02:26:33.666 -> Average Cal sensor Temp.: 104.67 Ambient Temp.: 88.32 F
02:26:35.256 -> Average Cal sensor Temp.: 105.18 Ambient Temp.: 88.32 F
02:26:36.811 -> Average Cal sensor Temp.: 105.10 Ambient Temp.: 88.32 F
02:26:38.367 -> Average Cal sensor Temp.: 105.49 Ambient Temp.: 88.32 F
02:26:39.927 -> Average Cal sensor Temp.: 105.44 Ambient Temp.: 88.36 F
02:26:41.522 -> Average Cal sensor Temp.: 105.37 Ambient Temp.: 88.32 F
02:26:43.082 -> Average Cal sensor Temp.: 104.59 Ambient Temp.: 88.36 F
02:26:44.640 -> Average Cal sensor Temp.: 105.03 Ambient Temp.: 88.32 F
02:26:46.194 -> Average Cal sensor Temp.: 105.02 Ambient Temp.: 88.39 F
02:26:47.750 -> Average Cal sensor Temp.: 104.78 Ambient Temp.: 88.32 F
02:26:49.335 -> Average Cal sensor Temp.: 106.29 Ambient Temp.: 88.39 F
02:26:50.887 -> Average Cal sensor Temp.: 105.16 Ambient Temp.: 88.39 F
02:26:52.475 -> Average Cal sensor Temp.: 105.25 Ambient Temp.: 88.32 F
02:26:54.035 -> Average Cal sensor Temp.: 104.96 Ambient Temp.: 88.36 F
02:26:55.594 -> Average Cal sensor Temp.: 104.80 Ambient Temp.: 88.36 F
02:26:57.148 -> Average Cal sensor Temp.: 104.98 Ambient Temp.: 88.36 F
02:26:58.735 -> Average Cal sensor Temp.: 105.19 Ambient Temp.: 88.39 F
02:27:00.287 -> Average Cal sensor Temp.: 105.06 Ambient Temp.: 88.32 F
02:27:01.841 -> Average Cal sensor Temp.: 105.52 Ambient Temp.: 88.39 F
02:27:03.399 -> Average Cal sensor Temp.: 104.67 Ambient Temp.: 88.39 F
02:27:04.988 -> Average Cal sensor Temp.: 104.59 Ambient Temp.: 88.36 F
02:27:06.542 -> Average Cal sensor Temp.: 104.51 Ambient Temp.: 88.29 F
02:27:08.095 -> Average Cal sensor Temp.: 104.63 Ambient Temp.: 88.29 F
02:27:09.686 -> Average Cal sensor Temp.: 104.54 Ambient Temp.: 88.29 F
02:27:11.242 -> Average Cal sensor Temp.: 104.89 Ambient Temp.: 88.25 F
02:27:12.796 -> Average Cal sensor Temp.: 105.31 Ambient Temp.: 88.29 F
02:27:14.389 -> Average Cal sensor Temp.: 105.64 Ambient Temp.: 88.32 F
02:27:15.939 -> Average Cal sensor Temp.: 105.20 Ambient Temp.: 88.25 F
02:27:17.491 -> Average Cal sensor Temp.: 105.81 Ambient Temp.: 88.29 F
02:27:19.072 -> Average Cal sensor Temp.: 105.65 Ambient Temp.: 88.32 F
02:27:20.619 -> Average Cal sensor Temp.: 104.95 Ambient Temp.: 88.32 F
02:27:22.205 -> Average Cal sensor Temp.: 104.83 Ambient Temp.: 88.36 F
02:27:23.760 -> Average Cal sensor Temp.: 105.66 Ambient Temp.: 88.32 F
02:27:25.319 -> Average Cal sensor Temp.: 105.40 Ambient Temp.: 88.32 F
02:27:26.902 -> Average Cal sensor Temp.: 105.19 Ambient Temp.: 88.32 F
02:27:28.447 -> Average Cal sensor Temp.: 104.81 Ambient Temp.: 88.32 F
02:27:30.035 -> Average Cal sensor Temp.: 104.71 Ambient Temp.: 88.25 F
02:27:31.595 -> Average Cal sensor Temp.: 105.04 Ambient Temp.: 88.25 F
02:27:33.147 -> Average Cal sensor Temp.: 105.30 Ambient Temp.: 88.29 F
02:27:34.700 -> Average Cal sensor Temp.: 104.76 Ambient Temp.: 88.29 F
02:27:36.286 -> Average Cal sensor Temp.: 105.00 Ambient Temp.: 88.21 F
02:27:37.840 -> Average Cal sensor Temp.: 104.78 Ambient Temp.: 88.18 F
02:27:39.424 -> Average Cal sensor Temp.: 105.06 Ambient Temp.: 88.14 F
02:27:40.975 -> Average Cal sensor Temp.: 104.91 Ambient Temp.: 88.14 F
02:27:42.523 -> Average Cal sensor Temp.: 105.02 Ambient Temp.: 88.14 F
02:27:44.098 -> Average Cal sensor Temp.: 105.59 Ambient Temp.: 88.14 F
02:27:45.674 -> Average Cal sensor Temp.: 105.30 Ambient Temp.: 88.14 F
02:27:47.218 -> Average Cal sensor Temp.: 105.58 Ambient Temp.: 88.21 F
02:27:48.807 -> Average Cal sensor Temp.: 105.33 Ambient Temp.: 88.14 F
02:27:50.361 -> Average Cal sensor Temp.: 105.12 Ambient Temp.: 88.14 F
02:27:51.921 -> Average Cal sensor Temp.: 105.19 Ambient Temp.: 88.21 F
02:27:53.476 -> Average Cal sensor Temp.: 105.47 Ambient Temp.: 88.21 F
02:27:55.060 -> Average Cal sensor Temp.: 104.95 Ambient Temp.: 88.18 F
02:27:56.616 -> Average Cal sensor Temp.: 105.20 Ambient Temp.: 88.14 F
02:27:58.174 -> Average Cal sensor Temp.: 105.00 Ambient Temp.: 88.21 F
02:27:59.730 -> Average Cal sensor Temp.: 105.08 Ambient Temp.: 88.25 F
02:28:01.321 -> Average Cal sensor Temp.: 105.29 Ambient Temp.: 88.25 F
02:28:02.876 -> Average Cal sensor Temp.: 105.10 Ambient Temp.: 88.25 F
02:28:04.432 -> Average Cal sensor Temp.: 104.90 Ambient Temp.: 88.25 F
02:28:06.018 -> Average Cal sensor Temp.: 105.11 Ambient Temp.: 88.29 F
02:28:07.567 -> Average Cal sensor Temp.: 104.76 Ambient Temp.: 88.25 F
02:28:09.117 -> Average Cal sensor Temp.: 104.86 Ambient Temp.: 88.21 F
02:28:10.711 -> Average Cal sensor Temp.: 105.41 Ambient Temp.: 88.25 F
02:28:12.258 -> Average Cal sensor Temp.: 105.62 Ambient Temp.: 88.25 F
02:28:13.835 -> Average Cal sensor Temp.: 106.13 Ambient Temp.: 88.25 F
02:28:15.380 -> Average Cal sensor Temp.: 106.41 Ambient Temp.: 88.29 F
02:28:16.941 -> Average Cal sensor Temp.: 105.39 Ambient Temp.: 88.32 F
02:28:18.526 -> Average Cal sensor Temp.: 104.96 Ambient Temp.: 88.25 F
02:28:20.080 -> Average Cal sensor Temp.: 105.30 Ambient Temp.: 88.21 F
02:28:21.632 -> Average Cal sensor Temp.: 104.67 Ambient Temp.: 88.21 F
02:28:23.218 -> Average Cal sensor Temp.: 104.77 Ambient Temp.: 88.21 F
02:28:24.770 -> Average Cal sensor Temp.: 105.11 Ambient Temp.: 88.21 F
02:28:26.328 -> Average Cal sensor Temp.: 104.62 Ambient Temp.: 88.18 F
02:28:27.918 -> Average Cal sensor Temp.: 104.15 Ambient Temp.: 88.14 F
02:28:29.470 -> Average Cal sensor Temp.: 104.55 Ambient Temp.: 88.11 F
02:28:31.026 -> Average Cal sensor Temp.: 104.76 Ambient Temp.: 88.07 F
02:28:32.613 -> Average Cal sensor Temp.: 104.56 Ambient Temp.: 88.07 F