## SET 2
Reference Clinical Thermometer Readout | Reference DIGI Thermometer Readout | Reference Water Temperature (Sensor) | Calibrating contactless sensor Readout (2-3 cm) | Calibrating contactless sensor AMBIENT Readout | DIFF bw clinical and sensor | DIFF bw DIGI and sensor | DIFF BETWEEN CLINICAL AND DIGI | DIFF bw clinical and sensor | DIFF bw DIGI and sensor | DIFF BETWEEN CLINICAL AND DIGI | DIFF bwtween Water temp and clinical
Read at (deg F) |  | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F)
94 | 94 | 92.86 | 91.79 | 84.18 |  |  |  | 2.21 | 2.21 | 0 | 1.14
94.4 | 94.2 | 92.97 | 92.26 |  |  |  |  | 2.14 | 1.94 | 0.2 | 1.43
94.8 | 94.8 | 93.54 | 92.65 |  |  |  |  | 2.15 | 2.15 | 0 | 1.26
95 | 94.7 | 93.76 | 92.77 |  |  |  |  | 2.23 | 1.93 | 0.3 | 1.24
95.4 | 95.5 | 94.55 | 93.25 |  |  |  |  | 2.15 | 2.25 | -0.1 | 0.85
95.8 | 95.8 | 94.55 | 93.45 |  |  |  |  | 2.35 | 2.35 | 0 | 1.25
96 | 95.8 | 94.66 | 93.52 |  |  |  |  | 2.48 | 2.28 | 0.2 | 1.34
96.4 | 96.3 | 94.89 | 93.93 |  |  |  |  | 2.47 | 2.37 | 0.1 | 1.51
96.8 | 96.5 | 95 | 94.22 | 84.58 |  |  |  | 2.58 | 2.28 | 0.3 | 1.8
97 | 96.8 | 95.22 | 94.46 |  |  |  |  | 2.54 | 2.34 | 0.2 | 1.78
97.4 | 97.4 | 95.56 | 94.56 |  |  |  |  | 2.84 | 2.84 | 0 | 1.84
97.8 | 97.5 | 95.9 | 94.95 |  |  |  |  | 2.85 | 2.55 | 0.3 | 1.9
98 | 97.5 | 95.79 | 94.99 |  |  |  |  | 3.01 | 2.51 | 0.5 | 2.21
98.2 | 97.6 | 96.01 | 95.23 |  |  |  |  | 2.97 | 2.37 | 0.6 | 2.19
98.4 | 98.2 | 96.24 | 95.65 |  |  |  |  | 2.75 | 2.55 | 0.2 | 2.16
98.6 | 98.6 | 96.35 | 95.75 |  |  |  |  | 2.85 | 2.85 | 0 | 2.25
98.8 | 98.9 | 96.8 | 96.01 |  |  |  |  | 2.79 | 2.89 | -0.1 | 2
99 | 99 | 96.91 | 96.11 |  |  |  |  | 2.89 | 2.89 | 0 | 2.09
99.2 | 98.8 | 97.02 | 96.4 | 84.72 |  |  |  | 2.8 | 2.4 | 0.4 | 2.18
99.4 | 99.1 | 97.47 | 96.75 |  |  |  |  | 2.65 | 2.35 | 0.3 | 1.93
99.6 | 99.4 | 97.65 | 96.96 |  |  |  |  | 2.64 | 2.44 | 0.2 | 1.95
99.8 | 99.6 | 97.92 | 97.2 |  |  |  |  | 2.6 | 2.4 | 0.2 | 1.88
100 | 99.9 | 98.15 | 97.45 |  |  |  |  | 2.55 | 2.45 | 0.1 | 1.85
100.2 | 100 | 98.37 | 97.7 | 87.06 |  |  |  | 2.5 | 2.3 | 0.2 | 1.83
100.4 | 100.2 | 98.71 | 97.78 |  |  |  |  | 2.62 | 2.42 | 0.2 | 1.69
100.6 | 100.4 | 98.82 | 97.83 |  |  |  |  | 2.77 | 2.57 | 0.2 | 1.78
100.8 | 100.6 | 99.27 | 98.25 |  |  |  |  | 2.55 | 2.35 | 0.2 | 1.53
101 | 100.9 | 99.5 | 98.6 |  |  |  |  | 2.4 | 2.3 | 0.1 | 1.5
101.2 | 101.1 | 99.61 | 99.02 |  |  |  |  | 2.18 | 2.08 | 0.1 | 1.59
101.4 | 101.3 | 99.72 | 99.45 |  |  |  |  | 1.95 | 1.85 | 0.1 | 1.68
101.6 | 101.6 | 99.95 | 99.6 |  |  |  |  | 2 | 2 | 0 | 1.65
101.8 | 101.6 | 100.3 | 99.32 |  |  |  |  | 2.48 | 2.28 | 0.2 | 1.5
102 | 101.8 | 100.17 | 99.5 |  |  |  |  | 2.5 | 2.3 | 0.2 | 1.83
102.2 | 102 | 100.4 | 99.69 |  |  |  |  | 2.51 | 2.31 | 0.2 | 1.8
102.4 | 102.1 | 100.51 | 99.84 |  |  |  |  | 2.56 | 2.26 | 0.3 | 1.89
102.6 | 102.4 | 100.96 | 100.1 | 88.14 |  |  |  | 2.5 | 2.3 | 0.2 | 1.64
102.8 | 102.6 | 101.19 | 100.26 |  |  |  |  | 2.54 | 2.34 | 0.2 | 1.61
103 | 102.9 | 101.3 | 100.45 |  |  |  |  | 2.55 | 2.45 | 0.1 | 1.7
103.4 | 103.4 | 101.52 | 100.57 |  |  |  |  | 2.83 | 2.83 | 0 | 1.88
103.8 | 103.7 | 102.2 | 100.85 |  |  |  |  | 2.95 | 2.85 | 0.1 | 1.6
104 | 103.9 | 102.31 | 101.29 |  |  |  |  | 2.71 | 2.61 | 0.1 | 1.69
104.4 | 104.2 | 102.54 | 101.43 | 88.5 |  |  |  | 2.97 | 2.77 | 0.2 | 1.86
104.8 | 104.8 | 103.21 | 101.85 |  |  |  |  | 2.95 | 2.95 | 0 | 1.59
105 | 104.9 | 103.32 | 102.36 |  |  |  |  | 2.64 | 2.54 | 0.1 | 1.68
105.4 | 105.6 | 103.77 | 102.74 |  |  |  |  | 2.66 | 2.86 | -0.2 | 1.63
105.8 | 105.7 | 104.05 | 103.15 |  |  |  |  | 2.65 | 2.55 | 0.1 | 1.75
106 | 106 | 104.22 | 103.22 |  |  |  |  | 2.78 | 2.78 | 0 | 1.78
106.4 | 106.4 | 104.56 | 103.75 |  |  |  |  | 2.65 | 2.65 | 0 | 1.84
106.8 | 106.7 | 104.67 | 103.99 |  |  |  |  | 2.81 | 2.71 | 0.1 | 2.13
107 | 106.9 | 104.9 | 104.25 |  |  |  |  | 2.75 | 2.65 | 0.1 | 2.1
107.4 | 107.3 | 105.46 | 105.13 |  |  |  |  | 2.27 | 2.17 | 0.1 | 1.94
107.8 | 107.6 | 105.57 | 105.58 |  |  |  |  | 2.22 | 2.02 | 0.2 | 2.23
108 | 108 | 106.14 | 105.85 | 90 |  |  |  | 2.15 | 2.15 | 0 | 1.86

## RAW_HEATER-2
14:19:00.052 -> Water_Temp.:91.62, PWM:0
14:19:00.808 -> Water_Temp.:91.62, PWM:0
14:19:01.595 -> Water_Temp.:91.62, PWM:0
14:19:02.390 -> Water_Temp.:91.62, PWM:0
14:19:03.146 -> Water_Temp.:91.62, PWM:0
14:19:03.936 -> Water_Temp.:91.62, PWM:0
14:19:04.731 -> Water_Temp.:91.62, PWM:0
14:19:05.489 -> Water_Temp.:91.62, PWM:0
14:19:06.286 -> Water_Temp.:91.62, PWM:0
14:19:07.045 -> Water_Temp.:91.62, PWM:0
14:19:07.838 -> Water_Temp.:91.62, PWM:0
14:19:08.631 -> Water_Temp.:91.62, PWM:0
14:19:09.391 -> Water_Temp.:91.51, PWM:0
14:19:10.185 -> Water_Temp.:91.51, PWM:0
14:19:10.976 -> Water_Temp.:91.51, PWM:0
14:19:11.736 -> Water_Temp.:91.51, PWM:0
14:19:12.531 -> Water_Temp.:91.51, PWM:0
14:19:13.295 -> Water_Temp.:91.51, PWM:0
14:19:14.086 -> Water_Temp.:91.51, PWM:0
14:19:14.843 -> Water_Temp.:91.51, PWM:0
14:19:15.640 -> Water_Temp.:91.51, PWM:0
14:19:16.436 -> Water_Temp.:91.51, PWM:0
14:19:17.199 -> Water_Temp.:91.51, PWM:0
14:19:17.995 -> Water_Temp.:91.51, PWM:0
14:19:18.755 -> Water_Temp.:91.51, PWM:0
14:19:19.550 -> Water_Temp.:91.51, PWM:0
14:19:20.313 -> Water_Temp.:91.51, PWM:0
14:19:21.105 -> Water_Temp.:91.51, PWM:0
14:19:21.859 -> Water_Temp.:91.51, PWM:0
14:19:22.656 -> Water_Temp.:91.51, PWM:0
14:19:23.451 -> Water_Temp.:91.51, PWM:0
14:19:24.203 -> Water_Temp.:91.40, PWM:0
14:19:24.996 -> Water_Temp.:91.40, PWM:0
14:19:25.790 -> Water_Temp.:91.40, PWM:0
14:19:26.541 -> Water_Temp.:91.40, PWM:0
14:19:27.337 -> Water_Temp.:91.40, PWM:0
14:19:28.127 -> Water_Temp.:91.40, PWM:0
14:19:28.887 -> Water_Temp.:91.40, PWM:0
14:19:29.681 -> Water_Temp.:91.40, PWM:0
14:19:30.444 -> Water_Temp.:91.40, PWM:0
14:19:31.235 -> Water_Temp.:91.40, PWM:0
14:19:32.028 -> Water_Temp.:91.40, PWM:0
14:19:32.787 -> Water_Temp.:91.40, PWM:0
14:19:33.579 -> Water_Temp.:91.40, PWM:0
14:19:34.372 -> Water_Temp.:91.40, PWM:0
14:19:35.130 -> Water_Temp.:91.40, PWM:0
14:19:35.923 -> Water_Temp.:91.40, PWM:0
14:19:36.679 -> Water_Temp.:91.40, PWM:0
14:19:37.474 -> Water_Temp.:91.40, PWM:0
14:19:38.264 -> Water_Temp.:91.40, PWM:0
14:19:39.027 -> Water_Temp.:91.40, PWM:0
14:19:39.819 -> Water_Temp.:91.40, PWM:0
14:19:40.580 -> Water_Temp.:91.40, PWM:0
14:19:41.371 -> Water_Temp.:91.40, PWM:0
14:19:42.168 -> Water_Temp.:91.40, PWM:0
14:19:42.925 -> Water_Temp.:91.40, PWM:0
14:19:43.722 -> Water_Temp.:91.40, PWM:0
14:19:44.481 -> Water_Temp.:91.40, PWM:0
14:19:45.276 -> Water_Temp.:91.40, PWM:0
14:19:46.039 -> Water_Temp.:91.40, PWM:0
14:19:46.833 -> Water_Temp.:91.29, PWM:0
14:19:47.630 -> Water_Temp.:91.29, PWM:0
14:19:48.390 -> Water_Temp.:91.29, PWM:0
14:19:49.184 -> Water_Temp.:91.29, PWM:0
14:19:49.946 -> Water_Temp.:91.29, PWM:0
14:19:50.741 -> Water_Temp.:91.29, PWM:0
14:19:51.499 -> Water_Temp.:91.29, PWM:0
14:19:52.291 -> Water_Temp.:91.29, PWM:0
14:19:53.055 -> Water_Temp.:91.29, PWM:0
14:19:53.853 -> Water_Temp.:91.29, PWM:0
14:19:54.645 -> Water_Temp.:91.29, PWM:0
14:19:55.408 -> Water_Temp.:91.29, PWM:0
14:19:56.205 -> Water_Temp.:91.29, PWM:0
14:19:56.961 -> Water_Temp.:91.29, PWM:8
14:19:57.755 -> Water_Temp.:91.29, PWM:13
14:19:58.544 -> Water_Temp.:91.29, PWM:14
14:19:59.303 -> Water_Temp.:91.29, PWM:14
14:20:00.855 -> Water_Temp.:91.29, PWM:13
14:20:02.440 -> Water_Temp.:91.29, PWM:13
14:20:03.985 -> Water_Temp.:91.18, PWM:14
14:20:04.779 -> Water_Temp.:91.18, PWM:13
14:20:05.538 -> Water_Temp.:91.18, PWM:16
14:20:06.334 -> Water_Temp.:91.18, PWM:20
14:20:07.884 -> Water_Temp.:91.18, PWM:20
14:20:08.681 -> Water_Temp.:91.18, PWM:20
14:20:09.444 -> Water_Temp.:91.18, PWM:20
14:20:10.239 -> Water_Temp.:91.18, PWM:20
14:20:11.787 -> Water_Temp.:91.18, PWM:20
14:20:13.346 -> Water_Temp.:91.18, PWM:20
14:20:14.102 -> Water_Temp.:91.06, PWM:20
14:20:15.654 -> Water_Temp.:91.06, PWM:20
14:20:16.407 -> Water_Temp.:91.06, PWM:20
14:20:17.202 -> Water_Temp.:91.06, PWM:20
14:20:17.994 -> Water_Temp.:91.06, PWM:20
14:20:19.516 -> Water_Temp.:91.06, PWM:20
14:20:20.311 -> Water_Temp.:91.06, PWM:20
14:20:21.105 -> Water_Temp.:91.06, PWM:20
14:20:21.868 -> Water_Temp.:91.06, PWM:20
14:20:23.426 -> Water_Temp.:91.06, PWM:20
14:20:24.221 -> Water_Temp.:91.06, PWM:20
14:20:24.981 -> Water_Temp.:90.95, PWM:20
14:20:26.530 -> Water_Temp.:90.95, PWM:20
14:20:28.844 -> Water_Temp.:90.95, PWM:21
14:20:30.403 -> Water_Temp.:90.95, PWM:20
14:20:31.201 -> Water_Temp.:90.95, PWM:20
14:20:31.956 -> Water_Temp.:90.95, PWM:20
14:20:33.513 -> Water_Temp.:90.95, PWM:20
14:20:35.099 -> Water_Temp.:90.95, PWM:20
14:20:36.617 -> Water_Temp.:90.95, PWM:20
14:20:37.412 -> Water_Temp.:90.95, PWM:20
14:20:38.172 -> Water_Temp.:90.95, PWM:20
14:20:38.969 -> Water_Temp.:90.95, PWM:20
14:20:39.728 -> Water_Temp.:90.95, PWM:20
14:20:40.522 -> Water_Temp.:90.84, PWM:20
14:20:41.319 -> Water_Temp.:90.95, PWM:20
14:20:42.081 -> Water_Temp.:90.84, PWM:20
14:20:42.879 -> Water_Temp.:90.84, PWM:20
14:20:43.641 -> Water_Temp.:90.84, PWM:20
14:20:44.436 -> Water_Temp.:90.84, PWM:20
14:20:46.709 -> Water_Temp.:90.84, PWM:20
14:20:47.500 -> Water_Temp.:90.84, PWM:20
14:20:48.264 -> Water_Temp.:90.84, PWM:20
14:20:49.055 -> Water_Temp.:90.84, PWM:20
14:20:49.849 -> Water_Temp.:90.84, PWM:20
14:20:50.607 -> Water_Temp.:90.84, PWM:20
14:20:51.399 -> Water_Temp.:90.84, PWM:20
14:20:52.193 -> Water_Temp.:90.84, PWM:20
14:20:53.712 -> Water_Temp.:90.84, PWM:20
14:20:54.508 -> Water_Temp.:90.84, PWM:20
14:20:55.268 -> Water_Temp.:90.84, PWM:20
14:20:56.059 -> Water_Temp.:90.84, PWM:20
14:20:56.854 -> Water_Temp.:90.72, PWM:20
14:20:57.612 -> Water_Temp.:90.72, PWM:20
14:20:58.405 -> Water_Temp.:90.84, PWM:20
14:20:59.167 -> Water_Temp.:90.84, PWM:20
14:21:00.719 -> Water_Temp.:90.72, PWM:20
14:21:03.070 -> Water_Temp.:90.72, PWM:20
14:21:06.106 -> Water_Temp.:90.72, PWM:20
14:21:06.905 -> Water_Temp.:90.72, PWM:20
14:21:10.010 -> Water_Temp.:90.61, PWM:20
14:21:10.803 -> Water_Temp.:90.61, PWM:20
14:21:11.563 -> Water_Temp.:90.61, PWM:20
14:21:12.354 -> Water_Temp.:90.61, PWM:20
14:21:13.908 -> Water_Temp.:90.61, PWM:20
14:21:14.704 -> Water_Temp.:90.61, PWM:28
14:21:15.468 -> Water_Temp.:90.61, PWM:28
14:21:17.024 -> Water_Temp.:90.61, PWM:28
14:21:17.784 -> Water_Temp.:90.61, PWM:29
14:21:18.579 -> Water_Temp.:90.61, PWM:29
14:21:20.884 -> Water_Temp.:90.61, PWM:30
14:21:21.646 -> Water_Temp.:90.61, PWM:30
14:21:22.441 -> Water_Temp.:90.50, PWM:30
14:21:24.753 -> Water_Temp.:90.50, PWM:31
14:21:26.305 -> Water_Temp.:90.50, PWM:30
14:21:27.067 -> Water_Temp.:90.50, PWM:31
14:21:27.858 -> Water_Temp.:90.50, PWM:30
14:21:28.623 -> Water_Temp.:90.50, PWM:33
14:21:30.185 -> Water_Temp.:90.50, PWM:33
14:21:30.942 -> Water_Temp.:90.50, PWM:33
14:21:34.811 -> Water_Temp.:90.50, PWM:33
14:21:36.366 -> Water_Temp.:90.50, PWM:33
14:21:37.123 -> Water_Temp.:90.50, PWM:33
14:21:39.437 -> Water_Temp.:90.50, PWM:33
14:21:40.198 -> Water_Temp.:90.50, PWM:33
14:21:40.994 -> Water_Temp.:90.50, PWM:33
14:21:41.754 -> Water_Temp.:90.50, PWM:33
14:21:42.553 -> Water_Temp.:90.50, PWM:33
14:21:44.107 -> Water_Temp.:90.50, PWM:33
14:21:45.630 -> Water_Temp.:90.50, PWM:33
14:21:47.153 -> Water_Temp.:90.50, PWM:33
14:21:47.947 -> Water_Temp.:90.50, PWM:33
14:21:49.497 -> Water_Temp.:90.50, PWM:33
14:21:50.297 -> Water_Temp.:90.50, PWM:33
14:21:51.057 -> Water_Temp.:90.50, PWM:33
14:21:51.853 -> Water_Temp.:90.50, PWM:33
14:21:53.408 -> Water_Temp.:90.50, PWM:33
14:21:55.717 -> Water_Temp.:90.50, PWM:33
14:21:56.508 -> Water_Temp.:90.50, PWM:33
14:21:58.064 -> Water_Temp.:90.50, PWM:33
14:21:59.588 -> Water_Temp.:90.50, PWM:33
14:22:00.379 -> Water_Temp.:90.50, PWM:33
14:22:01.928 -> Water_Temp.:90.50, PWM:33
14:22:02.723 -> Water_Temp.:90.50, PWM:33
14:22:03.483 -> Water_Temp.:90.39, PWM:33
14:22:05.069 -> Water_Temp.:90.50, PWM:33
14:22:05.829 -> Water_Temp.:90.39, PWM:33
14:22:06.625 -> Water_Temp.:90.50, PWM:33
14:22:08.148 -> Water_Temp.:90.50, PWM:33
14:22:08.946 -> Water_Temp.:90.50, PWM:33
14:22:11.259 -> Water_Temp.:90.50, PWM:33
14:22:12.836 -> Water_Temp.:90.50, PWM:33
14:22:13.595 -> Water_Temp.:90.50, PWM:33
14:22:14.388 -> Water_Temp.:90.50, PWM:33
14:22:15.908 -> Water_Temp.:90.50, PWM:33
14:22:17.497 -> Water_Temp.:90.50, PWM:33
14:22:18.260 -> Water_Temp.:90.50, PWM:33
14:22:19.053 -> Water_Temp.:90.50, PWM:33
14:22:19.818 -> Water_Temp.:90.50, PWM:33
14:22:20.616 -> Water_Temp.:90.50, PWM:33
14:22:21.374 -> Water_Temp.:90.50, PWM:33
14:22:22.930 -> Water_Temp.:90.50, PWM:33
14:22:24.476 -> Water_Temp.:90.61, PWM:33
14:22:25.238 -> Water_Temp.:90.50, PWM:33
14:22:27.556 -> Water_Temp.:90.61, PWM:33
14:22:29.108 -> Water_Temp.:90.61, PWM:33
14:22:29.899 -> Water_Temp.:90.61, PWM:33
14:22:30.659 -> Water_Temp.:90.61, PWM:33
14:22:31.452 -> Water_Temp.:90.61, PWM:33
14:22:33.733 -> Water_Temp.:90.61, PWM:33
14:22:34.526 -> Water_Temp.:90.61, PWM:33
14:22:35.287 -> Water_Temp.:90.61, PWM:33
14:22:36.081 -> Water_Temp.:90.61, PWM:33
14:22:37.603 -> Water_Temp.:90.61, PWM:33
14:22:38.398 -> Water_Temp.:90.61, PWM:33
14:22:39.951 -> Water_Temp.:90.61, PWM:33
14:22:40.714 -> Water_Temp.:90.61, PWM:33
14:22:42.266 -> Water_Temp.:90.72, PWM:33
14:22:43.817 -> Water_Temp.:90.72, PWM:33
14:22:44.579 -> Water_Temp.:90.72, PWM:33
14:22:45.373 -> Water_Temp.:90.72, PWM:33
14:22:46.929 -> Water_Temp.:90.61, PWM:33
14:22:47.690 -> Water_Temp.:90.72, PWM:33
14:22:49.244 -> Water_Temp.:90.72, PWM:33
14:22:50.036 -> Water_Temp.:90.72, PWM:33
14:22:51.590 -> Water_Temp.:90.72, PWM:33
14:22:52.385 -> Water_Temp.:90.72, PWM:33
14:22:53.141 -> Water_Temp.:90.72, PWM:33
14:22:53.929 -> Water_Temp.:90.72, PWM:33
14:22:55.482 -> Water_Temp.:90.72, PWM:33
14:22:56.273 -> Water_Temp.:90.72, PWM:33
14:22:57.030 -> Water_Temp.:90.72, PWM:33
14:22:57.821 -> Water_Temp.:90.72, PWM:33
14:23:00.896 -> Water_Temp.:90.72, PWM:33
14:23:01.656 -> Water_Temp.:90.72, PWM:33
14:23:02.450 -> Water_Temp.:90.72, PWM:33
14:23:03.210 -> Water_Temp.:90.72, PWM:33
14:23:04.005 -> Water_Temp.:90.72, PWM:33
14:23:04.767 -> Water_Temp.:90.72, PWM:33
14:23:06.358 -> Water_Temp.:90.72, PWM:33
14:23:08.675 -> Water_Temp.:90.72, PWM:33
14:23:12.510 -> Water_Temp.:90.72, PWM:33
14:23:14.854 -> Water_Temp.:90.72, PWM:33
14:23:15.613 -> Water_Temp.:90.72, PWM:33
14:23:16.408 -> Water_Temp.:90.72, PWM:33
14:23:17.169 -> Water_Temp.:90.72, PWM:33
14:23:18.727 -> Water_Temp.:90.72, PWM:33
14:23:19.522 -> Water_Temp.:90.72, PWM:33
14:23:21.038 -> Water_Temp.:90.72, PWM:33
14:23:21.833 -> Water_Temp.:90.72, PWM:33
14:23:22.597 -> Water_Temp.:90.72, PWM:33
14:23:23.396 -> Water_Temp.:90.72, PWM:33
14:23:24.159 -> Water_Temp.:90.72, PWM:33
14:23:24.948 -> Water_Temp.:90.72, PWM:33
14:23:26.499 -> Water_Temp.:90.72, PWM:33
14:23:28.785 -> Water_Temp.:90.72, PWM:33
14:23:29.574 -> Water_Temp.:90.72, PWM:33
14:23:30.365 -> Water_Temp.:90.72, PWM:33
14:23:31.125 -> Water_Temp.:90.72, PWM:33
14:23:31.919 -> Water_Temp.:90.72, PWM:33
14:23:32.677 -> Water_Temp.:90.72, PWM:33
14:23:33.468 -> Water_Temp.:90.72, PWM:33
14:23:34.259 -> Water_Temp.:90.72, PWM:33
14:23:35.815 -> Water_Temp.:90.72, PWM:33
14:23:38.126 -> Water_Temp.:90.84, PWM:33
14:23:38.889 -> Water_Temp.:90.84, PWM:33
14:23:39.679 -> Water_Temp.:90.84, PWM:33
14:23:40.474 -> Water_Temp.:90.84, PWM:33
14:23:41.234 -> Water_Temp.:90.84, PWM:33
14:23:44.279 -> Water_Temp.:90.84, PWM:33
14:23:45.833 -> Water_Temp.:90.84, PWM:33
14:23:46.630 -> Water_Temp.:90.84, PWM:33
14:23:47.389 -> Water_Temp.:90.84, PWM:33
14:23:48.180 -> Water_Temp.:90.84, PWM:33
14:23:48.975 -> Water_Temp.:90.84, PWM:33
14:23:49.730 -> Water_Temp.:90.84, PWM:33
14:23:50.529 -> Water_Temp.:90.84, PWM:33
14:23:53.600 -> Water_Temp.:90.84, PWM:33
14:23:54.358 -> Water_Temp.:90.84, PWM:33
14:23:55.151 -> Water_Temp.:90.84, PWM:33
14:23:55.947 -> Water_Temp.:90.84, PWM:33
14:23:56.706 -> Water_Temp.:90.84, PWM:33
14:23:57.498 -> Water_Temp.:90.84, PWM:33
14:23:59.808 -> Water_Temp.:90.84, PWM:33
14:24:02.883 -> Water_Temp.:90.84, PWM:33
14:24:03.675 -> Water_Temp.:90.84, PWM:33
14:24:04.431 -> Water_Temp.:90.84, PWM:33
14:24:05.986 -> Water_Temp.:90.84, PWM:33
14:24:07.504 -> Water_Temp.:90.84, PWM:33
14:24:08.300 -> Water_Temp.:90.84, PWM:33
14:24:09.091 -> Water_Temp.:90.84, PWM:33
14:24:09.849 -> Water_Temp.:90.95, PWM:33
14:24:10.645 -> Water_Temp.:90.95, PWM:33
14:24:11.401 -> Water_Temp.:90.95, PWM:33
14:24:12.189 -> Water_Temp.:90.95, PWM:33
14:24:12.986 -> Water_Temp.:90.95, PWM:33
14:24:13.741 -> Water_Temp.:90.95, PWM:33
14:24:15.298 -> Water_Temp.:90.95, PWM:33
14:24:16.094 -> Water_Temp.:90.95, PWM:33
14:24:16.891 -> Water_Temp.:90.95, PWM:33
14:24:17.654 -> Water_Temp.:90.95, PWM:33
14:24:18.416 -> Water_Temp.:90.95, PWM:33
14:24:19.212 -> Water_Temp.:90.95, PWM:33
14:24:19.975 -> Water_Temp.:90.95, PWM:33
14:24:21.537 -> Water_Temp.:91.06, PWM:33
14:24:22.327 -> Water_Temp.:91.06, PWM:33
14:24:24.635 -> Water_Temp.:91.06, PWM:33
14:24:26.951 -> Water_Temp.:91.06, PWM:33
14:24:29.268 -> Water_Temp.:91.06, PWM:33
14:24:30.825 -> Water_Temp.:91.06, PWM:33
14:24:32.380 -> Water_Temp.:91.06, PWM:33
14:24:33.172 -> Water_Temp.:91.06, PWM:33
14:24:37.038 -> Water_Temp.:91.06, PWM:33
14:24:37.796 -> Water_Temp.:91.06, PWM:33
14:24:39.348 -> Water_Temp.:91.06, PWM:33
14:24:40.142 -> Water_Temp.:91.06, PWM:33
14:24:42.457 -> Water_Temp.:91.06, PWM:33
14:24:43.249 -> Water_Temp.:91.06, PWM:33
14:24:44.045 -> Water_Temp.:90.95, PWM:33
14:24:44.808 -> Water_Temp.:91.06, PWM:33
14:24:45.604 -> Water_Temp.:91.06, PWM:33
14:24:47.127 -> Water_Temp.:91.06, PWM:33
14:24:49.438 -> Water_Temp.:91.06, PWM:33
14:24:50.197 -> Water_Temp.:91.06, PWM:33
14:24:50.990 -> Water_Temp.:91.06, PWM:33
14:24:51.784 -> Water_Temp.:91.06, PWM:33
14:24:52.541 -> Water_Temp.:91.06, PWM:33
14:24:53.332 -> Water_Temp.:91.06, PWM:33
14:24:54.126 -> Water_Temp.:91.06, PWM:33
14:24:54.887 -> Water_Temp.:91.06, PWM:33
14:24:57.195 -> Water_Temp.:91.18, PWM:33
14:24:57.958 -> Water_Temp.:91.18, PWM:33
14:24:58.751 -> Water_Temp.:91.18, PWM:33
14:24:59.544 -> Water_Temp.:91.18, PWM:33
14:25:01.099 -> Water_Temp.:91.18, PWM:33
14:25:03.419 -> Water_Temp.:91.18, PWM:33
14:25:04.961 -> Water_Temp.:91.18, PWM:33
14:25:06.514 -> Water_Temp.:91.18, PWM:33
14:25:07.309 -> Water_Temp.:91.18, PWM:33
14:25:08.069 -> Water_Temp.:91.18, PWM:33
14:25:08.860 -> Water_Temp.:91.18, PWM:33
14:25:10.421 -> Water_Temp.:91.18, PWM:33
14:25:11.179 -> Water_Temp.:91.18, PWM:33
14:25:13.487 -> Water_Temp.:91.18, PWM:33
14:25:14.279 -> Water_Temp.:91.18, PWM:33
14:25:16.559 -> Water_Temp.:91.18, PWM:33
14:25:18.114 -> Water_Temp.:91.18, PWM:33
14:25:18.908 -> Water_Temp.:91.18, PWM:33
14:25:19.669 -> Water_Temp.:91.18, PWM:33
14:25:20.468 -> Water_Temp.:91.18, PWM:33
14:25:21.228 -> Water_Temp.:91.18, PWM:33
14:25:22.025 -> Water_Temp.:91.18, PWM:33
14:25:22.786 -> Water_Temp.:91.29, PWM:33
14:25:23.580 -> Water_Temp.:91.29, PWM:33
14:25:24.335 -> Water_Temp.:91.29, PWM:33
14:25:25.127 -> Water_Temp.:91.29, PWM:33
14:25:25.922 -> Water_Temp.:91.29, PWM:33
14:25:26.683 -> Water_Temp.:91.29, PWM:33
14:25:27.477 -> Water_Temp.:91.29, PWM:33
14:25:28.236 -> Water_Temp.:91.29, PWM:33
14:25:29.028 -> Water_Temp.:91.29, PWM:33
14:25:29.822 -> Water_Temp.:91.29, PWM:33
14:25:30.585 -> Water_Temp.:91.29, PWM:33
14:25:31.382 -> Water_Temp.:91.29, PWM:33
14:25:32.141 -> Water_Temp.:91.29, PWM:33
14:25:32.935 -> Water_Temp.:91.29, PWM:33
14:25:33.695 -> Water_Temp.:91.29, PWM:33
14:25:34.480 -> Water_Temp.:91.29, PWM:33
14:25:36.033 -> Water_Temp.:91.29, PWM:33
14:25:38.381 -> Water_Temp.:92.97, PWM:33
14:25:40.701 -> Water_Temp.:92.19, PWM:33
14:25:41.495 -> Water_Temp.:91.96, PWM:33
14:25:42.257 -> Water_Temp.:91.85, PWM:33
14:25:43.050 -> Water_Temp.:91.74, PWM:33
14:25:45.361 -> Water_Temp.:91.62, PWM:33
14:25:46.120 -> Water_Temp.:91.62, PWM:33
14:25:46.915 -> Water_Temp.:91.51, PWM:33
14:25:51.533 -> Water_Temp.:91.51, PWM:33
14:25:53.084 -> Water_Temp.:91.51, PWM:33
14:25:55.435 -> Water_Temp.:91.40, PWM:33
14:25:56.233 -> Water_Temp.:91.51, PWM:33
14:25:56.991 -> Water_Temp.:91.40, PWM:33
14:25:57.788 -> Water_Temp.:91.51, PWM:33
14:25:58.549 -> Water_Temp.:91.51, PWM:33
14:25:59.342 -> Water_Temp.:91.51, PWM:33
14:26:00.103 -> Water_Temp.:91.51, PWM:33
14:26:00.897 -> Water_Temp.:91.51, PWM:33
14:26:03.969 -> Water_Temp.:91.51, PWM:33
14:26:04.729 -> Water_Temp.:91.51, PWM:33
14:26:06.280 -> Water_Temp.:91.51, PWM:33
14:26:07.072 -> Water_Temp.:91.51, PWM:33
14:26:07.826 -> Water_Temp.:91.51, PWM:33
14:26:08.624 -> Water_Temp.:91.51, PWM:33
14:26:10.181 -> Water_Temp.:91.51, PWM:33
14:26:11.704 -> Water_Temp.:91.51, PWM:33
14:26:12.496 -> Water_Temp.:91.51, PWM:33
14:26:13.291 -> Water_Temp.:91.51, PWM:33
14:26:15.574 -> Water_Temp.:91.51, PWM:33
14:26:17.883 -> Water_Temp.:91.51, PWM:33
14:26:18.681 -> Water_Temp.:91.51, PWM:33
14:26:19.442 -> Water_Temp.:91.51, PWM:33
14:26:20.239 -> Water_Temp.:91.51, PWM:33
14:26:21.001 -> Water_Temp.:91.51, PWM:33
14:26:22.557 -> Water_Temp.:91.51, PWM:33
14:26:23.349 -> Water_Temp.:91.51, PWM:33
14:26:24.146 -> Water_Temp.:91.51, PWM:33
14:26:24.906 -> Water_Temp.:91.51, PWM:33
14:26:26.462 -> Water_Temp.:91.51, PWM:33
14:26:27.257 -> Water_Temp.:91.51, PWM:33
14:26:28.772 -> Water_Temp.:91.51, PWM:33
14:26:31.095 -> Water_Temp.:91.51, PWM:33
14:26:31.886 -> Water_Temp.:91.51, PWM:33
14:26:32.645 -> Water_Temp.:91.51, PWM:33
14:26:33.438 -> Water_Temp.:91.51, PWM:33
14:26:34.230 -> Water_Temp.:91.51, PWM:33
14:26:35.778 -> Water_Temp.:91.51, PWM:33
14:26:36.572 -> Water_Temp.:91.51, PWM:33
14:26:37.333 -> Water_Temp.:91.62, PWM:33
14:26:38.125 -> Water_Temp.:91.62, PWM:33
14:26:38.887 -> Water_Temp.:91.62, PWM:33
14:26:39.674 -> Water_Temp.:91.62, PWM:33
14:26:40.473 -> Water_Temp.:91.62, PWM:33
14:26:41.992 -> Water_Temp.:91.62, PWM:33
14:26:42.786 -> Water_Temp.:91.62, PWM:33
14:26:43.547 -> Water_Temp.:91.74, PWM:33
14:26:44.342 -> Water_Temp.:91.74, PWM:33
14:26:45.093 -> Water_Temp.:91.74, PWM:33
14:26:46.644 -> Water_Temp.:91.74, PWM:33
14:26:47.436 -> Water_Temp.:91.74, PWM:33
14:26:48.196 -> Water_Temp.:91.74, PWM:33
14:26:48.988 -> Water_Temp.:91.74, PWM:33
14:26:49.782 -> Water_Temp.:91.74, PWM:33
14:26:51.308 -> Water_Temp.:91.74, PWM:33
14:26:53.629 -> Water_Temp.:91.74, PWM:33
14:26:55.145 -> Water_Temp.:91.74, PWM:33
14:26:56.733 -> Water_Temp.:91.74, PWM:33
14:26:57.491 -> Water_Temp.:91.74, PWM:33
14:26:59.047 -> Water_Temp.:91.74, PWM:33
14:26:59.841 -> Water_Temp.:93.43, PWM:33
14:27:00.636 -> Water_Temp.:95.56, PWM:33
14:27:02.193 -> Water_Temp.:92.64, PWM:33
14:27:02.952 -> Water_Temp.:92.41, PWM:33
14:27:05.296 -> Water_Temp.:92.07, PWM:33
14:27:06.817 -> Water_Temp.:91.96, PWM:33
14:27:07.611 -> Water_Temp.:91.96, PWM:33
14:27:08.371 -> Water_Temp.:91.96, PWM:33
14:27:09.162 -> Water_Temp.:91.85, PWM:33
14:27:10.715 -> Water_Temp.:91.85, PWM:33
14:27:12.273 -> Water_Temp.:91.85, PWM:33
14:27:13.032 -> Water_Temp.:91.85, PWM:33
14:27:14.582 -> Water_Temp.:91.85, PWM:33
14:27:15.378 -> Water_Temp.:91.85, PWM:33
14:27:17.701 -> Water_Temp.:91.85, PWM:33
14:27:20.017 -> Water_Temp.:91.85, PWM:33
14:27:23.106 -> Water_Temp.:91.96, PWM:33
14:27:24.661 -> Water_Temp.:91.96, PWM:33
14:27:25.455 -> Water_Temp.:91.96, PWM:33
14:27:27.013 -> Water_Temp.:91.96, PWM:33
14:27:27.774 -> Water_Temp.:91.96, PWM:33
14:27:28.566 -> Water_Temp.:91.96, PWM:33
14:27:30.123 -> Water_Temp.:91.96, PWM:33
14:27:30.918 -> Water_Temp.:91.96, PWM:33
14:27:31.679 -> Water_Temp.:91.96, PWM:33
14:27:32.473 -> Water_Temp.:91.96, PWM:33
14:27:34.026 -> Water_Temp.:91.96, PWM:33
14:27:34.787 -> Water_Temp.:91.85, PWM:33
14:27:37.136 -> Water_Temp.:91.85, PWM:33
14:27:38.689 -> Water_Temp.:91.85, PWM:33
14:27:39.485 -> Water_Temp.:91.85, PWM:33
14:27:41.006 -> Water_Temp.:91.85, PWM:33
14:27:42.595 -> Water_Temp.:91.85, PWM:33
14:27:43.355 -> Water_Temp.:91.85, PWM:33
14:27:44.151 -> Water_Temp.:91.85, PWM:33
14:27:44.907 -> Water_Temp.:91.85, PWM:33
14:27:45.701 -> Water_Temp.:91.85, PWM:33
14:27:46.495 -> Water_Temp.:91.85, PWM:33
14:27:47.256 -> Water_Temp.:91.85, PWM:33
14:27:48.811 -> Water_Temp.:91.85, PWM:33
14:27:49.568 -> Water_Temp.:91.85, PWM:33
14:27:50.362 -> Water_Temp.:91.85, PWM:33
14:27:51.155 -> Water_Temp.:91.85, PWM:33
14:27:52.676 -> Water_Temp.:91.85, PWM:33
14:27:53.468 -> Water_Temp.:91.85, PWM:33
14:27:55.020 -> Water_Temp.:91.85, PWM:33
14:27:55.815 -> Water_Temp.:91.85, PWM:33
14:27:56.578 -> Water_Temp.:91.96, PWM:33
14:27:58.136 -> Water_Temp.:91.96, PWM:33
14:27:58.923 -> Water_Temp.:91.96, PWM:33
14:28:01.235 -> Water_Temp.:91.96, PWM:33
14:28:02.029 -> Water_Temp.:91.96, PWM:33
14:28:02.788 -> Water_Temp.:91.96, PWM:33
14:28:05.110 -> Water_Temp.:91.96, PWM:33
14:28:07.426 -> Water_Temp.:91.96, PWM:33
14:28:08.221 -> Water_Temp.:91.96, PWM:33
14:28:10.541 -> Water_Temp.:91.96, PWM:33
14:28:11.299 -> Water_Temp.:91.85, PWM:33
14:28:12.091 -> Water_Temp.:91.85, PWM:33
14:28:12.857 -> Water_Temp.:91.85, PWM:33
14:28:13.654 -> Water_Temp.:91.85, PWM:33
14:28:14.413 -> Water_Temp.:91.96, PWM:33
14:28:15.206 -> Water_Temp.:91.96, PWM:33
14:28:16.000 -> Water_Temp.:91.96, PWM:33
14:28:16.758 -> Water_Temp.:91.96, PWM:33
14:28:17.551 -> Water_Temp.:91.96, PWM:33
14:28:19.104 -> Water_Temp.:91.96, PWM:33
14:28:20.659 -> Water_Temp.:91.96, PWM:33
14:28:21.421 -> Water_Temp.:91.96, PWM:33
14:28:22.217 -> Water_Temp.:91.85, PWM:33
14:28:22.980 -> Water_Temp.:91.96, PWM:33
14:28:23.773 -> Water_Temp.:91.85, PWM:33
14:28:24.565 -> Water_Temp.:91.85, PWM:33
14:28:25.329 -> Water_Temp.:91.85, PWM:33
14:28:26.848 -> Water_Temp.:91.96, PWM:33
14:28:27.643 -> Water_Temp.:91.96, PWM:33
14:28:28.405 -> Water_Temp.:91.96, PWM:33
14:28:29.205 -> Water_Temp.:91.96, PWM:33
14:28:30.762 -> Water_Temp.:91.96, PWM:33
14:28:32.319 -> Water_Temp.:91.96, PWM:33
14:28:33.833 -> Water_Temp.:91.96, PWM:33
14:28:35.391 -> Water_Temp.:91.96, PWM:33
14:28:36.154 -> Water_Temp.:91.96, PWM:33
14:28:37.711 -> Water_Temp.:92.07, PWM:33
14:28:38.504 -> Water_Temp.:91.96, PWM:33
14:28:40.059 -> Water_Temp.:92.07, PWM:33
14:28:40.817 -> Water_Temp.:92.07, PWM:33
14:28:43.162 -> Water_Temp.:92.07, PWM:33
14:28:43.955 -> Water_Temp.:92.07, PWM:33
14:28:44.718 -> Water_Temp.:92.07, PWM:33
14:28:46.275 -> Water_Temp.:92.07, PWM:33
14:28:47.070 -> Water_Temp.:92.07, PWM:33
14:28:49.382 -> Water_Temp.:92.07, PWM:33
14:28:50.177 -> Water_Temp.:92.07, PWM:33
14:28:50.937 -> Water_Temp.:92.07, PWM:33
14:28:51.729 -> Water_Temp.:92.07, PWM:33
14:28:53.283 -> Water_Temp.:92.07, PWM:33
14:28:54.837 -> Water_Temp.:92.07, PWM:33
14:28:55.598 -> Water_Temp.:91.96, PWM:33
14:28:56.389 -> Water_Temp.:91.96, PWM:33
14:28:57.940 -> Water_Temp.:91.96, PWM:33
14:28:59.489 -> Water_Temp.:91.96, PWM:33
14:29:00.285 -> Water_Temp.:91.96, PWM:33
14:29:01.043 -> Water_Temp.:91.96, PWM:33
14:29:01.833 -> Water_Temp.:91.96, PWM:33
14:29:03.392 -> Water_Temp.:91.96, PWM:33
14:29:04.184 -> Water_Temp.:91.96, PWM:33
14:29:04.943 -> Water_Temp.:91.96, PWM:33
14:29:06.498 -> Water_Temp.:91.96, PWM:33
14:29:07.257 -> Water_Temp.:91.96, PWM:33
14:29:08.053 -> Water_Temp.:91.96, PWM:33
14:29:08.812 -> Water_Temp.:91.96, PWM:33
14:29:09.602 -> Water_Temp.:91.96, PWM:33
14:29:10.398 -> Water_Temp.:91.96, PWM:33
14:29:11.156 -> Water_Temp.:91.96, PWM:33
14:29:11.947 -> Water_Temp.:91.96, PWM:33
14:29:12.709 -> Water_Temp.:91.96, PWM:33
14:29:13.501 -> Water_Temp.:91.96, PWM:33
14:29:14.294 -> Water_Temp.:91.96, PWM:33
14:29:15.053 -> Water_Temp.:91.96, PWM:33
14:29:15.848 -> Water_Temp.:91.96, PWM:33
14:29:16.611 -> Water_Temp.:91.96, PWM:33
14:29:17.409 -> Water_Temp.:91.96, PWM:33
14:29:22.002 -> Water_Temp.:92.07, PWM:33
14:29:24.314 -> Water_Temp.:92.07, PWM:33
14:29:25.106 -> Water_Temp.:92.07, PWM:33
14:29:26.624 -> Water_Temp.:92.07, PWM:33
14:29:27.421 -> Water_Temp.:92.07, PWM:33
14:29:28.205 -> Water_Temp.:92.07, PWM:33
14:29:28.998 -> Water_Temp.:92.07, PWM:33
14:29:29.757 -> Water_Temp.:92.07, PWM:33
14:29:30.554 -> Water_Temp.:92.07, PWM:33
14:29:31.313 -> Water_Temp.:92.07, PWM:33
14:29:33.627 -> Water_Temp.:92.07, PWM:33
14:29:34.425 -> Water_Temp.:92.07, PWM:33
14:29:35.185 -> Water_Temp.:92.07, PWM:33
14:29:36.740 -> Water_Temp.:92.07, PWM:33
14:29:37.535 -> Water_Temp.:92.07, PWM:33
14:29:38.295 -> Water_Temp.:92.07, PWM:33
14:29:39.091 -> Water_Temp.:92.19, PWM:33
14:29:39.881 -> Water_Temp.:92.19, PWM:33
14:29:41.401 -> Water_Temp.:92.19, PWM:33
14:29:42.195 -> Water_Temp.:92.19, PWM:33
14:29:42.953 -> Water_Temp.:92.19, PWM:33
14:29:44.509 -> Water_Temp.:92.19, PWM:33
14:29:45.272 -> Water_Temp.:92.19, PWM:33
14:29:46.070 -> Water_Temp.:92.19, PWM:33
14:29:46.829 -> Water_Temp.:92.19, PWM:33
14:29:47.624 -> Water_Temp.:92.19, PWM:33
14:29:48.382 -> Water_Temp.:92.19, PWM:33
14:29:49.176 -> Water_Temp.:92.19, PWM:33
14:29:50.730 -> Water_Temp.:92.19, PWM:33
14:29:51.526 -> Water_Temp.:92.19, PWM:33
14:29:53.836 -> Water_Temp.:92.30, PWM:33
14:29:54.631 -> Water_Temp.:92.30, PWM:33
14:29:56.184 -> Water_Temp.:92.30, PWM:33
14:29:56.973 -> Water_Temp.:92.30, PWM:33
14:29:58.523 -> Water_Temp.:92.41, PWM:33
14:29:59.278 -> Water_Temp.:92.41, PWM:33
14:30:00.827 -> Water_Temp.:92.41, PWM:33
14:30:03.152 -> Water_Temp.:92.30, PWM:33
14:30:05.472 -> Water_Temp.:92.41, PWM:33
14:30:07.786 -> Water_Temp.:92.41, PWM:33
14:30:10.894 -> Water_Temp.:92.41, PWM:33
14:30:11.659 -> Water_Temp.:92.41, PWM:33
14:30:13.210 -> Water_Temp.:92.41, PWM:33
14:30:14.003 -> Water_Temp.:92.41, PWM:33
14:30:14.797 -> Water_Temp.:92.41, PWM:33
14:30:16.352 -> Water_Temp.:92.30, PWM:33
14:30:17.910 -> Water_Temp.:92.41, PWM:33
14:30:19.431 -> Water_Temp.:92.30, PWM:33
14:30:20.983 -> Water_Temp.:92.30, PWM:33
14:30:21.776 -> Water_Temp.:92.30, PWM:33
14:30:22.537 -> Water_Temp.:92.30, PWM:33
14:30:23.331 -> Water_Temp.:92.30, PWM:33
14:30:24.882 -> Water_Temp.:92.30, PWM:33
14:30:26.433 -> Water_Temp.:92.41, PWM:33
14:30:27.192 -> Water_Temp.:92.30, PWM:33
14:30:27.982 -> Water_Temp.:92.30, PWM:33
14:30:28.778 -> Water_Temp.:92.30, PWM:33
14:30:29.537 -> Water_Temp.:92.41, PWM:33
14:30:30.332 -> Water_Temp.:92.41, PWM:33
14:30:31.122 -> Water_Temp.:92.41, PWM:33
14:30:31.881 -> Water_Temp.:92.41, PWM:33
14:30:33.438 -> Water_Temp.:92.41, PWM:33
14:30:34.236 -> Water_Temp.:92.41, PWM:33
14:30:34.992 -> Water_Temp.:92.41, PWM:33
14:30:36.548 -> Water_Temp.:92.41, PWM:33
14:30:38.859 -> Water_Temp.:92.41, PWM:33
14:30:41.923 -> Water_Temp.:92.52, PWM:33
14:30:42.720 -> Water_Temp.:92.52, PWM:33
14:30:43.514 -> Water_Temp.:92.52, PWM:32
14:30:44.272 -> Water_Temp.:92.52, PWM:33
14:30:45.826 -> Water_Temp.:92.41, PWM:33
14:30:47.380 -> Water_Temp.:92.41, PWM:33
14:30:48.176 -> Water_Temp.:92.41, PWM:33
14:30:48.970 -> Water_Temp.:92.52, PWM:33
14:30:49.727 -> Water_Temp.:92.52, PWM:33
14:30:50.523 -> Water_Temp.:92.52, PWM:33
14:30:51.284 -> Water_Temp.:92.52, PWM:33
14:30:52.073 -> Water_Temp.:92.52, PWM:33
14:30:52.869 -> Water_Temp.:92.52, PWM:33
14:30:53.630 -> Water_Temp.:92.52, PWM:33
14:30:54.428 -> Water_Temp.:92.52, PWM:33
14:30:55.190 -> Water_Temp.:92.52, PWM:33
14:30:55.985 -> Water_Temp.:92.52, PWM:33
14:30:57.533 -> Water_Temp.:92.52, PWM:33
14:30:58.332 -> Water_Temp.:92.52, PWM:33
14:30:59.847 -> Water_Temp.:92.52, PWM:33
14:31:00.640 -> Water_Temp.:92.52, PWM:33
14:31:02.186 -> Water_Temp.:92.52, PWM:33
14:31:02.946 -> Water_Temp.:92.52, PWM:33
14:31:05.254 -> Water_Temp.:92.52, PWM:33
14:31:06.807 -> Water_Temp.:92.41, PWM:33
14:31:07.567 -> Water_Temp.:92.41, PWM:33
14:31:08.365 -> Water_Temp.:92.41, PWM:33
14:31:09.126 -> Water_Temp.:92.41, PWM:33
14:31:09.919 -> Water_Temp.:92.41, PWM:33
14:31:10.681 -> Water_Temp.:92.41, PWM:33
14:31:11.473 -> Water_Temp.:92.41, PWM:33
14:31:13.028 -> Water_Temp.:92.41, PWM:33
14:31:14.589 -> Water_Temp.:92.41, PWM:33
14:31:16.904 -> Water_Temp.:94.10, PWM:33
14:31:17.698 -> Water_Temp.:96.24, PWM:33
14:31:20.769 -> Water_Temp.:92.97, PWM:33
14:31:21.532 -> Water_Temp.:92.86, PWM:33
14:31:22.325 -> Water_Temp.:92.86, PWM:33
14:31:23.089 -> Water_Temp.:92.75, PWM:33
14:31:24.643 -> Water_Temp.:92.64, PWM:33
14:31:25.402 -> Water_Temp.:92.64, PWM:33
14:31:26.956 -> Water_Temp.:92.64, PWM:33
14:31:27.712 -> Water_Temp.:92.64, PWM:33
14:31:29.263 -> Water_Temp.:92.64, PWM:33
14:31:30.059 -> Water_Temp.:92.52, PWM:33
14:31:30.817 -> Water_Temp.:92.64, PWM:33
14:31:31.608 -> Water_Temp.:92.64, PWM:33
14:31:32.401 -> Water_Temp.:92.52, PWM:33
14:31:33.163 -> Water_Temp.:92.52, PWM:33
14:31:33.958 -> Water_Temp.:92.52, PWM:33
14:31:34.721 -> Water_Temp.:92.52, PWM:33
14:31:36.276 -> Water_Temp.:92.52, PWM:33
14:31:37.038 -> Water_Temp.:92.52, PWM:33
14:31:37.832 -> Water_Temp.:92.52, PWM:33
14:31:40.879 -> Water_Temp.:92.52, PWM:33
14:31:41.673 -> Water_Temp.:92.52, PWM:33
14:31:42.433 -> Water_Temp.:92.52, PWM:33
14:31:44.783 -> Water_Temp.:92.64, PWM:33
14:31:45.576 -> Water_Temp.:92.64, PWM:33
14:31:46.333 -> Water_Temp.:92.64, PWM:33
14:31:47.126 -> Water_Temp.:92.64, PWM:33
14:31:47.921 -> Water_Temp.:92.64, PWM:33
14:31:50.235 -> Water_Temp.:92.64, PWM:33
14:31:50.995 -> Water_Temp.:92.64, PWM:33
14:31:52.550 -> Water_Temp.:92.64, PWM:33
14:31:53.339 -> Water_Temp.:92.64, PWM:33
14:31:54.134 -> Water_Temp.:92.64, PWM:33
14:31:54.894 -> Water_Temp.:92.64, PWM:33
14:31:56.452 -> Water_Temp.:92.64, PWM:33
14:31:57.240 -> Water_Temp.:92.64, PWM:33
14:31:58.035 -> Water_Temp.:92.64, PWM:33
14:32:00.350 -> Water_Temp.:92.64, PWM:33
14:32:01.110 -> Water_Temp.:92.64, PWM:33
14:32:01.904 -> Water_Temp.:92.64, PWM:33
14:32:02.663 -> Water_Temp.:92.64, PWM:33
14:32:05.769 -> Water_Temp.:92.75, PWM:33
14:32:08.076 -> Water_Temp.:92.75, PWM:33
14:32:08.834 -> Water_Temp.:92.75, PWM:33
14:32:09.628 -> Water_Temp.:92.75, PWM:33
14:32:11.178 -> Water_Temp.:92.75, PWM:33
14:32:11.975 -> Water_Temp.:92.75, PWM:33
14:32:13.495 -> Water_Temp.:92.75, PWM:33
14:32:15.835 -> Water_Temp.:94.44, PWM:33
14:32:16.594 -> Water_Temp.:96.46, PWM:33
14:32:17.389 -> Water_Temp.:94.10, PWM:33
14:32:18.146 -> Water_Temp.:93.65, PWM:33
14:32:18.941 -> Water_Temp.:93.43, PWM:33
14:32:19.737 -> Water_Temp.:93.31, PWM:33
14:32:20.497 -> Water_Temp.:93.20, PWM:33
14:32:22.055 -> Water_Temp.:93.09, PWM:33
14:32:22.848 -> Water_Temp.:93.09, PWM:33
14:32:23.610 -> Water_Temp.:93.09, PWM:33
14:32:26.745 -> Water_Temp.:92.97, PWM:33
14:32:28.295 -> Water_Temp.:92.97, PWM:33
14:32:29.057 -> Water_Temp.:92.86, PWM:33
14:32:29.855 -> Water_Temp.:92.97, PWM:33
14:32:31.408 -> Water_Temp.:92.97, PWM:33
14:32:32.166 -> Water_Temp.:92.97, PWM:33
14:32:33.716 -> Water_Temp.:92.97, PWM:33
14:32:34.476 -> Water_Temp.:92.97, PWM:33
14:32:36.067 -> Water_Temp.:92.97, PWM:33
14:32:36.826 -> Water_Temp.:92.97, PWM:33
14:32:37.619 -> Water_Temp.:92.97, PWM:33
14:32:39.169 -> Water_Temp.:92.97, PWM:33
14:32:40.691 -> Water_Temp.:92.97, PWM:33
14:32:41.485 -> Water_Temp.:92.97, PWM:33
14:32:43.038 -> Water_Temp.:92.97, PWM:33
14:32:43.834 -> Water_Temp.:92.97, PWM:33
14:32:46.150 -> Water_Temp.:92.97, PWM:33
14:32:46.909 -> Water_Temp.:92.97, PWM:33
14:32:47.704 -> Water_Temp.:92.97, PWM:33
14:32:49.257 -> Water_Temp.:92.86, PWM:33
14:32:51.572 -> Water_Temp.:92.86, PWM:33
14:32:53.089 -> Water_Temp.:92.86, PWM:33
14:32:53.883 -> Water_Temp.:92.86, PWM:33
14:32:54.642 -> Water_Temp.:92.86, PWM:33
14:32:55.430 -> Water_Temp.:92.86, PWM:33
14:32:56.222 -> Water_Temp.:92.86, PWM:33
14:32:57.777 -> Water_Temp.:92.86, PWM:33
14:32:59.291 -> Water_Temp.:92.86, PWM:33
14:33:00.087 -> Water_Temp.:92.86, PWM:33
14:33:00.881 -> Water_Temp.:92.86, PWM:33
14:33:01.635 -> Water_Temp.:94.55, PWM:33
14:33:03.185 -> Water_Temp.:94.21, PWM:33
14:33:03.947 -> Water_Temp.:93.76, PWM:33
14:33:05.533 -> Water_Temp.:93.43, PWM:33
14:33:06.293 -> Water_Temp.:93.31, PWM:33
14:33:07.086 -> Water_Temp.:93.31, PWM:33
14:33:10.155 -> Water_Temp.:93.09, PWM:33
14:33:10.947 -> Water_Temp.:93.09, PWM:33
14:33:12.497 -> Water_Temp.:93.09, PWM:33
14:33:13.286 -> Water_Temp.:92.97, PWM:33
14:33:14.041 -> Water_Temp.:92.97, PWM:33
14:33:16.352 -> Water_Temp.:92.97, PWM:33
14:33:17.145 -> Water_Temp.:92.97, PWM:33
14:33:17.937 -> Water_Temp.:92.97, PWM:33
14:33:19.486 -> Water_Temp.:92.97, PWM:33
14:33:20.246 -> Water_Temp.:92.97, PWM:33
14:33:21.042 -> Water_Temp.:92.97, PWM:33
14:33:21.836 -> Water_Temp.:92.97, PWM:33
14:33:23.386 -> Water_Temp.:92.97, PWM:33
14:33:24.146 -> Water_Temp.:92.97, PWM:33
14:33:26.485 -> Water_Temp.:92.97, PWM:33
14:33:27.276 -> Water_Temp.:92.97, PWM:33
14:33:28.034 -> Water_Temp.:92.97, PWM:33
14:33:28.828 -> Water_Temp.:92.97, PWM:33
14:33:29.622 -> Water_Temp.:92.97, PWM:33
14:33:30.383 -> Water_Temp.:92.97, PWM:33
14:33:31.175 -> Water_Temp.:92.97, PWM:33
14:33:31.936 -> Water_Temp.:92.97, PWM:33
14:33:32.731 -> Water_Temp.:92.97, PWM:33
14:33:35.047 -> Water_Temp.:92.97, PWM:33
14:33:35.837 -> Water_Temp.:92.97, PWM:33
14:33:37.391 -> Water_Temp.:92.97, PWM:33
14:33:38.188 -> Water_Temp.:92.97, PWM:33
14:33:39.709 -> Water_Temp.:92.97, PWM:33
14:33:40.503 -> Water_Temp.:92.97, PWM:33
14:33:41.290 -> Water_Temp.:92.97, PWM:33
14:33:44.361 -> Water_Temp.:92.97, PWM:33
14:33:45.950 -> Water_Temp.:92.97, PWM:33
14:33:48.267 -> Water_Temp.:92.97, PWM:33
14:33:49.820 -> Water_Temp.:92.97, PWM:33
14:33:50.609 -> Water_Temp.:92.97, PWM:33
14:33:51.367 -> Water_Temp.:92.97, PWM:33
14:33:52.918 -> Water_Temp.:92.97, PWM:33
14:33:53.713 -> Water_Temp.:92.97, PWM:33
14:33:55.235 -> Water_Temp.:93.09, PWM:33
14:33:56.027 -> Water_Temp.:93.09, PWM:33
14:33:58.335 -> Water_Temp.:93.09, PWM:33
14:34:02.153 -> Water_Temp.:93.09, PWM:33
14:34:04.494 -> Water_Temp.:93.09, PWM:33
14:34:06.041 -> Water_Temp.:93.20, PWM:33
14:34:07.599 -> Water_Temp.:93.20, PWM:33
14:34:08.391 -> Water_Temp.:93.20, PWM:33
14:34:09.147 -> Water_Temp.:93.20, PWM:33
14:34:09.940 -> Water_Temp.:93.20, PWM:33
14:34:10.701 -> Water_Temp.:93.20, PWM:33
14:34:12.258 -> Water_Temp.:93.20, PWM:33
14:34:13.050 -> Water_Temp.:93.20, PWM:33
14:34:16.920 -> Water_Temp.:93.31, PWM:33
14:34:17.710 -> Water_Temp.:93.31, PWM:33
14:34:18.470 -> Water_Temp.:93.31, PWM:33
14:34:19.266 -> Water_Temp.:93.31, PWM:33
14:34:20.790 -> Water_Temp.:93.20, PWM:33
14:34:22.343 -> Water_Temp.:93.20, PWM:33
14:34:23.138 -> Water_Temp.:93.20, PWM:33
14:34:23.898 -> Water_Temp.:93.20, PWM:33
14:34:24.692 -> Water_Temp.:93.20, PWM:33
14:34:25.456 -> Water_Temp.:93.20, PWM:33
14:34:27.776 -> Water_Temp.:93.20, PWM:33
14:34:29.330 -> Water_Temp.:93.20, PWM:33
14:34:30.883 -> Water_Temp.:93.20, PWM:33
14:34:31.643 -> Water_Temp.:93.20, PWM:33
14:34:33.994 -> Water_Temp.:93.20, PWM:33
14:34:35.549 -> Water_Temp.:93.20, PWM:33
14:34:36.311 -> Water_Temp.:93.20, PWM:33
14:34:37.100 -> Water_Temp.:93.20, PWM:33
14:34:37.896 -> Water_Temp.:93.09, PWM:33
14:34:38.656 -> Water_Temp.:93.20, PWM:33
14:34:39.452 -> Water_Temp.:93.20, PWM:33
14:34:40.214 -> Water_Temp.:93.20, PWM:33
14:34:41.007 -> Water_Temp.:93.20, PWM:33
14:34:41.800 -> Water_Temp.:93.20, PWM:33
14:34:42.563 -> Water_Temp.:93.20, PWM:33
14:34:44.114 -> Water_Temp.:93.20, PWM:33
14:34:44.875 -> Water_Temp.:93.20, PWM:33
14:34:47.219 -> Water_Temp.:93.20, PWM:33
14:34:48.776 -> Water_Temp.:93.09, PWM:33
14:34:50.298 -> Water_Temp.:93.20, PWM:33
14:34:51.092 -> Water_Temp.:93.20, PWM:33
14:34:51.852 -> Water_Temp.:93.20, PWM:33
14:34:52.649 -> Water_Temp.:93.20, PWM:33
14:34:54.173 -> Water_Temp.:93.20, PWM:33
14:34:54.966 -> Water_Temp.:93.20, PWM:33
14:34:55.728 -> Water_Temp.:93.20, PWM:33
14:34:58.833 -> Water_Temp.:93.31, PWM:33
14:35:00.354 -> Water_Temp.:93.31, PWM:33
14:35:01.147 -> Water_Temp.:93.31, PWM:33
14:35:01.906 -> Water_Temp.:93.31, PWM:33
14:35:02.703 -> Water_Temp.:93.31, PWM:33
14:35:03.466 -> Water_Temp.:93.31, PWM:33
14:35:04.261 -> Water_Temp.:93.31, PWM:33
14:35:05.784 -> Water_Temp.:93.31, PWM:33
14:35:06.577 -> Water_Temp.:93.31, PWM:33
14:35:07.369 -> Water_Temp.:93.31, PWM:33
14:35:08.133 -> Water_Temp.:93.31, PWM:33
14:35:08.925 -> Water_Temp.:93.31, PWM:33
14:35:09.685 -> Water_Temp.:93.31, PWM:33
14:35:10.480 -> Water_Temp.:93.31, PWM:33
14:35:11.271 -> Water_Temp.:93.31, PWM:33
14:35:13.587 -> Water_Temp.:93.31, PWM:33
14:35:14.346 -> Water_Temp.:93.31, PWM:33
14:35:15.135 -> Water_Temp.:93.43, PWM:33
14:35:15.928 -> Water_Temp.:93.43, PWM:33
14:35:17.485 -> Water_Temp.:93.43, PWM:33
14:35:19.792 -> Water_Temp.:93.43, PWM:33
14:35:22.111 -> Water_Temp.:93.43, PWM:33
14:35:22.869 -> Water_Temp.:93.43, PWM:33
14:35:23.662 -> Water_Temp.:93.43, PWM:33
14:35:24.455 -> Water_Temp.:93.43, PWM:33
14:35:28.326 -> Water_Temp.:93.43, PWM:33
14:35:29.087 -> Water_Temp.:93.43, PWM:33
14:35:31.435 -> Water_Temp.:93.43, PWM:33
14:35:33.710 -> Water_Temp.:93.43, PWM:33
14:35:35.298 -> Water_Temp.:93.43, PWM:33
14:35:36.853 -> Water_Temp.:93.43, PWM:33
14:35:38.410 -> Water_Temp.:93.43, PWM:33
14:35:39.171 -> Water_Temp.:93.43, PWM:33
14:35:39.964 -> Water_Temp.:93.43, PWM:33
14:35:41.521 -> Water_Temp.:93.43, PWM:33
14:35:42.313 -> Water_Temp.:93.43, PWM:33
14:35:43.074 -> Water_Temp.:93.43, PWM:33
14:35:43.870 -> Water_Temp.:93.43, PWM:33
14:35:44.630 -> Water_Temp.:93.43, PWM:33
14:35:46.973 -> Water_Temp.:93.43, PWM:33
14:35:47.731 -> Water_Temp.:93.43, PWM:33
14:35:48.529 -> Water_Temp.:93.43, PWM:33
14:35:49.289 -> Water_Temp.:93.43, PWM:33
14:35:50.084 -> Water_Temp.:93.43, PWM:33
14:35:50.843 -> Water_Temp.:93.43, PWM:33
14:35:53.186 -> Water_Temp.:93.54, PWM:33
14:35:53.982 -> Water_Temp.:93.54, PWM:33
14:35:54.740 -> Water_Temp.:93.54, PWM:33
14:35:55.535 -> Water_Temp.:93.54, PWM:33
14:35:56.329 -> Water_Temp.:93.54, PWM:33
14:35:57.845 -> Water_Temp.:93.54, PWM:33
14:35:59.394 -> Water_Temp.:93.54, PWM:33
14:36:00.190 -> Water_Temp.:93.54, PWM:33
14:36:01.737 -> Water_Temp.:93.54, PWM:33
14:36:04.084 -> Water_Temp.:93.54, PWM:33
14:36:04.846 -> Water_Temp.:93.54, PWM:33
14:36:05.639 -> Water_Temp.:93.54, PWM:33
14:36:06.431 -> Water_Temp.:93.54, PWM:33
14:36:07.189 -> Water_Temp.:93.54, PWM:33
14:36:08.742 -> Water_Temp.:93.54, PWM:33
14:36:09.537 -> Water_Temp.:93.43, PWM:33
14:36:10.303 -> Water_Temp.:93.43, PWM:33
14:36:11.095 -> Water_Temp.:93.43, PWM:33
14:36:11.855 -> Water_Temp.:93.43, PWM:33
14:36:12.651 -> Water_Temp.:93.43, PWM:33
14:36:14.167 -> Water_Temp.:93.43, PWM:33
14:36:14.966 -> Water_Temp.:93.54, PWM:33
14:36:16.520 -> Water_Temp.:93.43, PWM:33
14:36:19.625 -> Water_Temp.:93.43, PWM:33
14:36:22.692 -> Water_Temp.:93.43, PWM:33
14:36:23.485 -> Water_Temp.:93.43, PWM:33
14:36:24.241 -> Water_Temp.:93.43, PWM:33
14:36:25.035 -> Water_Temp.:93.43, PWM:33
14:36:25.827 -> Water_Temp.:93.43, PWM:33
14:36:26.586 -> Water_Temp.:93.43, PWM:33
14:36:27.382 -> Water_Temp.:93.43, PWM:33
14:36:28.171 -> Water_Temp.:93.43, PWM:33
14:36:29.725 -> Water_Temp.:93.43, PWM:33
14:36:30.485 -> Water_Temp.:93.43, PWM:33
14:36:32.036 -> Water_Temp.:93.43, PWM:33
14:36:34.344 -> Water_Temp.:93.43, PWM:33
14:36:35.131 -> Water_Temp.:93.43, PWM:33
14:36:35.891 -> Water_Temp.:93.43, PWM:33
14:36:37.443 -> Water_Temp.:93.43, PWM:33
14:36:38.236 -> Water_Temp.:93.43, PWM:33
14:36:39.789 -> Water_Temp.:93.43, PWM:33
14:36:42.106 -> Water_Temp.:93.43, PWM:33
14:36:42.901 -> Water_Temp.:93.43, PWM:33
14:36:44.452 -> Water_Temp.:93.43, PWM:33
14:36:45.241 -> Water_Temp.:93.43, PWM:33
14:36:46.798 -> Water_Temp.:93.43, PWM:33
14:36:49.108 -> Water_Temp.:93.43, PWM:33
14:36:49.905 -> Water_Temp.:93.43, PWM:33
14:36:50.667 -> Water_Temp.:93.43, PWM:33
14:36:52.212 -> Water_Temp.:93.43, PWM:33
14:36:53.005 -> Water_Temp.:93.54, PWM:33
14:36:53.794 -> Water_Temp.:93.54, PWM:33
14:36:55.347 -> Water_Temp.:93.54, PWM:33
14:36:56.112 -> Water_Temp.:93.54, PWM:33
14:36:56.905 -> Water_Temp.:93.54, PWM:33
14:36:58.457 -> Water_Temp.:93.54, PWM:33
14:37:00.013 -> Water_Temp.:93.54, PWM:33
14:37:03.118 -> Water_Temp.:95.22, PWM:33
14:37:04.671 -> Water_Temp.:94.89, PWM:33
14:37:06.975 -> Water_Temp.:94.21, PWM:33
14:37:07.770 -> Water_Temp.:94.10, PWM:33
14:37:08.564 -> Water_Temp.:93.99, PWM:33
14:37:09.327 -> Water_Temp.:93.99, PWM:33
14:37:10.888 -> Water_Temp.:93.87, PWM:33
14:37:12.441 -> Water_Temp.:93.76, PWM:33
14:37:13.235 -> Water_Temp.:93.76, PWM:33
14:37:14.793 -> Water_Temp.:93.76, PWM:33
14:37:17.111 -> Water_Temp.:93.76, PWM:33
14:37:17.871 -> Water_Temp.:93.65, PWM:33
14:37:18.663 -> Water_Temp.:93.65, PWM:33
14:37:19.459 -> Water_Temp.:93.65, PWM:33
14:37:20.218 -> Water_Temp.:93.65, PWM:33
14:37:22.532 -> Water_Temp.:93.65, PWM:33
14:37:23.326 -> Water_Temp.:93.65, PWM:33
14:37:24.089 -> Water_Temp.:93.65, PWM:33
14:37:24.884 -> Water_Temp.:93.65, PWM:33
14:37:26.436 -> Water_Temp.:93.76, PWM:33
14:37:30.296 -> Water_Temp.:93.76, PWM:33
14:37:32.605 -> Water_Temp.:93.65, PWM:33
14:37:33.362 -> Water_Temp.:93.76, PWM:33
14:37:34.156 -> Water_Temp.:93.76, PWM:33
14:37:35.712 -> Water_Temp.:93.76, PWM:33
14:37:36.466 -> Water_Temp.:93.76, PWM:33
14:37:38.049 -> Water_Temp.:93.76, PWM:33
14:37:40.366 -> Water_Temp.:93.76, PWM:33
14:37:44.250 -> Water_Temp.:93.76, PWM:33
14:37:45.044 -> Water_Temp.:93.65, PWM:33
14:37:45.802 -> Water_Temp.:93.65, PWM:33
14:37:49.640 -> Water_Temp.:93.65, PWM:33
14:37:50.436 -> Water_Temp.:93.76, PWM:33
14:37:51.198 -> Water_Temp.:93.76, PWM:33
14:37:53.505 -> Water_Temp.:93.76, PWM:33
14:37:54.262 -> Water_Temp.:93.76, PWM:33
14:37:55.058 -> Water_Temp.:93.76, PWM:33
14:37:55.851 -> Water_Temp.:93.76, PWM:33
14:37:56.611 -> Water_Temp.:93.76, PWM:33
14:37:57.406 -> Water_Temp.:93.76, PWM:33
14:37:58.170 -> Water_Temp.:93.87, PWM:33
14:37:58.959 -> Water_Temp.:93.87, PWM:33
14:37:59.750 -> Water_Temp.:93.87, PWM:33
14:38:00.512 -> Water_Temp.:93.87, PWM:33
14:38:01.306 -> Water_Temp.:93.87, PWM:33
14:38:02.063 -> Water_Temp.:93.87, PWM:33
14:38:02.861 -> Water_Temp.:93.87, PWM:33
14:38:04.413 -> Water_Temp.:93.87, PWM:33
14:38:05.967 -> Water_Temp.:93.87, PWM:33
14:38:06.727 -> Water_Temp.:93.87, PWM:33
14:38:08.285 -> Water_Temp.:93.87, PWM:33
14:38:09.042 -> Water_Temp.:93.87, PWM:33
14:38:09.832 -> Water_Temp.:93.87, PWM:33
14:38:11.385 -> Water_Temp.:93.87, PWM:33
14:38:12.940 -> Water_Temp.:93.76, PWM:33
14:38:18.317 -> Water_Temp.:93.87, PWM:33
14:38:20.668 -> Water_Temp.:93.76, PWM:33
14:38:21.429 -> Water_Temp.:93.87, PWM:33
14:38:22.222 -> Water_Temp.:93.87, PWM:33
14:38:22.982 -> Water_Temp.:93.87, PWM:33
14:38:23.775 -> Water_Temp.:93.87, PWM:33
14:38:25.299 -> Water_Temp.:93.87, PWM:33
14:38:26.846 -> Water_Temp.:93.87, PWM:33
14:38:27.636 -> Water_Temp.:93.87, PWM:33
14:38:28.431 -> Water_Temp.:93.87, PWM:33
14:38:29.189 -> Water_Temp.:93.87, PWM:33
14:38:29.986 -> Water_Temp.:93.87, PWM:33
14:38:30.747 -> Water_Temp.:93.87, PWM:33
14:38:32.296 -> Water_Temp.:93.87, PWM:33
14:38:33.852 -> Water_Temp.:93.87, PWM:33
14:38:34.649 -> Water_Temp.:93.87, PWM:33
14:38:35.410 -> Water_Temp.:93.87, PWM:33
14:38:36.959 -> Water_Temp.:93.87, PWM:33
14:38:37.754 -> Water_Temp.:93.87, PWM:33
14:38:38.512 -> Water_Temp.:93.99, PWM:33
14:38:39.304 -> Water_Temp.:93.87, PWM:33
14:38:40.097 -> Water_Temp.:93.99, PWM:33
14:38:41.647 -> Water_Temp.:93.99, PWM:33
14:38:42.407 -> Water_Temp.:93.87, PWM:33
14:38:44.753 -> Water_Temp.:93.87, PWM:33
14:38:46.307 -> Water_Temp.:93.87, PWM:33
14:38:47.068 -> Water_Temp.:93.99, PWM:33
14:38:47.864 -> Water_Temp.:93.87, PWM:33
14:38:48.625 -> Water_Temp.:93.87, PWM:33
14:38:50.180 -> Water_Temp.:93.87, PWM:33
14:38:50.974 -> Water_Temp.:93.87, PWM:33
14:38:52.531 -> Water_Temp.:93.99, PWM:33
14:38:53.322 -> Water_Temp.:93.99, PWM:33
14:38:54.840 -> Water_Temp.:93.99, PWM:33
14:38:55.638 -> Water_Temp.:93.99, PWM:33
14:38:56.400 -> Water_Temp.:93.99, PWM:33
14:38:57.196 -> Water_Temp.:93.99, PWM:33
14:38:57.954 -> Water_Temp.:93.99, PWM:33
14:38:58.747 -> Water_Temp.:93.99, PWM:33
14:38:59.540 -> Water_Temp.:93.99, PWM:33
14:39:02.606 -> Water_Temp.:93.99, PWM:33
14:39:03.363 -> Water_Temp.:93.99, PWM:33
14:39:04.913 -> Water_Temp.:93.99, PWM:33
14:39:06.466 -> Water_Temp.:93.99, PWM:33
14:39:07.226 -> Water_Temp.:93.99, PWM:33
14:39:08.014 -> Water_Temp.:93.99, PWM:33
14:39:08.806 -> Water_Temp.:93.99, PWM:33
14:39:10.325 -> Water_Temp.:93.99, PWM:33
14:39:11.906 -> Water_Temp.:93.99, PWM:33
14:39:12.666 -> Water_Temp.:93.99, PWM:33
14:39:15.013 -> Water_Temp.:93.99, PWM:33
14:39:15.808 -> Water_Temp.:93.99, PWM:33
14:39:17.328 -> Water_Temp.:93.99, PWM:33
14:39:18.089 -> Water_Temp.:93.99, PWM:33
14:39:18.881 -> Water_Temp.:93.99, PWM:33
14:39:19.673 -> Water_Temp.:93.99, PWM:33
14:39:20.436 -> Water_Temp.:93.99, PWM:33
14:39:21.991 -> Water_Temp.:93.99, PWM:33
14:39:22.751 -> Water_Temp.:93.99, PWM:33
14:39:24.338 -> Water_Temp.:93.99, PWM:33
14:39:26.650 -> Water_Temp.:93.99, PWM:33
14:39:28.926 -> Water_Temp.:93.99, PWM:33
14:39:29.722 -> Water_Temp.:93.99, PWM:33
14:39:31.270 -> Water_Temp.:93.99, PWM:33
14:39:32.026 -> Water_Temp.:93.99, PWM:33
14:39:32.820 -> Water_Temp.:93.99, PWM:33
14:39:33.583 -> Water_Temp.:93.99, PWM:33
14:39:34.374 -> Water_Temp.:93.99, PWM:33
14:39:35.136 -> Water_Temp.:93.99, PWM:33
14:39:35.931 -> Water_Temp.:93.99, PWM:33
14:39:37.482 -> Water_Temp.:94.10, PWM:33
14:39:38.276 -> Water_Temp.:93.99, PWM:33
14:39:39.036 -> Water_Temp.:94.10, PWM:33
14:39:39.828 -> Water_Temp.:94.10, PWM:33
14:39:40.626 -> Water_Temp.:94.10, PWM:33
14:39:41.382 -> Water_Temp.:94.10, PWM:33
14:39:42.171 -> Water_Temp.:94.10, PWM:33
14:39:44.488 -> Water_Temp.:94.10, PWM:33
14:39:45.247 -> Water_Temp.:94.10, PWM:33
14:39:46.804 -> Water_Temp.:94.10, PWM:33
14:39:47.595 -> Water_Temp.:94.10, PWM:33
14:39:48.389 -> Water_Temp.:94.10, PWM:33
14:39:49.944 -> Water_Temp.:94.10, PWM:33
14:39:50.703 -> Water_Temp.:94.10, PWM:33
14:39:51.497 -> Water_Temp.:94.10, PWM:33
14:39:53.048 -> Water_Temp.:94.21, PWM:33
14:39:54.601 -> Water_Temp.:94.21, PWM:33
14:39:55.393 -> Water_Temp.:94.21, PWM:33
14:39:56.154 -> Water_Temp.:94.10, PWM:33
14:39:56.949 -> Water_Temp.:94.21, PWM:33
14:39:58.509 -> Water_Temp.:94.21, PWM:33
14:39:59.266 -> Water_Temp.:94.21, PWM:33
14:40:00.057 -> Water_Temp.:94.21, PWM:33
14:40:00.853 -> Water_Temp.:94.21, PWM:33
14:40:01.613 -> Water_Temp.:94.21, PWM:33
14:40:03.196 -> Water_Temp.:94.21, PWM:33
14:40:05.505 -> Water_Temp.:94.10, PWM:33
14:40:06.296 -> Water_Temp.:94.10, PWM:33
14:40:08.644 -> Water_Temp.:94.10, PWM:33
14:40:10.203 -> Water_Temp.:94.21, PWM:33
14:40:10.963 -> Water_Temp.:94.10, PWM:33
14:40:11.758 -> Water_Temp.:94.10, PWM:33
14:40:12.519 -> Water_Temp.:94.21, PWM:33
14:40:13.315 -> Water_Temp.:94.21, PWM:33
14:40:14.077 -> Water_Temp.:94.10, PWM:33
14:40:14.872 -> Water_Temp.:94.10, PWM:33
14:40:15.665 -> Water_Temp.:94.10, PWM:33
14:40:16.425 -> Water_Temp.:94.10, PWM:33
14:40:17.221 -> Water_Temp.:94.10, PWM:33
14:40:19.504 -> Water_Temp.:94.10, PWM:33
14:40:21.060 -> Water_Temp.:94.10, PWM:33
14:40:21.821 -> Water_Temp.:94.21, PWM:33
14:40:22.614 -> Water_Temp.:94.21, PWM:33
14:40:23.404 -> Water_Temp.:94.21, PWM:33
14:40:25.722 -> Water_Temp.:94.21, PWM:33
14:40:26.513 -> Water_Temp.:94.21, PWM:33
14:40:27.274 -> Water_Temp.:94.21, PWM:33
14:40:30.384 -> Water_Temp.:94.21, PWM:33
14:40:31.142 -> Water_Temp.:94.21, PWM:33
14:40:31.939 -> Water_Temp.:94.32, PWM:33
14:40:34.247 -> Water_Temp.:94.32, PWM:33
14:40:35.044 -> Water_Temp.:94.21, PWM:33
14:40:35.808 -> Water_Temp.:94.32, PWM:33
14:40:36.602 -> Water_Temp.:95.90, PWM:33
14:40:37.356 -> Water_Temp.:98.04, PWM:33
14:40:38.149 -> Water_Temp.:95.56, PWM:33
14:40:38.946 -> Water_Temp.:95.22, PWM:33
14:40:40.471 -> Water_Temp.:94.89, PWM:33
14:40:41.267 -> Water_Temp.:94.77, PWM:33
14:40:42.060 -> Water_Temp.:94.66, PWM:33
14:40:42.819 -> Water_Temp.:94.66, PWM:33
14:40:43.615 -> Water_Temp.:94.66, PWM:33
14:40:45.161 -> Water_Temp.:94.55, PWM:33
14:40:47.506 -> Water_Temp.:94.55, PWM:33
14:40:48.260 -> Water_Temp.:94.44, PWM:33
14:40:52.156 -> Water_Temp.:94.44, PWM:33
14:40:52.920 -> Water_Temp.:94.32, PWM:33
14:40:53.714 -> Water_Temp.:94.32, PWM:33
14:40:56.025 -> Water_Temp.:94.32, PWM:33
14:40:57.576 -> Water_Temp.:94.32, PWM:33
14:40:59.125 -> Water_Temp.:94.32, PWM:33
14:40:59.915 -> Water_Temp.:94.32, PWM:33
14:41:02.217 -> Water_Temp.:94.32, PWM:33
14:41:03.009 -> Water_Temp.:94.32, PWM:33
14:41:04.558 -> Water_Temp.:94.32, PWM:33
14:41:06.082 -> Water_Temp.:94.32, PWM:33
14:41:07.634 -> Water_Temp.:94.32, PWM:33
14:41:08.427 -> Water_Temp.:94.32, PWM:33
14:41:09.181 -> Water_Temp.:94.32, PWM:33
14:41:09.974 -> Water_Temp.:94.32, PWM:33
14:41:10.770 -> Water_Temp.:94.32, PWM:33
14:41:12.285 -> Water_Temp.:94.32, PWM:33
14:41:13.080 -> Water_Temp.:94.32, PWM:33
14:41:14.630 -> Water_Temp.:94.32, PWM:33
14:41:16.159 -> Water_Temp.:94.44, PWM:33
14:41:17.708 -> Water_Temp.:94.44, PWM:33
14:41:18.504 -> Water_Temp.:94.44, PWM:33
14:41:19.289 -> Water_Temp.:94.44, PWM:33
14:41:20.051 -> Water_Temp.:94.44, PWM:33
14:41:21.608 -> Water_Temp.:94.44, PWM:33
14:41:22.362 -> Water_Temp.:94.44, PWM:33
14:41:23.158 -> Water_Temp.:94.44, PWM:33
14:41:23.920 -> Water_Temp.:94.44, PWM:33
14:41:26.229 -> Water_Temp.:94.44, PWM:33
14:41:27.020 -> Water_Temp.:94.44, PWM:33
14:41:27.810 -> Water_Temp.:94.44, PWM:33
14:41:28.572 -> Water_Temp.:94.44, PWM:33
14:41:29.365 -> Water_Temp.:94.44, PWM:33
14:41:30.129 -> Water_Temp.:94.44, PWM:33
14:41:30.922 -> Water_Temp.:94.44, PWM:33
14:41:32.477 -> Water_Temp.:94.44, PWM:33
14:41:33.274 -> Water_Temp.:94.44, PWM:33
14:41:34.036 -> Water_Temp.:94.44, PWM:33
14:41:34.798 -> Water_Temp.:94.44, PWM:33
14:41:36.349 -> Water_Temp.:94.44, PWM:33
14:41:38.700 -> Water_Temp.:94.44, PWM:33
14:41:39.455 -> Water_Temp.:94.44, PWM:33
14:41:40.253 -> Water_Temp.:94.44, PWM:33
14:41:41.012 -> Water_Temp.:94.44, PWM:33
14:41:41.805 -> Water_Temp.:94.44, PWM:33
14:41:43.329 -> Water_Temp.:94.44, PWM:33
14:41:44.125 -> Water_Temp.:94.44, PWM:33
14:41:45.646 -> Water_Temp.:94.44, PWM:33
14:41:46.440 -> Water_Temp.:94.44, PWM:33
14:41:48.754 -> Water_Temp.:94.44, PWM:33
14:41:49.548 -> Water_Temp.:94.44, PWM:33
14:41:52.593 -> Water_Temp.:94.44, PWM:33
14:41:54.909 -> Water_Temp.:94.44, PWM:33
14:41:56.460 -> Water_Temp.:94.44, PWM:33
14:41:59.532 -> Water_Temp.:94.55, PWM:33
14:42:00.327 -> Water_Temp.:94.55, PWM:33
14:42:01.082 -> Water_Temp.:94.55, PWM:33
14:42:01.871 -> Water_Temp.:94.55, PWM:33
14:42:04.151 -> Water_Temp.:94.55, PWM:33
14:42:04.940 -> Water_Temp.:94.55, PWM:33
14:42:06.495 -> Water_Temp.:94.55, PWM:33
14:42:07.253 -> Water_Temp.:94.55, PWM:33
14:42:08.805 -> Water_Temp.:94.55, PWM:33
14:42:09.595 -> Water_Temp.:94.55, PWM:33
14:42:11.902 -> Water_Temp.:94.55, PWM:33
14:42:12.695 -> Water_Temp.:94.55, PWM:33
14:42:13.449 -> Water_Temp.:94.55, PWM:33
14:42:14.244 -> Water_Temp.:94.55, PWM:33
14:42:15.038 -> Water_Temp.:94.55, PWM:33
14:42:15.798 -> Water_Temp.:94.55, PWM:33
14:42:16.586 -> Water_Temp.:94.55, PWM:33
14:42:18.134 -> Water_Temp.:94.55, PWM:33
14:42:18.925 -> Water_Temp.:94.55, PWM:33
14:42:19.682 -> Water_Temp.:94.55, PWM:33
14:42:20.476 -> Water_Temp.:94.55, PWM:33
14:42:21.992 -> Water_Temp.:94.55, PWM:33
14:42:23.569 -> Water_Temp.:94.55, PWM:33
14:42:24.328 -> Water_Temp.:94.55, PWM:33
14:42:25.125 -> Water_Temp.:94.55, PWM:33
14:42:27.408 -> Water_Temp.:94.55, PWM:33
14:42:28.200 -> Water_Temp.:94.55, PWM:33
14:42:29.755 -> Water_Temp.:94.55, PWM:33
14:42:30.551 -> Water_Temp.:94.55, PWM:33
14:42:33.622 -> Water_Temp.:94.55, PWM:33
14:42:34.415 -> Water_Temp.:94.55, PWM:33
14:42:35.173 -> Water_Temp.:94.55, PWM:33
14:42:36.757 -> Water_Temp.:94.55, PWM:33
14:42:38.310 -> Water_Temp.:94.55, PWM:32
14:42:39.069 -> Water_Temp.:94.55, PWM:33
14:42:39.866 -> Water_Temp.:94.55, PWM:33
14:42:41.417 -> Water_Temp.:94.55, PWM:33
14:42:47.588 -> Water_Temp.:94.55, PWM:33
14:42:48.343 -> Water_Temp.:94.55, PWM:33
14:42:49.138 -> Water_Temp.:94.55, PWM:33
14:42:49.899 -> Water_Temp.:94.55, PWM:33
14:42:51.459 -> Water_Temp.:94.55, PWM:33
14:42:52.250 -> Water_Temp.:94.55, PWM:33
14:42:54.569 -> Water_Temp.:94.55, PWM:33
14:42:55.331 -> Water_Temp.:94.55, PWM:33
14:42:56.128 -> Water_Temp.:94.55, PWM:33
14:42:56.919 -> Water_Temp.:94.55, PWM:33
14:42:57.678 -> Water_Temp.:94.55, PWM:33
14:42:58.471 -> Water_Temp.:94.55, PWM:33
14:43:00.023 -> Water_Temp.:94.55, PWM:33
14:43:01.575 -> Water_Temp.:94.44, PWM:33
14:43:02.366 -> Water_Temp.:94.44, PWM:33
14:43:04.677 -> Water_Temp.:94.44, PWM:33
14:43:05.472 -> Water_Temp.:94.44, PWM:33
14:43:06.233 -> Water_Temp.:94.44, PWM:33
14:43:07.024 -> Water_Temp.:94.44, PWM:33
14:43:07.821 -> Water_Temp.:94.44, PWM:33
14:43:08.583 -> Water_Temp.:94.44, PWM:33
14:43:09.378 -> Water_Temp.:94.44, PWM:33
14:43:10.134 -> Water_Temp.:94.44, PWM:33
14:43:10.923 -> Water_Temp.:94.44, PWM:33
14:43:11.714 -> Water_Temp.:94.44, PWM:33
14:43:12.475 -> Water_Temp.:94.44, PWM:33
14:43:14.024 -> Water_Temp.:94.44, PWM:33
14:43:14.814 -> Water_Temp.:94.44, PWM:33
14:43:15.608 -> Water_Temp.:94.44, PWM:33
14:43:16.367 -> Water_Temp.:94.44, PWM:33
14:43:17.164 -> Water_Temp.:94.44, PWM:33
14:43:17.926 -> Water_Temp.:94.44, PWM:33
14:43:18.718 -> Water_Temp.:94.44, PWM:33
14:43:20.236 -> Water_Temp.:94.44, PWM:33
14:43:21.023 -> Water_Temp.:94.44, PWM:33
14:43:21.820 -> Water_Temp.:94.44, PWM:33
14:43:22.580 -> Water_Temp.:94.32, PWM:33
14:43:23.376 -> Water_Temp.:94.32, PWM:33
14:43:24.142 -> Water_Temp.:94.32, PWM:33
14:43:24.936 -> Water_Temp.:94.32, PWM:33
14:43:27.254 -> Water_Temp.:94.32, PWM:33
14:43:28.011 -> Water_Temp.:94.32, PWM:33
14:43:30.329 -> Water_Temp.:94.32, PWM:33
14:43:31.887 -> Water_Temp.:94.32, PWM:33
14:43:33.442 -> Water_Temp.:94.32, PWM:33
14:43:34.235 -> Water_Temp.:94.32, PWM:33
14:43:35.021 -> Water_Temp.:94.32, PWM:33
14:43:35.782 -> Water_Temp.:94.32, PWM:33
14:43:36.574 -> Water_Temp.:94.32, PWM:33
14:43:37.334 -> Water_Temp.:94.32, PWM:33
14:43:38.130 -> Water_Temp.:94.32, PWM:33
14:43:38.891 -> Water_Temp.:94.32, PWM:33
14:43:40.442 -> Water_Temp.:94.32, PWM:33
14:43:41.233 -> Water_Temp.:94.32, PWM:33
14:43:41.991 -> Water_Temp.:94.32, PWM:33
14:43:42.785 -> Water_Temp.:96.01, PWM:33
14:43:44.309 -> Water_Temp.:95.68, PWM:33
14:43:46.661 -> Water_Temp.:94.89, PWM:33
14:43:47.452 -> Water_Temp.:94.77, PWM:33
14:43:48.208 -> Water_Temp.:94.77, PWM:33
14:43:49.002 -> Water_Temp.:94.77, PWM:33
14:43:49.761 -> Water_Temp.:94.66, PWM:33
14:43:51.353 -> Water_Temp.:94.66, PWM:33
14:43:52.115 -> Water_Temp.:94.55, PWM:33
14:43:52.908 -> Water_Temp.:94.55, PWM:33
14:43:53.670 -> Water_Temp.:94.55, PWM:33
14:43:54.465 -> Water_Temp.:94.55, PWM:33
14:43:55.986 -> Water_Temp.:94.55, PWM:33
14:43:56.783 -> Water_Temp.:94.55, PWM:33
14:43:58.341 -> Water_Temp.:94.55, PWM:33
14:43:59.892 -> Water_Temp.:94.55, PWM:33
14:44:01.443 -> Water_Temp.:94.55, PWM:33
14:44:02.200 -> Water_Temp.:94.55, PWM:33
14:44:03.749 -> Water_Temp.:94.55, PWM:33
14:44:04.508 -> Water_Temp.:94.55, PWM:33
14:44:05.300 -> Water_Temp.:94.55, PWM:33
14:44:06.097 -> Water_Temp.:94.55, PWM:33
14:44:06.859 -> Water_Temp.:94.55, PWM:33
14:44:07.652 -> Water_Temp.:94.55, PWM:33
14:44:08.412 -> Water_Temp.:94.55, PWM:33
14:44:09.964 -> Water_Temp.:94.55, PWM:33
14:44:10.755 -> Water_Temp.:94.55, PWM:33
14:44:11.514 -> Water_Temp.:94.55, PWM:33
14:44:13.833 -> Water_Temp.:94.44, PWM:33
14:44:14.594 -> Water_Temp.:94.44, PWM:33
14:44:16.150 -> Water_Temp.:94.44, PWM:33
14:44:17.733 -> Water_Temp.:94.55, PWM:33
14:44:18.494 -> Water_Temp.:94.44, PWM:33
14:44:19.289 -> Water_Temp.:94.55, PWM:33
14:44:20.046 -> Water_Temp.:94.55, PWM:33
14:44:20.842 -> Water_Temp.:94.55, PWM:33
14:44:23.153 -> Water_Temp.:94.55, PWM:33
14:44:23.948 -> Water_Temp.:94.55, PWM:33
14:44:24.707 -> Water_Temp.:94.55, PWM:33
14:44:25.503 -> Water_Temp.:94.55, PWM:33
14:44:26.262 -> Water_Temp.:94.66, PWM:33
14:44:27.057 -> Water_Temp.:94.66, PWM:33
14:44:28.613 -> Water_Temp.:94.66, PWM:33
14:44:29.374 -> Water_Temp.:94.66, PWM:33
14:44:30.169 -> Water_Temp.:94.66, PWM:33
14:44:31.716 -> Water_Temp.:94.66, PWM:33
14:44:32.505 -> Water_Temp.:94.66, PWM:33
14:44:34.055 -> Water_Temp.:94.77, PWM:33
14:44:34.816 -> Water_Temp.:94.77, PWM:33
14:44:36.365 -> Water_Temp.:94.77, PWM:33
14:44:37.159 -> Water_Temp.:94.77, PWM:33
14:44:37.922 -> Water_Temp.:94.77, PWM:33
14:44:38.719 -> Water_Temp.:94.77, PWM:33
14:44:40.998 -> Water_Temp.:94.77, PWM:33
14:44:43.344 -> Water_Temp.:94.77, PWM:33
14:44:44.105 -> Water_Temp.:94.77, PWM:33
14:44:44.900 -> Water_Temp.:94.77, PWM:33
14:44:45.662 -> Water_Temp.:94.77, PWM:33
14:44:47.252 -> Water_Temp.:94.77, PWM:33
14:44:48.773 -> Water_Temp.:94.77, PWM:33
14:44:49.568 -> Water_Temp.:94.77, PWM:33
14:44:50.329 -> Water_Temp.:94.77, PWM:33
14:44:51.884 -> Water_Temp.:94.66, PWM:33
14:44:52.678 -> Water_Temp.:94.77, PWM:33
14:44:53.437 -> Water_Temp.:94.66, PWM:33
14:44:55.751 -> Water_Temp.:94.66, PWM:33
14:44:56.545 -> Water_Temp.:94.66, PWM:33
14:44:57.339 -> Water_Temp.:94.77, PWM:33
14:44:58.101 -> Water_Temp.:94.77, PWM:33
14:44:59.654 -> Water_Temp.:94.77, PWM:33
14:45:00.447 -> Water_Temp.:94.66, PWM:33
14:45:01.210 -> Water_Temp.:94.77, PWM:33
14:45:01.997 -> Water_Temp.:94.77, PWM:33
14:45:02.754 -> Water_Temp.:94.77, PWM:33
14:45:05.100 -> Water_Temp.:94.77, PWM:33
14:45:06.658 -> Water_Temp.:94.66, PWM:33
14:45:08.974 -> Water_Temp.:94.66, PWM:33
14:45:09.741 -> Water_Temp.:94.66, PWM:33
14:45:11.273 -> Water_Temp.:94.66, PWM:33
14:45:12.827 -> Water_Temp.:94.66, PWM:33
14:45:13.616 -> Water_Temp.:94.66, PWM:33
14:45:15.891 -> Water_Temp.:94.66, PWM:33
14:45:16.685 -> Water_Temp.:94.66, PWM:33
14:45:18.239 -> Water_Temp.:94.77, PWM:33
14:45:19.032 -> Water_Temp.:94.66, PWM:33
14:45:19.792 -> Water_Temp.:94.66, PWM:33
14:45:21.350 -> Water_Temp.:94.77, PWM:33
14:45:22.140 -> Water_Temp.:94.77, PWM:33
14:45:22.897 -> Water_Temp.:94.77, PWM:33
14:45:23.691 -> Water_Temp.:94.77, PWM:33
14:45:25.245 -> Water_Temp.:94.77, PWM:33
14:45:29.072 -> Water_Temp.:94.77, PWM:33
14:45:29.865 -> Water_Temp.:94.77, PWM:33
14:45:30.629 -> Water_Temp.:94.77, PWM:33
14:45:31.423 -> Water_Temp.:94.77, PWM:33
14:45:34.558 -> Water_Temp.:94.77, PWM:33
14:45:36.079 -> Water_Temp.:94.77, PWM:33
14:45:36.837 -> Water_Temp.:94.77, PWM:33
14:45:37.630 -> Water_Temp.:94.77, PWM:33
14:45:39.180 -> Water_Temp.:94.77, PWM:33
14:45:42.251 -> Water_Temp.:94.77, PWM:33
14:45:43.804 -> Water_Temp.:94.77, PWM:33
14:45:44.601 -> Water_Temp.:94.66, PWM:33
14:45:45.359 -> Water_Temp.:94.77, PWM:33
14:45:46.157 -> Water_Temp.:94.66, PWM:33
14:45:46.917 -> Water_Temp.:94.66, PWM:33
14:45:47.709 -> Water_Temp.:94.66, PWM:33
14:45:48.502 -> Water_Temp.:94.77, PWM:33
14:45:49.258 -> Water_Temp.:94.77, PWM:33
14:45:51.575 -> Water_Temp.:94.77, PWM:33
14:45:52.371 -> Water_Temp.:94.66, PWM:33
14:45:53.133 -> Water_Temp.:94.66, PWM:33
14:45:56.241 -> Water_Temp.:98.49, PWM:33
14:45:57.793 -> Water_Temp.:95.56, PWM:33
14:45:59.346 -> Water_Temp.:95.34, PWM:33
14:46:00.108 -> Water_Temp.:95.11, PWM:33
14:46:00.902 -> Water_Temp.:95.11, PWM:33
14:46:01.697 -> Water_Temp.:95.00, PWM:33
14:46:02.457 -> Water_Temp.:95.00, PWM:33
14:46:04.006 -> Water_Temp.:94.89, PWM:33
14:46:04.767 -> Water_Temp.:94.89, PWM:33
14:46:05.562 -> Water_Temp.:94.89, PWM:33
14:46:06.355 -> Water_Temp.:94.89, PWM:33
14:46:07.873 -> Water_Temp.:94.77, PWM:33
14:46:08.634 -> Water_Temp.:94.77, PWM:33
14:46:10.955 -> Water_Temp.:94.77, PWM:33
14:46:11.713 -> Water_Temp.:94.77, PWM:33
14:46:12.505 -> Water_Temp.:94.77, PWM:33
14:46:13.299 -> Water_Temp.:94.77, PWM:33
14:46:14.855 -> Water_Temp.:94.77, PWM:33
14:46:15.615 -> Water_Temp.:94.77, PWM:33
14:46:16.406 -> Water_Temp.:94.77, PWM:33
14:46:17.161 -> Water_Temp.:94.77, PWM:33
14:46:17.952 -> Water_Temp.:94.77, PWM:33
14:46:18.746 -> Water_Temp.:94.77, PWM:33
14:46:19.510 -> Water_Temp.:94.77, PWM:33
14:46:21.068 -> Water_Temp.:94.77, PWM:33
14:46:22.617 -> Water_Temp.:94.77, PWM:33
14:46:23.378 -> Water_Temp.:94.77, PWM:33
14:46:24.170 -> Water_Temp.:94.66, PWM:33
14:46:24.963 -> Water_Temp.:94.77, PWM:33
14:46:25.720 -> Water_Temp.:94.66, PWM:33
14:46:26.513 -> Water_Temp.:94.66, PWM:33
14:46:30.386 -> Water_Temp.:94.66, PWM:33
14:46:31.942 -> Water_Temp.:94.66, PWM:33
14:46:33.493 -> Water_Temp.:94.66, PWM:33
14:46:34.256 -> Water_Temp.:94.77, PWM:33
14:46:35.817 -> Water_Temp.:94.77, PWM:33
14:46:38.127 -> Water_Temp.:94.66, PWM:33
14:46:38.882 -> Water_Temp.:94.66, PWM:33
14:46:39.680 -> Water_Temp.:94.66, PWM:33
14:46:40.440 -> Water_Temp.:94.66, PWM:33
14:46:41.233 -> Water_Temp.:94.66, PWM:33
14:46:42.030 -> Water_Temp.:94.66, PWM:33
14:46:42.790 -> Water_Temp.:94.66, PWM:33
14:46:43.585 -> Water_Temp.:94.66, PWM:33
14:46:44.345 -> Water_Temp.:94.66, PWM:33
14:46:45.141 -> Water_Temp.:94.66, PWM:33
14:46:45.903 -> Water_Temp.:94.66, PWM:33
14:46:47.451 -> Water_Temp.:94.66, PWM:33
14:46:48.247 -> Water_Temp.:94.66, PWM:33
14:46:49.004 -> Water_Temp.:94.66, PWM:33
14:46:49.796 -> Water_Temp.:94.66, PWM:33
14:46:50.557 -> Water_Temp.:94.55, PWM:33
14:46:51.353 -> Water_Temp.:94.55, PWM:33
14:46:52.143 -> Water_Temp.:94.55, PWM:33
14:46:52.901 -> Water_Temp.:94.55, PWM:33
14:46:54.456 -> Water_Temp.:94.55, PWM:33
14:46:55.250 -> Water_Temp.:94.55, PWM:33
14:46:56.048 -> Water_Temp.:94.55, PWM:33
14:46:56.806 -> Water_Temp.:94.55, PWM:33
14:46:57.601 -> Water_Temp.:94.55, PWM:33
14:46:59.123 -> Water_Temp.:94.55, PWM:33
14:46:59.917 -> Water_Temp.:94.55, PWM:33
14:47:00.679 -> Water_Temp.:94.55, PWM:33
14:47:01.474 -> Water_Temp.:94.55, PWM:33
14:47:02.265 -> Water_Temp.:94.55, PWM:33
14:47:03.808 -> Water_Temp.:94.55, PWM:33
14:47:05.364 -> Water_Temp.:94.55, PWM:33
14:47:06.919 -> Water_Temp.:94.55, PWM:33
14:47:07.712 -> Water_Temp.:94.55, PWM:33
14:47:08.473 -> Water_Temp.:94.55, PWM:33
14:47:09.271 -> Water_Temp.:94.55, PWM:33
14:47:10.031 -> Water_Temp.:94.55, PWM:33
14:47:11.588 -> Water_Temp.:94.55, PWM:33
14:47:12.345 -> Water_Temp.:94.55, PWM:33
14:47:13.141 -> Water_Temp.:94.55, PWM:33
14:47:14.692 -> Water_Temp.:94.55, PWM:33
14:47:17.045 -> Water_Temp.:94.55, PWM:33
14:47:17.806 -> Water_Temp.:94.55, PWM:33
14:47:19.363 -> Water_Temp.:94.55, PWM:33
14:47:20.917 -> Water_Temp.:94.55, PWM:33
14:47:21.713 -> Water_Temp.:94.55, PWM:33
14:47:23.272 -> Water_Temp.:94.55, PWM:33
14:47:24.067 -> Water_Temp.:94.66, PWM:33
14:47:24.821 -> Water_Temp.:94.66, PWM:33
14:47:25.619 -> Water_Temp.:94.66, PWM:33
14:47:27.172 -> Water_Temp.:94.66, PWM:33
14:47:29.474 -> Water_Temp.:94.66, PWM:33
14:47:30.233 -> Water_Temp.:94.66, PWM:33
14:47:31.804 -> Water_Temp.:94.55, PWM:33
14:47:32.588 -> Water_Temp.:94.66, PWM:33
14:47:34.892 -> Water_Temp.:94.66, PWM:33
14:47:35.687 -> Water_Temp.:94.66, PWM:33
14:47:36.479 -> Water_Temp.:94.66, PWM:33
14:47:37.240 -> Water_Temp.:94.66, PWM:33
14:47:38.036 -> Water_Temp.:94.66, PWM:33
14:47:38.796 -> Water_Temp.:94.66, PWM:33
14:47:39.590 -> Water_Temp.:94.66, PWM:33
14:47:41.142 -> Water_Temp.:94.66, PWM:33
14:47:41.903 -> Water_Temp.:94.66, PWM:33
14:47:42.693 -> Water_Temp.:94.66, PWM:33
14:47:43.451 -> Water_Temp.:94.66, PWM:33
14:47:44.247 -> Water_Temp.:94.66, PWM:33
14:47:45.039 -> Water_Temp.:94.66, PWM:33
14:47:45.800 -> Water_Temp.:94.66, PWM:33
14:47:47.353 -> Water_Temp.:94.66, PWM:33
14:47:48.899 -> Water_Temp.:94.66, PWM:33
14:47:50.456 -> Water_Temp.:94.66, PWM:33
14:47:51.248 -> Water_Temp.:94.66, PWM:33
14:47:52.799 -> Water_Temp.:94.66, PWM:33
14:47:53.561 -> Water_Temp.:94.66, PWM:33
14:47:55.114 -> Water_Temp.:94.66, PWM:33
14:47:55.911 -> Water_Temp.:94.66, PWM:33
14:47:56.673 -> Water_Temp.:94.66, PWM:33
14:47:57.463 -> Water_Temp.:94.66, PWM:33
14:47:58.255 -> Water_Temp.:94.66, PWM:33
14:47:59.014 -> Water_Temp.:94.66, PWM:33
14:47:59.799 -> Water_Temp.:94.66, PWM:33
14:48:01.351 -> Water_Temp.:94.66, PWM:33
14:48:02.145 -> Water_Temp.:94.66, PWM:33
14:48:03.663 -> Water_Temp.:94.66, PWM:33
14:48:04.457 -> Water_Temp.:94.66, PWM:33
14:48:05.213 -> Water_Temp.:94.66, PWM:33
14:48:06.008 -> Water_Temp.:94.66, PWM:33
14:48:06.800 -> Water_Temp.:94.66, PWM:33
14:48:08.316 -> Water_Temp.:94.66, PWM:33
14:48:09.109 -> Water_Temp.:94.66, PWM:33
14:48:09.897 -> Water_Temp.:94.66, PWM:33
14:48:10.655 -> Water_Temp.:94.66, PWM:33
14:48:11.449 -> Water_Temp.:94.66, PWM:33
14:48:12.241 -> Water_Temp.:94.66, PWM:33
14:48:13.001 -> Water_Temp.:94.66, PWM:33
14:48:13.792 -> Water_Temp.:94.66, PWM:33
14:48:14.551 -> Water_Temp.:94.66, PWM:33
14:48:17.658 -> Water_Temp.:94.55, PWM:33
14:48:19.174 -> Water_Temp.:94.55, PWM:33
14:48:20.751 -> Water_Temp.:94.55, PWM:33
14:48:22.271 -> Water_Temp.:94.55, PWM:33
14:48:23.820 -> Water_Temp.:94.55, PWM:33
14:48:25.372 -> Water_Temp.:94.55, PWM:33
14:48:26.169 -> Water_Temp.:94.55, PWM:33
14:48:26.926 -> Water_Temp.:94.55, PWM:33
14:48:27.722 -> Water_Temp.:94.55, PWM:33
14:48:28.515 -> Water_Temp.:94.55, PWM:33
14:48:29.274 -> Water_Temp.:94.55, PWM:33
14:48:30.067 -> Water_Temp.:94.55, PWM:33
14:48:32.367 -> Water_Temp.:96.24, PWM:33
14:48:34.709 -> Water_Temp.:95.45, PWM:33
14:48:36.264 -> Water_Temp.:95.11, PWM:33
14:48:37.058 -> Water_Temp.:95.11, PWM:33
14:48:38.607 -> Water_Temp.:94.89, PWM:33
14:48:40.159 -> Water_Temp.:94.89, PWM:33
14:48:40.954 -> Water_Temp.:94.77, PWM:33
14:48:42.472 -> Water_Temp.:94.77, PWM:33
14:48:43.264 -> Water_Temp.:94.77, PWM:33
14:48:44.815 -> Water_Temp.:94.66, PWM:33
14:48:45.578 -> Water_Temp.:94.66, PWM:33
14:48:46.375 -> Water_Temp.:94.66, PWM:33
14:48:47.135 -> Water_Temp.:94.66, PWM:33
14:48:48.685 -> Water_Temp.:94.66, PWM:33
14:48:51.030 -> Water_Temp.:94.66, PWM:33
14:48:53.338 -> Water_Temp.:94.66, PWM:33
14:48:54.132 -> Water_Temp.:94.55, PWM:33
14:48:54.891 -> Water_Temp.:94.55, PWM:33
14:48:55.690 -> Water_Temp.:94.55, PWM:33
14:48:57.211 -> Water_Temp.:94.55, PWM:33
14:48:58.007 -> Water_Temp.:94.55, PWM:33
14:48:58.766 -> Water_Temp.:94.55, PWM:33
14:49:00.316 -> Water_Temp.:94.55, PWM:33
14:49:01.870 -> Water_Temp.:94.55, PWM:33
14:49:02.627 -> Water_Temp.:94.55, PWM:33
14:49:04.975 -> Water_Temp.:94.55, PWM:33
14:49:07.293 -> Water_Temp.:94.55, PWM:33
14:49:08.852 -> Water_Temp.:94.55, PWM:33
14:49:11.934 -> Water_Temp.:94.44, PWM:33
14:49:12.730 -> Water_Temp.:94.44, PWM:33
14:49:13.493 -> Water_Temp.:94.44, PWM:33
14:49:14.284 -> Water_Temp.:94.44, PWM:33
14:49:15.046 -> Water_Temp.:94.44, PWM:35
14:49:16.634 -> Water_Temp.:94.44, PWM:35
14:49:17.388 -> Water_Temp.:94.44, PWM:34
14:49:18.182 -> Water_Temp.:94.44, PWM:35
14:49:18.973 -> Water_Temp.:94.44, PWM:34
14:49:20.494 -> Water_Temp.:94.44, PWM:34
14:49:21.289 -> Water_Temp.:94.44, PWM:34
14:49:22.048 -> Water_Temp.:94.44, PWM:35
14:49:22.837 -> Water_Temp.:94.44, PWM:35
14:49:23.632 -> Water_Temp.:94.44, PWM:35
14:49:25.950 -> Water_Temp.:94.44, PWM:35
14:49:26.711 -> Water_Temp.:94.44, PWM:34
14:49:27.509 -> Water_Temp.:94.44, PWM:35
14:49:28.272 -> Water_Temp.:94.44, PWM:34
14:49:29.068 -> Water_Temp.:94.44, PWM:34
14:49:29.828 -> Water_Temp.:94.44, PWM:34
14:49:30.622 -> Water_Temp.:94.44, PWM:35
14:49:31.383 -> Water_Temp.:94.44, PWM:34
14:49:32.181 -> Water_Temp.:94.44, PWM:35
14:49:33.739 -> Water_Temp.:94.44, PWM:34
14:49:35.293 -> Water_Temp.:94.44, PWM:34
14:49:36.088 -> Water_Temp.:94.44, PWM:34
14:49:39.961 -> Water_Temp.:94.44, PWM:34
14:49:40.718 -> Water_Temp.:94.44, PWM:35
14:49:41.514 -> Water_Temp.:94.44, PWM:35
14:49:42.276 -> Water_Temp.:94.44, PWM:34
14:49:43.831 -> Water_Temp.:94.44, PWM:34
14:49:44.627 -> Water_Temp.:94.44, PWM:34
14:49:45.391 -> Water_Temp.:94.44, PWM:34
14:49:46.186 -> Water_Temp.:94.44, PWM:35
14:49:46.979 -> Water_Temp.:94.44, PWM:34
14:49:47.743 -> Water_Temp.:94.44, PWM:34
14:49:50.814 -> Water_Temp.:94.44, PWM:34
14:49:52.365 -> Water_Temp.:94.44, PWM:35
14:49:53.160 -> Water_Temp.:94.44, PWM:34
14:49:53.921 -> Water_Temp.:94.44, PWM:34
14:49:54.714 -> Water_Temp.:94.44, PWM:34
14:49:57.028 -> Water_Temp.:94.44, PWM:34
14:49:59.364 -> Water_Temp.:94.44, PWM:35
14:50:00.910 -> Water_Temp.:94.44, PWM:34
14:50:01.707 -> Water_Temp.:94.44, PWM:34
14:50:02.464 -> Water_Temp.:94.44, PWM:35
14:50:03.255 -> Water_Temp.:94.44, PWM:35
14:50:04.804 -> Water_Temp.:94.44, PWM:34
14:50:07.883 -> Water_Temp.:94.44, PWM:34
14:50:08.681 -> Water_Temp.:94.44, PWM:35
14:50:10.201 -> Water_Temp.:94.44, PWM:34
14:50:10.994 -> Water_Temp.:94.44, PWM:35
14:50:11.752 -> Water_Temp.:94.44, PWM:35
14:50:12.545 -> Water_Temp.:94.44, PWM:34
14:50:13.341 -> Water_Temp.:94.44, PWM:34
14:50:14.103 -> Water_Temp.:94.44, PWM:35
14:50:14.898 -> Water_Temp.:94.44, PWM:34
14:50:15.658 -> Water_Temp.:94.44, PWM:34
14:50:17.215 -> Water_Temp.:94.55, PWM:34
14:50:18.767 -> Water_Temp.:94.55, PWM:34
14:50:19.529 -> Water_Temp.:94.55, PWM:35
14:50:21.082 -> Water_Temp.:94.44, PWM:35
14:50:21.875 -> Water_Temp.:94.44, PWM:34
14:50:22.672 -> Water_Temp.:94.44, PWM:35
14:50:24.196 -> Water_Temp.:94.44, PWM:34
14:50:24.957 -> Water_Temp.:94.44, PWM:34
14:50:25.751 -> Water_Temp.:94.44, PWM:34
14:50:27.277 -> Water_Temp.:94.44, PWM:34
14:50:28.832 -> Water_Temp.:94.44, PWM:34
14:50:29.625 -> Water_Temp.:94.44, PWM:34
14:50:31.183 -> Water_Temp.:94.44, PWM:35
14:50:31.943 -> Water_Temp.:94.44, PWM:34
14:50:32.739 -> Water_Temp.:94.44, PWM:35
14:50:34.292 -> Water_Temp.:94.32, PWM:35
14:50:35.086 -> Water_Temp.:94.32, PWM:34
14:50:35.845 -> Water_Temp.:94.44, PWM:34
14:50:37.397 -> Water_Temp.:94.44, PWM:34
14:50:38.947 -> Water_Temp.:94.44, PWM:35
14:50:39.706 -> Water_Temp.:94.44, PWM:34
14:50:40.501 -> Water_Temp.:94.44, PWM:34
14:50:41.261 -> Water_Temp.:94.44, PWM:34
14:50:42.812 -> Water_Temp.:94.44, PWM:35
14:50:43.607 -> Water_Temp.:94.44, PWM:34
14:50:44.367 -> Water_Temp.:94.44, PWM:34
14:50:46.716 -> Water_Temp.:94.44, PWM:34
14:50:47.513 -> Water_Temp.:94.44, PWM:34
14:50:48.275 -> Water_Temp.:94.44, PWM:35
14:50:49.826 -> Water_Temp.:94.44, PWM:35
14:50:51.386 -> Water_Temp.:94.44, PWM:34
14:50:52.180 -> Water_Temp.:94.44, PWM:35
14:50:52.939 -> Water_Temp.:94.44, PWM:35
14:50:53.731 -> Water_Temp.:94.44, PWM:35
14:50:54.518 -> Water_Temp.:94.44, PWM:35
14:50:55.277 -> Water_Temp.:94.44, PWM:34
14:50:56.830 -> Water_Temp.:94.44, PWM:35
14:50:57.626 -> Water_Temp.:94.44, PWM:35
14:50:59.182 -> Water_Temp.:94.44, PWM:35
14:50:59.976 -> Water_Temp.:94.44, PWM:35
14:51:00.739 -> Water_Temp.:94.44, PWM:34
14:51:01.531 -> Water_Temp.:94.44, PWM:35
14:51:02.285 -> Water_Temp.:94.44, PWM:34
14:51:03.865 -> Water_Temp.:94.55, PWM:35
14:51:05.420 -> Water_Temp.:94.55, PWM:35
14:51:06.180 -> Water_Temp.:94.55, PWM:35
14:51:06.974 -> Water_Temp.:94.55, PWM:34
14:51:07.736 -> Water_Temp.:94.55, PWM:34
14:51:08.525 -> Water_Temp.:94.55, PWM:34
14:51:10.079 -> Water_Temp.:94.55, PWM:35
14:51:10.871 -> Water_Temp.:94.55, PWM:34
14:51:11.632 -> Water_Temp.:94.55, PWM:34
14:51:12.426 -> Water_Temp.:94.44, PWM:34
14:51:13.983 -> Water_Temp.:94.44, PWM:35
14:51:14.738 -> Water_Temp.:94.44, PWM:34
14:51:15.531 -> Water_Temp.:94.44, PWM:34
14:51:16.291 -> Water_Temp.:96.12, PWM:35
14:51:17.083 -> Water_Temp.:98.26, PWM:35
14:51:18.630 -> Water_Temp.:95.34, PWM:34
14:51:19.392 -> Water_Temp.:95.11, PWM:35
14:51:20.944 -> Water_Temp.:95.00, PWM:34
14:51:23.287 -> Water_Temp.:94.77, PWM:35
14:51:24.841 -> Water_Temp.:94.66, PWM:35
14:51:25.607 -> Water_Temp.:94.66, PWM:34
14:51:27.161 -> Water_Temp.:94.55, PWM:35
14:51:27.953 -> Water_Temp.:94.55, PWM:35
14:51:28.711 -> Water_Temp.:94.55, PWM:35
14:51:29.509 -> Water_Temp.:94.55, PWM:34
14:51:31.057 -> Water_Temp.:94.55, PWM:35
14:51:31.853 -> Water_Temp.:94.55, PWM:34
14:51:32.615 -> Water_Temp.:94.55, PWM:34
14:51:33.411 -> Water_Temp.:94.55, PWM:35
14:51:34.174 -> Water_Temp.:94.55, PWM:34
14:51:35.726 -> Water_Temp.:94.44, PWM:35
14:51:36.520 -> Water_Temp.:94.44, PWM:34
14:51:37.279 -> Water_Temp.:94.44, PWM:34
14:51:38.832 -> Water_Temp.:94.44, PWM:34
14:51:40.385 -> Water_Temp.:94.44, PWM:34
14:51:41.176 -> Water_Temp.:94.44, PWM:34
14:51:41.933 -> Water_Temp.:94.44, PWM:35
14:51:42.724 -> Water_Temp.:94.44, PWM:35
14:51:45.063 -> Water_Temp.:94.55, PWM:35
14:51:46.585 -> Water_Temp.:94.55, PWM:34
14:51:47.381 -> Water_Temp.:94.55, PWM:35
14:51:48.929 -> Water_Temp.:94.55, PWM:35
14:51:49.724 -> Water_Temp.:94.55, PWM:35
14:51:50.487 -> Water_Temp.:94.55, PWM:35
14:51:52.043 -> Water_Temp.:94.55, PWM:34
14:51:53.558 -> Water_Temp.:94.55, PWM:34
14:51:57.392 -> Water_Temp.:94.55, PWM:34
14:52:00.509 -> Water_Temp.:94.55, PWM:34
14:52:01.269 -> Water_Temp.:94.55, PWM:34
14:52:02.065 -> Water_Temp.:94.55, PWM:35
14:52:02.823 -> Water_Temp.:94.55, PWM:34
14:52:03.619 -> Water_Temp.:94.55, PWM:34
14:52:04.413 -> Water_Temp.:94.55, PWM:35
14:52:05.175 -> Water_Temp.:94.55, PWM:34
14:52:05.970 -> Water_Temp.:94.55, PWM:34
14:52:07.486 -> Water_Temp.:94.66, PWM:34
14:52:08.276 -> Water_Temp.:94.55, PWM:34
14:52:09.072 -> Water_Temp.:94.66, PWM:34
14:52:09.834 -> Water_Temp.:94.55, PWM:37
14:52:11.383 -> Water_Temp.:94.66, PWM:38
14:52:12.178 -> Water_Temp.:94.66, PWM:38
14:52:14.452 -> Water_Temp.:94.66, PWM:38
14:52:15.247 -> Water_Temp.:94.66, PWM:38
14:52:16.041 -> Water_Temp.:94.66, PWM:38
14:52:17.598 -> Water_Temp.:94.66, PWM:38
14:52:18.359 -> Water_Temp.:94.66, PWM:38
14:52:19.916 -> Water_Temp.:94.66, PWM:38
14:52:20.711 -> Water_Temp.:94.66, PWM:38
14:52:21.471 -> Water_Temp.:94.66, PWM:38
14:52:22.268 -> Water_Temp.:94.66, PWM:38
14:52:23.059 -> Water_Temp.:94.66, PWM:38
14:52:23.818 -> Water_Temp.:94.66, PWM:38
14:52:24.615 -> Water_Temp.:94.66, PWM:38
14:52:25.374 -> Water_Temp.:94.66, PWM:38
14:52:26.171 -> Water_Temp.:94.66, PWM:38
14:52:26.935 -> Water_Temp.:94.66, PWM:38
14:52:27.731 -> Water_Temp.:94.66, PWM:38
14:52:28.491 -> Water_Temp.:94.66, PWM:38
14:52:30.807 -> Water_Temp.:94.66, PWM:38
14:52:33.121 -> Water_Temp.:94.66, PWM:38
14:52:33.879 -> Water_Temp.:94.66, PWM:38
14:52:34.671 -> Water_Temp.:94.66, PWM:38
14:52:35.465 -> Water_Temp.:94.66, PWM:38
14:52:36.988 -> Water_Temp.:94.66, PWM:38
14:52:37.780 -> Water_Temp.:94.66, PWM:38
14:52:39.293 -> Water_Temp.:94.66, PWM:38
14:52:40.090 -> Water_Temp.:94.66, PWM:38
14:52:40.878 -> Water_Temp.:94.66, PWM:38
14:52:41.640 -> Water_Temp.:94.66, PWM:38
14:52:43.193 -> Water_Temp.:94.66, PWM:38
14:52:43.955 -> Water_Temp.:94.66, PWM:38
14:52:44.748 -> Water_Temp.:94.66, PWM:38
14:52:45.508 -> Water_Temp.:94.66, PWM:38
14:52:47.100 -> Water_Temp.:94.66, PWM:38
14:52:47.861 -> Water_Temp.:94.55, PWM:38
14:52:48.653 -> Water_Temp.:94.66, PWM:38
14:52:49.411 -> Water_Temp.:94.66, PWM:38
14:52:50.965 -> Water_Temp.:94.66, PWM:38
14:52:52.520 -> Water_Temp.:94.55, PWM:38
14:52:53.314 -> Water_Temp.:94.66, PWM:38
14:52:54.872 -> Water_Temp.:94.66, PWM:38
14:52:55.636 -> Water_Temp.:94.66, PWM:38
14:52:56.434 -> Water_Temp.:94.66, PWM:38
14:52:57.192 -> Water_Temp.:94.66, PWM:38
14:52:57.988 -> Water_Temp.:94.66, PWM:38
14:52:58.749 -> Water_Temp.:94.66, PWM:38
14:53:00.306 -> Water_Temp.:94.66, PWM:38
14:53:01.102 -> Water_Temp.:94.55, PWM:38
14:53:01.861 -> Water_Temp.:94.55, PWM:38
14:53:02.656 -> Water_Temp.:94.55, PWM:38
14:53:03.416 -> Water_Temp.:94.55, PWM:38
14:53:04.974 -> Water_Temp.:94.55, PWM:38
14:53:05.735 -> Water_Temp.:94.55, PWM:38
14:53:06.530 -> Water_Temp.:94.55, PWM:38
14:53:07.296 -> Water_Temp.:94.55, PWM:38
14:53:08.090 -> Water_Temp.:94.55, PWM:38
14:53:08.885 -> Water_Temp.:94.55, PWM:38
14:53:09.646 -> Water_Temp.:94.55, PWM:38
14:53:10.409 -> Water_Temp.:94.55, PWM:38
14:53:12.758 -> Water_Temp.:94.55, PWM:38
14:53:13.553 -> Water_Temp.:94.55, PWM:38
14:53:15.870 -> Water_Temp.:94.55, PWM:38
14:53:16.630 -> Water_Temp.:94.66, PWM:38
14:53:17.426 -> Water_Temp.:94.66, PWM:38
14:53:18.187 -> Water_Temp.:94.66, PWM:38
14:53:18.979 -> Water_Temp.:94.66, PWM:38
14:53:19.738 -> Water_Temp.:94.66, PWM:38
14:53:21.329 -> Water_Temp.:94.66, PWM:38
14:53:23.637 -> Water_Temp.:94.66, PWM:38
14:53:24.431 -> Water_Temp.:94.66, PWM:38
14:53:25.222 -> Water_Temp.:94.66, PWM:38
14:53:25.982 -> Water_Temp.:94.77, PWM:38
14:53:26.777 -> Water_Temp.:94.77, PWM:38
14:53:27.537 -> Water_Temp.:94.77, PWM:38
14:53:28.335 -> Water_Temp.:94.77, PWM:38
14:53:29.100 -> Water_Temp.:94.77, PWM:38
14:53:29.895 -> Water_Temp.:94.77, PWM:38
14:53:30.688 -> Water_Temp.:94.77, PWM:38
14:53:31.451 -> Water_Temp.:94.77, PWM:38
14:53:33.766 -> Water_Temp.:94.77, PWM:38
14:53:34.524 -> Water_Temp.:94.77, PWM:38
14:53:35.323 -> Water_Temp.:94.77, PWM:38
14:53:37.607 -> Water_Temp.:94.77, PWM:38
14:53:38.402 -> Water_Temp.:94.89, PWM:38
14:53:39.161 -> Water_Temp.:94.77, PWM:38
14:53:41.476 -> Water_Temp.:94.89, PWM:38
14:53:42.267 -> Water_Temp.:94.89, PWM:38
14:53:43.063 -> Water_Temp.:94.89, PWM:38
14:53:43.822 -> Water_Temp.:94.89, PWM:38
14:53:45.376 -> Water_Temp.:94.89, PWM:38
14:53:46.172 -> Water_Temp.:94.89, PWM:38
14:53:46.937 -> Water_Temp.:94.89, PWM:38
14:53:47.729 -> Water_Temp.:94.89, PWM:38
14:53:48.486 -> Water_Temp.:94.89, PWM:38
14:53:49.272 -> Water_Temp.:94.89, PWM:38
14:53:50.060 -> Water_Temp.:94.89, PWM:38
14:53:51.616 -> Water_Temp.:94.89, PWM:38
14:53:55.441 -> Water_Temp.:94.89, PWM:38
14:53:56.234 -> Water_Temp.:94.89, PWM:38
14:53:57.023 -> Water_Temp.:94.89, PWM:38
14:53:58.569 -> Water_Temp.:94.89, PWM:38
14:53:59.366 -> Water_Temp.:94.89, PWM:38
14:54:00.126 -> Water_Temp.:94.89, PWM:38
14:54:00.918 -> Water_Temp.:94.89, PWM:38
14:54:01.677 -> Water_Temp.:94.89, PWM:38
14:54:05.572 -> Water_Temp.:94.89, PWM:38
14:54:07.887 -> Water_Temp.:94.89, PWM:38
14:54:08.677 -> Water_Temp.:94.89, PWM:38
14:54:10.226 -> Water_Temp.:94.89, PWM:38
14:54:11.782 -> Water_Temp.:94.89, PWM:38
14:54:13.340 -> Water_Temp.:94.89, PWM:38
14:54:14.130 -> Water_Temp.:94.89, PWM:38
14:54:14.923 -> Water_Temp.:94.89, PWM:38
14:54:16.438 -> Water_Temp.:94.89, PWM:38
14:54:17.995 -> Water_Temp.:94.89, PWM:38
14:54:19.547 -> Water_Temp.:94.89, PWM:38
14:54:20.307 -> Water_Temp.:94.89, PWM:38
14:54:21.098 -> Water_Temp.:94.89, PWM:38
14:54:21.894 -> Water_Temp.:94.89, PWM:38
14:54:22.658 -> Water_Temp.:94.89, PWM:38
14:54:25.012 -> Water_Temp.:94.89, PWM:38
14:54:25.772 -> Water_Temp.:94.89, PWM:38
14:54:26.563 -> Water_Temp.:94.89, PWM:38
14:54:27.324 -> Water_Temp.:94.89, PWM:38
14:54:28.122 -> Water_Temp.:94.89, PWM:38
14:54:28.913 -> Water_Temp.:94.89, PWM:38
14:54:29.676 -> Water_Temp.:94.89, PWM:38
14:54:30.468 -> Water_Temp.:94.89, PWM:38
14:54:31.227 -> Water_Temp.:94.89, PWM:38
14:54:33.539 -> Water_Temp.:94.89, PWM:38
14:54:35.092 -> Water_Temp.:94.89, PWM:38
14:54:37.408 -> Water_Temp.:94.89, PWM:38
14:54:38.962 -> Water_Temp.:94.89, PWM:38
14:54:39.752 -> Water_Temp.:94.89, PWM:38
14:54:40.509 -> Water_Temp.:94.89, PWM:38
14:54:42.091 -> Water_Temp.:95.00, PWM:38
14:54:42.847 -> Water_Temp.:95.00, PWM:38
14:54:44.405 -> Water_Temp.:95.00, PWM:38
14:54:45.167 -> Water_Temp.:95.00, PWM:38
14:54:45.965 -> Water_Temp.:95.00, PWM:38
14:54:47.516 -> Water_Temp.:95.00, PWM:38
14:54:48.312 -> Water_Temp.:95.00, PWM:38
14:54:50.620 -> Water_Temp.:95.00, PWM:38
14:54:52.173 -> Water_Temp.:95.00, PWM:38
14:54:52.965 -> Water_Temp.:95.00, PWM:38
14:54:53.723 -> Water_Temp.:95.11, PWM:38
14:54:55.276 -> Water_Temp.:95.00, PWM:38
14:54:57.589 -> Water_Temp.:95.11, PWM:38
14:54:58.381 -> Water_Temp.:95.11, PWM:38
14:54:59.173 -> Water_Temp.:95.11, PWM:38
14:55:00.729 -> Water_Temp.:95.11, PWM:38
14:55:03.009 -> Water_Temp.:95.11, PWM:38
14:55:03.767 -> Water_Temp.:95.11, PWM:32
14:55:04.564 -> Water_Temp.:95.11, PWM:32
14:55:05.324 -> Water_Temp.:95.00, PWM:32
14:55:06.913 -> Water_Temp.:95.11, PWM:32
14:55:07.673 -> Water_Temp.:95.11, PWM:32
14:55:08.466 -> Water_Temp.:95.11, PWM:32
14:55:10.022 -> Water_Temp.:95.00, PWM:32
14:55:10.784 -> Water_Temp.:95.11, PWM:32
14:55:11.576 -> Water_Temp.:95.00, PWM:33
14:55:13.123 -> Water_Temp.:95.11, PWM:32
14:55:15.436 -> Water_Temp.:95.00, PWM:33
14:55:16.990 -> Water_Temp.:95.00, PWM:33
14:55:17.782 -> Water_Temp.:95.00, PWM:32
14:55:18.539 -> Water_Temp.:95.00, PWM:32
14:55:19.336 -> Water_Temp.:95.00, PWM:32
14:55:20.889 -> Water_Temp.:95.00, PWM:32
14:55:21.680 -> Water_Temp.:95.00, PWM:32
14:55:22.469 -> Water_Temp.:95.00, PWM:33
14:55:23.229 -> Water_Temp.:95.11, PWM:33
14:55:24.792 -> Water_Temp.:95.11, PWM:32
14:55:25.589 -> Water_Temp.:95.11, PWM:32
14:55:26.345 -> Water_Temp.:95.11, PWM:32
14:55:27.142 -> Water_Temp.:95.11, PWM:32
14:55:29.452 -> Water_Temp.:95.22, PWM:33
14:55:30.246 -> Water_Temp.:95.22, PWM:33
14:55:31.041 -> Water_Temp.:95.22, PWM:32
14:55:31.799 -> Water_Temp.:95.22, PWM:32
14:55:32.595 -> Water_Temp.:95.22, PWM:33
14:55:34.145 -> Water_Temp.:95.22, PWM:33
14:55:34.905 -> Water_Temp.:95.22, PWM:33
14:55:35.700 -> Water_Temp.:95.22, PWM:32
14:55:37.217 -> Water_Temp.:95.22, PWM:32
14:55:38.772 -> Water_Temp.:95.22, PWM:32
14:55:39.564 -> Water_Temp.:95.22, PWM:32
14:55:40.325 -> Water_Temp.:95.22, PWM:32
14:55:41.118 -> Water_Temp.:95.22, PWM:32
14:55:41.878 -> Water_Temp.:95.22, PWM:32
14:55:43.428 -> Water_Temp.:95.22, PWM:32
14:55:44.224 -> Water_Temp.:95.22, PWM:32
14:55:44.986 -> Water_Temp.:95.34, PWM:32
14:55:45.779 -> Water_Temp.:95.22, PWM:32
14:55:46.539 -> Water_Temp.:95.34, PWM:32
14:55:47.330 -> Water_Temp.:95.22, PWM:33
14:55:48.124 -> Water_Temp.:95.22, PWM:33
14:55:49.643 -> Water_Temp.:95.22, PWM:33
14:55:51.201 -> Water_Temp.:95.22, PWM:33
14:55:51.997 -> Water_Temp.:95.22, PWM:32
14:55:52.787 -> Water_Temp.:95.22, PWM:33
14:55:53.544 -> Water_Temp.:95.22, PWM:33
14:55:54.339 -> Water_Temp.:95.22, PWM:32
14:55:57.414 -> Water_Temp.:95.22, PWM:33
14:55:59.721 -> Water_Temp.:95.22, PWM:33
14:56:00.512 -> Water_Temp.:95.22, PWM:33
14:56:02.028 -> Water_Temp.:95.22, PWM:33
14:56:03.613 -> Water_Temp.:95.22, PWM:33
14:56:05.168 -> Water_Temp.:95.22, PWM:32
14:56:05.930 -> Water_Temp.:95.22, PWM:33
14:56:06.722 -> Water_Temp.:95.22, PWM:32
14:56:07.512 -> Water_Temp.:95.22, PWM:33
14:56:08.268 -> Water_Temp.:95.22, PWM:32
14:56:09.057 -> Water_Temp.:95.22, PWM:32
14:56:09.851 -> Water_Temp.:95.22, PWM:32
14:56:11.372 -> Water_Temp.:95.22, PWM:33
14:56:12.165 -> Water_Temp.:95.22, PWM:32
14:56:12.920 -> Water_Temp.:95.11, PWM:32
14:56:13.705 -> Water_Temp.:95.11, PWM:33
14:56:14.501 -> Water_Temp.:95.11, PWM:32
14:56:16.057 -> Water_Temp.:95.11, PWM:33
14:56:17.608 -> Water_Temp.:95.11, PWM:33
14:56:19.955 -> Water_Temp.:95.11, PWM:33
14:56:20.717 -> Water_Temp.:95.11, PWM:32
14:56:23.031 -> Water_Temp.:95.22, PWM:32
14:56:23.822 -> Water_Temp.:95.22, PWM:33
14:56:25.338 -> Water_Temp.:95.22, PWM:32
14:56:26.133 -> Water_Temp.:95.22, PWM:32
14:56:27.653 -> Water_Temp.:95.22, PWM:33
14:56:28.446 -> Water_Temp.:95.22, PWM:32
14:56:29.209 -> Water_Temp.:95.22, PWM:32
14:56:30.006 -> Water_Temp.:96.91, PWM:33
14:56:31.525 -> Water_Temp.:96.57, PWM:33
14:56:32.320 -> Water_Temp.:96.12, PWM:33
14:56:33.875 -> Water_Temp.:95.79, PWM:32
14:56:34.631 -> Water_Temp.:95.68, PWM:33
14:56:35.427 -> Water_Temp.:95.68, PWM:33
14:56:36.976 -> Water_Temp.:95.56, PWM:33
14:56:38.531 -> Water_Temp.:95.45, PWM:33
14:56:40.088 -> Water_Temp.:95.34, PWM:32
14:56:40.847 -> Water_Temp.:95.34, PWM:32
14:56:41.641 -> Water_Temp.:95.34, PWM:33
14:56:42.435 -> Water_Temp.:95.34, PWM:32
14:56:43.197 -> Water_Temp.:95.34, PWM:32
14:56:43.991 -> Water_Temp.:95.34, PWM:32
14:56:44.750 -> Water_Temp.:95.34, PWM:32
14:56:47.891 -> Water_Temp.:95.22, PWM:33
14:56:48.653 -> Water_Temp.:95.22, PWM:32
14:56:49.445 -> Water_Temp.:95.22, PWM:32
14:56:51.002 -> Water_Temp.:95.34, PWM:33
14:56:51.761 -> Water_Temp.:95.22, PWM:33
14:56:53.347 -> Water_Temp.:95.34, PWM:32
14:56:54.103 -> Water_Temp.:95.22, PWM:32
14:56:54.897 -> Water_Temp.:95.34, PWM:32
14:56:56.421 -> Water_Temp.:95.34, PWM:33
14:56:57.217 -> Water_Temp.:95.22, PWM:33
14:56:57.979 -> Water_Temp.:95.22, PWM:32
14:56:59.531 -> Water_Temp.:95.34, PWM:32
14:57:00.328 -> Water_Temp.:95.22, PWM:32
14:57:01.120 -> Water_Temp.:95.22, PWM:32
14:57:02.676 -> Water_Temp.:95.22, PWM:32
14:57:04.197 -> Water_Temp.:95.22, PWM:33
14:57:04.991 -> Water_Temp.:95.34, PWM:32
14:57:05.784 -> Water_Temp.:95.34, PWM:33
14:57:06.543 -> Water_Temp.:95.34, PWM:32
14:57:07.340 -> Water_Temp.:95.34, PWM:33
14:57:08.128 -> Water_Temp.:95.22, PWM:33
14:57:08.887 -> Water_Temp.:95.22, PWM:33
14:57:10.445 -> Water_Temp.:95.22, PWM:32
14:57:11.239 -> Water_Temp.:95.22, PWM:32
14:57:12.755 -> Water_Temp.:95.22, PWM:32
14:57:16.624 -> Water_Temp.:95.34, PWM:32
14:57:17.419 -> Water_Temp.:95.34, PWM:33
14:57:18.180 -> Water_Temp.:95.34, PWM:32
14:57:18.975 -> Water_Temp.:95.34, PWM:32
14:57:19.768 -> Water_Temp.:95.34, PWM:33
14:57:21.286 -> Water_Temp.:95.22, PWM:32
14:57:22.083 -> Water_Temp.:95.22, PWM:32
14:57:22.845 -> Water_Temp.:95.22, PWM:33
14:57:25.150 -> Water_Temp.:95.22, PWM:32
14:57:26.709 -> Water_Temp.:95.22, PWM:32
14:57:27.473 -> Water_Temp.:95.22, PWM:32
14:57:29.056 -> Water_Temp.:95.22, PWM:32
14:57:29.815 -> Water_Temp.:95.22, PWM:32
14:57:30.609 -> Water_Temp.:95.22, PWM:33
14:57:32.126 -> Water_Temp.:95.22, PWM:33
14:57:34.468 -> Water_Temp.:95.22, PWM:32
14:57:36.026 -> Water_Temp.:95.22, PWM:32
14:57:36.820 -> Water_Temp.:95.22, PWM:33
14:57:38.375 -> Water_Temp.:95.22, PWM:32
14:57:39.923 -> Water_Temp.:95.22, PWM:32
14:57:40.686 -> Water_Temp.:95.22, PWM:33
14:57:41.484 -> Water_Temp.:95.22, PWM:32
14:57:42.245 -> Water_Temp.:95.22, PWM:33
14:57:44.583 -> Water_Temp.:95.22, PWM:32
14:57:46.930 -> Water_Temp.:95.22, PWM:33
14:57:47.692 -> Water_Temp.:95.22, PWM:33
14:57:49.275 -> Water_Temp.:95.22, PWM:32
14:57:50.831 -> Water_Temp.:95.22, PWM:33
14:57:51.592 -> Water_Temp.:95.22, PWM:32
14:57:53.138 -> Water_Temp.:95.22, PWM:33
14:57:53.933 -> Water_Temp.:95.22, PWM:32
14:57:54.689 -> Water_Temp.:95.22, PWM:32
14:57:57.007 -> Water_Temp.:95.22, PWM:32
14:57:57.798 -> Water_Temp.:95.22, PWM:32
14:57:58.560 -> Water_Temp.:95.22, PWM:32
14:57:59.355 -> Water_Temp.:95.22, PWM:33
14:58:00.151 -> Water_Temp.:95.22, PWM:32
14:58:01.708 -> Water_Temp.:95.22, PWM:32
14:58:02.464 -> Water_Temp.:95.22, PWM:33
14:58:03.257 -> Water_Temp.:95.22, PWM:32
14:58:04.018 -> Water_Temp.:95.22, PWM:32
14:58:04.816 -> Water_Temp.:95.22, PWM:32
14:58:06.373 -> Water_Temp.:95.34, PWM:33
14:58:07.165 -> Water_Temp.:95.34, PWM:33
14:58:09.473 -> Water_Temp.:95.34, PWM:32
14:58:11.026 -> Water_Temp.:95.34, PWM:33
14:58:11.821 -> Water_Temp.:95.34, PWM:33
14:58:12.579 -> Water_Temp.:95.34, PWM:33
14:58:13.376 -> Water_Temp.:95.34, PWM:32
14:58:14.136 -> Water_Temp.:95.34, PWM:33
14:58:14.926 -> Water_Temp.:95.34, PWM:32
14:58:17.238 -> Water_Temp.:95.34, PWM:32
14:58:19.553 -> Water_Temp.:95.45, PWM:32
14:58:20.314 -> Water_Temp.:95.45, PWM:33
14:58:21.868 -> Water_Temp.:95.45, PWM:33
14:58:22.662 -> Water_Temp.:95.45, PWM:32
14:58:24.181 -> Water_Temp.:95.45, PWM:32
14:58:24.971 -> Water_Temp.:95.45, PWM:32
14:58:25.768 -> Water_Temp.:95.45, PWM:32
14:58:26.528 -> Water_Temp.:95.45, PWM:33
14:58:27.318 -> Water_Temp.:95.45, PWM:32
14:58:28.112 -> Water_Temp.:95.45, PWM:32
14:58:29.667 -> Water_Temp.:95.45, PWM:32
14:58:31.190 -> Water_Temp.:95.45, PWM:33
14:58:31.990 -> Water_Temp.:95.45, PWM:33
14:58:32.755 -> Water_Temp.:95.45, PWM:33
14:58:33.553 -> Water_Temp.:95.45, PWM:32
14:58:35.083 -> Water_Temp.:95.34, PWM:33
14:58:35.881 -> Water_Temp.:95.34, PWM:33
14:58:36.641 -> Water_Temp.:95.34, PWM:32
14:58:37.427 -> Water_Temp.:95.34, PWM:33
14:58:38.948 -> Water_Temp.:95.34, PWM:33
14:58:40.503 -> Water_Temp.:95.34, PWM:33
14:58:41.297 -> Water_Temp.:95.34, PWM:32
14:58:42.059 -> Water_Temp.:95.34, PWM:32
14:58:42.853 -> Water_Temp.:95.34, PWM:33
14:58:43.649 -> Water_Temp.:95.22, PWM:32
14:58:44.409 -> Water_Temp.:95.22, PWM:33
14:58:45.203 -> Water_Temp.:95.22, PWM:32
14:58:45.966 -> Water_Temp.:95.22, PWM:32
14:58:46.760 -> Water_Temp.:95.22, PWM:32
14:58:47.551 -> Water_Temp.:95.22, PWM:32
14:58:48.304 -> Water_Temp.:95.22, PWM:32
14:58:49.094 -> Water_Temp.:95.22, PWM:32
14:58:50.613 -> Water_Temp.:95.22, PWM:32
14:58:51.406 -> Water_Temp.:95.22, PWM:33
14:58:55.285 -> Water_Temp.:95.22, PWM:32
14:58:56.838 -> Water_Temp.:95.22, PWM:32
14:58:57.599 -> Water_Temp.:95.22, PWM:32
14:58:58.388 -> Water_Temp.:95.22, PWM:32
14:58:59.148 -> Water_Temp.:95.22, PWM:32
14:59:00.736 -> Water_Temp.:95.22, PWM:32
14:59:02.261 -> Water_Temp.:95.22, PWM:33
14:59:03.056 -> Water_Temp.:95.22, PWM:32
14:59:05.374 -> Water_Temp.:95.22, PWM:32
14:59:07.649 -> Water_Temp.:95.11, PWM:32
14:59:09.205 -> Water_Temp.:95.11, PWM:33
14:59:09.966 -> Water_Temp.:95.11, PWM:32
14:59:10.761 -> Water_Temp.:95.11, PWM:32
14:59:13.084 -> Water_Temp.:96.80, PWM:32
14:59:14.602 -> Water_Temp.:96.46, PWM:33
14:59:15.396 -> Water_Temp.:96.01, PWM:32
14:59:16.191 -> Water_Temp.:95.79, PWM:32
14:59:16.954 -> Water_Temp.:95.68, PWM:32
14:59:18.507 -> Water_Temp.:95.56, PWM:32
14:59:20.821 -> Water_Temp.:95.45, PWM:32
14:59:21.581 -> Water_Temp.:95.34, PWM:32
14:59:22.378 -> Water_Temp.:95.34, PWM:33
14:59:23.171 -> Water_Temp.:95.34, PWM:33
14:59:23.934 -> Water_Temp.:95.34, PWM:33
14:59:24.731 -> Water_Temp.:95.34, PWM:33
14:59:26.247 -> Water_Temp.:95.34, PWM:32
14:59:28.552 -> Water_Temp.:95.34, PWM:32
14:59:29.341 -> Water_Temp.:95.34, PWM:33
14:59:30.102 -> Water_Temp.:95.22, PWM:32
14:59:31.657 -> Water_Temp.:95.34, PWM:32
14:59:32.415 -> Water_Temp.:95.34, PWM:32
14:59:33.213 -> Water_Temp.:95.22, PWM:32
14:59:34.004 -> Water_Temp.:95.34, PWM:33
14:59:34.758 -> Water_Temp.:95.22, PWM:32
14:59:35.554 -> Water_Temp.:95.22, PWM:32
14:59:36.344 -> Water_Temp.:95.34, PWM:33
14:59:37.862 -> Water_Temp.:95.34, PWM:32
14:59:38.655 -> Water_Temp.:95.34, PWM:33
14:59:39.412 -> Water_Temp.:95.34, PWM:32
14:59:40.206 -> Water_Temp.:95.22, PWM:32
14:59:41.761 -> Water_Temp.:95.34, PWM:32
14:59:42.554 -> Water_Temp.:95.34, PWM:32
14:59:44.079 -> Water_Temp.:95.34, PWM:32
14:59:45.629 -> Water_Temp.:95.22, PWM:33
14:59:46.422 -> Water_Temp.:95.34, PWM:33
14:59:47.179 -> Water_Temp.:95.34, PWM:33
14:59:47.973 -> Water_Temp.:95.34, PWM:32
14:59:48.732 -> Water_Temp.:95.34, PWM:33
14:59:49.528 -> Water_Temp.:95.34, PWM:33
14:59:50.291 -> Water_Temp.:95.34, PWM:32
14:59:51.084 -> Water_Temp.:95.34, PWM:32
14:59:54.164 -> Water_Temp.:95.34, PWM:32
14:59:54.925 -> Water_Temp.:95.45, PWM:33
14:59:59.552 -> Water_Temp.:95.45, PWM:33
15:00:02.596 -> Water_Temp.:95.45, PWM:33
15:00:03.389 -> Water_Temp.:95.45, PWM:33
15:00:04.941 -> Water_Temp.:95.45, PWM:33
15:00:05.704 -> Water_Temp.:95.45, PWM:32
15:00:06.497 -> Water_Temp.:95.45, PWM:33
15:00:07.289 -> Water_Temp.:95.45, PWM:33
15:00:08.050 -> Water_Temp.:95.45, PWM:32
15:00:08.814 -> Water_Temp.:95.45, PWM:33
15:00:09.609 -> Water_Temp.:95.45, PWM:32
15:00:10.404 -> Water_Temp.:95.45, PWM:33
15:00:11.163 -> Water_Temp.:95.45, PWM:33
15:00:11.959 -> Water_Temp.:95.45, PWM:33
15:00:14.272 -> Water_Temp.:95.45, PWM:32
15:00:15.032 -> Water_Temp.:95.56, PWM:32
15:00:15.823 -> Water_Temp.:95.56, PWM:32
15:00:16.613 -> Water_Temp.:95.56, PWM:33
15:00:18.137 -> Water_Temp.:95.56, PWM:33
15:00:19.691 -> Water_Temp.:95.56, PWM:33
15:00:20.485 -> Water_Temp.:95.56, PWM:33
15:00:21.248 -> Water_Temp.:95.56, PWM:32
15:00:22.042 -> Water_Temp.:95.56, PWM:33
15:00:23.563 -> Water_Temp.:95.68, PWM:32
15:00:24.356 -> Water_Temp.:95.68, PWM:32
15:00:25.147 -> Water_Temp.:95.68, PWM:33
15:00:25.905 -> Water_Temp.:97.36, PWM:32
15:00:26.702 -> Water_Temp.:99.39, PWM:33
15:00:27.462 -> Water_Temp.:97.02, PWM:33
15:00:28.260 -> Water_Temp.:96.57, PWM:33
15:00:29.025 -> Water_Temp.:96.35, PWM:32
15:00:29.820 -> Water_Temp.:96.24, PWM:33
15:00:31.375 -> Water_Temp.:96.12, PWM:32
15:00:32.929 -> Water_Temp.:96.01, PWM:33
15:00:33.692 -> Water_Temp.:95.90, PWM:32
15:00:35.241 -> Water_Temp.:95.90, PWM:32
15:00:36.040 -> Water_Temp.:95.90, PWM:33
15:00:36.801 -> Water_Temp.:95.90, PWM:32
15:00:38.353 -> Water_Temp.:95.90, PWM:32
15:00:39.147 -> Water_Temp.:95.90, PWM:33
15:00:41.456 -> Water_Temp.:95.79, PWM:32
15:00:42.249 -> Water_Temp.:95.79, PWM:33
15:00:43.007 -> Water_Temp.:95.79, PWM:32
15:00:43.805 -> Water_Temp.:95.79, PWM:32
15:00:46.089 -> Water_Temp.:95.79, PWM:33
15:00:46.885 -> Water_Temp.:95.79, PWM:32
15:00:49.993 -> Water_Temp.:95.79, PWM:32
15:00:50.755 -> Water_Temp.:95.79, PWM:32
15:00:52.309 -> Water_Temp.:95.79, PWM:33
15:00:53.897 -> Water_Temp.:95.79, PWM:32
15:00:54.657 -> Water_Temp.:95.79, PWM:32
15:00:56.204 -> Water_Temp.:95.79, PWM:32
15:00:57.001 -> Water_Temp.:95.79, PWM:33
15:00:58.549 -> Water_Temp.:95.79, PWM:32
15:00:59.337 -> Water_Temp.:95.79, PWM:32
15:01:00.133 -> Water_Temp.:95.79, PWM:33
15:01:00.892 -> Water_Temp.:95.79, PWM:32
15:01:01.683 -> Water_Temp.:95.68, PWM:32
15:01:02.442 -> Water_Temp.:95.68, PWM:32
15:01:04.750 -> Water_Temp.:95.68, PWM:32
15:01:06.306 -> Water_Temp.:95.68, PWM:32
15:01:07.067 -> Water_Temp.:95.68, PWM:32
15:01:07.862 -> Water_Temp.:95.68, PWM:32
15:01:08.651 -> Water_Temp.:95.68, PWM:33
15:01:09.411 -> Water_Temp.:95.68, PWM:32
15:01:10.967 -> Water_Temp.:95.68, PWM:33
15:01:12.522 -> Water_Temp.:95.68, PWM:33
15:01:13.285 -> Water_Temp.:95.68, PWM:33
15:01:14.077 -> Water_Temp.:95.68, PWM:32
15:01:14.869 -> Water_Temp.:95.68, PWM:33
15:01:15.634 -> Water_Temp.:95.68, PWM:32
15:01:17.185 -> Water_Temp.:95.68, PWM:32
15:01:17.979 -> Water_Temp.:95.68, PWM:32
15:01:18.739 -> Water_Temp.:95.68, PWM:32
15:01:19.531 -> Water_Temp.:95.68, PWM:32
15:01:21.081 -> Water_Temp.:95.68, PWM:33
15:01:21.873 -> Water_Temp.:95.68, PWM:33
15:01:22.666 -> Water_Temp.:95.56, PWM:33
15:01:23.426 -> Water_Temp.:95.68, PWM:32
15:01:24.219 -> Water_Temp.:95.68, PWM:32
15:01:24.977 -> Water_Temp.:95.56, PWM:33
15:01:25.772 -> Water_Temp.:95.56, PWM:32
15:01:26.564 -> Water_Temp.:95.68, PWM:33
15:01:27.321 -> Water_Temp.:95.68, PWM:33
15:01:28.117 -> Water_Temp.:95.56, PWM:33
15:01:29.665 -> Water_Temp.:95.56, PWM:32
15:01:30.427 -> Water_Temp.:95.56, PWM:32
15:01:31.222 -> Water_Temp.:95.56, PWM:33
15:01:31.984 -> Water_Temp.:95.56, PWM:32
15:01:33.538 -> Water_Temp.:95.56, PWM:32
15:01:34.333 -> Water_Temp.:95.56, PWM:33
15:01:36.622 -> Water_Temp.:95.56, PWM:33
15:01:38.139 -> Water_Temp.:95.56, PWM:32
15:01:38.928 -> Water_Temp.:95.56, PWM:32
15:01:40.483 -> Water_Temp.:95.56, PWM:33
15:01:41.272 -> Water_Temp.:95.56, PWM:32
15:01:42.030 -> Water_Temp.:95.56, PWM:32
15:01:42.821 -> Water_Temp.:95.56, PWM:32
15:01:43.610 -> Water_Temp.:95.56, PWM:33
15:01:44.403 -> Water_Temp.:95.45, PWM:33
15:01:45.163 -> Water_Temp.:95.45, PWM:32
15:01:45.961 -> Water_Temp.:95.45, PWM:32
15:01:47.514 -> Water_Temp.:95.45, PWM:33
15:01:48.274 -> Water_Temp.:95.45, PWM:33
15:01:49.068 -> Water_Temp.:95.45, PWM:32
15:01:49.827 -> Water_Temp.:95.45, PWM:32
15:01:50.623 -> Water_Temp.:95.45, PWM:32
15:01:52.180 -> Water_Temp.:95.45, PWM:33
15:01:52.977 -> Water_Temp.:95.45, PWM:32
15:01:53.737 -> Water_Temp.:95.45, PWM:32
15:01:54.532 -> Water_Temp.:95.45, PWM:32
15:01:55.295 -> Water_Temp.:95.45, PWM:33
15:01:56.089 -> Water_Temp.:95.45, PWM:33
15:01:57.647 -> Water_Temp.:95.45, PWM:33
15:01:58.436 -> Water_Temp.:95.45, PWM:32
15:01:59.196 -> Water_Temp.:95.45, PWM:32
15:01:59.993 -> Water_Temp.:95.45, PWM:32
15:02:00.754 -> Water_Temp.:95.45, PWM:33
15:02:02.302 -> Water_Temp.:95.45, PWM:33
15:02:03.057 -> Water_Temp.:95.45, PWM:33
15:02:04.609 -> Water_Temp.:95.45, PWM:32
15:02:05.404 -> Water_Temp.:95.45, PWM:32
15:02:06.959 -> Water_Temp.:95.45, PWM:32
15:02:07.721 -> Water_Temp.:95.45, PWM:32
15:02:08.514 -> Water_Temp.:95.45, PWM:33
15:02:09.308 -> Water_Temp.:95.45, PWM:32
15:02:10.064 -> Water_Temp.:95.56, PWM:33
15:02:11.619 -> Water_Temp.:95.56, PWM:32
15:02:13.135 -> Water_Temp.:95.56, PWM:33
15:02:14.688 -> Water_Temp.:95.56, PWM:33
15:02:15.477 -> Water_Temp.:95.56, PWM:32
15:02:16.998 -> Water_Temp.:95.56, PWM:33
15:02:18.551 -> Water_Temp.:95.56, PWM:32
15:02:19.345 -> Water_Temp.:95.56, PWM:32
15:02:20.102 -> Water_Temp.:95.56, PWM:32
15:02:20.895 -> Water_Temp.:95.56, PWM:32
15:02:21.658 -> Water_Temp.:95.56, PWM:32
15:02:22.449 -> Water_Temp.:95.56, PWM:33
15:02:23.243 -> Water_Temp.:95.56, PWM:33
15:02:24.003 -> Water_Temp.:95.56, PWM:32
15:02:25.561 -> Water_Temp.:95.45, PWM:32
15:02:26.352 -> Water_Temp.:95.45, PWM:32
15:02:27.904 -> Water_Temp.:95.45, PWM:32
15:02:30.256 -> Water_Temp.:95.45, PWM:33
15:02:31.806 -> Water_Temp.:95.45, PWM:32
15:02:32.569 -> Water_Temp.:95.45, PWM:32
15:02:33.363 -> Water_Temp.:95.56, PWM:32
15:02:36.463 -> Water_Temp.:95.56, PWM:32
15:02:38.018 -> Water_Temp.:95.56, PWM:32
15:02:38.781 -> Water_Temp.:95.56, PWM:32
15:02:39.571 -> Water_Temp.:95.56, PWM:32
15:02:40.361 -> Water_Temp.:95.56, PWM:32
15:02:41.123 -> Water_Temp.:95.56, PWM:32
15:02:42.671 -> Water_Temp.:95.56, PWM:33
15:02:43.429 -> Water_Temp.:95.56, PWM:33
15:02:45.014 -> Water_Temp.:95.56, PWM:32
15:02:45.774 -> Water_Temp.:95.56, PWM:32
15:02:46.563 -> Water_Temp.:95.56, PWM:33
15:02:48.078 -> Water_Temp.:95.56, PWM:33
15:02:48.871 -> Water_Temp.:95.56, PWM:32
15:02:49.665 -> Water_Temp.:95.56, PWM:32
15:02:51.180 -> Water_Temp.:95.56, PWM:33
15:02:55.809 -> Water_Temp.:95.56, PWM:32
15:02:57.358 -> Water_Temp.:95.56, PWM:33
15:02:58.120 -> Water_Temp.:95.56, PWM:33
15:03:00.439 -> Water_Temp.:95.56, PWM:32
15:03:01.231 -> Water_Temp.:95.56, PWM:32
15:03:02.024 -> Water_Temp.:95.56, PWM:33
15:03:02.788 -> Water_Temp.:97.25, PWM:32
15:03:03.583 -> Water_Temp.:99.27, PWM:33
15:03:04.339 -> Water_Temp.:96.80, PWM:33
15:03:05.134 -> Water_Temp.:96.46, PWM:32
15:03:05.895 -> Water_Temp.:96.24, PWM:32
15:03:07.448 -> Water_Temp.:96.01, PWM:33
15:03:08.244 -> Water_Temp.:95.90, PWM:32
15:03:09.002 -> Water_Temp.:95.90, PWM:32
15:03:09.796 -> Water_Temp.:95.79, PWM:32
15:03:10.555 -> Water_Temp.:95.79, PWM:32
15:03:11.351 -> Water_Temp.:95.68, PWM:33
15:03:12.143 -> Water_Temp.:95.68, PWM:33
15:03:12.904 -> Water_Temp.:95.68, PWM:32
15:03:14.457 -> Water_Temp.:95.68, PWM:33
15:03:15.211 -> Water_Temp.:95.68, PWM:32
15:03:16.768 -> Water_Temp.:95.68, PWM:32
15:03:17.562 -> Water_Temp.:95.56, PWM:33
15:03:18.355 -> Water_Temp.:95.56, PWM:32
15:03:19.116 -> Water_Temp.:95.56, PWM:32
15:03:22.951 -> Water_Temp.:95.56, PWM:33
15:03:24.501 -> Water_Temp.:95.56, PWM:32
15:03:26.059 -> Water_Temp.:95.56, PWM:32
15:03:26.853 -> Water_Temp.:95.56, PWM:32
15:03:29.163 -> Water_Temp.:95.56, PWM:33
15:03:29.956 -> Water_Temp.:95.56, PWM:33
15:03:30.749 -> Water_Temp.:95.56, PWM:33
15:03:31.506 -> Water_Temp.:95.56, PWM:32
15:03:32.296 -> Water_Temp.:95.56, PWM:32
15:03:33.089 -> Water_Temp.:95.56, PWM:33
15:03:35.430 -> Water_Temp.:95.56, PWM:32
15:03:36.189 -> Water_Temp.:95.56, PWM:33
15:03:37.766 -> Water_Temp.:95.56, PWM:32
15:03:38.522 -> Water_Temp.:95.56, PWM:32
15:03:39.315 -> Water_Temp.:95.56, PWM:33
15:03:40.108 -> Water_Temp.:95.56, PWM:32
15:03:40.865 -> Water_Temp.:95.56, PWM:33
15:03:41.653 -> Water_Temp.:95.56, PWM:32
15:03:44.745 -> Water_Temp.:95.45, PWM:33
15:03:45.542 -> Water_Temp.:95.45, PWM:33
15:03:46.300 -> Water_Temp.:95.45, PWM:33
15:03:47.092 -> Water_Temp.:95.45, PWM:32
15:03:47.881 -> Water_Temp.:95.45, PWM:32
15:03:52.462 -> Water_Temp.:95.45, PWM:32
15:03:53.253 -> Water_Temp.:95.45, PWM:32
15:03:54.007 -> Water_Temp.:95.45, PWM:32
15:03:54.801 -> Water_Temp.:95.45, PWM:32
15:03:55.593 -> Water_Temp.:95.45, PWM:32
15:03:57.145 -> Water_Temp.:95.45, PWM:33
15:03:58.698 -> Water_Temp.:95.45, PWM:33
15:04:01.763 -> Water_Temp.:95.45, PWM:33
15:04:02.557 -> Water_Temp.:95.45, PWM:32
15:04:04.109 -> Water_Temp.:95.45, PWM:33
15:04:04.900 -> Water_Temp.:95.45, PWM:32
15:04:05.660 -> Water_Temp.:95.45, PWM:33
15:04:06.454 -> Water_Temp.:95.45, PWM:33
15:04:08.766 -> Water_Temp.:95.45, PWM:33
15:04:09.558 -> Water_Temp.:95.45, PWM:32
15:04:11.082 -> Water_Temp.:95.45, PWM:33
15:04:11.876 -> Water_Temp.:95.45, PWM:32
15:04:12.630 -> Water_Temp.:95.45, PWM:32
15:04:13.426 -> Water_Temp.:95.45, PWM:33
15:04:15.739 -> Water_Temp.:95.45, PWM:33
15:04:16.501 -> Water_Temp.:95.45, PWM:32
15:04:17.292 -> Water_Temp.:95.45, PWM:32
15:04:18.841 -> Water_Temp.:95.45, PWM:33
15:04:19.596 -> Water_Temp.:95.45, PWM:32
15:04:20.389 -> Water_Temp.:95.45, PWM:33
15:04:21.181 -> Water_Temp.:95.45, PWM:33
15:04:23.492 -> Water_Temp.:95.45, PWM:33
15:04:25.074 -> Water_Temp.:95.45, PWM:32
15:04:25.832 -> Water_Temp.:95.45, PWM:32
15:04:26.627 -> Water_Temp.:95.45, PWM:32
15:04:28.145 -> Water_Temp.:95.45, PWM:33
15:04:28.940 -> Water_Temp.:95.45, PWM:32
15:04:30.496 -> Water_Temp.:95.45, PWM:32
15:04:31.292 -> Water_Temp.:95.45, PWM:32
15:04:32.053 -> Water_Temp.:95.45, PWM:33
15:04:32.844 -> Water_Temp.:95.45, PWM:32
15:04:34.369 -> Water_Temp.:95.45, PWM:32
15:04:36.715 -> Water_Temp.:95.45, PWM:33
15:04:37.506 -> Water_Temp.:95.45, PWM:36
15:04:38.268 -> Water_Temp.:95.45, PWM:36
15:04:39.065 -> Water_Temp.:95.45, PWM:37
15:04:40.588 -> Water_Temp.:95.45, PWM:36
15:04:41.348 -> Water_Temp.:95.45, PWM:34
15:04:42.140 -> Water_Temp.:95.45, PWM:34
15:04:42.937 -> Water_Temp.:95.45, PWM:34
15:04:43.694 -> Water_Temp.:95.45, PWM:34
15:04:45.244 -> Water_Temp.:95.45, PWM:34
15:04:47.588 -> Water_Temp.:95.45, PWM:34
15:04:48.382 -> Water_Temp.:95.45, PWM:34
15:04:49.142 -> Water_Temp.:95.45, PWM:34
15:04:49.934 -> Water_Temp.:95.45, PWM:34
15:04:51.490 -> Water_Temp.:95.45, PWM:34
15:04:52.250 -> Water_Temp.:95.45, PWM:34
15:04:53.043 -> Water_Temp.:95.45, PWM:34
15:04:53.837 -> Water_Temp.:95.45, PWM:34
15:04:56.115 -> Water_Temp.:95.34, PWM:34
15:04:56.908 -> Water_Temp.:95.34, PWM:34
15:04:57.667 -> Water_Temp.:95.34, PWM:34
15:04:58.462 -> Water_Temp.:95.34, PWM:34
15:04:59.251 -> Water_Temp.:95.34, PWM:34
15:05:00.008 -> Water_Temp.:95.34, PWM:34
15:05:01.559 -> Water_Temp.:95.34, PWM:34
15:05:03.874 -> Water_Temp.:95.34, PWM:34
15:05:06.952 -> Water_Temp.:95.34, PWM:34
15:05:07.714 -> Water_Temp.:95.34, PWM:34
15:05:08.508 -> Water_Temp.:95.34, PWM:34
15:05:09.267 -> Water_Temp.:95.34, PWM:34
15:05:10.061 -> Water_Temp.:95.34, PWM:34
15:05:10.856 -> Water_Temp.:95.34, PWM:34
15:05:11.615 -> Water_Temp.:95.34, PWM:34
15:05:12.410 -> Water_Temp.:95.34, PWM:34
15:05:14.721 -> Water_Temp.:95.45, PWM:34
15:05:16.274 -> Water_Temp.:95.45, PWM:34
15:05:18.590 -> Water_Temp.:95.45, PWM:34
15:05:20.906 -> Water_Temp.:95.45, PWM:34
15:05:22.460 -> Water_Temp.:95.45, PWM:34
15:05:23.252 -> Water_Temp.:95.45, PWM:34
15:05:24.801 -> Water_Temp.:95.45, PWM:34
15:05:26.392 -> Water_Temp.:95.45, PWM:34
15:05:27.908 -> Water_Temp.:95.45, PWM:34
15:05:28.708 -> Water_Temp.:95.45, PWM:34
15:05:31.048 -> Water_Temp.:95.45, PWM:34
15:05:32.563 -> Water_Temp.:95.45, PWM:34
15:05:33.359 -> Water_Temp.:95.45, PWM:34
15:05:34.118 -> Water_Temp.:95.45, PWM:34
15:05:34.915 -> Water_Temp.:95.45, PWM:34
15:05:36.469 -> Water_Temp.:95.45, PWM:34
15:05:37.230 -> Water_Temp.:95.45, PWM:34
15:05:38.780 -> Water_Temp.:95.45, PWM:34
15:05:39.571 -> Water_Temp.:95.45, PWM:34
15:05:41.124 -> Water_Temp.:95.45, PWM:34
15:05:41.885 -> Water_Temp.:95.45, PWM:34
15:05:43.435 -> Water_Temp.:95.45, PWM:34
15:05:44.230 -> Water_Temp.:95.45, PWM:34
15:05:44.988 -> Water_Temp.:95.45, PWM:34
15:05:47.338 -> Water_Temp.:95.45, PWM:34
15:05:48.130 -> Water_Temp.:95.45, PWM:34
15:05:50.405 -> Water_Temp.:95.45, PWM:34
15:05:51.199 -> Water_Temp.:95.45, PWM:34
15:05:51.958 -> Water_Temp.:95.45, PWM:34
15:05:52.751 -> Water_Temp.:95.45, PWM:34
15:05:53.511 -> Water_Temp.:95.45, PWM:34
15:05:54.306 -> Water_Temp.:95.45, PWM:34
15:05:55.862 -> Water_Temp.:95.45, PWM:34
15:05:56.620 -> Water_Temp.:95.56, PWM:34
15:05:58.971 -> Water_Temp.:95.56, PWM:34
15:06:00.523 -> Water_Temp.:95.56, PWM:34
15:06:01.284 -> Water_Temp.:95.56, PWM:34
15:06:02.079 -> Water_Temp.:95.56, PWM:34
15:06:02.875 -> Water_Temp.:95.56, PWM:34
15:06:03.635 -> Water_Temp.:95.56, PWM:34
15:06:04.429 -> Water_Temp.:95.56, PWM:34
15:06:05.187 -> Water_Temp.:95.56, PWM:34
15:06:05.981 -> Water_Temp.:95.56, PWM:34
15:06:07.536 -> Water_Temp.:95.56, PWM:34
15:06:10.647 -> Water_Temp.:95.56, PWM:34
15:06:12.959 -> Water_Temp.:95.56, PWM:34
15:06:13.754 -> Water_Temp.:95.56, PWM:34
15:06:14.512 -> Water_Temp.:95.56, PWM:34
15:06:15.305 -> Water_Temp.:95.56, PWM:34
15:06:16.063 -> Water_Temp.:95.56, PWM:34
15:06:17.618 -> Water_Temp.:95.56, PWM:34
15:06:18.413 -> Water_Temp.:95.56, PWM:34
15:06:19.173 -> Water_Temp.:95.56, PWM:34
15:06:19.966 -> Water_Temp.:95.56, PWM:34
15:06:21.517 -> Water_Temp.:95.56, PWM:34
15:06:22.276 -> Water_Temp.:95.56, PWM:34
15:06:24.621 -> Water_Temp.:95.56, PWM:34
15:06:25.379 -> Water_Temp.:95.56, PWM:34
15:06:26.170 -> Water_Temp.:95.56, PWM:34
15:06:27.726 -> Water_Temp.:95.56, PWM:34
15:06:28.520 -> Water_Temp.:95.56, PWM:34
15:06:29.278 -> Water_Temp.:95.56, PWM:34
15:06:30.068 -> Water_Temp.:95.56, PWM:34
15:06:30.826 -> Water_Temp.:95.56, PWM:34
15:06:35.460 -> Water_Temp.:95.68, PWM:34
15:06:36.256 -> Water_Temp.:95.68, PWM:34
15:06:37.016 -> Water_Temp.:95.68, PWM:34
15:06:37.811 -> Water_Temp.:95.68, PWM:34
15:06:38.606 -> Water_Temp.:95.56, PWM:34
15:06:39.368 -> Water_Temp.:95.56, PWM:34
15:06:40.163 -> Water_Temp.:95.56, PWM:34
15:06:41.720 -> Water_Temp.:95.56, PWM:34
15:06:42.481 -> Water_Temp.:95.56, PWM:34
15:06:43.276 -> Water_Temp.:95.56, PWM:34
15:06:44.039 -> Water_Temp.:95.56, PWM:34
15:06:46.347 -> Water_Temp.:95.56, PWM:34
15:06:47.936 -> Water_Temp.:95.56, PWM:34
15:06:49.491 -> Water_Temp.:95.56, PWM:34
15:06:51.010 -> Water_Temp.:95.56, PWM:34
15:06:51.802 -> Water_Temp.:95.56, PWM:34
15:06:53.354 -> Water_Temp.:95.56, PWM:34
15:06:55.670 -> Water_Temp.:95.56, PWM:34
15:06:58.768 -> Water_Temp.:95.45, PWM:34
15:07:00.290 -> Water_Temp.:95.45, PWM:34
15:07:01.054 -> Water_Temp.:95.45, PWM:34
15:07:02.607 -> Water_Temp.:95.45, PWM:34
15:07:03.365 -> Water_Temp.:95.45, PWM:34
15:07:04.158 -> Water_Temp.:95.45, PWM:34
15:07:04.951 -> Water_Temp.:95.45, PWM:34
15:07:05.711 -> Water_Temp.:95.45, PWM:34
15:07:06.502 -> Water_Temp.:95.45, PWM:34
15:07:07.259 -> Water_Temp.:95.45, PWM:34
15:07:08.055 -> Water_Temp.:95.45, PWM:34
15:07:08.849 -> Water_Temp.:95.45, PWM:34
15:07:09.606 -> Water_Temp.:95.45, PWM:34
15:07:10.403 -> Water_Temp.:95.45, PWM:34
15:07:12.719 -> Water_Temp.:95.45, PWM:34
15:07:13.510 -> Water_Temp.:95.45, PWM:34
15:07:14.267 -> Water_Temp.:95.45, PWM:34
15:07:17.345 -> Water_Temp.:95.56, PWM:34
15:07:18.141 -> Water_Temp.:95.45, PWM:34
15:07:18.899 -> Water_Temp.:95.56, PWM:34
15:07:21.970 -> Water_Temp.:95.45, PWM:34
15:07:22.770 -> Water_Temp.:95.45, PWM:34
15:07:25.814 -> Water_Temp.:95.45, PWM:34
15:07:28.132 -> Water_Temp.:95.56, PWM:34
15:07:28.893 -> Water_Temp.:95.56, PWM:34
15:07:29.689 -> Water_Temp.:95.56, PWM:34
15:07:31.210 -> Water_Temp.:95.56, PWM:34
15:07:32.006 -> Water_Temp.:95.56, PWM:34
15:07:34.290 -> Water_Temp.:95.56, PWM:34
15:07:35.086 -> Water_Temp.:95.56, PWM:34
15:07:35.847 -> Water_Temp.:95.56, PWM:34
15:07:37.402 -> Water_Temp.:95.56, PWM:34
15:07:38.194 -> Water_Temp.:95.56, PWM:34
15:07:38.952 -> Water_Temp.:95.56, PWM:34
15:07:39.749 -> Water_Temp.:95.56, PWM:34
15:07:41.273 -> Water_Temp.:95.56, PWM:34
15:07:42.065 -> Water_Temp.:95.56, PWM:34
15:07:42.825 -> Water_Temp.:95.56, PWM:34
15:07:43.620 -> Water_Temp.:95.56, PWM:34
15:07:44.411 -> Water_Temp.:95.56, PWM:34
15:07:45.170 -> Water_Temp.:95.56, PWM:34
15:07:47.514 -> Water_Temp.:95.56, PWM:34
15:07:48.305 -> Water_Temp.:95.56, PWM:34
15:07:49.064 -> Water_Temp.:95.56, PWM:34
15:07:49.860 -> Water_Temp.:95.56, PWM:34
15:07:50.621 -> Water_Temp.:95.56, PWM:34
15:07:52.173 -> Water_Temp.:95.56, PWM:34
15:07:52.967 -> Water_Temp.:95.56, PWM:34
15:07:53.725 -> Water_Temp.:95.56, PWM:34
15:07:54.519 -> Water_Temp.:95.56, PWM:34
15:07:55.311 -> Water_Temp.:95.56, PWM:34
15:07:56.073 -> Water_Temp.:95.56, PWM:34
15:07:56.870 -> Water_Temp.:95.68, PWM:34
15:07:57.627 -> Water_Temp.:95.56, PWM:34
15:07:59.181 -> Water_Temp.:95.68, PWM:34
15:07:59.979 -> Water_Temp.:95.68, PWM:34
15:08:01.535 -> Water_Temp.:95.68, PWM:34
15:08:02.297 -> Water_Temp.:95.68, PWM:34
15:08:03.092 -> Water_Temp.:95.68, PWM:34
15:08:05.405 -> Water_Temp.:95.68, PWM:34
15:08:06.200 -> Water_Temp.:95.68, PWM:34
15:08:07.749 -> Water_Temp.:95.68, PWM:34
15:08:08.536 -> Water_Temp.:95.68, PWM:34
15:08:09.299 -> Water_Temp.:95.68, PWM:34
15:08:10.094 -> Water_Temp.:95.68, PWM:34
15:08:13.164 -> Water_Temp.:95.68, PWM:34
15:08:13.925 -> Water_Temp.:95.68, PWM:34
15:08:14.717 -> Water_Temp.:95.68, PWM:34
15:08:15.509 -> Water_Temp.:95.68, PWM:34
15:08:16.270 -> Water_Temp.:95.68, PWM:34
15:08:17.063 -> Water_Temp.:95.56, PWM:34
15:08:17.827 -> Water_Temp.:95.56, PWM:34
15:08:20.173 -> Water_Temp.:95.56, PWM:34
15:08:20.931 -> Water_Temp.:95.56, PWM:34
15:08:21.727 -> Water_Temp.:95.56, PWM:34
15:08:23.276 -> Water_Temp.:95.56, PWM:34
15:08:24.068 -> Water_Temp.:95.56, PWM:34
15:08:25.590 -> Water_Temp.:95.68, PWM:34
15:08:26.384 -> Water_Temp.:95.68, PWM:34
15:08:27.945 -> Water_Temp.:95.68, PWM:34
15:08:28.705 -> Water_Temp.:95.68, PWM:34
15:08:29.502 -> Water_Temp.:95.68, PWM:34
15:08:30.297 -> Water_Temp.:95.68, PWM:34
15:08:31.059 -> Water_Temp.:95.68, PWM:34
15:08:32.615 -> Water_Temp.:95.68, PWM:34
15:08:33.411 -> Water_Temp.:95.68, PWM:34
15:08:34.171 -> Water_Temp.:95.68, PWM:34
15:08:34.968 -> Water_Temp.:95.68, PWM:34
15:08:36.526 -> Water_Temp.:95.68, PWM:34
15:08:37.283 -> Water_Temp.:95.68, PWM:34
15:08:38.869 -> Water_Temp.:95.68, PWM:34
15:08:40.424 -> Water_Temp.:95.68, PWM:34
15:08:41.979 -> Water_Temp.:95.79, PWM:34
15:08:43.534 -> Water_Temp.:95.79, PWM:34
15:08:44.294 -> Water_Temp.:95.79, PWM:34
15:08:45.845 -> Water_Temp.:95.79, PWM:34
15:08:46.639 -> Water_Temp.:95.79, PWM:34
15:08:48.161 -> Water_Temp.:95.79, PWM:34
15:08:48.920 -> Water_Temp.:95.79, PWM:34
15:08:50.503 -> Water_Temp.:95.68, PWM:34
15:08:52.058 -> Water_Temp.:95.79, PWM:34
15:08:52.852 -> Water_Temp.:95.79, PWM:34
15:08:58.237 -> Water_Temp.:95.79, PWM:34
15:08:58.995 -> Water_Temp.:95.79, PWM:34
15:09:01.335 -> Water_Temp.:95.79, PWM:34
15:09:02.127 -> Water_Temp.:95.79, PWM:34
15:09:04.450 -> Water_Temp.:95.79, PWM:34
15:09:06.037 -> Water_Temp.:95.79, PWM:34
15:09:07.586 -> Water_Temp.:95.79, PWM:34
15:09:08.376 -> Water_Temp.:95.79, PWM:34
15:09:09.137 -> Water_Temp.:95.79, PWM:34
15:09:11.446 -> Water_Temp.:95.79, PWM:34
15:09:12.206 -> Water_Temp.:95.79, PWM:34
15:09:13.755 -> Water_Temp.:95.79, PWM:34
15:09:14.546 -> Water_Temp.:95.79, PWM:34
15:09:16.102 -> Water_Temp.:95.79, PWM:34
15:09:16.859 -> Water_Temp.:95.79, PWM:34
15:09:17.654 -> Water_Temp.:95.79, PWM:34
15:09:18.415 -> Water_Temp.:95.79, PWM:34
15:09:20.000 -> Water_Temp.:95.79, PWM:34
15:09:20.762 -> Water_Temp.:95.79, PWM:34
15:09:23.067 -> Water_Temp.:95.79, PWM:34
15:09:25.386 -> Water_Temp.:95.79, PWM:34
15:09:26.181 -> Water_Temp.:95.79, PWM:34
15:09:26.939 -> Water_Temp.:95.79, PWM:34
15:09:27.732 -> Water_Temp.:95.79, PWM:34
15:09:28.522 -> Water_Temp.:95.79, PWM:34
15:09:29.283 -> Water_Temp.:95.79, PWM:34
15:09:30.077 -> Water_Temp.:95.79, PWM:34
15:09:30.839 -> Water_Temp.:95.79, PWM:34
15:09:32.422 -> Water_Temp.:95.79, PWM:34
15:09:33.180 -> Water_Temp.:95.79, PWM:34
15:09:34.734 -> Water_Temp.:95.79, PWM:34
15:09:37.837 -> Water_Temp.:95.79, PWM:34
15:09:38.597 -> Water_Temp.:95.79, PWM:34
15:09:40.908 -> Water_Temp.:95.79, PWM:34
15:09:41.701 -> Water_Temp.:95.79, PWM:34
15:09:42.496 -> Water_Temp.:95.90, PWM:34
15:09:43.258 -> Water_Temp.:95.90, PWM:34
15:09:44.055 -> Water_Temp.:95.90, PWM:34
15:09:45.575 -> Water_Temp.:95.90, PWM:34
15:09:47.128 -> Water_Temp.:95.90, PWM:34
15:09:47.887 -> Water_Temp.:95.90, PWM:34
15:09:48.683 -> Water_Temp.:95.90, PWM:34
15:09:50.239 -> Water_Temp.:95.90, PWM:34
15:09:51.003 -> Water_Temp.:95.90, PWM:34
15:09:54.077 -> Water_Temp.:95.90, PWM:34
15:09:54.835 -> Water_Temp.:95.90, PWM:34
15:09:57.952 -> Water_Temp.:95.90, PWM:34
15:09:58.712 -> Water_Temp.:95.90, PWM:34
15:09:59.504 -> Water_Temp.:95.90, PWM:34
15:10:01.813 -> Water_Temp.:95.90, PWM:34
15:10:02.607 -> Water_Temp.:95.90, PWM:34
15:10:04.127 -> Water_Temp.:95.90, PWM:34
15:10:06.448 -> Water_Temp.:95.90, PWM:34
15:10:08.007 -> Water_Temp.:95.90, PWM:34
15:10:08.770 -> Water_Temp.:95.90, PWM:34
15:10:11.088 -> Water_Temp.:95.90, PWM:34
15:10:11.848 -> Water_Temp.:95.90, PWM:34
15:10:14.953 -> Water_Temp.:95.90, PWM:34
15:10:16.477 -> Water_Temp.:95.90, PWM:34
15:10:17.272 -> Water_Temp.:95.90, PWM:34
15:10:18.032 -> Water_Temp.:95.90, PWM:34
15:10:19.589 -> Water_Temp.:95.90, PWM:34
15:10:21.936 -> Water_Temp.:95.90, PWM:34
15:10:22.693 -> Water_Temp.:95.90, PWM:34
15:10:23.490 -> Water_Temp.:95.90, PWM:34
15:10:24.250 -> Water_Temp.:95.90, PWM:34
15:10:25.044 -> Water_Temp.:95.90, PWM:34
15:10:25.841 -> Water_Temp.:95.90, PWM:34
15:10:27.392 -> Water_Temp.:95.90, PWM:34
15:10:28.154 -> Water_Temp.:95.90, PWM:34
15:10:28.948 -> Water_Temp.:95.90, PWM:34
15:10:29.709 -> Water_Temp.:95.90, PWM:34
15:10:30.502 -> Water_Temp.:95.90, PWM:34
15:10:31.294 -> Water_Temp.:95.90, PWM:34
15:10:34.369 -> Water_Temp.:95.90, PWM:34
15:10:35.161 -> Water_Temp.:95.90, PWM:34
15:10:36.715 -> Water_Temp.:95.90, PWM:34
15:10:38.274 -> Water_Temp.:96.01, PWM:34
15:10:41.344 -> Water_Temp.:96.01, PWM:34
15:10:43.690 -> Water_Temp.:96.01, PWM:34
15:10:45.244 -> Water_Temp.:96.01, PWM:34
15:10:46.794 -> Water_Temp.:96.01, PWM:34
15:10:47.588 -> Water_Temp.:96.01, PWM:34
15:10:48.344 -> Water_Temp.:96.01, PWM:34
15:10:49.136 -> Water_Temp.:96.01, PWM:34
15:10:50.655 -> Water_Temp.:96.01, PWM:34
15:10:51.446 -> Water_Temp.:96.01, PWM:34
15:10:52.993 -> Water_Temp.:96.01, PWM:34
15:10:53.787 -> Water_Temp.:96.01, PWM:34
15:10:55.336 -> Water_Temp.:95.90, PWM:34
15:10:56.130 -> Water_Temp.:95.90, PWM:34
15:10:56.894 -> Water_Temp.:95.90, PWM:34
15:10:57.683 -> Water_Temp.:95.90, PWM:34
15:10:59.962 -> Water_Temp.:95.90, PWM:34
15:11:00.722 -> Water_Temp.:95.90, PWM:34
15:11:02.313 -> Water_Temp.:95.90, PWM:34
15:11:03.076 -> Water_Temp.:95.90, PWM:34
15:11:03.872 -> Water_Temp.:96.01, PWM:34
15:11:06.154 -> Water_Temp.:96.01, PWM:34
15:11:06.948 -> Water_Temp.:96.01, PWM:34
15:11:08.502 -> Water_Temp.:96.01, PWM:34
15:11:09.265 -> Water_Temp.:96.01, PWM:34
15:11:10.060 -> Water_Temp.:96.01, PWM:34
15:11:10.823 -> Water_Temp.:96.01, PWM:34
15:11:11.621 -> Water_Temp.:96.01, PWM:34
15:11:13.143 -> Water_Temp.:96.01, PWM:34
15:11:14.694 -> Water_Temp.:96.01, PWM:34
15:11:15.490 -> Water_Temp.:96.01, PWM:34
15:11:17.047 -> Water_Temp.:95.90, PWM:34
15:11:17.809 -> Water_Temp.:96.01, PWM:34
15:11:19.362 -> Water_Temp.:95.90, PWM:34
15:11:20.123 -> Water_Temp.:95.90, PWM:34
15:11:20.917 -> Water_Temp.:95.90, PWM:34
15:11:22.472 -> Water_Temp.:96.01, PWM:34
15:11:27.868 -> Water_Temp.:96.01, PWM:34
15:11:30.182 -> Water_Temp.:96.01, PWM:34
15:11:30.945 -> Water_Temp.:96.01, PWM:34
15:11:31.735 -> Water_Temp.:96.01, PWM:34
15:11:32.528 -> Water_Temp.:96.01, PWM:34
15:11:36.329 -> Water_Temp.:96.01, PWM:34
15:11:37.124 -> Water_Temp.:96.01, PWM:34
15:11:37.919 -> Water_Temp.:96.01, PWM:34
15:11:38.683 -> Water_Temp.:96.01, PWM:34
15:11:39.474 -> Water_Temp.:96.01, PWM:34
15:11:40.237 -> Water_Temp.:96.01, PWM:34
15:11:41.036 -> Water_Temp.:96.01, PWM:34
15:11:41.796 -> Water_Temp.:96.01, PWM:34
15:11:42.586 -> Water_Temp.:96.01, PWM:34
15:11:44.933 -> Water_Temp.:96.01, PWM:34
15:11:46.489 -> Water_Temp.:95.90, PWM:34
15:11:47.251 -> Water_Temp.:95.90, PWM:34
15:11:48.043 -> Water_Temp.:95.90, PWM:34
15:11:49.556 -> Water_Temp.:95.90, PWM:34
15:11:51.908 -> Water_Temp.:95.90, PWM:34
15:11:52.705 -> Water_Temp.:95.90, PWM:34
15:11:53.465 -> Water_Temp.:95.90, PWM:34
15:11:54.261 -> Water_Temp.:95.90, PWM:34
15:11:55.022 -> Water_Temp.:95.90, PWM:34
15:11:56.575 -> Water_Temp.:95.90, PWM:34
15:11:57.368 -> Water_Temp.:95.90, PWM:34
15:11:58.161 -> Water_Temp.:95.90, PWM:34
15:11:59.674 -> Water_Temp.:95.90, PWM:34
15:12:00.470 -> Water_Temp.:95.90, PWM:34
15:12:01.231 -> Water_Temp.:95.90, PWM:34
15:12:02.789 -> Water_Temp.:95.90, PWM:34
15:12:04.344 -> Water_Temp.:95.79, PWM:34
15:12:05.138 -> Water_Temp.:95.79, PWM:34
15:12:05.897 -> Water_Temp.:95.79, PWM:34
15:12:06.693 -> Water_Temp.:95.79, PWM:34
15:12:07.455 -> Water_Temp.:95.90, PWM:34
15:12:08.249 -> Water_Temp.:95.79, PWM:34
15:12:09.809 -> Water_Temp.:95.79, PWM:34
15:12:10.571 -> Water_Temp.:95.90, PWM:34
15:12:11.372 -> Water_Temp.:95.79, PWM:34
15:12:14.440 -> Water_Temp.:95.79, PWM:34
15:12:15.230 -> Water_Temp.:95.79, PWM:34
15:12:15.991 -> Water_Temp.:95.79, PWM:34
15:12:17.544 -> Water_Temp.:95.79, PWM:34
15:12:18.332 -> Water_Temp.:95.79, PWM:34
15:12:19.125 -> Water_Temp.:95.79, PWM:34
15:12:20.646 -> Water_Temp.:95.79, PWM:34
15:12:21.443 -> Water_Temp.:95.79, PWM:34
15:12:22.202 -> Water_Temp.:95.79, PWM:34
15:12:23.001 -> Water_Temp.:95.79, PWM:34
15:12:23.794 -> Water_Temp.:95.79, PWM:34
15:12:24.558 -> Water_Temp.:95.79, PWM:34
15:12:25.352 -> Water_Temp.:95.79, PWM:34
15:12:26.112 -> Water_Temp.:95.79, PWM:34
15:12:26.908 -> Water_Temp.:95.79, PWM:34
15:12:27.665 -> Water_Temp.:95.79, PWM:34
15:12:28.456 -> Water_Temp.:95.79, PWM:34
15:12:29.251 -> Water_Temp.:95.79, PWM:34
15:12:30.007 -> Water_Temp.:95.79, PWM:34
15:12:32.292 -> Water_Temp.:95.79, PWM:34
15:12:33.088 -> Water_Temp.:95.79, PWM:34
15:12:34.639 -> Water_Temp.:95.79, PWM:34
15:12:35.431 -> Water_Temp.:95.79, PWM:34
15:12:36.191 -> Water_Temp.:95.79, PWM:34
15:12:37.744 -> Water_Temp.:95.79, PWM:34
15:12:38.506 -> Water_Temp.:95.79, PWM:34
15:12:39.299 -> Water_Temp.:95.79, PWM:34
15:12:40.096 -> Water_Temp.:95.79, PWM:34
15:12:41.654 -> Water_Temp.:95.79, PWM:34
15:12:42.416 -> Water_Temp.:95.79, PWM:34
15:12:43.208 -> Water_Temp.:95.79, PWM:34
15:12:43.971 -> Water_Temp.:95.79, PWM:34
15:12:44.765 -> Water_Temp.:95.79, PWM:34
15:12:46.316 -> Water_Temp.:95.79, PWM:34
15:12:47.112 -> Water_Temp.:95.79, PWM:34
15:12:47.872 -> Water_Temp.:95.79, PWM:34
15:12:48.665 -> Water_Temp.:95.79, PWM:34
15:12:50.971 -> Water_Temp.:95.79, PWM:34
15:12:51.767 -> Water_Temp.:95.79, PWM:34
15:12:52.561 -> Water_Temp.:95.79, PWM:34
15:12:53.319 -> Water_Temp.:95.79, PWM:34
15:12:54.118 -> Water_Temp.:95.79, PWM:34
15:12:57.193 -> Water_Temp.:95.79, PWM:34
15:12:57.950 -> Water_Temp.:95.79, PWM:34
15:12:58.743 -> Water_Temp.:95.79, PWM:34
15:12:59.502 -> Water_Temp.:95.79, PWM:34
15:13:01.056 -> Water_Temp.:95.79, PWM:34
15:13:01.850 -> Water_Temp.:95.79, PWM:34
15:13:02.641 -> Water_Temp.:95.79, PWM:34
15:13:03.430 -> Water_Temp.:95.68, PWM:34
15:13:05.738 -> Water_Temp.:95.68, PWM:34
15:13:07.286 -> Water_Temp.:95.68, PWM:34
15:13:08.081 -> Water_Temp.:95.68, PWM:34
15:13:08.842 -> Water_Temp.:95.68, PWM:34
15:13:09.637 -> Water_Temp.:95.68, PWM:34
15:13:11.156 -> Water_Temp.:95.68, PWM:34
15:13:11.950 -> Water_Temp.:95.68, PWM:34
15:13:15.812 -> Water_Temp.:95.68, PWM:34
15:13:16.607 -> Water_Temp.:95.68, PWM:34
15:13:18.131 -> Water_Temp.:95.68, PWM:34
15:13:18.929 -> Water_Temp.:95.68, PWM:34
15:13:19.691 -> Water_Temp.:95.68, PWM:34
15:13:20.487 -> Water_Temp.:95.68, PWM:34
15:13:21.245 -> Water_Temp.:95.68, PWM:34
15:13:23.562 -> Water_Temp.:95.68, PWM:34
15:13:24.351 -> Water_Temp.:95.68, PWM:34
15:13:25.908 -> Water_Temp.:95.68, PWM:34
15:13:26.705 -> Water_Temp.:95.68, PWM:34
15:13:27.465 -> Water_Temp.:95.68, PWM:34
15:13:28.255 -> Water_Temp.:95.68, PWM:34
15:13:29.809 -> Water_Temp.:95.68, PWM:34
15:13:31.367 -> Water_Temp.:95.68, PWM:34
15:13:32.128 -> Water_Temp.:95.68, PWM:34
15:13:32.919 -> Water_Temp.:95.68, PWM:34
15:13:33.712 -> Water_Temp.:95.68, PWM:34
15:13:34.472 -> Water_Temp.:95.68, PWM:34
15:13:35.260 -> Water_Temp.:95.79, PWM:34
15:13:36.022 -> Water_Temp.:95.79, PWM:34
15:13:36.813 -> Water_Temp.:95.79, PWM:34
15:13:39.884 -> Water_Temp.:95.79, PWM:34
15:13:40.678 -> Water_Temp.:95.79, PWM:34
15:13:41.437 -> Water_Temp.:95.79, PWM:34
15:13:42.989 -> Water_Temp.:95.79, PWM:34
15:13:43.749 -> Water_Temp.:95.79, PWM:34
15:13:44.543 -> Water_Temp.:95.79, PWM:34
15:13:45.335 -> Water_Temp.:95.79, PWM:34
15:13:46.859 -> Water_Temp.:95.79, PWM:34
15:13:47.654 -> Water_Temp.:95.79, PWM:34
15:13:48.417 -> Water_Temp.:95.79, PWM:34
15:13:52.286 -> Water_Temp.:95.79, PWM:34
15:13:54.600 -> Water_Temp.:95.79, PWM:34
15:13:55.398 -> Water_Temp.:95.79, PWM:34
15:13:56.155 -> Water_Temp.:95.79, PWM:34
15:13:57.719 -> Water_Temp.:95.79, PWM:34
15:13:58.510 -> Water_Temp.:95.79, PWM:34
15:13:59.307 -> Water_Temp.:95.79, PWM:34
15:14:00.067 -> Water_Temp.:95.79, PWM:34
15:14:00.863 -> Water_Temp.:95.79, PWM:34
15:14:02.379 -> Water_Temp.:95.79, PWM:34
15:14:03.937 -> Water_Temp.:95.79, PWM:34
15:14:06.290 -> Water_Temp.:95.79, PWM:34
15:14:07.045 -> Water_Temp.:95.79, PWM:34
15:14:07.840 -> Water_Temp.:95.79, PWM:34
15:14:12.463 -> Water_Temp.:95.79, PWM:34
15:14:13.986 -> Water_Temp.:95.79, PWM:34
15:14:14.782 -> Water_Temp.:95.79, PWM:34
15:14:16.302 -> Water_Temp.:95.90, PWM:34
15:14:17.094 -> Water_Temp.:95.90, PWM:34
15:14:17.886 -> Water_Temp.:95.90, PWM:34
15:14:18.646 -> Water_Temp.:95.90, PWM:34
15:14:19.446 -> Water_Temp.:95.90, PWM:34
15:14:20.209 -> Water_Temp.:95.90, PWM:34
15:14:21.006 -> Water_Temp.:95.90, PWM:34
15:14:21.765 -> Water_Temp.:95.90, PWM:34
15:14:23.349 -> Water_Temp.:95.90, PWM:34
15:14:24.107 -> Water_Temp.:95.90, PWM:34
15:14:25.663 -> Water_Temp.:95.90, PWM:34
15:14:26.457 -> Water_Temp.:95.90, PWM:34
15:14:27.251 -> Water_Temp.:95.90, PWM:34
15:14:28.769 -> Water_Temp.:95.90, PWM:34
15:14:29.566 -> Water_Temp.:95.90, PWM:34
15:14:30.324 -> Water_Temp.:95.90, PWM:34
15:14:31.882 -> Water_Temp.:95.90, PWM:34
15:14:32.643 -> Water_Temp.:95.90, PWM:34
15:14:34.201 -> Water_Temp.:95.90, PWM:34
15:14:34.960 -> Water_Temp.:95.90, PWM:34
15:14:35.755 -> Water_Temp.:95.90, PWM:34
15:14:36.518 -> Water_Temp.:95.90, PWM:34
15:14:37.315 -> Water_Temp.:95.90, PWM:34
15:14:38.077 -> Water_Temp.:95.90, PWM:34
15:14:38.872 -> Water_Temp.:95.90, PWM:34
15:14:40.420 -> Water_Temp.:95.90, PWM:34
15:14:41.180 -> Water_Temp.:95.90, PWM:34
15:14:42.735 -> Water_Temp.:95.90, PWM:34
15:14:43.528 -> Water_Temp.:95.90, PWM:34
15:14:44.324 -> Water_Temp.:95.90, PWM:34
15:14:45.851 -> Water_Temp.:95.90, PWM:34
15:14:47.407 -> Water_Temp.:95.90, PWM:34
15:14:49.724 -> Water_Temp.:95.90, PWM:34
15:14:50.518 -> Water_Temp.:95.90, PWM:34
15:14:52.830 -> Water_Temp.:95.90, PWM:34
15:14:53.624 -> Water_Temp.:95.90, PWM:34
15:14:54.420 -> Water_Temp.:95.90, PWM:34
15:14:55.179 -> Water_Temp.:95.90, PWM:34
15:14:57.492 -> Water_Temp.:95.90, PWM:34
15:14:59.800 -> Water_Temp.:95.90, PWM:34
15:15:01.355 -> Water_Temp.:95.90, PWM:34
15:15:02.900 -> Water_Temp.:95.90, PWM:34
15:15:03.692 -> Water_Temp.:95.90, PWM:34
15:15:05.240 -> Water_Temp.:95.90, PWM:34
15:15:06.001 -> Water_Temp.:95.90, PWM:34
15:15:06.794 -> Water_Temp.:95.90, PWM:34
15:15:08.348 -> Water_Temp.:95.90, PWM:34
15:15:09.141 -> Water_Temp.:95.90, PWM:34
15:15:09.900 -> Water_Temp.:95.90, PWM:34
15:15:10.691 -> Water_Temp.:95.90, PWM:34
15:15:11.475 -> Water_Temp.:95.90, PWM:34
15:15:13.014 -> Water_Temp.:95.90, PWM:34
15:15:13.803 -> Water_Temp.:95.90, PWM:34
15:15:15.348 -> Water_Temp.:95.90, PWM:34
15:15:16.144 -> Water_Temp.:95.90, PWM:34
15:15:17.662 -> Water_Temp.:95.90, PWM:34
15:15:18.451 -> Water_Temp.:95.90, PWM:34
15:15:21.557 -> Water_Temp.:95.90, PWM:34
15:15:22.347 -> Water_Temp.:95.90, PWM:34
15:15:25.454 -> Water_Temp.:95.90, PWM:34
15:15:26.246 -> Water_Temp.:95.90, PWM:34
15:15:27.005 -> Water_Temp.:95.90, PWM:34
15:15:28.562 -> Water_Temp.:95.90, PWM:34
15:15:29.355 -> Water_Temp.:97.59, PWM:34
15:15:31.673 -> Water_Temp.:97.25, PWM:34
15:15:32.430 -> Water_Temp.:96.69, PWM:34
15:15:33.222 -> Water_Temp.:96.46, PWM:34
15:15:34.778 -> Water_Temp.:96.35, PWM:34
15:15:35.576 -> Water_Temp.:96.24, PWM:34
15:15:36.341 -> Water_Temp.:96.24, PWM:34
15:15:37.137 -> Water_Temp.:96.24, PWM:34
15:15:37.901 -> Water_Temp.:96.12, PWM:34
15:15:39.455 -> Water_Temp.:96.12, PWM:34
15:15:40.244 -> Water_Temp.:96.12, PWM:34
15:15:41.003 -> Water_Temp.:96.12, PWM:34
15:15:42.556 -> Water_Temp.:96.12, PWM:34
15:15:43.315 -> Water_Temp.:96.12, PWM:34
15:15:44.109 -> Water_Temp.:96.12, PWM:34
15:15:45.621 -> Water_Temp.:96.12, PWM:34
15:15:46.412 -> Water_Temp.:96.12, PWM:34
15:15:47.198 -> Water_Temp.:96.12, PWM:34
15:15:51.816 -> Water_Temp.:96.12, PWM:34
15:15:52.576 -> Water_Temp.:96.12, PWM:34
15:15:53.364 -> Water_Temp.:96.12, PWM:34
15:15:54.154 -> Water_Temp.:96.12, PWM:34
15:15:55.714 -> Water_Temp.:96.12, PWM:34
15:15:57.269 -> Water_Temp.:96.12, PWM:34
15:15:58.028 -> Water_Temp.:96.12, PWM:34
15:15:58.824 -> Water_Temp.:96.12, PWM:34
15:15:59.620 -> Water_Temp.:96.01, PWM:34
15:16:01.138 -> Water_Temp.:96.12, PWM:34
15:16:02.683 -> Water_Temp.:96.01, PWM:34
15:16:03.476 -> Water_Temp.:96.12, PWM:34
15:16:04.269 -> Water_Temp.:96.12, PWM:34
15:16:05.027 -> Water_Temp.:96.01, PWM:34
15:16:05.817 -> Water_Temp.:96.01, PWM:34
15:16:07.336 -> Water_Temp.:96.01, PWM:34
15:16:08.124 -> Water_Temp.:96.01, PWM:34
15:16:08.914 -> Water_Temp.:96.01, PWM:34
15:16:09.671 -> Water_Temp.:96.01, PWM:34
15:16:11.227 -> Water_Temp.:96.01, PWM:34
15:16:12.021 -> Water_Temp.:96.01, PWM:34
15:16:13.572 -> Water_Temp.:96.01, PWM:34
15:16:14.331 -> Water_Temp.:96.01, PWM:34
15:16:15.882 -> Water_Temp.:96.01, PWM:34
15:16:16.677 -> Water_Temp.:96.01, PWM:34
15:16:18.232 -> Water_Temp.:96.01, PWM:34
15:16:21.303 -> Water_Temp.:97.70, PWM:34
15:16:22.842 -> Water_Temp.:97.25, PWM:34
15:16:23.633 -> Water_Temp.:96.91, PWM:34
15:16:24.421 -> Water_Temp.:96.91, PWM:34
15:16:25.180 -> Water_Temp.:96.46, PWM:34
15:16:25.974 -> Water_Temp.:96.46, PWM:34
15:16:27.513 -> Water_Temp.:96.35, PWM:34
15:16:29.060 -> Water_Temp.:96.24, PWM:34
15:16:29.846 -> Water_Temp.:96.24, PWM:34
15:16:30.644 -> Water_Temp.:96.12, PWM:34
15:16:31.401 -> Water_Temp.:96.12, PWM:34
15:16:32.192 -> Water_Temp.:96.12, PWM:34
15:16:33.711 -> Water_Temp.:96.12, PWM:34
15:16:35.260 -> Water_Temp.:96.01, PWM:34
15:16:36.807 -> Water_Temp.:96.01, PWM:34
15:16:38.355 -> Water_Temp.:96.01, PWM:34
15:16:39.905 -> Water_Temp.:96.01, PWM:34
15:16:40.696 -> Water_Temp.:96.01, PWM:34
15:16:41.454 -> Water_Temp.:96.01, PWM:34
15:16:42.243 -> Water_Temp.:96.01, PWM:34
15:16:43.789 -> Water_Temp.:96.01, PWM:34
15:16:44.580 -> Water_Temp.:96.01, PWM:34
15:16:46.131 -> Water_Temp.:96.01, PWM:34
15:16:46.920 -> Water_Temp.:96.01, PWM:34
15:16:47.712 -> Water_Temp.:96.01, PWM:34
15:16:50.044 -> Water_Temp.:96.01, PWM:34
15:16:50.803 -> Water_Temp.:96.01, PWM:34
15:16:52.379 -> Water_Temp.:96.01, PWM:34
15:16:53.133 -> Water_Temp.:96.01, PWM:34
15:16:53.917 -> Water_Temp.:96.01, PWM:34
15:16:55.454 -> Water_Temp.:96.01, PWM:34
15:16:57.031 -> Water_Temp.:96.01, PWM:34
15:16:59.330 -> Water_Temp.:96.01, PWM:34
15:17:00.117 -> Water_Temp.:96.01, PWM:34
15:17:03.198 -> Water_Temp.:99.72, PWM:34
15:17:03.985 -> Water_Temp.:97.36, PWM:34
15:17:06.267 -> Water_Temp.:96.69, PWM:34
15:17:07.061 -> Water_Temp.:96.46, PWM:34
15:17:07.855 -> Water_Temp.:96.35, PWM:34
15:17:09.377 -> Water_Temp.:96.35, PWM:34
15:17:11.688 -> Water_Temp.:96.24, PWM:34
15:17:12.478 -> Water_Temp.:96.24, PWM:34
15:17:13.239 -> Water_Temp.:96.24, PWM:34
15:17:14.795 -> Water_Temp.:96.24, PWM:34
15:17:16.370 -> Water_Temp.:96.12, PWM:34
15:17:18.653 -> Water_Temp.:96.12, PWM:34
15:17:20.212 -> Water_Temp.:96.12, PWM:34
15:17:21.763 -> Water_Temp.:96.12, PWM:34
15:17:22.522 -> Water_Temp.:96.12, PWM:34
15:17:23.315 -> Water_Temp.:96.12, PWM:34
15:17:24.109 -> Water_Temp.:96.12, PWM:34
15:17:24.867 -> Water_Temp.:96.12, PWM:34
15:17:25.661 -> Water_Temp.:96.12, PWM:34
15:17:26.418 -> Water_Temp.:96.12, PWM:34
15:17:27.217 -> Water_Temp.:96.12, PWM:34
15:17:28.779 -> Water_Temp.:96.01, PWM:34
15:17:29.543 -> Water_Temp.:96.01, PWM:34
15:17:30.337 -> Water_Temp.:96.01, PWM:34
15:17:32.659 -> Water_Temp.:96.01, PWM:34
15:17:33.453 -> Water_Temp.:96.01, PWM:34
15:17:34.245 -> Water_Temp.:97.59, PWM:34
15:17:35.001 -> Water_Temp.:99.72, PWM:34
15:17:38.110 -> Water_Temp.:96.57, PWM:34
15:17:38.872 -> Water_Temp.:96.46, PWM:34
15:17:39.669 -> Water_Temp.:96.35, PWM:34
15:17:40.462 -> Water_Temp.:97.92, PWM:34
15:17:42.786 -> Water_Temp.:97.14, PWM:34
15:17:44.344 -> Water_Temp.:96.69, PWM:34
15:17:46.650 -> Water_Temp.:96.46, PWM:34
15:17:47.446 -> Water_Temp.:96.35, PWM:34
15:17:48.205 -> Water_Temp.:96.35, PWM:34
15:17:49.727 -> Water_Temp.:96.24, PWM:34
15:17:50.524 -> Water_Temp.:96.24, PWM:34
15:17:51.319 -> Water_Temp.:96.24, PWM:34
15:17:52.079 -> Water_Temp.:96.24, PWM:34
15:17:54.397 -> Water_Temp.:96.12, PWM:34
15:17:55.188 -> Water_Temp.:96.12, PWM:34
15:17:57.503 -> Water_Temp.:96.12, PWM:34
15:17:58.295 -> Water_Temp.:96.12, PWM:34
15:18:02.135 -> Water_Temp.:96.12, PWM:34
15:18:02.929 -> Water_Temp.:96.12, PWM:34
15:18:03.693 -> Water_Temp.:96.12, PWM:34
15:18:04.486 -> Water_Temp.:96.12, PWM:34
15:18:05.247 -> Water_Temp.:96.12, PWM:34
15:18:06.795 -> Water_Temp.:96.12, PWM:34
15:18:07.589 -> Water_Temp.:96.12, PWM:34
15:18:08.345 -> Water_Temp.:96.12, PWM:34
15:18:09.139 -> Water_Temp.:96.12, PWM:34
15:18:09.902 -> Water_Temp.:96.12, PWM:34
15:18:10.696 -> Water_Temp.:96.12, PWM:34
15:18:11.458 -> Water_Temp.:96.12, PWM:34
15:18:12.256 -> Water_Temp.:96.12, PWM:34
15:18:13.018 -> Water_Temp.:96.01, PWM:34
15:18:13.814 -> Water_Temp.:96.12, PWM:34
15:18:15.331 -> Water_Temp.:96.12, PWM:34
15:18:17.643 -> Water_Temp.:96.12, PWM:34
15:18:23.027 -> Water_Temp.:96.12, PWM:34
15:18:26.093 -> Water_Temp.:96.12, PWM:34
15:18:26.888 -> Water_Temp.:96.12, PWM:34
15:18:28.445 -> Water_Temp.:96.12, PWM:34
15:18:29.208 -> Water_Temp.:96.12, PWM:34
15:18:30.005 -> Water_Temp.:96.12, PWM:34
15:18:30.765 -> Water_Temp.:96.12, PWM:34
15:18:33.908 -> Water_Temp.:96.12, PWM:34
15:18:34.668 -> Water_Temp.:96.12, PWM:34
15:18:36.228 -> Water_Temp.:96.12, PWM:34
15:18:37.779 -> Water_Temp.:96.12, PWM:34
15:18:38.573 -> Water_Temp.:96.12, PWM:34
15:18:40.090 -> Water_Temp.:96.12, PWM:34
15:18:40.883 -> Water_Temp.:96.12, PWM:34
15:18:43.229 -> Water_Temp.:96.12, PWM:34
15:18:45.550 -> Water_Temp.:96.12, PWM:34
15:18:47.069 -> Water_Temp.:96.01, PWM:34
15:18:47.861 -> Water_Temp.:96.12, PWM:34
15:18:49.387 -> Water_Temp.:96.12, PWM:34
15:18:50.185 -> Water_Temp.:96.01, PWM:34
15:18:50.943 -> Water_Temp.:96.12, PWM:34
15:18:51.733 -> Water_Temp.:96.01, PWM:34
15:18:54.013 -> Water_Temp.:96.12, PWM:34
15:18:54.800 -> Water_Temp.:96.12, PWM:34
15:18:56.359 -> Water_Temp.:96.12, PWM:34
15:19:00.197 -> Water_Temp.:96.12, PWM:34
15:19:00.957 -> Water_Temp.:96.12, PWM:34
15:19:01.753 -> Water_Temp.:96.12, PWM:34
15:19:03.279 -> Water_Temp.:96.12, PWM:34
15:19:04.068 -> Water_Temp.:96.12, PWM:34
15:19:06.380 -> Water_Temp.:96.12, PWM:34
15:19:07.171 -> Water_Temp.:96.12, PWM:34
15:19:08.727 -> Water_Temp.:96.24, PWM:34
15:19:10.274 -> Water_Temp.:96.12, PWM:34
15:19:11.033 -> Water_Temp.:96.12, PWM:34
15:19:12.583 -> Water_Temp.:96.24, PWM:34
15:19:13.348 -> Water_Temp.:96.24, PWM:34
15:19:14.143 -> Water_Temp.:96.24, PWM:34
15:19:14.903 -> Water_Temp.:96.24, PWM:34
15:19:16.461 -> Water_Temp.:96.24, PWM:34
15:19:17.254 -> Water_Temp.:96.24, PWM:34
15:19:18.012 -> Water_Temp.:96.24, PWM:34
15:19:18.802 -> Water_Temp.:96.24, PWM:34
15:19:21.149 -> Water_Temp.:96.24, PWM:34
15:19:21.907 -> Water_Temp.:96.24, PWM:34
15:19:23.466 -> Water_Temp.:96.24, PWM:34
15:19:24.262 -> Water_Temp.:96.24, PWM:34
15:19:25.816 -> Water_Temp.:96.24, PWM:34
15:19:26.580 -> Water_Temp.:96.24, PWM:34
15:19:27.378 -> Water_Temp.:96.24, PWM:34
15:19:28.139 -> Water_Temp.:96.24, PWM:34
15:19:28.931 -> Water_Temp.:96.24, PWM:34
15:19:29.692 -> Water_Temp.:96.24, PWM:34
15:19:30.491 -> Water_Temp.:96.24, PWM:34
15:19:31.250 -> Water_Temp.:96.24, PWM:34
15:19:32.043 -> Water_Temp.:96.24, PWM:34
15:19:32.833 -> Water_Temp.:96.24, PWM:34
15:19:33.595 -> Water_Temp.:96.24, PWM:34
15:19:35.143 -> Water_Temp.:96.24, PWM:34
15:19:35.934 -> Water_Temp.:96.24, PWM:34
15:19:36.697 -> Water_Temp.:96.24, PWM:34
15:19:37.494 -> Water_Temp.:96.24, PWM:34
15:19:39.797 -> Water_Temp.:96.24, PWM:34
15:19:40.585 -> Water_Temp.:96.24, PWM:34
15:19:42.133 -> Water_Temp.:96.24, PWM:34
15:19:42.930 -> Water_Temp.:96.24, PWM:34
15:19:43.689 -> Water_Temp.:96.24, PWM:34
15:19:45.237 -> Water_Temp.:96.12, PWM:34
15:19:46.035 -> Water_Temp.:96.12, PWM:34
15:19:46.797 -> Water_Temp.:96.12, PWM:34
15:19:47.590 -> Water_Temp.:96.12, PWM:34
15:19:48.343 -> Water_Temp.:96.12, PWM:34
15:19:50.684 -> Water_Temp.:96.12, PWM:34
15:19:52.995 -> Water_Temp.:96.12, PWM:34
15:19:53.775 -> Water_Temp.:96.12, PWM:34
15:19:55.355 -> Water_Temp.:96.12, PWM:34
15:19:56.120 -> Water_Temp.:96.12, PWM:34
15:19:56.903 -> Water_Temp.:96.12, PWM:34
15:19:57.689 -> Water_Temp.:96.12, PWM:34
15:19:58.480 -> Water_Temp.:96.12, PWM:34
15:20:00.781 -> Water_Temp.:96.01, PWM:34
15:20:02.329 -> Water_Temp.:96.12, PWM:34
15:20:03.117 -> Water_Temp.:96.12, PWM:34
15:20:03.878 -> Water_Temp.:96.12, PWM:34
15:20:04.668 -> Water_Temp.:96.12, PWM:34
15:20:05.460 -> Water_Temp.:96.12, PWM:34
15:20:06.224 -> Water_Temp.:96.12, PWM:34
15:20:07.015 -> Water_Temp.:96.12, PWM:34
15:20:07.772 -> Water_Temp.:96.12, PWM:34
15:20:08.565 -> Water_Temp.:96.12, PWM:34
15:20:09.360 -> Water_Temp.:96.12, PWM:34
15:20:10.120 -> Water_Temp.:96.12, PWM:34
15:20:11.675 -> Water_Temp.:96.12, PWM:34
15:20:13.234 -> Water_Temp.:96.12, PWM:34
15:20:15.518 -> Water_Temp.:96.12, PWM:34
15:20:17.070 -> Water_Temp.:96.12, PWM:34
15:20:17.865 -> Water_Temp.:96.12, PWM:34
15:20:20.148 -> Water_Temp.:96.12, PWM:34
15:20:20.940 -> Water_Temp.:96.12, PWM:34
15:20:21.700 -> Water_Temp.:96.12, PWM:34
15:20:22.497 -> Water_Temp.:96.12, PWM:34
15:20:23.256 -> Water_Temp.:96.12, PWM:34
15:20:24.047 -> Water_Temp.:96.12, PWM:34
15:20:24.840 -> Water_Temp.:96.12, PWM:34
15:20:26.363 -> Water_Temp.:96.12, PWM:34
15:20:27.159 -> Water_Temp.:96.12, PWM:34
15:20:27.917 -> Water_Temp.:96.12, PWM:34
15:20:28.714 -> Water_Temp.:96.12, PWM:34
15:20:29.471 -> Water_Temp.:96.12, PWM:34
15:20:31.024 -> Water_Temp.:96.12, PWM:34
15:20:32.580 -> Water_Temp.:96.12, PWM:34
15:20:33.342 -> Water_Temp.:96.12, PWM:34
15:20:34.139 -> Water_Temp.:96.12, PWM:34
15:20:34.898 -> Water_Temp.:96.12, PWM:34
15:20:35.693 -> Water_Temp.:96.12, PWM:34
15:20:36.454 -> Water_Temp.:96.12, PWM:34
15:20:37.248 -> Water_Temp.:96.12, PWM:34
15:20:38.041 -> Water_Temp.:96.12, PWM:34
15:20:39.559 -> Water_Temp.:96.12, PWM:34
15:20:40.356 -> Water_Temp.:96.12, PWM:34
15:20:41.114 -> Water_Temp.:96.12, PWM:34
15:20:41.906 -> Water_Temp.:96.12, PWM:34
15:20:43.461 -> Water_Temp.:96.12, PWM:34
15:20:46.566 -> Water_Temp.:96.12, PWM:34
15:20:48.117 -> Water_Temp.:96.12, PWM:34
15:20:48.877 -> Water_Temp.:96.12, PWM:34
15:20:50.463 -> Water_Temp.:96.12, PWM:34
15:20:51.225 -> Water_Temp.:96.12, PWM:34
15:20:52.016 -> Water_Temp.:96.12, PWM:34
15:20:52.777 -> Water_Temp.:96.12, PWM:34
15:20:54.330 -> Water_Temp.:96.12, PWM:34
15:20:55.090 -> Water_Temp.:96.12, PWM:34
15:20:55.884 -> Water_Temp.:96.12, PWM:34
15:20:57.435 -> Water_Temp.:96.12, PWM:36
15:20:58.231 -> Water_Temp.:96.12, PWM:37
15:20:58.986 -> Water_Temp.:96.12, PWM:37
15:20:59.774 -> Water_Temp.:96.12, PWM:37
15:21:00.568 -> Water_Temp.:96.12, PWM:39
15:21:02.085 -> Water_Temp.:96.12, PWM:40
15:21:04.430 -> Water_Temp.:96.12, PWM:40
15:21:07.540 -> Water_Temp.:96.12, PWM:39
15:21:09.088 -> Water_Temp.:96.12, PWM:40
15:21:09.849 -> Water_Temp.:96.12, PWM:39
15:21:10.641 -> Water_Temp.:96.12, PWM:39
15:21:12.957 -> Water_Temp.:96.12, PWM:39
15:21:14.511 -> Water_Temp.:96.12, PWM:40
15:21:16.830 -> Water_Temp.:96.12, PWM:39
15:21:17.591 -> Water_Temp.:96.24, PWM:40
15:21:18.387 -> Water_Temp.:96.24, PWM:40
15:21:19.147 -> Water_Temp.:96.24, PWM:39
15:21:19.937 -> Water_Temp.:96.24, PWM:40
15:21:20.730 -> Water_Temp.:96.24, PWM:39
15:21:21.487 -> Water_Temp.:96.24, PWM:40
15:21:22.277 -> Water_Temp.:96.24, PWM:40
15:21:23.038 -> Water_Temp.:96.24, PWM:40
15:21:24.595 -> Water_Temp.:96.24, PWM:40
15:21:25.351 -> Water_Temp.:96.24, PWM:40
15:21:28.479 -> Water_Temp.:96.12, PWM:40
15:21:30.025 -> Water_Temp.:96.12, PWM:40
15:21:30.822 -> Water_Temp.:96.12, PWM:40
15:21:31.579 -> Water_Temp.:96.12, PWM:40
15:21:33.165 -> Water_Temp.:96.12, PWM:40
15:21:33.922 -> Water_Temp.:96.12, PWM:39
15:21:36.231 -> Water_Temp.:96.12, PWM:40
15:21:36.989 -> Water_Temp.:96.24, PWM:40
15:21:37.784 -> Water_Temp.:96.24, PWM:39
15:21:38.543 -> Water_Temp.:96.24, PWM:39
15:21:39.341 -> Water_Temp.:96.24, PWM:39
15:21:40.102 -> Water_Temp.:96.24, PWM:40
15:21:41.659 -> Water_Temp.:96.24, PWM:40
15:21:42.450 -> Water_Temp.:96.24, PWM:39
15:21:43.211 -> Water_Temp.:96.24, PWM:39
15:21:44.004 -> Water_Temp.:97.92, PWM:39
15:21:45.556 -> Water_Temp.:97.47, PWM:39
15:21:47.076 -> Water_Temp.:96.91, PWM:40
15:21:47.869 -> Water_Temp.:96.80, PWM:39
15:21:49.426 -> Water_Temp.:96.57, PWM:40
15:21:50.215 -> Water_Temp.:96.57, PWM:40
15:21:51.765 -> Water_Temp.:96.46, PWM:39
15:21:53.313 -> Water_Temp.:96.35, PWM:39
15:21:54.863 -> Water_Temp.:96.35, PWM:40
15:21:55.626 -> Water_Temp.:96.35, PWM:39
15:21:56.425 -> Water_Temp.:96.35, PWM:40
15:21:57.183 -> Water_Temp.:96.35, PWM:40
15:21:58.734 -> Water_Temp.:96.35, PWM:39
15:21:59.527 -> Water_Temp.:96.24, PWM:39
15:22:00.322 -> Water_Temp.:96.24, PWM:40
15:22:02.666 -> Water_Temp.:96.24, PWM:40
15:22:03.426 -> Water_Temp.:96.24, PWM:40
15:22:07.290 -> Water_Temp.:96.24, PWM:39
15:22:08.840 -> Water_Temp.:96.24, PWM:39
15:22:09.635 -> Water_Temp.:96.24, PWM:39
15:22:10.396 -> Water_Temp.:96.24, PWM:39
15:22:13.493 -> Water_Temp.:96.24, PWM:39
15:22:14.249 -> Water_Temp.:96.24, PWM:39
15:22:15.043 -> Water_Temp.:96.24, PWM:39
15:22:15.839 -> Water_Temp.:96.24, PWM:39
15:22:16.598 -> Water_Temp.:96.24, PWM:39
15:22:17.394 -> Water_Temp.:96.24, PWM:39
15:22:18.154 -> Water_Temp.:96.24, PWM:39
15:22:18.949 -> Water_Temp.:96.24, PWM:40
15:22:20.470 -> Water_Temp.:96.24, PWM:40
15:22:21.263 -> Water_Temp.:96.24, PWM:39
15:22:22.057 -> Water_Temp.:96.24, PWM:40
15:22:22.816 -> Water_Temp.:96.24, PWM:40
15:22:23.606 -> Water_Temp.:96.24, PWM:39
15:22:24.400 -> Water_Temp.:96.24, PWM:39
15:22:25.157 -> Water_Temp.:96.24, PWM:40
15:22:25.952 -> Water_Temp.:96.24, PWM:39
15:22:28.262 -> Water_Temp.:96.24, PWM:40
15:22:29.053 -> Water_Temp.:96.24, PWM:39
15:22:29.813 -> Water_Temp.:96.24, PWM:39
15:22:32.120 -> Water_Temp.:96.24, PWM:39
15:22:33.675 -> Water_Temp.:96.24, PWM:39
15:22:34.433 -> Water_Temp.:96.24, PWM:40
15:22:35.231 -> Water_Temp.:96.24, PWM:40
15:22:36.751 -> Water_Temp.:96.24, PWM:39
15:22:37.547 -> Water_Temp.:96.24, PWM:39
15:22:38.306 -> Water_Temp.:96.24, PWM:39
15:22:39.098 -> Water_Temp.:96.24, PWM:39
15:22:39.892 -> Water_Temp.:96.24, PWM:39
15:22:40.652 -> Water_Temp.:96.24, PWM:40
15:22:41.446 -> Water_Temp.:96.24, PWM:39
15:22:42.207 -> Water_Temp.:96.24, PWM:40
15:22:43.000 -> Water_Temp.:96.24, PWM:39
15:22:43.791 -> Water_Temp.:96.24, PWM:39
15:22:44.548 -> Water_Temp.:96.24, PWM:40
15:22:45.343 -> Water_Temp.:96.24, PWM:40
15:22:46.104 -> Water_Temp.:96.24, PWM:39
15:22:46.901 -> Water_Temp.:96.24, PWM:39
15:22:47.693 -> Water_Temp.:96.24, PWM:40
15:22:48.449 -> Water_Temp.:96.24, PWM:39
15:22:50.007 -> Water_Temp.:96.35, PWM:39
15:22:52.317 -> Water_Temp.:96.35, PWM:40
15:22:53.111 -> Water_Temp.:96.35, PWM:40
15:22:55.390 -> Water_Temp.:96.35, PWM:39
15:22:56.185 -> Water_Temp.:96.35, PWM:40
15:22:56.976 -> Water_Temp.:96.35, PWM:40
15:22:57.739 -> Water_Temp.:96.35, PWM:40
15:22:58.531 -> Water_Temp.:96.35, PWM:39
15:22:59.292 -> Water_Temp.:96.35, PWM:40
15:23:00.085 -> Water_Temp.:96.35, PWM:40
15:23:00.848 -> Water_Temp.:96.46, PWM:39
15:23:02.406 -> Water_Temp.:96.46, PWM:40
15:23:03.955 -> Water_Temp.:96.46, PWM:40
15:23:04.748 -> Water_Temp.:96.46, PWM:39
15:23:06.269 -> Water_Temp.:96.46, PWM:39
15:23:07.819 -> Water_Temp.:96.46, PWM:39
15:23:08.614 -> Water_Temp.:96.46, PWM:39
15:23:09.371 -> Water_Temp.:96.46, PWM:40
15:23:10.164 -> Water_Temp.:96.46, PWM:40
15:23:11.716 -> Water_Temp.:96.46, PWM:40
15:23:14.026 -> Water_Temp.:96.46, PWM:39
15:23:16.341 -> Water_Temp.:96.46, PWM:39
15:23:17.136 -> Water_Temp.:96.46, PWM:40
15:23:18.688 -> Water_Temp.:96.46, PWM:40
15:23:19.450 -> Water_Temp.:96.46, PWM:39
15:23:20.244 -> Water_Temp.:96.46, PWM:39
15:23:21.764 -> Water_Temp.:96.46, PWM:39
15:23:24.074 -> Water_Temp.:96.46, PWM:39
15:23:26.390 -> Water_Temp.:96.46, PWM:40
15:23:27.152 -> Water_Temp.:96.57, PWM:39
15:23:29.463 -> Water_Temp.:96.57, PWM:39
15:23:30.259 -> Water_Temp.:96.57, PWM:40
15:23:31.022 -> Water_Temp.:96.57, PWM:39
15:23:31.818 -> Water_Temp.:96.57, PWM:40
15:23:32.579 -> Water_Temp.:96.57, PWM:40
15:23:33.376 -> Water_Temp.:96.57, PWM:39
15:23:34.923 -> Water_Temp.:96.69, PWM:39
15:23:35.685 -> Water_Temp.:96.57, PWM:39
15:23:36.478 -> Water_Temp.:96.69, PWM:40
15:23:37.271 -> Water_Temp.:96.69, PWM:40
15:23:38.032 -> Water_Temp.:96.69, PWM:39
15:23:39.586 -> Water_Temp.:96.69, PWM:40
15:23:40.347 -> Water_Temp.:96.69, PWM:39
15:23:41.138 -> Water_Temp.:96.69, PWM:39
15:23:43.409 -> Water_Temp.:96.69, PWM:40
15:23:44.205 -> Water_Temp.:96.69, PWM:40
15:23:45.728 -> Water_Temp.:96.69, PWM:39
15:23:48.083 -> Water_Temp.:96.69, PWM:39
15:23:48.841 -> Water_Temp.:96.69, PWM:39
15:23:50.435 -> Water_Temp.:96.69, PWM:39
15:23:51.193 -> Water_Temp.:96.69, PWM:40
15:23:51.987 -> Water_Temp.:96.69, PWM:39
15:23:52.747 -> Water_Temp.:96.69, PWM:40
15:23:54.303 -> Water_Temp.:96.69, PWM:39
15:23:57.387 -> Water_Temp.:96.69, PWM:40
15:23:58.142 -> Water_Temp.:96.69, PWM:40
15:23:59.689 -> Water_Temp.:96.69, PWM:40
15:24:00.484 -> Water_Temp.:96.69, PWM:39
15:24:01.242 -> Water_Temp.:96.69, PWM:40
15:24:02.032 -> Water_Temp.:96.69, PWM:40
15:24:02.822 -> Water_Temp.:96.69, PWM:40
15:24:03.579 -> Water_Temp.:96.69, PWM:39
15:24:04.374 -> Water_Temp.:96.69, PWM:39
15:24:05.166 -> Water_Temp.:96.69, PWM:40
15:24:05.928 -> Water_Temp.:96.69, PWM:39
15:24:06.720 -> Water_Temp.:96.69, PWM:39
15:24:08.239 -> Water_Temp.:96.69, PWM:39
15:24:09.034 -> Water_Temp.:96.69, PWM:40
15:24:09.831 -> Water_Temp.:96.69, PWM:39
15:24:10.595 -> Water_Temp.:96.69, PWM:39
15:24:11.388 -> Water_Temp.:96.69, PWM:40
15:24:12.147 -> Water_Temp.:96.69, PWM:39
15:24:12.943 -> Water_Temp.:96.69, PWM:39
15:24:13.703 -> Water_Temp.:96.69, PWM:39
15:24:15.255 -> Water_Temp.:96.69, PWM:40
15:24:16.017 -> Water_Temp.:96.69, PWM:39
15:24:16.810 -> Water_Temp.:96.69, PWM:39
15:24:18.325 -> Water_Temp.:96.69, PWM:39
15:24:19.114 -> Water_Temp.:96.69, PWM:39
15:24:19.907 -> Water_Temp.:96.69, PWM:40
15:24:20.670 -> Water_Temp.:96.69, PWM:40
15:24:21.462 -> Water_Temp.:96.69, PWM:40
15:24:22.258 -> Water_Temp.:96.69, PWM:39
15:24:23.015 -> Water_Temp.:96.69, PWM:40
15:24:23.810 -> Water_Temp.:96.69, PWM:39
15:24:24.573 -> Water_Temp.:96.80, PWM:40
15:24:25.363 -> Water_Temp.:96.80, PWM:40
15:24:26.913 -> Water_Temp.:96.80, PWM:39
15:24:27.707 -> Water_Temp.:96.80, PWM:40
15:24:31.545 -> Water_Temp.:96.80, PWM:39
15:24:32.306 -> Water_Temp.:96.80, PWM:39
15:24:34.655 -> Water_Temp.:96.80, PWM:39
15:24:36.213 -> Water_Temp.:96.80, PWM:39
15:24:37.006 -> Water_Temp.:96.80, PWM:34
15:24:37.770 -> Water_Temp.:96.80, PWM:35
15:24:38.562 -> Water_Temp.:96.80, PWM:35
15:24:42.387 -> Water_Temp.:96.80, PWM:34
15:24:43.937 -> Water_Temp.:96.80, PWM:34
15:24:44.727 -> Water_Temp.:96.80, PWM:34
15:24:47.070 -> Water_Temp.:96.80, PWM:34
15:24:47.826 -> Water_Temp.:96.80, PWM:34
15:24:48.622 -> Water_Temp.:98.49, PWM:35
15:24:49.384 -> Water_Temp.:100.62, PWM:35
15:24:52.483 -> Water_Temp.:97.47, PWM:35
15:24:53.274 -> Water_Temp.:97.36, PWM:35
15:24:54.033 -> Water_Temp.:97.36, PWM:34
15:24:54.825 -> Water_Temp.:97.25, PWM:35
15:24:55.618 -> Water_Temp.:97.14, PWM:34
15:24:56.375 -> Water_Temp.:97.14, PWM:35
15:25:01.034 -> Water_Temp.:97.02, PWM:34
15:25:01.829 -> Water_Temp.:97.02, PWM:35
15:25:02.590 -> Water_Temp.:97.02, PWM:34
15:25:03.385 -> Water_Temp.:97.02, PWM:34
15:25:04.178 -> Water_Temp.:97.02, PWM:34
15:25:04.934 -> Water_Temp.:97.02, PWM:34
15:25:05.731 -> Water_Temp.:97.02, PWM:34
15:25:07.254 -> Water_Temp.:97.02, PWM:34
15:25:08.047 -> Water_Temp.:96.91, PWM:35
15:25:08.805 -> Water_Temp.:96.91, PWM:35
15:25:10.364 -> Water_Temp.:96.91, PWM:34
15:25:11.920 -> Water_Temp.:96.91, PWM:35
15:25:13.462 -> Water_Temp.:96.91, PWM:34
15:25:14.220 -> Water_Temp.:96.91, PWM:35
15:25:20.360 -> Water_Temp.:96.91, PWM:34
15:25:21.156 -> Water_Temp.:96.91, PWM:35
15:25:21.945 -> Water_Temp.:96.91, PWM:34
15:25:22.702 -> Water_Temp.:98.60, PWM:34
15:25:24.261 -> Water_Temp.:98.26, PWM:34
15:25:25.022 -> Water_Temp.:97.81, PWM:35
15:25:25.819 -> Water_Temp.:97.59, PWM:35
15:25:26.577 -> Water_Temp.:97.47, PWM:34
15:25:28.121 -> Water_Temp.:97.25, PWM:35
15:25:28.911 -> Water_Temp.:97.25, PWM:35
15:25:30.468 -> Water_Temp.:97.14, PWM:34
15:25:31.228 -> Water_Temp.:97.14, PWM:34
15:25:32.022 -> Water_Temp.:97.14, PWM:35
15:25:32.816 -> Water_Temp.:97.02, PWM:34
15:25:33.576 -> Water_Temp.:97.02, PWM:34
15:25:34.371 -> Water_Temp.:97.02, PWM:34
15:25:35.130 -> Water_Temp.:97.02, PWM:35
15:25:35.927 -> Water_Temp.:97.02, PWM:34
15:25:37.469 -> Water_Temp.:97.02, PWM:35
15:25:38.262 -> Water_Temp.:97.02, PWM:35
15:25:40.597 -> Water_Temp.:97.02, PWM:34
15:25:42.114 -> Water_Temp.:97.02, PWM:34
15:25:42.907 -> Water_Temp.:97.02, PWM:34
15:25:44.463 -> Water_Temp.:97.02, PWM:34
15:25:45.257 -> Water_Temp.:97.02, PWM:34
15:25:46.779 -> Water_Temp.:97.02, PWM:35
15:25:47.576 -> Water_Temp.:97.02, PWM:34
15:25:48.337 -> Water_Temp.:97.02, PWM:35
15:25:49.132 -> Water_Temp.:97.02, PWM:35
15:25:49.892 -> Water_Temp.:97.02, PWM:34
15:25:50.683 -> Water_Temp.:97.02, PWM:35
15:25:51.480 -> Water_Temp.:97.02, PWM:34
15:25:53.002 -> Water_Temp.:97.02, PWM:34
15:25:53.794 -> Water_Temp.:97.02, PWM:35
15:25:54.558 -> Water_Temp.:97.02, PWM:34
15:25:55.352 -> Water_Temp.:97.02, PWM:35
15:25:56.113 -> Water_Temp.:97.02, PWM:35
15:25:56.903 -> Water_Temp.:97.02, PWM:35
15:25:57.694 -> Water_Temp.:97.02, PWM:35
15:25:58.456 -> Water_Temp.:97.02, PWM:35
15:26:00.013 -> Water_Temp.:97.02, PWM:34
15:26:00.772 -> Water_Temp.:97.02, PWM:35
15:26:01.566 -> Water_Temp.:97.02, PWM:34
15:26:02.327 -> Water_Temp.:97.02, PWM:35
15:26:03.123 -> Water_Temp.:97.02, PWM:35
15:26:05.439 -> Water_Temp.:97.02, PWM:34
15:26:06.200 -> Water_Temp.:97.02, PWM:35
15:26:06.995 -> Water_Temp.:97.02, PWM:34
15:26:07.790 -> Water_Temp.:97.02, PWM:34
15:26:08.551 -> Water_Temp.:97.02, PWM:34
15:26:09.341 -> Water_Temp.:97.02, PWM:35
15:26:10.890 -> Water_Temp.:97.02, PWM:34
15:26:13.205 -> Water_Temp.:97.02, PWM:35
15:26:14.753 -> Water_Temp.:97.02, PWM:34
15:26:16.307 -> Water_Temp.:97.02, PWM:35
15:26:18.652 -> Water_Temp.:97.02, PWM:34
15:26:19.416 -> Water_Temp.:97.02, PWM:35
15:26:20.206 -> Water_Temp.:97.02, PWM:34
15:26:20.965 -> Water_Temp.:97.02, PWM:34
15:26:21.757 -> Water_Temp.:96.91, PWM:34
15:26:22.549 -> Water_Temp.:97.02, PWM:34
15:26:23.309 -> Water_Temp.:97.02, PWM:34
15:26:24.869 -> Water_Temp.:96.91, PWM:34
15:26:25.629 -> Water_Temp.:96.91, PWM:35
15:26:27.213 -> Water_Temp.:96.91, PWM:34
15:26:27.970 -> Water_Temp.:96.91, PWM:34
15:26:28.765 -> Water_Temp.:96.91, PWM:35
15:26:29.524 -> Water_Temp.:96.91, PWM:34
15:26:30.321 -> Water_Temp.:96.91, PWM:34
15:26:31.080 -> Water_Temp.:96.91, PWM:35
15:26:34.153 -> Water_Temp.:96.91, PWM:34
15:26:34.946 -> Water_Temp.:96.91, PWM:34
15:26:35.707 -> Water_Temp.:96.91, PWM:35
15:26:36.498 -> Water_Temp.:96.91, PWM:34
15:26:37.259 -> Water_Temp.:96.91, PWM:34
15:26:38.814 -> Water_Temp.:96.91, PWM:34
15:26:40.369 -> Water_Temp.:96.91, PWM:34
15:26:41.128 -> Water_Temp.:96.91, PWM:34
15:26:41.921 -> Water_Temp.:96.91, PWM:35
15:26:42.681 -> Water_Temp.:96.91, PWM:34
15:26:43.475 -> Water_Temp.:96.91, PWM:34
15:26:45.022 -> Water_Temp.:96.91, PWM:34
15:26:47.331 -> Water_Temp.:96.91, PWM:34
15:26:48.125 -> Water_Temp.:96.91, PWM:34
15:26:48.920 -> Water_Temp.:96.91, PWM:34
15:26:50.476 -> Water_Temp.:96.91, PWM:34
15:26:53.560 -> Water_Temp.:96.91, PWM:34
15:26:54.354 -> Water_Temp.:96.91, PWM:34
15:26:55.115 -> Water_Temp.:96.91, PWM:35
15:26:57.394 -> Water_Temp.:97.02, PWM:35
15:26:58.192 -> Water_Temp.:97.02, PWM:34
15:26:58.986 -> Water_Temp.:97.02, PWM:34
15:26:59.744 -> Water_Temp.:98.71, PWM:34
15:27:00.533 -> Water_Temp.:100.74, PWM:35
15:27:01.325 -> Water_Temp.:98.37, PWM:34
15:27:02.842 -> Water_Temp.:97.70, PWM:35
15:27:03.637 -> Water_Temp.:97.59, PWM:34
15:27:04.394 -> Water_Temp.:97.47, PWM:35
15:27:05.184 -> Water_Temp.:97.47, PWM:35
15:27:06.704 -> Water_Temp.:97.36, PWM:35
15:27:07.497 -> Water_Temp.:97.25, PWM:35
15:27:08.293 -> Water_Temp.:97.25, PWM:34
15:27:09.055 -> Water_Temp.:97.25, PWM:35
15:27:09.850 -> Water_Temp.:97.25, PWM:34
15:27:11.397 -> Water_Temp.:97.14, PWM:34
15:27:12.160 -> Water_Temp.:97.14, PWM:35
15:27:12.952 -> Water_Temp.:98.82, PWM:35
15:27:13.744 -> Water_Temp.:100.85, PWM:34
15:27:14.505 -> Water_Temp.:98.37, PWM:34
15:27:16.056 -> Water_Temp.:97.81, PWM:35
15:27:16.819 -> Water_Temp.:97.59, PWM:35
15:27:17.613 -> Water_Temp.:97.59, PWM:34
15:27:18.375 -> Water_Temp.:97.47, PWM:35
15:27:19.928 -> Water_Temp.:97.36, PWM:34
15:27:20.724 -> Water_Temp.:97.25, PWM:34
15:27:21.485 -> Water_Temp.:97.25, PWM:35
15:27:22.278 -> Water_Temp.:97.25, PWM:34
15:27:23.070 -> Water_Temp.:97.25, PWM:34
15:27:26.146 -> Water_Temp.:98.71, PWM:35
15:27:26.941 -> Water_Temp.:100.85, PWM:35
15:27:27.699 -> Water_Temp.:98.37, PWM:35
15:27:28.496 -> Water_Temp.:97.92, PWM:35
15:27:30.057 -> Water_Temp.:97.59, PWM:35
15:27:30.853 -> Water_Temp.:97.47, PWM:35
15:27:32.373 -> Water_Temp.:97.25, PWM:34
15:27:33.925 -> Water_Temp.:97.25, PWM:34
15:27:34.685 -> Water_Temp.:97.14, PWM:34
15:27:35.482 -> Water_Temp.:97.14, PWM:34
15:27:36.245 -> Water_Temp.:97.14, PWM:34
15:27:37.798 -> Water_Temp.:97.14, PWM:35
15:27:38.559 -> Water_Temp.:97.14, PWM:34
15:27:40.109 -> Water_Temp.:97.02, PWM:34
15:27:40.906 -> Water_Temp.:97.02, PWM:34
15:27:41.702 -> Water_Temp.:97.02, PWM:34
15:27:45.534 -> Water_Temp.:97.02, PWM:34
15:27:46.295 -> Water_Temp.:97.02, PWM:34
15:27:47.088 -> Water_Temp.:97.02, PWM:35
15:27:47.884 -> Water_Temp.:97.02, PWM:34
15:27:49.404 -> Water_Temp.:97.02, PWM:35
15:27:50.196 -> Water_Temp.:97.02, PWM:34
15:27:51.752 -> Water_Temp.:97.02, PWM:35
15:27:52.545 -> Water_Temp.:97.02, PWM:35
15:27:54.059 -> Water_Temp.:97.02, PWM:35
15:27:55.614 -> Water_Temp.:97.02, PWM:35
15:27:57.166 -> Water_Temp.:97.02, PWM:35
15:27:59.482 -> Water_Temp.:97.14, PWM:35
15:28:01.067 -> Water_Temp.:97.14, PWM:34
15:28:01.829 -> Water_Temp.:97.14, PWM:35
15:28:02.621 -> Water_Temp.:97.14, PWM:35
15:28:03.377 -> Water_Temp.:97.14, PWM:34
15:28:04.170 -> Water_Temp.:97.14, PWM:34
15:28:05.723 -> Water_Temp.:97.14, PWM:34
15:28:06.519 -> Water_Temp.:97.14, PWM:35
15:28:07.279 -> Water_Temp.:97.14, PWM:34
15:28:08.075 -> Water_Temp.:97.14, PWM:34
15:28:08.831 -> Water_Temp.:97.14, PWM:34
15:28:09.626 -> Water_Temp.:97.14, PWM:34
15:28:10.422 -> Water_Temp.:97.14, PWM:34
15:28:11.185 -> Water_Temp.:97.14, PWM:35
15:28:13.493 -> Water_Temp.:97.14, PWM:34
15:28:14.285 -> Water_Temp.:97.14, PWM:34
15:28:15.835 -> Water_Temp.:97.14, PWM:35
15:28:16.629 -> Water_Temp.:97.14, PWM:34
15:28:18.182 -> Water_Temp.:97.14, PWM:34
15:28:18.942 -> Water_Temp.:97.14, PWM:34
15:28:19.736 -> Water_Temp.:97.14, PWM:34
15:28:21.262 -> Water_Temp.:97.14, PWM:34
15:28:22.814 -> Water_Temp.:97.02, PWM:35
15:28:23.608 -> Water_Temp.:97.02, PWM:35
15:28:24.367 -> Water_Temp.:97.02, PWM:34
15:28:25.161 -> Water_Temp.:97.02, PWM:34
15:28:25.925 -> Water_Temp.:97.02, PWM:34
15:28:26.721 -> Water_Temp.:97.02, PWM:35
15:28:27.484 -> Water_Temp.:97.02, PWM:34
15:28:28.280 -> Water_Temp.:97.02, PWM:34
15:28:29.040 -> Water_Temp.:97.02, PWM:34
15:28:29.833 -> Water_Temp.:97.02, PWM:35
15:28:30.595 -> Water_Temp.:97.02, PWM:34
15:28:31.391 -> Water_Temp.:97.02, PWM:34
15:28:32.910 -> Water_Temp.:97.02, PWM:34
15:28:33.702 -> Water_Temp.:97.02, PWM:34
15:28:36.775 -> Water_Temp.:97.02, PWM:34
15:28:37.566 -> Water_Temp.:97.02, PWM:34
15:28:38.323 -> Water_Temp.:97.02, PWM:35
15:28:39.114 -> Water_Temp.:97.02, PWM:35
15:28:39.911 -> Water_Temp.:97.02, PWM:35
15:28:40.672 -> Water_Temp.:97.02, PWM:35
15:28:41.464 -> Water_Temp.:97.02, PWM:34
15:28:45.290 -> Water_Temp.:97.02, PWM:34
15:28:48.358 -> Water_Temp.:97.02, PWM:35
15:28:49.146 -> Water_Temp.:97.02, PWM:35
15:28:49.901 -> Water_Temp.:97.02, PWM:34
15:28:50.697 -> Water_Temp.:97.02, PWM:35
15:28:52.245 -> Water_Temp.:97.02, PWM:34
15:28:53.004 -> Water_Temp.:97.02, PWM:35
15:28:53.791 -> Water_Temp.:97.02, PWM:34
15:28:55.338 -> Water_Temp.:97.02, PWM:34
15:28:56.855 -> Water_Temp.:97.14, PWM:35
15:28:57.648 -> Water_Temp.:97.14, PWM:34
15:28:58.411 -> Water_Temp.:97.14, PWM:35
15:28:59.202 -> Water_Temp.:97.14, PWM:34
15:29:00.723 -> Water_Temp.:97.14, PWM:34
15:29:01.513 -> Water_Temp.:97.14, PWM:35
15:29:03.065 -> Water_Temp.:97.14, PWM:34
15:29:03.861 -> Water_Temp.:97.14, PWM:34
15:29:05.381 -> Water_Temp.:97.14, PWM:34
15:29:06.177 -> Water_Temp.:97.14, PWM:35
15:29:06.936 -> Water_Temp.:97.14, PWM:34
15:29:07.732 -> Water_Temp.:97.14, PWM:35
15:29:08.526 -> Water_Temp.:97.14, PWM:34
15:29:10.810 -> Water_Temp.:97.14, PWM:35
15:29:11.608 -> Water_Temp.:97.14, PWM:34
15:29:13.155 -> Water_Temp.:97.14, PWM:35
15:29:13.914 -> Water_Temp.:97.14, PWM:34
15:29:14.710 -> Water_Temp.:97.14, PWM:35
15:29:15.472 -> Water_Temp.:97.14, PWM:35
15:29:16.271 -> Water_Temp.:97.14, PWM:35
15:29:18.579 -> Water_Temp.:98.82, PWM:35
15:29:21.646 -> Water_Temp.:97.92, PWM:34
15:29:22.399 -> Water_Temp.:97.81, PWM:35
15:29:23.191 -> Water_Temp.:97.70, PWM:34
15:29:24.741 -> Water_Temp.:97.59, PWM:34
15:29:27.080 -> Water_Temp.:97.47, PWM:35
15:29:27.870 -> Water_Temp.:97.36, PWM:35
15:29:28.629 -> Water_Temp.:97.36, PWM:35
15:29:30.174 -> Water_Temp.:97.36, PWM:34
15:29:30.972 -> Water_Temp.:97.36, PWM:34
15:29:32.486 -> Water_Temp.:97.36, PWM:34
15:29:33.280 -> Water_Temp.:97.25, PWM:34
15:29:34.073 -> Water_Temp.:97.25, PWM:34
15:29:34.835 -> Water_Temp.:97.25, PWM:34
15:29:35.632 -> Water_Temp.:97.25, PWM:34
15:29:36.392 -> Water_Temp.:97.25, PWM:34
15:29:37.186 -> Water_Temp.:97.25, PWM:34
15:29:37.943 -> Water_Temp.:97.25, PWM:35
15:29:40.254 -> Water_Temp.:97.25, PWM:34
15:29:41.045 -> Water_Temp.:97.25, PWM:34
15:29:42.593 -> Water_Temp.:97.25, PWM:34
15:29:43.384 -> Water_Temp.:97.25, PWM:35
15:29:44.141 -> Water_Temp.:97.25, PWM:34
15:29:47.247 -> Water_Temp.:97.14, PWM:34
15:29:51.834 -> Water_Temp.:97.25, PWM:34
15:29:52.597 -> Water_Temp.:97.25, PWM:34
15:29:53.391 -> Water_Temp.:97.25, PWM:35
15:29:54.946 -> Water_Temp.:97.14, PWM:34
15:29:55.742 -> Water_Temp.:97.14, PWM:34
15:29:57.259 -> Water_Temp.:97.14, PWM:34
15:29:58.052 -> Water_Temp.:97.14, PWM:34
15:30:00.356 -> Water_Temp.:97.14, PWM:35
15:30:01.907 -> Water_Temp.:97.14, PWM:35
15:30:02.698 -> Water_Temp.:97.14, PWM:34
15:30:04.245 -> Water_Temp.:97.14, PWM:34
15:30:05.006 -> Water_Temp.:97.14, PWM:34
15:30:06.563 -> Water_Temp.:97.14, PWM:34
15:30:10.426 -> Water_Temp.:97.14, PWM:34
15:30:11.218 -> Water_Temp.:97.14, PWM:35
15:30:12.009 -> Water_Temp.:97.14, PWM:34
15:30:14.293 -> Water_Temp.:97.25, PWM:34
15:30:15.085 -> Water_Temp.:97.25, PWM:34
15:30:16.638 -> Water_Temp.:97.25, PWM:35
15:30:17.435 -> Water_Temp.:97.25, PWM:34
15:30:18.195 -> Water_Temp.:97.25, PWM:35
15:30:18.992 -> Water_Temp.:97.25, PWM:34
15:30:20.546 -> Water_Temp.:97.25, PWM:35
15:30:21.303 -> Water_Temp.:97.25, PWM:35
15:30:22.857 -> Water_Temp.:97.25, PWM:34
15:30:23.649 -> Water_Temp.:97.14, PWM:34
15:30:24.440 -> Water_Temp.:97.14, PWM:35
15:30:25.200 -> Water_Temp.:97.14, PWM:34
15:30:26.723 -> Water_Temp.:97.14, PWM:35
15:30:27.514 -> Water_Temp.:97.14, PWM:34
15:30:28.308 -> Water_Temp.:97.14, PWM:35
15:30:29.065 -> Water_Temp.:97.14, PWM:35
15:30:30.621 -> Water_Temp.:97.14, PWM:35
15:30:32.170 -> Water_Temp.:97.14, PWM:35
15:30:32.959 -> Water_Temp.:97.14, PWM:35
15:30:33.719 -> Water_Temp.:97.14, PWM:34
15:30:34.510 -> Water_Temp.:97.14, PWM:34
15:30:35.297 -> Water_Temp.:97.14, PWM:35
15:30:36.092 -> Water_Temp.:97.14, PWM:34
15:30:37.648 -> Water_Temp.:97.14, PWM:34
15:30:38.409 -> Water_Temp.:97.14, PWM:34
15:30:39.965 -> Water_Temp.:97.14, PWM:34
15:30:40.755 -> Water_Temp.:97.14, PWM:34
15:30:42.279 -> Water_Temp.:97.14, PWM:35
15:30:43.828 -> Water_Temp.:97.14, PWM:35
15:30:44.621 -> Water_Temp.:97.14, PWM:34
15:30:46.967 -> Water_Temp.:97.14, PWM:35
15:30:47.728 -> Water_Temp.:97.14, PWM:34
15:30:48.525 -> Water_Temp.:97.14, PWM:35
15:30:49.283 -> Water_Temp.:97.14, PWM:34
15:30:50.080 -> Water_Temp.:97.25, PWM:34
15:30:50.842 -> Water_Temp.:97.14, PWM:35
15:30:51.638 -> Water_Temp.:97.14, PWM:34
15:30:53.188 -> Water_Temp.:97.14, PWM:34
15:30:54.746 -> Water_Temp.:97.25, PWM:35
15:30:56.300 -> Water_Temp.:97.25, PWM:34
15:30:57.059 -> Water_Temp.:97.25, PWM:34
15:30:58.608 -> Water_Temp.:97.25, PWM:35
15:30:59.401 -> Water_Temp.:97.25, PWM:35
15:31:00.195 -> Water_Temp.:97.25, PWM:35
15:31:01.709 -> Water_Temp.:97.25, PWM:34
15:31:03.294 -> Water_Temp.:97.25, PWM:34
15:31:04.056 -> Water_Temp.:97.25, PWM:35
15:31:04.849 -> Water_Temp.:97.25, PWM:35
15:31:05.611 -> Water_Temp.:97.25, PWM:35
15:31:07.956 -> Water_Temp.:97.25, PWM:34
15:31:09.477 -> Water_Temp.:97.25, PWM:34
15:31:11.028 -> Water_Temp.:97.25, PWM:35
15:31:11.825 -> Water_Temp.:97.25, PWM:35
15:31:12.587 -> Water_Temp.:97.25, PWM:35
15:31:13.377 -> Water_Temp.:97.25, PWM:35
15:31:14.898 -> Water_Temp.:97.25, PWM:34
15:31:16.456 -> Water_Temp.:97.25, PWM:34
15:31:17.253 -> Water_Temp.:97.25, PWM:34
15:31:18.013 -> Water_Temp.:97.25, PWM:35
15:31:21.093 -> Water_Temp.:97.25, PWM:34
15:31:22.634 -> Water_Temp.:97.25, PWM:34
15:31:24.218 -> Water_Temp.:97.25, PWM:34
15:31:26.527 -> Water_Temp.:97.25, PWM:35
15:31:28.842 -> Water_Temp.:97.25, PWM:35
15:31:29.604 -> Water_Temp.:97.25, PWM:34
15:31:31.923 -> Water_Temp.:97.25, PWM:35
15:31:32.720 -> Water_Temp.:97.25, PWM:34
15:31:33.516 -> Water_Temp.:97.25, PWM:34
15:31:34.276 -> Water_Temp.:98.94, PWM:34
15:31:36.625 -> Water_Temp.:98.15, PWM:35
15:31:38.176 -> Water_Temp.:97.81, PWM:34
15:31:42.009 -> Water_Temp.:97.47, PWM:34
15:31:43.566 -> Water_Temp.:97.47, PWM:34
15:31:44.326 -> Water_Temp.:97.47, PWM:34
15:31:45.917 -> Water_Temp.:97.36, PWM:34
15:31:46.680 -> Water_Temp.:97.36, PWM:34
15:31:48.205 -> Water_Temp.:97.36, PWM:35
15:31:48.999 -> Water_Temp.:97.36, PWM:34
15:31:49.762 -> Water_Temp.:97.36, PWM:34
15:31:50.555 -> Water_Temp.:97.36, PWM:34
15:31:51.344 -> Water_Temp.:97.36, PWM:34
15:31:55.979 -> Water_Temp.:97.36, PWM:35
15:31:59.054 -> Water_Temp.:97.36, PWM:35
15:32:00.577 -> Water_Temp.:97.25, PWM:34
15:32:02.884 -> Water_Temp.:97.36, PWM:34
15:32:06.718 -> Water_Temp.:97.36, PWM:34
15:32:09.820 -> Water_Temp.:97.36, PWM:34
15:32:11.371 -> Water_Temp.:97.36, PWM:34
15:32:12.917 -> Water_Temp.:97.36, PWM:34
15:32:14.472 -> Water_Temp.:97.36, PWM:34
15:32:16.023 -> Water_Temp.:97.36, PWM:34
15:32:16.780 -> Water_Temp.:97.36, PWM:35
15:32:17.578 -> Water_Temp.:97.36, PWM:34
15:32:18.371 -> Water_Temp.:97.36, PWM:34
15:32:19.888 -> Water_Temp.:97.36, PWM:34
15:32:20.684 -> Water_Temp.:97.36, PWM:35
15:32:21.479 -> Water_Temp.:97.36, PWM:34
15:32:22.992 -> Water_Temp.:97.36, PWM:35
15:32:25.339 -> Water_Temp.:97.36, PWM:35
15:32:27.661 -> Water_Temp.:97.36, PWM:35
15:32:28.457 -> Water_Temp.:97.36, PWM:34
15:32:30.770 -> Water_Temp.:97.36, PWM:34
15:32:33.120 -> Water_Temp.:97.36, PWM:34
15:32:33.880 -> Water_Temp.:97.36, PWM:34
15:32:34.678 -> Water_Temp.:97.36, PWM:34
15:32:36.226 -> Water_Temp.:97.36, PWM:34
15:32:37.018 -> Water_Temp.:97.36, PWM:34
15:32:37.780 -> Water_Temp.:97.36, PWM:35
15:32:38.577 -> Water_Temp.:97.36, PWM:35
15:32:40.135 -> Water_Temp.:97.36, PWM:34
15:32:41.691 -> Water_Temp.:97.36, PWM:34
15:32:42.448 -> Water_Temp.:97.36, PWM:35
15:32:44.793 -> Water_Temp.:97.36, PWM:35
15:32:45.556 -> Water_Temp.:97.47, PWM:34
15:32:47.908 -> Water_Temp.:97.36, PWM:35
15:32:49.463 -> Water_Temp.:97.47, PWM:35
15:32:50.257 -> Water_Temp.:97.47, PWM:35
15:32:52.568 -> Water_Temp.:97.47, PWM:34
15:32:54.853 -> Water_Temp.:97.47, PWM:35
15:32:55.647 -> Water_Temp.:97.47, PWM:35
15:32:56.441 -> Water_Temp.:97.47, PWM:34
15:32:57.959 -> Water_Temp.:97.47, PWM:35
15:32:58.753 -> Water_Temp.:97.47, PWM:35
15:32:59.517 -> Water_Temp.:97.47, PWM:35
15:33:00.307 -> Water_Temp.:97.47, PWM:34
15:33:03.415 -> Water_Temp.:97.47, PWM:34
15:33:04.176 -> Water_Temp.:97.47, PWM:35
15:33:04.971 -> Water_Temp.:97.47, PWM:34
15:33:06.527 -> Water_Temp.:97.47, PWM:34
15:33:07.319 -> Water_Temp.:97.47, PWM:34
15:33:08.081 -> Water_Temp.:97.47, PWM:34
15:33:08.870 -> Water_Temp.:97.47, PWM:35
15:33:11.180 -> Water_Temp.:97.36, PWM:34
15:33:11.976 -> Water_Temp.:97.36, PWM:34
15:33:14.284 -> Water_Temp.:97.36, PWM:34
15:33:15.044 -> Water_Temp.:97.36, PWM:34
15:33:15.841 -> Water_Temp.:97.36, PWM:35
15:33:18.149 -> Water_Temp.:97.36, PWM:35
15:33:18.906 -> Water_Temp.:97.36, PWM:34
15:33:19.703 -> Water_Temp.:97.36, PWM:34
15:33:20.463 -> Water_Temp.:97.36, PWM:35
15:33:22.804 -> Water_Temp.:97.36, PWM:35
15:33:23.566 -> Water_Temp.:97.36, PWM:34
15:33:25.906 -> Water_Temp.:97.25, PWM:34
15:33:26.698 -> Water_Temp.:97.25, PWM:35
15:33:29.008 -> Water_Temp.:97.25, PWM:34
15:33:29.804 -> Water_Temp.:97.25, PWM:34
15:33:30.565 -> Water_Temp.:97.25, PWM:35
15:33:31.360 -> Water_Temp.:97.25, PWM:34
15:33:33.644 -> Water_Temp.:97.25, PWM:34
15:33:37.522 -> Water_Temp.:97.25, PWM:35
15:33:38.285 -> Water_Temp.:97.36, PWM:34
15:33:39.080 -> Water_Temp.:97.36, PWM:34
15:33:39.840 -> Water_Temp.:97.36, PWM:34
15:33:40.636 -> Water_Temp.:97.25, PWM:34
15:33:42.184 -> Water_Temp.:97.25, PWM:35
15:33:42.977 -> Water_Temp.:97.25, PWM:34
15:33:43.736 -> Water_Temp.:97.25, PWM:34
15:33:44.534 -> Water_Temp.:97.25, PWM:34
15:33:45.293 -> Water_Temp.:97.36, PWM:34
15:33:46.849 -> Water_Temp.:97.36, PWM:34
15:33:47.644 -> Water_Temp.:97.36, PWM:34
15:33:48.439 -> Water_Temp.:97.25, PWM:35
15:33:49.959 -> Water_Temp.:97.36, PWM:34
15:33:50.715 -> Water_Temp.:97.36, PWM:35
15:33:51.511 -> Water_Temp.:97.36, PWM:35
15:33:52.309 -> Water_Temp.:97.25, PWM:34
15:33:53.068 -> Water_Temp.:97.36, PWM:34
15:33:53.860 -> Water_Temp.:97.25, PWM:34
15:33:55.385 -> Water_Temp.:97.36, PWM:34
15:33:56.184 -> Water_Temp.:97.25, PWM:34
15:33:57.735 -> Water_Temp.:97.25, PWM:35
15:33:59.286 -> Water_Temp.:97.36, PWM:34
15:34:00.081 -> Water_Temp.:97.36, PWM:34
15:34:02.360 -> Water_Temp.:97.25, PWM:35
15:34:03.915 -> Water_Temp.:97.25, PWM:34
15:34:04.709 -> Water_Temp.:97.25, PWM:35
15:34:06.261 -> Water_Temp.:97.25, PWM:34
15:34:07.779 -> Water_Temp.:97.25, PWM:34
15:34:12.408 -> Water_Temp.:97.25, PWM:34
15:34:13.199 -> Water_Temp.:97.25, PWM:34
15:34:13.958 -> Water_Temp.:97.25, PWM:34
15:34:16.272 -> Water_Temp.:97.25, PWM:35
15:34:17.029 -> Water_Temp.:97.25, PWM:35
15:34:17.824 -> Water_Temp.:97.25, PWM:34
15:34:18.614 -> Water_Temp.:97.25, PWM:35
15:34:19.373 -> Water_Temp.:97.25, PWM:35
15:34:20.164 -> Water_Temp.:97.25, PWM:34
15:34:21.710 -> Water_Temp.:97.25, PWM:34
15:34:22.471 -> Water_Temp.:97.25, PWM:34
15:34:23.262 -> Water_Temp.:97.25, PWM:34
15:34:24.056 -> Water_Temp.:97.25, PWM:34
15:34:24.812 -> Water_Temp.:97.25, PWM:35
15:34:25.611 -> Water_Temp.:97.36, PWM:34
15:34:26.402 -> Water_Temp.:97.36, PWM:35
15:34:27.162 -> Water_Temp.:97.25, PWM:35
15:34:28.714 -> Water_Temp.:97.36, PWM:34
15:34:30.266 -> Water_Temp.:97.36, PWM:34
15:34:34.102 -> Water_Temp.:97.25, PWM:34
15:34:35.658 -> Water_Temp.:97.25, PWM:39
15:34:37.213 -> Water_Temp.:97.25, PWM:39
15:34:38.008 -> Water_Temp.:97.25, PWM:37
15:34:39.565 -> Water_Temp.:97.25, PWM:36
15:34:41.084 -> Water_Temp.:97.25, PWM:36
15:34:41.878 -> Water_Temp.:97.25, PWM:36
15:34:42.633 -> Water_Temp.:97.25, PWM:36
15:34:43.427 -> Water_Temp.:97.25, PWM:36
15:34:44.220 -> Water_Temp.:97.25, PWM:36
15:34:44.979 -> Water_Temp.:97.25, PWM:36
15:34:45.774 -> Water_Temp.:97.25, PWM:36
15:34:47.328 -> Water_Temp.:97.25, PWM:36
15:34:48.087 -> Water_Temp.:97.14, PWM:36
15:34:48.883 -> Water_Temp.:97.25, PWM:36
15:34:53.482 -> Water_Temp.:97.14, PWM:36
15:34:54.278 -> Water_Temp.:97.14, PWM:36
15:34:55.042 -> Water_Temp.:97.14, PWM:36
15:34:55.837 -> Water_Temp.:97.14, PWM:36
15:34:57.387 -> Water_Temp.:97.14, PWM:36
15:34:58.145 -> Water_Temp.:97.14, PWM:36
15:34:59.705 -> Water_Temp.:97.14, PWM:36
15:35:00.500 -> Water_Temp.:97.14, PWM:36
15:35:01.259 -> Water_Temp.:97.14, PWM:36
15:35:02.807 -> Water_Temp.:97.14, PWM:36
15:35:03.599 -> Water_Temp.:97.14, PWM:36
15:35:04.359 -> Water_Temp.:97.14, PWM:36
15:35:05.159 -> Water_Temp.:97.14, PWM:36
15:35:05.923 -> Water_Temp.:97.14, PWM:36
15:35:08.206 -> Water_Temp.:97.14, PWM:36
15:35:09.756 -> Water_Temp.:97.14, PWM:36
15:35:10.551 -> Water_Temp.:97.14, PWM:36
15:35:12.895 -> Water_Temp.:97.14, PWM:36
15:35:15.209 -> Water_Temp.:97.14, PWM:36
15:35:15.999 -> Water_Temp.:97.25, PWM:36
15:35:16.759 -> Water_Temp.:97.14, PWM:36
15:35:17.554 -> Water_Temp.:97.25, PWM:36
15:35:19.865 -> Water_Temp.:97.25, PWM:36
15:35:20.658 -> Water_Temp.:97.14, PWM:36
15:35:21.448 -> Water_Temp.:97.25, PWM:36
15:35:22.998 -> Water_Temp.:97.14, PWM:36
15:35:23.761 -> Water_Temp.:97.14, PWM:36
15:35:25.313 -> Water_Temp.:97.14, PWM:36
15:35:26.109 -> Water_Temp.:97.14, PWM:36
15:35:27.663 -> Water_Temp.:97.14, PWM:36
15:35:28.422 -> Water_Temp.:97.14, PWM:36
15:35:29.218 -> Water_Temp.:97.14, PWM:36
15:35:29.981 -> Water_Temp.:97.14, PWM:36
15:35:30.774 -> Water_Temp.:97.14, PWM:36
15:35:32.327 -> Water_Temp.:97.14, PWM:36
15:35:33.119 -> Water_Temp.:97.14, PWM:36
15:35:33.880 -> Water_Temp.:97.25, PWM:36
15:35:35.433 -> Water_Temp.:97.14, PWM:36
15:35:36.193 -> Water_Temp.:97.25, PWM:36
15:35:36.986 -> Water_Temp.:97.25, PWM:36
15:35:37.780 -> Water_Temp.:98.82, PWM:36
15:35:40.892 -> Water_Temp.:97.92, PWM:36
15:35:41.650 -> Water_Temp.:97.70, PWM:36
15:35:42.442 -> Water_Temp.:97.70, PWM:36
15:35:43.203 -> Water_Temp.:97.59, PWM:36
15:35:44.749 -> Water_Temp.:97.47, PWM:36
15:35:47.091 -> Water_Temp.:97.36, PWM:36
15:35:47.848 -> Water_Temp.:97.36, PWM:36
15:35:48.641 -> Water_Temp.:97.36, PWM:36
15:35:49.437 -> Water_Temp.:97.36, PWM:36
15:35:50.955 -> Water_Temp.:97.36, PWM:36
15:35:52.515 -> Water_Temp.:97.36, PWM:36
15:35:54.070 -> Water_Temp.:97.25, PWM:36
15:35:54.863 -> Water_Temp.:97.25, PWM:36
15:35:55.625 -> Water_Temp.:97.25, PWM:36
15:35:56.420 -> Water_Temp.:97.25, PWM:36
15:35:57.939 -> Water_Temp.:97.25, PWM:36
15:35:58.736 -> Water_Temp.:97.25, PWM:36
15:36:00.259 -> Water_Temp.:97.25, PWM:36
15:36:01.052 -> Water_Temp.:97.25, PWM:36
15:36:01.811 -> Water_Temp.:97.25, PWM:36
15:36:02.607 -> Water_Temp.:97.25, PWM:36
15:36:03.400 -> Water_Temp.:97.25, PWM:36
15:36:04.161 -> Water_Temp.:97.25, PWM:36
15:36:06.509 -> Water_Temp.:97.25, PWM:36
15:36:07.270 -> Water_Temp.:97.25, PWM:36
15:36:08.829 -> Water_Temp.:97.25, PWM:36
15:36:09.591 -> Water_Temp.:97.25, PWM:36
15:36:10.385 -> Water_Temp.:97.25, PWM:36
15:36:11.177 -> Water_Temp.:97.25, PWM:36
15:36:14.218 -> Water_Temp.:97.25, PWM:36
15:36:16.521 -> Water_Temp.:97.14, PWM:36
15:36:17.313 -> Water_Temp.:97.14, PWM:36
15:36:18.104 -> Water_Temp.:97.14, PWM:36
15:36:18.865 -> Water_Temp.:97.14, PWM:36
15:36:21.912 -> Water_Temp.:97.25, PWM:36
15:36:23.499 -> Water_Temp.:97.25, PWM:36
15:36:24.254 -> Water_Temp.:97.25, PWM:36
15:36:25.046 -> Water_Temp.:97.25, PWM:36
15:36:26.572 -> Water_Temp.:97.25, PWM:36
15:36:28.153 -> Water_Temp.:97.25, PWM:36
15:36:28.914 -> Water_Temp.:97.25, PWM:36
15:36:30.465 -> Water_Temp.:97.25, PWM:36
15:36:31.261 -> Water_Temp.:97.25, PWM:36
15:36:32.054 -> Water_Temp.:97.25, PWM:36
15:36:33.571 -> Water_Temp.:97.25, PWM:36
15:36:35.883 -> Water_Temp.:97.14, PWM:36
15:36:36.676 -> Water_Temp.:97.25, PWM:36
15:36:37.472 -> Water_Temp.:97.25, PWM:36
15:36:38.232 -> Water_Temp.:97.25, PWM:36
15:36:39.784 -> Water_Temp.:97.25, PWM:36
15:36:42.092 -> Water_Temp.:97.25, PWM:36
15:36:42.886 -> Water_Temp.:97.25, PWM:36
15:36:43.681 -> Water_Temp.:97.25, PWM:36
15:36:46.719 -> Water_Temp.:97.25, PWM:36
15:36:48.280 -> Water_Temp.:97.25, PWM:36
15:36:49.073 -> Water_Temp.:97.25, PWM:36
15:36:50.590 -> Water_Temp.:97.25, PWM:36
15:36:51.383 -> Water_Temp.:97.25, PWM:36
15:36:52.932 -> Water_Temp.:97.25, PWM:36
15:36:53.730 -> Water_Temp.:97.25, PWM:36
15:36:55.277 -> Water_Temp.:97.25, PWM:36
15:36:57.598 -> Water_Temp.:98.94, PWM:36
15:36:59.950 -> Water_Temp.:98.26, PWM:36
15:37:00.710 -> Water_Temp.:98.04, PWM:36
15:37:01.506 -> Water_Temp.:97.92, PWM:36
15:37:02.261 -> Water_Temp.:97.81, PWM:36
15:37:03.053 -> Water_Temp.:97.70, PWM:36
15:37:03.848 -> Water_Temp.:97.59, PWM:36
15:37:04.610 -> Water_Temp.:97.59, PWM:36
15:37:07.717 -> Water_Temp.:97.47, PWM:36
15:37:09.270 -> Water_Temp.:97.47, PWM:36
15:37:10.031 -> Water_Temp.:97.47, PWM:36
15:37:11.586 -> Water_Temp.:97.36, PWM:36
15:37:12.377 -> Water_Temp.:97.36, PWM:36
15:37:13.926 -> Water_Temp.:97.36, PWM:36
15:37:14.720 -> Water_Temp.:97.36, PWM:36
15:37:15.478 -> Water_Temp.:97.36, PWM:36
15:37:17.788 -> Water_Temp.:97.36, PWM:36
15:37:20.859 -> Water_Temp.:97.36, PWM:36
15:37:22.418 -> Water_Temp.:97.47, PWM:36
15:37:23.177 -> Water_Temp.:97.47, PWM:36
15:37:23.972 -> Water_Temp.:97.47, PWM:36
15:37:24.733 -> Water_Temp.:97.47, PWM:36
15:37:25.530 -> Water_Temp.:97.47, PWM:36
15:37:26.288 -> Water_Temp.:97.47, PWM:36
15:37:27.876 -> Water_Temp.:97.47, PWM:36
15:37:28.636 -> Water_Temp.:97.47, PWM:36
15:37:29.431 -> Water_Temp.:97.47, PWM:36
15:37:30.989 -> Water_Temp.:97.47, PWM:36
15:37:31.753 -> Water_Temp.:97.47, PWM:36
15:37:32.548 -> Water_Temp.:97.47, PWM:36
15:37:33.311 -> Water_Temp.:97.47, PWM:36
15:37:34.102 -> Water_Temp.:97.47, PWM:36
15:37:34.862 -> Water_Temp.:97.47, PWM:36
15:37:36.419 -> Water_Temp.:97.47, PWM:36
15:37:37.942 -> Water_Temp.:97.59, PWM:36
15:37:38.740 -> Water_Temp.:97.59, PWM:36
15:37:39.498 -> Water_Temp.:97.59, PWM:36
15:37:40.293 -> Water_Temp.:97.59, PWM:36
15:37:41.086 -> Water_Temp.:97.59, PWM:36
15:37:42.641 -> Water_Temp.:97.59, PWM:36
15:37:44.195 -> Water_Temp.:97.59, PWM:36
15:37:44.958 -> Water_Temp.:97.59, PWM:36
15:37:45.752 -> Water_Temp.:97.59, PWM:36
15:37:46.543 -> Water_Temp.:97.59, PWM:36
15:37:48.850 -> Water_Temp.:97.59, PWM:36
15:37:49.638 -> Water_Temp.:97.59, PWM:36
15:37:51.954 -> Water_Temp.:97.59, PWM:36
15:37:53.469 -> Water_Temp.:97.59, PWM:36
15:37:54.264 -> Water_Temp.:97.59, PWM:36
15:37:55.025 -> Water_Temp.:97.59, PWM:36
15:37:55.818 -> Water_Temp.:97.59, PWM:36
15:37:56.608 -> Water_Temp.:97.59, PWM:36
15:37:57.367 -> Water_Temp.:97.59, PWM:36
15:37:58.163 -> Water_Temp.:97.59, PWM:36
15:37:58.956 -> Water_Temp.:97.59, PWM:36
15:37:59.711 -> Water_Temp.:97.47, PWM:36
15:38:00.507 -> Water_Temp.:97.47, PWM:36
15:38:01.296 -> Water_Temp.:97.47, PWM:36
15:38:02.845 -> Water_Temp.:97.47, PWM:36
15:38:03.604 -> Water_Temp.:97.47, PWM:36
15:38:04.392 -> Water_Temp.:97.47, PWM:36
15:38:05.184 -> Water_Temp.:97.47, PWM:36
15:38:05.943 -> Water_Temp.:97.47, PWM:36
15:38:06.737 -> Water_Temp.:97.47, PWM:36
15:38:07.527 -> Water_Temp.:97.47, PWM:36
15:38:09.082 -> Water_Temp.:97.47, PWM:36
15:38:10.637 -> Water_Temp.:97.47, PWM:36
15:38:12.974 -> Water_Temp.:97.47, PWM:36
15:38:14.529 -> Water_Temp.:97.47, PWM:36
15:38:15.285 -> Water_Temp.:97.47, PWM:36
15:38:16.082 -> Water_Temp.:97.47, PWM:36
15:38:16.842 -> Water_Temp.:97.47, PWM:36
15:38:17.637 -> Water_Temp.:97.47, PWM:36
15:38:18.431 -> Water_Temp.:97.47, PWM:36
15:38:20.720 -> Water_Temp.:97.47, PWM:36
15:38:21.518 -> Water_Temp.:97.47, PWM:36
15:38:23.064 -> Water_Temp.:97.59, PWM:36
15:38:23.827 -> Water_Temp.:97.59, PWM:36
15:38:24.618 -> Water_Temp.:97.59, PWM:36
15:38:25.412 -> Water_Temp.:97.59, PWM:36
15:38:27.694 -> Water_Temp.:97.59, PWM:36
15:38:29.213 -> Water_Temp.:97.59, PWM:36
15:38:30.007 -> Water_Temp.:97.59, PWM:36
15:38:31.561 -> Water_Temp.:97.59, PWM:36
15:38:34.659 -> Water_Temp.:97.59, PWM:36
15:38:35.423 -> Water_Temp.:97.59, PWM:36
15:38:36.980 -> Water_Temp.:97.59, PWM:36
15:38:37.774 -> Water_Temp.:97.59, PWM:36
15:38:38.533 -> Water_Temp.:97.59, PWM:36
15:38:39.325 -> Water_Temp.:97.70, PWM:36
15:38:40.086 -> Water_Temp.:97.59, PWM:36
15:38:40.878 -> Water_Temp.:99.27, PWM:36
15:38:41.674 -> Water_Temp.:101.41, PWM:36
15:38:42.430 -> Water_Temp.:98.94, PWM:36
15:38:44.015 -> Water_Temp.:98.37, PWM:36
15:38:44.780 -> Water_Temp.:98.15, PWM:36
15:38:46.336 -> Water_Temp.:98.04, PWM:36
15:38:47.096 -> Water_Temp.:98.04, PWM:36
15:38:47.888 -> Water_Temp.:97.92, PWM:36
15:38:48.680 -> Water_Temp.:97.92, PWM:36
15:38:49.435 -> Water_Temp.:97.92, PWM:36
15:38:51.019 -> Water_Temp.:97.81, PWM:36
15:38:51.773 -> Water_Temp.:97.81, PWM:36
15:38:55.607 -> Water_Temp.:97.81, PWM:36
15:38:56.401 -> Water_Temp.:97.70, PWM:36
15:38:57.191 -> Water_Temp.:97.70, PWM:36
15:38:57.948 -> Water_Temp.:97.70, PWM:36
15:38:58.736 -> Water_Temp.:97.70, PWM:36
15:38:59.528 -> Water_Temp.:97.70, PWM:36
15:39:00.284 -> Water_Temp.:97.70, PWM:36
15:39:01.874 -> Water_Temp.:97.70, PWM:36
15:39:02.630 -> Water_Temp.:97.70, PWM:36
15:39:04.980 -> Water_Temp.:97.70, PWM:36
15:39:05.739 -> Water_Temp.:97.70, PWM:36
15:39:08.083 -> Water_Temp.:97.70, PWM:36
15:39:09.629 -> Water_Temp.:97.70, PWM:36
15:39:10.389 -> Water_Temp.:97.70, PWM:36
15:39:11.935 -> Water_Temp.:97.70, PWM:36
15:39:12.727 -> Water_Temp.:97.81, PWM:36
15:39:15.032 -> Water_Temp.:97.81, PWM:36
15:39:15.823 -> Water_Temp.:97.81, PWM:36
15:39:17.375 -> Water_Temp.:97.81, PWM:36
15:39:18.895 -> Water_Temp.:97.81, PWM:36
15:39:19.689 -> Water_Temp.:97.81, PWM:36
15:39:21.204 -> Water_Temp.:97.81, PWM:36
15:39:21.996 -> Water_Temp.:97.81, PWM:36
15:39:23.548 -> Water_Temp.:97.81, PWM:36
15:39:24.309 -> Water_Temp.:97.81, PWM:36
15:39:25.102 -> Water_Temp.:97.81, PWM:36
15:39:25.896 -> Water_Temp.:97.81, PWM:36
15:39:26.657 -> Water_Temp.:97.81, PWM:36
15:39:27.450 -> Water_Temp.:97.81, PWM:36
15:39:28.207 -> Water_Temp.:97.81, PWM:36
15:39:29.000 -> Water_Temp.:97.81, PWM:36
15:39:29.796 -> Water_Temp.:97.81, PWM:36
15:39:31.351 -> Water_Temp.:97.81, PWM:36
15:39:32.111 -> Water_Temp.:97.81, PWM:36
15:39:34.418 -> Water_Temp.:97.81, PWM:36
15:39:35.207 -> Water_Temp.:97.81, PWM:36
15:39:35.999 -> Water_Temp.:97.81, PWM:36
15:39:36.755 -> Water_Temp.:97.81, PWM:36
15:39:39.892 -> Water_Temp.:97.81, PWM:36
15:39:41.409 -> Water_Temp.:97.81, PWM:33
15:39:42.992 -> Water_Temp.:97.81, PWM:31
15:39:43.748 -> Water_Temp.:97.81, PWM:31
15:39:44.542 -> Water_Temp.:97.81, PWM:30
15:39:46.061 -> Water_Temp.:97.70, PWM:31
15:39:46.852 -> Water_Temp.:97.70, PWM:31
15:39:47.643 -> Water_Temp.:97.70, PWM:30
15:39:49.912 -> Water_Temp.:97.70, PWM:31
15:39:50.703 -> Water_Temp.:97.70, PWM:30
15:39:51.459 -> Water_Temp.:97.70, PWM:30
15:39:52.247 -> Water_Temp.:97.70, PWM:31
15:39:53.032 -> Water_Temp.:97.70, PWM:30
15:39:53.823 -> Water_Temp.:97.70, PWM:30
15:39:54.578 -> Water_Temp.:97.70, PWM:31
15:39:55.362 -> Water_Temp.:97.70, PWM:31
15:39:56.155 -> Water_Temp.:97.70, PWM:31
15:39:56.951 -> Water_Temp.:97.70, PWM:31
15:39:57.709 -> Water_Temp.:97.81, PWM:30
15:39:59.252 -> Water_Temp.:97.81, PWM:31
15:40:00.795 -> Water_Temp.:97.81, PWM:31
15:40:02.379 -> Water_Temp.:97.81, PWM:31
15:40:03.938 -> Water_Temp.:97.92, PWM:30
15:40:04.697 -> Water_Temp.:97.92, PWM:31
15:40:05.492 -> Water_Temp.:97.92, PWM:31
15:40:07.802 -> Water_Temp.:97.92, PWM:31
15:40:09.352 -> Water_Temp.:97.92, PWM:31
15:40:10.144 -> Water_Temp.:97.92, PWM:31
15:40:11.687 -> Water_Temp.:97.92, PWM:31
15:40:12.472 -> Water_Temp.:97.92, PWM:30
15:40:13.263 -> Water_Temp.:97.92, PWM:31
15:40:14.021 -> Water_Temp.:97.92, PWM:30
15:40:14.816 -> Water_Temp.:97.92, PWM:31
15:40:15.572 -> Water_Temp.:97.92, PWM:31
15:40:16.366 -> Water_Temp.:97.92, PWM:31
15:40:17.159 -> Water_Temp.:97.92, PWM:31
15:40:17.915 -> Water_Temp.:97.92, PWM:31
15:40:18.708 -> Water_Temp.:97.92, PWM:31
15:40:19.500 -> Water_Temp.:97.92, PWM:31
15:40:20.258 -> Water_Temp.:97.92, PWM:31
15:40:21.804 -> Water_Temp.:97.92, PWM:31
15:40:23.390 -> Water_Temp.:97.92, PWM:31
15:40:24.150 -> Water_Temp.:97.92, PWM:31
15:40:25.696 -> Water_Temp.:98.04, PWM:30
15:40:26.492 -> Water_Temp.:97.92, PWM:31
15:40:27.285 -> Water_Temp.:98.04, PWM:31
15:40:28.046 -> Water_Temp.:98.04, PWM:31
15:40:28.838 -> Water_Temp.:98.04, PWM:31
15:40:29.628 -> Water_Temp.:98.04, PWM:31
15:40:30.392 -> Water_Temp.:98.04, PWM:31
15:40:31.188 -> Water_Temp.:98.04, PWM:31
15:40:33.496 -> Water_Temp.:98.04, PWM:31
15:40:36.602 -> Water_Temp.:98.04, PWM:31
15:40:37.359 -> Water_Temp.:98.04, PWM:31
15:40:40.463 -> Water_Temp.:98.04, PWM:31
15:40:41.254 -> Water_Temp.:98.04, PWM:31
15:40:42.009 -> Water_Temp.:98.04, PWM:31
15:40:44.356 -> Water_Temp.:98.04, PWM:30
15:40:45.149 -> Water_Temp.:98.04, PWM:31
15:40:45.911 -> Water_Temp.:98.04, PWM:31
15:40:46.700 -> Water_Temp.:98.04, PWM:31
15:40:49.010 -> Water_Temp.:99.61, PWM:31
15:40:50.593 -> Water_Temp.:99.27, PWM:30
15:40:52.866 -> Water_Temp.:98.49, PWM:31
15:40:54.451 -> Water_Temp.:98.26, PWM:31
15:40:55.206 -> Water_Temp.:98.26, PWM:31
15:40:55.997 -> Water_Temp.:98.26, PWM:31
15:40:56.791 -> Water_Temp.:98.15, PWM:31
15:40:57.547 -> Water_Temp.:98.15, PWM:31
15:40:58.336 -> Water_Temp.:98.15, PWM:31
15:40:59.127 -> Water_Temp.:98.15, PWM:31
15:41:02.186 -> Water_Temp.:98.15, PWM:31
15:41:03.735 -> Water_Temp.:98.15, PWM:31
15:41:04.533 -> Water_Temp.:98.15, PWM:31
15:41:05.293 -> Water_Temp.:98.15, PWM:31
15:41:07.607 -> Water_Temp.:98.15, PWM:31
15:41:08.399 -> Water_Temp.:98.15, PWM:31
15:41:09.913 -> Water_Temp.:98.15, PWM:31
15:41:11.471 -> Water_Temp.:98.15, PWM:31
15:41:12.261 -> Water_Temp.:98.04, PWM:31
15:41:13.058 -> Water_Temp.:98.04, PWM:31
15:41:13.817 -> Water_Temp.:98.04, PWM:31
15:41:14.615 -> Water_Temp.:98.04, PWM:31
15:41:15.375 -> Water_Temp.:98.04, PWM:31
15:41:16.171 -> Water_Temp.:98.04, PWM:31
15:41:17.725 -> Water_Temp.:98.04, PWM:31
15:41:19.282 -> Water_Temp.:98.04, PWM:31
15:41:20.042 -> Water_Temp.:98.04, PWM:31
15:41:20.836 -> Water_Temp.:98.04, PWM:31
15:41:21.595 -> Water_Temp.:98.04, PWM:31
15:41:24.701 -> Water_Temp.:98.04, PWM:31
15:41:27.052 -> Water_Temp.:98.15, PWM:31
15:41:29.367 -> Water_Temp.:98.15, PWM:31
15:41:30.919 -> Water_Temp.:98.15, PWM:30
15:41:31.717 -> Water_Temp.:98.15, PWM:26
15:41:34.792 -> Water_Temp.:98.15, PWM:26
15:41:37.102 -> Water_Temp.:98.04, PWM:34
15:41:37.861 -> Water_Temp.:98.15, PWM:34
15:41:38.653 -> Water_Temp.:98.04, PWM:33
15:41:40.203 -> Water_Temp.:98.15, PWM:34
15:41:40.997 -> Water_Temp.:98.04, PWM:34
15:41:41.757 -> Water_Temp.:98.15, PWM:34
15:41:43.309 -> Water_Temp.:98.04, PWM:34
15:41:44.100 -> Water_Temp.:98.04, PWM:34
15:41:46.379 -> Water_Temp.:98.15, PWM:34
15:41:47.171 -> Water_Temp.:98.15, PWM:34
15:41:47.930 -> Water_Temp.:98.15, PWM:34
15:41:48.724 -> Water_Temp.:98.15, PWM:34
15:41:49.483 -> Water_Temp.:98.15, PWM:34
15:41:50.276 -> Water_Temp.:98.15, PWM:34
15:41:51.069 -> Water_Temp.:98.15, PWM:34
15:41:51.831 -> Water_Temp.:98.15, PWM:34
15:41:54.935 -> Water_Temp.:98.15, PWM:34
15:41:56.489 -> Water_Temp.:98.15, PWM:34
15:41:57.279 -> Water_Temp.:98.15, PWM:34
15:41:58.833 -> Water_Temp.:98.15, PWM:34
15:41:59.592 -> Water_Temp.:98.15, PWM:34
15:42:01.143 -> Water_Temp.:98.15, PWM:34
15:42:01.937 -> Water_Temp.:98.15, PWM:34
15:42:03.457 -> Water_Temp.:98.15, PWM:34
15:42:04.250 -> Water_Temp.:98.15, PWM:34
15:42:05.801 -> Water_Temp.:98.26, PWM:34
15:42:06.557 -> Water_Temp.:98.15, PWM:34
15:42:07.348 -> Water_Temp.:98.15, PWM:34
15:42:08.108 -> Water_Temp.:98.15, PWM:34
15:42:08.896 -> Water_Temp.:98.15, PWM:34
15:42:09.687 -> Water_Temp.:98.15, PWM:34
15:42:10.447 -> Water_Temp.:98.15, PWM:34
15:42:11.239 -> Water_Temp.:98.15, PWM:34
15:42:12.032 -> Water_Temp.:98.15, PWM:34
15:42:14.339 -> Water_Temp.:99.84, PWM:34
15:42:16.655 -> Water_Temp.:99.05, PWM:34
15:42:18.975 -> Water_Temp.:98.60, PWM:34
15:42:19.768 -> Water_Temp.:98.60, PWM:34
15:42:20.564 -> Water_Temp.:98.49, PWM:34
15:42:21.326 -> Water_Temp.:98.49, PWM:34
15:42:22.880 -> Water_Temp.:98.37, PWM:34
15:42:23.639 -> Water_Temp.:98.37, PWM:34
15:42:24.433 -> Water_Temp.:98.37, PWM:34
15:42:25.195 -> Water_Temp.:98.37, PWM:34
15:42:26.781 -> Water_Temp.:98.26, PWM:34
15:42:28.298 -> Water_Temp.:98.26, PWM:34
15:42:29.094 -> Water_Temp.:98.26, PWM:34
15:42:31.405 -> Water_Temp.:98.26, PWM:34
15:42:32.158 -> Water_Temp.:98.26, PWM:34
15:42:34.477 -> Water_Temp.:98.26, PWM:34
15:42:35.236 -> Water_Temp.:98.26, PWM:34
15:42:36.032 -> Water_Temp.:98.26, PWM:34
15:42:36.822 -> Water_Temp.:98.26, PWM:34
15:42:37.576 -> Water_Temp.:98.26, PWM:34
15:42:39.125 -> Water_Temp.:98.26, PWM:34
15:42:40.673 -> Water_Temp.:98.26, PWM:34
15:42:43.747 -> Water_Temp.:98.26, PWM:34
15:42:44.511 -> Water_Temp.:98.26, PWM:34
15:42:45.299 -> Water_Temp.:98.26, PWM:34
15:42:46.092 -> Water_Temp.:98.37, PWM:34
15:42:46.849 -> Water_Temp.:98.37, PWM:34
15:42:48.426 -> Water_Temp.:98.37, PWM:34
15:42:49.977 -> Water_Temp.:98.37, PWM:34
15:42:50.772 -> Water_Temp.:98.37, PWM:34
15:42:51.534 -> Water_Temp.:98.37, PWM:34
15:42:52.326 -> Water_Temp.:98.37, PWM:34
15:42:53.085 -> Water_Temp.:98.37, PWM:34
15:42:53.878 -> Water_Temp.:98.37, PWM:34
15:42:54.675 -> Water_Temp.:98.37, PWM:34
15:42:56.990 -> Water_Temp.:98.37, PWM:34
15:42:58.544 -> Water_Temp.:98.37, PWM:34
15:42:59.332 -> Water_Temp.:98.37, PWM:34
15:43:00.087 -> Water_Temp.:98.37, PWM:34
15:43:00.882 -> Water_Temp.:98.37, PWM:34
15:43:02.437 -> Water_Temp.:98.37, PWM:34
15:43:03.229 -> Water_Temp.:98.37, PWM:34
15:43:04.749 -> Water_Temp.:98.37, PWM:34
15:43:07.842 -> Water_Temp.:98.37, PWM:34
15:43:08.636 -> Water_Temp.:98.37, PWM:34
15:43:10.190 -> Water_Temp.:98.37, PWM:34
15:43:12.497 -> Water_Temp.:98.37, PWM:34
15:43:13.285 -> Water_Temp.:98.37, PWM:34
15:43:14.837 -> Water_Temp.:98.37, PWM:34
15:43:15.595 -> Water_Temp.:98.37, PWM:34
15:43:17.911 -> Water_Temp.:98.37, PWM:34
15:43:18.702 -> Water_Temp.:98.37, PWM:34
15:43:19.491 -> Water_Temp.:98.37, PWM:34
15:43:21.795 -> Water_Temp.:98.37, PWM:34
15:43:23.346 -> Water_Temp.:98.37, PWM:34
15:43:24.138 -> Water_Temp.:98.37, PWM:34
15:43:25.687 -> Water_Temp.:98.37, PWM:34
15:43:26.480 -> Water_Temp.:98.37, PWM:34
15:43:27.240 -> Water_Temp.:98.37, PWM:34
15:43:28.031 -> Water_Temp.:98.37, PWM:34
15:43:28.819 -> Water_Temp.:98.37, PWM:34
15:43:29.580 -> Water_Temp.:98.37, PWM:34
15:43:31.135 -> Water_Temp.:98.37, PWM:34
15:43:31.896 -> Water_Temp.:98.37, PWM:34
15:43:33.449 -> Water_Temp.:98.37, PWM:34
15:43:34.243 -> Water_Temp.:98.37, PWM:34
15:43:35.794 -> Water_Temp.:98.37, PWM:34
15:43:36.554 -> Water_Temp.:98.49, PWM:34
15:43:38.110 -> Water_Temp.:98.49, PWM:34
15:43:38.872 -> Water_Temp.:98.49, PWM:34
15:43:39.666 -> Water_Temp.:98.49, PWM:34
15:43:40.426 -> Water_Temp.:98.49, PWM:34
15:43:41.975 -> Water_Temp.:98.49, PWM:34
15:43:42.770 -> Water_Temp.:98.49, PWM:34
15:43:43.560 -> Water_Temp.:98.49, PWM:34
15:43:44.318 -> Water_Temp.:100.17, PWM:34
15:43:45.112 -> Water_Temp.:102.20, PWM:34
15:43:46.663 -> Water_Temp.:99.39, PWM:34
15:43:47.458 -> Water_Temp.:99.16, PWM:34
15:43:48.213 -> Water_Temp.:99.05, PWM:34
15:43:49.004 -> Water_Temp.:98.94, PWM:34
15:43:49.802 -> Water_Temp.:98.94, PWM:34
15:43:50.559 -> Water_Temp.:98.82, PWM:34
15:43:51.349 -> Water_Temp.:98.82, PWM:34
15:43:52.108 -> Water_Temp.:98.82, PWM:34
15:43:52.900 -> Water_Temp.:98.71, PWM:34
15:43:53.693 -> Water_Temp.:98.71, PWM:34
15:43:55.238 -> Water_Temp.:98.71, PWM:34
15:43:56.035 -> Water_Temp.:98.71, PWM:34
15:43:56.791 -> Water_Temp.:98.71, PWM:34
15:43:57.583 -> Water_Temp.:98.71, PWM:34
15:43:59.128 -> Water_Temp.:98.60, PWM:34
15:44:00.677 -> Water_Temp.:98.71, PWM:34
15:44:01.463 -> Water_Temp.:98.60, PWM:34
15:44:02.981 -> Water_Temp.:98.71, PWM:34
15:44:03.775 -> Water_Temp.:98.71, PWM:34
15:44:04.569 -> Water_Temp.:98.71, PWM:34
15:44:05.328 -> Water_Temp.:100.29, PWM:34
15:44:06.123 -> Water_Temp.:102.42, PWM:34
15:44:07.673 -> Water_Temp.:99.95, PWM:34
15:44:09.983 -> Water_Temp.:99.16, PWM:34
15:44:11.538 -> Water_Temp.:98.94, PWM:34
15:44:12.330 -> Water_Temp.:98.94, PWM:34
15:44:13.874 -> Water_Temp.:98.82, PWM:34
15:44:14.667 -> Water_Temp.:100.51, PWM:34
15:44:15.456 -> Water_Temp.:102.54, PWM:34
15:44:16.221 -> Water_Temp.:100.06, PWM:34
15:44:17.018 -> Water_Temp.:99.72, PWM:34
15:44:18.537 -> Water_Temp.:99.27, PWM:34
15:44:19.331 -> Water_Temp.:99.16, PWM:34
15:44:20.091 -> Water_Temp.:99.16, PWM:34
15:44:20.886 -> Water_Temp.:99.05, PWM:34
15:44:22.406 -> Water_Temp.:98.94, PWM:34
15:44:23.198 -> Water_Temp.:98.94, PWM:34
15:44:23.991 -> Water_Temp.:98.94, PWM:34
15:44:24.750 -> Water_Temp.:98.94, PWM:34
15:44:25.542 -> Water_Temp.:98.82, PWM:34
15:44:27.091 -> Water_Temp.:98.82, PWM:34
15:44:27.850 -> Water_Temp.:98.82, PWM:34
15:44:28.640 -> Water_Temp.:98.82, PWM:34
15:44:29.435 -> Water_Temp.:100.40, PWM:34
15:44:30.991 -> Water_Temp.:100.06, PWM:34
15:44:31.745 -> Water_Temp.:99.61, PWM:34
15:44:33.298 -> Water_Temp.:99.27, PWM:34
15:44:34.095 -> Water_Temp.:99.16, PWM:34
15:44:34.855 -> Water_Temp.:99.05, PWM:34
15:44:35.652 -> Water_Temp.:99.05, PWM:34
15:44:36.411 -> Water_Temp.:98.94, PWM:34
15:44:37.203 -> Water_Temp.:98.94, PWM:34
15:44:37.963 -> Water_Temp.:98.94, PWM:34
15:44:38.753 -> Water_Temp.:98.82, PWM:34
15:44:41.819 -> Water_Temp.:98.82, PWM:34
15:44:42.580 -> Water_Temp.:98.82, PWM:34
15:44:44.137 -> Water_Temp.:98.82, PWM:34
15:44:44.896 -> Water_Temp.:98.82, PWM:34
15:44:45.694 -> Water_Temp.:98.82, PWM:34
15:44:47.252 -> Water_Temp.:98.82, PWM:34
15:44:48.776 -> Water_Temp.:98.82, PWM:34
15:44:49.571 -> Water_Temp.:98.82, PWM:34
15:44:51.121 -> Water_Temp.:98.82, PWM:34
15:44:51.887 -> Water_Temp.:98.82, PWM:28
15:44:52.682 -> Water_Temp.:98.82, PWM:28
15:44:53.475 -> Water_Temp.:98.82, PWM:28
15:44:54.233 -> Water_Temp.:98.82, PWM:29
15:44:55.029 -> Water_Temp.:98.82, PWM:29
15:44:55.790 -> Water_Temp.:98.82, PWM:28
15:44:59.661 -> Water_Temp.:98.82, PWM:29
15:45:00.422 -> Water_Temp.:98.82, PWM:29
15:45:01.216 -> Water_Temp.:98.82, PWM:29
15:45:01.978 -> Water_Temp.:98.82, PWM:28
15:45:05.093 -> Water_Temp.:98.82, PWM:28
15:45:06.653 -> Water_Temp.:98.82, PWM:28
15:45:07.411 -> Water_Temp.:98.82, PWM:29
15:45:08.203 -> Water_Temp.:98.82, PWM:29
15:45:11.303 -> Water_Temp.:98.82, PWM:28
15:45:12.063 -> Water_Temp.:98.82, PWM:29
15:45:12.854 -> Water_Temp.:98.82, PWM:29
15:45:15.964 -> Water_Temp.:100.17, PWM:28
15:45:16.756 -> Water_Temp.:99.72, PWM:29
15:45:17.516 -> Water_Temp.:99.50, PWM:28
15:45:18.308 -> Water_Temp.:99.39, PWM:28
15:45:19.105 -> Water_Temp.:99.27, PWM:28
15:45:19.865 -> Water_Temp.:99.27, PWM:29
15:45:22.173 -> Water_Temp.:99.05, PWM:29
15:45:22.961 -> Water_Temp.:99.05, PWM:28
15:45:24.510 -> Water_Temp.:99.05, PWM:28
15:45:26.815 -> Water_Temp.:98.94, PWM:29
15:45:27.605 -> Water_Temp.:98.94, PWM:28
15:45:28.363 -> Water_Temp.:98.94, PWM:29
15:45:29.160 -> Water_Temp.:98.94, PWM:28
15:45:29.921 -> Water_Temp.:98.94, PWM:29
15:45:30.717 -> Water_Temp.:98.94, PWM:28
15:45:31.505 -> Water_Temp.:98.94, PWM:29
15:45:34.546 -> Water_Temp.:98.94, PWM:28
15:45:35.338 -> Water_Temp.:98.94, PWM:29
15:45:36.898 -> Water_Temp.:98.82, PWM:28
15:45:39.212 -> Water_Temp.:98.82, PWM:28
15:45:39.975 -> Water_Temp.:98.82, PWM:28
15:45:40.768 -> Water_Temp.:98.82, PWM:28
15:45:43.077 -> Water_Temp.:98.82, PWM:29
15:45:45.354 -> Water_Temp.:98.82, PWM:28
15:45:46.913 -> Water_Temp.:98.82, PWM:29
15:45:49.231 -> Water_Temp.:98.82, PWM:29
15:45:50.778 -> Water_Temp.:98.71, PWM:28
15:45:51.537 -> Water_Temp.:98.71, PWM:28
15:45:52.335 -> Water_Temp.:98.71, PWM:28
15:45:53.891 -> Water_Temp.:98.71, PWM:28
15:45:54.681 -> Water_Temp.:98.71, PWM:28
15:45:55.440 -> Water_Temp.:98.71, PWM:29
15:45:56.996 -> Water_Temp.:98.71, PWM:28
15:45:57.759 -> Water_Temp.:98.71, PWM:28
15:45:58.553 -> Water_Temp.:98.71, PWM:29
15:45:59.311 -> Water_Temp.:98.71, PWM:28
15:46:00.101 -> Water_Temp.:98.71, PWM:29
15:46:02.447 -> Water_Temp.:98.71, PWM:29
15:46:03.242 -> Water_Temp.:98.71, PWM:29
15:46:04.004 -> Water_Temp.:98.71, PWM:29
15:46:05.558 -> Water_Temp.:98.71, PWM:28
15:46:07.119 -> Water_Temp.:98.71, PWM:28
15:46:07.880 -> Water_Temp.:98.71, PWM:29
15:46:08.671 -> Water_Temp.:98.71, PWM:29
15:46:10.955 -> Water_Temp.:98.82, PWM:28
15:46:11.748 -> Water_Temp.:98.82, PWM:29
15:46:12.504 -> Water_Temp.:98.82, PWM:29
15:46:13.298 -> Water_Temp.:98.82, PWM:29
15:46:14.844 -> Water_Temp.:98.82, PWM:28
15:46:15.638 -> Water_Temp.:98.82, PWM:28
15:46:16.401 -> Water_Temp.:98.82, PWM:29
15:46:17.194 -> Water_Temp.:98.82, PWM:28
15:46:17.953 -> Water_Temp.:98.82, PWM:28
15:46:18.745 -> Water_Temp.:98.82, PWM:29
15:46:19.539 -> Water_Temp.:98.82, PWM:29
15:46:20.300 -> Water_Temp.:98.82, PWM:28
15:46:21.099 -> Water_Temp.:98.94, PWM:28
15:46:21.857 -> Water_Temp.:98.94, PWM:28
15:46:23.406 -> Water_Temp.:98.94, PWM:28
15:46:24.197 -> Water_Temp.:98.94, PWM:28
15:46:24.992 -> Water_Temp.:98.94, PWM:28
15:46:25.753 -> Water_Temp.:98.94, PWM:28
15:46:26.547 -> Water_Temp.:98.94, PWM:29
15:46:27.306 -> Water_Temp.:98.94, PWM:29
15:46:28.103 -> Water_Temp.:98.94, PWM:28
15:46:31.218 -> Water_Temp.:98.82, PWM:28
15:46:32.736 -> Water_Temp.:98.94, PWM:28
15:46:34.290 -> Water_Temp.:98.94, PWM:28
15:46:35.842 -> Water_Temp.:98.94, PWM:29
15:46:36.637 -> Water_Temp.:98.94, PWM:29
15:46:37.392 -> Water_Temp.:98.94, PWM:28
15:46:38.187 -> Water_Temp.:98.94, PWM:29
15:46:39.707 -> Water_Temp.:98.94, PWM:28
15:46:42.055 -> Water_Temp.:98.94, PWM:29
15:46:43.610 -> Water_Temp.:98.94, PWM:29
15:46:44.370 -> Water_Temp.:98.94, PWM:29
15:46:45.163 -> Water_Temp.:98.94, PWM:28
15:46:47.473 -> Water_Temp.:98.94, PWM:29
15:46:48.266 -> Water_Temp.:98.94, PWM:29
15:46:49.060 -> Water_Temp.:98.94, PWM:28
15:46:49.820 -> Water_Temp.:98.94, PWM:28
15:46:50.616 -> Water_Temp.:98.94, PWM:28
15:46:51.378 -> Water_Temp.:98.94, PWM:28
15:46:52.172 -> Water_Temp.:98.94, PWM:28
15:46:53.724 -> Water_Temp.:98.94, PWM:29
15:46:54.483 -> Water_Temp.:98.94, PWM:28
15:46:56.040 -> Water_Temp.:98.82, PWM:28
15:46:56.836 -> Water_Temp.:98.82, PWM:28
15:46:57.628 -> Water_Temp.:98.82, PWM:28
15:46:59.909 -> Water_Temp.:98.94, PWM:28
15:47:01.459 -> Water_Temp.:98.94, PWM:28
15:47:04.530 -> Water_Temp.:98.94, PWM:29
15:47:05.323 -> Water_Temp.:98.94, PWM:28
15:47:06.876 -> Water_Temp.:98.94, PWM:28
15:47:07.675 -> Water_Temp.:98.94, PWM:28
15:47:08.435 -> Water_Temp.:98.94, PWM:28
15:47:09.229 -> Water_Temp.:98.94, PWM:28
15:47:10.022 -> Water_Temp.:98.94, PWM:28
15:47:10.782 -> Water_Temp.:98.94, PWM:28
15:47:11.581 -> Water_Temp.:98.94, PWM:28
15:47:12.340 -> Water_Temp.:98.94, PWM:29
15:47:13.891 -> Water_Temp.:98.94, PWM:28
15:47:14.681 -> Water_Temp.:98.94, PWM:29
15:47:15.473 -> Water_Temp.:98.94, PWM:28
15:47:18.578 -> Water_Temp.:98.94, PWM:28
15:47:19.336 -> Water_Temp.:98.94, PWM:28
15:47:20.889 -> Water_Temp.:98.94, PWM:28
15:47:22.444 -> Water_Temp.:98.94, PWM:28
15:47:23.207 -> Water_Temp.:98.94, PWM:28
15:47:24.004 -> Water_Temp.:98.94, PWM:28
15:47:24.766 -> Water_Temp.:98.94, PWM:28
15:47:26.324 -> Water_Temp.:98.94, PWM:28
15:47:27.083 -> Water_Temp.:98.94, PWM:29
15:47:27.875 -> Water_Temp.:98.94, PWM:28
15:47:28.638 -> Water_Temp.:98.94, PWM:29
15:47:29.431 -> Water_Temp.:98.94, PWM:28
15:47:30.225 -> Water_Temp.:98.94, PWM:28
15:47:30.984 -> Water_Temp.:98.94, PWM:28
15:47:34.062 -> Water_Temp.:98.82, PWM:28
15:47:34.859 -> Water_Temp.:98.82, PWM:28
15:47:36.409 -> Water_Temp.:98.82, PWM:28
15:47:37.204 -> Water_Temp.:98.82, PWM:28
15:47:38.762 -> Water_Temp.:98.82, PWM:28
15:47:41.069 -> Water_Temp.:98.82, PWM:28
15:47:41.826 -> Water_Temp.:98.82, PWM:28
15:47:43.412 -> Water_Temp.:98.82, PWM:28
15:47:44.171 -> Water_Temp.:98.82, PWM:28
15:47:44.963 -> Water_Temp.:98.82, PWM:28
15:47:45.723 -> Water_Temp.:98.82, PWM:28
15:47:46.514 -> Water_Temp.:98.82, PWM:28
15:47:47.308 -> Water_Temp.:98.82, PWM:28
15:47:48.064 -> Water_Temp.:98.82, PWM:28
15:47:48.859 -> Water_Temp.:98.82, PWM:28
15:47:49.652 -> Water_Temp.:98.82, PWM:28
15:47:50.406 -> Water_Temp.:98.82, PWM:28
15:47:51.189 -> Water_Temp.:98.82, PWM:28
15:47:52.772 -> Water_Temp.:98.82, PWM:28
15:47:53.539 -> Water_Temp.:98.82, PWM:28
15:47:54.331 -> Water_Temp.:98.82, PWM:28
15:47:58.232 -> Water_Temp.:98.82, PWM:28
15:47:58.993 -> Water_Temp.:98.82, PWM:28
15:47:59.786 -> Water_Temp.:98.82, PWM:28
15:48:00.542 -> Water_Temp.:98.82, PWM:28
15:48:01.332 -> Water_Temp.:98.82, PWM:29
15:48:02.122 -> Water_Temp.:98.82, PWM:28
15:48:02.881 -> Water_Temp.:98.82, PWM:28
15:48:05.194 -> Water_Temp.:98.82, PWM:28
15:48:05.986 -> Water_Temp.:98.82, PWM:28
15:48:06.779 -> Water_Temp.:98.82, PWM:28
15:48:09.090 -> Water_Temp.:98.82, PWM:28
15:48:09.848 -> Water_Temp.:98.82, PWM:28
15:48:10.644 -> Water_Temp.:98.82, PWM:28
15:48:11.401 -> Water_Temp.:98.82, PWM:28
15:48:12.931 -> Water_Temp.:98.82, PWM:28
15:48:13.726 -> Water_Temp.:98.82, PWM:28
15:48:16.038 -> Water_Temp.:98.82, PWM:28
15:48:16.796 -> Water_Temp.:98.82, PWM:28
15:48:18.354 -> Water_Temp.:98.82, PWM:28
15:48:19.111 -> Water_Temp.:98.82, PWM:28
15:48:20.665 -> Water_Temp.:98.82, PWM:28
15:48:21.423 -> Water_Temp.:98.82, PWM:28
15:48:22.218 -> Water_Temp.:98.82, PWM:28
15:48:24.563 -> Water_Temp.:98.82, PWM:28
15:48:25.322 -> Water_Temp.:98.82, PWM:28
15:48:26.875 -> Water_Temp.:98.82, PWM:28
15:48:29.191 -> Water_Temp.:98.82, PWM:28
15:48:30.745 -> Water_Temp.:98.82, PWM:28
15:48:32.294 -> Water_Temp.:98.82, PWM:28
15:48:35.367 -> Water_Temp.:98.82, PWM:28
15:48:36.919 -> Water_Temp.:98.71, PWM:28
15:48:37.681 -> Water_Temp.:98.71, PWM:28
15:48:39.998 -> Water_Temp.:98.71, PWM:28
15:48:41.546 -> Water_Temp.:98.71, PWM:28
15:48:43.886 -> Water_Temp.:98.71, PWM:28
15:48:46.977 -> Water_Temp.:98.71, PWM:28
15:48:49.320 -> Water_Temp.:98.71, PWM:28
15:48:50.079 -> Water_Temp.:98.71, PWM:28
15:48:51.634 -> Water_Temp.:98.71, PWM:28
15:48:53.190 -> Water_Temp.:98.71, PWM:28
15:48:53.950 -> Water_Temp.:98.71, PWM:28
15:48:54.746 -> Water_Temp.:98.71, PWM:28
15:48:55.506 -> Water_Temp.:98.71, PWM:28
15:48:57.823 -> Water_Temp.:98.71, PWM:28
15:48:58.622 -> Water_Temp.:98.71, PWM:28
15:48:59.381 -> Water_Temp.:98.71, PWM:28
15:49:00.937 -> Water_Temp.:98.71, PWM:28
15:49:04.777 -> Water_Temp.:98.71, PWM:28
15:49:05.539 -> Water_Temp.:98.71, PWM:28
15:49:06.334 -> Water_Temp.:98.71, PWM:28
15:49:07.124 -> Water_Temp.:98.71, PWM:28
15:49:08.654 -> Water_Temp.:98.71, PWM:28
15:49:09.419 -> Water_Temp.:98.71, PWM:28
15:49:10.210 -> Water_Temp.:98.71, PWM:28
15:49:11.006 -> Water_Temp.:98.71, PWM:28
15:49:11.768 -> Water_Temp.:98.71, PWM:28
15:49:13.323 -> Water_Temp.:98.71, PWM:28
15:49:14.084 -> Water_Temp.:98.71, PWM:28
15:49:14.878 -> Water_Temp.:98.71, PWM:28
15:49:15.670 -> Water_Temp.:98.71, PWM:28
15:49:16.431 -> Water_Temp.:98.71, PWM:28
15:49:17.226 -> Water_Temp.:98.71, PWM:29
15:49:17.985 -> Water_Temp.:98.71, PWM:23
15:49:18.778 -> Water_Temp.:100.40, PWM:23
15:49:19.569 -> Water_Temp.:102.54, PWM:23
15:49:20.332 -> Water_Temp.:100.06, PWM:23
15:49:21.883 -> Water_Temp.:99.39, PWM:21
15:49:22.673 -> Water_Temp.:99.27, PWM:28
15:49:24.994 -> Water_Temp.:99.05, PWM:28
15:49:25.754 -> Water_Temp.:99.05, PWM:29
15:49:26.548 -> Water_Temp.:98.94, PWM:29
15:49:27.343 -> Water_Temp.:98.94, PWM:29
15:49:29.659 -> Water_Temp.:98.94, PWM:31
15:49:31.213 -> Water_Temp.:98.82, PWM:31
15:49:31.971 -> Water_Temp.:98.82, PWM:31
15:49:32.768 -> Water_Temp.:98.82, PWM:31
15:49:33.526 -> Water_Temp.:98.82, PWM:31
15:49:34.319 -> Water_Temp.:98.82, PWM:31
15:49:35.875 -> Water_Temp.:98.82, PWM:31
15:49:36.631 -> Water_Temp.:98.82, PWM:31
15:49:37.425 -> Water_Temp.:98.82, PWM:31
15:49:38.987 -> Water_Temp.:98.82, PWM:31
15:49:43.643 -> Water_Temp.:98.82, PWM:31
15:49:45.956 -> Water_Temp.:98.82, PWM:31
15:49:46.750 -> Water_Temp.:98.82, PWM:31
15:49:47.510 -> Water_Temp.:98.82, PWM:31
15:49:48.302 -> Water_Temp.:98.82, PWM:31
15:49:50.650 -> Water_Temp.:98.82, PWM:31
15:49:51.409 -> Water_Temp.:98.82, PWM:31
15:49:52.206 -> Water_Temp.:98.82, PWM:31
15:49:52.996 -> Water_Temp.:98.82, PWM:31
15:49:54.515 -> Water_Temp.:98.82, PWM:31
15:49:55.312 -> Water_Temp.:98.82, PWM:31
15:49:56.071 -> Water_Temp.:98.82, PWM:31
15:49:56.866 -> Water_Temp.:98.82, PWM:31
15:49:57.659 -> Water_Temp.:98.82, PWM:31
15:49:59.183 -> Water_Temp.:98.82, PWM:31
15:49:59.980 -> Water_Temp.:98.82, PWM:31
15:50:02.289 -> Water_Temp.:98.82, PWM:31
15:50:03.050 -> Water_Temp.:98.82, PWM:31
15:50:03.843 -> Water_Temp.:98.82, PWM:31
15:50:04.635 -> Water_Temp.:98.82, PWM:31
15:50:05.393 -> Water_Temp.:98.82, PWM:31
15:50:06.191 -> Water_Temp.:98.82, PWM:31
15:50:06.953 -> Water_Temp.:98.82, PWM:31
15:50:08.506 -> Water_Temp.:98.82, PWM:31
15:50:09.303 -> Water_Temp.:98.82, PWM:31
15:50:13.142 -> Water_Temp.:98.82, PWM:31
15:50:13.901 -> Water_Temp.:98.82, PWM:31
15:50:14.696 -> Water_Temp.:98.82, PWM:31
15:50:15.488 -> Water_Temp.:98.82, PWM:31
15:50:16.251 -> Water_Temp.:98.82, PWM:31
15:50:17.045 -> Water_Temp.:98.82, PWM:31
15:50:17.803 -> Water_Temp.:98.82, PWM:31
15:50:18.601 -> Water_Temp.:98.71, PWM:31
15:50:20.918 -> Water_Temp.:98.71, PWM:31
15:50:21.715 -> Water_Temp.:98.71, PWM:31
15:50:22.472 -> Water_Temp.:98.71, PWM:31
15:50:23.267 -> Water_Temp.:98.71, PWM:31
15:50:24.818 -> Water_Temp.:98.71, PWM:31
15:50:25.578 -> Water_Temp.:98.71, PWM:31
15:50:26.372 -> Water_Temp.:98.71, PWM:31
15:50:27.133 -> Water_Temp.:100.29, PWM:31
15:50:27.927 -> Water_Temp.:102.42, PWM:31
15:50:29.480 -> Water_Temp.:99.61, PWM:31
15:50:31.033 -> Water_Temp.:99.27, PWM:31
15:50:34.936 -> Water_Temp.:98.94, PWM:31
15:50:36.454 -> Water_Temp.:98.94, PWM:31
15:50:37.250 -> Water_Temp.:98.94, PWM:31
15:50:38.043 -> Water_Temp.:98.94, PWM:31
15:50:38.801 -> Water_Temp.:98.94, PWM:31
15:50:39.595 -> Water_Temp.:98.94, PWM:31
15:50:40.358 -> Water_Temp.:98.94, PWM:31
15:50:41.152 -> Water_Temp.:98.94, PWM:31
15:50:41.913 -> Water_Temp.:98.94, PWM:31
15:50:42.714 -> Water_Temp.:98.94, PWM:31
15:50:44.274 -> Water_Temp.:98.94, PWM:31
15:50:45.034 -> Water_Temp.:98.94, PWM:31
15:50:45.829 -> Water_Temp.:98.94, PWM:31
15:50:46.626 -> Water_Temp.:98.94, PWM:31
15:50:47.387 -> Water_Temp.:98.94, PWM:31
15:50:48.151 -> Water_Temp.:98.94, PWM:31
15:50:48.947 -> Water_Temp.:98.94, PWM:31
15:50:49.710 -> Water_Temp.:98.94, PWM:31
15:50:54.335 -> Water_Temp.:98.94, PWM:31
15:50:55.095 -> Water_Temp.:98.94, PWM:31
15:50:55.888 -> Water_Temp.:98.94, PWM:31
15:50:57.441 -> Water_Temp.:98.94, PWM:31
15:50:59.753 -> Water_Temp.:98.94, PWM:31
15:51:02.103 -> Water_Temp.:98.94, PWM:31
15:51:02.864 -> Water_Temp.:98.94, PWM:31
15:51:06.730 -> Water_Temp.:98.82, PWM:31
15:51:08.281 -> Water_Temp.:98.82, PWM:31
15:51:09.074 -> Water_Temp.:98.82, PWM:31
15:51:09.833 -> Water_Temp.:98.82, PWM:31
15:51:10.626 -> Water_Temp.:98.82, PWM:31
15:51:12.145 -> Water_Temp.:98.82, PWM:31
15:51:12.932 -> Water_Temp.:98.82, PWM:31
15:51:13.727 -> Water_Temp.:98.82, PWM:31
15:51:14.519 -> Water_Temp.:98.82, PWM:31
15:51:16.803 -> Water_Temp.:98.82, PWM:31
15:51:17.600 -> Water_Temp.:98.82, PWM:31
15:51:19.150 -> Water_Temp.:98.82, PWM:31
15:51:19.911 -> Water_Temp.:98.82, PWM:31
15:51:20.704 -> Water_Temp.:98.82, PWM:31
15:51:21.466 -> Water_Temp.:98.82, PWM:31
15:51:23.782 -> Water_Temp.:98.82, PWM:31
15:51:24.572 -> Water_Temp.:98.82, PWM:31
15:51:25.333 -> Water_Temp.:98.82, PWM:31
15:51:26.131 -> Water_Temp.:98.82, PWM:31
15:51:27.687 -> Water_Temp.:98.82, PWM:31
15:51:28.445 -> Water_Temp.:98.82, PWM:31
15:51:30.000 -> Water_Temp.:98.94, PWM:31
15:51:32.351 -> Water_Temp.:98.94, PWM:31
15:51:33.110 -> Water_Temp.:98.94, PWM:31
15:51:33.901 -> Water_Temp.:98.94, PWM:31
15:51:34.694 -> Water_Temp.:98.94, PWM:31
15:51:36.217 -> Water_Temp.:98.94, PWM:31
15:51:37.010 -> Water_Temp.:98.94, PWM:31
15:51:37.767 -> Water_Temp.:98.94, PWM:31
15:51:38.561 -> Water_Temp.:98.94, PWM:36
15:51:39.317 -> Water_Temp.:98.94, PWM:35
15:51:40.906 -> Water_Temp.:98.94, PWM:36
15:51:41.663 -> Water_Temp.:98.94, PWM:32
15:51:42.456 -> Water_Temp.:98.94, PWM:28
15:51:43.245 -> Water_Temp.:98.94, PWM:28
15:51:44.001 -> Water_Temp.:98.94, PWM:28
15:51:44.794 -> Water_Temp.:98.94, PWM:31
15:51:45.587 -> Water_Temp.:98.94, PWM:31
15:51:46.348 -> Water_Temp.:98.94, PWM:31
15:51:47.140 -> Water_Temp.:98.94, PWM:31
15:51:47.902 -> Water_Temp.:98.94, PWM:31
15:51:48.701 -> Water_Temp.:98.94, PWM:31
15:51:51.018 -> Water_Temp.:98.94, PWM:31
15:51:51.777 -> Water_Temp.:98.82, PWM:31
15:51:52.570 -> Water_Temp.:98.94, PWM:31
15:51:53.331 -> Water_Temp.:98.94, PWM:31
15:51:54.126 -> Water_Temp.:98.82, PWM:31
15:51:55.646 -> Water_Temp.:98.82, PWM:32
15:51:56.442 -> Water_Temp.:98.82, PWM:33
15:51:58.751 -> Water_Temp.:98.82, PWM:32
15:51:59.542 -> Water_Temp.:98.82, PWM:32
15:52:00.301 -> Water_Temp.:98.82, PWM:32
15:52:01.096 -> Water_Temp.:98.82, PWM:33
15:52:02.650 -> Water_Temp.:98.82, PWM:32
15:52:03.407 -> Water_Temp.:98.82, PWM:32
15:52:04.201 -> Water_Temp.:98.82, PWM:32
15:52:04.994 -> Water_Temp.:98.82, PWM:32
15:52:05.756 -> Water_Temp.:98.82, PWM:32
15:52:08.078 -> Water_Temp.:98.82, PWM:32
15:52:08.843 -> Water_Temp.:98.82, PWM:32
15:52:09.637 -> Water_Temp.:98.82, PWM:32
15:52:10.433 -> Water_Temp.:98.82, PWM:32
15:52:11.198 -> Water_Temp.:98.82, PWM:33
15:52:11.989 -> Water_Temp.:98.82, PWM:33
15:52:12.751 -> Water_Temp.:98.82, PWM:32
15:52:15.824 -> Water_Temp.:98.82, PWM:32
15:52:16.623 -> Water_Temp.:98.82, PWM:32
15:52:18.176 -> Water_Temp.:98.82, PWM:33
15:52:20.489 -> Water_Temp.:98.82, PWM:33
15:52:22.075 -> Water_Temp.:98.82, PWM:32
15:52:23.594 -> Water_Temp.:98.71, PWM:32
15:52:25.943 -> Water_Temp.:98.71, PWM:32
15:52:28.226 -> Water_Temp.:98.71, PWM:32
15:52:29.781 -> Water_Temp.:98.71, PWM:32
15:52:30.578 -> Water_Temp.:98.71, PWM:32
15:52:32.133 -> Water_Temp.:98.71, PWM:32
15:52:32.889 -> Water_Temp.:98.71, PWM:33
15:52:33.683 -> Water_Temp.:98.71, PWM:33
15:52:34.480 -> Water_Temp.:98.71, PWM:34
15:52:36.789 -> Water_Temp.:98.71, PWM:34
15:52:38.346 -> Water_Temp.:98.71, PWM:33
15:52:39.105 -> Water_Temp.:98.71, PWM:33
15:52:39.898 -> Water_Temp.:98.71, PWM:34
15:52:40.659 -> Water_Temp.:98.71, PWM:34
15:52:41.455 -> Water_Temp.:98.71, PWM:33
15:52:43.008 -> Water_Temp.:98.71, PWM:34
15:52:43.804 -> Water_Temp.:98.71, PWM:34
15:52:45.359 -> Water_Temp.:98.71, PWM:34
15:52:47.670 -> Water_Temp.:98.71, PWM:34
15:52:48.430 -> Water_Temp.:98.71, PWM:34
15:52:49.227 -> Water_Temp.:98.71, PWM:34
15:52:51.512 -> Water_Temp.:98.71, PWM:34
15:52:52.307 -> Water_Temp.:98.71, PWM:34
15:52:53.829 -> Water_Temp.:98.71, PWM:34
15:52:55.384 -> Water_Temp.:98.71, PWM:34
15:52:56.181 -> Water_Temp.:98.71, PWM:34
15:52:57.731 -> Water_Temp.:98.71, PWM:34
15:53:00.079 -> Water_Temp.:98.71, PWM:34
15:53:00.844 -> Water_Temp.:98.71, PWM:34
15:53:02.399 -> Water_Temp.:98.71, PWM:34
15:53:03.952 -> Water_Temp.:98.71, PWM:33
15:53:04.706 -> Water_Temp.:98.71, PWM:34
15:53:07.014 -> Water_Temp.:98.71, PWM:34
15:53:07.808 -> Water_Temp.:98.71, PWM:34
15:53:09.330 -> Water_Temp.:98.71, PWM:34
15:53:10.122 -> Water_Temp.:98.71, PWM:33
15:53:10.914 -> Water_Temp.:98.71, PWM:34
15:53:11.672 -> Water_Temp.:98.71, PWM:34
15:53:13.230 -> Water_Temp.:98.71, PWM:34
15:53:13.990 -> Water_Temp.:98.71, PWM:34
15:53:14.779 -> Water_Temp.:98.71, PWM:33
15:53:15.573 -> Water_Temp.:98.71, PWM:34
15:53:16.335 -> Water_Temp.:98.71, PWM:34
15:53:17.128 -> Water_Temp.:98.71, PWM:34
15:53:18.649 -> Water_Temp.:98.71, PWM:33
15:53:19.445 -> Water_Temp.:98.71, PWM:34
15:53:20.237 -> Water_Temp.:98.71, PWM:34
15:53:20.998 -> Water_Temp.:98.71, PWM:34
15:53:21.796 -> Water_Temp.:98.71, PWM:34
15:53:22.558 -> Water_Temp.:98.71, PWM:34
15:53:23.353 -> Water_Temp.:98.71, PWM:34
15:53:24.110 -> Water_Temp.:98.71, PWM:34
15:53:27.225 -> Water_Temp.:98.71, PWM:34
15:53:28.746 -> Water_Temp.:98.71, PWM:34
15:53:29.544 -> Water_Temp.:98.71, PWM:34
15:53:31.061 -> Water_Temp.:98.71, PWM:34
15:53:31.851 -> Water_Temp.:98.71, PWM:34
15:53:33.375 -> Water_Temp.:98.71, PWM:34
15:53:34.968 -> Water_Temp.:98.71, PWM:33
15:53:37.248 -> Water_Temp.:98.82, PWM:34
15:53:38.043 -> Water_Temp.:98.71, PWM:34
15:53:38.806 -> Water_Temp.:98.71, PWM:34
15:53:41.126 -> Water_Temp.:98.82, PWM:34
15:53:41.920 -> Water_Temp.:98.82, PWM:33
15:53:42.683 -> Water_Temp.:98.82, PWM:34
15:53:43.475 -> Water_Temp.:98.82, PWM:34
15:53:45.032 -> Water_Temp.:98.82, PWM:34
15:53:45.790 -> Water_Temp.:98.82, PWM:34
15:53:46.587 -> Water_Temp.:98.82, PWM:34
15:53:47.378 -> Water_Temp.:98.82, PWM:34
15:53:48.137 -> Water_Temp.:98.82, PWM:34
15:53:48.931 -> Water_Temp.:98.82, PWM:33
15:53:49.690 -> Water_Temp.:98.82, PWM:34
15:53:52.007 -> Water_Temp.:98.82, PWM:34
15:53:53.566 -> Water_Temp.:98.82, PWM:34
15:53:55.118 -> Water_Temp.:98.82, PWM:34
15:53:55.916 -> Water_Temp.:98.82, PWM:34
15:53:58.239 -> Water_Temp.:100.51, PWM:34
15:53:59.000 -> Water_Temp.:102.54, PWM:34
15:53:59.795 -> Water_Temp.:100.17, PWM:34
15:54:00.559 -> Water_Temp.:100.17, PWM:34
15:54:02.147 -> Water_Temp.:99.39, PWM:34
15:54:02.909 -> Water_Temp.:99.27, PWM:34
15:54:03.704 -> Water_Temp.:99.16, PWM:34
15:54:04.462 -> Water_Temp.:99.16, PWM:34
15:54:06.015 -> Water_Temp.:99.05, PWM:34
15:54:07.571 -> Water_Temp.:98.94, PWM:34
15:54:08.363 -> Water_Temp.:98.94, PWM:34
15:54:09.125 -> Water_Temp.:98.94, PWM:34
15:54:10.683 -> Water_Temp.:98.94, PWM:34
15:54:11.477 -> Water_Temp.:98.94, PWM:34
15:54:13.749 -> Water_Temp.:98.82, PWM:33
15:54:14.544 -> Water_Temp.:98.82, PWM:34
15:54:18.406 -> Water_Temp.:98.82, PWM:34
15:54:19.165 -> Water_Temp.:98.82, PWM:34
15:54:19.955 -> Water_Temp.:98.82, PWM:34
15:54:20.750 -> Water_Temp.:98.82, PWM:34
15:54:21.512 -> Water_Temp.:98.82, PWM:34
15:54:23.073 -> Water_Temp.:98.82, PWM:34
15:54:23.869 -> Water_Temp.:98.82, PWM:34
15:54:24.630 -> Water_Temp.:98.82, PWM:34
15:54:25.424 -> Water_Temp.:98.82, PWM:34
15:54:26.188 -> Water_Temp.:98.82, PWM:34
15:54:26.981 -> Water_Temp.:98.82, PWM:34
15:54:27.773 -> Water_Temp.:98.82, PWM:34
15:54:29.330 -> Water_Temp.:98.82, PWM:34
15:54:30.853 -> Water_Temp.:98.71, PWM:34
15:54:31.649 -> Water_Temp.:98.82, PWM:34
15:54:32.411 -> Water_Temp.:98.82, PWM:34
15:54:33.208 -> Water_Temp.:98.71, PWM:34
15:54:33.970 -> Water_Temp.:98.71, PWM:34
15:54:34.765 -> Water_Temp.:98.82, PWM:34
15:54:35.558 -> Water_Temp.:98.71, PWM:34
15:54:36.318 -> Water_Temp.:98.71, PWM:34
15:54:37.110 -> Water_Temp.:98.82, PWM:34
15:54:37.870 -> Water_Temp.:98.82, PWM:34
15:54:39.426 -> Water_Temp.:98.82, PWM:34
15:54:40.979 -> Water_Temp.:98.82, PWM:34
15:54:43.292 -> Water_Temp.:98.82, PWM:34
15:54:44.055 -> Water_Temp.:98.82, PWM:34
15:54:44.848 -> Water_Temp.:98.82, PWM:34
15:54:45.610 -> Water_Temp.:98.82, PWM:34
15:54:47.925 -> Water_Temp.:98.82, PWM:34
15:54:48.684 -> Water_Temp.:98.82, PWM:34
15:54:49.482 -> Water_Temp.:98.82, PWM:34
15:54:50.242 -> Water_Temp.:98.82, PWM:34
15:54:51.792 -> Water_Temp.:98.82, PWM:34
15:54:52.555 -> Water_Temp.:98.82, PWM:34
15:54:53.350 -> Water_Temp.:98.82, PWM:34
15:54:54.143 -> Water_Temp.:98.82, PWM:34
15:54:54.902 -> Water_Temp.:98.82, PWM:34
15:54:55.697 -> Water_Temp.:98.82, PWM:34
15:54:56.457 -> Water_Temp.:98.82, PWM:34
15:54:57.254 -> Water_Temp.:98.82, PWM:34
15:54:58.047 -> Water_Temp.:98.82, PWM:34
15:54:58.807 -> Water_Temp.:98.82, PWM:34
15:55:00.363 -> Water_Temp.:98.82, PWM:34
15:55:02.678 -> Water_Temp.:98.82, PWM:34
15:55:04.237 -> Water_Temp.:98.82, PWM:34
15:55:04.992 -> Water_Temp.:98.82, PWM:34
15:55:06.549 -> Water_Temp.:98.82, PWM:34
15:55:07.308 -> Water_Temp.:98.82, PWM:34
15:55:08.853 -> Water_Temp.:98.82, PWM:34
15:55:09.648 -> Water_Temp.:98.82, PWM:34
15:55:10.409 -> Water_Temp.:98.82, PWM:34
15:55:11.203 -> Water_Temp.:98.82, PWM:34
15:55:12.718 -> Water_Temp.:98.82, PWM:34
15:55:13.509 -> Water_Temp.:98.82, PWM:34
15:55:14.298 -> Water_Temp.:98.82, PWM:34
15:55:15.059 -> Water_Temp.:98.82, PWM:34
15:55:15.854 -> Water_Temp.:98.82, PWM:34
15:55:16.647 -> Water_Temp.:98.82, PWM:34
15:55:17.406 -> Water_Temp.:98.82, PWM:34
15:55:20.520 -> Water_Temp.:98.82, PWM:34
15:55:22.075 -> Water_Temp.:98.82, PWM:34
15:55:23.629 -> Water_Temp.:98.82, PWM:34
15:55:24.387 -> Water_Temp.:98.94, PWM:34
15:55:25.946 -> Water_Temp.:98.94, PWM:34
15:55:27.500 -> Water_Temp.:98.94, PWM:34
15:55:28.261 -> Water_Temp.:98.94, PWM:34
15:55:30.540 -> Water_Temp.:98.94, PWM:34
15:55:31.335 -> Water_Temp.:98.94, PWM:34
15:55:32.095 -> Water_Temp.:98.94, PWM:34
15:55:33.648 -> Water_Temp.:98.94, PWM:34
15:55:35.957 -> Water_Temp.:98.94, PWM:34
15:55:36.719 -> Water_Temp.:98.94, PWM:34
15:55:37.516 -> Water_Temp.:98.94, PWM:34
15:55:38.278 -> Water_Temp.:98.94, PWM:34
15:55:39.075 -> Water_Temp.:98.94, PWM:34
15:55:39.870 -> Water_Temp.:98.94, PWM:34
15:55:40.632 -> Water_Temp.:98.94, PWM:34
15:55:42.186 -> Water_Temp.:98.94, PWM:34
15:55:42.941 -> Water_Temp.:98.94, PWM:34
15:55:43.734 -> Water_Temp.:98.94, PWM:34
15:55:44.527 -> Water_Temp.:98.94, PWM:34
15:55:45.289 -> Water_Temp.:98.94, PWM:34
15:55:46.082 -> Water_Temp.:98.94, PWM:34
15:55:48.401 -> Water_Temp.:98.94, PWM:34
15:55:50.683 -> Water_Temp.:98.94, PWM:34
15:55:51.479 -> Water_Temp.:98.94, PWM:34
15:55:52.241 -> Water_Temp.:98.94, PWM:34
15:55:53.036 -> Water_Temp.:98.94, PWM:34
15:55:56.144 -> Water_Temp.:98.94, PWM:34
15:55:56.903 -> Water_Temp.:98.94, PWM:34
15:55:58.457 -> Water_Temp.:98.94, PWM:34
15:55:59.250 -> Water_Temp.:98.94, PWM:34
15:56:00.042 -> Water_Temp.:98.94, PWM:34
15:56:03.880 -> Water_Temp.:99.05, PWM:34
15:56:06.192 -> Water_Temp.:100.74, PWM:34
15:56:06.977 -> Water_Temp.:102.76, PWM:34
15:56:08.516 -> Water_Temp.:99.95, PWM:34
15:56:09.303 -> Water_Temp.:99.72, PWM:34
15:56:10.094 -> Water_Temp.:99.50, PWM:34
15:56:11.647 -> Water_Temp.:99.39, PWM:34
15:56:12.405 -> Water_Temp.:99.27, PWM:34
15:56:13.960 -> Water_Temp.:99.27, PWM:34
15:56:16.282 -> Water_Temp.:99.16, PWM:34
15:56:17.076 -> Water_Temp.:99.16, PWM:34
15:56:17.836 -> Water_Temp.:99.16, PWM:34
15:56:18.630 -> Water_Temp.:99.05, PWM:34
15:56:19.426 -> Water_Temp.:99.05, PWM:34
15:56:20.946 -> Water_Temp.:99.05, PWM:34
15:56:21.743 -> Water_Temp.:99.05, PWM:34
15:56:22.504 -> Water_Temp.:99.05, PWM:34
15:56:23.301 -> Water_Temp.:99.05, PWM:34
15:56:24.820 -> Water_Temp.:99.05, PWM:34
15:56:25.616 -> Water_Temp.:99.05, PWM:34
15:56:26.375 -> Water_Temp.:99.05, PWM:34
15:56:27.925 -> Water_Temp.:99.05, PWM:34
15:56:31.766 -> Water_Temp.:99.05, PWM:34
15:56:33.324 -> Water_Temp.:99.05, PWM:34
15:56:34.082 -> Water_Temp.:99.05, PWM:34
15:56:34.878 -> Water_Temp.:99.05, PWM:34
15:56:36.436 -> Water_Temp.:99.05, PWM:34
15:56:38.745 -> Water_Temp.:99.05, PWM:34
15:56:39.539 -> Water_Temp.:99.05, PWM:34
15:56:40.334 -> Water_Temp.:99.16, PWM:34
15:56:41.093 -> Water_Temp.:99.16, PWM:34
15:56:41.888 -> Water_Temp.:99.16, PWM:34
15:56:42.652 -> Water_Temp.:99.16, PWM:34
15:56:43.443 -> Water_Temp.:99.16, PWM:34
15:56:44.203 -> Water_Temp.:99.16, PWM:34
15:56:45.756 -> Water_Temp.:99.16, PWM:34
15:56:47.311 -> Water_Temp.:99.16, PWM:34
15:56:48.070 -> Water_Temp.:99.16, PWM:34
15:56:48.864 -> Water_Temp.:99.16, PWM:34
15:56:49.659 -> Water_Temp.:99.16, PWM:34
15:56:51.180 -> Water_Temp.:99.16, PWM:34
15:56:51.971 -> Water_Temp.:99.16, PWM:34
15:56:53.523 -> Water_Temp.:99.16, PWM:34
15:56:55.071 -> Water_Temp.:99.16, PWM:34
15:56:56.589 -> Water_Temp.:99.16, PWM:34
15:56:58.137 -> Water_Temp.:99.16, PWM:34
15:56:59.687 -> Water_Temp.:99.16, PWM:34
15:57:00.483 -> Water_Temp.:99.16, PWM:34
15:57:02.035 -> Water_Temp.:99.16, PWM:34
15:57:03.588 -> Water_Temp.:99.16, PWM:34
15:57:05.138 -> Water_Temp.:99.16, PWM:34
15:57:06.694 -> Water_Temp.:99.16, PWM:34
15:57:08.249 -> Water_Temp.:99.16, PWM:34
15:57:09.805 -> Water_Temp.:99.16, PWM:34
15:57:10.572 -> Water_Temp.:99.27, PWM:34
15:57:11.366 -> Water_Temp.:99.16, PWM:34
15:57:12.126 -> Water_Temp.:99.27, PWM:34
15:57:12.918 -> Water_Temp.:99.16, PWM:34
15:57:13.680 -> Water_Temp.:99.16, PWM:34
15:57:14.472 -> Water_Temp.:99.16, PWM:34
15:57:16.024 -> Water_Temp.:99.16, PWM:34
15:57:16.816 -> Water_Temp.:99.16, PWM:34
15:57:18.373 -> Water_Temp.:99.16, PWM:34
15:57:19.929 -> Water_Temp.:99.16, PWM:34
15:57:21.493 -> Water_Temp.:99.16, PWM:34
15:57:22.254 -> Water_Temp.:99.16, PWM:34
15:57:23.812 -> Water_Temp.:99.16, PWM:34
15:57:26.891 -> Water_Temp.:99.16, PWM:34
15:57:28.415 -> Water_Temp.:99.16, PWM:34
15:57:29.212 -> Water_Temp.:99.16, PWM:34
15:57:29.974 -> Water_Temp.:99.16, PWM:34
15:57:30.766 -> Water_Temp.:99.16, PWM:34
15:57:31.523 -> Water_Temp.:99.16, PWM:34
15:57:33.111 -> Water_Temp.:99.16, PWM:34
15:57:33.869 -> Water_Temp.:99.16, PWM:34
15:57:34.660 -> Water_Temp.:99.16, PWM:34
15:57:36.972 -> Water_Temp.:99.16, PWM:34
15:57:37.731 -> Water_Temp.:99.16, PWM:34
15:57:38.532 -> Water_Temp.:99.16, PWM:34
15:57:41.612 -> Water_Temp.:99.16, PWM:34
15:57:42.375 -> Water_Temp.:99.16, PWM:34
15:57:43.167 -> Water_Temp.:99.16, PWM:34
15:57:44.677 -> Water_Temp.:99.16, PWM:34
15:57:46.229 -> Water_Temp.:99.16, PWM:34
15:57:47.020 -> Water_Temp.:99.16, PWM:34
15:57:47.782 -> Water_Temp.:99.27, PWM:34
15:57:48.574 -> Water_Temp.:99.27, PWM:34
15:57:50.132 -> Water_Temp.:99.27, PWM:34
15:57:51.649 -> Water_Temp.:99.27, PWM:34
15:57:53.238 -> Water_Temp.:99.27, PWM:34
15:57:53.994 -> Water_Temp.:99.27, PWM:34
15:57:54.789 -> Water_Temp.:99.27, PWM:34
15:57:55.547 -> Water_Temp.:99.27, PWM:34
15:57:58.656 -> Water_Temp.:99.27, PWM:34
15:58:00.179 -> Water_Temp.:99.27, PWM:34
15:58:00.971 -> Water_Temp.:99.27, PWM:34
15:58:01.732 -> Water_Temp.:99.27, PWM:34
15:58:02.526 -> Water_Temp.:99.27, PWM:34
15:58:03.316 -> Water_Temp.:99.27, PWM:34
15:58:04.840 -> Water_Temp.:99.27, PWM:34
15:58:05.632 -> Water_Temp.:99.27, PWM:34
15:58:07.199 -> Water_Temp.:99.27, PWM:34
15:58:07.957 -> Water_Temp.:99.27, PWM:34
15:58:08.748 -> Water_Temp.:99.27, PWM:34
15:58:09.508 -> Water_Temp.:99.27, PWM:34
15:58:10.298 -> Water_Temp.:99.27, PWM:34
15:58:12.646 -> Water_Temp.:99.27, PWM:34
15:58:13.407 -> Water_Temp.:99.27, PWM:34
15:58:14.202 -> Water_Temp.:99.27, PWM:34
15:58:15.754 -> Water_Temp.:100.96, PWM:34
15:58:16.549 -> Water_Temp.:103.10, PWM:34
15:58:18.099 -> Water_Temp.:100.17, PWM:34
15:58:18.890 -> Water_Temp.:100.06, PWM:34
15:58:20.443 -> Water_Temp.:99.84, PWM:34
15:58:21.231 -> Water_Temp.:99.72, PWM:34
15:58:21.990 -> Water_Temp.:99.61, PWM:34
15:58:23.543 -> Water_Temp.:99.61, PWM:34
15:58:26.650 -> Water_Temp.:99.50, PWM:34
15:58:27.413 -> Water_Temp.:99.50, PWM:34
15:58:28.206 -> Water_Temp.:99.50, PWM:34
15:58:29.758 -> Water_Temp.:99.39, PWM:34
15:58:31.313 -> Water_Temp.:99.39, PWM:34
15:58:32.864 -> Water_Temp.:99.39, PWM:34
15:58:33.662 -> Water_Temp.:99.39, PWM:34
15:58:35.977 -> Water_Temp.:99.39, PWM:34
15:58:36.736 -> Water_Temp.:99.39, PWM:34
15:58:37.529 -> Water_Temp.:99.39, PWM:34
15:58:38.326 -> Water_Temp.:99.39, PWM:34
15:58:39.087 -> Water_Temp.:99.39, PWM:34
15:58:40.610 -> Water_Temp.:99.39, PWM:34
15:58:43.725 -> Water_Temp.:99.27, PWM:34
15:58:44.487 -> Water_Temp.:99.27, PWM:34
15:58:45.278 -> Water_Temp.:99.27, PWM:34
15:58:47.595 -> Water_Temp.:99.39, PWM:34
15:58:48.358 -> Water_Temp.:99.27, PWM:34
15:58:49.151 -> Water_Temp.:99.27, PWM:34
15:58:49.942 -> Water_Temp.:99.27, PWM:34
15:58:51.462 -> Water_Temp.:99.27, PWM:34
15:58:52.222 -> Water_Temp.:99.27, PWM:34
15:58:54.538 -> Water_Temp.:99.27, PWM:34
15:58:55.332 -> Water_Temp.:99.27, PWM:34
15:58:56.089 -> Water_Temp.:99.27, PWM:34
15:58:59.169 -> Water_Temp.:99.27, PWM:34
15:59:00.725 -> Water_Temp.:99.27, PWM:34
15:59:01.486 -> Water_Temp.:99.27, PWM:34
15:59:02.281 -> Water_Temp.:99.27, PWM:34
15:59:03.044 -> Water_Temp.:99.27, PWM:34
15:59:03.840 -> Water_Temp.:99.27, PWM:34
15:59:05.360 -> Water_Temp.:99.27, PWM:34
15:59:06.156 -> Water_Temp.:99.27, PWM:34
15:59:06.951 -> Water_Temp.:99.27, PWM:34
15:59:07.714 -> Water_Temp.:99.27, PWM:34
15:59:09.239 -> Water_Temp.:99.27, PWM:34
15:59:10.031 -> Water_Temp.:99.27, PWM:34
15:59:11.581 -> Water_Temp.:99.27, PWM:34
15:59:12.371 -> Water_Temp.:99.27, PWM:34
15:59:13.131 -> Water_Temp.:99.27, PWM:34
15:59:14.685 -> Water_Temp.:99.27, PWM:34
15:59:15.477 -> Water_Temp.:99.27, PWM:34
15:59:16.242 -> Water_Temp.:99.27, PWM:34
15:59:17.803 -> Water_Temp.:99.27, PWM:34
15:59:19.360 -> Water_Temp.:99.27, PWM:34
15:59:21.678 -> Water_Temp.:99.27, PWM:34
15:59:22.470 -> Water_Temp.:99.27, PWM:34
15:59:23.236 -> Water_Temp.:99.27, PWM:34
15:59:24.029 -> Water_Temp.:99.27, PWM:34
15:59:25.548 -> Water_Temp.:99.27, PWM:34
15:59:26.350 -> Water_Temp.:99.27, PWM:34
15:59:27.900 -> Water_Temp.:99.27, PWM:34
15:59:29.416 -> Water_Temp.:99.27, PWM:34
15:59:30.206 -> Water_Temp.:99.27, PWM:34
15:59:31.000 -> Water_Temp.:99.27, PWM:34
15:59:31.760 -> Water_Temp.:99.27, PWM:34
15:59:32.553 -> Water_Temp.:99.27, PWM:34
15:59:33.346 -> Water_Temp.:99.27, PWM:34
15:59:35.660 -> Water_Temp.:99.27, PWM:34
15:59:36.421 -> Water_Temp.:99.27, PWM:34
15:59:37.218 -> Water_Temp.:99.27, PWM:34
15:59:37.979 -> Water_Temp.:99.27, PWM:34
15:59:38.775 -> Water_Temp.:99.27, PWM:34
15:59:39.537 -> Water_Temp.:99.27, PWM:34
15:59:41.092 -> Water_Temp.:99.27, PWM:34
15:59:42.676 -> Water_Temp.:99.27, PWM:34
15:59:44.990 -> Water_Temp.:99.27, PWM:34
15:59:45.751 -> Water_Temp.:99.27, PWM:34
15:59:46.551 -> Water_Temp.:99.27, PWM:34
15:59:47.313 -> Water_Temp.:99.27, PWM:34
15:59:48.104 -> Water_Temp.:99.27, PWM:34
15:59:48.898 -> Water_Temp.:99.27, PWM:34
15:59:49.661 -> Water_Temp.:99.27, PWM:34
15:59:50.457 -> Water_Temp.:99.27, PWM:34
15:59:51.218 -> Water_Temp.:99.27, PWM:34
15:59:52.013 -> Water_Temp.:99.27, PWM:34
15:59:53.564 -> Water_Temp.:99.27, PWM:34
15:59:54.357 -> Water_Temp.:99.27, PWM:34
15:59:55.912 -> Water_Temp.:99.27, PWM:34
15:59:56.675 -> Water_Temp.:99.27, PWM:34
15:59:58.987 -> Water_Temp.:99.27, PWM:34
16:00:00.542 -> Water_Temp.:99.16, PWM:34
16:00:01.334 -> Water_Temp.:99.27, PWM:34
16:00:02.095 -> Water_Temp.:99.27, PWM:34
16:00:02.891 -> Water_Temp.:99.16, PWM:34
16:00:04.449 -> Water_Temp.:99.16, PWM:34
16:00:05.245 -> Water_Temp.:99.16, PWM:34
16:00:06.006 -> Water_Temp.:99.16, PWM:34
16:00:06.798 -> Water_Temp.:99.16, PWM:34
16:00:07.559 -> Water_Temp.:99.16, PWM:34
16:00:08.352 -> Water_Temp.:99.16, PWM:34
16:00:09.145 -> Water_Temp.:100.85, PWM:34
16:00:09.908 -> Water_Temp.:102.99, PWM:34
16:00:11.463 -> Water_Temp.:100.06, PWM:34
16:00:12.253 -> Water_Temp.:99.84, PWM:34
16:00:13.807 -> Water_Temp.:99.61, PWM:34
16:00:15.329 -> Water_Temp.:99.50, PWM:34
16:00:16.125 -> Water_Temp.:99.50, PWM:34
16:00:16.888 -> Water_Temp.:99.39, PWM:34
16:00:17.684 -> Water_Temp.:99.39, PWM:34
16:00:18.444 -> Water_Temp.:99.39, PWM:34
16:00:19.241 -> Water_Temp.:99.39, PWM:34
16:00:20.001 -> Water_Temp.:99.27, PWM:34
16:00:20.797 -> Water_Temp.:99.27, PWM:34
16:00:21.557 -> Water_Temp.:99.27, PWM:34
16:00:23.144 -> Water_Temp.:99.27, PWM:34
16:00:23.902 -> Water_Temp.:99.27, PWM:34
16:00:24.694 -> Water_Temp.:99.27, PWM:34
16:00:26.252 -> Water_Temp.:99.16, PWM:34
16:00:27.047 -> Water_Temp.:99.16, PWM:34
16:00:27.805 -> Water_Temp.:99.16, PWM:34
16:00:28.598 -> Water_Temp.:99.16, PWM:34
16:00:29.355 -> Water_Temp.:99.16, PWM:34
16:00:30.149 -> Water_Temp.:99.16, PWM:34
16:00:30.946 -> Water_Temp.:99.16, PWM:34
16:00:34.049 -> Water_Temp.:99.16, PWM:34
16:00:34.809 -> Water_Temp.:99.16, PWM:34
16:00:35.604 -> Water_Temp.:99.16, PWM:34
16:00:36.368 -> Water_Temp.:99.16, PWM:34
16:00:37.163 -> Water_Temp.:99.16, PWM:34
16:00:37.956 -> Water_Temp.:99.16, PWM:34
16:00:38.715 -> Water_Temp.:99.16, PWM:34
16:00:40.272 -> Water_Temp.:99.16, PWM:34
16:00:41.066 -> Water_Temp.:99.16, PWM:34
16:00:43.382 -> Water_Temp.:99.16, PWM:34
16:00:44.171 -> Water_Temp.:99.16, PWM:34
16:00:45.687 -> Water_Temp.:99.16, PWM:34
16:00:46.484 -> Water_Temp.:99.16, PWM:34
16:00:48.000 -> Water_Temp.:99.16, PWM:34
16:00:48.794 -> Water_Temp.:99.16, PWM:34
16:00:49.590 -> Water_Temp.:99.16, PWM:34
16:00:50.350 -> Water_Temp.:99.16, PWM:34
16:00:51.910 -> Water_Temp.:99.16, PWM:34
16:00:52.667 -> Water_Temp.:99.16, PWM:34
16:00:53.460 -> Water_Temp.:99.16, PWM:34
16:00:54.255 -> Water_Temp.:99.16, PWM:34
16:00:55.014 -> Water_Temp.:99.16, PWM:34
16:00:55.807 -> Water_Temp.:99.16, PWM:34
16:00:57.326 -> Water_Temp.:99.16, PWM:34
16:00:58.912 -> Water_Temp.:99.16, PWM:34
16:00:59.674 -> Water_Temp.:99.16, PWM:34
16:01:02.777 -> Water_Temp.:99.16, PWM:34
16:01:03.540 -> Water_Temp.:99.16, PWM:34
16:01:04.332 -> Water_Temp.:99.16, PWM:34
16:01:05.095 -> Water_Temp.:99.16, PWM:34
16:01:05.889 -> Water_Temp.:99.16, PWM:34
16:01:07.444 -> Water_Temp.:99.16, PWM:33
16:01:08.242 -> Water_Temp.:99.16, PWM:34
16:01:09.799 -> Water_Temp.:99.16, PWM:34
16:01:10.566 -> Water_Temp.:99.16, PWM:34
16:01:11.363 -> Water_Temp.:99.16, PWM:34
16:01:12.122 -> Water_Temp.:99.16, PWM:34
16:01:12.916 -> Water_Temp.:99.16, PWM:34
16:01:13.675 -> Water_Temp.:99.16, PWM:34
16:01:14.465 -> Water_Temp.:99.16, PWM:34
16:01:16.023 -> Water_Temp.:99.16, PWM:34
16:01:19.136 -> Water_Temp.:99.16, PWM:34
16:01:19.894 -> Water_Temp.:99.16, PWM:34
16:01:22.210 -> Water_Temp.:99.16, PWM:34
16:01:24.521 -> Water_Temp.:99.16, PWM:34
16:01:25.320 -> Water_Temp.:99.16, PWM:34
16:01:26.082 -> Water_Temp.:99.16, PWM:34
16:01:26.878 -> Water_Temp.:99.16, PWM:34
16:01:27.636 -> Water_Temp.:99.16, PWM:34
16:01:28.430 -> Water_Temp.:99.16, PWM:34
16:01:29.982 -> Water_Temp.:99.16, PWM:34
16:01:31.537 -> Water_Temp.:99.16, PWM:34
16:01:33.864 -> Water_Temp.:99.16, PWM:34
16:01:35.420 -> Water_Temp.:99.16, PWM:34
16:01:36.980 -> Water_Temp.:99.16, PWM:34
16:01:37.742 -> Water_Temp.:99.16, PWM:34
16:01:39.295 -> Water_Temp.:99.16, PWM:34
16:01:40.848 -> Water_Temp.:99.16, PWM:34
16:01:41.611 -> Water_Temp.:99.16, PWM:34
16:01:42.405 -> Water_Temp.:99.16, PWM:34
16:01:43.166 -> Water_Temp.:99.16, PWM:34
16:01:44.724 -> Water_Temp.:99.16, PWM:34
16:01:45.483 -> Water_Temp.:99.16, PWM:34
16:01:47.038 -> Water_Temp.:99.16, PWM:34
16:01:47.829 -> Water_Temp.:99.16, PWM:34
16:01:50.935 -> Water_Temp.:99.16, PWM:34
16:01:51.696 -> Water_Temp.:99.16, PWM:34
16:01:52.490 -> Water_Temp.:99.16, PWM:34
16:01:53.252 -> Water_Temp.:99.16, PWM:34
16:01:55.533 -> Water_Temp.:99.16, PWM:34
16:01:56.332 -> Water_Temp.:99.16, PWM:34
16:01:57.122 -> Water_Temp.:99.16, PWM:34
16:01:57.884 -> Water_Temp.:99.16, PWM:34
16:01:58.679 -> Water_Temp.:99.16, PWM:34
16:02:00.196 -> Water_Temp.:99.16, PWM:34
16:02:03.308 -> Water_Temp.:99.16, PWM:34
16:02:04.864 -> Water_Temp.:99.16, PWM:34
16:02:05.661 -> Water_Temp.:99.16, PWM:34
16:02:08.733 -> Water_Temp.:99.16, PWM:34
16:02:09.531 -> Water_Temp.:99.16, PWM:34
16:02:11.838 -> Water_Temp.:99.16, PWM:34
16:02:12.630 -> Water_Temp.:99.16, PWM:34
16:02:14.176 -> Water_Temp.:99.16, PWM:34
16:02:15.697 -> Water_Temp.:99.16, PWM:34
16:02:17.283 -> Water_Temp.:99.16, PWM:34
16:02:18.044 -> Water_Temp.:99.16, PWM:34
16:02:19.569 -> Water_Temp.:99.16, PWM:34
16:02:20.363 -> Water_Temp.:99.16, PWM:34
16:02:21.158 -> Water_Temp.:99.16, PWM:34
16:02:21.916 -> Water_Temp.:99.16, PWM:34
16:02:23.473 -> Water_Temp.:99.16, PWM:34
16:02:24.228 -> Water_Temp.:99.16, PWM:34
16:02:25.024 -> Water_Temp.:99.16, PWM:34
16:02:25.788 -> Water_Temp.:99.16, PWM:34
16:02:26.582 -> Water_Temp.:99.16, PWM:34
16:02:27.370 -> Water_Temp.:99.16, PWM:34
16:02:28.890 -> Water_Temp.:99.16, PWM:34
16:02:29.680 -> Water_Temp.:99.16, PWM:34
16:02:35.866 -> Water_Temp.:99.16, PWM:34
16:02:36.626 -> Water_Temp.:99.16, PWM:34
16:02:37.420 -> Water_Temp.:99.16, PWM:34
16:02:38.947 -> Water_Temp.:99.16, PWM:34
16:02:39.744 -> Water_Temp.:99.16, PWM:34
16:02:40.507 -> Water_Temp.:99.27, PWM:34
16:02:42.060 -> Water_Temp.:99.27, PWM:34
16:02:42.856 -> Water_Temp.:99.27, PWM:34
16:02:43.620 -> Water_Temp.:99.27, PWM:34
16:02:44.412 -> Water_Temp.:99.27, PWM:34
16:02:45.173 -> Water_Temp.:99.27, PWM:34
16:02:46.725 -> Water_Temp.:99.27, PWM:34
16:02:47.519 -> Water_Temp.:99.27, PWM:34
16:02:48.310 -> Water_Temp.:99.27, PWM:34
16:02:49.069 -> Water_Temp.:99.27, PWM:34
16:02:49.866 -> Water_Temp.:99.27, PWM:34
16:02:50.627 -> Water_Temp.:99.27, PWM:34
16:02:51.421 -> Water_Temp.:99.27, PWM:34
16:02:53.736 -> Water_Temp.:99.27, PWM:34
16:02:54.527 -> Water_Temp.:99.27, PWM:34
16:02:56.840 -> Water_Temp.:99.27, PWM:34
16:02:58.398 -> Water_Temp.:99.27, PWM:34
16:02:59.158 -> Water_Temp.:99.27, PWM:34
16:03:00.718 -> Water_Temp.:99.27, PWM:34
16:03:02.261 -> Water_Temp.:99.27, PWM:34
16:03:03.016 -> Water_Temp.:99.27, PWM:34
16:03:03.808 -> Water_Temp.:99.27, PWM:34
16:03:04.596 -> Water_Temp.:99.27, PWM:34
16:03:06.116 -> Water_Temp.:99.27, PWM:34
16:03:06.906 -> Water_Temp.:99.27, PWM:34
16:03:07.699 -> Water_Temp.:99.27, PWM:34
16:03:09.250 -> Water_Temp.:99.27, PWM:34
16:03:10.035 -> Water_Temp.:99.27, PWM:34
16:03:11.588 -> Water_Temp.:99.27, PWM:34
16:03:13.139 -> Water_Temp.:99.27, PWM:34
16:03:14.680 -> Water_Temp.:99.27, PWM:34
16:03:15.438 -> Water_Temp.:99.27, PWM:34
16:03:17.020 -> Water_Temp.:99.27, PWM:34
16:03:18.564 -> Water_Temp.:99.27, PWM:34
16:03:23.930 -> Water_Temp.:99.27, PWM:34
16:03:24.721 -> Water_Temp.:99.27, PWM:34
16:03:26.269 -> Water_Temp.:99.27, PWM:34
16:03:27.063 -> Water_Temp.:99.27, PWM:34
16:03:27.853 -> Water_Temp.:99.39, PWM:34
16:03:30.937 -> Water_Temp.:99.39, PWM:34
16:03:32.484 -> Water_Temp.:99.39, PWM:34
16:03:33.274 -> Water_Temp.:99.39, PWM:34
16:03:34.062 -> Water_Temp.:99.39, PWM:34
16:03:34.850 -> Water_Temp.:99.39, PWM:34
16:03:35.605 -> Water_Temp.:99.39, PWM:34
16:03:37.187 -> Water_Temp.:99.39, PWM:34
16:03:37.942 -> Water_Temp.:99.39, PWM:34
16:03:38.731 -> Water_Temp.:99.39, PWM:34
16:03:41.820 -> Water_Temp.:99.39, PWM:34
16:03:42.612 -> Water_Temp.:99.39, PWM:34
16:03:43.404 -> Water_Temp.:99.39, PWM:34
16:03:44.161 -> Water_Temp.:99.39, PWM:34
16:03:44.959 -> Water_Temp.:99.39, PWM:34
16:03:45.746 -> Water_Temp.:99.39, PWM:34
16:03:46.500 -> Water_Temp.:99.39, PWM:34
16:03:48.808 -> Water_Temp.:99.39, PWM:34
16:03:50.365 -> Water_Temp.:99.27, PWM:34
16:03:51.158 -> Water_Temp.:99.39, PWM:34
16:03:51.913 -> Water_Temp.:99.39, PWM:34
16:03:52.702 -> Water_Temp.:99.27, PWM:34
16:03:53.495 -> Water_Temp.:99.27, PWM:34
16:03:54.253 -> Water_Temp.:99.39, PWM:34
16:03:55.039 -> Water_Temp.:99.27, PWM:34
16:03:56.594 -> Water_Temp.:99.27, PWM:34
16:03:57.388 -> Water_Temp.:99.27, PWM:34
16:03:58.178 -> Water_Temp.:99.27, PWM:34
16:03:59.700 -> Water_Temp.:99.27, PWM:34
16:04:01.256 -> Water_Temp.:99.27, PWM:34
16:04:03.569 -> Water_Temp.:99.27, PWM:34
16:04:05.877 -> Water_Temp.:99.27, PWM:34
16:04:07.432 -> Water_Temp.:99.27, PWM:34
16:04:08.190 -> Water_Temp.:99.27, PWM:34
16:04:08.984 -> Water_Temp.:99.27, PWM:34
16:04:09.775 -> Water_Temp.:99.27, PWM:34
16:04:11.328 -> Water_Temp.:99.27, PWM:34
16:04:12.842 -> Water_Temp.:99.27, PWM:34
16:04:15.957 -> Water_Temp.:99.27, PWM:34
16:04:16.716 -> Water_Temp.:99.27, PWM:34
16:04:18.301 -> Water_Temp.:99.27, PWM:34
16:04:19.063 -> Water_Temp.:99.27, PWM:34
16:04:20.614 -> Water_Temp.:99.27, PWM:34
16:04:21.376 -> Water_Temp.:99.27, PWM:34
16:04:22.171 -> Water_Temp.:99.27, PWM:34
16:04:22.933 -> Water_Temp.:99.27, PWM:34
16:04:24.487 -> Water_Temp.:99.27, PWM:34
16:04:25.277 -> Water_Temp.:99.27, PWM:34
16:04:26.039 -> Water_Temp.:99.27, PWM:34
16:04:26.833 -> Water_Temp.:99.27, PWM:34
16:04:29.110 -> Water_Temp.:99.27, PWM:34
16:04:29.871 -> Water_Temp.:99.27, PWM:34
16:04:31.429 -> Water_Temp.:99.27, PWM:34
16:04:32.224 -> Water_Temp.:99.27, PWM:34
16:04:33.743 -> Water_Temp.:99.27, PWM:34
16:04:35.295 -> Water_Temp.:99.27, PWM:39
16:04:36.091 -> Water_Temp.:99.27, PWM:39
16:04:36.850 -> Water_Temp.:99.27, PWM:39
16:04:38.441 -> Water_Temp.:99.27, PWM:39
16:04:39.198 -> Water_Temp.:99.27, PWM:37
16:04:39.993 -> Water_Temp.:99.27, PWM:37
16:04:40.752 -> Water_Temp.:99.27, PWM:36
16:04:41.549 -> Water_Temp.:99.39, PWM:37
16:04:42.308 -> Water_Temp.:99.27, PWM:37
16:04:43.104 -> Water_Temp.:99.27, PWM:37
16:04:46.219 -> Water_Temp.:99.27, PWM:37
16:04:46.983 -> Water_Temp.:99.27, PWM:35
16:04:48.541 -> Water_Temp.:99.27, PWM:35
16:04:49.337 -> Water_Temp.:99.27, PWM:35
16:04:50.132 -> Water_Temp.:99.27, PWM:35
16:04:51.649 -> Water_Temp.:99.27, PWM:35
16:04:53.207 -> Water_Temp.:99.27, PWM:36
16:04:54.002 -> Water_Temp.:99.27, PWM:36
16:04:54.760 -> Water_Temp.:99.27, PWM:36
16:04:55.551 -> Water_Temp.:99.27, PWM:36
16:04:57.863 -> Water_Temp.:99.27, PWM:36
16:04:58.657 -> Water_Temp.:99.27, PWM:36
16:04:59.416 -> Water_Temp.:99.27, PWM:36
16:05:00.211 -> Water_Temp.:99.27, PWM:36
16:05:00.972 -> Water_Temp.:99.27, PWM:36
16:05:01.763 -> Water_Temp.:99.27, PWM:36
16:05:02.524 -> Water_Temp.:99.27, PWM:36
16:05:03.319 -> Water_Temp.:99.27, PWM:36
16:05:04.112 -> Water_Temp.:99.27, PWM:36
16:05:04.871 -> Water_Temp.:99.27, PWM:36
16:05:05.665 -> Water_Temp.:99.27, PWM:36
16:05:06.429 -> Water_Temp.:99.27, PWM:36
16:05:07.224 -> Water_Temp.:99.39, PWM:36
16:05:07.986 -> Water_Temp.:99.39, PWM:36
16:05:09.538 -> Water_Temp.:99.39, PWM:36
16:05:11.096 -> Water_Temp.:99.39, PWM:36
16:05:11.887 -> Water_Temp.:99.39, PWM:36
16:05:13.405 -> Water_Temp.:99.39, PWM:36
16:05:14.168 -> Water_Temp.:99.39, PWM:36
16:05:14.960 -> Water_Temp.:99.39, PWM:36
16:05:15.754 -> Water_Temp.:99.39, PWM:36
16:05:16.515 -> Water_Temp.:99.39, PWM:36
16:05:20.384 -> Water_Temp.:99.39, PWM:36
16:05:21.934 -> Water_Temp.:99.39, PWM:36
16:05:22.727 -> Water_Temp.:99.39, PWM:36
16:05:25.041 -> Water_Temp.:99.39, PWM:36
16:05:25.803 -> Water_Temp.:99.39, PWM:36
16:05:26.600 -> Water_Temp.:99.39, PWM:36
16:05:28.913 -> Water_Temp.:99.39, PWM:36
16:05:29.709 -> Water_Temp.:99.39, PWM:36
16:05:31.259 -> Water_Temp.:99.39, PWM:36
16:05:32.017 -> Water_Temp.:99.39, PWM:36
16:05:32.806 -> Water_Temp.:99.39, PWM:36
16:05:33.603 -> Water_Temp.:99.39, PWM:36
16:05:34.365 -> Water_Temp.:99.39, PWM:35
16:05:35.158 -> Water_Temp.:99.39, PWM:36
16:05:35.922 -> Water_Temp.:99.39, PWM:36
16:05:37.472 -> Water_Temp.:99.39, PWM:36
16:05:41.307 -> Water_Temp.:99.39, PWM:36
16:05:42.862 -> Water_Temp.:99.39, PWM:36
16:05:43.661 -> Water_Temp.:99.39, PWM:36
16:05:44.420 -> Water_Temp.:99.39, PWM:36
16:05:46.004 -> Water_Temp.:99.39, PWM:36
16:05:47.525 -> Water_Temp.:99.39, PWM:36
16:05:49.869 -> Water_Temp.:99.50, PWM:36
16:05:50.630 -> Water_Temp.:99.50, PWM:36
16:05:51.425 -> Water_Temp.:99.39, PWM:36
16:05:53.773 -> Water_Temp.:99.39, PWM:36
16:05:54.535 -> Water_Temp.:99.39, PWM:36
16:05:55.323 -> Water_Temp.:99.39, PWM:36
16:05:56.848 -> Water_Temp.:99.39, PWM:36
16:05:59.922 -> Water_Temp.:99.39, PWM:36
16:06:00.684 -> Water_Temp.:99.39, PWM:36
16:06:02.238 -> Water_Temp.:99.39, PWM:36
16:06:04.553 -> Water_Temp.:99.39, PWM:36
16:06:06.103 -> Water_Temp.:99.39, PWM:36
16:06:06.867 -> Water_Temp.:99.39, PWM:36
16:06:08.419 -> Water_Temp.:99.39, PWM:35
16:06:09.211 -> Water_Temp.:99.39, PWM:36
16:06:10.768 -> Water_Temp.:99.39, PWM:35
16:06:12.323 -> Water_Temp.:99.39, PWM:36
16:06:13.081 -> Water_Temp.:99.50, PWM:36
16:06:13.872 -> Water_Temp.:99.39, PWM:36
16:06:16.953 -> Water_Temp.:99.50, PWM:36
16:06:21.622 -> Water_Temp.:99.50, PWM:36
16:06:22.385 -> Water_Temp.:99.50, PWM:36
16:06:23.181 -> Water_Temp.:99.50, PWM:36
16:06:26.252 -> Water_Temp.:99.50, PWM:36
16:06:27.838 -> Water_Temp.:99.50, PWM:36
16:06:29.397 -> Water_Temp.:99.50, PWM:36
16:06:30.155 -> Water_Temp.:99.50, PWM:36
16:06:30.950 -> Water_Temp.:99.50, PWM:36
16:06:33.260 -> Water_Temp.:99.50, PWM:36
16:06:34.052 -> Water_Temp.:99.50, PWM:36
16:06:34.815 -> Water_Temp.:99.50, PWM:36
16:06:35.605 -> Water_Temp.:99.50, PWM:36
16:06:37.153 -> Water_Temp.:99.50, PWM:36
16:06:38.704 -> Water_Temp.:99.50, PWM:36
16:06:39.464 -> Water_Temp.:99.50, PWM:36
16:06:41.016 -> Water_Temp.:99.50, PWM:36
16:06:42.565 -> Water_Temp.:99.50, PWM:36
16:06:44.840 -> Water_Temp.:99.50, PWM:36
16:06:47.905 -> Water_Temp.:99.50, PWM:36
16:06:49.454 -> Water_Temp.:99.50, PWM:36
16:06:50.246 -> Water_Temp.:99.50, PWM:36
16:06:51.005 -> Water_Temp.:99.50, PWM:36
16:06:52.584 -> Water_Temp.:99.50, PWM:36
16:06:53.375 -> Water_Temp.:99.50, PWM:36
16:06:54.898 -> Water_Temp.:99.50, PWM:36
16:06:55.687 -> Water_Temp.:99.50, PWM:36
16:06:56.440 -> Water_Temp.:99.50, PWM:36
16:06:57.234 -> Water_Temp.:99.50, PWM:36
16:06:58.789 -> Water_Temp.:99.50, PWM:36
16:07:00.304 -> Water_Temp.:99.50, PWM:36
16:07:01.097 -> Water_Temp.:99.50, PWM:36
16:07:02.645 -> Water_Temp.:99.50, PWM:36
16:07:04.226 -> Water_Temp.:99.50, PWM:36
16:07:05.745 -> Water_Temp.:99.50, PWM:36
16:07:06.542 -> Water_Temp.:99.50, PWM:36
16:07:08.095 -> Water_Temp.:99.50, PWM:36
16:07:08.854 -> Water_Temp.:99.50, PWM:36
16:07:09.648 -> Water_Temp.:99.50, PWM:36
16:07:10.409 -> Water_Temp.:99.50, PWM:36
16:07:13.506 -> Water_Temp.:99.50, PWM:36
16:07:15.859 -> Water_Temp.:99.50, PWM:36
16:07:16.620 -> Water_Temp.:99.50, PWM:36
16:07:17.416 -> Water_Temp.:99.50, PWM:36
16:07:18.175 -> Water_Temp.:99.50, PWM:36
16:07:18.970 -> Water_Temp.:99.50, PWM:36
16:07:22.080 -> Water_Temp.:99.50, PWM:36
16:07:23.634 -> Water_Temp.:99.61, PWM:36
16:07:24.425 -> Water_Temp.:99.61, PWM:36
16:07:25.182 -> Water_Temp.:99.61, PWM:36
16:07:25.975 -> Water_Temp.:99.61, PWM:36
16:07:27.523 -> Water_Temp.:99.61, PWM:36
16:07:28.277 -> Water_Temp.:99.61, PWM:36
16:07:29.866 -> Water_Temp.:99.61, PWM:36
16:07:30.626 -> Water_Temp.:99.61, PWM:36
16:07:31.418 -> Water_Temp.:99.61, PWM:36
16:07:32.172 -> Water_Temp.:99.61, PWM:36
16:07:33.754 -> Water_Temp.:99.61, PWM:36
16:07:34.514 -> Water_Temp.:99.61, PWM:36
16:07:35.310 -> Water_Temp.:99.61, PWM:36
16:07:36.871 -> Water_Temp.:99.61, PWM:36
16:07:37.634 -> Water_Temp.:99.61, PWM:36
16:07:38.429 -> Water_Temp.:99.61, PWM:36
16:07:39.976 -> Water_Temp.:99.61, PWM:36
16:07:40.738 -> Water_Temp.:99.61, PWM:36
16:07:42.286 -> Water_Temp.:99.61, PWM:36
16:07:43.047 -> Water_Temp.:99.61, PWM:36
16:07:45.360 -> Water_Temp.:99.61, PWM:36
16:07:46.154 -> Water_Temp.:99.61, PWM:36
16:07:47.709 -> Water_Temp.:99.61, PWM:36
16:07:49.252 -> Water_Temp.:99.61, PWM:36
16:07:50.048 -> Water_Temp.:99.72, PWM:36
16:07:51.569 -> Water_Temp.:99.72, PWM:36
16:07:52.359 -> Water_Temp.:99.72, PWM:36
16:07:55.395 -> Water_Temp.:99.72, PWM:36
16:07:56.187 -> Water_Temp.:99.72, PWM:36
16:07:56.985 -> Water_Temp.:99.72, PWM:36
16:07:57.745 -> Water_Temp.:99.72, PWM:36
16:07:58.542 -> Water_Temp.:99.72, PWM:36
16:07:59.303 -> Water_Temp.:99.72, PWM:36
16:08:00.856 -> Water_Temp.:99.72, PWM:36
16:08:01.651 -> Water_Temp.:99.72, PWM:36
16:08:02.448 -> Water_Temp.:99.72, PWM:36
16:08:03.967 -> Water_Temp.:99.72, PWM:36
16:08:05.515 -> Water_Temp.:99.72, PWM:36
16:08:06.312 -> Water_Temp.:99.72, PWM:36
16:08:07.074 -> Water_Temp.:99.72, PWM:36
16:08:08.628 -> Water_Temp.:99.72, PWM:36
16:08:09.424 -> Water_Temp.:99.72, PWM:36
16:08:13.292 -> Water_Temp.:99.72, PWM:36
16:08:14.844 -> Water_Temp.:99.72, PWM:36
16:08:15.639 -> Water_Temp.:99.72, PWM:36
16:08:17.161 -> Water_Temp.:99.72, PWM:36
16:08:17.950 -> Water_Temp.:99.72, PWM:36
16:08:20.273 -> Water_Temp.:99.72, PWM:36
16:08:21.032 -> Water_Temp.:99.72, PWM:36
16:08:21.829 -> Water_Temp.:99.72, PWM:36
16:08:22.619 -> Water_Temp.:99.72, PWM:36
16:08:23.378 -> Water_Temp.:99.72, PWM:36
16:08:25.727 -> Water_Temp.:99.84, PWM:36
16:08:26.488 -> Water_Temp.:99.84, PWM:36
16:08:28.806 -> Water_Temp.:99.84, PWM:36
16:08:29.568 -> Water_Temp.:99.84, PWM:36
16:08:30.361 -> Water_Temp.:99.84, PWM:36
16:08:31.151 -> Water_Temp.:99.84, PWM:36
16:08:32.665 -> Water_Temp.:99.84, PWM:36
16:08:33.457 -> Water_Temp.:99.84, PWM:36
16:08:34.248 -> Water_Temp.:99.84, PWM:36
16:08:35.764 -> Water_Temp.:99.84, PWM:36
16:08:36.555 -> Water_Temp.:99.84, PWM:36
16:08:37.353 -> Water_Temp.:99.84, PWM:36
16:08:40.398 -> Water_Temp.:99.84, PWM:36
16:08:41.194 -> Water_Temp.:99.84, PWM:36
16:08:41.954 -> Water_Temp.:99.84, PWM:36
16:08:42.748 -> Water_Temp.:99.84, PWM:36
16:08:44.268 -> Water_Temp.:99.84, PWM:36
16:08:45.822 -> Water_Temp.:99.84, PWM:36
16:08:47.371 -> Water_Temp.:99.72, PWM:36
16:08:48.128 -> Water_Temp.:99.72, PWM:36
16:08:48.922 -> Water_Temp.:99.72, PWM:36
16:08:49.715 -> Water_Temp.:99.72, PWM:36
16:08:52.031 -> Water_Temp.:99.72, PWM:36
16:08:53.584 -> Water_Temp.:99.72, PWM:36
16:08:55.134 -> Water_Temp.:99.72, PWM:36
16:08:56.693 -> Water_Temp.:99.72, PWM:36
16:08:59.001 -> Water_Temp.:99.84, PWM:36
16:08:59.797 -> Water_Temp.:99.84, PWM:36
16:09:00.551 -> Water_Temp.:99.84, PWM:36
16:09:01.341 -> Water_Temp.:99.84, PWM:36
16:09:02.136 -> Water_Temp.:99.84, PWM:36
16:09:03.655 -> Water_Temp.:99.84, PWM:36
16:09:05.967 -> Water_Temp.:99.84, PWM:36
16:09:07.512 -> Water_Temp.:99.84, PWM:36
16:09:08.267 -> Water_Temp.:99.84, PWM:36
16:09:09.848 -> Water_Temp.:99.95, PWM:36
16:09:10.610 -> Water_Temp.:99.84, PWM:36
16:09:11.405 -> Water_Temp.:99.95, PWM:36
16:09:12.166 -> Water_Temp.:99.95, PWM:36
16:09:13.750 -> Water_Temp.:99.95, PWM:36
16:09:16.061 -> Water_Temp.:99.95, PWM:36
16:09:17.619 -> Water_Temp.:99.95, PWM:36
16:09:18.380 -> Water_Temp.:99.95, PWM:36
16:09:19.172 -> Water_Temp.:99.95, PWM:36
16:09:19.961 -> Water_Temp.:99.95, PWM:36
16:09:20.724 -> Water_Temp.:99.84, PWM:36
16:09:22.246 -> Water_Temp.:99.84, PWM:36
16:09:23.039 -> Water_Temp.:99.84, PWM:36
16:09:24.596 -> Water_Temp.:99.84, PWM:36
16:09:26.118 -> Water_Temp.:99.84, PWM:36
16:09:26.910 -> Water_Temp.:99.84, PWM:36
16:09:27.704 -> Water_Temp.:99.84, PWM:36
16:09:28.466 -> Water_Temp.:99.84, PWM:36
16:09:29.257 -> Water_Temp.:99.84, PWM:36
16:09:31.577 -> Water_Temp.:99.84, PWM:36
16:09:32.339 -> Water_Temp.:99.84, PWM:36
16:09:33.133 -> Water_Temp.:99.84, PWM:36
16:09:33.931 -> Water_Temp.:99.84, PWM:36
16:09:34.690 -> Water_Temp.:99.84, PWM:36
16:09:35.482 -> Water_Temp.:99.84, PWM:36
16:09:36.244 -> Water_Temp.:99.84, PWM:36
16:09:37.793 -> Water_Temp.:99.84, PWM:36
16:09:38.583 -> Water_Temp.:99.84, PWM:36
16:09:40.134 -> Water_Temp.:99.84, PWM:36
16:09:40.925 -> Water_Temp.:99.84, PWM:36
16:09:43.997 -> Water_Temp.:99.84, PWM:36
16:09:44.752 -> Water_Temp.:99.84, PWM:36
16:09:45.550 -> Water_Temp.:99.84, PWM:36
16:09:46.310 -> Water_Temp.:99.84, PWM:36
16:09:47.864 -> Water_Temp.:99.84, PWM:36
16:09:50.969 -> Water_Temp.:99.84, PWM:36
16:09:52.526 -> Water_Temp.:99.84, PWM:36
16:09:53.286 -> Water_Temp.:99.84, PWM:36
16:09:54.080 -> Water_Temp.:99.84, PWM:36
16:09:54.840 -> Water_Temp.:99.95, PWM:36
16:09:55.636 -> Water_Temp.:99.95, PWM:36
16:09:58.680 -> Water_Temp.:99.95, PWM:36
16:09:59.476 -> Water_Temp.:99.95, PWM:36
16:10:00.233 -> Water_Temp.:99.95, PWM:36
16:10:01.027 -> Water_Temp.:99.95, PWM:36
16:10:01.825 -> Water_Temp.:99.95, PWM:36
16:10:03.343 -> Water_Temp.:99.95, PWM:36
16:10:04.131 -> Water_Temp.:99.95, PWM:36
16:10:05.680 -> Water_Temp.:99.95, PWM:36
16:10:06.472 -> Water_Temp.:99.95, PWM:36
16:10:08.024 -> Water_Temp.:99.95, PWM:36
16:10:11.859 -> Water_Temp.:99.95, PWM:36
16:10:14.144 -> Water_Temp.:99.95, PWM:36
16:10:14.937 -> Water_Temp.:99.95, PWM:36
16:10:15.697 -> Water_Temp.:99.95, PWM:36
16:10:16.494 -> Water_Temp.:99.95, PWM:36
16:10:17.252 -> Water_Temp.:99.95, PWM:36
16:10:18.047 -> Water_Temp.:99.95, PWM:36
16:10:18.839 -> Water_Temp.:99.95, PWM:36
16:10:19.599 -> Water_Temp.:99.95, PWM:36
16:10:20.391 -> Water_Temp.:99.95, PWM:36
16:10:21.150 -> Water_Temp.:99.84, PWM:36
16:10:22.708 -> Water_Temp.:99.84, PWM:36
16:10:23.502 -> Water_Temp.:99.84, PWM:36
16:10:24.255 -> Water_Temp.:99.84, PWM:36
16:10:25.051 -> Water_Temp.:99.84, PWM:36
16:10:25.844 -> Water_Temp.:99.95, PWM:36
16:10:26.599 -> Water_Temp.:99.95, PWM:36
16:10:28.181 -> Water_Temp.:99.95, PWM:36
16:10:28.941 -> Water_Temp.:99.95, PWM:36
16:10:29.733 -> Water_Temp.:99.95, PWM:36
16:10:30.494 -> Water_Temp.:99.95, PWM:36
16:10:31.284 -> Water_Temp.:99.95, PWM:36
16:10:32.806 -> Water_Temp.:99.95, PWM:36
16:10:33.597 -> Water_Temp.:99.95, PWM:36
16:10:36.640 -> Water_Temp.:99.95, PWM:36
16:10:37.433 -> Water_Temp.:99.95, PWM:36
16:10:38.223 -> Water_Temp.:99.95, PWM:36
16:10:38.983 -> Water_Temp.:99.95, PWM:36
16:10:40.541 -> Water_Temp.:99.95, PWM:36
16:10:41.299 -> Water_Temp.:99.95, PWM:36
16:10:42.096 -> Water_Temp.:99.95, PWM:36
16:10:42.856 -> Water_Temp.:99.95, PWM:36
16:10:43.653 -> Water_Temp.:99.95, PWM:36
16:10:44.446 -> Water_Temp.:99.95, PWM:36
16:10:45.209 -> Water_Temp.:99.95, PWM:36
16:10:47.516 -> Water_Temp.:99.95, PWM:36
16:10:49.066 -> Water_Temp.:99.95, PWM:36
16:10:49.864 -> Water_Temp.:99.95, PWM:36
16:10:50.625 -> Water_Temp.:99.95, PWM:36
16:10:52.209 -> Water_Temp.:99.95, PWM:36
16:10:52.964 -> Water_Temp.:99.95, PWM:36
16:10:54.550 -> Water_Temp.:99.95, PWM:36
16:10:56.834 -> Water_Temp.:99.95, PWM:36
16:10:58.390 -> Water_Temp.:99.95, PWM:36
16:10:59.187 -> Water_Temp.:99.95, PWM:36
16:10:59.946 -> Water_Temp.:99.95, PWM:36
16:11:00.742 -> Water_Temp.:99.95, PWM:36
16:11:01.501 -> Water_Temp.:100.06, PWM:36
16:11:02.296 -> Water_Temp.:100.06, PWM:36
16:11:03.087 -> Water_Temp.:100.06, PWM:36
16:11:04.640 -> Water_Temp.:100.06, PWM:36
16:11:06.200 -> Water_Temp.:100.06, PWM:36
16:11:06.958 -> Water_Temp.:100.06, PWM:36
16:11:07.755 -> Water_Temp.:100.06, PWM:36
16:11:09.276 -> Water_Temp.:100.06, PWM:36
16:11:10.072 -> Water_Temp.:100.06, PWM:36
16:11:10.833 -> Water_Temp.:100.06, PWM:36
16:11:11.627 -> Water_Temp.:100.06, PWM:36
16:11:12.423 -> Water_Temp.:100.06, PWM:36
16:11:13.187 -> Water_Temp.:100.06, PWM:36
16:11:13.983 -> Water_Temp.:100.06, PWM:36
16:11:15.498 -> Water_Temp.:100.06, PWM:36
16:11:16.292 -> Water_Temp.:100.06, PWM:36
16:11:17.085 -> Water_Temp.:100.06, PWM:36
16:11:18.605 -> Water_Temp.:100.06, PWM:36
16:11:19.398 -> Water_Temp.:100.06, PWM:36
16:11:20.157 -> Water_Temp.:100.06, PWM:36
16:11:21.713 -> Water_Temp.:100.06, PWM:36
16:11:22.501 -> Water_Temp.:100.06, PWM:36
16:11:24.813 -> Water_Temp.:100.06, PWM:36
16:11:25.574 -> Water_Temp.:100.06, PWM:36
16:11:27.129 -> Water_Temp.:100.06, PWM:36
16:11:27.924 -> Water_Temp.:100.06, PWM:36
16:11:31.000 -> Water_Temp.:99.95, PWM:36
16:11:31.797 -> Water_Temp.:99.95, PWM:36
16:11:32.552 -> Water_Temp.:99.95, PWM:36
16:11:34.106 -> Water_Temp.:99.95, PWM:36
16:11:34.865 -> Water_Temp.:99.95, PWM:36
16:11:35.660 -> Water_Temp.:99.95, PWM:36
16:11:36.455 -> Water_Temp.:99.95, PWM:36
16:11:37.211 -> Water_Temp.:99.95, PWM:36
16:11:38.763 -> Water_Temp.:99.95, PWM:36
16:11:39.522 -> Water_Temp.:99.95, PWM:36
16:11:40.316 -> Water_Temp.:99.95, PWM:36
16:11:41.109 -> Water_Temp.:101.64, PWM:36
16:11:41.874 -> Water_Temp.:103.66, PWM:36
16:11:42.670 -> Water_Temp.:101.30, PWM:36
16:11:43.432 -> Water_Temp.:100.85, PWM:36
16:11:44.226 -> Water_Temp.:100.62, PWM:36
16:11:44.982 -> Water_Temp.:100.51, PWM:36
16:11:45.775 -> Water_Temp.:100.40, PWM:36
16:11:47.323 -> Water_Temp.:100.29, PWM:36
16:11:48.083 -> Water_Temp.:100.29, PWM:36
16:11:48.876 -> Water_Temp.:100.29, PWM:36
16:11:49.670 -> Water_Temp.:100.29, PWM:36
16:11:51.226 -> Water_Temp.:100.17, PWM:36
16:11:52.782 -> Water_Temp.:100.17, PWM:36
16:11:53.542 -> Water_Temp.:100.06, PWM:36
16:11:54.338 -> Water_Temp.:100.06, PWM:36
16:11:55.863 -> Water_Temp.:100.06, PWM:36
16:11:56.660 -> Water_Temp.:100.06, PWM:36
16:11:57.419 -> Water_Temp.:100.06, PWM:36
16:11:58.211 -> Water_Temp.:100.06, PWM:36
16:11:59.004 -> Water_Temp.:100.06, PWM:36
16:11:59.763 -> Water_Temp.:100.06, PWM:36
16:12:01.317 -> Water_Temp.:100.06, PWM:36
16:12:02.108 -> Water_Temp.:100.06, PWM:36
16:12:02.872 -> Water_Temp.:100.06, PWM:36
16:12:03.671 -> Water_Temp.:100.06, PWM:36
16:12:04.434 -> Water_Temp.:100.06, PWM:36
16:12:05.984 -> Water_Temp.:100.06, PWM:36
16:12:07.536 -> Water_Temp.:100.06, PWM:36
16:12:08.331 -> Water_Temp.:100.06, PWM:36
16:12:09.093 -> Water_Temp.:100.06, PWM:36
16:12:10.649 -> Water_Temp.:100.06, PWM:36
16:12:12.226 -> Water_Temp.:99.95, PWM:36
16:12:12.990 -> Water_Temp.:99.95, PWM:36
16:12:13.784 -> Water_Temp.:99.95, PWM:36
16:12:14.542 -> Water_Temp.:99.95, PWM:36
16:12:16.897 -> Water_Temp.:99.95, PWM:36
16:12:19.183 -> Water_Temp.:99.95, PWM:36
16:12:19.977 -> Water_Temp.:99.95, PWM:32
16:12:22.280 -> Water_Temp.:99.95, PWM:31
16:12:23.070 -> Water_Temp.:99.95, PWM:31
16:12:23.864 -> Water_Temp.:99.95, PWM:32
16:12:24.622 -> Water_Temp.:99.95, PWM:31
16:12:26.934 -> Water_Temp.:99.95, PWM:31
16:12:27.728 -> Water_Temp.:99.95, PWM:31
16:12:28.487 -> Water_Temp.:99.95, PWM:31
16:12:30.043 -> Water_Temp.:99.95, PWM:31
16:12:30.802 -> Water_Temp.:99.95, PWM:32
16:12:31.600 -> Water_Temp.:99.95, PWM:31
16:12:32.360 -> Water_Temp.:99.95, PWM:31
16:12:33.151 -> Water_Temp.:99.95, PWM:31
16:12:34.701 -> Water_Temp.:99.95, PWM:32
16:12:35.495 -> Water_Temp.:99.95, PWM:31
16:12:37.804 -> Water_Temp.:99.95, PWM:31
16:12:38.562 -> Water_Temp.:99.95, PWM:31
16:12:39.354 -> Water_Temp.:99.95, PWM:31
16:12:40.905 -> Water_Temp.:99.95, PWM:31
16:12:42.461 -> Water_Temp.:99.95, PWM:31
16:12:43.217 -> Water_Temp.:99.95, PWM:31
16:12:44.012 -> Water_Temp.:99.95, PWM:31
16:12:45.564 -> Water_Temp.:99.95, PWM:32
16:12:46.325 -> Water_Temp.:99.95, PWM:31
16:12:47.123 -> Water_Temp.:99.95, PWM:31
16:12:49.434 -> Water_Temp.:99.95, PWM:31
16:12:50.226 -> Water_Temp.:100.06, PWM:31
16:12:51.780 -> Water_Temp.:100.06, PWM:32
16:12:52.535 -> Water_Temp.:100.06, PWM:31
16:12:53.328 -> Water_Temp.:100.06, PWM:31
16:12:54.121 -> Water_Temp.:100.06, PWM:31
16:12:55.671 -> Water_Temp.:100.06, PWM:32
16:12:56.431 -> Water_Temp.:100.06, PWM:31
16:12:57.223 -> Water_Temp.:100.06, PWM:32
16:12:57.983 -> Water_Temp.:100.06, PWM:31
16:12:59.571 -> Water_Temp.:100.06, PWM:32
16:13:00.333 -> Water_Temp.:100.06, PWM:31
16:13:01.882 -> Water_Temp.:100.06, PWM:31
16:13:02.678 -> Water_Temp.:100.06, PWM:31
16:13:04.994 -> Water_Temp.:100.06, PWM:31
16:13:05.790 -> Water_Temp.:100.06, PWM:31
16:13:06.552 -> Water_Temp.:100.06, PWM:32
16:13:07.347 -> Water_Temp.:100.06, PWM:32
16:13:08.106 -> Water_Temp.:100.06, PWM:31
16:13:09.661 -> Water_Temp.:100.06, PWM:31
16:13:11.210 -> Water_Temp.:100.06, PWM:31
16:13:12.007 -> Water_Temp.:100.06, PWM:35
16:13:12.764 -> Water_Temp.:100.06, PWM:35
16:13:14.352 -> Water_Temp.:100.06, PWM:35
16:13:15.111 -> Water_Temp.:100.06, PWM:35
16:13:16.664 -> Water_Temp.:100.06, PWM:36
16:13:17.425 -> Water_Temp.:100.17, PWM:36
16:13:18.216 -> Water_Temp.:100.17, PWM:36
16:13:19.011 -> Water_Temp.:100.06, PWM:35
16:13:19.771 -> Water_Temp.:100.06, PWM:36
16:13:20.567 -> Water_Temp.:100.06, PWM:36
16:13:22.122 -> Water_Temp.:100.06, PWM:36
16:13:22.885 -> Water_Temp.:100.06, PWM:36
16:13:24.435 -> Water_Temp.:100.06, PWM:36
16:13:25.229 -> Water_Temp.:100.06, PWM:36
16:13:25.992 -> Water_Temp.:100.06, PWM:36
16:13:28.309 -> Water_Temp.:99.95, PWM:36
16:13:29.108 -> Water_Temp.:99.95, PWM:35
16:13:29.866 -> Water_Temp.:99.95, PWM:35
16:13:30.664 -> Water_Temp.:99.95, PWM:35
16:13:31.425 -> Water_Temp.:99.95, PWM:35
16:13:32.974 -> Water_Temp.:99.95, PWM:35
16:13:36.849 -> Water_Temp.:100.06, PWM:35
16:13:37.640 -> Water_Temp.:100.06, PWM:35
16:13:38.395 -> Water_Temp.:100.06, PWM:35
16:13:39.953 -> Water_Temp.:100.06, PWM:35
16:13:40.749 -> Water_Temp.:100.06, PWM:35
16:13:41.544 -> Water_Temp.:100.06, PWM:35
16:13:42.302 -> Water_Temp.:100.06, PWM:35
16:13:44.584 -> Water_Temp.:100.06, PWM:35
16:13:45.374 -> Water_Temp.:100.06, PWM:35
16:13:46.136 -> Water_Temp.:100.06, PWM:35
16:13:46.932 -> Water_Temp.:100.06, PWM:35
16:13:48.482 -> Water_Temp.:99.95, PWM:35
16:13:49.282 -> Water_Temp.:99.95, PWM:35
16:13:50.799 -> Water_Temp.:99.95, PWM:35
16:13:51.597 -> Water_Temp.:99.95, PWM:35
16:13:52.359 -> Water_Temp.:99.95, PWM:35
16:13:53.154 -> Water_Temp.:99.95, PWM:35
16:13:53.914 -> Water_Temp.:99.95, PWM:35
16:13:54.709 -> Water_Temp.:99.95, PWM:35
16:13:55.496 -> Water_Temp.:99.95, PWM:35
16:13:57.023 -> Water_Temp.:100.06, PWM:35
16:13:57.817 -> Water_Temp.:100.06, PWM:35
16:13:58.579 -> Water_Temp.:100.06, PWM:35
16:13:59.373 -> Water_Temp.:100.06, PWM:35
16:14:00.923 -> Water_Temp.:100.06, PWM:35
16:14:04.751 -> Water_Temp.:100.06, PWM:35
16:14:05.538 -> Water_Temp.:100.06, PWM:35
16:14:06.332 -> Water_Temp.:100.06, PWM:35
16:14:07.090 -> Water_Temp.:100.06, PWM:35
16:14:08.681 -> Water_Temp.:100.06, PWM:35
16:14:09.438 -> Water_Temp.:100.06, PWM:35
16:14:10.232 -> Water_Temp.:100.06, PWM:35
16:14:10.991 -> Water_Temp.:100.06, PWM:35
16:14:11.786 -> Water_Temp.:100.06, PWM:35
16:14:13.340 -> Water_Temp.:100.06, PWM:35
16:14:14.097 -> Water_Temp.:100.17, PWM:35
16:14:14.885 -> Water_Temp.:100.17, PWM:35
16:14:15.677 -> Water_Temp.:100.17, PWM:35
16:14:16.438 -> Water_Temp.:100.17, PWM:35
16:14:19.545 -> Water_Temp.:100.17, PWM:35
16:14:22.652 -> Water_Temp.:100.17, PWM:35
16:14:23.411 -> Water_Temp.:100.17, PWM:35
16:14:25.756 -> Water_Temp.:100.29, PWM:35
16:14:26.553 -> Water_Temp.:100.17, PWM:35
16:14:27.310 -> Water_Temp.:100.17, PWM:35
16:14:28.099 -> Water_Temp.:100.17, PWM:35
16:14:29.651 -> Water_Temp.:100.17, PWM:35
16:14:30.446 -> Water_Temp.:100.17, PWM:35
16:14:31.970 -> Water_Temp.:100.17, PWM:35
16:14:32.762 -> Water_Temp.:100.17, PWM:35
16:14:33.557 -> Water_Temp.:100.17, PWM:35
16:14:35.116 -> Water_Temp.:100.17, PWM:35
16:14:35.875 -> Water_Temp.:100.17, PWM:35
16:14:36.666 -> Water_Temp.:100.17, PWM:35
16:14:37.428 -> Water_Temp.:100.17, PWM:35
16:14:39.020 -> Water_Temp.:100.17, PWM:35
16:14:39.779 -> Water_Temp.:100.17, PWM:35
16:14:40.574 -> Water_Temp.:100.17, PWM:35
16:14:42.892 -> Water_Temp.:100.17, PWM:35
16:14:45.167 -> Water_Temp.:100.17, PWM:35
16:14:45.963 -> Water_Temp.:100.17, PWM:35
16:14:46.757 -> Water_Temp.:100.17, PWM:35
16:14:48.316 -> Water_Temp.:100.17, PWM:35
16:14:50.635 -> Water_Temp.:100.29, PWM:35
16:14:51.429 -> Water_Temp.:100.29, PWM:35
16:14:52.191 -> Water_Temp.:100.29, PWM:35
16:14:52.984 -> Water_Temp.:100.29, PWM:35
16:14:54.540 -> Water_Temp.:101.86, PWM:35
16:14:55.334 -> Water_Temp.:104.00, PWM:35
16:14:56.862 -> Water_Temp.:101.19, PWM:35
16:14:57.655 -> Water_Temp.:100.96, PWM:35
16:14:58.446 -> Water_Temp.:100.85, PWM:35
16:14:59.204 -> Water_Temp.:100.74, PWM:35
16:15:03.073 -> Water_Temp.:100.51, PWM:35
16:15:03.870 -> Water_Temp.:100.40, PWM:35
16:15:05.423 -> Water_Temp.:100.40, PWM:35
16:15:06.186 -> Water_Temp.:100.29, PWM:35
16:15:07.745 -> Water_Temp.:100.29, PWM:35
16:15:08.536 -> Water_Temp.:100.29, PWM:35
16:15:13.194 -> Water_Temp.:100.17, PWM:35
16:15:13.989 -> Water_Temp.:100.17, PWM:35
16:15:17.090 -> Water_Temp.:100.17, PWM:35
16:15:17.884 -> Water_Temp.:100.17, PWM:35
16:15:18.675 -> Water_Temp.:100.17, PWM:35
16:15:19.434 -> Water_Temp.:100.29, PWM:35
16:15:20.226 -> Water_Temp.:100.29, PWM:35
16:15:20.982 -> Water_Temp.:100.29, PWM:35
16:15:21.777 -> Water_Temp.:100.29, PWM:35
16:15:22.575 -> Water_Temp.:100.29, PWM:35
16:15:24.128 -> Water_Temp.:100.40, PWM:35
16:15:27.956 -> Water_Temp.:100.40, PWM:35
16:15:28.751 -> Water_Temp.:100.40, PWM:35
16:15:29.510 -> Water_Temp.:100.40, PWM:35
16:15:30.302 -> Water_Temp.:100.40, PWM:35
16:15:31.096 -> Water_Temp.:100.40, PWM:35
16:15:31.855 -> Water_Temp.:100.40, PWM:35
16:15:32.649 -> Water_Temp.:100.40, PWM:35
16:15:34.199 -> Water_Temp.:100.40, PWM:35
16:15:35.752 -> Water_Temp.:100.40, PWM:35
16:15:36.549 -> Water_Temp.:100.40, PWM:35
16:15:39.619 -> Water_Temp.:100.40, PWM:35
16:15:40.410 -> Water_Temp.:100.40, PWM:35
16:15:41.167 -> Water_Temp.:100.40, PWM:35
16:15:41.959 -> Water_Temp.:100.51, PWM:35
16:15:42.754 -> Water_Temp.:100.40, PWM:35
16:15:44.270 -> Water_Temp.:100.51, PWM:35
16:15:45.058 -> Water_Temp.:100.51, PWM:35
16:15:45.851 -> Water_Temp.:100.51, PWM:35
16:15:46.643 -> Water_Temp.:100.51, PWM:35
16:15:47.406 -> Water_Temp.:100.51, PWM:35
16:15:48.965 -> Water_Temp.:100.51, PWM:35
16:15:49.759 -> Water_Temp.:100.51, PWM:35
16:15:50.518 -> Water_Temp.:100.51, PWM:35
16:15:51.315 -> Water_Temp.:100.51, PWM:35
16:15:52.077 -> Water_Temp.:100.51, PWM:35
16:15:52.874 -> Water_Temp.:100.51, PWM:35
16:15:54.425 -> Water_Temp.:100.51, PWM:35
16:15:55.184 -> Water_Temp.:100.51, PWM:35
16:15:55.981 -> Water_Temp.:100.51, PWM:35
16:15:57.533 -> Water_Temp.:100.51, PWM:35
16:15:59.090 -> Water_Temp.:100.51, PWM:35
16:16:02.129 -> Water_Temp.:100.51, PWM:35
16:16:04.442 -> Water_Temp.:100.40, PWM:35
16:16:05.994 -> Water_Temp.:100.51, PWM:35
16:16:08.339 -> Water_Temp.:100.40, PWM:35
16:16:09.099 -> Water_Temp.:100.40, PWM:35
16:16:09.891 -> Water_Temp.:100.40, PWM:35
16:16:10.684 -> Water_Temp.:100.40, PWM:35
16:16:11.444 -> Water_Temp.:100.40, PWM:35
16:16:12.237 -> Water_Temp.:100.40, PWM:35
16:16:12.996 -> Water_Temp.:100.40, PWM:35
16:16:14.582 -> Water_Temp.:100.40, PWM:35
16:16:16.136 -> Water_Temp.:100.40, PWM:35
16:16:17.687 -> Water_Temp.:100.40, PWM:35
16:16:18.479 -> Water_Temp.:100.40, PWM:35
16:16:19.241 -> Water_Temp.:100.40, PWM:35
16:16:20.033 -> Water_Temp.:100.40, PWM:35
16:16:20.792 -> Water_Temp.:100.40, PWM:35
16:16:21.588 -> Water_Temp.:100.40, PWM:35
16:16:22.382 -> Water_Temp.:100.40, PWM:35
16:16:23.938 -> Water_Temp.:100.40, PWM:35
16:16:25.493 -> Water_Temp.:100.40, PWM:35
16:16:26.252 -> Water_Temp.:100.40, PWM:35
16:16:27.044 -> Water_Temp.:100.51, PWM:35
16:16:27.832 -> Water_Temp.:100.51, PWM:35
16:16:29.387 -> Water_Temp.:100.51, PWM:35
16:16:30.149 -> Water_Temp.:100.51, PWM:35
16:16:30.943 -> Water_Temp.:100.51, PWM:35
16:16:32.489 -> Water_Temp.:100.51, PWM:35
16:16:33.281 -> Water_Temp.:100.51, PWM:35
16:16:34.076 -> Water_Temp.:100.51, PWM:35
16:16:34.840 -> Water_Temp.:100.51, PWM:35
16:16:35.634 -> Water_Temp.:100.51, PWM:35
16:16:37.188 -> Water_Temp.:100.51, PWM:35
16:16:37.944 -> Water_Temp.:100.51, PWM:35
16:16:38.739 -> Water_Temp.:100.51, PWM:35
16:16:39.532 -> Water_Temp.:100.51, PWM:35
16:16:40.289 -> Water_Temp.:100.51, PWM:35
16:16:41.082 -> Water_Temp.:100.51, PWM:35
16:16:42.630 -> Water_Temp.:100.51, PWM:35
16:16:43.421 -> Water_Temp.:100.51, PWM:35
16:16:44.177 -> Water_Temp.:100.51, PWM:35
16:16:45.729 -> Water_Temp.:100.51, PWM:35
16:16:49.582 -> Water_Temp.:100.51, PWM:35
16:16:50.376 -> Water_Temp.:100.51, PWM:35
16:16:51.136 -> Water_Temp.:100.51, PWM:35
16:16:51.933 -> Water_Temp.:100.51, PWM:35
16:16:52.689 -> Water_Temp.:100.51, PWM:35
16:16:54.277 -> Water_Temp.:100.62, PWM:35
16:16:56.590 -> Water_Temp.:100.62, PWM:35
16:16:57.347 -> Water_Temp.:100.62, PWM:35
16:16:58.136 -> Water_Temp.:100.51, PWM:35
16:16:58.931 -> Water_Temp.:100.51, PWM:35
16:16:59.687 -> Water_Temp.:100.51, PWM:35
16:17:00.485 -> Water_Temp.:100.51, PWM:35
16:17:01.241 -> Water_Temp.:100.51, PWM:35
16:17:02.033 -> Water_Temp.:100.51, PWM:35
16:17:04.341 -> Water_Temp.:100.51, PWM:35
16:17:08.208 -> Water_Temp.:100.62, PWM:35
16:17:09.003 -> Water_Temp.:100.62, PWM:35
16:17:09.760 -> Water_Temp.:100.62, PWM:35
16:17:10.555 -> Water_Temp.:100.62, PWM:35
16:17:12.114 -> Water_Temp.:100.62, PWM:35
16:17:12.871 -> Water_Temp.:100.62, PWM:35
16:17:13.667 -> Water_Temp.:100.51, PWM:35
16:17:14.462 -> Water_Temp.:100.62, PWM:35
16:17:15.223 -> Water_Temp.:100.62, PWM:35
16:17:16.776 -> Water_Temp.:100.62, PWM:35
16:17:17.568 -> Water_Temp.:100.51, PWM:35
16:17:18.331 -> Water_Temp.:100.62, PWM:35
16:17:19.124 -> Water_Temp.:100.62, PWM:35
16:17:21.437 -> Water_Temp.:100.51, PWM:35
16:17:22.985 -> Water_Temp.:100.51, PWM:35
16:17:23.779 -> Water_Temp.:100.62, PWM:35
16:17:24.537 -> Water_Temp.:100.51, PWM:35
16:17:25.335 -> Water_Temp.:100.51, PWM:35
16:17:26.098 -> Water_Temp.:100.62, PWM:35
16:17:26.892 -> Water_Temp.:100.62, PWM:35
16:17:28.407 -> Water_Temp.:100.62, PWM:35
16:17:29.199 -> Water_Temp.:100.62, PWM:35
16:17:29.997 -> Water_Temp.:100.62, PWM:35
16:17:30.755 -> Water_Temp.:100.62, PWM:35
16:17:31.549 -> Water_Temp.:100.62, PWM:35
16:17:32.307 -> Water_Temp.:100.62, PWM:35
16:17:33.101 -> Water_Temp.:100.62, PWM:35
16:17:33.897 -> Water_Temp.:100.62, PWM:35
16:17:35.446 -> Water_Temp.:100.62, PWM:35
16:17:36.205 -> Water_Temp.:100.62, PWM:35
16:17:38.552 -> Water_Temp.:100.62, PWM:35
16:17:39.312 -> Water_Temp.:100.62, PWM:35
16:17:40.109 -> Water_Temp.:100.62, PWM:35
16:17:41.628 -> Water_Temp.:100.62, PWM:35
16:17:42.418 -> Water_Temp.:100.62, PWM:35
16:17:43.978 -> Water_Temp.:100.62, PWM:35
16:17:44.736 -> Water_Temp.:100.62, PWM:35
16:17:46.291 -> Water_Temp.:100.62, PWM:35
16:17:47.084 -> Water_Temp.:100.62, PWM:35
16:17:47.846 -> Water_Temp.:100.62, PWM:35
16:17:48.645 -> Water_Temp.:100.62, PWM:35
16:17:49.405 -> Water_Temp.:100.62, PWM:35
16:17:50.196 -> Water_Temp.:100.62, PWM:35
16:17:50.991 -> Water_Temp.:100.62, PWM:35
16:17:52.539 -> Water_Temp.:100.62, PWM:35
16:17:53.297 -> Water_Temp.:100.62, PWM:35
16:17:54.095 -> Water_Temp.:100.62, PWM:35
16:17:55.644 -> Water_Temp.:100.62, PWM:35
16:17:57.194 -> Water_Temp.:100.62, PWM:35
16:17:57.991 -> Water_Temp.:100.62, PWM:35
16:17:58.751 -> Water_Temp.:100.62, PWM:35
16:17:59.551 -> Water_Temp.:100.62, PWM:35
16:18:00.309 -> Water_Temp.:100.62, PWM:35
16:18:01.105 -> Water_Temp.:100.62, PWM:35
16:18:04.208 -> Water_Temp.:100.62, PWM:35
16:18:04.970 -> Water_Temp.:100.62, PWM:35
16:18:05.766 -> Water_Temp.:100.62, PWM:35
16:18:06.524 -> Water_Temp.:100.62, PWM:35
16:18:08.077 -> Water_Temp.:100.62, PWM:35
16:18:09.594 -> Water_Temp.:100.62, PWM:35
16:18:10.389 -> Water_Temp.:100.62, PWM:35
16:18:11.151 -> Water_Temp.:100.62, PWM:35
16:18:11.948 -> Water_Temp.:100.62, PWM:35
16:18:12.742 -> Water_Temp.:100.62, PWM:35
16:18:14.293 -> Water_Temp.:100.62, PWM:35
16:18:15.053 -> Water_Temp.:100.62, PWM:35
16:18:17.396 -> Water_Temp.:100.62, PWM:35
16:18:18.150 -> Water_Temp.:100.62, PWM:35
16:18:18.946 -> Water_Temp.:100.62, PWM:35
16:18:19.742 -> Water_Temp.:100.62, PWM:35
16:18:20.500 -> Water_Temp.:100.62, PWM:35
16:18:21.295 -> Water_Temp.:100.62, PWM:35
16:18:22.055 -> Water_Temp.:100.62, PWM:35
16:18:24.369 -> Water_Temp.:100.62, PWM:35
16:18:25.163 -> Water_Temp.:100.62, PWM:35
16:18:31.380 -> Water_Temp.:100.62, PWM:35
16:18:32.930 -> Water_Temp.:100.51, PWM:35
16:18:34.488 -> Water_Temp.:100.62, PWM:35
16:18:35.280 -> Water_Temp.:100.62, PWM:35
16:18:36.042 -> Water_Temp.:100.62, PWM:35
16:18:37.566 -> Water_Temp.:100.62, PWM:35
16:18:38.364 -> Water_Temp.:100.62, PWM:35
16:18:39.129 -> Water_Temp.:100.62, PWM:35
16:18:39.925 -> Water_Temp.:100.62, PWM:35
16:18:41.486 -> Water_Temp.:100.62, PWM:35
16:18:42.249 -> Water_Temp.:100.62, PWM:35
16:18:43.044 -> Water_Temp.:100.62, PWM:35
16:18:43.805 -> Water_Temp.:100.62, PWM:35
16:18:44.600 -> Water_Temp.:100.62, PWM:35
16:18:45.391 -> Water_Temp.:100.62, PWM:35
16:18:46.151 -> Water_Temp.:100.62, PWM:35
16:18:47.699 -> Water_Temp.:100.62, PWM:35
16:18:48.489 -> Water_Temp.:100.62, PWM:35
16:18:49.248 -> Water_Temp.:100.62, PWM:35
16:18:50.042 -> Water_Temp.:100.62, PWM:35
16:18:50.806 -> Water_Temp.:100.62, PWM:35
16:18:51.602 -> Water_Temp.:100.62, PWM:35
16:18:52.361 -> Water_Temp.:100.62, PWM:35
16:18:53.151 -> Water_Temp.:100.62, PWM:35
16:18:54.713 -> Water_Temp.:100.62, PWM:35
16:18:55.471 -> Water_Temp.:100.62, PWM:35
16:18:56.268 -> Water_Temp.:100.62, PWM:35
16:18:57.032 -> Water_Temp.:100.62, PWM:35
16:18:59.381 -> Water_Temp.:100.62, PWM:35
16:19:02.459 -> Water_Temp.:100.62, PWM:35
16:19:03.252 -> Water_Temp.:100.62, PWM:35
16:19:04.807 -> Water_Temp.:100.62, PWM:35
16:19:06.359 -> Water_Temp.:100.62, PWM:35
16:19:07.120 -> Water_Temp.:100.62, PWM:35
16:19:07.915 -> Water_Temp.:100.62, PWM:35
16:19:08.676 -> Water_Temp.:100.62, PWM:35
16:19:09.473 -> Water_Temp.:100.62, PWM:35
16:19:10.238 -> Water_Temp.:100.62, PWM:35
16:19:11.796 -> Water_Temp.:100.62, PWM:35
16:19:12.554 -> Water_Temp.:100.62, PWM:35
16:19:13.349 -> Water_Temp.:100.62, PWM:35
16:19:14.141 -> Water_Temp.:100.62, PWM:35
16:19:14.903 -> Water_Temp.:100.62, PWM:35
16:19:17.247 -> Water_Temp.:100.62, PWM:35
16:19:18.041 -> Water_Temp.:100.62, PWM:35
16:19:18.797 -> Water_Temp.:100.62, PWM:35
16:19:19.591 -> Water_Temp.:100.62, PWM:35
16:19:20.385 -> Water_Temp.:100.62, PWM:35
16:19:23.457 -> Water_Temp.:100.74, PWM:35
16:19:25.766 -> Water_Temp.:100.62, PWM:35
16:19:26.526 -> Water_Temp.:100.62, PWM:35
16:19:28.084 -> Water_Temp.:100.62, PWM:35
16:19:28.847 -> Water_Temp.:100.62, PWM:35
16:19:29.635 -> Water_Temp.:100.62, PWM:35
16:19:30.397 -> Water_Temp.:100.62, PWM:35
16:19:32.747 -> Water_Temp.:100.74, PWM:35
16:19:34.305 -> Water_Temp.:100.74, PWM:35
16:19:35.066 -> Water_Temp.:100.74, PWM:35
16:19:36.625 -> Water_Temp.:104.45, PWM:35
16:19:37.414 -> Water_Temp.:101.97, PWM:35
16:19:38.209 -> Water_Temp.:101.64, PWM:35
16:19:38.967 -> Water_Temp.:101.41, PWM:35
16:19:39.765 -> Water_Temp.:101.41, PWM:35
16:19:41.323 -> Water_Temp.:101.07, PWM:35
16:19:42.081 -> Water_Temp.:100.96, PWM:35
16:19:45.194 -> Water_Temp.:100.85, PWM:35
16:19:45.955 -> Water_Temp.:100.85, PWM:35
16:19:46.747 -> Water_Temp.:100.85, PWM:35
16:19:47.542 -> Water_Temp.:100.85, PWM:35
16:19:48.302 -> Water_Temp.:100.74, PWM:35
16:19:49.098 -> Water_Temp.:100.74, PWM:35
16:19:50.653 -> Water_Temp.:100.74, PWM:35
16:19:51.418 -> Water_Temp.:100.74, PWM:35
16:19:52.181 -> Water_Temp.:100.74, PWM:35
16:19:52.972 -> Water_Temp.:100.74, PWM:35
16:19:53.769 -> Water_Temp.:100.74, PWM:35
16:19:54.530 -> Water_Temp.:100.74, PWM:35
16:19:55.323 -> Water_Temp.:100.74, PWM:35
16:19:56.847 -> Water_Temp.:100.74, PWM:35
16:19:57.641 -> Water_Temp.:100.74, PWM:35
16:19:58.434 -> Water_Temp.:100.74, PWM:35
16:19:59.196 -> Water_Temp.:100.74, PWM:35
16:19:59.989 -> Water_Temp.:100.74, PWM:35
16:20:00.746 -> Water_Temp.:100.74, PWM:35
16:20:02.303 -> Water_Temp.:100.74, PWM:35
16:20:03.889 -> Water_Temp.:100.74, PWM:35
16:20:04.649 -> Water_Temp.:100.74, PWM:35
16:20:06.200 -> Water_Temp.:100.74, PWM:35
16:20:06.962 -> Water_Temp.:100.74, PWM:35
16:20:07.752 -> Water_Temp.:100.74, PWM:35
16:20:10.094 -> Water_Temp.:100.74, PWM:35
16:20:10.857 -> Water_Temp.:100.74, PWM:35
16:20:11.656 -> Water_Temp.:100.74, PWM:35
16:20:12.417 -> Water_Temp.:100.74, PWM:35
16:20:14.005 -> Water_Temp.:100.74, PWM:35
16:20:14.765 -> Water_Temp.:100.74, PWM:35
16:20:15.560 -> Water_Temp.:100.74, PWM:35
16:20:17.872 -> Water_Temp.:100.74, PWM:35
16:20:18.667 -> Water_Temp.:100.74, PWM:35
16:20:22.494 -> Water_Temp.:100.74, PWM:35
16:20:23.251 -> Water_Temp.:100.74, PWM:35
16:20:24.043 -> Water_Temp.:100.74, PWM:35
16:20:25.591 -> Water_Temp.:100.74, PWM:35
16:20:27.146 -> Water_Temp.:100.74, PWM:35
16:20:28.695 -> Water_Temp.:100.74, PWM:35
16:20:29.488 -> Water_Temp.:100.74, PWM:35
16:20:30.247 -> Water_Temp.:100.74, PWM:35
16:20:31.043 -> Water_Temp.:100.85, PWM:35
16:20:31.802 -> Water_Temp.:100.85, PWM:35
16:20:32.595 -> Water_Temp.:100.85, PWM:35
16:20:33.389 -> Water_Temp.:100.85, PWM:35
16:20:34.149 -> Water_Temp.:100.85, PWM:35
16:20:34.945 -> Water_Temp.:100.85, PWM:35
16:20:36.464 -> Water_Temp.:100.85, PWM:35
16:20:38.052 -> Water_Temp.:100.85, PWM:35
16:20:38.810 -> Water_Temp.:100.85, PWM:35
16:20:40.392 -> Water_Temp.:100.85, PWM:35
16:20:41.912 -> Water_Temp.:100.85, PWM:35
16:20:42.701 -> Water_Temp.:100.85, PWM:35
16:20:43.492 -> Water_Temp.:100.85, PWM:35
16:20:44.250 -> Water_Temp.:100.85, PWM:35
16:20:45.047 -> Water_Temp.:100.85, PWM:35
16:20:45.843 -> Water_Temp.:100.85, PWM:35
16:20:46.606 -> Water_Temp.:100.85, PWM:35
16:20:47.399 -> Water_Temp.:100.85, PWM:35
16:20:48.156 -> Water_Temp.:100.85, PWM:35
16:20:48.951 -> Water_Temp.:100.85, PWM:35
16:20:50.505 -> Water_Temp.:100.85, PWM:35
16:20:51.297 -> Water_Temp.:100.74, PWM:35
16:20:52.817 -> Water_Temp.:100.74, PWM:35
16:20:53.611 -> Water_Temp.:100.74, PWM:35
16:20:54.376 -> Water_Temp.:100.74, PWM:35
16:20:55.173 -> Water_Temp.:100.74, PWM:35
16:20:56.725 -> Water_Temp.:100.85, PWM:35
16:20:57.520 -> Water_Temp.:100.74, PWM:35
16:20:58.281 -> Water_Temp.:100.74, PWM:35
16:20:59.838 -> Water_Temp.:100.74, PWM:35
16:21:00.600 -> Water_Temp.:100.74, PWM:35
16:21:01.397 -> Water_Temp.:100.74, PWM:35
16:21:03.746 -> Water_Temp.:100.74, PWM:35
16:21:04.507 -> Water_Temp.:100.74, PWM:35
16:21:05.302 -> Water_Temp.:100.74, PWM:35
16:21:06.063 -> Water_Temp.:100.74, PWM:35
16:21:06.849 -> Water_Temp.:100.85, PWM:35
16:21:07.613 -> Water_Temp.:100.85, PWM:35
16:21:08.409 -> Water_Temp.:100.85, PWM:35
16:21:09.172 -> Water_Temp.:100.85, PWM:35
16:21:11.488 -> Water_Temp.:100.85, PWM:35
16:21:12.285 -> Water_Temp.:100.85, PWM:35
16:21:14.563 -> Water_Temp.:100.85, PWM:35
16:21:17.675 -> Water_Temp.:100.85, PWM:35
16:21:18.437 -> Water_Temp.:100.85, PWM:35
16:21:19.230 -> Water_Temp.:100.85, PWM:35
16:21:19.993 -> Water_Temp.:100.96, PWM:35
16:21:20.786 -> Water_Temp.:100.96, PWM:35
16:21:23.129 -> Water_Temp.:100.96, PWM:35
16:21:24.678 -> Water_Temp.:100.96, PWM:35
16:21:25.468 -> Water_Temp.:100.96, PWM:35
16:21:27.024 -> Water_Temp.:100.96, PWM:35
16:21:27.815 -> Water_Temp.:100.96, PWM:35
16:21:28.576 -> Water_Temp.:100.96, PWM:35
16:21:29.365 -> Water_Temp.:100.96, PWM:35
16:21:30.123 -> Water_Temp.:100.96, PWM:35
16:21:31.672 -> Water_Temp.:100.96, PWM:35
16:21:32.466 -> Water_Temp.:100.85, PWM:35
16:21:33.228 -> Water_Temp.:100.85, PWM:35
16:21:34.019 -> Water_Temp.:100.85, PWM:35
16:21:34.813 -> Water_Temp.:100.85, PWM:35
16:21:35.573 -> Water_Temp.:100.85, PWM:35
16:21:36.368 -> Water_Temp.:100.85, PWM:35
16:21:40.151 -> Water_Temp.:100.85, PWM:35
16:21:40.946 -> Water_Temp.:100.85, PWM:35
16:21:41.738 -> Water_Temp.:100.85, PWM:35
16:21:42.492 -> Water_Temp.:100.85, PWM:35
16:21:43.285 -> Water_Temp.:100.85, PWM:35
16:21:44.834 -> Water_Temp.:100.85, PWM:35
16:21:46.388 -> Water_Temp.:100.96, PWM:35
16:21:47.178 -> Water_Temp.:100.96, PWM:35
16:21:47.969 -> Water_Temp.:100.96, PWM:35
16:21:48.727 -> Water_Temp.:100.96, PWM:35
16:21:51.075 -> Water_Temp.:100.96, PWM:35
16:21:53.349 -> Water_Temp.:100.96, PWM:35
16:21:54.141 -> Water_Temp.:100.96, PWM:35
16:21:54.898 -> Water_Temp.:100.96, PWM:35
16:21:56.453 -> Water_Temp.:100.96, PWM:35
16:21:57.245 -> Water_Temp.:100.96, PWM:35
16:21:59.596 -> Water_Temp.:100.96, PWM:35
16:22:00.358 -> Water_Temp.:100.96, PWM:35
16:22:01.155 -> Water_Temp.:100.96, PWM:35
16:22:01.914 -> Water_Temp.:100.96, PWM:35
16:22:02.703 -> Water_Temp.:100.96, PWM:35
16:22:03.495 -> Water_Temp.:100.96, PWM:35
16:22:05.047 -> Water_Temp.:100.96, PWM:35
16:22:05.807 -> Water_Temp.:100.96, PWM:35
16:22:06.602 -> Water_Temp.:100.96, PWM:35
16:22:07.392 -> Water_Temp.:100.96, PWM:35
16:22:10.426 -> Water_Temp.:100.96, PWM:35
16:22:11.220 -> Water_Temp.:100.96, PWM:35
16:22:11.978 -> Water_Temp.:100.96, PWM:35
16:22:13.530 -> Water_Temp.:100.96, PWM:35
16:22:14.322 -> Water_Temp.:100.96, PWM:35
16:22:15.081 -> Water_Temp.:100.96, PWM:35
16:22:17.399 -> Water_Temp.:100.96, PWM:35
16:22:19.710 -> Water_Temp.:100.96, PWM:35
16:22:20.509 -> Water_Temp.:100.96, PWM:35
16:22:21.267 -> Water_Temp.:101.07, PWM:35
16:22:22.063 -> Water_Temp.:101.07, PWM:35
16:22:22.823 -> Water_Temp.:100.96, PWM:35
16:22:23.617 -> Water_Temp.:101.07, PWM:35
16:22:25.930 -> Water_Temp.:101.07, PWM:35
16:22:26.691 -> Water_Temp.:101.07, PWM:35
16:22:28.238 -> Water_Temp.:101.07, PWM:35
16:22:29.794 -> Water_Temp.:101.07, PWM:35
16:22:30.555 -> Water_Temp.:101.07, PWM:35
16:22:31.344 -> Water_Temp.:101.07, PWM:35
16:22:32.887 -> Water_Temp.:101.07, PWM:35
16:22:33.682 -> Water_Temp.:101.07, PWM:35
16:22:34.439 -> Water_Temp.:101.07, PWM:35
16:22:35.995 -> Water_Temp.:101.07, PWM:35
16:22:38.308 -> Water_Temp.:101.07, PWM:35
16:22:39.859 -> Water_Temp.:101.07, PWM:35
16:22:41.411 -> Water_Temp.:101.07, PWM:35
16:22:42.173 -> Water_Temp.:101.07, PWM:35
16:22:42.968 -> Water_Temp.:101.07, PWM:35
16:22:43.729 -> Water_Temp.:101.07, PWM:35
16:22:45.282 -> Water_Temp.:101.07, PWM:35
16:22:46.080 -> Water_Temp.:101.07, PWM:35
16:22:49.188 -> Water_Temp.:101.07, PWM:35
16:22:49.950 -> Water_Temp.:101.07, PWM:35
16:22:51.502 -> Water_Temp.:101.07, PWM:35
16:22:53.848 -> Water_Temp.:101.07, PWM:35
16:22:54.636 -> Water_Temp.:101.07, PWM:35
16:22:55.431 -> Water_Temp.:101.07, PWM:35
16:22:57.708 -> Water_Temp.:101.07, PWM:35
16:22:58.504 -> Water_Temp.:101.19, PWM:35
16:22:59.299 -> Water_Temp.:101.19, PWM:35
16:23:00.058 -> Water_Temp.:101.19, PWM:35
16:23:00.853 -> Water_Temp.:101.19, PWM:35
16:23:01.616 -> Water_Temp.:101.19, PWM:35
16:23:04.689 -> Water_Temp.:101.19, PWM:35
16:23:05.478 -> Water_Temp.:101.07, PWM:35
16:23:06.275 -> Water_Temp.:101.07, PWM:35
16:23:07.034 -> Water_Temp.:101.07, PWM:35
16:23:07.825 -> Water_Temp.:101.07, PWM:35
16:23:10.932 -> Water_Temp.:101.07, PWM:35
16:23:11.693 -> Water_Temp.:101.07, PWM:35
16:23:12.486 -> Water_Temp.:101.07, PWM:35
16:23:13.246 -> Water_Temp.:101.07, PWM:35
16:23:14.042 -> Water_Temp.:101.07, PWM:35
16:23:15.597 -> Water_Temp.:101.07, PWM:35
16:23:16.390 -> Water_Temp.:101.07, PWM:35
16:23:17.149 -> Water_Temp.:101.07, PWM:35
16:23:17.940 -> Water_Temp.:101.07, PWM:35
16:23:19.484 -> Water_Temp.:101.07, PWM:35
16:23:21.831 -> Water_Temp.:101.07, PWM:35
16:23:24.141 -> Water_Temp.:101.19, PWM:35
16:23:25.688 -> Water_Temp.:101.19, PWM:35
16:23:26.447 -> Water_Temp.:101.19, PWM:35
16:23:27.238 -> Water_Temp.:101.19, PWM:35
16:23:29.577 -> Water_Temp.:101.19, PWM:35
16:23:30.339 -> Water_Temp.:101.19, PWM:35
16:23:31.132 -> Water_Temp.:101.19, PWM:35
16:23:31.895 -> Water_Temp.:101.19, PWM:35
16:23:32.690 -> Water_Temp.:101.19, PWM:35
16:23:35.009 -> Water_Temp.:101.19, PWM:35
16:23:37.323 -> Water_Temp.:101.19, PWM:35
16:23:38.119 -> Water_Temp.:101.19, PWM:35
16:23:38.881 -> Water_Temp.:101.19, PWM:35
16:23:39.679 -> Water_Temp.:101.19, PWM:35
16:23:41.238 -> Water_Temp.:101.19, PWM:35
16:23:41.997 -> Water_Temp.:101.19, PWM:35
16:23:42.795 -> Water_Temp.:101.19, PWM:35
16:23:44.348 -> Water_Temp.:101.19, PWM:35
16:23:45.107 -> Water_Temp.:101.19, PWM:35
16:23:45.905 -> Water_Temp.:101.19, PWM:35
16:23:46.667 -> Water_Temp.:101.19, PWM:35
16:23:47.463 -> Water_Temp.:101.07, PWM:35
16:23:48.221 -> Water_Temp.:101.07, PWM:35
16:23:49.016 -> Water_Temp.:101.07, PWM:35
16:23:49.809 -> Water_Temp.:101.07, PWM:35
16:23:50.565 -> Water_Temp.:101.07, PWM:35
16:23:51.350 -> Water_Temp.:101.19, PWM:35
16:23:52.135 -> Water_Temp.:101.19, PWM:35
16:23:52.928 -> Water_Temp.:101.19, PWM:35
16:23:55.237 -> Water_Temp.:101.19, PWM:35
16:23:56.034 -> Water_Temp.:101.19, PWM:35
16:23:57.589 -> Water_Temp.:101.19, PWM:35
16:23:58.384 -> Water_Temp.:101.19, PWM:35
16:23:59.935 -> Water_Temp.:101.07, PWM:35
16:24:00.697 -> Water_Temp.:101.07, PWM:35
16:24:01.485 -> Water_Temp.:101.07, PWM:35
16:24:04.584 -> Water_Temp.:101.19, PWM:35
16:24:06.105 -> Water_Temp.:101.07, PWM:35
16:24:06.903 -> Water_Temp.:101.19, PWM:35
16:24:07.664 -> Water_Temp.:101.19, PWM:35
16:24:08.458 -> Water_Temp.:101.07, PWM:35
16:24:09.220 -> Water_Temp.:101.19, PWM:35
16:24:10.015 -> Water_Temp.:101.19, PWM:35
16:24:11.576 -> Water_Temp.:101.07, PWM:35
16:24:12.335 -> Water_Temp.:101.07, PWM:35
16:24:13.134 -> Water_Temp.:101.19, PWM:35
16:24:14.651 -> Water_Temp.:101.19, PWM:35
16:24:15.439 -> Water_Temp.:101.19, PWM:35
16:24:16.237 -> Water_Temp.:101.19, PWM:35
16:24:16.998 -> Water_Temp.:101.19, PWM:35
16:24:18.545 -> Water_Temp.:101.19, PWM:35
16:24:20.098 -> Water_Temp.:101.19, PWM:35
16:24:22.447 -> Water_Temp.:101.19, PWM:35
16:24:25.513 -> Water_Temp.:101.19, PWM:35
16:24:26.274 -> Water_Temp.:101.19, PWM:35
16:24:29.354 -> Water_Temp.:101.07, PWM:35
16:24:31.699 -> Water_Temp.:101.07, PWM:35
16:24:32.461 -> Water_Temp.:101.07, PWM:35
16:24:33.253 -> Water_Temp.:101.07, PWM:35
16:24:34.015 -> Water_Temp.:101.07, PWM:35
16:24:35.576 -> Water_Temp.:101.07, PWM:35
16:24:36.340 -> Water_Temp.:101.07, PWM:35
16:24:37.135 -> Water_Temp.:102.76, PWM:35
16:24:37.928 -> Water_Temp.:104.90, PWM:35
16:24:38.689 -> Water_Temp.:102.42, PWM:35
16:24:39.484 -> Water_Temp.:101.97, PWM:35
16:24:40.244 -> Water_Temp.:101.75, PWM:35
16:24:41.039 -> Water_Temp.:101.64, PWM:35
16:24:41.828 -> Water_Temp.:101.64, PWM:35
16:24:42.584 -> Water_Temp.:101.52, PWM:35
16:24:43.381 -> Water_Temp.:101.41, PWM:35
16:24:44.902 -> Water_Temp.:101.41, PWM:35
16:24:45.698 -> Water_Temp.:101.30, PWM:35
16:24:46.457 -> Water_Temp.:101.30, PWM:35
16:24:47.254 -> Water_Temp.:101.30, PWM:35
16:24:48.049 -> Water_Temp.:101.30, PWM:35
16:24:48.808 -> Water_Temp.:101.30, PWM:35
16:24:49.603 -> Water_Temp.:101.30, PWM:35
16:24:51.123 -> Water_Temp.:101.30, PWM:35
16:24:54.224 -> Water_Temp.:101.19, PWM:35
16:24:54.985 -> Water_Temp.:101.30, PWM:35
16:24:56.535 -> Water_Temp.:101.19, PWM:35
16:24:58.882 -> Water_Temp.:101.19, PWM:35
16:25:01.228 -> Water_Temp.:101.19, PWM:35
16:25:02.751 -> Water_Temp.:101.30, PWM:35
16:25:03.549 -> Water_Temp.:101.30, PWM:35
16:25:04.312 -> Water_Temp.:101.30, PWM:35
16:25:05.106 -> Water_Temp.:101.30, PWM:35
16:25:06.625 -> Water_Temp.:101.30, PWM:35
16:25:08.215 -> Water_Temp.:101.30, PWM:35
16:25:08.970 -> Water_Temp.:101.30, PWM:35
16:25:10.522 -> Water_Temp.:101.30, PWM:35
16:25:12.076 -> Water_Temp.:101.30, PWM:35
16:25:12.872 -> Water_Temp.:101.30, PWM:35
16:25:13.633 -> Water_Temp.:101.30, PWM:35
16:25:14.429 -> Water_Temp.:101.19, PWM:35
16:25:15.189 -> Water_Temp.:101.30, PWM:35
16:25:15.985 -> Water_Temp.:101.19, PWM:35
16:25:16.739 -> Water_Temp.:101.30, PWM:35
16:25:17.527 -> Water_Temp.:101.30, PWM:35
16:25:18.308 -> Water_Temp.:101.30, PWM:35
16:25:19.097 -> Water_Temp.:101.19, PWM:35
16:25:19.884 -> Water_Temp.:101.19, PWM:35
16:25:20.644 -> Water_Temp.:101.19, PWM:35
16:25:21.437 -> Water_Temp.:101.19, PWM:35
16:25:22.230 -> Water_Temp.:102.87, PWM:35
16:25:22.991 -> Water_Temp.:105.01, PWM:35
16:25:23.781 -> Water_Temp.:102.54, PWM:35
16:25:24.542 -> Water_Temp.:102.09, PWM:35
16:25:26.860 -> Water_Temp.:101.64, PWM:35
16:25:27.621 -> Water_Temp.:101.52, PWM:35
16:25:29.178 -> Water_Temp.:101.52, PWM:35
16:25:29.970 -> Water_Temp.:101.41, PWM:35
16:25:30.761 -> Water_Temp.:101.41, PWM:35
16:25:31.553 -> Water_Temp.:101.41, PWM:35
16:25:33.103 -> Water_Temp.:101.30, PWM:35
16:25:36.211 -> Water_Temp.:101.30, PWM:35
16:25:37.009 -> Water_Temp.:101.30, PWM:35
16:25:37.768 -> Water_Temp.:101.30, PWM:35
16:25:38.564 -> Water_Temp.:101.30, PWM:35
16:25:40.874 -> Water_Temp.:101.30, PWM:35
16:25:41.638 -> Water_Temp.:101.30, PWM:35
16:25:42.429 -> Water_Temp.:101.30, PWM:35
16:25:43.189 -> Water_Temp.:101.30, PWM:35
16:25:45.501 -> Water_Temp.:101.30, PWM:35
16:25:46.295 -> Water_Temp.:101.30, PWM:35
16:25:47.089 -> Water_Temp.:101.30, PWM:35
16:25:47.848 -> Water_Temp.:101.19, PWM:35
16:25:48.640 -> Water_Temp.:101.19, PWM:35
16:25:49.431 -> Water_Temp.:101.19, PWM:35
16:25:50.192 -> Water_Temp.:101.19, PWM:35
16:25:51.746 -> Water_Temp.:101.19, PWM:35
16:25:52.505 -> Water_Temp.:101.19, PWM:35
16:25:53.298 -> Water_Temp.:101.19, PWM:35
16:25:54.820 -> Water_Temp.:101.19, PWM:35
16:25:55.608 -> Water_Temp.:101.19, PWM:35
16:25:56.368 -> Water_Temp.:101.19, PWM:35
16:25:57.159 -> Water_Temp.:101.19, PWM:35
16:25:58.716 -> Water_Temp.:101.19, PWM:35
16:25:59.508 -> Water_Temp.:101.19, PWM:35
16:26:00.269 -> Water_Temp.:101.19, PWM:35
16:26:02.574 -> Water_Temp.:101.19, PWM:35
16:26:03.370 -> Water_Temp.:101.19, PWM:35
16:26:04.129 -> Water_Temp.:101.19, PWM:35
16:26:04.926 -> Water_Temp.:101.19, PWM:35
16:26:05.686 -> Water_Temp.:101.19, PWM:35
16:26:07.245 -> Water_Temp.:101.19, PWM:35
16:26:08.037 -> Water_Temp.:101.19, PWM:35
16:26:08.795 -> Water_Temp.:101.19, PWM:35
16:26:09.584 -> Water_Temp.:101.19, PWM:35
16:26:10.378 -> Water_Temp.:101.19, PWM:35
16:26:11.904 -> Water_Temp.:101.19, PWM:35
16:26:12.693 -> Water_Temp.:101.07, PWM:35
16:26:14.247 -> Water_Temp.:101.07, PWM:35
16:26:15.042 -> Water_Temp.:101.07, PWM:35
16:26:16.597 -> Water_Temp.:101.07, PWM:35
16:26:18.111 -> Water_Temp.:101.07, PWM:35
16:26:18.908 -> Water_Temp.:101.07, PWM:35
16:26:19.670 -> Water_Temp.:101.07, PWM:35
16:26:21.254 -> Water_Temp.:102.76, PWM:35
16:26:22.014 -> Water_Temp.:104.79, PWM:35
16:26:22.808 -> Water_Temp.:102.42, PWM:35
16:26:23.569 -> Water_Temp.:102.42, PWM:35
16:26:25.122 -> Water_Temp.:101.64, PWM:35
16:26:27.436 -> Water_Temp.:101.41, PWM:35
16:26:30.552 -> Water_Temp.:101.19, PWM:35
16:26:32.071 -> Water_Temp.:101.19, PWM:35
16:26:32.861 -> Water_Temp.:101.19, PWM:35
16:26:34.414 -> Water_Temp.:101.19, PWM:35
16:26:35.210 -> Water_Temp.:101.07, PWM:35
16:26:35.968 -> Water_Temp.:101.07, PWM:35
16:26:36.764 -> Water_Temp.:101.07, PWM:35
16:26:37.527 -> Water_Temp.:101.07, PWM:35
16:26:39.841 -> Water_Temp.:101.07, PWM:35
16:26:40.636 -> Water_Temp.:101.07, PWM:35
16:26:41.429 -> Water_Temp.:101.07, PWM:35
16:26:42.190 -> Water_Temp.:101.07, PWM:35
16:26:43.746 -> Water_Temp.:101.07, PWM:35
16:26:44.502 -> Water_Temp.:101.07, PWM:35
16:26:45.292 -> Water_Temp.:101.07, PWM:35
16:26:46.089 -> Water_Temp.:101.07, PWM:35
16:26:47.602 -> Water_Temp.:101.07, PWM:35
16:26:49.158 -> Water_Temp.:101.07, PWM:35
16:26:49.947 -> Water_Temp.:101.07, PWM:35
16:26:51.503 -> Water_Temp.:101.07, PWM:35
16:26:52.301 -> Water_Temp.:101.07, PWM:35
16:26:53.852 -> Water_Temp.:101.07, PWM:35
16:26:54.612 -> Water_Temp.:101.07, PWM:35
16:26:55.404 -> Water_Temp.:101.07, PWM:35
16:26:56.197 -> Water_Temp.:101.19, PWM:35
16:26:57.752 -> Water_Temp.:101.07, PWM:35
16:27:01.620 -> Water_Temp.:101.07, PWM:35
16:27:02.379 -> Water_Temp.:101.07, PWM:35
16:27:03.175 -> Water_Temp.:101.07, PWM:35
16:27:04.723 -> Water_Temp.:101.07, PWM:35
16:27:05.516 -> Water_Temp.:101.07, PWM:35
16:27:06.280 -> Water_Temp.:101.07, PWM:35
16:27:07.077 -> Water_Temp.:101.07, PWM:35
16:27:09.390 -> Water_Temp.:101.07, PWM:35
16:27:10.184 -> Water_Temp.:101.07, PWM:36
16:27:10.946 -> Water_Temp.:101.07, PWM:35
16:27:11.740 -> Water_Temp.:101.07, PWM:35
16:27:12.501 -> Water_Temp.:101.07, PWM:35
16:27:13.295 -> Water_Temp.:100.96, PWM:35
16:27:14.058 -> Water_Temp.:101.07, PWM:35
16:27:14.849 -> Water_Temp.:101.07, PWM:35
16:27:15.646 -> Water_Temp.:101.07, PWM:35
16:27:17.163 -> Water_Temp.:101.07, PWM:35
16:27:17.958 -> Water_Temp.:101.07, PWM:35
16:27:18.752 -> Water_Temp.:101.07, PWM:35
16:27:19.507 -> Water_Temp.:101.07, PWM:35
16:27:21.059 -> Water_Temp.:101.07, PWM:35
16:27:21.817 -> Water_Temp.:101.07, PWM:35
16:27:22.613 -> Water_Temp.:101.07, PWM:35
16:27:23.408 -> Water_Temp.:101.07, PWM:35
16:27:24.169 -> Water_Temp.:101.07, PWM:35
16:27:28.032 -> Water_Temp.:101.07, PWM:35
16:27:29.585 -> Water_Temp.:101.07, PWM:35
16:27:30.380 -> Water_Temp.:101.07, PWM:35
16:27:31.927 -> Water_Temp.:101.07, PWM:35
16:27:32.717 -> Water_Temp.:101.07, PWM:35
16:27:33.472 -> Water_Temp.:101.07, PWM:35
16:27:34.263 -> Water_Temp.:101.07, PWM:35
16:27:35.050 -> Water_Temp.:101.07, PWM:35
16:27:36.602 -> Water_Temp.:101.19, PWM:35
16:27:37.362 -> Water_Temp.:101.07, PWM:35
16:27:38.151 -> Water_Temp.:101.19, PWM:35
16:27:38.943 -> Water_Temp.:101.07, PWM:35
16:27:39.738 -> Water_Temp.:101.19, PWM:35
16:27:40.500 -> Water_Temp.:101.19, PWM:35
16:27:42.054 -> Water_Temp.:101.19, PWM:35
16:27:42.850 -> Water_Temp.:101.19, PWM:35
16:27:43.607 -> Water_Temp.:101.19, PWM:35
16:27:44.395 -> Water_Temp.:101.19, PWM:35
16:27:45.189 -> Water_Temp.:101.19, PWM:35
16:27:46.734 -> Water_Temp.:101.19, PWM:35
16:27:47.525 -> Water_Temp.:101.19, PWM:35
16:27:48.315 -> Water_Temp.:101.19, PWM:36
16:27:49.835 -> Water_Temp.:101.19, PWM:35
16:27:50.630 -> Water_Temp.:101.19, PWM:35
16:27:51.388 -> Water_Temp.:101.19, PWM:35
16:27:52.180 -> Water_Temp.:101.19, PWM:35
16:27:52.975 -> Water_Temp.:101.19, PWM:36
16:27:53.732 -> Water_Temp.:101.19, PWM:43
16:27:54.530 -> Water_Temp.:101.19, PWM:43
16:27:56.082 -> Water_Temp.:101.19, PWM:43
16:27:58.394 -> Water_Temp.:101.19, PWM:34
16:27:59.187 -> Water_Temp.:101.19, PWM:33
16:27:59.947 -> Water_Temp.:101.19, PWM:33
16:28:01.499 -> Water_Temp.:101.19, PWM:33
16:28:03.088 -> Water_Temp.:101.19, PWM:35
16:28:04.615 -> Water_Temp.:101.19, PWM:36
16:28:05.408 -> Water_Temp.:101.30, PWM:38
16:28:06.173 -> Water_Temp.:101.19, PWM:38
16:28:06.968 -> Water_Temp.:101.19, PWM:38
16:28:09.314 -> Water_Temp.:101.19, PWM:38
16:28:10.871 -> Water_Temp.:101.19, PWM:38
16:28:11.635 -> Water_Temp.:101.30, PWM:40
16:28:12.429 -> Water_Temp.:101.30, PWM:40
16:28:13.191 -> Water_Temp.:101.30, PWM:40
16:28:13.987 -> Water_Temp.:101.30, PWM:40
16:28:15.534 -> Water_Temp.:101.30, PWM:40
16:28:17.088 -> Water_Temp.:101.30, PWM:40
16:28:19.434 -> Water_Temp.:101.30, PWM:40
16:28:20.991 -> Water_Temp.:101.30, PWM:39
16:28:21.785 -> Water_Temp.:101.30, PWM:40
16:28:22.543 -> Water_Temp.:101.30, PWM:40
16:28:23.342 -> Water_Temp.:101.30, PWM:39
16:28:24.103 -> Water_Temp.:101.30, PWM:40
16:28:24.895 -> Water_Temp.:101.30, PWM:39
16:28:25.654 -> Water_Temp.:101.30, PWM:39
16:28:27.214 -> Water_Temp.:101.30, PWM:38
16:28:27.973 -> Water_Temp.:101.30, PWM:38
16:28:28.768 -> Water_Temp.:101.30, PWM:39
16:28:29.558 -> Water_Temp.:101.30, PWM:38
16:28:30.317 -> Water_Temp.:101.30, PWM:39
16:28:31.111 -> Water_Temp.:101.19, PWM:38
16:28:31.906 -> Water_Temp.:101.19, PWM:38
16:28:33.426 -> Water_Temp.:101.30, PWM:38
16:28:35.771 -> Water_Temp.:101.19, PWM:38
16:28:36.536 -> Water_Temp.:101.30, PWM:38
16:28:37.334 -> Water_Temp.:101.30, PWM:39
16:28:39.645 -> Water_Temp.:101.30, PWM:38
16:28:40.443 -> Water_Temp.:102.99, PWM:38
16:28:41.202 -> Water_Temp.:105.01, PWM:38
16:28:41.999 -> Water_Temp.:102.65, PWM:38
16:28:42.792 -> Water_Temp.:102.20, PWM:38
16:28:43.553 -> Water_Temp.:101.97, PWM:39
16:28:44.351 -> Water_Temp.:101.86, PWM:38
16:28:45.871 -> Water_Temp.:101.75, PWM:39
16:28:49.734 -> Water_Temp.:101.52, PWM:38
16:28:50.526 -> Water_Temp.:101.52, PWM:38
16:28:51.318 -> Water_Temp.:101.52, PWM:39
16:28:52.075 -> Water_Temp.:101.41, PWM:39
16:28:52.867 -> Water_Temp.:101.41, PWM:38
16:28:53.661 -> Water_Temp.:101.41, PWM:38
16:28:54.424 -> Water_Temp.:101.41, PWM:38
16:28:55.220 -> Water_Temp.:101.41, PWM:38
16:28:56.772 -> Water_Temp.:101.41, PWM:38
16:28:57.534 -> Water_Temp.:101.41, PWM:39
16:28:58.328 -> Water_Temp.:101.41, PWM:38
16:28:59.088 -> Water_Temp.:101.41, PWM:39
16:28:59.880 -> Water_Temp.:101.41, PWM:38
16:29:02.195 -> Water_Temp.:102.99, PWM:38
16:29:02.986 -> Water_Temp.:105.12, PWM:39
16:29:03.747 -> Water_Temp.:102.65, PWM:38
16:29:04.541 -> Water_Temp.:102.20, PWM:38
16:29:06.094 -> Water_Temp.:101.86, PWM:39
16:29:06.883 -> Water_Temp.:101.75, PWM:38
16:29:07.648 -> Water_Temp.:101.75, PWM:39
16:29:09.201 -> Water_Temp.:101.64, PWM:39
16:29:09.997 -> Water_Temp.:101.64, PWM:39
16:29:11.551 -> Water_Temp.:101.52, PWM:39
16:29:12.316 -> Water_Temp.:101.52, PWM:38
16:29:13.897 -> Water_Temp.:101.52, PWM:38
16:29:16.208 -> Water_Temp.:101.52, PWM:39
16:29:18.519 -> Water_Temp.:101.52, PWM:38
16:29:19.313 -> Water_Temp.:101.52, PWM:38
16:29:20.073 -> Water_Temp.:101.41, PWM:38
16:29:20.872 -> Water_Temp.:101.41, PWM:39
16:29:21.632 -> Water_Temp.:101.41, PWM:39
16:29:22.428 -> Water_Temp.:101.41, PWM:38
16:29:23.982 -> Water_Temp.:101.41, PWM:38
16:29:25.539 -> Water_Temp.:101.41, PWM:38
16:29:26.333 -> Water_Temp.:101.41, PWM:38
16:29:27.096 -> Water_Temp.:101.41, PWM:38
16:29:28.618 -> Water_Temp.:101.41, PWM:38
16:29:30.941 -> Water_Temp.:101.52, PWM:39
16:29:31.737 -> Water_Temp.:101.52, PWM:39
16:29:32.499 -> Water_Temp.:101.52, PWM:38
16:29:34.087 -> Water_Temp.:101.52, PWM:39
16:29:35.609 -> Water_Temp.:101.52, PWM:39
16:29:36.400 -> Water_Temp.:101.52, PWM:38
16:29:37.192 -> Water_Temp.:103.10, PWM:39
16:29:37.948 -> Water_Temp.:105.24, PWM:38
16:29:39.533 -> Water_Temp.:102.42, PWM:38
16:29:40.292 -> Water_Temp.:102.20, PWM:38
16:29:41.088 -> Water_Temp.:102.09, PWM:38
16:29:41.845 -> Water_Temp.:101.97, PWM:38
16:29:42.640 -> Water_Temp.:101.86, PWM:39
16:29:43.433 -> Water_Temp.:101.75, PWM:38
16:29:45.741 -> Water_Temp.:101.64, PWM:38
16:29:46.533 -> Water_Temp.:101.64, PWM:39
16:29:47.323 -> Water_Temp.:101.64, PWM:38
16:29:48.877 -> Water_Temp.:101.64, PWM:38
16:29:49.636 -> Water_Temp.:101.64, PWM:39
16:29:51.192 -> Water_Temp.:101.64, PWM:39
16:29:51.986 -> Water_Temp.:101.52, PWM:39
16:29:52.777 -> Water_Temp.:101.52, PWM:38
16:29:54.294 -> Water_Temp.:101.52, PWM:38
16:29:55.086 -> Water_Temp.:101.52, PWM:38
16:29:55.878 -> Water_Temp.:101.52, PWM:38
16:29:56.640 -> Water_Temp.:101.52, PWM:38
16:29:57.430 -> Water_Temp.:101.52, PWM:38
16:29:58.192 -> Water_Temp.:101.52, PWM:39
16:29:58.987 -> Water_Temp.:101.52, PWM:39
16:29:59.781 -> Water_Temp.:101.52, PWM:39
16:30:00.538 -> Water_Temp.:101.52, PWM:38
16:30:02.861 -> Water_Temp.:101.52, PWM:38
16:30:04.411 -> Water_Temp.:101.52, PWM:38
16:30:05.967 -> Water_Temp.:101.52, PWM:38
16:30:08.284 -> Water_Temp.:101.52, PWM:39
16:30:09.075 -> Water_Temp.:101.52, PWM:38
16:30:09.832 -> Water_Temp.:101.52, PWM:38
16:30:10.626 -> Water_Temp.:101.52, PWM:38
16:30:12.903 -> Water_Temp.:101.52, PWM:38
16:30:13.699 -> Water_Temp.:101.52, PWM:38
16:30:14.458 -> Water_Temp.:101.52, PWM:38
16:30:15.247 -> Water_Temp.:101.52, PWM:38
16:30:16.772 -> Water_Temp.:101.52, PWM:39
16:30:17.568 -> Water_Temp.:101.52, PWM:39
16:30:18.332 -> Water_Temp.:101.52, PWM:38
16:30:19.123 -> Water_Temp.:101.52, PWM:39
16:30:19.916 -> Water_Temp.:101.52, PWM:39
16:30:21.438 -> Water_Temp.:101.52, PWM:38
16:30:22.229 -> Water_Temp.:101.52, PWM:38
16:30:26.857 -> Water_Temp.:101.52, PWM:39
16:30:27.649 -> Water_Temp.:101.52, PWM:39
16:30:28.412 -> Water_Temp.:101.52, PWM:38
16:30:29.208 -> Water_Temp.:101.52, PWM:38
16:30:31.518 -> Water_Temp.:101.64, PWM:39
16:30:33.105 -> Water_Temp.:101.52, PWM:38
16:30:34.624 -> Water_Temp.:101.52, PWM:39
16:30:35.420 -> Water_Temp.:101.52, PWM:38
16:30:36.178 -> Water_Temp.:101.52, PWM:38
16:30:36.975 -> Water_Temp.:101.52, PWM:38
16:30:37.735 -> Water_Temp.:101.52, PWM:38
16:30:38.528 -> Water_Temp.:101.52, PWM:38
16:30:39.319 -> Water_Temp.:101.52, PWM:39
16:30:42.360 -> Water_Temp.:101.52, PWM:39
16:30:43.158 -> Water_Temp.:101.52, PWM:38
16:30:43.918 -> Water_Temp.:101.52, PWM:38
16:30:44.709 -> Water_Temp.:101.52, PWM:38
16:30:46.229 -> Water_Temp.:101.64, PWM:39
16:30:47.778 -> Water_Temp.:101.52, PWM:38
16:30:48.570 -> Water_Temp.:101.52, PWM:38
16:30:50.124 -> Water_Temp.:101.52, PWM:39
16:30:50.887 -> Water_Temp.:101.52, PWM:38
16:30:52.444 -> Water_Temp.:101.52, PWM:39
16:30:54.001 -> Water_Temp.:101.52, PWM:39
16:30:54.795 -> Water_Temp.:101.52, PWM:38
16:30:57.082 -> Water_Temp.:101.52, PWM:39
16:30:59.439 -> Water_Temp.:101.52, PWM:38
16:31:00.201 -> Water_Temp.:101.52, PWM:38
16:31:01.000 -> Water_Temp.:101.52, PWM:38
16:31:04.903 -> Water_Temp.:101.52, PWM:38
16:31:05.661 -> Water_Temp.:101.52, PWM:38
16:31:08.796 -> Water_Temp.:101.52, PWM:38
16:31:09.556 -> Water_Temp.:101.52, PWM:38
16:31:10.349 -> Water_Temp.:101.52, PWM:38
16:31:11.110 -> Water_Temp.:101.52, PWM:39
16:31:11.906 -> Water_Temp.:101.52, PWM:38
16:31:12.703 -> Water_Temp.:101.52, PWM:39
16:31:13.467 -> Water_Temp.:101.52, PWM:38
16:31:15.020 -> Water_Temp.:101.52, PWM:38
16:31:18.094 -> Water_Temp.:101.52, PWM:38
16:31:20.409 -> Water_Temp.:101.52, PWM:38
16:31:21.201 -> Water_Temp.:101.52, PWM:39
16:31:25.062 -> Water_Temp.:101.52, PWM:38
16:31:25.819 -> Water_Temp.:101.52, PWM:39
16:31:26.615 -> Water_Temp.:101.52, PWM:38
16:31:28.171 -> Water_Temp.:101.41, PWM:39
16:31:32.756 -> Water_Temp.:101.41, PWM:38
16:31:33.549 -> Water_Temp.:101.41, PWM:38
16:31:34.339 -> Water_Temp.:101.41, PWM:38
16:31:35.098 -> Water_Temp.:101.41, PWM:38
16:31:36.649 -> Water_Temp.:101.52, PWM:38
16:31:39.776 -> Water_Temp.:103.21, PWM:38
16:31:40.537 -> Water_Temp.:105.24, PWM:38
16:31:41.328 -> Water_Temp.:102.87, PWM:38
16:31:42.120 -> Water_Temp.:102.42, PWM:39
16:31:43.670 -> Water_Temp.:102.09, PWM:38
16:31:45.219 -> Water_Temp.:101.97, PWM:39
16:31:46.013 -> Water_Temp.:101.97, PWM:39
16:31:46.773 -> Water_Temp.:101.86, PWM:38
16:31:47.568 -> Water_Temp.:101.86, PWM:39
16:31:48.328 -> Water_Temp.:101.86, PWM:39
16:31:49.120 -> Water_Temp.:101.86, PWM:38
16:31:49.881 -> Water_Temp.:101.75, PWM:38
16:31:51.468 -> Water_Temp.:101.75, PWM:38
16:31:52.984 -> Water_Temp.:101.75, PWM:39
16:31:53.778 -> Water_Temp.:101.75, PWM:38
16:31:56.127 -> Water_Temp.:101.64, PWM:38
16:31:56.887 -> Water_Temp.:101.64, PWM:39
16:31:57.677 -> Water_Temp.:101.64, PWM:39
16:31:58.468 -> Water_Temp.:101.64, PWM:39
16:31:59.223 -> Water_Temp.:101.64, PWM:38
16:32:00.783 -> Water_Temp.:101.64, PWM:38
16:32:01.574 -> Water_Temp.:101.64, PWM:39
16:32:02.363 -> Water_Temp.:101.64, PWM:38
16:32:03.909 -> Water_Temp.:101.64, PWM:39
16:32:04.667 -> Water_Temp.:101.64, PWM:38
16:32:05.463 -> Water_Temp.:101.64, PWM:39
16:32:06.227 -> Water_Temp.:101.64, PWM:38
16:32:07.022 -> Water_Temp.:101.64, PWM:38
16:32:07.813 -> Water_Temp.:101.64, PWM:39
16:32:08.570 -> Water_Temp.:101.64, PWM:38
16:32:10.125 -> Water_Temp.:101.64, PWM:38
16:32:10.923 -> Water_Temp.:101.64, PWM:39
16:32:11.681 -> Water_Temp.:101.64, PWM:39
16:32:12.474 -> Water_Temp.:101.64, PWM:38
16:32:13.235 -> Water_Temp.:101.64, PWM:39
16:32:14.818 -> Water_Temp.:101.64, PWM:38
16:32:15.577 -> Water_Temp.:101.64, PWM:39
16:32:17.130 -> Water_Temp.:101.64, PWM:39
16:32:17.927 -> Water_Temp.:101.64, PWM:38
16:32:18.687 -> Water_Temp.:101.64, PWM:39
16:32:20.240 -> Water_Temp.:101.64, PWM:39
16:32:21.030 -> Water_Temp.:101.64, PWM:38
16:32:21.791 -> Water_Temp.:101.64, PWM:39
16:32:22.584 -> Water_Temp.:101.64, PWM:39
16:32:23.380 -> Water_Temp.:101.64, PWM:38
16:32:24.138 -> Water_Temp.:101.64, PWM:39
16:32:26.460 -> Water_Temp.:101.64, PWM:38
16:32:30.317 -> Water_Temp.:101.64, PWM:39
16:32:31.079 -> Water_Temp.:101.64, PWM:38
16:32:31.875 -> Water_Temp.:101.64, PWM:38
16:32:34.191 -> Water_Temp.:101.64, PWM:38
16:32:35.776 -> Water_Temp.:101.64, PWM:38
16:32:38.089 -> Water_Temp.:101.64, PWM:39
16:32:41.198 -> Water_Temp.:101.64, PWM:38
16:32:41.955 -> Water_Temp.:101.64, PWM:38
16:32:42.746 -> Water_Temp.:101.64, PWM:39
16:32:44.297 -> Water_Temp.:101.64, PWM:38
16:32:45.055 -> Water_Temp.:101.64, PWM:38
16:32:45.853 -> Water_Temp.:101.64, PWM:38
16:32:46.615 -> Water_Temp.:101.64, PWM:39
16:32:47.410 -> Water_Temp.:101.64, PWM:38
16:32:48.965 -> Water_Temp.:101.64, PWM:38
16:32:49.727 -> Water_Temp.:101.64, PWM:38
16:32:50.522 -> Water_Temp.:101.64, PWM:38
16:32:51.281 -> Water_Temp.:101.75, PWM:39
16:32:52.837 -> Water_Temp.:101.75, PWM:38
16:32:53.630 -> Water_Temp.:101.75, PWM:38
16:32:54.421 -> Water_Temp.:101.75, PWM:38
16:32:55.978 -> Water_Temp.:101.75, PWM:38
16:32:57.496 -> Water_Temp.:101.75, PWM:39
16:32:58.290 -> Water_Temp.:101.75, PWM:39
16:32:59.086 -> Water_Temp.:101.86, PWM:38
16:32:59.847 -> Water_Temp.:101.86, PWM:38
16:33:00.611 -> Water_Temp.:101.86, PWM:38
16:33:01.406 -> Water_Temp.:101.86, PWM:38
16:33:02.196 -> Water_Temp.:101.86, PWM:39
16:33:02.955 -> Water_Temp.:101.86, PWM:39
16:33:03.745 -> Water_Temp.:101.86, PWM:39
16:33:05.295 -> Water_Temp.:101.86, PWM:39
16:33:06.849 -> Water_Temp.:101.86, PWM:38
16:33:07.611 -> Water_Temp.:101.86, PWM:38
16:33:09.164 -> Water_Temp.:101.86, PWM:38
16:33:12.238 -> Water_Temp.:101.97, PWM:38
16:33:15.348 -> Water_Temp.:101.97, PWM:39
16:33:17.653 -> Water_Temp.:101.97, PWM:39
16:33:18.412 -> Water_Temp.:101.97, PWM:38
16:33:19.208 -> Water_Temp.:101.97, PWM:39
16:33:20.762 -> Water_Temp.:101.97, PWM:39
16:33:23.074 -> Water_Temp.:101.97, PWM:38
16:33:25.388 -> Water_Temp.:101.97, PWM:38
16:33:27.736 -> Water_Temp.:101.97, PWM:38
16:33:28.497 -> Water_Temp.:101.97, PWM:38
16:33:29.290 -> Water_Temp.:101.97, PWM:38
16:33:30.052 -> Water_Temp.:101.97, PWM:39
16:33:33.127 -> Water_Temp.:101.97, PWM:38
16:33:33.916 -> Water_Temp.:101.97, PWM:39
16:33:34.709 -> Water_Temp.:101.97, PWM:38
16:33:35.469 -> Water_Temp.:101.97, PWM:39
16:33:36.267 -> Water_Temp.:102.09, PWM:38
16:33:37.820 -> Water_Temp.:102.09, PWM:38
16:33:39.374 -> Water_Temp.:102.09, PWM:39
16:33:40.168 -> Water_Temp.:102.09, PWM:38
16:33:40.926 -> Water_Temp.:102.09, PWM:38
16:33:43.275 -> Water_Temp.:102.09, PWM:39
16:33:44.037 -> Water_Temp.:102.09, PWM:38
16:33:44.831 -> Water_Temp.:102.09, PWM:38
16:33:48.700 -> Water_Temp.:102.09, PWM:38
16:33:49.496 -> Water_Temp.:102.09, PWM:38
16:33:50.255 -> Water_Temp.:102.09, PWM:38
16:33:51.049 -> Water_Temp.:101.97, PWM:39
16:33:51.835 -> Water_Temp.:101.97, PWM:38
16:33:52.595 -> Water_Temp.:101.97, PWM:38
16:33:53.387 -> Water_Temp.:101.97, PWM:39
16:33:54.146 -> Water_Temp.:101.97, PWM:38
16:33:55.732 -> Water_Temp.:101.97, PWM:38
16:33:56.494 -> Water_Temp.:101.97, PWM:39
16:33:57.291 -> Water_Temp.:101.97, PWM:38
16:33:58.052 -> Water_Temp.:101.97, PWM:39
16:34:00.393 -> Water_Temp.:101.97, PWM:38
16:34:01.148 -> Water_Temp.:101.97, PWM:39
16:34:01.940 -> Water_Temp.:101.97, PWM:38
16:34:02.733 -> Water_Temp.:101.97, PWM:38
16:34:05.808 -> Water_Temp.:102.09, PWM:38
16:34:07.364 -> Water_Temp.:102.09, PWM:38
16:34:08.924 -> Water_Temp.:102.09, PWM:38
16:34:10.476 -> Water_Temp.:102.09, PWM:38
16:34:11.267 -> Water_Temp.:102.09, PWM:38
16:34:12.023 -> Water_Temp.:102.20, PWM:39
16:34:15.124 -> Water_Temp.:102.20, PWM:38
16:34:15.884 -> Water_Temp.:102.20, PWM:39
16:34:16.673 -> Water_Temp.:102.20, PWM:38
16:34:18.233 -> Water_Temp.:102.20, PWM:38
16:34:19.025 -> Water_Temp.:102.20, PWM:38
16:34:22.856 -> Water_Temp.:102.20, PWM:38
16:34:23.648 -> Water_Temp.:102.20, PWM:38
16:34:25.165 -> Water_Temp.:102.20, PWM:38
16:34:25.959 -> Water_Temp.:102.20, PWM:38
16:34:26.715 -> Water_Temp.:102.20, PWM:38
16:34:29.028 -> Water_Temp.:102.20, PWM:39
16:34:29.820 -> Water_Temp.:102.20, PWM:38
16:34:31.371 -> Water_Temp.:102.20, PWM:38
16:34:33.716 -> Water_Temp.:102.20, PWM:38
16:34:34.472 -> Water_Temp.:102.20, PWM:39
16:34:35.264 -> Water_Temp.:102.20, PWM:39
16:34:36.817 -> Water_Temp.:102.20, PWM:39
16:34:37.611 -> Water_Temp.:102.20, PWM:38
16:34:39.159 -> Water_Temp.:102.20, PWM:38
16:34:39.954 -> Water_Temp.:102.20, PWM:39
16:34:42.271 -> Water_Temp.:102.20, PWM:38
16:34:43.062 -> Water_Temp.:102.20, PWM:38
16:34:43.851 -> Water_Temp.:102.20, PWM:38
16:34:45.395 -> Water_Temp.:102.20, PWM:38
16:34:46.155 -> Water_Temp.:102.20, PWM:38
16:34:46.949 -> Water_Temp.:102.20, PWM:39
16:34:47.712 -> Water_Temp.:102.20, PWM:38
16:34:50.051 -> Water_Temp.:102.20, PWM:38
16:34:50.844 -> Water_Temp.:102.20, PWM:38
16:34:53.147 -> Water_Temp.:102.20, PWM:38
16:34:55.495 -> Water_Temp.:102.20, PWM:38
16:34:56.260 -> Water_Temp.:102.20, PWM:38
16:34:57.058 -> Water_Temp.:102.20, PWM:38
16:34:59.334 -> Water_Temp.:102.20, PWM:39
16:35:00.927 -> Water_Temp.:102.09, PWM:38
16:35:02.450 -> Water_Temp.:102.20, PWM:38
16:35:04.002 -> Water_Temp.:102.20, PWM:39
16:35:05.552 -> Water_Temp.:102.20, PWM:39
16:35:07.107 -> Water_Temp.:102.20, PWM:39
16:35:08.690 -> Water_Temp.:102.20, PWM:39
16:35:11.007 -> Water_Temp.:102.20, PWM:38
16:35:11.800 -> Water_Temp.:102.20, PWM:38
16:35:12.588 -> Water_Temp.:102.20, PWM:38
16:35:14.905 -> Water_Temp.:102.20, PWM:39
16:35:15.666 -> Water_Temp.:102.20, PWM:38
16:35:16.465 -> Water_Temp.:102.31, PWM:39
16:35:17.989 -> Water_Temp.:102.31, PWM:38
16:35:19.540 -> Water_Temp.:102.31, PWM:38
16:35:20.299 -> Water_Temp.:102.31, PWM:39
16:35:21.096 -> Water_Temp.:102.31, PWM:38
16:35:24.964 -> Water_Temp.:102.31, PWM:38
16:35:25.723 -> Water_Temp.:102.31, PWM:39
16:35:26.517 -> Water_Temp.:102.31, PWM:39
16:35:27.311 -> Water_Temp.:102.31, PWM:38
16:35:28.830 -> Water_Temp.:102.20, PWM:38
16:35:29.621 -> Water_Temp.:102.31, PWM:39
16:35:30.415 -> Water_Temp.:102.20, PWM:38
16:35:31.936 -> Water_Temp.:102.31, PWM:39
16:35:32.729 -> Water_Temp.:102.31, PWM:38
16:35:34.238 -> Water_Temp.:102.20, PWM:38
16:35:35.825 -> Water_Temp.:102.31, PWM:39
16:35:36.579 -> Water_Temp.:102.31, PWM:39
16:35:37.375 -> Water_Temp.:102.31, PWM:39
16:35:38.893 -> Water_Temp.:102.31, PWM:39
16:35:39.684 -> Water_Temp.:102.31, PWM:38
16:35:40.481 -> Water_Temp.:102.31, PWM:39
16:35:42.032 -> Water_Temp.:102.31, PWM:39
16:35:42.792 -> Water_Temp.:102.31, PWM:38
16:35:43.590 -> Water_Temp.:102.31, PWM:38
16:35:44.378 -> Water_Temp.:102.31, PWM:39
16:35:45.900 -> Water_Temp.:102.31, PWM:38
16:35:47.457 -> Water_Temp.:102.31, PWM:38
16:35:48.249 -> Water_Temp.:102.31, PWM:38
16:35:49.010 -> Water_Temp.:102.31, PWM:39
16:35:49.808 -> Water_Temp.:102.31, PWM:38
16:35:50.565 -> Water_Temp.:102.31, PWM:38
16:35:51.359 -> Water_Temp.:102.31, PWM:38
16:35:52.146 -> Water_Temp.:102.31, PWM:39
16:35:52.902 -> Water_Temp.:102.31, PWM:38
16:35:54.489 -> Water_Temp.:102.31, PWM:39
16:35:55.250 -> Water_Temp.:102.31, PWM:38
16:35:56.046 -> Water_Temp.:102.31, PWM:38
16:35:57.602 -> Water_Temp.:102.42, PWM:38
16:35:58.358 -> Water_Temp.:102.42, PWM:38
16:35:59.903 -> Water_Temp.:102.31, PWM:39
16:36:00.698 -> Water_Temp.:102.42, PWM:38
16:36:02.252 -> Water_Temp.:102.42, PWM:39
16:36:03.008 -> Water_Temp.:102.42, PWM:38
16:36:04.556 -> Water_Temp.:102.42, PWM:38
16:36:05.345 -> Water_Temp.:102.42, PWM:38
16:36:07.670 -> Water_Temp.:102.42, PWM:39
16:36:08.464 -> Water_Temp.:102.42, PWM:39
16:36:09.222 -> Water_Temp.:102.42, PWM:39
16:36:10.015 -> Water_Temp.:102.42, PWM:39
16:36:12.362 -> Water_Temp.:102.42, PWM:38
16:36:13.153 -> Water_Temp.:102.42, PWM:38
16:36:13.910 -> Water_Temp.:102.42, PWM:39
16:36:14.705 -> Water_Temp.:102.42, PWM:39
16:36:16.258 -> Water_Temp.:102.42, PWM:38
16:36:17.018 -> Water_Temp.:102.42, PWM:39
16:36:17.811 -> Water_Temp.:102.54, PWM:39
16:36:18.574 -> Water_Temp.:102.42, PWM:39
16:36:19.369 -> Water_Temp.:102.54, PWM:38
16:36:20.130 -> Water_Temp.:102.42, PWM:38
16:36:20.927 -> Water_Temp.:102.42, PWM:39
16:36:21.686 -> Water_Temp.:102.42, PWM:39
16:36:22.481 -> Water_Temp.:102.42, PWM:38
16:36:23.242 -> Water_Temp.:102.42, PWM:39
16:36:26.315 -> Water_Temp.:102.54, PWM:38
16:36:27.110 -> Water_Temp.:102.42, PWM:38
16:36:29.390 -> Water_Temp.:102.42, PWM:39
16:36:30.181 -> Water_Temp.:102.42, PWM:38
16:36:30.946 -> Water_Temp.:102.42, PWM:38
16:36:31.745 -> Water_Temp.:102.42, PWM:39
16:36:32.503 -> Water_Temp.:102.42, PWM:39
16:36:33.298 -> Water_Temp.:102.42, PWM:39
16:36:34.058 -> Water_Temp.:102.42, PWM:39
16:36:34.853 -> Water_Temp.:102.42, PWM:38
16:36:35.650 -> Water_Temp.:102.42, PWM:39
16:36:37.203 -> Water_Temp.:102.42, PWM:39
16:36:38.754 -> Water_Temp.:102.42, PWM:38
16:36:39.548 -> Water_Temp.:102.42, PWM:39
16:36:41.863 -> Water_Temp.:102.42, PWM:39
16:36:42.624 -> Water_Temp.:102.42, PWM:39
16:36:43.420 -> Water_Temp.:102.54, PWM:39
16:36:44.175 -> Water_Temp.:102.42, PWM:38
16:36:44.967 -> Water_Temp.:102.54, PWM:38
16:36:47.279 -> Water_Temp.:102.54, PWM:38
16:36:48.829 -> Water_Temp.:102.54, PWM:39
16:36:49.624 -> Water_Temp.:102.54, PWM:38
16:36:50.383 -> Water_Temp.:102.54, PWM:38
16:36:51.179 -> Water_Temp.:102.54, PWM:39
16:36:51.971 -> Water_Temp.:102.54, PWM:39
16:36:52.730 -> Water_Temp.:102.54, PWM:39
16:36:53.526 -> Water_Temp.:102.54, PWM:38
16:36:54.312 -> Water_Temp.:102.54, PWM:38
16:36:57.352 -> Water_Temp.:102.54, PWM:38
16:36:58.147 -> Water_Temp.:102.54, PWM:38
16:36:58.910 -> Water_Temp.:102.54, PWM:39
16:36:59.706 -> Water_Temp.:102.54, PWM:38
16:37:00.463 -> Water_Temp.:102.54, PWM:38
16:37:01.256 -> Water_Temp.:102.54, PWM:38
16:37:02.020 -> Water_Temp.:102.54, PWM:38
16:37:03.570 -> Water_Temp.:102.54, PWM:38
16:37:04.363 -> Water_Temp.:102.54, PWM:38
16:37:05.127 -> Water_Temp.:102.54, PWM:39
16:37:05.923 -> Water_Temp.:102.54, PWM:38
16:37:06.681 -> Water_Temp.:102.54, PWM:38
16:37:07.475 -> Water_Temp.:102.54, PWM:39
16:37:08.234 -> Water_Temp.:102.54, PWM:39
16:37:09.822 -> Water_Temp.:102.54, PWM:38
16:37:11.335 -> Water_Temp.:102.54, PWM:38
16:37:12.926 -> Water_Temp.:102.54, PWM:39
16:37:13.687 -> Water_Temp.:102.42, PWM:38
16:37:14.479 -> Water_Temp.:102.54, PWM:38
16:37:15.238 -> Water_Temp.:102.54, PWM:39
16:37:16.037 -> Water_Temp.:102.54, PWM:38
16:37:16.793 -> Water_Temp.:102.54, PWM:38
16:37:18.382 -> Water_Temp.:102.54, PWM:38
16:37:19.144 -> Water_Temp.:102.54, PWM:38
16:37:20.696 -> Water_Temp.:102.54, PWM:38
16:37:22.253 -> Water_Temp.:102.54, PWM:38
16:37:23.812 -> Water_Temp.:102.54, PWM:39
16:37:25.365 -> Water_Temp.:102.54, PWM:38
16:37:26.130 -> Water_Temp.:102.54, PWM:38
16:37:27.723 -> Water_Temp.:102.54, PWM:38
16:37:28.483 -> Water_Temp.:102.54, PWM:39
16:37:30.033 -> Water_Temp.:102.54, PWM:38
16:37:30.830 -> Water_Temp.:102.54, PWM:38
16:37:31.586 -> Water_Temp.:102.54, PWM:38
16:37:33.166 -> Water_Temp.:102.54, PWM:38
16:37:33.925 -> Water_Temp.:102.54, PWM:39
16:37:34.717 -> Water_Temp.:102.54, PWM:38
16:37:35.474 -> Water_Temp.:102.54, PWM:38
16:37:37.066 -> Water_Temp.:102.54, PWM:38
16:37:37.821 -> Water_Temp.:102.54, PWM:39
16:37:42.445 -> Water_Temp.:102.54, PWM:39
16:37:43.235 -> Water_Temp.:102.54, PWM:39
16:37:43.996 -> Water_Temp.:102.54, PWM:38
16:37:46.335 -> Water_Temp.:102.54, PWM:38
16:37:47.096 -> Water_Temp.:102.54, PWM:38
16:37:47.887 -> Water_Temp.:102.54, PWM:38
16:37:48.685 -> Water_Temp.:102.54, PWM:38
16:37:50.200 -> Water_Temp.:102.54, PWM:38
16:37:52.510 -> Water_Temp.:102.54, PWM:39
16:37:53.302 -> Water_Temp.:102.54, PWM:38
16:37:54.853 -> Water_Temp.:102.54, PWM:39
16:37:55.607 -> Water_Temp.:102.54, PWM:38
16:37:56.406 -> Water_Temp.:102.54, PWM:38
16:37:57.167 -> Water_Temp.:102.54, PWM:39
16:37:57.956 -> Water_Temp.:102.54, PWM:38
16:37:59.504 -> Water_Temp.:102.54, PWM:39
16:38:01.061 -> Water_Temp.:102.54, PWM:38
16:38:01.859 -> Water_Temp.:102.54, PWM:38
16:38:02.650 -> Water_Temp.:102.54, PWM:38
16:38:04.169 -> Water_Temp.:102.54, PWM:38
16:38:04.962 -> Water_Temp.:102.42, PWM:39
16:38:05.753 -> Water_Temp.:102.42, PWM:38
16:38:06.513 -> Water_Temp.:102.42, PWM:38
16:38:07.307 -> Water_Temp.:102.54, PWM:38
16:38:10.390 -> Water_Temp.:102.54, PWM:39
16:38:11.152 -> Water_Temp.:102.54, PWM:38
16:38:11.948 -> Water_Temp.:102.54, PWM:38
16:38:12.711 -> Water_Temp.:102.54, PWM:38
16:38:13.505 -> Water_Temp.:102.54, PWM:38
16:38:15.852 -> Water_Temp.:102.54, PWM:39
16:38:17.409 -> Water_Temp.:102.54, PWM:38
16:38:18.170 -> Water_Temp.:102.54, PWM:38
16:38:18.968 -> Water_Temp.:102.54, PWM:39
16:38:19.759 -> Water_Temp.:102.54, PWM:38
16:38:22.036 -> Water_Temp.:102.54, PWM:38
16:38:23.593 -> Water_Temp.:102.54, PWM:39
16:38:24.384 -> Water_Temp.:102.54, PWM:38
16:38:25.143 -> Water_Temp.:102.54, PWM:39
16:38:25.939 -> Water_Temp.:102.54, PWM:38
16:38:27.489 -> Water_Temp.:102.54, PWM:38
16:38:28.285 -> Water_Temp.:102.54, PWM:38
16:38:29.046 -> Water_Temp.:102.54, PWM:38
16:38:29.842 -> Water_Temp.:102.54, PWM:39
16:38:32.154 -> Water_Temp.:102.54, PWM:39
16:38:33.710 -> Water_Temp.:102.42, PWM:38
16:38:36.021 -> Water_Temp.:102.42, PWM:38
16:38:37.568 -> Water_Temp.:102.42, PWM:39
16:38:38.362 -> Water_Temp.:102.42, PWM:38
16:38:39.917 -> Water_Temp.:102.42, PWM:38
16:38:40.673 -> Water_Temp.:102.42, PWM:38
16:38:41.465 -> Water_Temp.:102.42, PWM:39
16:38:42.260 -> Water_Temp.:102.42, PWM:38
16:38:43.018 -> Water_Temp.:102.42, PWM:38
16:38:44.570 -> Water_Temp.:102.42, PWM:39
16:38:45.361 -> Water_Temp.:102.42, PWM:38
16:38:46.124 -> Water_Temp.:102.42, PWM:38
16:38:46.920 -> Water_Temp.:102.42, PWM:38
16:38:47.680 -> Water_Temp.:102.42, PWM:38
16:38:48.475 -> Water_Temp.:102.42, PWM:39
16:38:50.032 -> Water_Temp.:102.54, PWM:39
16:38:50.826 -> Water_Temp.:102.54, PWM:39
16:38:51.582 -> Water_Temp.:102.54, PWM:38
16:38:53.136 -> Water_Temp.:102.54, PWM:39
16:38:55.487 -> Water_Temp.:102.65, PWM:39
16:38:56.249 -> Water_Temp.:102.65, PWM:38
16:38:57.043 -> Water_Temp.:102.65, PWM:38
16:38:57.802 -> Water_Temp.:102.65, PWM:38
16:38:58.597 -> Water_Temp.:102.65, PWM:39
16:38:59.356 -> Water_Temp.:102.65, PWM:38
16:39:00.144 -> Water_Temp.:102.65, PWM:39
16:39:00.939 -> Water_Temp.:102.65, PWM:38
16:39:01.728 -> Water_Temp.:102.65, PWM:39
16:39:02.486 -> Water_Temp.:102.65, PWM:38
16:39:03.278 -> Water_Temp.:102.65, PWM:39
16:39:04.039 -> Water_Temp.:102.65, PWM:38
16:39:06.382 -> Water_Temp.:102.65, PWM:39
16:39:07.141 -> Water_Temp.:102.65, PWM:38
16:39:07.941 -> Water_Temp.:102.65, PWM:38
16:39:08.698 -> Water_Temp.:102.65, PWM:38
16:39:10.250 -> Water_Temp.:102.65, PWM:38
16:39:11.011 -> Water_Temp.:102.65, PWM:38
16:39:11.806 -> Water_Temp.:102.76, PWM:39
16:39:13.326 -> Water_Temp.:102.65, PWM:38
16:39:14.116 -> Water_Temp.:102.65, PWM:39
16:39:14.908 -> Water_Temp.:102.65, PWM:39
16:39:15.663 -> Water_Temp.:102.65, PWM:39
16:39:17.253 -> Water_Temp.:102.65, PWM:38
16:39:18.009 -> Water_Temp.:102.65, PWM:38
16:39:18.808 -> Water_Temp.:102.65, PWM:38
16:39:19.570 -> Water_Temp.:102.65, PWM:39
16:39:21.127 -> Water_Temp.:102.65, PWM:38
16:39:21.883 -> Water_Temp.:102.65, PWM:38
16:39:24.234 -> Water_Temp.:102.65, PWM:38
16:39:25.790 -> Water_Temp.:102.65, PWM:38
16:39:28.137 -> Water_Temp.:102.65, PWM:38
16:39:30.475 -> Water_Temp.:102.76, PWM:38
16:39:31.271 -> Water_Temp.:102.76, PWM:39
16:39:32.026 -> Water_Temp.:102.76, PWM:38
16:39:33.584 -> Water_Temp.:102.76, PWM:38
16:39:35.893 -> Water_Temp.:102.76, PWM:39
16:39:38.209 -> Water_Temp.:102.76, PWM:39
16:39:39.000 -> Water_Temp.:102.76, PWM:38
16:39:39.785 -> Water_Temp.:102.76, PWM:38
16:39:40.547 -> Water_Temp.:102.76, PWM:39
16:39:41.345 -> Water_Temp.:102.76, PWM:38
16:39:43.656 -> Water_Temp.:102.76, PWM:39
16:39:44.452 -> Water_Temp.:102.76, PWM:38
16:39:45.213 -> Water_Temp.:102.76, PWM:39
16:39:46.011 -> Water_Temp.:102.76, PWM:39
16:39:46.794 -> Water_Temp.:102.76, PWM:38
16:39:47.548 -> Water_Temp.:102.76, PWM:38
16:39:48.334 -> Water_Temp.:102.76, PWM:38
16:39:50.648 -> Water_Temp.:102.76, PWM:38
16:39:51.430 -> Water_Temp.:102.76, PWM:39
16:39:52.220 -> Water_Temp.:102.76, PWM:38
16:39:53.738 -> Water_Temp.:102.76, PWM:39
16:39:54.526 -> Water_Temp.:102.76, PWM:38
16:39:55.310 -> Water_Temp.:102.87, PWM:38
16:39:56.102 -> Water_Temp.:102.87, PWM:39
16:39:56.854 -> Water_Temp.:102.76, PWM:38
16:39:58.437 -> Water_Temp.:102.87, PWM:39
16:40:01.501 -> Water_Temp.:102.76, PWM:39
16:40:03.047 -> Water_Temp.:102.76, PWM:38
16:40:04.628 -> Water_Temp.:102.76, PWM:38
16:40:05.412 -> Water_Temp.:102.76, PWM:39
16:40:06.174 -> Water_Temp.:102.76, PWM:38
16:40:07.722 -> Water_Temp.:102.76, PWM:39
16:40:09.266 -> Water_Temp.:102.87, PWM:38
16:40:10.059 -> Water_Temp.:102.87, PWM:38
16:40:10.850 -> Water_Temp.:102.87, PWM:38
16:40:11.601 -> Water_Temp.:102.87, PWM:39
16:40:12.388 -> Water_Temp.:102.87, PWM:38
16:40:13.178 -> Water_Temp.:102.87, PWM:38
16:40:13.975 -> Water_Temp.:102.87, PWM:38
16:40:16.255 -> Water_Temp.:102.99, PWM:38
16:40:17.046 -> Water_Temp.:102.99, PWM:38
16:40:17.807 -> Water_Temp.:102.99, PWM:38
16:40:18.601 -> Water_Temp.:102.99, PWM:38
16:40:19.362 -> Water_Temp.:102.99, PWM:38
16:40:21.673 -> Water_Temp.:102.99, PWM:39
16:40:24.022 -> Water_Temp.:102.99, PWM:39
16:40:25.570 -> Water_Temp.:102.99, PWM:38
16:40:26.364 -> Water_Temp.:102.99, PWM:39
16:40:30.205 -> Water_Temp.:102.99, PWM:38
16:40:31.759 -> Water_Temp.:102.99, PWM:39
16:40:34.859 -> Water_Temp.:102.99, PWM:38
16:40:37.968 -> Water_Temp.:102.99, PWM:39
16:40:38.730 -> Water_Temp.:102.99, PWM:38
16:40:39.523 -> Water_Temp.:102.99, PWM:38
16:40:40.280 -> Water_Temp.:102.99, PWM:39
16:40:41.070 -> Water_Temp.:102.99, PWM:38
16:40:43.422 -> Water_Temp.:102.99, PWM:38
16:40:44.183 -> Water_Temp.:102.99, PWM:38
16:40:44.979 -> Water_Temp.:102.99, PWM:39
16:40:46.531 -> Water_Temp.:102.99, PWM:38
16:40:48.079 -> Water_Temp.:102.99, PWM:38
16:40:50.390 -> Water_Temp.:102.99, PWM:39
16:40:51.147 -> Water_Temp.:102.99, PWM:38
16:40:51.943 -> Water_Temp.:102.99, PWM:39
16:40:52.737 -> Water_Temp.:102.99, PWM:38
16:40:53.496 -> Water_Temp.:102.99, PWM:39
16:40:54.286 -> Water_Temp.:102.99, PWM:39
16:40:55.078 -> Water_Temp.:103.10, PWM:39
16:40:55.841 -> Water_Temp.:103.10, PWM:38
16:40:56.638 -> Water_Temp.:103.10, PWM:39
16:40:58.190 -> Water_Temp.:103.10, PWM:38
16:40:58.953 -> Water_Temp.:103.10, PWM:39
16:41:00.506 -> Water_Temp.:103.10, PWM:39
16:41:01.301 -> Water_Temp.:103.10, PWM:38
16:41:02.857 -> Water_Temp.:103.10, PWM:39
16:41:03.621 -> Water_Temp.:103.10, PWM:39
16:41:05.935 -> Water_Temp.:103.10, PWM:38
16:41:06.729 -> Water_Temp.:103.10, PWM:39
16:41:09.037 -> Water_Temp.:103.10, PWM:39
16:41:09.832 -> Water_Temp.:103.10, PWM:39
16:41:10.591 -> Water_Temp.:103.10, PWM:38
16:41:11.384 -> Water_Temp.:103.10, PWM:38
16:41:13.659 -> Water_Temp.:103.10, PWM:38
16:41:14.455 -> Water_Temp.:103.10, PWM:38
16:41:15.253 -> Water_Temp.:103.10, PWM:39
16:41:17.567 -> Water_Temp.:103.10, PWM:38
16:41:18.363 -> Water_Temp.:103.21, PWM:38
16:41:19.915 -> Water_Temp.:103.21, PWM:38
16:41:20.676 -> Water_Temp.:103.21, PWM:38
16:41:21.468 -> Water_Temp.:103.21, PWM:39
16:41:22.259 -> Water_Temp.:103.21, PWM:38
16:41:23.777 -> Water_Temp.:103.21, PWM:39
16:41:25.332 -> Water_Temp.:103.21, PWM:38
16:41:26.130 -> Water_Temp.:103.21, PWM:38
16:41:26.890 -> Water_Temp.:103.21, PWM:38
16:41:28.479 -> Water_Temp.:103.21, PWM:38
16:41:30.034 -> Water_Temp.:103.21, PWM:38
16:41:31.591 -> Water_Temp.:103.21, PWM:38
16:41:32.351 -> Water_Temp.:103.21, PWM:38
16:41:33.141 -> Water_Temp.:103.21, PWM:39
16:41:35.491 -> Water_Temp.:103.21, PWM:39
16:41:36.253 -> Water_Temp.:103.21, PWM:38
16:41:38.565 -> Water_Temp.:103.21, PWM:38
16:41:39.323 -> Water_Temp.:103.21, PWM:38
16:41:40.912 -> Water_Temp.:103.21, PWM:38
16:41:41.669 -> Water_Temp.:103.21, PWM:39
16:41:42.463 -> Water_Temp.:103.21, PWM:38
16:41:43.224 -> Water_Temp.:103.21, PWM:38
16:41:45.570 -> Water_Temp.:103.21, PWM:39
16:41:46.328 -> Water_Temp.:103.21, PWM:38
16:41:47.121 -> Water_Temp.:103.21, PWM:39
16:41:47.884 -> Water_Temp.:103.21, PWM:38
16:41:50.187 -> Water_Temp.:103.21, PWM:38
16:41:51.746 -> Water_Temp.:103.21, PWM:39
16:41:52.540 -> Water_Temp.:103.21, PWM:38
16:41:53.329 -> Water_Temp.:103.21, PWM:39
16:41:54.883 -> Water_Temp.:103.21, PWM:38
16:41:55.676 -> Water_Temp.:103.21, PWM:38
16:41:56.433 -> Water_Temp.:103.21, PWM:38
16:41:58.740 -> Water_Temp.:103.21, PWM:38
16:41:59.537 -> Water_Temp.:103.21, PWM:38
16:42:00.295 -> Water_Temp.:103.21, PWM:38
16:42:01.875 -> Water_Temp.:103.21, PWM:38
16:42:02.634 -> Water_Temp.:103.21, PWM:38
16:42:04.186 -> Water_Temp.:103.21, PWM:38
16:42:04.978 -> Water_Temp.:103.21, PWM:38
16:42:05.768 -> Water_Temp.:103.21, PWM:39
16:42:07.312 -> Water_Temp.:103.21, PWM:39
16:42:08.103 -> Water_Temp.:103.21, PWM:38
16:42:08.861 -> Water_Temp.:103.21, PWM:38
16:42:09.655 -> Water_Temp.:103.21, PWM:39
16:42:10.444 -> Water_Temp.:103.21, PWM:38
16:42:11.963 -> Water_Temp.:103.21, PWM:39
16:42:13.553 -> Water_Temp.:103.32, PWM:38
16:42:15.102 -> Water_Temp.:103.32, PWM:38
16:42:15.867 -> Water_Temp.:103.32, PWM:39
16:42:16.664 -> Water_Temp.:103.32, PWM:39
16:42:17.422 -> Water_Temp.:103.32, PWM:38
16:42:18.217 -> Water_Temp.:103.32, PWM:38
16:42:19.012 -> Water_Temp.:103.32, PWM:39
16:42:19.776 -> Water_Temp.:103.32, PWM:39
16:42:20.570 -> Water_Temp.:103.32, PWM:39
16:42:24.369 -> Water_Temp.:103.32, PWM:39
16:42:25.924 -> Water_Temp.:103.32, PWM:39
16:42:26.721 -> Water_Temp.:103.32, PWM:39
16:42:27.477 -> Water_Temp.:103.32, PWM:38
16:42:29.035 -> Water_Temp.:103.21, PWM:38
16:42:29.798 -> Water_Temp.:103.32, PWM:38
16:42:30.594 -> Water_Temp.:103.21, PWM:39
16:42:31.353 -> Water_Temp.:103.21, PWM:38
16:42:32.148 -> Water_Temp.:103.21, PWM:38
16:42:32.934 -> Water_Temp.:103.21, PWM:38
16:42:34.485 -> Water_Temp.:103.21, PWM:38
16:42:35.244 -> Water_Temp.:103.21, PWM:39
16:42:36.828 -> Water_Temp.:103.21, PWM:38
16:42:37.583 -> Water_Temp.:103.21, PWM:38
16:42:39.130 -> Water_Temp.:103.21, PWM:38
16:42:39.920 -> Water_Temp.:103.21, PWM:39
16:42:40.677 -> Water_Temp.:103.21, PWM:39
16:42:41.467 -> Water_Temp.:103.21, PWM:39
16:42:42.257 -> Water_Temp.:103.21, PWM:38
16:42:43.014 -> Water_Temp.:103.21, PWM:38
16:42:43.806 -> Water_Temp.:103.21, PWM:39
16:42:44.596 -> Water_Temp.:103.21, PWM:38
16:42:46.913 -> Water_Temp.:103.21, PWM:38
16:42:47.668 -> Water_Temp.:103.21, PWM:39
16:42:48.457 -> Water_Temp.:103.21, PWM:39
16:42:49.244 -> Water_Temp.:103.21, PWM:39
16:42:50.791 -> Water_Temp.:103.21, PWM:38
16:42:51.582 -> Water_Temp.:103.21, PWM:39
16:42:52.376 -> Water_Temp.:103.21, PWM:38
16:42:53.136 -> Water_Temp.:103.32, PWM:39
16:42:54.689 -> Water_Temp.:103.21, PWM:38
16:42:57.760 -> Water_Temp.:103.21, PWM:38
16:42:59.315 -> Water_Temp.:103.21, PWM:39
16:43:00.105 -> Water_Temp.:103.21, PWM:39
16:43:00.864 -> Water_Temp.:103.21, PWM:38
16:43:01.655 -> Water_Temp.:103.21, PWM:39
16:43:02.444 -> Water_Temp.:103.21, PWM:38
16:43:03.201 -> Water_Temp.:103.21, PWM:39
16:43:03.993 -> Water_Temp.:103.21, PWM:38
16:43:04.786 -> Water_Temp.:103.21, PWM:39
16:43:05.538 -> Water_Temp.:103.21, PWM:39
16:43:06.334 -> Water_Temp.:103.21, PWM:38
16:43:07.885 -> Water_Temp.:103.21, PWM:38
16:43:09.437 -> Water_Temp.:103.21, PWM:39
16:43:10.195 -> Water_Temp.:103.21, PWM:38
16:43:10.986 -> Water_Temp.:103.21, PWM:38
16:43:11.780 -> Water_Temp.:103.21, PWM:39
16:43:14.089 -> Water_Temp.:103.21, PWM:38
16:43:17.160 -> Water_Temp.:103.21, PWM:38
16:43:17.952 -> Water_Temp.:103.21, PWM:39
16:43:21.784 -> Water_Temp.:103.32, PWM:38
16:43:22.580 -> Water_Temp.:103.32, PWM:38
16:43:24.137 -> Water_Temp.:103.32, PWM:39
16:43:24.899 -> Water_Temp.:103.32, PWM:39
16:43:25.694 -> Water_Temp.:103.32, PWM:39
16:43:26.453 -> Water_Temp.:103.32, PWM:38
16:43:28.793 -> Water_Temp.:103.32, PWM:38
16:43:29.590 -> Water_Temp.:103.32, PWM:38
16:43:30.348 -> Water_Temp.:103.32, PWM:38
16:43:31.904 -> Water_Temp.:103.32, PWM:39
16:43:34.182 -> Water_Temp.:103.32, PWM:38
16:43:34.977 -> Water_Temp.:103.32, PWM:38
16:43:35.734 -> Water_Temp.:103.32, PWM:39
16:43:37.286 -> Water_Temp.:103.32, PWM:39
16:43:38.842 -> Water_Temp.:103.32, PWM:38
16:43:39.600 -> Water_Temp.:103.32, PWM:38
16:43:40.391 -> Water_Temp.:103.32, PWM:38
16:43:41.181 -> Water_Temp.:103.32, PWM:39
16:43:41.942 -> Water_Temp.:103.32, PWM:38
16:43:42.736 -> Water_Temp.:103.32, PWM:38
16:43:43.497 -> Water_Temp.:103.32, PWM:38
16:43:44.294 -> Water_Temp.:103.32, PWM:38
16:43:45.054 -> Water_Temp.:103.32, PWM:38
16:43:45.850 -> Water_Temp.:103.32, PWM:38
16:43:46.609 -> Water_Temp.:103.32, PWM:39
16:43:48.954 -> Water_Temp.:105.01, PWM:39
16:43:49.746 -> Water_Temp.:107.04, PWM:39
16:43:50.507 -> Water_Temp.:104.56, PWM:38
16:43:51.299 -> Water_Temp.:104.22, PWM:38
16:43:52.815 -> Water_Temp.:103.89, PWM:39
16:43:55.159 -> Water_Temp.:103.66, PWM:39
16:43:56.715 -> Water_Temp.:103.55, PWM:38
16:43:57.512 -> Water_Temp.:103.55, PWM:38
16:43:58.273 -> Water_Temp.:103.44, PWM:38
16:44:00.624 -> Water_Temp.:103.44, PWM:38
16:44:02.905 -> Water_Temp.:103.32, PWM:38
16:44:03.703 -> Water_Temp.:103.32, PWM:38
16:44:04.466 -> Water_Temp.:103.32, PWM:38
16:44:05.258 -> Water_Temp.:103.32, PWM:38
16:44:06.048 -> Water_Temp.:103.32, PWM:38
16:44:06.804 -> Water_Temp.:103.32, PWM:38
16:44:07.595 -> Water_Temp.:103.32, PWM:38
16:44:08.391 -> Water_Temp.:103.32, PWM:38
16:44:09.148 -> Water_Temp.:103.32, PWM:39
16:44:11.457 -> Water_Temp.:103.32, PWM:39
16:44:13.015 -> Water_Temp.:103.32, PWM:38
16:44:13.812 -> Water_Temp.:103.32, PWM:38
16:44:15.363 -> Water_Temp.:103.32, PWM:39
16:44:16.124 -> Water_Temp.:103.32, PWM:39
16:44:16.922 -> Water_Temp.:103.21, PWM:38
16:44:17.677 -> Water_Temp.:103.21, PWM:39
16:44:18.470 -> Water_Temp.:103.21, PWM:38
16:44:20.024 -> Water_Temp.:103.21, PWM:39
16:44:20.818 -> Water_Temp.:103.21, PWM:38
16:44:21.580 -> Water_Temp.:103.21, PWM:39
16:44:22.369 -> Water_Temp.:103.21, PWM:38
16:44:23.130 -> Water_Temp.:103.21, PWM:38
16:44:23.924 -> Water_Temp.:103.21, PWM:38
16:44:25.473 -> Water_Temp.:103.21, PWM:38
16:44:26.239 -> Water_Temp.:103.32, PWM:38
16:44:27.037 -> Water_Temp.:103.32, PWM:38
16:44:27.796 -> Water_Temp.:103.32, PWM:38
16:44:28.588 -> Water_Temp.:103.32, PWM:39
16:44:29.375 -> Water_Temp.:103.32, PWM:38
16:44:30.926 -> Water_Temp.:103.32, PWM:39
16:44:31.683 -> Water_Temp.:103.32, PWM:38
16:44:34.032 -> Water_Temp.:103.21, PWM:38
16:44:34.789 -> Water_Temp.:103.32, PWM:39
16:44:36.374 -> Water_Temp.:103.21, PWM:38
16:44:37.922 -> Water_Temp.:103.32, PWM:38
16:44:38.682 -> Water_Temp.:103.21, PWM:39
16:44:39.473 -> Water_Temp.:103.21, PWM:39
16:44:41.817 -> Water_Temp.:103.32, PWM:38
16:44:43.376 -> Water_Temp.:103.32, PWM:38
16:44:44.895 -> Water_Temp.:103.21, PWM:38
16:44:45.690 -> Water_Temp.:103.21, PWM:38
16:44:46.451 -> Water_Temp.:103.21, PWM:38
16:44:48.764 -> Water_Temp.:103.21, PWM:39
16:44:49.526 -> Water_Temp.:103.21, PWM:38
16:44:50.319 -> Water_Temp.:103.21, PWM:39
16:44:51.115 -> Water_Temp.:103.21, PWM:38
16:44:52.633 -> Water_Temp.:103.21, PWM:39
16:44:54.981 -> Water_Temp.:103.21, PWM:38
16:44:55.742 -> Water_Temp.:103.21, PWM:38
16:44:56.539 -> Water_Temp.:103.21, PWM:38
16:44:57.302 -> Water_Temp.:103.21, PWM:38
16:44:58.098 -> Water_Temp.:103.21, PWM:38
16:45:00.416 -> Water_Temp.:103.21, PWM:38
16:45:02.766 -> Water_Temp.:103.21, PWM:39
16:45:03.562 -> Water_Temp.:103.21, PWM:38
16:45:04.324 -> Water_Temp.:103.21, PWM:39
16:45:05.114 -> Water_Temp.:103.21, PWM:39
16:45:05.874 -> Water_Temp.:103.21, PWM:38
16:45:06.664 -> Water_Temp.:103.21, PWM:38
16:45:08.220 -> Water_Temp.:103.21, PWM:39
16:45:08.980 -> Water_Temp.:103.21, PWM:38
16:45:09.772 -> Water_Temp.:103.21, PWM:38
16:45:12.086 -> Water_Temp.:103.10, PWM:38
16:45:12.880 -> Water_Temp.:103.21, PWM:38
16:45:13.675 -> Water_Temp.:103.10, PWM:38
16:45:15.228 -> Water_Temp.:103.21, PWM:38
16:45:16.747 -> Water_Temp.:103.21, PWM:39
16:45:18.305 -> Water_Temp.:103.21, PWM:39
16:45:19.099 -> Water_Temp.:103.21, PWM:39
16:45:19.861 -> Water_Temp.:103.21, PWM:38
16:45:21.414 -> Water_Temp.:103.21, PWM:38
16:45:22.968 -> Water_Temp.:103.21, PWM:39
16:45:23.763 -> Water_Temp.:103.21, PWM:38
16:45:24.523 -> Water_Temp.:103.21, PWM:38
16:45:25.317 -> Water_Temp.:103.21, PWM:39
16:45:26.080 -> Water_Temp.:103.21, PWM:38
16:45:26.872 -> Water_Temp.:103.21, PWM:38
16:45:27.658 -> Water_Temp.:103.21, PWM:38
16:45:29.971 -> Water_Temp.:104.79, PWM:38
16:45:30.766 -> Water_Temp.:106.92, PWM:38
16:45:31.527 -> Water_Temp.:104.45, PWM:38
16:45:32.321 -> Water_Temp.:104.11, PWM:38
16:45:33.079 -> Water_Temp.:103.89, PWM:39
16:45:33.874 -> Water_Temp.:103.77, PWM:39
16:45:35.396 -> Water_Temp.:103.55, PWM:38
16:45:36.191 -> Water_Temp.:103.55, PWM:39
16:45:36.982 -> Water_Temp.:103.55, PWM:38
16:45:37.740 -> Water_Temp.:103.44, PWM:38
16:45:39.298 -> Water_Temp.:103.44, PWM:38
16:45:40.096 -> Water_Temp.:103.44, PWM:38
16:45:40.858 -> Water_Temp.:103.44, PWM:39
16:45:41.654 -> Water_Temp.:103.32, PWM:39
16:45:44.006 -> Water_Temp.:103.32, PWM:39
16:45:44.764 -> Water_Temp.:103.32, PWM:38
16:45:45.560 -> Water_Temp.:103.32, PWM:39
16:45:47.874 -> Water_Temp.:103.32, PWM:38
16:45:48.667 -> Water_Temp.:103.32, PWM:39
16:45:49.427 -> Water_Temp.:103.32, PWM:38
16:45:50.221 -> Water_Temp.:103.32, PWM:38
16:45:50.982 -> Water_Temp.:103.32, PWM:39
16:45:51.775 -> Water_Temp.:103.32, PWM:39
16:45:52.567 -> Water_Temp.:103.32, PWM:39
16:45:55.679 -> Water_Temp.:103.44, PWM:38
16:45:56.441 -> Water_Temp.:103.44, PWM:39
16:45:57.238 -> Water_Temp.:103.44, PWM:39
16:45:58.759 -> Water_Temp.:103.44, PWM:39
16:45:59.554 -> Water_Temp.:103.44, PWM:38
16:46:00.349 -> Water_Temp.:103.44, PWM:38
16:46:01.112 -> Water_Temp.:103.44, PWM:38
16:46:02.659 -> Water_Temp.:103.44, PWM:38
16:46:03.449 -> Water_Temp.:103.44, PWM:38
16:46:04.240 -> Water_Temp.:103.44, PWM:39
16:46:05.791 -> Water_Temp.:103.44, PWM:39
16:46:07.343 -> Water_Temp.:103.44, PWM:39
16:46:08.901 -> Water_Temp.:103.44, PWM:39
16:46:09.662 -> Water_Temp.:103.44, PWM:38
16:46:10.458 -> Water_Temp.:105.12, PWM:38
16:46:15.110 -> Water_Temp.:105.57, PWM:39
16:46:15.872 -> Water_Temp.:107.60, PWM:38
16:46:17.430 -> Water_Temp.:104.67, PWM:39
16:46:18.224 -> Water_Temp.:104.45, PWM:39
16:46:19.774 -> Water_Temp.:104.11, PWM:39
16:46:20.570 -> Water_Temp.:104.11, PWM:38
16:46:22.126 -> Water_Temp.:104.00, PWM:38
16:46:22.888 -> Water_Temp.:103.89, PWM:39
16:46:25.938 -> Water_Temp.:103.77, PWM:38
16:46:26.730 -> Water_Temp.:103.77, PWM:38
16:46:27.521 -> Water_Temp.:103.66, PWM:38
16:46:28.279 -> Water_Temp.:103.66, PWM:38
16:46:29.074 -> Water_Temp.:103.66, PWM:38
16:46:30.627 -> Water_Temp.:103.66, PWM:39
16:46:31.421 -> Water_Temp.:103.66, PWM:38
16:46:32.182 -> Water_Temp.:103.66, PWM:39
16:46:32.978 -> Water_Temp.:103.66, PWM:39
16:46:33.739 -> Water_Temp.:103.66, PWM:38
16:46:34.533 -> Water_Temp.:103.66, PWM:39
16:46:36.051 -> Water_Temp.:103.55, PWM:38
16:46:37.635 -> Water_Temp.:103.55, PWM:38
16:46:38.394 -> Water_Temp.:103.55, PWM:39
16:46:39.949 -> Water_Temp.:103.55, PWM:39
16:46:40.747 -> Water_Temp.:103.55, PWM:38
16:46:42.306 -> Water_Temp.:103.55, PWM:39
16:46:43.068 -> Water_Temp.:103.55, PWM:39
16:46:43.866 -> Water_Temp.:103.55, PWM:39
16:46:46.180 -> Water_Temp.:103.55, PWM:39
16:46:47.733 -> Water_Temp.:103.55, PWM:38
16:46:50.050 -> Water_Temp.:103.55, PWM:39
16:46:50.810 -> Water_Temp.:103.55, PWM:38
16:46:52.362 -> Water_Temp.:103.55, PWM:38
16:46:53.916 -> Water_Temp.:103.55, PWM:38
16:46:55.472 -> Water_Temp.:103.55, PWM:39
16:46:56.236 -> Water_Temp.:103.55, PWM:39
16:46:57.032 -> Water_Temp.:103.55, PWM:38
16:46:58.588 -> Water_Temp.:103.55, PWM:38
16:47:00.108 -> Water_Temp.:103.55, PWM:38
16:47:00.902 -> Water_Temp.:103.55, PWM:39
16:47:01.695 -> Water_Temp.:103.55, PWM:38
16:47:02.455 -> Water_Temp.:103.55, PWM:38
16:47:03.247 -> Water_Temp.:103.55, PWM:38
16:47:04.004 -> Water_Temp.:103.55, PWM:39
16:47:04.800 -> Water_Temp.:103.55, PWM:39
16:47:06.354 -> Water_Temp.:103.55, PWM:39
16:47:07.148 -> Water_Temp.:103.55, PWM:39
16:47:07.908 -> Water_Temp.:103.55, PWM:38
16:47:08.699 -> Water_Temp.:103.55, PWM:39
16:47:09.460 -> Water_Temp.:103.55, PWM:39
16:47:10.256 -> Water_Temp.:103.44, PWM:39
16:47:11.054 -> Water_Temp.:103.44, PWM:39
16:47:11.817 -> Water_Temp.:103.44, PWM:38
16:47:12.605 -> Water_Temp.:103.44, PWM:38
16:47:14.122 -> Water_Temp.:103.44, PWM:38
16:47:14.919 -> Water_Temp.:103.44, PWM:38
16:47:15.714 -> Water_Temp.:103.44, PWM:38
16:47:16.474 -> Water_Temp.:103.44, PWM:38
16:47:17.262 -> Water_Temp.:103.55, PWM:38
16:47:18.050 -> Water_Temp.:103.55, PWM:38
16:47:19.572 -> Water_Temp.:103.55, PWM:38
16:47:20.367 -> Water_Temp.:103.55, PWM:39
16:47:21.923 -> Water_Temp.:103.55, PWM:38
16:47:22.711 -> Water_Temp.:103.55, PWM:39
16:47:23.473 -> Water_Temp.:103.55, PWM:38
16:47:24.270 -> Water_Temp.:103.55, PWM:39
16:47:25.031 -> Water_Temp.:103.55, PWM:38
16:47:25.829 -> Water_Temp.:103.55, PWM:38
16:47:26.589 -> Water_Temp.:103.55, PWM:39
16:47:27.380 -> Water_Temp.:103.55, PWM:38
16:47:28.177 -> Water_Temp.:103.66, PWM:39
16:47:29.694 -> Water_Temp.:103.66, PWM:38
16:47:30.488 -> Water_Temp.:103.66, PWM:39
16:47:31.252 -> Water_Temp.:103.66, PWM:38
16:47:33.601 -> Water_Temp.:103.66, PWM:39
16:47:34.360 -> Water_Temp.:103.66, PWM:39
16:47:35.153 -> Water_Temp.:103.66, PWM:38
16:47:38.256 -> Water_Temp.:103.66, PWM:39
16:47:39.018 -> Water_Temp.:103.66, PWM:38
16:47:39.807 -> Water_Temp.:103.66, PWM:39
16:47:40.569 -> Water_Temp.:103.66, PWM:38
16:47:41.361 -> Water_Temp.:103.66, PWM:39
16:47:42.916 -> Water_Temp.:103.66, PWM:39
16:47:43.711 -> Water_Temp.:103.77, PWM:39
16:47:45.267 -> Water_Temp.:103.77, PWM:38
16:47:46.025 -> Water_Temp.:103.77, PWM:38
16:47:46.823 -> Water_Temp.:103.77, PWM:38
16:47:47.586 -> Water_Temp.:103.77, PWM:38
16:47:48.379 -> Water_Temp.:103.77, PWM:38
16:47:49.171 -> Water_Temp.:103.77, PWM:38
16:47:49.932 -> Water_Temp.:103.77, PWM:38
16:47:52.248 -> Water_Temp.:103.77, PWM:38
16:47:53.802 -> Water_Temp.:103.77, PWM:38
16:47:54.557 -> Water_Temp.:103.77, PWM:39
16:47:55.352 -> Water_Temp.:103.77, PWM:38
16:47:56.110 -> Water_Temp.:103.77, PWM:39
16:47:57.696 -> Water_Temp.:103.77, PWM:39
16:47:58.457 -> Water_Temp.:103.77, PWM:38
16:47:59.254 -> Water_Temp.:103.77, PWM:38
16:48:02.321 -> Water_Temp.:103.77, PWM:39
16:48:03.080 -> Water_Temp.:103.77, PWM:38
16:48:04.632 -> Water_Temp.:103.89, PWM:38
16:48:05.424 -> Water_Temp.:103.89, PWM:38
16:48:06.184 -> Water_Temp.:103.77, PWM:38
16:48:06.981 -> Water_Temp.:103.77, PWM:38
16:48:08.502 -> Water_Temp.:103.77, PWM:38
16:48:09.296 -> Water_Temp.:103.77, PWM:38
16:48:10.844 -> Water_Temp.:103.77, PWM:38
16:48:11.641 -> Water_Temp.:103.77, PWM:39
16:48:12.398 -> Water_Temp.:103.77, PWM:39
16:48:13.192 -> Water_Temp.:103.77, PWM:39
16:48:13.984 -> Water_Temp.:103.77, PWM:39
16:48:14.746 -> Water_Temp.:103.77, PWM:39
16:48:15.538 -> Water_Temp.:103.77, PWM:38
16:48:17.064 -> Water_Temp.:103.66, PWM:38
16:48:17.860 -> Water_Temp.:103.66, PWM:38
16:48:18.651 -> Water_Temp.:103.66, PWM:38
16:48:19.413 -> Water_Temp.:103.66, PWM:38
16:48:20.206 -> Water_Temp.:103.66, PWM:38
16:48:20.970 -> Water_Temp.:103.66, PWM:38
16:48:21.764 -> Water_Temp.:103.66, PWM:38
16:48:22.523 -> Water_Temp.:103.66, PWM:38
16:48:24.078 -> Water_Temp.:103.66, PWM:38
16:48:24.871 -> Water_Temp.:103.66, PWM:38
16:48:25.632 -> Water_Temp.:103.66, PWM:38
16:48:26.430 -> Water_Temp.:103.66, PWM:39
16:48:28.747 -> Water_Temp.:103.66, PWM:39
16:48:30.289 -> Water_Temp.:103.66, PWM:38
16:48:31.049 -> Water_Temp.:103.77, PWM:38
16:48:31.845 -> Water_Temp.:103.77, PWM:39
16:48:32.639 -> Water_Temp.:103.66, PWM:38
16:48:33.396 -> Water_Temp.:103.77, PWM:38
16:48:34.948 -> Water_Temp.:103.77, PWM:38
16:48:35.706 -> Water_Temp.:103.77, PWM:39
16:48:36.500 -> Water_Temp.:103.77, PWM:39
16:48:38.022 -> Water_Temp.:103.77, PWM:39
16:48:38.817 -> Water_Temp.:103.77, PWM:39
16:48:40.372 -> Water_Temp.:103.77, PWM:38
16:48:41.164 -> Water_Temp.:103.77, PWM:39
16:48:45.756 -> Water_Temp.:103.89, PWM:38
16:48:48.832 -> Water_Temp.:103.89, PWM:38
16:48:49.591 -> Water_Temp.:103.89, PWM:39
16:48:50.382 -> Water_Temp.:103.89, PWM:38
16:48:51.177 -> Water_Temp.:103.89, PWM:38
16:48:52.734 -> Water_Temp.:103.89, PWM:38
16:48:53.492 -> Water_Temp.:103.77, PWM:38
16:48:54.287 -> Water_Temp.:103.77, PWM:38
16:48:55.078 -> Water_Temp.:103.77, PWM:38
16:48:55.839 -> Water_Temp.:103.77, PWM:38
16:49:00.536 -> Water_Temp.:107.37, PWM:38
16:49:02.086 -> Water_Temp.:104.56, PWM:38
16:49:02.844 -> Water_Temp.:104.45, PWM:38
16:49:03.636 -> Water_Temp.:104.45, PWM:38
16:49:05.954 -> Water_Temp.:104.11, PWM:38
16:49:06.714 -> Water_Temp.:104.11, PWM:38
16:49:08.267 -> Water_Temp.:104.00, PWM:38
16:49:09.034 -> Water_Temp.:104.00, PWM:39
16:49:09.830 -> Water_Temp.:104.00, PWM:38
16:49:10.621 -> Water_Temp.:104.00, PWM:38
16:49:11.381 -> Water_Temp.:104.00, PWM:38
16:49:12.931 -> Water_Temp.:104.00, PWM:38
16:49:13.725 -> Water_Temp.:104.00, PWM:38
16:49:14.514 -> Water_Temp.:104.00, PWM:38
16:49:16.039 -> Water_Temp.:103.89, PWM:39
16:49:16.835 -> Water_Temp.:103.89, PWM:39
16:49:18.391 -> Water_Temp.:103.89, PWM:38
16:49:19.151 -> Water_Temp.:103.89, PWM:38
16:49:20.707 -> Water_Temp.:103.89, PWM:38
16:49:21.499 -> Water_Temp.:103.89, PWM:38
16:49:22.261 -> Water_Temp.:103.89, PWM:39
16:49:23.817 -> Water_Temp.:103.77, PWM:39
16:49:25.374 -> Water_Temp.:103.77, PWM:39
16:49:26.169 -> Water_Temp.:103.77, PWM:38
16:49:26.929 -> Water_Temp.:103.77, PWM:38
16:49:27.729 -> Water_Temp.:103.77, PWM:38
16:49:28.491 -> Water_Temp.:103.77, PWM:39
16:49:29.287 -> Water_Temp.:103.66, PWM:39
16:49:30.048 -> Water_Temp.:103.66, PWM:39
16:49:31.604 -> Water_Temp.:103.66, PWM:38
16:49:32.398 -> Water_Temp.:103.66, PWM:38
16:49:33.160 -> Water_Temp.:103.66, PWM:38
16:49:33.953 -> Water_Temp.:103.66, PWM:38
16:49:35.509 -> Water_Temp.:103.66, PWM:38
16:49:37.063 -> Water_Temp.:103.77, PWM:38
16:49:37.857 -> Water_Temp.:103.77, PWM:39
16:49:38.615 -> Water_Temp.:103.77, PWM:39
16:49:39.407 -> Water_Temp.:103.77, PWM:39
16:49:40.203 -> Water_Temp.:103.77, PWM:38
16:49:40.962 -> Water_Temp.:103.77, PWM:38
16:49:43.282 -> Water_Temp.:103.77, PWM:39
16:49:45.639 -> Water_Temp.:103.77, PWM:38
16:49:46.403 -> Water_Temp.:103.77, PWM:39
16:49:47.198 -> Water_Temp.:103.77, PWM:38
16:49:47.959 -> Water_Temp.:103.77, PWM:39
16:49:48.746 -> Water_Temp.:103.77, PWM:38
16:49:49.544 -> Water_Temp.:103.77, PWM:38
16:49:51.098 -> Water_Temp.:103.77, PWM:38
16:49:52.612 -> Water_Temp.:103.89, PWM:39
16:49:53.403 -> Water_Temp.:103.89, PWM:39
16:49:54.197 -> Water_Temp.:103.89, PWM:39
16:49:54.957 -> Water_Temp.:103.89, PWM:39
16:49:55.748 -> Water_Temp.:103.89, PWM:38
16:49:56.541 -> Water_Temp.:103.89, PWM:39
16:49:58.094 -> Water_Temp.:103.89, PWM:38
16:49:58.853 -> Water_Temp.:103.89, PWM:39
16:49:59.649 -> Water_Temp.:103.77, PWM:39
16:50:00.441 -> Water_Temp.:103.89, PWM:39
16:50:01.994 -> Water_Temp.:103.89, PWM:38
16:50:03.545 -> Water_Temp.:103.77, PWM:39
16:50:05.100 -> Water_Temp.:103.77, PWM:39
16:50:06.659 -> Water_Temp.:103.77, PWM:38
16:50:07.419 -> Water_Temp.:103.77, PWM:38
16:50:08.212 -> Water_Temp.:103.77, PWM:38
16:50:08.975 -> Water_Temp.:103.77, PWM:38
16:50:09.768 -> Water_Temp.:103.77, PWM:38
16:50:12.078 -> Water_Temp.:103.77, PWM:39
16:50:12.839 -> Water_Temp.:103.77, PWM:39
16:50:15.190 -> Water_Temp.:103.66, PWM:39
16:50:16.742 -> Water_Temp.:103.66, PWM:39
16:50:17.505 -> Water_Temp.:103.66, PWM:39
16:50:18.298 -> Water_Temp.:103.77, PWM:38
16:50:19.820 -> Water_Temp.:103.66, PWM:39
16:50:22.162 -> Water_Temp.:103.66, PWM:39
16:50:22.957 -> Water_Temp.:103.66, PWM:38
16:50:23.715 -> Water_Temp.:103.66, PWM:39
16:50:25.302 -> Water_Temp.:103.66, PWM:38
16:50:26.064 -> Water_Temp.:103.66, PWM:38
16:50:26.855 -> Water_Temp.:103.66, PWM:39
16:50:27.618 -> Water_Temp.:103.66, PWM:39
16:50:28.414 -> Water_Temp.:103.66, PWM:38
16:50:30.736 -> Water_Temp.:103.66, PWM:38
16:50:31.497 -> Water_Temp.:103.66, PWM:39
16:50:32.289 -> Water_Temp.:103.66, PWM:39
16:50:33.808 -> Water_Temp.:103.66, PWM:38
16:50:34.603 -> Water_Temp.:103.66, PWM:38
16:50:35.358 -> Water_Temp.:103.66, PWM:38
16:50:36.153 -> Water_Temp.:103.66, PWM:39
16:50:39.268 -> Water_Temp.:103.66, PWM:38
16:50:40.064 -> Water_Temp.:103.66, PWM:38
16:50:41.620 -> Water_Temp.:103.66, PWM:38
16:50:42.381 -> Water_Temp.:103.66, PWM:39
16:50:43.179 -> Water_Temp.:103.66, PWM:39
16:50:44.729 -> Water_Temp.:103.66, PWM:38
16:50:46.253 -> Water_Temp.:103.66, PWM:38
16:50:47.050 -> Water_Temp.:103.66, PWM:39
16:50:50.158 -> Water_Temp.:103.66, PWM:39
16:50:50.949 -> Water_Temp.:103.66, PWM:39
16:50:51.741 -> Water_Temp.:103.66, PWM:38
16:50:54.057 -> Water_Temp.:103.66, PWM:39
16:50:54.849 -> Water_Temp.:103.66, PWM:38
16:50:55.641 -> Water_Temp.:103.66, PWM:38
16:50:57.169 -> Water_Temp.:103.55, PWM:39
16:50:57.962 -> Water_Temp.:103.55, PWM:39
16:50:58.756 -> Water_Temp.:103.55, PWM:39
16:51:00.272 -> Water_Temp.:103.55, PWM:39
16:51:01.066 -> Water_Temp.:103.55, PWM:39
16:51:01.860 -> Water_Temp.:103.55, PWM:39
16:51:02.621 -> Water_Temp.:103.55, PWM:39
16:51:03.415 -> Water_Temp.:103.55, PWM:38
16:51:04.173 -> Water_Temp.:103.66, PWM:38
16:51:04.966 -> Water_Temp.:103.66, PWM:39
16:51:08.076 -> Water_Temp.:103.66, PWM:39
16:51:09.637 -> Water_Temp.:103.66, PWM:38
16:51:10.432 -> Water_Temp.:103.66, PWM:38
16:51:11.194 -> Water_Temp.:103.66, PWM:38
16:51:11.990 -> Water_Temp.:103.66, PWM:38
16:51:13.508 -> Water_Temp.:103.66, PWM:39
16:51:14.306 -> Water_Temp.:103.66, PWM:38
16:51:15.097 -> Water_Temp.:103.77, PWM:38
16:51:15.861 -> Water_Temp.:103.77, PWM:39
16:51:16.652 -> Water_Temp.:103.66, PWM:38
16:51:18.202 -> Water_Temp.:103.66, PWM:38
16:51:18.996 -> Water_Temp.:103.77, PWM:38
16:51:19.759 -> Water_Temp.:103.77, PWM:38
16:51:21.280 -> Water_Temp.:103.77, PWM:38
16:51:22.069 -> Water_Temp.:103.77, PWM:39
16:51:22.859 -> Water_Temp.:103.77, PWM:38
16:51:25.933 -> Water_Temp.:103.77, PWM:39
16:51:26.728 -> Water_Temp.:105.46, PWM:38
16:51:27.520 -> Water_Temp.:107.49, PWM:38
16:51:32.113 -> Water_Temp.:104.22, PWM:39
16:51:32.875 -> Water_Temp.:104.11, PWM:39
16:51:33.671 -> Water_Temp.:104.11, PWM:38
16:51:36.019 -> Water_Temp.:104.00, PWM:39
16:51:36.776 -> Water_Temp.:103.89, PWM:39
16:51:37.570 -> Water_Temp.:104.00, PWM:38
16:51:38.329 -> Water_Temp.:103.89, PWM:39
16:51:39.124 -> Water_Temp.:103.89, PWM:38
16:51:40.647 -> Water_Temp.:103.89, PWM:38
16:51:42.962 -> Water_Temp.:103.89, PWM:39
16:51:43.727 -> Water_Temp.:103.89, PWM:39
16:51:44.519 -> Water_Temp.:103.89, PWM:39
16:51:45.314 -> Water_Temp.:103.89, PWM:39
16:51:46.072 -> Water_Temp.:103.89, PWM:38
16:51:46.867 -> Water_Temp.:103.89, PWM:38
16:51:49.149 -> Water_Temp.:103.89, PWM:38
16:51:49.944 -> Water_Temp.:103.89, PWM:38
16:51:50.740 -> Water_Temp.:103.89, PWM:39
16:51:51.503 -> Water_Temp.:103.89, PWM:39
16:51:54.587 -> Water_Temp.:103.89, PWM:38
16:51:55.379 -> Water_Temp.:103.89, PWM:38
16:51:56.141 -> Water_Temp.:103.89, PWM:39
16:51:56.937 -> Water_Temp.:103.89, PWM:38
16:51:57.701 -> Water_Temp.:103.89, PWM:38
16:51:58.496 -> Water_Temp.:103.89, PWM:39
16:52:00.046 -> Water_Temp.:103.89, PWM:39
16:52:02.326 -> Water_Temp.:103.89, PWM:38
16:52:03.122 -> Water_Temp.:103.89, PWM:39
16:52:03.884 -> Water_Temp.:103.89, PWM:38
16:52:04.678 -> Water_Temp.:103.89, PWM:38
16:52:05.469 -> Water_Temp.:103.89, PWM:39
16:52:07.790 -> Water_Temp.:103.89, PWM:38
16:52:12.421 -> Water_Temp.:103.77, PWM:38
16:52:13.978 -> Water_Temp.:103.77, PWM:38
16:52:15.534 -> Water_Temp.:103.77, PWM:38
16:52:19.399 -> Water_Temp.:103.77, PWM:39
16:52:20.161 -> Water_Temp.:103.77, PWM:38
16:52:21.712 -> Water_Temp.:103.77, PWM:38
16:52:22.504 -> Water_Temp.:103.77, PWM:39
16:52:23.261 -> Water_Temp.:103.77, PWM:39
16:52:24.813 -> Water_Temp.:103.77, PWM:39
16:52:25.606 -> Water_Temp.:103.77, PWM:39
16:52:26.399 -> Water_Temp.:103.77, PWM:39
16:52:27.161 -> Water_Temp.:103.77, PWM:39
16:52:27.958 -> Water_Temp.:103.77, PWM:39
16:52:29.478 -> Water_Temp.:103.77, PWM:39
16:52:30.269 -> Water_Temp.:103.77, PWM:40
16:52:31.030 -> Water_Temp.:103.77, PWM:40
16:52:31.825 -> Water_Temp.:103.77, PWM:39
16:52:32.582 -> Water_Temp.:103.77, PWM:40
16:52:34.898 -> Water_Temp.:103.89, PWM:39
16:52:35.692 -> Water_Temp.:103.89, PWM:39
16:52:36.453 -> Water_Temp.:103.89, PWM:39
16:52:38.004 -> Water_Temp.:103.89, PWM:39
16:52:38.798 -> Water_Temp.:103.89, PWM:39
16:52:39.560 -> Water_Temp.:103.89, PWM:40
16:52:40.357 -> Water_Temp.:103.89, PWM:40
16:52:42.677 -> Water_Temp.:103.89, PWM:39
16:52:43.439 -> Water_Temp.:103.89, PWM:40
16:52:44.230 -> Water_Temp.:103.89, PWM:39
16:52:44.991 -> Water_Temp.:103.89, PWM:40
16:52:45.787 -> Water_Temp.:103.89, PWM:40
16:52:46.546 -> Water_Temp.:103.89, PWM:39
16:52:48.104 -> Water_Temp.:103.89, PWM:40
16:52:48.901 -> Water_Temp.:103.89, PWM:39
16:52:49.662 -> Water_Temp.:103.89, PWM:40
16:52:50.458 -> Water_Temp.:103.89, PWM:39
16:52:51.215 -> Water_Temp.:103.89, PWM:39
16:52:52.011 -> Water_Temp.:103.89, PWM:40
16:52:53.566 -> Water_Temp.:103.89, PWM:40
16:52:54.358 -> Water_Temp.:103.89, PWM:39
16:52:56.680 -> Water_Temp.:103.89, PWM:40
16:52:57.477 -> Water_Temp.:103.89, PWM:39
16:52:58.239 -> Water_Temp.:103.89, PWM:39
16:52:59.799 -> Water_Temp.:103.89, PWM:40
16:53:00.594 -> Water_Temp.:103.89, PWM:40
16:53:01.357 -> Water_Temp.:103.89, PWM:40
16:53:02.148 -> Water_Temp.:103.89, PWM:39
16:53:03.674 -> Water_Temp.:103.89, PWM:39
16:53:04.470 -> Water_Temp.:103.89, PWM:40
16:53:05.232 -> Water_Temp.:103.89, PWM:39
16:53:06.028 -> Water_Temp.:103.89, PWM:39
16:53:06.821 -> Water_Temp.:103.89, PWM:39
16:53:09.132 -> Water_Temp.:103.89, PWM:39
16:53:09.892 -> Water_Temp.:103.89, PWM:40
16:53:10.684 -> Water_Temp.:103.89, PWM:39
16:53:11.478 -> Water_Temp.:103.89, PWM:40
16:53:12.242 -> Water_Temp.:103.89, PWM:40
16:53:13.001 -> Water_Temp.:103.89, PWM:40
16:53:14.549 -> Water_Temp.:103.89, PWM:40
16:53:15.333 -> Water_Temp.:103.89, PWM:40
16:53:17.684 -> Water_Temp.:104.00, PWM:40
16:53:18.478 -> Water_Temp.:104.00, PWM:40
16:53:19.235 -> Water_Temp.:104.00, PWM:40
16:53:20.027 -> Water_Temp.:104.00, PWM:40
16:53:20.821 -> Water_Temp.:103.89, PWM:40
16:53:21.584 -> Water_Temp.:103.89, PWM:40
16:53:23.108 -> Water_Temp.:103.89, PWM:40
16:53:23.900 -> Water_Temp.:103.89, PWM:40
16:53:24.693 -> Water_Temp.:103.89, PWM:40
16:53:25.453 -> Water_Temp.:103.89, PWM:40
16:53:27.010 -> Water_Temp.:103.89, PWM:40
16:53:27.798 -> Water_Temp.:104.00, PWM:40
16:53:28.592 -> Water_Temp.:104.00, PWM:40
16:53:29.354 -> Water_Temp.:104.00, PWM:40
16:53:30.148 -> Water_Temp.:104.00, PWM:40
16:53:31.702 -> Water_Temp.:104.00, PWM:40
16:53:34.047 -> Water_Temp.:104.00, PWM:40
16:53:34.802 -> Water_Temp.:104.00, PWM:40
16:53:35.593 -> Water_Temp.:104.00, PWM:40
16:53:36.386 -> Water_Temp.:104.00, PWM:40
16:53:37.143 -> Water_Temp.:104.00, PWM:40
16:53:37.932 -> Water_Temp.:104.00, PWM:40
16:53:41.795 -> Water_Temp.:104.00, PWM:40
16:53:43.313 -> Water_Temp.:104.00, PWM:40
16:53:44.864 -> Water_Temp.:104.00, PWM:40
16:53:46.455 -> Water_Temp.:104.11, PWM:40
16:53:47.971 -> Water_Temp.:104.00, PWM:40
16:53:48.759 -> Water_Temp.:104.11, PWM:40
16:53:49.520 -> Water_Temp.:104.11, PWM:40
16:53:50.310 -> Water_Temp.:104.11, PWM:40
16:53:52.657 -> Water_Temp.:104.11, PWM:40
16:53:53.417 -> Water_Temp.:104.11, PWM:40
16:53:54.211 -> Water_Temp.:104.11, PWM:40
16:53:55.002 -> Water_Temp.:104.11, PWM:40
16:53:57.342 -> Water_Temp.:104.11, PWM:40
16:53:58.098 -> Water_Temp.:104.11, PWM:40
16:54:00.439 -> Water_Temp.:104.11, PWM:40
16:54:01.235 -> Water_Temp.:104.11, PWM:40
16:54:01.988 -> Water_Temp.:104.11, PWM:40
16:54:02.784 -> Water_Temp.:104.11, PWM:40
16:54:03.542 -> Water_Temp.:104.11, PWM:40
16:54:04.339 -> Water_Temp.:104.11, PWM:40
16:54:08.172 -> Water_Temp.:104.11, PWM:40
16:54:10.512 -> Water_Temp.:104.11, PWM:40
16:54:11.306 -> Water_Temp.:104.11, PWM:40
16:54:12.095 -> Water_Temp.:104.11, PWM:40
16:54:12.857 -> Water_Temp.:104.11, PWM:40
16:54:13.650 -> Water_Temp.:104.11, PWM:40
16:54:14.410 -> Water_Temp.:104.11, PWM:40
16:54:16.761 -> Water_Temp.:104.11, PWM:40
16:54:17.554 -> Water_Temp.:104.11, PWM:40
16:54:18.314 -> Water_Temp.:104.11, PWM:40
16:54:19.110 -> Water_Temp.:104.11, PWM:40
16:54:19.870 -> Water_Temp.:104.11, PWM:40
16:54:20.667 -> Water_Temp.:104.11, PWM:40
16:54:22.216 -> Water_Temp.:104.22, PWM:40
16:54:22.974 -> Water_Temp.:104.22, PWM:40
16:54:23.770 -> Water_Temp.:104.22, PWM:40
16:54:24.564 -> Water_Temp.:104.22, PWM:40
16:54:25.319 -> Water_Temp.:104.22, PWM:40
16:54:26.118 -> Water_Temp.:104.22, PWM:40
16:54:26.878 -> Water_Temp.:104.22, PWM:40
16:54:27.674 -> Water_Temp.:104.22, PWM:40
16:54:28.432 -> Water_Temp.:104.22, PWM:40
16:54:32.301 -> Water_Temp.:104.22, PWM:40
16:54:33.089 -> Water_Temp.:104.22, PWM:40
16:54:33.879 -> Water_Temp.:104.11, PWM:40
16:54:34.641 -> Water_Temp.:104.11, PWM:40
16:54:36.951 -> Water_Temp.:104.11, PWM:40
16:54:38.503 -> Water_Temp.:104.11, PWM:40
16:54:40.056 -> Water_Temp.:104.11, PWM:40
16:54:41.609 -> Water_Temp.:104.22, PWM:40
16:54:42.403 -> Water_Temp.:104.11, PWM:40
16:54:43.158 -> Water_Temp.:104.22, PWM:40
16:54:43.948 -> Water_Temp.:104.22, PWM:40
16:54:45.499 -> Water_Temp.:104.22, PWM:40
16:54:46.293 -> Water_Temp.:104.11, PWM:40
16:54:47.055 -> Water_Temp.:104.22, PWM:40
16:54:49.364 -> Water_Temp.:104.22, PWM:40
16:54:50.917 -> Water_Temp.:104.22, PWM:40
16:54:51.712 -> Water_Temp.:104.22, PWM:40
16:54:52.473 -> Water_Temp.:104.22, PWM:40
16:54:53.266 -> Water_Temp.:104.22, PWM:40
16:54:54.025 -> Water_Temp.:104.22, PWM:40
16:54:54.821 -> Water_Temp.:104.22, PWM:40
16:54:56.376 -> Water_Temp.:104.22, PWM:40
16:54:58.688 -> Water_Temp.:104.22, PWM:40
16:54:59.483 -> Water_Temp.:104.22, PWM:40
16:55:00.239 -> Water_Temp.:104.11, PWM:40
16:55:02.551 -> Water_Temp.:104.11, PWM:40
16:55:03.345 -> Water_Temp.:104.11, PWM:40
16:55:04.893 -> Water_Temp.:104.11, PWM:40
16:55:09.478 -> Water_Temp.:104.22, PWM:40
16:55:11.032 -> Water_Temp.:104.22, PWM:40
16:55:11.828 -> Water_Temp.:104.22, PWM:40
16:55:12.623 -> Water_Temp.:104.22, PWM:40
16:55:14.174 -> Water_Temp.:104.22, PWM:40
16:55:14.929 -> Water_Temp.:104.22, PWM:40
16:55:16.521 -> Water_Temp.:104.22, PWM:40
16:55:18.046 -> Water_Temp.:104.22, PWM:40
16:55:18.837 -> Water_Temp.:104.22, PWM:40
16:55:20.394 -> Water_Temp.:104.22, PWM:40
16:55:21.158 -> Water_Temp.:104.22, PWM:40
16:55:22.709 -> Water_Temp.:104.34, PWM:40
16:55:24.263 -> Water_Temp.:104.22, PWM:40
16:55:25.058 -> Water_Temp.:104.34, PWM:40
16:55:25.818 -> Water_Temp.:104.34, PWM:40
16:55:26.614 -> Water_Temp.:104.34, PWM:40
16:55:27.375 -> Water_Temp.:104.34, PWM:40
16:55:28.168 -> Water_Temp.:104.34, PWM:40
16:55:28.961 -> Water_Temp.:104.34, PWM:40
16:55:29.717 -> Water_Temp.:104.34, PWM:40
16:55:30.514 -> Water_Temp.:104.34, PWM:40
16:55:31.306 -> Water_Temp.:104.34, PWM:40
16:55:32.064 -> Water_Temp.:104.34, PWM:40
16:55:34.377 -> Water_Temp.:104.34, PWM:40
16:55:35.171 -> Water_Temp.:104.34, PWM:40
16:55:35.934 -> Water_Temp.:104.34, PWM:40
16:55:36.731 -> Water_Temp.:104.34, PWM:40
16:55:37.522 -> Water_Temp.:104.34, PWM:40
16:55:38.276 -> Water_Temp.:104.34, PWM:40
16:55:39.068 -> Water_Temp.:104.34, PWM:40
16:55:39.865 -> Water_Temp.:104.34, PWM:40
16:55:40.623 -> Water_Temp.:104.34, PWM:40
16:55:42.177 -> Water_Temp.:104.34, PWM:40
16:55:42.968 -> Water_Temp.:104.34, PWM:40
16:55:43.728 -> Water_Temp.:104.34, PWM:40
16:55:45.275 -> Water_Temp.:104.34, PWM:40
16:55:46.071 -> Water_Temp.:104.34, PWM:40
16:55:46.869 -> Water_Temp.:104.22, PWM:40
16:55:47.628 -> Water_Temp.:104.34, PWM:40
16:55:48.423 -> Water_Temp.:104.34, PWM:40
16:55:49.184 -> Water_Temp.:104.34, PWM:40
16:55:50.739 -> Water_Temp.:104.22, PWM:40
16:55:51.499 -> Water_Temp.:104.22, PWM:40
16:55:53.849 -> Water_Temp.:104.22, PWM:40
16:55:54.644 -> Water_Temp.:104.22, PWM:40
16:55:55.401 -> Water_Temp.:104.22, PWM:40
16:55:56.200 -> Water_Temp.:104.22, PWM:40
16:55:56.959 -> Water_Temp.:104.22, PWM:40
16:55:59.276 -> Water_Temp.:104.34, PWM:40
16:56:00.035 -> Water_Temp.:104.34, PWM:40
16:56:01.621 -> Water_Temp.:104.34, PWM:40
16:56:02.380 -> Water_Temp.:104.34, PWM:40
16:56:03.168 -> Water_Temp.:104.34, PWM:40
16:56:03.961 -> Water_Temp.:104.34, PWM:40
16:56:04.715 -> Water_Temp.:104.34, PWM:40
16:56:07.054 -> Water_Temp.:104.34, PWM:40
16:56:07.808 -> Water_Temp.:104.34, PWM:40
16:56:09.361 -> Water_Temp.:104.34, PWM:40
16:56:10.148 -> Water_Temp.:104.34, PWM:40
16:56:11.704 -> Water_Temp.:104.34, PWM:40
16:56:12.463 -> Water_Temp.:104.34, PWM:40
16:56:13.257 -> Water_Temp.:104.34, PWM:40
16:56:14.802 -> Water_Temp.:104.22, PWM:40
16:56:15.590 -> Water_Temp.:104.22, PWM:40
16:56:17.139 -> Water_Temp.:104.22, PWM:40
16:56:17.897 -> Water_Temp.:104.22, PWM:40
16:56:18.688 -> Water_Temp.:104.22, PWM:40
16:56:19.483 -> Water_Temp.:104.22, PWM:40
16:56:20.240 -> Water_Temp.:104.22, PWM:40
16:56:21.031 -> Water_Temp.:104.22, PWM:40
16:56:21.819 -> Water_Temp.:104.22, PWM:40
16:56:22.613 -> Water_Temp.:104.22, PWM:40
16:56:24.157 -> Water_Temp.:104.22, PWM:40
16:56:25.678 -> Water_Temp.:104.34, PWM:40
16:56:27.266 -> Water_Temp.:104.34, PWM:40
16:56:28.022 -> Water_Temp.:104.34, PWM:40
16:56:29.608 -> Water_Temp.:104.22, PWM:40
16:56:31.128 -> Water_Temp.:104.34, PWM:40
16:56:32.710 -> Water_Temp.:104.22, PWM:40
16:56:33.469 -> Water_Temp.:104.22, PWM:40
16:56:34.261 -> Water_Temp.:104.22, PWM:40
16:56:35.784 -> Water_Temp.:104.22, PWM:40
16:56:36.580 -> Water_Temp.:104.22, PWM:40
16:56:38.132 -> Water_Temp.:104.22, PWM:40
16:56:38.893 -> Water_Temp.:104.22, PWM:40
16:56:39.686 -> Water_Temp.:104.22, PWM:40
16:56:40.445 -> Water_Temp.:104.22, PWM:40
16:56:41.243 -> Water_Temp.:104.11, PWM:40
16:56:42.791 -> Water_Temp.:104.11, PWM:40
16:56:43.556 -> Water_Temp.:104.11, PWM:40
16:56:45.907 -> Water_Temp.:104.11, PWM:40
16:56:46.667 -> Water_Temp.:104.11, PWM:40
16:56:47.460 -> Water_Temp.:104.11, PWM:40
16:56:48.221 -> Water_Temp.:104.11, PWM:40
16:56:49.016 -> Water_Temp.:104.11, PWM:40
16:56:49.808 -> Water_Temp.:104.11, PWM:40
16:56:50.567 -> Water_Temp.:104.11, PWM:40
16:56:52.882 -> Water_Temp.:104.11, PWM:40
16:56:53.676 -> Water_Temp.:104.11, PWM:40
16:56:56.024 -> Water_Temp.:104.22, PWM:40
16:56:56.782 -> Water_Temp.:104.22, PWM:40
16:56:57.572 -> Water_Temp.:104.22, PWM:40
16:56:58.365 -> Water_Temp.:104.22, PWM:40
16:57:00.681 -> Water_Temp.:104.22, PWM:40
16:57:01.440 -> Water_Temp.:104.22, PWM:40
16:57:02.230 -> Water_Temp.:104.22, PWM:40
16:57:03.022 -> Water_Temp.:104.22, PWM:40
16:57:06.875 -> Water_Temp.:104.22, PWM:40
16:57:09.186 -> Water_Temp.:104.11, PWM:40
16:57:09.944 -> Water_Temp.:104.22, PWM:40
16:57:11.529 -> Water_Temp.:104.22, PWM:40
16:57:13.077 -> Water_Temp.:104.22, PWM:40
16:57:15.384 -> Water_Temp.:104.22, PWM:40
16:57:19.242 -> Water_Temp.:104.11, PWM:40
16:57:20.791 -> Water_Temp.:104.11, PWM:40
16:57:22.337 -> Water_Temp.:104.11, PWM:40
16:57:24.647 -> Water_Temp.:104.11, PWM:40
16:57:27.714 -> Water_Temp.:104.11, PWM:40
16:57:28.503 -> Water_Temp.:104.22, PWM:40
16:57:30.020 -> Water_Temp.:104.22, PWM:40
16:57:30.814 -> Water_Temp.:104.22, PWM:40
16:57:32.370 -> Water_Temp.:104.22, PWM:40
16:57:33.925 -> Water_Temp.:104.22, PWM:40
16:57:34.717 -> Water_Temp.:104.22, PWM:40
16:57:35.509 -> Water_Temp.:104.22, PWM:40
16:57:36.268 -> Water_Temp.:104.22, PWM:40
16:57:37.059 -> Water_Temp.:104.22, PWM:40
16:57:40.167 -> Water_Temp.:104.34, PWM:40
16:57:40.928 -> Water_Temp.:104.34, PWM:40
16:57:42.513 -> Water_Temp.:104.34, PWM:40
16:57:43.276 -> Water_Temp.:104.34, PWM:40
16:57:45.625 -> Water_Temp.:104.34, PWM:40
16:57:47.179 -> Water_Temp.:104.34, PWM:40
16:57:47.939 -> Water_Temp.:104.34, PWM:40
16:57:48.734 -> Water_Temp.:104.34, PWM:40
16:57:49.527 -> Water_Temp.:104.34, PWM:40
16:57:50.287 -> Water_Temp.:104.34, PWM:40
16:57:52.631 -> Water_Temp.:104.34, PWM:40
16:57:53.391 -> Water_Temp.:104.34, PWM:40
16:57:54.967 -> Water_Temp.:104.34, PWM:40
16:57:57.286 -> Water_Temp.:104.34, PWM:40
16:57:58.076 -> Water_Temp.:104.34, PWM:40
16:57:58.870 -> Water_Temp.:104.34, PWM:40
16:58:01.186 -> Water_Temp.:104.34, PWM:40
16:58:01.979 -> Water_Temp.:104.34, PWM:40
16:58:05.089 -> Water_Temp.:104.22, PWM:40
16:58:07.436 -> Water_Temp.:104.22, PWM:40
16:58:09.729 -> Water_Temp.:104.22, PWM:40
16:58:10.489 -> Water_Temp.:104.22, PWM:40
16:58:11.282 -> Water_Temp.:104.22, PWM:40
16:58:12.075 -> Water_Temp.:104.22, PWM:40
16:58:13.627 -> Water_Temp.:104.22, PWM:40
16:58:14.383 -> Water_Temp.:104.22, PWM:40
16:58:15.176 -> Water_Temp.:104.34, PWM:40
16:58:15.941 -> Water_Temp.:104.22, PWM:40
16:58:17.501 -> Water_Temp.:104.22, PWM:40
16:58:19.091 -> Water_Temp.:104.22, PWM:40
16:58:19.851 -> Water_Temp.:104.22, PWM:40
16:58:20.648 -> Water_Temp.:104.22, PWM:40
16:58:21.409 -> Water_Temp.:104.22, PWM:40
16:58:22.203 -> Water_Temp.:104.11, PWM:40
16:58:22.963 -> Water_Temp.:104.11, PWM:40
16:58:23.753 -> Water_Temp.:104.11, PWM:40
16:58:25.304 -> Water_Temp.:104.11, PWM:40
16:58:26.064 -> Water_Temp.:104.11, PWM:40
16:58:26.861 -> Water_Temp.:104.11, PWM:40
16:58:27.622 -> Water_Temp.:104.11, PWM:40
16:58:29.172 -> Water_Temp.:104.00, PWM:40
16:58:29.966 -> Water_Temp.:104.00, PWM:40
16:58:30.726 -> Water_Temp.:104.00, PWM:40
16:58:31.518 -> Water_Temp.:104.00, PWM:40
16:58:32.282 -> Water_Temp.:104.00, PWM:40
16:58:33.078 -> Water_Temp.:104.00, PWM:40
16:58:33.837 -> Water_Temp.:104.00, PWM:40
16:58:34.631 -> Water_Temp.:104.00, PWM:40
16:58:35.425 -> Water_Temp.:104.00, PWM:40
16:58:36.187 -> Water_Temp.:104.00, PWM:40
16:58:38.501 -> Water_Temp.:104.00, PWM:40
16:58:39.294 -> Water_Temp.:104.11, PWM:40
16:58:40.054 -> Water_Temp.:104.11, PWM:40
16:58:40.847 -> Water_Temp.:104.11, PWM:40
16:58:41.641 -> Water_Temp.:104.11, PWM:40
16:58:43.160 -> Water_Temp.:104.11, PWM:40
16:58:43.951 -> Water_Temp.:104.11, PWM:40
16:58:44.741 -> Water_Temp.:104.11, PWM:40
16:58:45.504 -> Water_Temp.:104.22, PWM:40
16:58:47.057 -> Water_Temp.:104.22, PWM:40
16:58:47.853 -> Water_Temp.:104.22, PWM:40
16:58:48.646 -> Water_Temp.:104.22, PWM:40
16:58:49.408 -> Water_Temp.:104.22, PWM:40
16:58:50.172 -> Water_Temp.:104.22, PWM:40
16:58:51.762 -> Water_Temp.:104.34, PWM:40
16:58:53.319 -> Water_Temp.:104.34, PWM:40
16:58:54.838 -> Water_Temp.:104.34, PWM:40
16:58:55.633 -> Water_Temp.:104.34, PWM:40
16:58:56.393 -> Water_Temp.:104.34, PWM:40
16:58:57.185 -> Water_Temp.:104.34, PWM:40
16:58:57.980 -> Water_Temp.:104.34, PWM:40
16:58:58.740 -> Water_Temp.:104.34, PWM:40
16:59:01.810 -> Water_Temp.:104.34, PWM:40
16:59:02.610 -> Water_Temp.:104.34, PWM:40
16:59:03.376 -> Water_Temp.:104.34, PWM:40
16:59:04.177 -> Water_Temp.:104.34, PWM:40
16:59:04.942 -> Water_Temp.:104.34, PWM:40
16:59:05.734 -> Water_Temp.:104.34, PWM:40
16:59:08.806 -> Water_Temp.:104.22, PWM:40
16:59:09.567 -> Water_Temp.:104.22, PWM:40
16:59:10.360 -> Water_Temp.:104.22, PWM:40
16:59:11.121 -> Water_Temp.:104.22, PWM:40
16:59:11.915 -> Water_Temp.:104.22, PWM:40
16:59:13.471 -> Water_Temp.:104.22, PWM:40
16:59:14.232 -> Water_Temp.:104.22, PWM:40
16:59:15.029 -> Water_Temp.:104.22, PWM:40
16:59:15.786 -> Water_Temp.:104.22, PWM:40
16:59:16.580 -> Water_Temp.:104.22, PWM:40
16:59:19.656 -> Water_Temp.:104.34, PWM:40
16:59:21.976 -> Water_Temp.:104.34, PWM:40
16:59:22.775 -> Water_Temp.:104.34, PWM:40
16:59:23.564 -> Water_Temp.:104.34, PWM:40
16:59:24.322 -> Water_Temp.:104.22, PWM:40
16:59:25.118 -> Water_Temp.:104.22, PWM:40
16:59:26.673 -> Water_Temp.:104.34, PWM:40
16:59:28.960 -> Water_Temp.:104.34, PWM:40
16:59:29.753 -> Water_Temp.:104.34, PWM:40
16:59:30.513 -> Water_Temp.:104.34, PWM:40
16:59:34.415 -> Water_Temp.:104.34, PWM:40
16:59:35.212 -> Water_Temp.:104.34, PWM:40
16:59:36.769 -> Water_Temp.:104.34, PWM:40
16:59:37.530 -> Water_Temp.:104.34, PWM:40
16:59:39.082 -> Water_Temp.:104.34, PWM:40
16:59:40.638 -> Water_Temp.:104.45, PWM:40
16:59:41.432 -> Water_Temp.:104.45, PWM:40
16:59:42.192 -> Water_Temp.:104.45, PWM:40
16:59:43.747 -> Water_Temp.:104.45, PWM:40
16:59:46.091 -> Water_Temp.:104.45, PWM:40
16:59:46.851 -> Water_Temp.:104.45, PWM:40
16:59:47.649 -> Water_Temp.:104.45, PWM:40
16:59:48.440 -> Water_Temp.:104.45, PWM:40
16:59:49.201 -> Water_Temp.:104.45, PWM:40
16:59:49.997 -> Water_Temp.:104.45, PWM:40
16:59:50.759 -> Water_Temp.:104.45, PWM:40
16:59:51.554 -> Water_Temp.:104.45, PWM:40
16:59:52.312 -> Water_Temp.:104.45, PWM:40
16:59:53.103 -> Water_Temp.:104.45, PWM:40
16:59:53.897 -> Water_Temp.:104.45, PWM:40
16:59:54.656 -> Water_Temp.:104.45, PWM:40
16:59:55.453 -> Water_Temp.:104.45, PWM:40
16:59:56.973 -> Water_Temp.:104.45, PWM:40
16:59:57.768 -> Water_Temp.:104.45, PWM:40
16:59:58.563 -> Water_Temp.:104.45, PWM:40
17:00:02.428 -> Water_Temp.:104.45, PWM:40
17:00:03.186 -> Water_Temp.:104.45, PWM:40
17:00:04.769 -> Water_Temp.:104.45, PWM:40
17:00:05.526 -> Water_Temp.:104.45, PWM:40
17:00:07.074 -> Water_Temp.:104.45, PWM:40
17:00:07.863 -> Water_Temp.:104.45, PWM:40
17:00:09.413 -> Water_Temp.:104.45, PWM:40
17:00:10.200 -> Water_Temp.:104.45, PWM:40
17:00:12.512 -> Water_Temp.:104.45, PWM:40
17:00:15.576 -> Water_Temp.:104.45, PWM:40
17:00:16.338 -> Water_Temp.:104.45, PWM:40
17:00:17.132 -> Water_Temp.:104.45, PWM:40
17:00:18.684 -> Water_Temp.:104.45, PWM:40
17:00:20.234 -> Water_Temp.:104.45, PWM:40
17:00:21.758 -> Water_Temp.:104.45, PWM:40
17:00:22.520 -> Water_Temp.:104.45, PWM:40
17:00:23.315 -> Water_Temp.:104.45, PWM:40
17:00:24.109 -> Water_Temp.:104.45, PWM:40
17:00:25.655 -> Water_Temp.:104.45, PWM:40
17:00:27.206 -> Water_Temp.:104.56, PWM:40
17:00:29.522 -> Water_Temp.:104.56, PWM:40
17:00:30.283 -> Water_Temp.:104.56, PWM:40
17:00:33.362 -> Water_Temp.:104.56, PWM:40
17:00:34.151 -> Water_Temp.:104.56, PWM:40
17:00:34.913 -> Water_Temp.:104.56, PWM:40
17:00:36.465 -> Water_Temp.:104.56, PWM:40
17:00:37.221 -> Water_Temp.:104.56, PWM:40
17:00:38.010 -> Water_Temp.:104.56, PWM:40
17:00:38.799 -> Water_Temp.:104.56, PWM:40
17:00:39.588 -> Water_Temp.:104.56, PWM:40
17:00:40.343 -> Water_Temp.:104.56, PWM:40
17:00:41.135 -> Water_Temp.:104.56, PWM:40
17:00:41.928 -> Water_Temp.:104.56, PWM:40
17:00:44.229 -> Water_Temp.:104.56, PWM:40
17:00:45.021 -> Water_Temp.:104.56, PWM:40
17:00:46.571 -> Water_Temp.:104.56, PWM:40
17:00:48.123 -> Water_Temp.:104.56, PWM:40
17:00:48.916 -> Water_Temp.:104.56, PWM:40
17:00:49.709 -> Water_Temp.:104.67, PWM:40
17:00:51.270 -> Water_Temp.:104.67, PWM:40
17:00:52.032 -> Water_Temp.:104.67, PWM:40
17:00:58.961 -> Water_Temp.:104.67, PWM:40
17:00:59.720 -> Water_Temp.:104.67, PWM:40
17:01:00.512 -> Water_Temp.:104.67, PWM:40
17:01:01.309 -> Water_Temp.:104.67, PWM:40
17:01:02.071 -> Water_Temp.:104.67, PWM:40
17:01:03.623 -> Water_Temp.:104.67, PWM:40
17:01:04.418 -> Water_Temp.:104.67, PWM:40
17:01:06.735 -> Water_Temp.:104.67, PWM:40
17:01:07.529 -> Water_Temp.:104.67, PWM:40
17:01:08.325 -> Water_Temp.:104.67, PWM:40
17:01:09.885 -> Water_Temp.:104.67, PWM:40
17:01:10.643 -> Water_Temp.:104.67, PWM:40
17:01:11.436 -> Water_Temp.:104.67, PWM:40
17:01:12.196 -> Water_Temp.:104.67, PWM:40
17:01:16.070 -> Water_Temp.:104.56, PWM:40
17:01:16.831 -> Water_Temp.:104.56, PWM:40
17:01:17.628 -> Water_Temp.:104.56, PWM:40
17:01:18.387 -> Water_Temp.:104.56, PWM:40
17:01:19.180 -> Water_Temp.:104.56, PWM:40
17:01:19.944 -> Water_Temp.:104.56, PWM:40
17:01:23.052 -> Water_Temp.:104.56, PWM:40
17:01:23.847 -> Water_Temp.:104.56, PWM:40
17:01:24.608 -> Water_Temp.:104.56, PWM:40
17:01:26.956 -> Water_Temp.:104.56, PWM:40
17:01:27.750 -> Water_Temp.:104.56, PWM:40
17:01:28.511 -> Water_Temp.:104.56, PWM:40
17:01:30.819 -> Water_Temp.:104.56, PWM:40
17:01:31.619 -> Water_Temp.:104.67, PWM:40
17:01:32.379 -> Water_Temp.:104.67, PWM:40
17:01:33.169 -> Water_Temp.:104.67, PWM:40
17:01:33.961 -> Water_Temp.:104.67, PWM:40
17:01:34.717 -> Water_Temp.:104.67, PWM:40
17:01:35.512 -> Water_Temp.:104.67, PWM:40
17:01:36.275 -> Water_Temp.:104.67, PWM:40
17:01:37.065 -> Water_Temp.:104.67, PWM:40
17:01:37.857 -> Water_Temp.:104.67, PWM:40
17:01:38.615 -> Water_Temp.:104.67, PWM:40
17:01:39.411 -> Water_Temp.:104.67, PWM:40
17:01:40.173 -> Water_Temp.:104.67, PWM:40
17:01:40.969 -> Water_Temp.:104.67, PWM:40
17:01:42.519 -> Water_Temp.:104.67, PWM:40
17:01:43.275 -> Water_Temp.:104.67, PWM:40
17:01:44.068 -> Water_Temp.:104.67, PWM:40
17:01:44.862 -> Water_Temp.:104.67, PWM:40
17:01:46.384 -> Water_Temp.:104.67, PWM:40
17:01:47.940 -> Water_Temp.:104.67, PWM:40
17:01:48.733 -> Water_Temp.:104.67, PWM:40
17:01:49.494 -> Water_Temp.:104.67, PWM:40
17:01:50.288 -> Water_Temp.:104.67, PWM:40
17:01:51.080 -> Water_Temp.:104.67, PWM:40
17:01:51.843 -> Water_Temp.:104.67, PWM:40
17:01:52.635 -> Water_Temp.:104.67, PWM:40
17:01:54.907 -> Water_Temp.:104.67, PWM:40
17:01:55.699 -> Water_Temp.:104.67, PWM:40
17:01:56.460 -> Water_Temp.:104.67, PWM:40
17:01:57.255 -> Water_Temp.:104.67, PWM:40
17:01:58.806 -> Water_Temp.:104.67, PWM:40
17:01:59.602 -> Water_Temp.:104.67, PWM:40
17:02:00.361 -> Water_Temp.:104.67, PWM:40
17:02:03.479 -> Water_Temp.:104.79, PWM:40
17:02:05.764 -> Water_Temp.:104.79, PWM:40
17:02:07.324 -> Water_Temp.:104.79, PWM:40
17:02:08.877 -> Water_Temp.:104.79, PWM:40
17:02:09.636 -> Water_Temp.:104.79, PWM:40
17:02:11.189 -> Water_Temp.:104.79, PWM:40
17:02:11.978 -> Water_Temp.:104.79, PWM:40
17:02:12.769 -> Water_Temp.:104.79, PWM:40
17:02:13.529 -> Water_Temp.:104.79, PWM:40
17:02:14.324 -> Water_Temp.:104.79, PWM:40
17:02:15.875 -> Water_Temp.:104.79, PWM:40
17:02:17.425 -> Water_Temp.:104.79, PWM:40
17:02:19.004 -> Water_Temp.:104.79, PWM:40
17:02:19.766 -> Water_Temp.:104.79, PWM:40
17:02:20.561 -> Water_Temp.:104.79, PWM:40
17:02:21.319 -> Water_Temp.:104.79, PWM:40
17:02:22.866 -> Water_Temp.:104.79, PWM:40
17:02:23.663 -> Water_Temp.:104.79, PWM:40
17:02:25.214 -> Water_Temp.:104.79, PWM:40
17:02:25.975 -> Water_Temp.:104.79, PWM:40
17:02:26.767 -> Water_Temp.:104.79, PWM:40
17:02:27.555 -> Water_Temp.:104.67, PWM:40
17:02:28.315 -> Water_Temp.:104.67, PWM:40
17:02:29.111 -> Water_Temp.:104.67, PWM:40
17:02:29.902 -> Water_Temp.:104.67, PWM:40
17:02:31.422 -> Water_Temp.:104.67, PWM:40
17:02:32.941 -> Water_Temp.:104.56, PWM:40
17:02:33.736 -> Water_Temp.:104.56, PWM:40
17:02:34.531 -> Water_Temp.:104.56, PWM:40
17:02:35.289 -> Water_Temp.:104.56, PWM:40
17:02:36.081 -> Water_Temp.:104.56, PWM:40
17:02:38.404 -> Water_Temp.:104.67, PWM:40
17:02:39.188 -> Water_Temp.:104.67, PWM:40
17:02:42.316 -> Water_Temp.:104.67, PWM:40
17:02:43.073 -> Water_Temp.:104.67, PWM:40
17:02:48.499 -> Water_Temp.:104.67, PWM:40
17:02:49.260 -> Water_Temp.:104.67, PWM:40
17:02:50.056 -> Water_Temp.:104.67, PWM:40
17:02:50.820 -> Water_Temp.:104.67, PWM:40
17:02:52.369 -> Water_Temp.:104.67, PWM:40
17:02:53.163 -> Water_Temp.:104.67, PWM:40
17:02:53.923 -> Water_Temp.:104.67, PWM:40
17:02:54.710 -> Water_Temp.:104.67, PWM:40
17:02:55.501 -> Water_Temp.:104.67, PWM:40
17:02:56.261 -> Water_Temp.:104.67, PWM:40
17:02:57.054 -> Water_Temp.:104.67, PWM:40
17:02:57.846 -> Water_Temp.:104.67, PWM:40
17:02:58.606 -> Water_Temp.:104.67, PWM:40
17:03:00.948 -> Water_Temp.:104.67, PWM:40
17:03:01.707 -> Water_Temp.:104.67, PWM:40
17:03:03.288 -> Water_Temp.:104.67, PWM:40
17:03:04.043 -> Water_Temp.:104.56, PWM:40
17:03:04.836 -> Water_Temp.:104.56, PWM:40
17:03:05.628 -> Water_Temp.:104.56, PWM:40
17:03:06.390 -> Water_Temp.:104.56, PWM:40
17:03:07.178 -> Water_Temp.:104.56, PWM:40
17:03:08.733 -> Water_Temp.:104.56, PWM:40
17:03:11.047 -> Water_Temp.:104.56, PWM:40
17:03:11.805 -> Water_Temp.:104.56, PWM:40
17:03:12.600 -> Water_Temp.:104.56, PWM:40
17:03:13.393 -> Water_Temp.:104.56, PWM:40
17:03:15.693 -> Water_Temp.:104.45, PWM:40
17:03:17.247 -> Water_Temp.:104.45, PWM:40
17:03:18.004 -> Water_Temp.:104.45, PWM:40
17:03:18.795 -> Water_Temp.:104.45, PWM:40
17:03:20.344 -> Water_Temp.:104.45, PWM:40
17:03:21.902 -> Water_Temp.:104.45, PWM:40
17:03:22.698 -> Water_Temp.:104.45, PWM:40
17:03:23.460 -> Water_Temp.:106.14, PWM:40
17:03:24.258 -> Water_Temp.:108.16, PWM:40
17:03:25.813 -> Water_Temp.:105.35, PWM:40
17:03:26.575 -> Water_Temp.:105.24, PWM:40
17:03:27.370 -> Water_Temp.:105.01, PWM:40
17:03:31.202 -> Water_Temp.:104.79, PWM:40
17:03:32.754 -> Water_Temp.:104.79, PWM:40
17:03:33.511 -> Water_Temp.:104.79, PWM:40
17:03:34.307 -> Water_Temp.:104.79, PWM:40
17:03:35.861 -> Water_Temp.:104.79, PWM:40
17:03:37.417 -> Water_Temp.:104.67, PWM:40
17:03:38.973 -> Water_Temp.:104.67, PWM:40
17:03:39.726 -> Water_Temp.:104.67, PWM:40
17:03:40.521 -> Water_Temp.:104.67, PWM:40
17:03:41.318 -> Water_Temp.:104.67, PWM:40
17:03:43.627 -> Water_Temp.:104.67, PWM:40
17:03:45.945 -> Water_Temp.:104.67, PWM:40
17:03:46.706 -> Water_Temp.:104.67, PWM:40
17:03:47.500 -> Water_Temp.:104.67, PWM:40
17:03:49.059 -> Water_Temp.:104.67, PWM:40
17:03:51.376 -> Water_Temp.:104.67, PWM:40
17:03:52.171 -> Water_Temp.:104.67, PWM:40
17:03:53.723 -> Water_Temp.:104.56, PWM:40
17:03:55.277 -> Water_Temp.:104.56, PWM:40
17:03:56.035 -> Water_Temp.:104.56, PWM:40
17:03:56.829 -> Water_Temp.:104.56, PWM:40
17:03:57.619 -> Water_Temp.:104.56, PWM:40
17:04:00.691 -> Water_Temp.:104.56, PWM:40
17:04:01.485 -> Water_Temp.:104.56, PWM:40
17:04:02.242 -> Water_Temp.:104.56, PWM:40
17:04:03.037 -> Water_Temp.:104.67, PWM:40
17:04:03.795 -> Water_Temp.:104.67, PWM:40
17:04:04.589 -> Water_Temp.:104.67, PWM:40
17:04:07.689 -> Water_Temp.:104.67, PWM:40
17:04:08.483 -> Water_Temp.:104.67, PWM:40
17:04:11.592 -> Water_Temp.:104.67, PWM:40
17:04:13.148 -> Water_Temp.:104.67, PWM:40
17:04:13.940 -> Water_Temp.:104.67, PWM:40
17:04:15.494 -> Water_Temp.:104.67, PWM:40
17:04:16.256 -> Water_Temp.:104.67, PWM:40
17:04:17.048 -> Water_Temp.:104.67, PWM:40
17:04:17.841 -> Water_Temp.:104.67, PWM:40
17:04:18.599 -> Water_Temp.:104.67, PWM:40
17:04:19.391 -> Water_Temp.:104.67, PWM:40
17:04:20.148 -> Water_Temp.:104.67, PWM:40
17:04:20.944 -> Water_Temp.:104.67, PWM:40
17:04:23.259 -> Water_Temp.:104.67, PWM:40
17:04:24.019 -> Water_Temp.:104.67, PWM:40
17:04:24.815 -> Water_Temp.:104.67, PWM:40
17:04:25.608 -> Water_Temp.:104.67, PWM:40
17:04:27.925 -> Water_Temp.:104.67, PWM:40
17:04:28.717 -> Water_Temp.:104.67, PWM:40
17:04:30.266 -> Water_Temp.:104.67, PWM:40
17:04:31.826 -> Water_Temp.:104.67, PWM:40
17:04:32.621 -> Water_Temp.:104.67, PWM:40
17:04:34.937 -> Water_Temp.:108.39, PWM:40
17:04:35.729 -> Water_Temp.:105.91, PWM:40
17:04:36.494 -> Water_Temp.:105.91, PWM:40
17:04:37.285 -> Water_Temp.:105.35, PWM:40
17:04:38.050 -> Water_Temp.:105.24, PWM:40
17:04:39.635 -> Water_Temp.:105.01, PWM:40
17:04:40.398 -> Water_Temp.:105.01, PWM:40
17:04:41.954 -> Water_Temp.:104.90, PWM:40
17:04:42.746 -> Water_Temp.:104.90, PWM:40
17:04:43.507 -> Water_Temp.:104.90, PWM:40
17:04:44.304 -> Water_Temp.:104.90, PWM:40
17:04:45.860 -> Water_Temp.:104.90, PWM:40
17:04:46.624 -> Water_Temp.:104.90, PWM:40
17:04:47.413 -> Water_Temp.:104.90, PWM:40
17:04:48.961 -> Water_Temp.:104.90, PWM:40
17:04:52.025 -> Water_Temp.:104.79, PWM:40
17:04:53.573 -> Water_Temp.:104.79, PWM:40
17:04:55.124 -> Water_Temp.:104.79, PWM:40
17:04:55.882 -> Water_Temp.:104.79, PWM:40
17:04:56.676 -> Water_Temp.:104.79, PWM:40
17:04:57.471 -> Water_Temp.:104.79, PWM:40
17:04:58.230 -> Water_Temp.:104.79, PWM:40
17:04:59.022 -> Water_Temp.:104.79, PWM:40
17:05:01.334 -> Water_Temp.:104.79, PWM:40
17:05:02.886 -> Water_Temp.:106.36, PWM:40
17:05:03.676 -> Water_Temp.:108.39, PWM:40
17:05:05.224 -> Water_Temp.:105.57, PWM:40
17:05:06.021 -> Water_Temp.:105.35, PWM:40
17:05:06.784 -> Water_Temp.:105.24, PWM:40
17:05:08.338 -> Water_Temp.:105.01, PWM:40
17:05:09.093 -> Water_Temp.:104.90, PWM:40
17:05:09.886 -> Water_Temp.:104.90, PWM:40
17:05:12.199 -> Water_Temp.:104.79, PWM:40
17:05:12.993 -> Water_Temp.:104.79, PWM:40
17:05:13.757 -> Water_Temp.:104.79, PWM:40
17:05:14.520 -> Water_Temp.:104.79, PWM:40
17:05:15.314 -> Water_Temp.:104.79, PWM:40
17:05:16.865 -> Water_Temp.:104.79, PWM:40
17:05:17.660 -> Water_Temp.:104.79, PWM:40
17:05:18.422 -> Water_Temp.:104.79, PWM:40
17:05:19.219 -> Water_Temp.:104.79, PWM:40
17:05:20.740 -> Water_Temp.:104.79, PWM:40
17:05:21.537 -> Water_Temp.:104.79, PWM:40
17:05:22.330 -> Water_Temp.:104.79, PWM:40
17:05:23.092 -> Water_Temp.:104.79, PWM:40
17:05:23.888 -> Water_Temp.:104.79, PWM:40
17:05:24.649 -> Water_Temp.:104.79, PWM:40
17:05:25.446 -> Water_Temp.:104.79, PWM:40
17:05:26.972 -> Water_Temp.:104.79, PWM:40
17:05:27.733 -> Water_Temp.:104.79, PWM:40
17:05:28.527 -> Water_Temp.:104.79, PWM:40
17:05:29.320 -> Water_Temp.:104.79, PWM:40
17:05:30.871 -> Water_Temp.:104.79, PWM:40
17:05:32.429 -> Water_Temp.:104.79, PWM:40
17:05:33.191 -> Water_Temp.:104.79, PWM:40
17:05:33.986 -> Water_Temp.:104.79, PWM:40
17:05:35.504 -> Water_Temp.:104.79, PWM:40
17:05:37.064 -> Water_Temp.:104.79, PWM:40
17:05:40.174 -> Water_Temp.:104.79, PWM:40
17:05:40.938 -> Water_Temp.:104.79, PWM:40
17:05:41.733 -> Water_Temp.:104.79, PWM:40
17:05:42.492 -> Water_Temp.:104.79, PWM:40
17:05:43.289 -> Water_Temp.:104.79, PWM:40
17:05:44.051 -> Water_Temp.:104.79, PWM:40
17:05:46.373 -> Water_Temp.:104.79, PWM:40
17:05:47.168 -> Water_Temp.:104.79, PWM:40
17:05:47.930 -> Water_Temp.:104.79, PWM:40
17:05:48.726 -> Water_Temp.:104.79, PWM:40
17:05:49.490 -> Water_Temp.:104.79, PWM:40
17:05:50.285 -> Water_Temp.:104.79, PWM:40
17:05:51.046 -> Water_Temp.:104.79, PWM:40
17:05:52.618 -> Water_Temp.:104.79, PWM:40
17:05:54.170 -> Water_Temp.:104.79, PWM:40
17:05:56.513 -> Water_Temp.:104.79, PWM:40
17:05:58.054 -> Water_Temp.:104.79, PWM:40
17:05:58.868 -> Water_Temp.:104.79, PWM:40
17:06:01.923 -> Water_Temp.:104.90, PWM:40
17:06:02.712 -> Water_Temp.:104.90, PWM:40
17:06:03.500 -> Water_Temp.:104.90, PWM:40
17:06:04.285 -> Water_Temp.:104.90, PWM:40
17:06:05.831 -> Water_Temp.:104.90, PWM:40
17:06:06.625 -> Water_Temp.:104.90, PWM:40
17:06:07.412 -> Water_Temp.:104.90, PWM:40
17:06:08.170 -> Water_Temp.:104.90, PWM:40
17:06:08.960 -> Water_Temp.:104.90, PWM:40
17:06:09.724 -> Water_Temp.:104.90, PWM:40
17:06:10.515 -> Water_Temp.:104.90, PWM:40
17:06:12.859 -> Water_Temp.:104.90, PWM:40
17:06:14.414 -> Water_Temp.:104.90, PWM:40
17:06:15.205 -> Water_Temp.:104.79, PWM:40
17:06:15.964 -> Water_Temp.:104.90, PWM:40
17:06:16.757 -> Water_Temp.:104.90, PWM:40
17:06:20.583 -> Water_Temp.:104.90, PWM:40
17:06:21.376 -> Water_Temp.:104.90, PWM:40
17:06:22.168 -> Water_Temp.:104.90, PWM:40
17:06:22.926 -> Water_Temp.:104.90, PWM:40
17:06:23.720 -> Water_Temp.:104.90, PWM:40
17:06:26.035 -> Water_Temp.:104.90, PWM:40
17:06:26.827 -> Water_Temp.:104.90, PWM:40
17:06:27.622 -> Water_Temp.:104.90, PWM:40
17:06:28.383 -> Water_Temp.:104.90, PWM:40
17:06:29.178 -> Water_Temp.:104.90, PWM:40
17:06:29.938 -> Water_Temp.:104.90, PWM:40
17:06:31.492 -> Water_Temp.:104.90, PWM:40
17:06:32.282 -> Water_Temp.:104.90, PWM:40
17:06:34.597 -> Water_Temp.:104.90, PWM:40
17:06:35.392 -> Water_Temp.:104.90, PWM:40
17:06:36.151 -> Water_Temp.:104.90, PWM:40
17:06:38.464 -> Water_Temp.:104.90, PWM:40
17:06:39.258 -> Water_Temp.:104.90, PWM:40
17:06:40.018 -> Water_Temp.:104.90, PWM:40
17:06:40.814 -> Water_Temp.:104.90, PWM:40
17:06:42.332 -> Water_Temp.:104.90, PWM:40
17:06:43.884 -> Water_Temp.:104.90, PWM:40
17:06:44.644 -> Water_Temp.:104.90, PWM:40
17:06:45.435 -> Water_Temp.:104.90, PWM:40
17:06:46.224 -> Water_Temp.:104.90, PWM:40
17:06:48.561 -> Water_Temp.:104.90, PWM:40
17:06:49.319 -> Water_Temp.:104.90, PWM:40
17:06:50.114 -> Water_Temp.:104.90, PWM:40
17:06:52.401 -> Water_Temp.:104.90, PWM:40
17:06:53.196 -> Water_Temp.:104.90, PWM:40
17:06:54.717 -> Water_Temp.:104.90, PWM:40
17:06:55.510 -> Water_Temp.:104.90, PWM:40
17:06:56.307 -> Water_Temp.:104.90, PWM:40
17:06:57.067 -> Water_Temp.:104.90, PWM:40
17:06:58.627 -> Water_Temp.:104.90, PWM:40
17:06:59.421 -> Water_Temp.:104.90, PWM:40
17:07:00.968 -> Water_Temp.:104.90, PWM:40
17:07:01.728 -> Water_Temp.:104.90, PWM:40
17:07:02.521 -> Water_Temp.:104.90, PWM:40
17:07:03.287 -> Water_Temp.:104.90, PWM:40
17:07:04.836 -> Water_Temp.:104.90, PWM:40
17:07:06.358 -> Water_Temp.:104.90, PWM:40
17:07:07.149 -> Water_Temp.:104.90, PWM:40
17:07:10.253 -> Water_Temp.:104.90, PWM:40
17:07:11.011 -> Water_Temp.:104.90, PWM:40
17:07:13.323 -> Water_Temp.:104.90, PWM:40
17:07:14.080 -> Water_Temp.:104.90, PWM:40
17:07:14.872 -> Water_Temp.:106.59, PWM:40
17:07:15.669 -> Water_Temp.:108.61, PWM:40
17:07:16.431 -> Water_Temp.:106.25, PWM:40
17:07:17.224 -> Water_Temp.:105.80, PWM:40
17:07:18.750 -> Water_Temp.:105.46, PWM:40
17:07:19.540 -> Water_Temp.:105.35, PWM:40
17:07:20.300 -> Water_Temp.:105.35, PWM:40
17:07:21.097 -> Water_Temp.:105.24, PWM:40
17:07:21.856 -> Water_Temp.:105.24, PWM:40
17:07:22.652 -> Water_Temp.:105.24, PWM:40
17:07:23.413 -> Water_Temp.:105.12, PWM:40
17:07:24.205 -> Water_Temp.:105.12, PWM:40
17:07:25.000 -> Water_Temp.:105.12, PWM:40
17:07:25.759 -> Water_Temp.:105.12, PWM:40
17:07:26.550 -> Water_Temp.:105.01, PWM:40
17:07:28.105 -> Water_Temp.:105.01, PWM:40
17:07:28.868 -> Water_Temp.:105.01, PWM:40
17:07:29.665 -> Water_Temp.:105.01, PWM:40
17:07:30.424 -> Water_Temp.:105.01, PWM:40
17:07:31.218 -> Water_Temp.:105.01, PWM:40
17:07:31.980 -> Water_Temp.:105.01, PWM:40
17:07:35.054 -> Water_Temp.:105.01, PWM:40
17:07:35.844 -> Water_Temp.:105.01, PWM:40
17:07:36.633 -> Water_Temp.:105.01, PWM:40
17:07:37.388 -> Water_Temp.:105.01, PWM:40
17:07:38.942 -> Water_Temp.:105.01, PWM:40
17:07:40.501 -> Water_Temp.:105.01, PWM:40
17:07:41.295 -> Water_Temp.:105.01, PWM:40
17:07:42.055 -> Water_Temp.:105.01, PWM:40
17:07:45.194 -> Water_Temp.:105.01, PWM:40
17:07:45.953 -> Water_Temp.:105.01, PWM:40
17:07:47.514 -> Water_Temp.:105.01, PWM:40
17:07:48.311 -> Water_Temp.:105.01, PWM:40
17:07:49.075 -> Water_Temp.:105.01, PWM:40
17:07:50.658 -> Water_Temp.:105.01, PWM:40
17:07:51.417 -> Water_Temp.:105.01, PWM:40
17:07:52.972 -> Water_Temp.:105.01, PWM:40
17:07:56.047 -> Water_Temp.:105.01, PWM:40
17:07:56.803 -> Water_Temp.:105.01, PWM:40
17:07:57.595 -> Water_Temp.:105.01, PWM:40
17:07:59.908 -> Water_Temp.:105.01, PWM:40
17:08:00.670 -> Water_Temp.:105.01, PWM:40
17:08:01.467 -> Water_Temp.:105.01, PWM:40
17:08:02.230 -> Water_Temp.:105.01, PWM:40
17:08:03.022 -> Water_Temp.:105.01, PWM:40
17:08:04.573 -> Water_Temp.:105.01, PWM:40
17:08:05.364 -> Water_Temp.:105.01, PWM:40
17:08:06.157 -> Water_Temp.:105.01, PWM:40
17:08:07.673 -> Water_Temp.:105.01, PWM:40
17:08:08.464 -> Water_Temp.:105.01, PWM:40
17:08:09.220 -> Water_Temp.:105.01, PWM:40
17:08:11.565 -> Water_Temp.:105.01, PWM:40
17:08:13.847 -> Water_Temp.:105.01, PWM:40
17:08:14.642 -> Water_Temp.:105.01, PWM:40
17:08:15.436 -> Water_Temp.:105.01, PWM:40
17:08:16.200 -> Water_Temp.:105.01, PWM:40
17:08:16.994 -> Water_Temp.:105.01, PWM:40
17:08:17.753 -> Water_Temp.:105.01, PWM:40
17:08:19.310 -> Water_Temp.:105.01, PWM:40
17:08:20.103 -> Water_Temp.:105.01, PWM:40
17:08:20.867 -> Water_Temp.:105.01, PWM:40
17:08:22.422 -> Water_Temp.:105.01, PWM:40
17:08:23.976 -> Water_Temp.:105.01, PWM:40
17:08:24.734 -> Water_Temp.:105.01, PWM:40
17:08:25.529 -> Water_Temp.:105.01, PWM:40
17:08:27.054 -> Water_Temp.:105.01, PWM:40
17:08:27.847 -> Water_Temp.:105.01, PWM:44
17:08:28.641 -> Water_Temp.:105.01, PWM:44
17:08:30.193 -> Water_Temp.:105.01, PWM:44
17:08:30.955 -> Water_Temp.:105.01, PWM:44
17:08:32.511 -> Water_Temp.:105.01, PWM:44
17:08:33.307 -> Water_Temp.:105.01, PWM:44
17:08:34.068 -> Water_Temp.:105.01, PWM:44
17:08:35.618 -> Water_Temp.:105.01, PWM:44
17:08:37.169 -> Water_Temp.:105.01, PWM:44
17:08:37.965 -> Water_Temp.:105.01, PWM:44
17:08:38.761 -> Water_Temp.:105.01, PWM:44
17:08:41.076 -> Water_Temp.:105.01, PWM:44
17:08:41.836 -> Water_Temp.:105.01, PWM:44
17:08:43.393 -> Water_Temp.:105.01, PWM:44
17:08:44.188 -> Water_Temp.:105.01, PWM:44
17:08:44.948 -> Water_Temp.:105.01, PWM:44
17:08:45.743 -> Water_Temp.:105.01, PWM:44
17:08:46.539 -> Water_Temp.:105.01, PWM:44
17:08:48.058 -> Water_Temp.:105.01, PWM:44
17:08:48.849 -> Water_Temp.:105.01, PWM:44
17:08:49.641 -> Water_Temp.:105.01, PWM:44
17:08:50.397 -> Water_Temp.:105.01, PWM:44
17:08:51.192 -> Water_Temp.:105.01, PWM:44
17:08:52.709 -> Water_Temp.:105.01, PWM:44
17:08:53.507 -> Water_Temp.:105.01, PWM:44
17:08:54.293 -> Water_Temp.:105.01, PWM:44
17:08:55.055 -> Water_Temp.:105.01, PWM:44
17:08:55.851 -> Water_Temp.:105.01, PWM:44
17:08:57.401 -> Water_Temp.:105.01, PWM:44
17:08:58.194 -> Water_Temp.:105.01, PWM:44
17:08:58.952 -> Water_Temp.:105.01, PWM:44
17:08:59.748 -> Water_Temp.:105.01, PWM:44
17:09:00.507 -> Water_Temp.:105.01, PWM:44
17:09:02.064 -> Water_Temp.:105.01, PWM:44
17:09:02.827 -> Water_Temp.:105.01, PWM:44
17:09:03.622 -> Water_Temp.:105.01, PWM:44
17:09:05.172 -> Water_Temp.:105.01, PWM:44
17:09:05.964 -> Water_Temp.:105.01, PWM:44
17:09:06.722 -> Water_Temp.:104.90, PWM:44
17:09:07.513 -> Water_Temp.:105.01, PWM:44
17:09:09.061 -> Water_Temp.:105.01, PWM:44
17:09:09.853 -> Water_Temp.:105.01, PWM:44
17:09:10.616 -> Water_Temp.:104.90, PWM:44
17:09:11.411 -> Water_Temp.:104.90, PWM:44
17:09:14.480 -> Water_Temp.:104.90, PWM:44
17:09:16.033 -> Water_Temp.:104.90, PWM:44
17:09:17.582 -> Water_Temp.:104.90, PWM:44
17:09:19.893 -> Water_Temp.:104.90, PWM:44
17:09:20.689 -> Water_Temp.:104.90, PWM:44
17:09:21.450 -> Water_Temp.:104.90, PWM:44
17:09:23.031 -> Water_Temp.:104.90, PWM:44
17:09:23.789 -> Water_Temp.:104.90, PWM:44
17:09:25.342 -> Water_Temp.:104.90, PWM:44
17:09:26.139 -> Water_Temp.:104.90, PWM:44
17:09:27.697 -> Water_Temp.:108.61, PWM:44
17:09:28.453 -> Water_Temp.:106.14, PWM:44
17:09:29.245 -> Water_Temp.:105.80, PWM:44
17:09:30.036 -> Water_Temp.:105.57, PWM:44
17:09:30.798 -> Water_Temp.:105.46, PWM:44
17:09:31.591 -> Water_Temp.:105.35, PWM:44
17:09:32.383 -> Water_Temp.:105.35, PWM:44
17:09:33.143 -> Water_Temp.:105.24, PWM:44
17:09:34.700 -> Water_Temp.:105.24, PWM:44
17:09:35.457 -> Water_Temp.:105.12, PWM:44
17:09:36.254 -> Water_Temp.:105.12, PWM:44
17:09:37.049 -> Water_Temp.:105.01, PWM:44
17:09:38.602 -> Water_Temp.:105.01, PWM:44
17:09:39.361 -> Water_Temp.:105.01, PWM:44
17:09:40.920 -> Water_Temp.:105.01, PWM:44
17:09:41.713 -> Water_Temp.:105.01, PWM:44
17:09:44.023 -> Water_Temp.:105.01, PWM:44
17:09:45.546 -> Water_Temp.:105.01, PWM:44
17:09:47.900 -> Water_Temp.:105.01, PWM:44
17:09:49.449 -> Water_Temp.:105.01, PWM:44
17:09:50.209 -> Water_Temp.:105.01, PWM:44
17:09:51.006 -> Water_Temp.:105.01, PWM:44
17:09:51.766 -> Water_Temp.:105.01, PWM:44
17:09:54.086 -> Water_Temp.:104.90, PWM:44
17:09:54.880 -> Water_Temp.:104.90, PWM:44
17:09:55.641 -> Water_Temp.:104.90, PWM:44
17:09:57.197 -> Water_Temp.:104.90, PWM:44
17:09:57.993 -> Water_Temp.:104.90, PWM:44
17:09:59.549 -> Water_Temp.:104.90, PWM:44
17:10:00.309 -> Water_Temp.:104.90, PWM:44
17:10:01.107 -> Water_Temp.:104.90, PWM:44
17:10:03.454 -> Water_Temp.:104.90, PWM:44
17:10:06.531 -> Water_Temp.:104.90, PWM:44
17:10:08.813 -> Water_Temp.:104.90, PWM:44
17:10:09.608 -> Water_Temp.:104.90, PWM:44
17:10:10.368 -> Water_Temp.:104.90, PWM:44
17:10:11.164 -> Water_Temp.:105.01, PWM:44
17:10:11.955 -> Water_Temp.:105.01, PWM:44
17:10:13.511 -> Water_Temp.:105.01, PWM:44
17:10:14.272 -> Water_Temp.:105.01, PWM:44
17:10:15.067 -> Water_Temp.:105.01, PWM:44
17:10:16.619 -> Water_Temp.:105.01, PWM:44
17:10:17.378 -> Water_Temp.:105.01, PWM:44
17:10:18.171 -> Water_Temp.:105.01, PWM:44
17:10:18.932 -> Water_Temp.:105.01, PWM:44
17:10:19.727 -> Water_Temp.:105.01, PWM:44
17:10:20.523 -> Water_Temp.:105.01, PWM:44
17:10:21.282 -> Water_Temp.:105.01, PWM:44
17:10:22.077 -> Water_Temp.:105.01, PWM:44
17:10:25.909 -> Water_Temp.:105.01, PWM:44
17:10:26.707 -> Water_Temp.:105.01, PWM:44
17:10:27.467 -> Water_Temp.:105.01, PWM:44
17:10:28.260 -> Water_Temp.:105.01, PWM:44
17:10:29.811 -> Water_Temp.:105.01, PWM:44
17:10:30.571 -> Water_Temp.:104.90, PWM:44
17:10:32.126 -> Water_Temp.:105.01, PWM:44
17:10:32.917 -> Water_Temp.:105.01, PWM:44
17:10:33.679 -> Water_Temp.:105.01, PWM:44
17:10:34.471 -> Water_Temp.:105.01, PWM:44
17:10:35.264 -> Water_Temp.:105.01, PWM:44
17:10:36.783 -> Water_Temp.:105.01, PWM:44
17:10:40.660 -> Water_Temp.:105.01, PWM:44
17:10:41.419 -> Water_Temp.:105.01, PWM:44
17:10:42.214 -> Water_Temp.:105.01, PWM:44
17:10:43.004 -> Water_Temp.:105.01, PWM:44
17:10:44.556 -> Water_Temp.:105.01, PWM:44
17:10:45.345 -> Water_Temp.:105.01, PWM:44
17:10:48.449 -> Water_Temp.:105.01, PWM:44
17:10:49.211 -> Water_Temp.:105.01, PWM:44
17:10:51.565 -> Water_Temp.:105.01, PWM:44
17:10:53.124 -> Water_Temp.:105.12, PWM:44
17:10:53.883 -> Water_Temp.:105.12, PWM:44
17:10:54.678 -> Water_Temp.:105.12, PWM:44
17:10:55.437 -> Water_Temp.:105.12, PWM:44
17:10:56.233 -> Water_Temp.:105.12, PWM:44
17:10:56.997 -> Water_Temp.:105.12, PWM:44
17:10:58.550 -> Water_Temp.:105.12, PWM:44
17:10:59.341 -> Water_Temp.:105.12, PWM:44
17:11:00.096 -> Water_Temp.:105.12, PWM:44
17:11:00.888 -> Water_Temp.:105.12, PWM:44
17:11:03.232 -> Water_Temp.:105.12, PWM:44
17:11:05.543 -> Water_Temp.:105.12, PWM:44
17:11:06.335 -> Water_Temp.:105.12, PWM:44
17:11:07.095 -> Water_Temp.:105.12, PWM:44
17:11:07.890 -> Water_Temp.:105.12, PWM:44
17:11:08.654 -> Water_Temp.:105.12, PWM:44
17:11:10.206 -> Water_Temp.:105.12, PWM:44
17:11:11.005 -> Water_Temp.:105.12, PWM:45
17:11:11.761 -> Water_Temp.:105.12, PWM:45
17:11:12.558 -> Water_Temp.:105.01, PWM:45
17:11:13.354 -> Water_Temp.:105.01, PWM:46
17:11:14.881 -> Water_Temp.:105.01, PWM:45
17:11:15.675 -> Water_Temp.:105.01, PWM:45
17:11:16.470 -> Water_Temp.:105.01, PWM:45
17:11:21.103 -> Water_Temp.:105.01, PWM:45
17:11:21.896 -> Water_Temp.:105.01, PWM:45
17:11:22.651 -> Water_Temp.:105.01, PWM:45
17:11:25.788 -> Water_Temp.:106.70, PWM:45
17:11:27.307 -> Water_Temp.:106.36, PWM:45
17:11:28.102 -> Water_Temp.:105.91, PWM:45
17:11:28.858 -> Water_Temp.:105.69, PWM:46
17:11:33.452 -> Water_Temp.:105.35, PWM:45
17:11:34.211 -> Water_Temp.:105.35, PWM:45
17:11:35.007 -> Water_Temp.:105.35, PWM:45
17:11:35.763 -> Water_Temp.:105.35, PWM:46
17:11:36.556 -> Water_Temp.:105.35, PWM:45
17:11:37.352 -> Water_Temp.:105.35, PWM:45
17:11:38.112 -> Water_Temp.:105.35, PWM:45
17:11:39.634 -> Water_Temp.:105.35, PWM:45
17:11:40.430 -> Water_Temp.:105.35, PWM:45
17:11:41.223 -> Water_Temp.:105.35, PWM:45
17:11:41.986 -> Water_Temp.:105.35, PWM:45
17:11:42.781 -> Water_Temp.:105.35, PWM:46
17:11:43.538 -> Water_Temp.:105.35, PWM:46
17:11:45.121 -> Water_Temp.:105.35, PWM:45
17:11:45.881 -> Water_Temp.:105.35, PWM:46
17:11:47.433 -> Water_Temp.:105.35, PWM:45
17:11:51.302 -> Water_Temp.:105.35, PWM:46
17:11:52.096 -> Water_Temp.:105.35, PWM:45
17:11:52.850 -> Water_Temp.:105.35, PWM:46
17:11:54.401 -> Water_Temp.:105.35, PWM:46
17:11:55.195 -> Water_Temp.:105.35, PWM:45
17:11:55.952 -> Water_Temp.:105.35, PWM:46
17:11:56.747 -> Water_Temp.:105.35, PWM:45
17:11:58.303 -> Water_Temp.:105.35, PWM:45
17:11:59.097 -> Water_Temp.:105.35, PWM:46
17:11:59.856 -> Water_Temp.:105.35, PWM:45
17:12:01.417 -> Water_Temp.:105.35, PWM:46
17:12:02.208 -> Water_Temp.:105.24, PWM:45
17:12:02.970 -> Water_Temp.:105.35, PWM:45
17:12:03.764 -> Water_Temp.:105.35, PWM:45
17:12:04.556 -> Water_Temp.:105.24, PWM:45
17:12:06.078 -> Water_Temp.:105.24, PWM:46
17:12:09.150 -> Water_Temp.:105.24, PWM:46
17:12:09.944 -> Water_Temp.:105.24, PWM:46
17:12:10.704 -> Water_Temp.:105.24, PWM:46
17:12:11.499 -> Water_Temp.:105.24, PWM:46
17:12:12.292 -> Water_Temp.:105.24, PWM:46
17:12:13.055 -> Water_Temp.:105.24, PWM:46
17:12:14.615 -> Water_Temp.:105.35, PWM:45
17:12:15.372 -> Water_Temp.:105.35, PWM:45
17:12:16.172 -> Water_Temp.:105.35, PWM:46
17:12:16.931 -> Water_Temp.:105.35, PWM:45
17:12:17.729 -> Water_Temp.:105.35, PWM:45
17:12:18.523 -> Water_Temp.:105.35, PWM:46
17:12:19.281 -> Water_Temp.:105.35, PWM:45
17:12:20.841 -> Water_Temp.:105.35, PWM:46
17:12:22.360 -> Water_Temp.:105.35, PWM:45
17:12:23.158 -> Water_Temp.:105.35, PWM:45
17:12:23.923 -> Water_Temp.:105.35, PWM:45
17:12:24.713 -> Water_Temp.:105.35, PWM:45
17:12:28.587 -> Water_Temp.:105.35, PWM:45
17:12:29.381 -> Water_Temp.:105.35, PWM:45
17:12:30.938 -> Water_Temp.:105.35, PWM:45
17:12:31.697 -> Water_Temp.:105.35, PWM:46
17:12:34.838 -> Water_Temp.:105.46, PWM:46
17:12:35.597 -> Water_Temp.:105.46, PWM:46
17:12:36.395 -> Water_Temp.:105.46, PWM:45
17:12:40.232 -> Water_Temp.:105.46, PWM:45
17:12:40.995 -> Water_Temp.:105.46, PWM:45
17:12:41.790 -> Water_Temp.:105.46, PWM:46
17:12:42.552 -> Water_Temp.:105.46, PWM:46
17:12:44.138 -> Water_Temp.:105.46, PWM:46
17:12:44.900 -> Water_Temp.:105.46, PWM:45
17:12:45.692 -> Water_Temp.:105.46, PWM:45
17:12:46.455 -> Water_Temp.:105.46, PWM:45
17:12:47.253 -> Water_Temp.:105.46, PWM:45
17:12:48.013 -> Water_Temp.:105.46, PWM:45
17:12:48.810 -> Water_Temp.:105.46, PWM:46
17:12:49.568 -> Water_Temp.:105.46, PWM:45
17:12:50.359 -> Water_Temp.:105.46, PWM:45
17:12:51.154 -> Water_Temp.:105.46, PWM:46
17:12:53.471 -> Water_Temp.:105.46, PWM:45
17:12:54.262 -> Water_Temp.:105.46, PWM:45
17:12:55.020 -> Water_Temp.:105.46, PWM:46
17:12:55.807 -> Water_Temp.:105.46, PWM:45
17:12:56.599 -> Water_Temp.:105.46, PWM:45
17:12:57.358 -> Water_Temp.:105.46, PWM:46
17:12:58.142 -> Water_Temp.:105.46, PWM:45
17:12:58.938 -> Water_Temp.:105.46, PWM:45
17:12:59.694 -> Water_Temp.:105.46, PWM:45
17:13:00.489 -> Water_Temp.:105.46, PWM:45
17:13:02.797 -> Water_Temp.:105.46, PWM:45
17:13:04.350 -> Water_Temp.:105.46, PWM:46
17:13:05.144 -> Water_Temp.:105.46, PWM:45
17:13:05.904 -> Water_Temp.:105.46, PWM:46
17:13:06.694 -> Water_Temp.:105.46, PWM:45
17:13:08.249 -> Water_Temp.:105.46, PWM:45
17:13:09.834 -> Water_Temp.:105.46, PWM:45
17:13:10.597 -> Water_Temp.:105.46, PWM:45
17:13:11.394 -> Water_Temp.:105.46, PWM:45
17:13:12.155 -> Water_Temp.:105.46, PWM:45
17:13:12.946 -> Water_Temp.:105.46, PWM:45
17:13:13.705 -> Water_Temp.:105.46, PWM:46
17:13:16.769 -> Water_Temp.:105.35, PWM:45
17:13:17.562 -> Water_Temp.:105.35, PWM:46
17:13:19.120 -> Water_Temp.:105.35, PWM:45
17:13:22.190 -> Water_Temp.:105.35, PWM:45
17:13:22.952 -> Water_Temp.:105.35, PWM:46
17:13:23.750 -> Water_Temp.:105.46, PWM:45
17:13:24.509 -> Water_Temp.:105.46, PWM:45
17:13:26.850 -> Water_Temp.:105.46, PWM:46
17:13:27.640 -> Water_Temp.:105.46, PWM:45
17:13:28.398 -> Water_Temp.:105.46, PWM:45
17:13:29.191 -> Water_Temp.:105.46, PWM:45
17:13:29.951 -> Water_Temp.:105.35, PWM:45
17:13:30.746 -> Water_Temp.:105.35, PWM:45
17:13:31.509 -> Water_Temp.:105.46, PWM:45
17:13:32.305 -> Water_Temp.:105.46, PWM:46
17:13:34.625 -> Water_Temp.:105.46, PWM:46
17:13:35.421 -> Water_Temp.:105.46, PWM:45
17:13:36.218 -> Water_Temp.:105.46, PWM:45
17:13:37.739 -> Water_Temp.:105.46, PWM:46
17:13:38.532 -> Water_Temp.:105.46, PWM:45
17:13:39.291 -> Water_Temp.:105.46, PWM:46
17:13:40.088 -> Water_Temp.:105.46, PWM:45
17:13:40.884 -> Water_Temp.:105.46, PWM:46
17:13:41.646 -> Water_Temp.:105.46, PWM:45
17:13:42.437 -> Water_Temp.:105.46, PWM:45
17:13:43.960 -> Water_Temp.:105.46, PWM:45
17:13:44.754 -> Water_Temp.:105.46, PWM:45
17:13:45.513 -> Water_Temp.:105.46, PWM:46
17:13:47.072 -> Water_Temp.:105.46, PWM:45
17:13:47.834 -> Water_Temp.:105.46, PWM:46
17:13:48.630 -> Water_Temp.:105.46, PWM:45
17:13:49.390 -> Water_Temp.:105.46, PWM:45
17:13:51.741 -> Water_Temp.:105.46, PWM:46
17:13:52.502 -> Water_Temp.:105.57, PWM:45
17:13:53.297 -> Water_Temp.:105.57, PWM:45
17:13:55.607 -> Water_Temp.:105.57, PWM:45
17:13:57.134 -> Water_Temp.:105.57, PWM:46
17:13:57.928 -> Water_Temp.:105.57, PWM:45
17:13:59.447 -> Water_Temp.:105.57, PWM:45
17:14:01.037 -> Water_Temp.:105.69, PWM:45
17:14:04.113 -> Water_Temp.:105.69, PWM:45
17:14:05.698 -> Water_Temp.:105.69, PWM:45
17:14:06.457 -> Water_Temp.:105.69, PWM:46
17:14:07.255 -> Water_Temp.:105.69, PWM:45
17:14:08.777 -> Water_Temp.:105.69, PWM:46
17:14:09.574 -> Water_Temp.:105.69, PWM:45
17:14:10.334 -> Water_Temp.:105.69, PWM:45
17:14:11.131 -> Water_Temp.:105.69, PWM:45
17:14:11.890 -> Water_Temp.:105.69, PWM:45
17:14:12.681 -> Water_Temp.:105.57, PWM:45
17:14:14.987 -> Water_Temp.:105.57, PWM:45
17:14:15.748 -> Water_Temp.:105.57, PWM:45
17:14:17.299 -> Water_Temp.:105.57, PWM:46
17:14:18.095 -> Water_Temp.:105.57, PWM:45
17:14:18.888 -> Water_Temp.:105.57, PWM:46
17:14:20.409 -> Water_Temp.:105.57, PWM:45
17:14:21.204 -> Water_Temp.:105.57, PWM:45
17:14:21.995 -> Water_Temp.:105.57, PWM:45
17:14:22.756 -> Water_Temp.:105.57, PWM:46
17:14:24.281 -> Water_Temp.:105.57, PWM:45
17:14:25.078 -> Water_Temp.:105.57, PWM:46
17:14:25.838 -> Water_Temp.:105.57, PWM:46
17:14:26.632 -> Water_Temp.:105.57, PWM:45
17:14:28.944 -> Water_Temp.:105.69, PWM:46
17:14:29.738 -> Water_Temp.:105.69, PWM:45
17:14:30.498 -> Water_Temp.:105.69, PWM:45
17:14:31.293 -> Water_Temp.:105.69, PWM:46
17:14:32.849 -> Water_Temp.:105.69, PWM:46
17:14:33.642 -> Water_Temp.:105.69, PWM:46
17:14:35.191 -> Water_Temp.:105.69, PWM:45
17:14:35.948 -> Water_Temp.:105.69, PWM:45
17:14:40.564 -> Water_Temp.:105.80, PWM:45
17:14:41.321 -> Water_Temp.:105.91, PWM:45
17:14:42.110 -> Water_Temp.:105.91, PWM:45
17:14:43.665 -> Water_Temp.:105.91, PWM:46
17:14:45.221 -> Water_Temp.:105.91, PWM:45
17:14:45.985 -> Water_Temp.:105.91, PWM:45
17:14:46.782 -> Water_Temp.:105.91, PWM:45
17:14:48.344 -> Water_Temp.:106.02, PWM:46
17:14:49.103 -> Water_Temp.:106.02, PWM:46
17:14:51.429 -> Water_Temp.:106.02, PWM:46
17:14:53.705 -> Water_Temp.:106.02, PWM:45
17:14:55.259 -> Water_Temp.:106.14, PWM:45
17:14:56.054 -> Water_Temp.:106.14, PWM:45
17:14:58.362 -> Water_Temp.:106.14, PWM:45
17:14:59.150 -> Water_Temp.:106.14, PWM:45
17:14:59.940 -> Water_Temp.:106.14, PWM:45
17:15:00.699 -> Water_Temp.:106.14, PWM:45
17:15:01.485 -> Water_Temp.:106.14, PWM:36
17:15:02.283 -> Water_Temp.:106.14, PWM:0
17:15:03.042 -> Water_Temp.:106.14, PWM:0
17:15:03.835 -> Water_Temp.:106.14, PWM:0
17:15:04.597 -> Water_Temp.:106.14, PWM:0
17:15:05.389 -> Water_Temp.:106.14, PWM:0
17:15:06.185 -> Water_Temp.:106.14, PWM:0
17:15:06.940 -> Water_Temp.:106.14, PWM:0
17:15:07.735 -> Water_Temp.:106.14, PWM:0
17:15:08.497 -> Water_Temp.:106.14, PWM:0
17:15:09.290 -> Water_Temp.:106.14, PWM:0
17:15:10.088 -> Water_Temp.:106.14, PWM:0
17:15:10.847 -> Water_Temp.:106.14, PWM:0
17:15:11.641 -> Water_Temp.:106.14, PWM:0
17:15:12.400 -> Water_Temp.:106.14, PWM:0
17:15:13.197 -> Water_Temp.:106.25, PWM:0
17:15:13.988 -> Water_Temp.:106.25, PWM:0
17:15:14.746 -> Water_Temp.:106.25, PWM:0
17:15:15.541 -> Water_Temp.:106.25, PWM:0
17:15:16.302 -> Water_Temp.:106.25, PWM:0
17:15:17.096 -> Water_Temp.:106.25, PWM:0
17:15:17.888 -> Water_Temp.:106.25, PWM:0
17:15:18.647 -> Water_Temp.:106.25, PWM:0
17:15:19.436 -> Water_Temp.:106.25, PWM:0
17:15:20.225 -> Water_Temp.:106.25, PWM:0
17:15:20.981 -> Water_Temp.:106.25, PWM:0
17:15:21.771 -> Water_Temp.:106.25, PWM:0
17:15:22.531 -> Water_Temp.:106.25, PWM:0
17:15:23.321 -> Water_Temp.:106.25, PWM:0
17:15:24.117 -> Water_Temp.:106.25, PWM:0
17:15:24.881 -> Water_Temp.:106.25, PWM:0
17:15:25.672 -> Water_Temp.:106.36, PWM:0
17:15:26.432 -> Water_Temp.:106.36, PWM:0
17:15:27.220 -> Water_Temp.:106.36, PWM:0
17:15:28.013 -> Water_Temp.:106.36, PWM:0
17:15:28.800 -> Water_Temp.:106.25, PWM:0
17:15:29.556 -> Water_Temp.:106.25, PWM:0
17:15:30.347 -> Water_Temp.:106.36, PWM:0
17:15:31.138 -> Water_Temp.:106.25, PWM:0
17:15:31.898 -> Water_Temp.:106.25, PWM:0
17:15:32.689 -> Water_Temp.:106.25, PWM:0
17:15:33.479 -> Water_Temp.:106.36, PWM:0
17:15:34.239 -> Water_Temp.:106.25, PWM:0
17:15:35.034 -> Water_Temp.:106.36, PWM:0
17:15:35.792 -> Water_Temp.:106.36, PWM:0
17:15:36.587 -> Water_Temp.:106.25, PWM:0
17:15:37.378 -> Water_Temp.:106.36, PWM:0
17:15:38.139 -> Water_Temp.:106.25, PWM:0
17:15:38.933 -> Water_Temp.:106.25, PWM:0
17:15:39.694 -> Water_Temp.:106.25, PWM:0
17:15:40.487 -> Water_Temp.:106.25, PWM:0
17:15:41.282 -> Water_Temp.:106.25, PWM:0
17:15:42.038 -> Water_Temp.:106.25, PWM:0
17:15:42.830 -> Water_Temp.:106.25, PWM:0
17:15:43.621 -> Water_Temp.:106.25, PWM:0
17:15:44.377 -> Water_Temp.:106.25, PWM:0
17:15:45.171 -> Water_Temp.:106.25, PWM:0
17:15:45.964 -> Water_Temp.:106.25, PWM:0
17:15:46.718 -> Water_Temp.:106.25, PWM:0
17:15:47.512 -> Water_Temp.:106.25, PWM:0
17:15:48.271 -> Water_Temp.:106.25, PWM:0
17:15:49.066 -> Water_Temp.:106.25, PWM:0
17:15:49.853 -> Water_Temp.:106.25, PWM:0
17:15:50.643 -> Water_Temp.:106.25, PWM:0
17:15:51.400 -> Water_Temp.:106.25, PWM:0
17:15:52.191 -> Water_Temp.:106.25, PWM:0
17:15:52.978 -> Water_Temp.:106.25, PWM:0

## RAW_CAL_MLX-2
14:19:00.703 -> Average Cal sensor Temp.: 90.81 Ambient Temp.: 83.97 F
14:19:02.249 -> Average Cal sensor Temp.: 90.80 Ambient Temp.: 83.93 F
14:19:03.799 -> Average Cal sensor Temp.: 90.91 Ambient Temp.: 83.93 F
14:19:05.383 -> Average Cal sensor Temp.: 90.95 Ambient Temp.: 83.97 F
14:19:06.929 -> Average Cal sensor Temp.: 91.15 Ambient Temp.: 83.93 F
14:19:08.482 -> Average Cal sensor Temp.: 91.26 Ambient Temp.: 83.97 F
14:19:10.037 -> Average Cal sensor Temp.: 91.47 Ambient Temp.: 83.93 F
14:19:11.588 -> Average Cal sensor Temp.: 90.95 Ambient Temp.: 83.93 F
14:19:13.140 -> Average Cal sensor Temp.: 90.85 Ambient Temp.: 83.89 F
14:19:14.694 -> Average Cal sensor Temp.: 90.70 Ambient Temp.: 83.93 F
14:19:16.246 -> Average Cal sensor Temp.: 90.54 Ambient Temp.: 83.89 F
14:19:17.830 -> Average Cal sensor Temp.: 90.38 Ambient Temp.: 83.86 F
14:19:19.383 -> Average Cal sensor Temp.: 90.18 Ambient Temp.: 83.89 F
14:19:20.933 -> Average Cal sensor Temp.: 89.90 Ambient Temp.: 83.82 F
14:19:22.484 -> Average Cal sensor Temp.: 90.05 Ambient Temp.: 83.82 F
14:19:24.036 -> Average Cal sensor Temp.: 90.07 Ambient Temp.: 83.82 F
14:19:25.616 -> Average Cal sensor Temp.: 90.18 Ambient Temp.: 83.82 F
14:19:27.164 -> Average Cal sensor Temp.: 90.36 Ambient Temp.: 83.86 F
14:19:28.714 -> Average Cal sensor Temp.: 90.18 Ambient Temp.: 83.79 F
14:19:30.270 -> Average Cal sensor Temp.: 89.89 Ambient Temp.: 83.71 F
14:19:31.820 -> Average Cal sensor Temp.: 89.87 Ambient Temp.: 83.68 F
14:19:33.406 -> Average Cal sensor Temp.: 89.82 Ambient Temp.: 83.71 F
14:19:34.956 -> Average Cal sensor Temp.: 89.90 Ambient Temp.: 83.71 F
14:19:36.500 -> Average Cal sensor Temp.: 90.05 Ambient Temp.: 83.71 F
14:19:38.050 -> Average Cal sensor Temp.: 89.80 Ambient Temp.: 83.71 F
14:19:39.604 -> Average Cal sensor Temp.: 89.85 Ambient Temp.: 83.64 F
14:19:41.189 -> Average Cal sensor Temp.: 90.25 Ambient Temp.: 83.57 F
14:19:42.742 -> Average Cal sensor Temp.: 90.05 Ambient Temp.: 83.57 F
14:19:44.285 -> Average Cal sensor Temp.: 90.09 Ambient Temp.: 83.57 F
14:19:45.836 -> Average Cal sensor Temp.: 90.07 Ambient Temp.: 83.64 F
14:19:47.393 -> Average Cal sensor Temp.: 89.72 Ambient Temp.: 83.57 F
14:19:48.952 -> Average Cal sensor Temp.: 89.83 Ambient Temp.: 83.53 F
14:19:50.502 -> Average Cal sensor Temp.: 89.52 Ambient Temp.: 83.50 F
14:19:52.052 -> Average Cal sensor Temp.: 89.71 Ambient Temp.: 83.46 F
14:19:53.605 -> Average Cal sensor Temp.: 89.49 Ambient Temp.: 83.57 F
14:19:55.190 -> Average Cal sensor Temp.: 89.47 Ambient Temp.: 83.50 F
14:19:56.739 -> Average Cal sensor Temp.: 89.46 Ambient Temp.: 83.46 F
14:19:58.287 -> Average Cal sensor Temp.: 89.75 Ambient Temp.: 83.43 F
14:19:59.836 -> Average Cal sensor Temp.: 89.77 Ambient Temp.: 83.43 F
14:20:01.420 -> Average Cal sensor Temp.: 89.28 Ambient Temp.: 83.43 F
14:20:02.972 -> Average Cal sensor Temp.: 89.76 Ambient Temp.: 83.43 F
14:20:04.527 -> Average Cal sensor Temp.: 90.05 Ambient Temp.: 83.46 F
14:20:06.085 -> Average Cal sensor Temp.: 89.88 Ambient Temp.: 83.43 F
14:20:07.640 -> Average Cal sensor Temp.: 89.94 Ambient Temp.: 83.35 F
14:20:09.197 -> Average Cal sensor Temp.: 89.93 Ambient Temp.: 83.39 F
14:20:10.747 -> Average Cal sensor Temp.: 89.77 Ambient Temp.: 83.43 F
14:20:12.297 -> Average Cal sensor Temp.: 89.67 Ambient Temp.: 83.32 F
14:20:13.851 -> Average Cal sensor Temp.: 89.89 Ambient Temp.: 83.35 F
14:20:15.403 -> Average Cal sensor Temp.: 89.87 Ambient Temp.: 83.35 F
14:20:16.987 -> Average Cal sensor Temp.: 89.78 Ambient Temp.: 83.32 F
14:20:18.541 -> Average Cal sensor Temp.: 89.86 Ambient Temp.: 83.32 F
14:20:19.129 -> Ambient Temp.: 83.32 F
14:20:19.681 -> Ambient Temp.: 83.32 F
14:20:20.273 -> Ambient Temp.: 83.32 F
14:20:21.825 -> Average Cal sensor Temp.: 89.98 Ambient Temp.: 83.32 F
14:20:23.381 -> Average Cal sensor Temp.: 89.98 Ambient Temp.: 83.25 F
14:20:24.691 -> Ambient Temp.: 83.28 F
14:20:25.277 -> Ambient Temp.: 83.28 F
14:20:25.862 -> Ambient Temp.: 83.28 F
14:20:26.452 -> Ambient Temp.: 83.28 F
14:20:27.007 -> Ambient Temp.: 83.28 F
14:20:27.591 -> Ambient Temp.: 83.28 F
14:20:28.179 -> Ambient Temp.: 83.28 F
14:20:28.904 -> Ambient Temp.: 83.28 F
14:20:29.458 -> Ambient Temp.: 83.28 F
14:20:30.188 -> Ambient Temp.: 83.17 F
14:20:30.736 -> Ambient Temp.: 83.17 F
14:20:31.456 -> Ambient Temp.: 83.25 F
14:20:33.010 -> Average Cal sensor Temp.: 89.79 Ambient Temp.: 83.25 F
14:20:34.563 -> Average Cal sensor Temp.: 89.96 Ambient Temp.: 83.25 F
14:20:36.120 -> Average Cal sensor Temp.: 89.77 Ambient Temp.: 83.14 F
14:20:37.672 -> Average Cal sensor Temp.: 89.76 Ambient Temp.: 83.14 F
14:20:39.225 -> Average Cal sensor Temp.: 89.22 Ambient Temp.: 83.17 F
14:20:40.781 -> Average Cal sensor Temp.: 89.15 Ambient Temp.: 83.10 F
14:20:42.330 -> Average Cal sensor Temp.: 89.30 Ambient Temp.: 83.03 F
14:20:43.883 -> Average Cal sensor Temp.: 89.27 Ambient Temp.: 83.07 F
14:20:45.464 -> Average Cal sensor Temp.: 88.82 Ambient Temp.: 83.03 F
14:20:47.017 -> Average Cal sensor Temp.: 89.08 Ambient Temp.: 82.96 F
14:20:48.571 -> Average Cal sensor Temp.: 89.53 Ambient Temp.: 82.96 F
14:20:50.122 -> Average Cal sensor Temp.: 89.46 Ambient Temp.: 82.92 F
14:20:51.669 -> Average Cal sensor Temp.: 88.89 Ambient Temp.: 82.92 F
14:20:53.250 -> Average Cal sensor Temp.: 88.96 Ambient Temp.: 82.92 F
14:20:54.807 -> Average Cal sensor Temp.: 89.39 Ambient Temp.: 82.85 F
14:20:56.358 -> Average Cal sensor Temp.: 89.60 Ambient Temp.: 82.89 F
14:20:57.908 -> Average Cal sensor Temp.: 89.17 Ambient Temp.: 82.85 F
14:20:59.456 -> Average Cal sensor Temp.: 89.06 Ambient Temp.: 82.71 F
14:21:01.011 -> Average Cal sensor Temp.: 89.48 Ambient Temp.: 82.71 F
14:21:02.568 -> Average Cal sensor Temp.: 89.35 Ambient Temp.: 82.71 F
14:21:04.153 -> Average Cal sensor Temp.: 89.41 Ambient Temp.: 82.71 F
14:21:05.706 -> Average Cal sensor Temp.: 89.36 Ambient Temp.: 82.67 F
14:21:07.261 -> Average Cal sensor Temp.: 89.38 Ambient Temp.: 82.63 F
14:21:08.814 -> Average Cal sensor Temp.: 89.17 Ambient Temp.: 82.60 F
14:21:10.364 -> Average Cal sensor Temp.: 89.57 Ambient Temp.: 82.60 F
14:21:11.911 -> Average Cal sensor Temp.: 89.83 Ambient Temp.: 82.60 F
14:21:13.465 -> Average Cal sensor Temp.: 89.61 Ambient Temp.: 82.60 F
14:21:15.016 -> Average Cal sensor Temp.: 89.82 Ambient Temp.: 82.56 F
14:21:16.603 -> Average Cal sensor Temp.: 89.82 Ambient Temp.: 82.56 F
14:21:18.150 -> Average Cal sensor Temp.: 89.68 Ambient Temp.: 82.49 F
14:21:19.701 -> Average Cal sensor Temp.: 89.43 Ambient Temp.: 82.49 F
14:21:21.255 -> Average Cal sensor Temp.: 89.71 Ambient Temp.: 82.53 F
14:21:22.808 -> Average Cal sensor Temp.: 89.41 Ambient Temp.: 82.49 F
14:21:24.362 -> Average Cal sensor Temp.: 89.28 Ambient Temp.: 82.53 F
14:21:25.920 -> Average Cal sensor Temp.: 89.09 Ambient Temp.: 82.49 F
14:21:27.473 -> Average Cal sensor Temp.: 89.24 Ambient Temp.: 82.45 F
14:21:29.062 -> Average Cal sensor Temp.: 89.41 Ambient Temp.: 82.45 F
14:21:30.588 -> Average Cal sensor Temp.: 89.27 Ambient Temp.: 82.38 F
14:21:32.142 -> Average Cal sensor Temp.: 89.55 Ambient Temp.: 82.38 F
14:21:33.727 -> Average Cal sensor Temp.: 89.59 Ambient Temp.: 82.45 F
14:21:35.278 -> Average Cal sensor Temp.: 89.27 Ambient Temp.: 82.45 F
14:21:36.827 -> Average Cal sensor Temp.: 89.37 Ambient Temp.: 82.31 F
14:21:38.381 -> Average Cal sensor Temp.: 89.58 Ambient Temp.: 82.35 F
14:21:39.939 -> Average Cal sensor Temp.: 89.62 Ambient Temp.: 82.38 F
14:21:41.492 -> Average Cal sensor Temp.: 89.64 Ambient Temp.: 82.35 F
14:21:43.046 -> Average Cal sensor Temp.: 89.84 Ambient Temp.: 82.35 F
14:21:44.603 -> Average Cal sensor Temp.: 89.80 Ambient Temp.: 82.38 F
14:21:46.157 -> Average Cal sensor Temp.: 89.77 Ambient Temp.: 82.38 F
14:21:47.709 -> Average Cal sensor Temp.: 89.51 Ambient Temp.: 82.38 F
14:21:49.263 -> Average Cal sensor Temp.: 89.59 Ambient Temp.: 82.45 F
14:21:50.849 -> Average Cal sensor Temp.: 89.54 Ambient Temp.: 82.38 F
14:21:52.402 -> Average Cal sensor Temp.: 89.79 Ambient Temp.: 82.35 F
14:21:53.964 -> Average Cal sensor Temp.: 89.98 Ambient Temp.: 82.45 F
14:21:55.506 -> Average Cal sensor Temp.: 89.81 Ambient Temp.: 82.35 F
14:21:57.062 -> Average Cal sensor Temp.: 89.95 Ambient Temp.: 82.45 F
14:21:58.619 -> Average Cal sensor Temp.: 89.77 Ambient Temp.: 82.38 F
14:22:00.174 -> Average Cal sensor Temp.: 90.14 Ambient Temp.: 82.38 F
14:22:01.722 -> Average Cal sensor Temp.: 90.17 Ambient Temp.: 82.53 F
14:22:03.274 -> Average Cal sensor Temp.: 90.19 Ambient Temp.: 82.49 F
14:22:04.858 -> Average Cal sensor Temp.: 89.73 Ambient Temp.: 82.56 F
14:22:06.410 -> Average Cal sensor Temp.: 89.52 Ambient Temp.: 82.49 F
14:22:07.963 -> Average Cal sensor Temp.: 89.51 Ambient Temp.: 82.45 F
14:22:09.515 -> Average Cal sensor Temp.: 89.55 Ambient Temp.: 82.53 F
14:22:11.068 -> Average Cal sensor Temp.: 89.49 Ambient Temp.: 82.49 F
14:22:12.621 -> Average Cal sensor Temp.: 89.73 Ambient Temp.: 82.53 F
14:22:14.174 -> Average Cal sensor Temp.: 89.74 Ambient Temp.: 82.56 F
14:22:15.723 -> Average Cal sensor Temp.: 90.02 Ambient Temp.: 82.60 F
14:22:17.302 -> Average Cal sensor Temp.: 89.87 Ambient Temp.: 82.60 F
14:22:18.850 -> Average Cal sensor Temp.: 89.86 Ambient Temp.: 82.56 F
14:22:20.405 -> Average Cal sensor Temp.: 89.96 Ambient Temp.: 82.60 F
14:22:21.957 -> Average Cal sensor Temp.: 89.89 Ambient Temp.: 82.63 F
14:22:23.541 -> Average Cal sensor Temp.: 89.44 Ambient Temp.: 82.60 F
14:22:25.094 -> Average Cal sensor Temp.: 89.43 Ambient Temp.: 82.60 F
14:22:26.649 -> Average Cal sensor Temp.: 89.34 Ambient Temp.: 82.56 F
14:22:28.201 -> Average Cal sensor Temp.: 89.46 Ambient Temp.: 82.63 F
14:22:29.749 -> Average Cal sensor Temp.: 89.48 Ambient Temp.: 82.60 F
14:22:31.304 -> Average Cal sensor Temp.: 89.76 Ambient Temp.: 82.63 F
14:22:32.856 -> Average Cal sensor Temp.: 89.46 Ambient Temp.: 82.63 F
14:22:34.443 -> Average Cal sensor Temp.: 89.68 Ambient Temp.: 82.67 F
14:22:35.997 -> Average Cal sensor Temp.: 89.36 Ambient Temp.: 82.67 F
14:22:37.553 -> Average Cal sensor Temp.: 89.48 Ambient Temp.: 82.67 F
14:22:39.103 -> Average Cal sensor Temp.: 89.66 Ambient Temp.: 82.63 F
14:22:40.663 -> Average Cal sensor Temp.: 89.79 Ambient Temp.: 82.67 F
14:22:42.217 -> Average Cal sensor Temp.: 89.87 Ambient Temp.: 82.71 F
14:22:43.768 -> Average Cal sensor Temp.: 89.97 Ambient Temp.: 82.71 F
14:22:45.320 -> Average Cal sensor Temp.: 90.00 Ambient Temp.: 82.71 F
14:22:46.879 -> Average Cal sensor Temp.: 89.76 Ambient Temp.: 82.74 F
14:22:48.430 -> Average Cal sensor Temp.: 89.69 Ambient Temp.: 82.74 F
14:22:50.013 -> Average Cal sensor Temp.: 89.78 Ambient Temp.: 82.78 F
14:22:51.564 -> Average Cal sensor Temp.: 89.69 Ambient Temp.: 82.74 F
14:22:53.117 -> Average Cal sensor Temp.: 89.66 Ambient Temp.: 82.85 F
14:22:54.660 -> Average Cal sensor Temp.: 89.61 Ambient Temp.: 82.85 F
14:22:56.244 -> Average Cal sensor Temp.: 89.66 Ambient Temp.: 82.78 F
14:22:57.797 -> Average Cal sensor Temp.: 89.82 Ambient Temp.: 82.89 F
14:22:59.348 -> Average Cal sensor Temp.: 90.05 Ambient Temp.: 82.78 F
14:23:00.900 -> Average Cal sensor Temp.: 89.91 Ambient Temp.: 82.89 F
14:23:02.447 -> Average Cal sensor Temp.: 89.94 Ambient Temp.: 82.92 F
14:23:04.025 -> Average Cal sensor Temp.: 89.90 Ambient Temp.: 82.89 F
14:23:05.577 -> Average Cal sensor Temp.: 89.81 Ambient Temp.: 82.92 F
14:23:07.133 -> Average Cal sensor Temp.: 89.70 Ambient Temp.: 82.92 F
14:23:08.687 -> Average Cal sensor Temp.: 89.66 Ambient Temp.: 82.96 F
14:23:10.235 -> Average Cal sensor Temp.: 89.83 Ambient Temp.: 82.96 F
14:23:11.790 -> Average Cal sensor Temp.: 90.08 Ambient Temp.: 82.96 F
14:23:13.348 -> Average Cal sensor Temp.: 89.73 Ambient Temp.: 82.99 F
14:23:14.934 -> Average Cal sensor Temp.: 89.85 Ambient Temp.: 82.99 F
14:23:16.490 -> Average Cal sensor Temp.: 89.83 Ambient Temp.: 82.96 F
14:23:18.044 -> Average Cal sensor Temp.: 89.87 Ambient Temp.: 83.03 F
14:23:19.601 -> Average Cal sensor Temp.: 89.99 Ambient Temp.: 82.99 F
14:23:21.154 -> Average Cal sensor Temp.: 90.16 Ambient Temp.: 83.10 F
14:23:22.715 -> Average Cal sensor Temp.: 89.95 Ambient Temp.: 83.03 F
14:23:24.266 -> Average Cal sensor Temp.: 90.13 Ambient Temp.: 83.10 F
14:23:25.817 -> Average Cal sensor Temp.: 90.10 Ambient Temp.: 83.10 F
14:23:27.371 -> Average Cal sensor Temp.: 89.87 Ambient Temp.: 83.10 F
14:23:28.924 -> Average Cal sensor Temp.: 89.98 Ambient Temp.: 83.14 F
14:23:30.502 -> Average Cal sensor Temp.: 89.90 Ambient Temp.: 83.14 F
14:23:32.055 -> Average Cal sensor Temp.: 89.80 Ambient Temp.: 83.14 F
14:23:33.605 -> Average Cal sensor Temp.: 90.16 Ambient Temp.: 83.25 F
14:23:35.159 -> Average Cal sensor Temp.: 90.00 Ambient Temp.: 83.28 F
14:23:36.707 -> Average Cal sensor Temp.: 89.97 Ambient Temp.: 83.25 F
14:23:38.265 -> Average Cal sensor Temp.: 90.08 Ambient Temp.: 83.28 F
14:23:39.817 -> Average Cal sensor Temp.: 90.03 Ambient Temp.: 83.32 F
14:23:41.372 -> Average Cal sensor Temp.: 90.12 Ambient Temp.: 83.32 F
14:23:42.960 -> Average Cal sensor Temp.: 90.09 Ambient Temp.: 83.32 F
14:23:44.515 -> Average Cal sensor Temp.: 90.20 Ambient Temp.: 83.39 F
14:23:46.066 -> Average Cal sensor Temp.: 90.16 Ambient Temp.: 83.39 F
14:23:47.622 -> Average Cal sensor Temp.: 90.26 Ambient Temp.: 83.43 F
14:23:49.177 -> Average Cal sensor Temp.: 90.23 Ambient Temp.: 83.46 F
14:23:50.733 -> Average Cal sensor Temp.: 90.06 Ambient Temp.: 83.46 F
14:23:52.289 -> Average Cal sensor Temp.: 89.96 Ambient Temp.: 83.53 F
14:23:53.838 -> Average Cal sensor Temp.: 90.21 Ambient Temp.: 83.57 F
14:23:55.392 -> Average Cal sensor Temp.: 90.30 Ambient Temp.: 83.64 F
14:23:56.949 -> Average Cal sensor Temp.: 90.26 Ambient Temp.: 83.57 F
14:23:58.497 -> Average Cal sensor Temp.: 90.19 Ambient Temp.: 83.64 F
14:24:00.050 -> Average Cal sensor Temp.: 90.23 Ambient Temp.: 83.64 F
14:24:01.607 -> Average Cal sensor Temp.: 90.31 Ambient Temp.: 83.71 F
14:24:03.161 -> Average Cal sensor Temp.: 90.51 Ambient Temp.: 83.71 F
14:24:04.741 -> Average Cal sensor Temp.: 90.60 Ambient Temp.: 83.75 F
14:24:06.298 -> Average Cal sensor Temp.: 90.67 Ambient Temp.: 83.79 F
14:24:07.852 -> Average Cal sensor Temp.: 90.54 Ambient Temp.: 83.82 F
14:24:09.401 -> Average Cal sensor Temp.: 90.53 Ambient Temp.: 83.89 F
14:24:10.950 -> Average Cal sensor Temp.: 90.57 Ambient Temp.: 83.89 F
14:24:12.499 -> Average Cal sensor Temp.: 90.65 Ambient Temp.: 83.93 F
14:24:14.085 -> Average Cal sensor Temp.: 90.58 Ambient Temp.: 83.93 F
14:24:15.641 -> Average Cal sensor Temp.: 90.50 Ambient Temp.: 83.97 F
14:24:17.201 -> Average Cal sensor Temp.: 90.37 Ambient Temp.: 83.97 F
14:24:18.754 -> Average Cal sensor Temp.: 90.37 Ambient Temp.: 84.00 F
14:24:20.308 -> Average Cal sensor Temp.: 90.37 Ambient Temp.: 84.07 F
14:24:21.860 -> Average Cal sensor Temp.: 90.39 Ambient Temp.: 84.11 F
14:24:23.411 -> Average Cal sensor Temp.: 90.50 Ambient Temp.: 84.15 F
14:24:24.964 -> Average Cal sensor Temp.: 90.19 Ambient Temp.: 84.18 F
14:24:26.523 -> Average Cal sensor Temp.: 90.26 Ambient Temp.: 84.22 F
14:24:28.076 -> Average Cal sensor Temp.: 90.32 Ambient Temp.: 84.18 F
14:24:29.630 -> Average Cal sensor Temp.: 90.07 Ambient Temp.: 84.18 F
14:24:31.182 -> Average Cal sensor Temp.: 89.99 Ambient Temp.: 84.22 F
14:24:32.735 -> Average Cal sensor Temp.: 90.05 Ambient Temp.: 84.22 F
14:24:34.321 -> Average Cal sensor Temp.: 89.96 Ambient Temp.: 84.25 F
14:24:35.870 -> Average Cal sensor Temp.: 89.61 Ambient Temp.: 84.25 F
14:24:37.420 -> Average Cal sensor Temp.: 89.63 Ambient Temp.: 84.25 F
14:24:38.971 -> Average Cal sensor Temp.: 90.15 Ambient Temp.: 84.33 F
14:24:40.523 -> Average Cal sensor Temp.: 90.03 Ambient Temp.: 84.29 F
14:24:42.104 -> Average Cal sensor Temp.: 89.95 Ambient Temp.: 84.25 F
14:24:43.658 -> Average Cal sensor Temp.: 89.75 Ambient Temp.: 84.25 F
14:24:45.209 -> Average Cal sensor Temp.: 89.65 Ambient Temp.: 84.29 F
14:24:46.765 -> Average Cal sensor Temp.: 90.02 Ambient Temp.: 84.33 F
14:24:48.313 -> Average Cal sensor Temp.: 89.75 Ambient Temp.: 84.29 F
14:24:49.863 -> Average Cal sensor Temp.: 90.12 Ambient Temp.: 84.33 F
14:24:51.414 -> Average Cal sensor Temp.: 90.12 Ambient Temp.: 84.29 F
14:24:52.972 -> Average Cal sensor Temp.: 89.82 Ambient Temp.: 84.29 F
14:24:54.554 -> Average Cal sensor Temp.: 90.14 Ambient Temp.: 84.29 F
14:24:56.105 -> Average Cal sensor Temp.: 90.19 Ambient Temp.: 84.33 F
14:24:57.658 -> Average Cal sensor Temp.: 89.90 Ambient Temp.: 84.25 F
14:24:59.211 -> Average Cal sensor Temp.: 90.18 Ambient Temp.: 84.29 F
14:25:00.755 -> Average Cal sensor Temp.: 90.39 Ambient Temp.: 84.29 F
14:25:02.341 -> Average Cal sensor Temp.: 90.37 Ambient Temp.: 84.29 F
14:25:03.898 -> Average Cal sensor Temp.: 90.18 Ambient Temp.: 84.33 F
14:25:05.448 -> Average Cal sensor Temp.: 90.38 Ambient Temp.: 84.29 F
14:25:06.996 -> Average Cal sensor Temp.: 90.08 Ambient Temp.: 84.33 F
14:25:08.552 -> Average Cal sensor Temp.: 90.40 Ambient Temp.: 84.29 F
14:25:10.107 -> Average Cal sensor Temp.: 90.52 Ambient Temp.: 84.29 F
14:25:11.661 -> Average Cal sensor Temp.: 90.48 Ambient Temp.: 84.25 F
14:25:13.211 -> Average Cal sensor Temp.: 90.66 Ambient Temp.: 84.29 F
14:25:14.794 -> Average Cal sensor Temp.: 90.74 Ambient Temp.: 84.33 F
14:25:16.352 -> Average Cal sensor Temp.: 90.32 Ambient Temp.: 84.25 F
14:25:17.906 -> Average Cal sensor Temp.: 90.44 Ambient Temp.: 84.29 F
14:25:19.460 -> Average Cal sensor Temp.: 90.25 Ambient Temp.: 84.25 F
14:25:21.016 -> Average Cal sensor Temp.: 90.35 Ambient Temp.: 84.29 F
14:25:22.572 -> Average Cal sensor Temp.: 90.34 Ambient Temp.: 84.29 F
14:25:24.121 -> Average Cal sensor Temp.: 90.44 Ambient Temp.: 84.29 F
14:25:25.675 -> Average Cal sensor Temp.: 90.43 Ambient Temp.: 84.33 F
14:25:27.230 -> Average Cal sensor Temp.: 90.56 Ambient Temp.: 84.29 F
14:25:28.783 -> Average Cal sensor Temp.: 90.55 Ambient Temp.: 84.33 F
14:25:30.367 -> Average Cal sensor Temp.: 90.08 Ambient Temp.: 84.33 F
14:25:31.918 -> Average Cal sensor Temp.: 89.85 Ambient Temp.: 84.33 F
14:25:33.466 -> Average Cal sensor Temp.: 90.08 Ambient Temp.: 84.33 F
14:25:35.009 -> Average Cal sensor Temp.: 90.52 Ambient Temp.: 84.25 F
14:25:36.595 -> Average Cal sensor Temp.: 90.20 Ambient Temp.: 84.29 F
14:25:38.147 -> Average Cal sensor Temp.: 90.04 Ambient Temp.: 84.25 F
14:25:39.699 -> Average Cal sensor Temp.: 89.89 Ambient Temp.: 84.25 F
14:25:41.253 -> Average Cal sensor Temp.: 90.45 Ambient Temp.: 84.22 F
14:25:42.803 -> Average Cal sensor Temp.: 90.47 Ambient Temp.: 84.25 F
14:25:44.361 -> Average Cal sensor Temp.: 90.47 Ambient Temp.: 84.22 F
14:25:45.913 -> Average Cal sensor Temp.: 90.32 Ambient Temp.: 84.22 F
14:25:47.465 -> Average Cal sensor Temp.: 90.19 Ambient Temp.: 84.15 F
14:25:49.050 -> Average Cal sensor Temp.: 90.30 Ambient Temp.: 84.15 F
14:25:50.602 -> Average Cal sensor Temp.: 90.31 Ambient Temp.: 84.18 F
14:25:52.156 -> Average Cal sensor Temp.: 90.54 Ambient Temp.: 84.18 F
14:25:53.703 -> Average Cal sensor Temp.: 90.41 Ambient Temp.: 84.15 F
14:25:55.257 -> Average Cal sensor Temp.: 90.63 Ambient Temp.: 84.11 F
14:25:56.808 -> Average Cal sensor Temp.: 90.41 Ambient Temp.: 84.15 F
14:25:58.363 -> Average Cal sensor Temp.: 90.74 Ambient Temp.: 84.18 F
14:25:59.916 -> Average Cal sensor Temp.: 90.45 Ambient Temp.: 84.11 F
14:26:01.467 -> Average Cal sensor Temp.: 90.38 Ambient Temp.: 84.11 F
14:26:03.052 -> Average Cal sensor Temp.: 90.63 Ambient Temp.: 84.15 F
14:26:04.599 -> Average Cal sensor Temp.: 90.40 Ambient Temp.: 84.15 F
14:26:06.149 -> Average Cal sensor Temp.: 90.48 Ambient Temp.: 84.07 F
14:26:07.699 -> Average Cal sensor Temp.: 90.35 Ambient Temp.: 84.00 F
14:26:09.251 -> Average Cal sensor Temp.: 90.32 Ambient Temp.: 84.00 F
14:26:10.804 -> Average Cal sensor Temp.: 90.50 Ambient Temp.: 84.07 F
14:26:12.391 -> Average Cal sensor Temp.: 90.41 Ambient Temp.: 84.07 F
14:26:13.944 -> Average Cal sensor Temp.: 90.66 Ambient Temp.: 84.00 F
14:26:15.502 -> Average Cal sensor Temp.: 90.34 Ambient Temp.: 84.07 F
14:26:17.054 -> Average Cal sensor Temp.: 90.41 Ambient Temp.: 83.97 F
14:26:18.612 -> Average Cal sensor Temp.: 90.41 Ambient Temp.: 84.00 F
14:26:20.169 -> Average Cal sensor Temp.: 90.37 Ambient Temp.: 83.97 F
14:26:21.725 -> Average Cal sensor Temp.: 90.35 Ambient Temp.: 83.89 F
14:26:23.279 -> Average Cal sensor Temp.: 90.25 Ambient Temp.: 83.97 F
14:26:24.839 -> Average Cal sensor Temp.: 90.02 Ambient Temp.: 83.93 F
14:26:26.393 -> Average Cal sensor Temp.: 89.97 Ambient Temp.: 83.89 F
14:26:27.946 -> Average Cal sensor Temp.: 90.20 Ambient Temp.: 83.89 F
14:26:29.501 -> Average Cal sensor Temp.: 90.05 Ambient Temp.: 83.86 F
14:26:31.047 -> Average Cal sensor Temp.: 90.45 Ambient Temp.: 83.82 F
14:26:32.603 -> Average Cal sensor Temp.: 90.46 Ambient Temp.: 83.79 F
14:26:34.186 -> Average Cal sensor Temp.: 90.59 Ambient Temp.: 83.86 F
14:26:35.737 -> Average Cal sensor Temp.: 90.17 Ambient Temp.: 83.79 F
14:26:37.285 -> Average Cal sensor Temp.: 90.56 Ambient Temp.: 83.75 F
14:26:38.836 -> Average Cal sensor Temp.: 90.54 Ambient Temp.: 83.71 F
14:26:40.390 -> Average Cal sensor Temp.: 90.54 Ambient Temp.: 83.71 F
14:26:41.938 -> Average Cal sensor Temp.: 90.47 Ambient Temp.: 83.68 F
14:26:43.522 -> Average Cal sensor Temp.: 90.11 Ambient Temp.: 83.68 F
14:26:45.072 -> Average Cal sensor Temp.: 90.19 Ambient Temp.: 83.64 F
14:26:46.628 -> Average Cal sensor Temp.: 90.43 Ambient Temp.: 83.64 F
14:26:48.176 -> Average Cal sensor Temp.: 90.60 Ambient Temp.: 83.64 F
14:26:49.728 -> Average Cal sensor Temp.: 90.84 Ambient Temp.: 83.57 F
14:26:51.281 -> Average Cal sensor Temp.: 91.12 Ambient Temp.: 83.53 F
14:26:52.832 -> Average Cal sensor Temp.: 90.70 Ambient Temp.: 83.53 F
14:26:54.417 -> Average Cal sensor Temp.: 90.76 Ambient Temp.: 83.53 F
14:26:55.961 -> Average Cal sensor Temp.: 90.82 Ambient Temp.: 83.53 F
14:26:57.507 -> Average Cal sensor Temp.: 90.17 Ambient Temp.: 83.50 F
14:26:59.060 -> Average Cal sensor Temp.: 90.25 Ambient Temp.: 83.46 F
14:27:00.612 -> Average Cal sensor Temp.: 90.58 Ambient Temp.: 83.53 F
14:27:02.200 -> Average Cal sensor Temp.: 90.54 Ambient Temp.: 83.46 F
14:27:03.753 -> Average Cal sensor Temp.: 90.53 Ambient Temp.: 83.46 F
14:27:05.308 -> Average Cal sensor Temp.: 90.37 Ambient Temp.: 83.46 F
14:27:06.859 -> Average Cal sensor Temp.: 90.64 Ambient Temp.: 83.39 F
14:27:08.408 -> Average Cal sensor Temp.: 90.92 Ambient Temp.: 83.43 F
14:27:09.955 -> Average Cal sensor Temp.: 90.34 Ambient Temp.: 83.39 F
14:27:11.507 -> Average Cal sensor Temp.: 90.47 Ambient Temp.: 83.39 F
14:27:13.096 -> Average Cal sensor Temp.: 90.72 Ambient Temp.: 83.39 F
14:27:14.649 -> Average Cal sensor Temp.: 90.40 Ambient Temp.: 83.39 F
14:27:16.206 -> Average Cal sensor Temp.: 90.75 Ambient Temp.: 83.32 F
14:27:17.731 -> Average Cal sensor Temp.: 90.83 Ambient Temp.: 83.32 F
14:27:19.321 -> Average Cal sensor Temp.: 90.57 Ambient Temp.: 83.25 F
14:27:20.845 -> Average Cal sensor Temp.: 90.45 Ambient Temp.: 83.32 F
14:27:22.427 -> Average Cal sensor Temp.: 90.19 Ambient Temp.: 83.25 F
14:27:23.979 -> Average Cal sensor Temp.: 89.95 Ambient Temp.: 83.17 F
14:27:25.536 -> Average Cal sensor Temp.: 90.28 Ambient Temp.: 83.10 F
14:27:27.089 -> Average Cal sensor Temp.: 90.54 Ambient Temp.: 83.14 F
14:27:28.644 -> Average Cal sensor Temp.: 90.56 Ambient Temp.: 83.10 F
14:27:30.197 -> Average Cal sensor Temp.: 90.72 Ambient Temp.: 83.07 F
14:27:31.749 -> Average Cal sensor Temp.: 90.72 Ambient Temp.: 83.03 F
14:27:33.300 -> Average Cal sensor Temp.: 90.71 Ambient Temp.: 83.03 F
14:27:34.856 -> Average Cal sensor Temp.: 90.53 Ambient Temp.: 83.03 F
14:27:36.414 -> Average Cal sensor Temp.: 90.24 Ambient Temp.: 82.96 F
14:27:37.967 -> Average Cal sensor Temp.: 90.38 Ambient Temp.: 82.96 F
14:27:39.525 -> Average Cal sensor Temp.: 90.62 Ambient Temp.: 82.96 F
14:27:41.079 -> Average Cal sensor Temp.: 91.19 Ambient Temp.: 82.96 F
14:27:42.633 -> Average Cal sensor Temp.: 91.19 Ambient Temp.: 82.96 F
14:27:44.186 -> Average Cal sensor Temp.: 90.91 Ambient Temp.: 82.96 F
14:27:45.769 -> Average Cal sensor Temp.: 91.05 Ambient Temp.: 82.89 F
14:27:47.326 -> Average Cal sensor Temp.: 90.79 Ambient Temp.: 82.96 F
14:27:48.881 -> Average Cal sensor Temp.: 90.77 Ambient Temp.: 82.96 F
14:27:50.433 -> Average Cal sensor Temp.: 90.64 Ambient Temp.: 82.96 F
14:27:51.989 -> Average Cal sensor Temp.: 90.70 Ambient Temp.: 82.96 F
14:27:53.545 -> Average Cal sensor Temp.: 90.63 Ambient Temp.: 82.92 F
14:27:55.093 -> Average Cal sensor Temp.: 90.91 Ambient Temp.: 82.92 F
14:27:56.645 -> Average Cal sensor Temp.: 90.58 Ambient Temp.: 82.89 F
14:27:58.203 -> Average Cal sensor Temp.: 90.67 Ambient Temp.: 82.89 F
14:27:59.785 -> Average Cal sensor Temp.: 90.79 Ambient Temp.: 82.89 F
14:28:01.339 -> Average Cal sensor Temp.: 90.57 Ambient Temp.: 82.85 F
14:28:02.892 -> Average Cal sensor Temp.: 90.78 Ambient Temp.: 82.85 F
14:28:04.449 -> Average Cal sensor Temp.: 90.97 Ambient Temp.: 82.85 F
14:28:06.007 -> Average Cal sensor Temp.: 90.62 Ambient Temp.: 82.78 F
14:28:07.564 -> Average Cal sensor Temp.: 90.69 Ambient Temp.: 82.85 F
14:28:09.120 -> Average Cal sensor Temp.: 90.73 Ambient Temp.: 82.85 F
14:28:10.680 -> Average Cal sensor Temp.: 90.41 Ambient Temp.: 82.78 F
14:28:12.230 -> Average Cal sensor Temp.: 90.73 Ambient Temp.: 82.74 F
14:28:13.791 -> Average Cal sensor Temp.: 90.82 Ambient Temp.: 82.78 F
14:28:15.345 -> Average Cal sensor Temp.: 90.52 Ambient Temp.: 82.78 F
14:28:16.896 -> Average Cal sensor Temp.: 90.80 Ambient Temp.: 82.74 F
14:28:18.450 -> Average Cal sensor Temp.: 90.39 Ambient Temp.: 82.63 F
14:28:20.004 -> Average Cal sensor Temp.: 90.77 Ambient Temp.: 82.67 F
14:28:21.560 -> Average Cal sensor Temp.: 90.85 Ambient Temp.: 82.60 F
14:28:23.117 -> Average Cal sensor Temp.: 90.76 Ambient Temp.: 82.60 F
14:28:24.668 -> Average Cal sensor Temp.: 90.91 Ambient Temp.: 82.63 F
14:28:26.229 -> Average Cal sensor Temp.: 90.82 Ambient Temp.: 82.63 F
14:28:27.782 -> Average Cal sensor Temp.: 90.72 Ambient Temp.: 82.56 F
14:28:29.344 -> Average Cal sensor Temp.: 90.98 Ambient Temp.: 82.53 F
14:28:30.903 -> Average Cal sensor Temp.: 91.18 Ambient Temp.: 82.53 F
14:28:32.458 -> Average Cal sensor Temp.: 91.10 Ambient Temp.: 82.53 F
14:28:34.012 -> Average Cal sensor Temp.: 90.95 Ambient Temp.: 82.56 F
14:28:35.565 -> Average Cal sensor Temp.: 91.09 Ambient Temp.: 82.56 F
14:28:37.123 -> Average Cal sensor Temp.: 90.69 Ambient Temp.: 82.49 F
14:28:38.678 -> Average Cal sensor Temp.: 91.03 Ambient Temp.: 82.49 F
14:28:40.231 -> Average Cal sensor Temp.: 91.05 Ambient Temp.: 82.53 F
14:28:41.784 -> Average Cal sensor Temp.: 90.72 Ambient Temp.: 82.53 F
14:28:43.367 -> Average Cal sensor Temp.: 90.74 Ambient Temp.: 82.49 F
14:28:44.926 -> Average Cal sensor Temp.: 90.84 Ambient Temp.: 82.49 F
14:28:46.480 -> Average Cal sensor Temp.: 91.31 Ambient Temp.: 82.45 F
14:28:48.037 -> Average Cal sensor Temp.: 91.12 Ambient Temp.: 82.49 F
14:28:49.591 -> Average Cal sensor Temp.: 90.87 Ambient Temp.: 82.45 F
14:28:51.146 -> Average Cal sensor Temp.: 91.04 Ambient Temp.: 82.45 F
14:28:52.698 -> Average Cal sensor Temp.: 91.13 Ambient Temp.: 82.45 F
14:28:54.255 -> Average Cal sensor Temp.: 90.98 Ambient Temp.: 82.49 F
14:28:55.806 -> Average Cal sensor Temp.: 90.95 Ambient Temp.: 82.49 F
14:28:57.357 -> Average Cal sensor Temp.: 90.78 Ambient Temp.: 82.45 F
14:28:58.905 -> Average Cal sensor Temp.: 90.86 Ambient Temp.: 82.45 F
14:29:00.462 -> Average Cal sensor Temp.: 91.05 Ambient Temp.: 82.45 F
14:29:02.012 -> Average Cal sensor Temp.: 91.19 Ambient Temp.: 82.45 F
14:29:03.565 -> Average Cal sensor Temp.: 91.31 Ambient Temp.: 82.45 F
14:29:05.147 -> Average Cal sensor Temp.: 90.96 Ambient Temp.: 82.49 F
14:29:06.705 -> Average Cal sensor Temp.: 91.03 Ambient Temp.: 82.49 F
14:29:08.259 -> Average Cal sensor Temp.: 91.27 Ambient Temp.: 82.49 F
14:29:09.809 -> Average Cal sensor Temp.: 91.26 Ambient Temp.: 82.45 F
14:29:11.370 -> Average Cal sensor Temp.: 91.19 Ambient Temp.: 82.53 F
14:29:12.921 -> Average Cal sensor Temp.: 91.21 Ambient Temp.: 82.56 F
14:29:14.481 -> Average Cal sensor Temp.: 91.30 Ambient Temp.: 82.53 F
14:29:16.037 -> Average Cal sensor Temp.: 91.13 Ambient Temp.: 82.53 F
14:29:17.594 -> Average Cal sensor Temp.: 91.34 Ambient Temp.: 82.53 F
14:29:19.148 -> Average Cal sensor Temp.: 91.09 Ambient Temp.: 82.56 F
14:29:20.704 -> Average Cal sensor Temp.: 91.14 Ambient Temp.: 82.56 F
14:29:22.258 -> Average Cal sensor Temp.: 91.26 Ambient Temp.: 82.60 F
14:29:23.812 -> Average Cal sensor Temp.: 91.00 Ambient Temp.: 82.56 F
14:29:25.362 -> Average Cal sensor Temp.: 90.76 Ambient Temp.: 82.60 F
14:29:26.915 -> Average Cal sensor Temp.: 91.17 Ambient Temp.: 82.53 F
14:29:28.465 -> Average Cal sensor Temp.: 91.22 Ambient Temp.: 82.56 F
14:29:30.023 -> Average Cal sensor Temp.: 91.26 Ambient Temp.: 82.60 F
14:29:31.576 -> Average Cal sensor Temp.: 91.37 Ambient Temp.: 82.60 F
14:29:33.128 -> Average Cal sensor Temp.: 91.42 Ambient Temp.: 82.60 F
14:29:34.684 -> Average Cal sensor Temp.: 91.17 Ambient Temp.: 82.63 F
14:29:36.239 -> Average Cal sensor Temp.: 91.27 Ambient Temp.: 82.60 F
14:29:37.791 -> Average Cal sensor Temp.: 90.89 Ambient Temp.: 82.63 F
14:29:39.349 -> Average Cal sensor Temp.: 91.27 Ambient Temp.: 82.67 F
14:29:40.909 -> Average Cal sensor Temp.: 90.88 Ambient Temp.: 82.67 F
14:29:42.465 -> Average Cal sensor Temp.: 91.40 Ambient Temp.: 82.63 F
14:29:44.019 -> Average Cal sensor Temp.: 91.32 Ambient Temp.: 82.67 F
14:29:45.579 -> Average Cal sensor Temp.: 91.08 Ambient Temp.: 82.74 F
14:29:47.131 -> Average Cal sensor Temp.: 91.19 Ambient Temp.: 82.67 F
14:29:48.690 -> Average Cal sensor Temp.: 91.37 Ambient Temp.: 82.67 F
14:29:50.241 -> Average Cal sensor Temp.: 91.47 Ambient Temp.: 82.71 F
14:29:51.790 -> Average Cal sensor Temp.: 91.52 Ambient Temp.: 82.71 F
14:29:53.344 -> Average Cal sensor Temp.: 91.20 Ambient Temp.: 82.78 F
14:29:54.897 -> Average Cal sensor Temp.: 91.35 Ambient Temp.: 82.74 F
14:29:56.453 -> Average Cal sensor Temp.: 91.50 Ambient Temp.: 82.78 F
14:29:58.002 -> Average Cal sensor Temp.: 91.45 Ambient Temp.: 82.78 F
14:29:59.584 -> Average Cal sensor Temp.: 91.69 Ambient Temp.: 82.89 F
14:30:01.135 -> Average Cal sensor Temp.: 91.59 Ambient Temp.: 82.78 F
14:30:02.687 -> Average Cal sensor Temp.: 91.65 Ambient Temp.: 82.92 F
14:30:04.243 -> Average Cal sensor Temp.: 91.02 Ambient Temp.: 82.89 F
14:30:05.801 -> Average Cal sensor Temp.: 90.72 Ambient Temp.: 82.92 F
14:30:07.356 -> Average Cal sensor Temp.: 91.18 Ambient Temp.: 82.96 F
14:30:08.912 -> Average Cal sensor Temp.: 91.17 Ambient Temp.: 82.92 F
14:30:10.466 -> Average Cal sensor Temp.: 91.22 Ambient Temp.: 82.92 F
14:30:12.023 -> Average Cal sensor Temp.: 91.40 Ambient Temp.: 82.96 F
14:30:13.576 -> Average Cal sensor Temp.: 91.22 Ambient Temp.: 82.99 F
14:30:15.129 -> Average Cal sensor Temp.: 91.36 Ambient Temp.: 83.03 F
14:30:16.683 -> Average Cal sensor Temp.: 91.22 Ambient Temp.: 82.99 F
14:30:18.243 -> Average Cal sensor Temp.: 91.15 Ambient Temp.: 82.96 F
14:30:19.797 -> Average Cal sensor Temp.: 91.34 Ambient Temp.: 83.03 F
14:30:21.350 -> Average Cal sensor Temp.: 91.44 Ambient Temp.: 83.07 F
14:30:22.899 -> Average Cal sensor Temp.: 91.65 Ambient Temp.: 83.07 F
14:30:24.454 -> Average Cal sensor Temp.: 91.86 Ambient Temp.: 83.07 F
14:30:26.013 -> Average Cal sensor Temp.: 91.85 Ambient Temp.: 83.14 F
14:30:27.566 -> Average Cal sensor Temp.: 91.64 Ambient Temp.: 83.10 F
14:30:29.123 -> Average Cal sensor Temp.: 91.46 Ambient Temp.: 83.14 F
14:30:30.672 -> Average Cal sensor Temp.: 91.60 Ambient Temp.: 83.17 F
14:30:32.226 -> Average Cal sensor Temp.: 91.41 Ambient Temp.: 83.17 F
14:30:33.786 -> Average Cal sensor Temp.: 91.43 Ambient Temp.: 83.25 F
14:30:35.338 -> Average Cal sensor Temp.: 91.58 Ambient Temp.: 83.25 F
14:30:36.894 -> Average Cal sensor Temp.: 91.67 Ambient Temp.: 83.28 F
14:30:38.448 -> Average Cal sensor Temp.: 91.75 Ambient Temp.: 83.35 F
14:30:39.995 -> Average Cal sensor Temp.: 91.81 Ambient Temp.: 83.35 F
14:30:41.547 -> Average Cal sensor Temp.: 91.73 Ambient Temp.: 83.32 F
14:30:43.098 -> Average Cal sensor Temp.: 91.70 Ambient Temp.: 83.35 F
14:30:44.653 -> Average Cal sensor Temp.: 91.74 Ambient Temp.: 83.39 F
14:30:46.208 -> Average Cal sensor Temp.: 91.70 Ambient Temp.: 83.43 F
14:30:47.795 -> Average Cal sensor Temp.: 91.53 Ambient Temp.: 83.43 F
14:30:49.348 -> Average Cal sensor Temp.: 91.33 Ambient Temp.: 83.43 F
14:30:50.903 -> Average Cal sensor Temp.: 91.56 Ambient Temp.: 83.46 F
14:30:52.455 -> Average Cal sensor Temp.: 91.59 Ambient Temp.: 83.50 F
14:30:54.011 -> Average Cal sensor Temp.: 91.66 Ambient Temp.: 83.50 F
14:30:55.570 -> Average Cal sensor Temp.: 91.48 Ambient Temp.: 83.50 F
14:30:57.121 -> Average Cal sensor Temp.: 91.54 Ambient Temp.: 83.57 F
14:30:58.678 -> Average Cal sensor Temp.: 91.59 Ambient Temp.: 83.53 F
14:31:00.224 -> Average Cal sensor Temp.: 91.38 Ambient Temp.: 83.64 F
14:31:01.782 -> Average Cal sensor Temp.: 91.23 Ambient Temp.: 83.68 F
14:31:03.333 -> Average Cal sensor Temp.: 91.40 Ambient Temp.: 83.64 F
14:31:04.881 -> Average Cal sensor Temp.: 91.38 Ambient Temp.: 83.71 F
14:31:06.436 -> Average Cal sensor Temp.: 91.59 Ambient Temp.: 83.75 F
14:31:07.990 -> Average Cal sensor Temp.: 91.71 Ambient Temp.: 83.79 F
14:31:09.548 -> Average Cal sensor Temp.: 91.63 Ambient Temp.: 83.79 F
14:31:11.106 -> Average Cal sensor Temp.: 91.44 Ambient Temp.: 83.82 F
14:31:12.661 -> Average Cal sensor Temp.: 91.58 Ambient Temp.: 83.82 F
14:31:14.215 -> Average Cal sensor Temp.: 91.72 Ambient Temp.: 83.82 F
14:31:15.766 -> Average Cal sensor Temp.: 91.69 Ambient Temp.: 83.86 F
14:31:17.323 -> Average Cal sensor Temp.: 91.71 Ambient Temp.: 83.89 F
14:31:18.876 -> Average Cal sensor Temp.: 91.65 Ambient Temp.: 83.93 F
14:31:20.431 -> Average Cal sensor Temp.: 91.85 Ambient Temp.: 83.97 F
14:31:21.984 -> Average Cal sensor Temp.: 91.96 Ambient Temp.: 84.00 F
14:31:23.540 -> Average Cal sensor Temp.: 91.72 Ambient Temp.: 84.00 F
14:31:25.092 -> Average Cal sensor Temp.: 91.75 Ambient Temp.: 84.00 F
14:31:26.677 -> Average Cal sensor Temp.: 91.89 Ambient Temp.: 84.15 F
14:31:28.227 -> Average Cal sensor Temp.: 91.94 Ambient Temp.: 84.15 F
14:31:29.782 -> Average Cal sensor Temp.: 91.38 Ambient Temp.: 84.15 F
14:31:31.331 -> Average Cal sensor Temp.: 91.44 Ambient Temp.: 84.11 F
14:31:32.887 -> Average Cal sensor Temp.: 91.69 Ambient Temp.: 84.11 F
14:31:34.445 -> Average Cal sensor Temp.: 91.57 Ambient Temp.: 84.18 F
14:31:35.999 -> Average Cal sensor Temp.: 91.53 Ambient Temp.: 84.18 F
14:31:37.556 -> Average Cal sensor Temp.: 91.33 Ambient Temp.: 84.22 F
14:31:39.112 -> Average Cal sensor Temp.: 91.49 Ambient Temp.: 84.25 F
14:31:40.672 -> Average Cal sensor Temp.: 91.38 Ambient Temp.: 84.29 F
14:31:42.226 -> Average Cal sensor Temp.: 91.39 Ambient Temp.: 84.25 F
14:31:43.783 -> Average Cal sensor Temp.: 91.50 Ambient Temp.: 84.29 F
14:31:45.334 -> Average Cal sensor Temp.: 91.76 Ambient Temp.: 84.25 F
14:31:46.887 -> Average Cal sensor Temp.: 91.71 Ambient Temp.: 84.33 F
14:31:48.440 -> Average Cal sensor Temp.: 91.74 Ambient Temp.: 84.36 F
14:31:49.993 -> Average Cal sensor Temp.: 91.67 Ambient Temp.: 84.40 F
14:31:51.549 -> Average Cal sensor Temp.: 91.71 Ambient Temp.: 84.33 F
14:31:53.100 -> Average Cal sensor Temp.: 91.35 Ambient Temp.: 84.36 F
14:31:54.654 -> Average Cal sensor Temp.: 91.98 Ambient Temp.: 84.33 F
14:31:56.208 -> Average Cal sensor Temp.: 92.01 Ambient Temp.: 84.40 F
14:31:57.759 -> Average Cal sensor Temp.: 91.78 Ambient Temp.: 84.40 F
14:31:59.314 -> Average Cal sensor Temp.: 91.73 Ambient Temp.: 84.47 F
14:32:00.868 -> Average Cal sensor Temp.: 91.81 Ambient Temp.: 84.47 F
14:32:02.420 -> Average Cal sensor Temp.: 91.54 Ambient Temp.: 84.40 F
14:32:03.979 -> Average Cal sensor Temp.: 91.35 Ambient Temp.: 84.40 F
14:32:05.528 -> Average Cal sensor Temp.: 91.62 Ambient Temp.: 84.40 F
14:32:07.113 -> Average Cal sensor Temp.: 91.47 Ambient Temp.: 84.51 F
14:32:08.662 -> Average Cal sensor Temp.: 91.27 Ambient Temp.: 84.47 F
14:32:10.215 -> Average Cal sensor Temp.: 91.29 Ambient Temp.: 84.54 F
14:32:11.767 -> Average Cal sensor Temp.: 91.02 Ambient Temp.: 84.47 F
14:32:13.322 -> Average Cal sensor Temp.: 91.02 Ambient Temp.: 84.47 F
14:32:14.873 -> Average Cal sensor Temp.: 91.37 Ambient Temp.: 84.47 F
14:32:16.422 -> Average Cal sensor Temp.: 91.44 Ambient Temp.: 84.47 F
14:32:17.975 -> Average Cal sensor Temp.: 91.44 Ambient Temp.: 84.47 F
14:32:19.530 -> Average Cal sensor Temp.: 91.67 Ambient Temp.: 84.47 F
14:32:21.083 -> Average Cal sensor Temp.: 91.72 Ambient Temp.: 84.47 F
14:32:22.644 -> Average Cal sensor Temp.: 91.62 Ambient Temp.: 84.40 F
14:32:24.195 -> Average Cal sensor Temp.: 91.58 Ambient Temp.: 84.36 F
14:32:25.783 -> Average Cal sensor Temp.: 91.50 Ambient Temp.: 84.36 F
14:32:27.333 -> Average Cal sensor Temp.: 91.58 Ambient Temp.: 84.47 F
14:32:28.885 -> Average Cal sensor Temp.: 91.45 Ambient Temp.: 84.40 F
14:32:30.441 -> Average Cal sensor Temp.: 91.58 Ambient Temp.: 84.40 F
14:32:31.993 -> Average Cal sensor Temp.: 91.67 Ambient Temp.: 84.51 F
14:32:33.542 -> Average Cal sensor Temp.: 91.62 Ambient Temp.: 84.36 F
14:32:35.098 -> Average Cal sensor Temp.: 92.03 Ambient Temp.: 84.40 F
14:32:36.653 -> Average Cal sensor Temp.: 91.96 Ambient Temp.: 84.40 F
14:32:38.205 -> Average Cal sensor Temp.: 92.30 Ambient Temp.: 84.40 F
14:32:39.757 -> Average Cal sensor Temp.: 91.97 Ambient Temp.: 84.36 F
14:32:41.313 -> Average Cal sensor Temp.: 91.79 Ambient Temp.: 84.54 F
14:32:42.867 -> Average Cal sensor Temp.: 91.44 Ambient Temp.: 84.40 F
14:32:44.456 -> Average Cal sensor Temp.: 91.62 Ambient Temp.: 84.40 F
14:32:46.012 -> Average Cal sensor Temp.: 91.85 Ambient Temp.: 84.47 F
14:32:47.568 -> Average Cal sensor Temp.: 91.67 Ambient Temp.: 84.36 F
14:32:49.119 -> Average Cal sensor Temp.: 91.35 Ambient Temp.: 84.40 F
14:32:50.673 -> Average Cal sensor Temp.: 91.67 Ambient Temp.: 84.40 F
14:32:52.227 -> Average Cal sensor Temp.: 91.76 Ambient Temp.: 84.40 F
14:32:53.779 -> Average Cal sensor Temp.: 92.03 Ambient Temp.: 84.36 F
14:32:55.327 -> Average Cal sensor Temp.: 91.65 Ambient Temp.: 84.36 F
14:32:56.880 -> Average Cal sensor Temp.: 91.87 Ambient Temp.: 84.36 F
14:32:58.430 -> Average Cal sensor Temp.: 92.10 Ambient Temp.: 84.40 F
14:32:59.982 -> Average Cal sensor Temp.: 92.18 Ambient Temp.: 84.29 F
14:33:01.566 -> Average Cal sensor Temp.: 92.05 Ambient Temp.: 84.33 F
14:33:03.115 -> Average Cal sensor Temp.: 92.09 Ambient Temp.: 84.29 F
14:33:04.669 -> Average Cal sensor Temp.: 91.79 Ambient Temp.: 84.29 F
14:33:06.225 -> Average Cal sensor Temp.: 91.49 Ambient Temp.: 84.29 F
14:33:07.777 -> Average Cal sensor Temp.: 91.72 Ambient Temp.: 84.22 F
14:33:09.325 -> Average Cal sensor Temp.: 92.01 Ambient Temp.: 84.29 F
14:33:10.878 -> Average Cal sensor Temp.: 91.77 Ambient Temp.: 84.25 F
14:33:12.429 -> Average Cal sensor Temp.: 92.12 Ambient Temp.: 84.25 F
14:33:14.016 -> Average Cal sensor Temp.: 92.27 Ambient Temp.: 84.25 F
14:33:15.567 -> Average Cal sensor Temp.: 92.07 Ambient Temp.: 84.25 F
14:33:17.111 -> Average Cal sensor Temp.: 92.17 Ambient Temp.: 84.22 F
14:33:18.660 -> Average Cal sensor Temp.: 91.49 Ambient Temp.: 84.18 F
14:33:20.211 -> Average Cal sensor Temp.: 91.44 Ambient Temp.: 84.22 F
14:33:21.767 -> Average Cal sensor Temp.: 91.85 Ambient Temp.: 84.18 F
14:33:23.316 -> Average Cal sensor Temp.: 92.03 Ambient Temp.: 84.18 F
14:33:24.902 -> Average Cal sensor Temp.: 91.87 Ambient Temp.: 84.15 F
14:33:26.451 -> Average Cal sensor Temp.: 91.79 Ambient Temp.: 84.18 F
14:33:28.000 -> Average Cal sensor Temp.: 91.76 Ambient Temp.: 84.15 F
14:33:29.554 -> Average Cal sensor Temp.: 91.96 Ambient Temp.: 84.07 F
14:33:31.105 -> Average Cal sensor Temp.: 92.03 Ambient Temp.: 84.07 F
14:33:32.663 -> Average Cal sensor Temp.: 92.34 Ambient Temp.: 84.07 F
14:33:34.217 -> Average Cal sensor Temp.: 92.28 Ambient Temp.: 84.11 F
14:33:35.769 -> Average Cal sensor Temp.: 92.51 Ambient Temp.: 84.07 F
14:33:37.323 -> Average Cal sensor Temp.: 92.45 Ambient Temp.: 84.00 F
14:33:38.878 -> Average Cal sensor Temp.: 92.32 Ambient Temp.: 84.07 F
14:33:40.433 -> Average Cal sensor Temp.: 92.03 Ambient Temp.: 84.07 F
14:33:42.014 -> Average Cal sensor Temp.: 91.92 Ambient Temp.: 83.97 F
14:33:43.564 -> Average Cal sensor Temp.: 91.77 Ambient Temp.: 83.97 F
14:33:45.118 -> Average Cal sensor Temp.: 91.83 Ambient Temp.: 84.00 F
14:33:46.680 -> Average Cal sensor Temp.: 91.91 Ambient Temp.: 83.93 F
14:33:48.233 -> Average Cal sensor Temp.: 91.99 Ambient Temp.: 83.97 F
14:33:49.783 -> Average Cal sensor Temp.: 92.22 Ambient Temp.: 83.97 F
14:33:51.332 -> Average Cal sensor Temp.: 92.41 Ambient Temp.: 83.93 F
14:33:52.878 -> Average Cal sensor Temp.: 92.26 Ambient Temp.: 83.97 F
14:33:54.429 -> Average Cal sensor Temp.: 92.31 Ambient Temp.: 84.00 F
14:33:56.014 -> Average Cal sensor Temp.: 92.16 Ambient Temp.: 83.89 F
14:33:57.564 -> Average Cal sensor Temp.: 91.91 Ambient Temp.: 83.89 F
14:33:59.110 -> Average Cal sensor Temp.: 91.67 Ambient Temp.: 83.86 F
14:34:00.669 -> Average Cal sensor Temp.: 91.48 Ambient Temp.: 83.86 F
14:34:02.216 -> Average Cal sensor Temp.: 91.87 Ambient Temp.: 83.86 F
14:34:03.766 -> Average Cal sensor Temp.: 92.05 Ambient Temp.: 83.86 F
14:34:05.353 -> Average Cal sensor Temp.: 91.87 Ambient Temp.: 83.79 F
14:34:06.906 -> Average Cal sensor Temp.: 92.04 Ambient Temp.: 83.79 F
14:34:08.460 -> Average Cal sensor Temp.: 91.87 Ambient Temp.: 83.79 F
14:34:10.010 -> Average Cal sensor Temp.: 91.59 Ambient Temp.: 83.75 F
14:34:11.567 -> Average Cal sensor Temp.: 91.70 Ambient Temp.: 83.75 F
14:34:13.118 -> Average Cal sensor Temp.: 91.72 Ambient Temp.: 83.71 F
14:34:14.671 -> Average Cal sensor Temp.: 91.99 Ambient Temp.: 83.68 F
14:34:16.228 -> Average Cal sensor Temp.: 92.13 Ambient Temp.: 83.71 F
14:34:17.778 -> Average Cal sensor Temp.: 91.96 Ambient Temp.: 83.68 F
14:34:19.336 -> Average Cal sensor Temp.: 91.85 Ambient Temp.: 83.64 F
14:34:20.892 -> Average Cal sensor Temp.: 92.10 Ambient Temp.: 83.64 F
14:34:22.448 -> Average Cal sensor Temp.: 92.23 Ambient Temp.: 83.57 F
14:34:24.000 -> Average Cal sensor Temp.: 92.11 Ambient Temp.: 83.57 F
14:34:25.561 -> Average Cal sensor Temp.: 92.38 Ambient Temp.: 83.50 F
14:34:27.116 -> Average Cal sensor Temp.: 92.37 Ambient Temp.: 83.53 F
14:34:28.673 -> Average Cal sensor Temp.: 91.76 Ambient Temp.: 83.53 F
14:34:30.231 -> Average Cal sensor Temp.: 91.93 Ambient Temp.: 83.50 F
14:34:31.780 -> Average Cal sensor Temp.: 92.30 Ambient Temp.: 83.50 F
14:34:33.336 -> Average Cal sensor Temp.: 92.16 Ambient Temp.: 83.50 F
14:34:34.890 -> Average Cal sensor Temp.: 92.21 Ambient Temp.: 83.43 F
14:34:36.447 -> Average Cal sensor Temp.: 91.89 Ambient Temp.: 83.39 F
14:34:37.999 -> Average Cal sensor Temp.: 92.15 Ambient Temp.: 83.39 F
14:34:39.555 -> Average Cal sensor Temp.: 92.13 Ambient Temp.: 83.43 F
14:34:41.111 -> Average Cal sensor Temp.: 92.13 Ambient Temp.: 83.39 F
14:34:42.667 -> Average Cal sensor Temp.: 92.10 Ambient Temp.: 83.39 F
14:34:44.217 -> Average Cal sensor Temp.: 92.31 Ambient Temp.: 83.35 F
14:34:45.770 -> Average Cal sensor Temp.: 92.60 Ambient Temp.: 83.39 F
14:34:47.323 -> Average Cal sensor Temp.: 92.20 Ambient Temp.: 83.32 F
14:34:48.881 -> Average Cal sensor Temp.: 92.36 Ambient Temp.: 83.35 F
14:34:50.436 -> Average Cal sensor Temp.: 92.81 Ambient Temp.: 83.28 F
14:34:51.989 -> Average Cal sensor Temp.: 92.43 Ambient Temp.: 83.32 F
14:34:53.541 -> Average Cal sensor Temp.: 92.57 Ambient Temp.: 83.39 F
14:34:55.097 -> Average Cal sensor Temp.: 92.43 Ambient Temp.: 83.35 F
14:34:56.653 -> Average Cal sensor Temp.: 92.09 Ambient Temp.: 83.32 F
14:34:58.234 -> Average Cal sensor Temp.: 91.97 Ambient Temp.: 83.25 F
14:34:59.788 -> Average Cal sensor Temp.: 91.79 Ambient Temp.: 83.32 F
14:35:01.333 -> Average Cal sensor Temp.: 91.97 Ambient Temp.: 83.28 F
14:35:02.885 -> Average Cal sensor Temp.: 92.18 Ambient Temp.: 83.25 F
14:35:04.440 -> Average Cal sensor Temp.: 92.16 Ambient Temp.: 83.25 F
14:35:05.997 -> Average Cal sensor Temp.: 92.46 Ambient Temp.: 83.25 F
14:35:07.553 -> Average Cal sensor Temp.: 92.49 Ambient Temp.: 83.17 F
14:35:09.102 -> Average Cal sensor Temp.: 92.31 Ambient Temp.: 83.14 F
14:35:10.663 -> Average Cal sensor Temp.: 92.15 Ambient Temp.: 83.14 F
14:35:12.213 -> Average Cal sensor Temp.: 91.71 Ambient Temp.: 83.10 F
14:35:13.768 -> Average Cal sensor Temp.: 92.32 Ambient Temp.: 83.14 F
14:35:15.320 -> Average Cal sensor Temp.: 92.52 Ambient Temp.: 83.07 F
14:35:16.873 -> Average Cal sensor Temp.: 92.66 Ambient Temp.: 83.03 F
14:35:18.429 -> Average Cal sensor Temp.: 92.46 Ambient Temp.: 83.10 F
14:35:19.978 -> Average Cal sensor Temp.: 92.67 Ambient Temp.: 83.07 F
14:35:21.538 -> Average Cal sensor Temp.: 92.16 Ambient Temp.: 83.10 F
14:35:23.085 -> Average Cal sensor Temp.: 92.02 Ambient Temp.: 82.99 F
14:35:24.670 -> Average Cal sensor Temp.: 92.48 Ambient Temp.: 82.96 F
14:35:26.224 -> Average Cal sensor Temp.: 92.66 Ambient Temp.: 82.96 F
14:35:27.780 -> Average Cal sensor Temp.: 92.85 Ambient Temp.: 82.96 F
14:35:29.336 -> Average Cal sensor Temp.: 92.69 Ambient Temp.: 82.99 F
14:35:30.890 -> Average Cal sensor Temp.: 92.92 Ambient Temp.: 82.99 F
14:35:32.447 -> Average Cal sensor Temp.: 92.52 Ambient Temp.: 82.96 F
14:35:33.999 -> Average Cal sensor Temp.: 92.46 Ambient Temp.: 82.96 F
14:35:35.562 -> Average Cal sensor Temp.: 92.46 Ambient Temp.: 82.96 F
14:35:37.110 -> Average Cal sensor Temp.: 92.10 Ambient Temp.: 82.92 F
14:35:38.666 -> Average Cal sensor Temp.: 92.01 Ambient Temp.: 82.92 F
14:35:40.224 -> Average Cal sensor Temp.: 92.35 Ambient Temp.: 82.96 F
14:35:41.782 -> Average Cal sensor Temp.: 92.78 Ambient Temp.: 82.92 F
14:35:43.337 -> Average Cal sensor Temp.: 92.39 Ambient Temp.: 82.96 F
14:35:44.894 -> Average Cal sensor Temp.: 92.21 Ambient Temp.: 82.96 F
14:35:46.447 -> Average Cal sensor Temp.: 92.49 Ambient Temp.: 82.96 F
14:35:48.005 -> Average Cal sensor Temp.: 92.74 Ambient Temp.: 82.99 F
14:35:49.561 -> Average Cal sensor Temp.: 93.02 Ambient Temp.: 82.96 F
14:35:51.112 -> Average Cal sensor Temp.: 92.98 Ambient Temp.: 82.99 F
14:35:52.660 -> Average Cal sensor Temp.: 92.85 Ambient Temp.: 82.99 F
14:35:54.246 -> Average Cal sensor Temp.: 92.69 Ambient Temp.: 82.96 F
14:35:55.801 -> Average Cal sensor Temp.: 92.19 Ambient Temp.: 82.99 F
14:35:57.352 -> Average Cal sensor Temp.: 92.61 Ambient Temp.: 82.99 F
14:35:58.908 -> Average Cal sensor Temp.: 92.63 Ambient Temp.: 82.96 F
14:36:00.463 -> Average Cal sensor Temp.: 92.75 Ambient Temp.: 82.96 F
14:36:02.010 -> Average Cal sensor Temp.: 92.75 Ambient Temp.: 82.96 F
14:36:03.568 -> Average Cal sensor Temp.: 92.94 Ambient Temp.: 82.99 F
14:36:05.122 -> Average Cal sensor Temp.: 92.79 Ambient Temp.: 83.03 F
14:36:06.673 -> Average Cal sensor Temp.: 92.82 Ambient Temp.: 83.07 F
14:36:08.228 -> Average Cal sensor Temp.: 92.82 Ambient Temp.: 83.07 F
14:36:09.815 -> Average Cal sensor Temp.: 92.50 Ambient Temp.: 83.07 F
14:36:11.336 -> Average Cal sensor Temp.: 92.44 Ambient Temp.: 82.99 F
14:36:12.895 -> Average Cal sensor Temp.: 92.76 Ambient Temp.: 83.03 F
14:36:14.480 -> Average Cal sensor Temp.: 92.65 Ambient Temp.: 83.07 F
14:36:16.037 -> Average Cal sensor Temp.: 92.52 Ambient Temp.: 83.07 F
14:36:17.591 -> Average Cal sensor Temp.: 92.00 Ambient Temp.: 83.07 F
14:36:19.139 -> Average Cal sensor Temp.: 92.33 Ambient Temp.: 83.07 F
14:36:20.697 -> Average Cal sensor Temp.: 92.49 Ambient Temp.: 83.03 F
14:36:22.246 -> Average Cal sensor Temp.: 92.40 Ambient Temp.: 83.03 F
14:36:23.793 -> Average Cal sensor Temp.: 92.45 Ambient Temp.: 83.03 F
14:36:25.345 -> Average Cal sensor Temp.: 92.64 Ambient Temp.: 83.03 F
14:36:26.897 -> Average Cal sensor Temp.: 92.64 Ambient Temp.: 83.03 F
14:36:28.485 -> Average Cal sensor Temp.: 92.66 Ambient Temp.: 83.03 F
14:36:30.035 -> Average Cal sensor Temp.: 92.51 Ambient Temp.: 83.07 F
14:36:31.588 -> Average Cal sensor Temp.: 92.73 Ambient Temp.: 83.07 F
14:36:33.135 -> Average Cal sensor Temp.: 92.67 Ambient Temp.: 83.03 F
14:36:34.687 -> Average Cal sensor Temp.: 92.61 Ambient Temp.: 83.03 F
14:36:36.239 -> Average Cal sensor Temp.: 92.43 Ambient Temp.: 83.07 F
14:36:37.796 -> Average Cal sensor Temp.: 92.07 Ambient Temp.: 83.10 F
14:36:39.346 -> Average Cal sensor Temp.: 92.05 Ambient Temp.: 83.07 F
14:36:40.903 -> Average Cal sensor Temp.: 92.00 Ambient Temp.: 83.10 F
14:36:42.459 -> Average Cal sensor Temp.: 92.37 Ambient Temp.: 83.10 F
14:36:44.012 -> Average Cal sensor Temp.: 92.36 Ambient Temp.: 83.10 F
14:36:45.566 -> Average Cal sensor Temp.: 92.54 Ambient Temp.: 83.14 F
14:36:47.119 -> Average Cal sensor Temp.: 92.74 Ambient Temp.: 83.14 F
14:36:48.708 -> Average Cal sensor Temp.: 92.48 Ambient Temp.: 83.10 F
14:36:50.259 -> Average Cal sensor Temp.: 92.43 Ambient Temp.: 83.10 F
14:36:51.812 -> Average Cal sensor Temp.: 92.33 Ambient Temp.: 83.14 F
14:36:53.368 -> Average Cal sensor Temp.: 92.78 Ambient Temp.: 83.10 F
14:36:54.921 -> Average Cal sensor Temp.: 92.43 Ambient Temp.: 83.10 F
14:36:56.477 -> Average Cal sensor Temp.: 92.65 Ambient Temp.: 83.10 F
14:36:58.035 -> Average Cal sensor Temp.: 92.17 Ambient Temp.: 83.14 F
14:36:59.591 -> Average Cal sensor Temp.: 92.28 Ambient Temp.: 83.14 F
14:37:01.140 -> Average Cal sensor Temp.: 92.68 Ambient Temp.: 83.14 F
14:37:02.690 -> Average Cal sensor Temp.: 92.95 Ambient Temp.: 83.14 F
14:37:04.244 -> Average Cal sensor Temp.: 92.82 Ambient Temp.: 83.17 F
14:37:05.799 -> Average Cal sensor Temp.: 92.38 Ambient Temp.: 83.17 F
14:37:07.352 -> Average Cal sensor Temp.: 92.71 Ambient Temp.: 83.25 F
14:37:08.909 -> Average Cal sensor Temp.: 92.71 Ambient Temp.: 83.14 F
14:37:10.472 -> Average Cal sensor Temp.: 92.47 Ambient Temp.: 83.25 F
14:37:12.022 -> Average Cal sensor Temp.: 92.57 Ambient Temp.: 83.25 F
14:37:13.577 -> Average Cal sensor Temp.: 92.94 Ambient Temp.: 83.17 F
14:37:15.134 -> Average Cal sensor Temp.: 92.70 Ambient Temp.: 83.28 F
14:37:16.688 -> Average Cal sensor Temp.: 92.55 Ambient Temp.: 83.28 F
14:37:18.239 -> Average Cal sensor Temp.: 92.79 Ambient Temp.: 83.32 F
14:37:19.791 -> Average Cal sensor Temp.: 92.99 Ambient Temp.: 83.28 F
14:37:21.343 -> Average Cal sensor Temp.: 92.82 Ambient Temp.: 83.39 F
14:37:22.925 -> Average Cal sensor Temp.: 92.79 Ambient Temp.: 83.35 F
14:37:24.479 -> Average Cal sensor Temp.: 92.91 Ambient Temp.: 83.39 F
14:37:26.033 -> Average Cal sensor Temp.: 92.69 Ambient Temp.: 83.43 F
14:37:27.586 -> Average Cal sensor Temp.: 92.29 Ambient Temp.: 83.43 F
14:37:29.141 -> Average Cal sensor Temp.: 92.16 Ambient Temp.: 83.46 F
14:37:30.693 -> Average Cal sensor Temp.: 92.26 Ambient Temp.: 83.46 F
14:37:32.237 -> Average Cal sensor Temp.: 92.51 Ambient Temp.: 83.46 F
14:37:33.785 -> Average Cal sensor Temp.: 92.63 Ambient Temp.: 83.46 F
14:37:35.369 -> Average Cal sensor Temp.: 92.52 Ambient Temp.: 83.53 F
14:37:36.926 -> Average Cal sensor Temp.: 92.77 Ambient Temp.: 83.50 F
14:37:38.474 -> Average Cal sensor Temp.: 92.77 Ambient Temp.: 83.53 F
14:37:40.028 -> Average Cal sensor Temp.: 92.72 Ambient Temp.: 83.53 F
14:37:41.576 -> Average Cal sensor Temp.: 92.82 Ambient Temp.: 83.64 F
14:37:43.119 -> Average Cal sensor Temp.: 92.68 Ambient Temp.: 83.57 F
14:37:44.673 -> Average Cal sensor Temp.: 92.67 Ambient Temp.: 83.57 F
14:37:46.257 -> Average Cal sensor Temp.: 92.83 Ambient Temp.: 83.68 F
14:37:47.812 -> Average Cal sensor Temp.: 92.85 Ambient Temp.: 83.64 F
14:37:49.366 -> Average Cal sensor Temp.: 92.68 Ambient Temp.: 83.75 F
14:37:50.926 -> Average Cal sensor Temp.: 92.77 Ambient Temp.: 83.71 F
14:37:52.480 -> Average Cal sensor Temp.: 92.93 Ambient Temp.: 83.71 F
14:37:54.033 -> Average Cal sensor Temp.: 92.61 Ambient Temp.: 83.82 F
14:37:55.586 -> Average Cal sensor Temp.: 92.45 Ambient Temp.: 83.75 F
14:37:57.140 -> Average Cal sensor Temp.: 92.48 Ambient Temp.: 83.79 F
14:37:58.697 -> Average Cal sensor Temp.: 92.73 Ambient Temp.: 83.86 F
14:38:00.243 -> Average Cal sensor Temp.: 92.46 Ambient Temp.: 83.86 F
14:38:01.798 -> Average Cal sensor Temp.: 92.68 Ambient Temp.: 83.89 F
14:38:03.353 -> Average Cal sensor Temp.: 92.80 Ambient Temp.: 83.89 F
14:38:04.936 -> Average Cal sensor Temp.: 92.80 Ambient Temp.: 83.89 F
14:38:06.490 -> Average Cal sensor Temp.: 92.79 Ambient Temp.: 83.97 F
14:38:08.046 -> Average Cal sensor Temp.: 92.67 Ambient Temp.: 84.00 F
14:38:09.592 -> Average Cal sensor Temp.: 92.82 Ambient Temp.: 84.00 F
14:38:11.143 -> Average Cal sensor Temp.: 93.01 Ambient Temp.: 84.00 F
14:38:12.701 -> Average Cal sensor Temp.: 93.15 Ambient Temp.: 84.00 F
14:38:14.258 -> Average Cal sensor Temp.: 93.15 Ambient Temp.: 84.11 F
14:38:15.818 -> Average Cal sensor Temp.: 93.44 Ambient Temp.: 84.15 F
14:38:17.369 -> Average Cal sensor Temp.: 93.48 Ambient Temp.: 84.15 F
14:38:18.923 -> Average Cal sensor Temp.: 93.32 Ambient Temp.: 84.22 F
14:38:20.480 -> Average Cal sensor Temp.: 93.26 Ambient Temp.: 84.18 F
14:38:22.037 -> Average Cal sensor Temp.: 93.13 Ambient Temp.: 84.25 F
14:38:23.591 -> Average Cal sensor Temp.: 93.15 Ambient Temp.: 84.22 F
14:38:25.147 -> Average Cal sensor Temp.: 93.10 Ambient Temp.: 84.29 F
14:38:26.700 -> Average Cal sensor Temp.: 93.11 Ambient Temp.: 84.29 F
14:38:28.256 -> Average Cal sensor Temp.: 93.07 Ambient Temp.: 84.33 F
14:38:29.812 -> Average Cal sensor Temp.: 92.82 Ambient Temp.: 84.36 F
14:38:31.367 -> Average Cal sensor Temp.: 92.92 Ambient Temp.: 84.40 F
14:38:32.920 -> Average Cal sensor Temp.: 92.84 Ambient Temp.: 84.40 F
14:38:34.476 -> Average Cal sensor Temp.: 92.97 Ambient Temp.: 84.51 F
14:38:36.030 -> Average Cal sensor Temp.: 92.83 Ambient Temp.: 84.51 F
14:38:37.579 -> Average Cal sensor Temp.: 92.67 Ambient Temp.: 84.54 F
14:38:39.134 -> Average Cal sensor Temp.: 92.83 Ambient Temp.: 84.54 F
14:38:40.693 -> Average Cal sensor Temp.: 92.80 Ambient Temp.: 84.58 F
14:38:42.247 -> Average Cal sensor Temp.: 92.90 Ambient Temp.: 84.65 F
14:38:43.805 -> Average Cal sensor Temp.: 93.10 Ambient Temp.: 84.58 F
14:38:45.358 -> Average Cal sensor Temp.: 92.88 Ambient Temp.: 84.61 F
14:38:46.942 -> Average Cal sensor Temp.: 93.15 Ambient Temp.: 84.65 F
14:38:48.498 -> Average Cal sensor Temp.: 93.22 Ambient Temp.: 84.65 F
14:38:50.051 -> Average Cal sensor Temp.: 93.11 Ambient Temp.: 84.69 F
14:38:51.612 -> Average Cal sensor Temp.: 92.83 Ambient Temp.: 84.72 F
14:38:53.166 -> Average Cal sensor Temp.: 92.75 Ambient Temp.: 84.72 F
14:38:54.719 -> Average Cal sensor Temp.: 93.07 Ambient Temp.: 84.72 F
14:38:56.274 -> Average Cal sensor Temp.: 92.77 Ambient Temp.: 84.69 F
14:38:57.826 -> Average Cal sensor Temp.: 92.79 Ambient Temp.: 84.72 F
14:38:59.373 -> Average Cal sensor Temp.: 92.94 Ambient Temp.: 84.69 F
14:39:00.927 -> Average Cal sensor Temp.: 92.93 Ambient Temp.: 84.76 F
14:39:02.478 -> Average Cal sensor Temp.: 92.97 Ambient Temp.: 84.79 F
14:39:04.062 -> Average Cal sensor Temp.: 93.34 Ambient Temp.: 84.87 F
14:39:05.615 -> Average Cal sensor Temp.: 93.45 Ambient Temp.: 84.87 F
14:39:07.167 -> Average Cal sensor Temp.: 93.40 Ambient Temp.: 84.79 F
14:39:08.721 -> Average Cal sensor Temp.: 93.20 Ambient Temp.: 84.90 F
14:39:10.269 -> Average Cal sensor Temp.: 92.67 Ambient Temp.: 84.90 F
14:39:11.820 -> Average Cal sensor Temp.: 92.74 Ambient Temp.: 84.87 F
14:39:13.377 -> Average Cal sensor Temp.: 93.26 Ambient Temp.: 84.87 F
14:39:14.933 -> Average Cal sensor Temp.: 93.07 Ambient Temp.: 84.87 F
14:39:16.483 -> Average Cal sensor Temp.: 93.05 Ambient Temp.: 84.94 F
14:39:18.071 -> Average Cal sensor Temp.: 93.01 Ambient Temp.: 84.90 F
14:39:19.621 -> Average Cal sensor Temp.: 92.97 Ambient Temp.: 84.94 F
14:39:21.184 -> Average Cal sensor Temp.: 92.69 Ambient Temp.: 84.94 F
14:39:22.740 -> Average Cal sensor Temp.: 93.07 Ambient Temp.: 84.94 F
14:39:24.296 -> Average Cal sensor Temp.: 93.40 Ambient Temp.: 84.94 F
14:39:25.850 -> Average Cal sensor Temp.: 93.10 Ambient Temp.: 84.94 F
14:39:27.401 -> Average Cal sensor Temp.: 93.20 Ambient Temp.: 84.97 F
14:39:28.957 -> Average Cal sensor Temp.: 93.54 Ambient Temp.: 84.97 F
14:39:30.511 -> Average Cal sensor Temp.: 93.04 Ambient Temp.: 85.01 F
14:39:32.060 -> Average Cal sensor Temp.: 92.92 Ambient Temp.: 85.01 F
14:39:33.618 -> Average Cal sensor Temp.: 92.91 Ambient Temp.: 85.01 F
14:39:35.171 -> Average Cal sensor Temp.: 92.84 Ambient Temp.: 85.01 F
14:39:36.722 -> Average Cal sensor Temp.: 93.17 Ambient Temp.: 84.97 F
14:39:38.276 -> Average Cal sensor Temp.: 93.22 Ambient Temp.: 85.01 F
14:39:39.827 -> Average Cal sensor Temp.: 93.04 Ambient Temp.: 85.01 F
14:39:41.417 -> Average Cal sensor Temp.: 92.91 Ambient Temp.: 84.97 F
14:39:42.971 -> Average Cal sensor Temp.: 93.00 Ambient Temp.: 84.97 F
14:39:44.527 -> Average Cal sensor Temp.: 92.76 Ambient Temp.: 85.01 F
14:39:46.084 -> Average Cal sensor Temp.: 92.82 Ambient Temp.: 85.01 F
14:39:47.638 -> Average Cal sensor Temp.: 92.61 Ambient Temp.: 84.94 F
14:39:49.193 -> Average Cal sensor Temp.: 92.95 Ambient Temp.: 84.94 F
14:39:50.753 -> Average Cal sensor Temp.: 93.18 Ambient Temp.: 84.94 F
14:39:52.305 -> Average Cal sensor Temp.: 92.92 Ambient Temp.: 84.97 F
14:39:53.861 -> Average Cal sensor Temp.: 92.82 Ambient Temp.: 84.90 F
14:39:55.409 -> Average Cal sensor Temp.: 92.82 Ambient Temp.: 84.90 F
14:39:56.963 -> Average Cal sensor Temp.: 93.00 Ambient Temp.: 84.94 F
14:39:58.519 -> Average Cal sensor Temp.: 92.76 Ambient Temp.: 84.90 F
14:40:00.070 -> Average Cal sensor Temp.: 93.09 Ambient Temp.: 84.90 F
14:40:01.626 -> Average Cal sensor Temp.: 92.77 Ambient Temp.: 84.79 F
14:40:03.208 -> Average Cal sensor Temp.: 92.98 Ambient Temp.: 84.76 F
14:40:04.766 -> Average Cal sensor Temp.: 93.43 Ambient Temp.: 84.79 F
14:40:06.324 -> Average Cal sensor Temp.: 93.37 Ambient Temp.: 84.76 F
14:40:07.878 -> Average Cal sensor Temp.: 93.28 Ambient Temp.: 84.79 F
14:40:09.427 -> Average Cal sensor Temp.: 92.97 Ambient Temp.: 84.76 F
14:40:10.982 -> Average Cal sensor Temp.: 93.01 Ambient Temp.: 84.69 F
14:40:12.531 -> Average Cal sensor Temp.: 93.15 Ambient Temp.: 84.72 F
14:40:14.090 -> Average Cal sensor Temp.: 92.88 Ambient Temp.: 84.69 F
14:40:15.643 -> Average Cal sensor Temp.: 93.16 Ambient Temp.: 84.69 F
14:40:17.197 -> Average Cal sensor Temp.: 92.86 Ambient Temp.: 84.61 F
14:40:18.758 -> Average Cal sensor Temp.: 93.30 Ambient Temp.: 84.69 F
14:40:20.310 -> Average Cal sensor Temp.: 93.37 Ambient Temp.: 84.61 F
14:40:21.865 -> Average Cal sensor Temp.: 93.22 Ambient Temp.: 84.65 F
14:40:23.422 -> Average Cal sensor Temp.: 93.20 Ambient Temp.: 84.61 F
14:40:24.977 -> Average Cal sensor Temp.: 93.25 Ambient Temp.: 84.58 F
14:40:26.527 -> Average Cal sensor Temp.: 93.47 Ambient Temp.: 84.61 F
14:40:28.083 -> Average Cal sensor Temp.: 93.25 Ambient Temp.: 84.58 F
14:40:29.637 -> Average Cal sensor Temp.: 93.55 Ambient Temp.: 84.58 F
14:40:31.188 -> Average Cal sensor Temp.: 93.52 Ambient Temp.: 84.61 F
14:40:32.742 -> Average Cal sensor Temp.: 93.52 Ambient Temp.: 84.54 F
14:40:34.302 -> Average Cal sensor Temp.: 93.08 Ambient Temp.: 84.58 F
14:40:35.866 -> Average Cal sensor Temp.: 93.50 Ambient Temp.: 84.54 F
14:40:37.445 -> Average Cal sensor Temp.: 93.36 Ambient Temp.: 84.58 F
14:40:39.001 -> Average Cal sensor Temp.: 93.34 Ambient Temp.: 84.58 F
14:40:40.552 -> Average Cal sensor Temp.: 93.23 Ambient Temp.: 84.58 F
14:40:42.107 -> Average Cal sensor Temp.: 92.88 Ambient Temp.: 84.58 F
14:40:43.664 -> Average Cal sensor Temp.: 92.99 Ambient Temp.: 84.54 F
14:40:45.213 -> Average Cal sensor Temp.: 93.21 Ambient Temp.: 84.51 F
14:40:46.769 -> Average Cal sensor Temp.: 92.83 Ambient Temp.: 84.47 F
14:40:48.321 -> Average Cal sensor Temp.: 92.93 Ambient Temp.: 84.47 F
14:40:49.877 -> Average Cal sensor Temp.: 93.12 Ambient Temp.: 84.47 F
14:40:51.430 -> Average Cal sensor Temp.: 92.82 Ambient Temp.: 84.47 F
14:40:52.990 -> Average Cal sensor Temp.: 93.05 Ambient Temp.: 84.40 F
14:40:54.539 -> Average Cal sensor Temp.: 93.33 Ambient Temp.: 84.47 F
14:40:56.095 -> Average Cal sensor Temp.: 93.24 Ambient Temp.: 84.47 F
14:40:57.644 -> Average Cal sensor Temp.: 93.26 Ambient Temp.: 84.47 F
14:40:59.227 -> Average Cal sensor Temp.: 93.35 Ambient Temp.: 84.40 F
14:41:00.776 -> Average Cal sensor Temp.: 93.38 Ambient Temp.: 84.40 F
14:41:02.320 -> Average Cal sensor Temp.: 93.44 Ambient Temp.: 84.36 F
14:41:03.870 -> Average Cal sensor Temp.: 93.20 Ambient Temp.: 84.47 F
14:41:05.430 -> Average Cal sensor Temp.: 93.08 Ambient Temp.: 84.36 F
14:41:07.014 -> Average Cal sensor Temp.: 92.95 Ambient Temp.: 84.40 F
14:41:08.565 -> Average Cal sensor Temp.: 93.30 Ambient Temp.: 84.33 F
14:41:10.111 -> Average Cal sensor Temp.: 93.20 Ambient Temp.: 84.36 F
14:41:11.666 -> Average Cal sensor Temp.: 92.91 Ambient Temp.: 84.29 F
14:41:13.216 -> Average Cal sensor Temp.: 93.07 Ambient Temp.: 84.25 F
14:41:14.768 -> Average Cal sensor Temp.: 93.43 Ambient Temp.: 84.33 F
14:41:16.329 -> Average Cal sensor Temp.: 93.46 Ambient Temp.: 84.25 F
14:41:17.881 -> Average Cal sensor Temp.: 93.54 Ambient Temp.: 84.29 F
14:41:19.461 -> Average Cal sensor Temp.: 93.56 Ambient Temp.: 84.29 F
14:41:21.014 -> Average Cal sensor Temp.: 93.17 Ambient Temp.: 84.25 F
14:41:22.569 -> Average Cal sensor Temp.: 93.35 Ambient Temp.: 84.18 F
14:41:24.127 -> Average Cal sensor Temp.: 93.38 Ambient Temp.: 84.29 F
14:41:25.680 -> Average Cal sensor Temp.: 93.41 Ambient Temp.: 84.18 F
14:41:27.227 -> Average Cal sensor Temp.: 93.34 Ambient Temp.: 84.15 F
14:41:28.780 -> Average Cal sensor Temp.: 93.31 Ambient Temp.: 84.18 F
14:41:30.339 -> Average Cal sensor Temp.: 93.38 Ambient Temp.: 84.15 F
14:41:31.888 -> Average Cal sensor Temp.: 93.45 Ambient Temp.: 84.11 F
14:41:33.447 -> Average Cal sensor Temp.: 93.71 Ambient Temp.: 84.18 F
14:41:35.006 -> Average Cal sensor Temp.: 93.34 Ambient Temp.: 84.11 F
14:41:36.552 -> Average Cal sensor Temp.: 92.98 Ambient Temp.: 84.11 F
14:41:38.110 -> Average Cal sensor Temp.: 93.59 Ambient Temp.: 84.07 F
14:41:39.665 -> Average Cal sensor Temp.: 93.29 Ambient Temp.: 84.07 F
14:41:41.219 -> Average Cal sensor Temp.: 92.70 Ambient Temp.: 84.00 F
14:41:42.773 -> Average Cal sensor Temp.: 92.84 Ambient Temp.: 84.00 F
14:41:44.332 -> Average Cal sensor Temp.: 93.53 Ambient Temp.: 83.97 F
14:41:45.886 -> Average Cal sensor Temp.: 93.39 Ambient Temp.: 83.97 F
14:41:47.439 -> Average Cal sensor Temp.: 92.85 Ambient Temp.: 83.93 F
14:41:48.995 -> Average Cal sensor Temp.: 93.39 Ambient Temp.: 83.86 F
14:41:50.553 -> Average Cal sensor Temp.: 93.43 Ambient Temp.: 83.82 F
14:41:52.112 -> Average Cal sensor Temp.: 93.59 Ambient Temp.: 83.82 F
14:41:53.668 -> Average Cal sensor Temp.: 93.57 Ambient Temp.: 83.86 F
14:41:55.216 -> Average Cal sensor Temp.: 93.78 Ambient Temp.: 83.79 F
14:41:56.803 -> Average Cal sensor Temp.: 93.62 Ambient Temp.: 83.75 F
14:41:58.325 -> Average Cal sensor Temp.: 93.38 Ambient Temp.: 83.79 F
14:41:59.912 -> Average Cal sensor Temp.: 93.41 Ambient Temp.: 83.75 F
14:42:01.460 -> Average Cal sensor Temp.: 93.60 Ambient Temp.: 83.79 F
14:42:03.015 -> Average Cal sensor Temp.: 93.51 Ambient Temp.: 83.75 F
14:42:04.568 -> Average Cal sensor Temp.: 93.30 Ambient Temp.: 83.71 F
14:42:06.124 -> Average Cal sensor Temp.: 93.42 Ambient Temp.: 83.71 F
14:42:07.675 -> Average Cal sensor Temp.: 93.29 Ambient Temp.: 83.71 F
14:42:09.216 -> Average Cal sensor Temp.: 93.25 Ambient Temp.: 83.68 F
14:42:10.766 -> Average Cal sensor Temp.: 93.08 Ambient Temp.: 83.68 F
14:42:12.349 -> Average Cal sensor Temp.: 93.45 Ambient Temp.: 83.68 F
14:42:13.899 -> Average Cal sensor Temp.: 93.18 Ambient Temp.: 83.57 F
14:42:15.453 -> Average Cal sensor Temp.: 93.53 Ambient Temp.: 83.53 F
14:42:16.997 -> Average Cal sensor Temp.: 93.21 Ambient Temp.: 83.57 F
14:42:18.544 -> Average Cal sensor Temp.: 93.24 Ambient Temp.: 83.57 F
14:42:20.131 -> Average Cal sensor Temp.: 93.35 Ambient Temp.: 83.57 F
14:42:21.684 -> Average Cal sensor Temp.: 93.57 Ambient Temp.: 83.50 F
14:42:23.226 -> Average Cal sensor Temp.: 93.58 Ambient Temp.: 83.46 F
14:42:24.778 -> Average Cal sensor Temp.: 93.30 Ambient Temp.: 83.53 F
14:42:26.332 -> Average Cal sensor Temp.: 93.23 Ambient Temp.: 83.43 F
14:42:27.882 -> Average Cal sensor Temp.: 93.20 Ambient Temp.: 83.46 F
14:42:29.437 -> Average Cal sensor Temp.: 93.61 Ambient Temp.: 83.46 F
14:42:30.988 -> Average Cal sensor Temp.: 93.58 Ambient Temp.: 83.46 F
14:42:32.572 -> Average Cal sensor Temp.: 93.73 Ambient Temp.: 83.43 F
14:42:34.126 -> Average Cal sensor Temp.: 93.67 Ambient Temp.: 83.43 F
14:42:35.681 -> Average Cal sensor Temp.: 93.47 Ambient Temp.: 83.43 F
14:42:37.235 -> Average Cal sensor Temp.: 93.33 Ambient Temp.: 83.39 F
14:42:38.794 -> Average Cal sensor Temp.: 93.29 Ambient Temp.: 83.35 F
14:42:40.350 -> Average Cal sensor Temp.: 93.26 Ambient Temp.: 83.39 F
14:42:41.897 -> Average Cal sensor Temp.: 93.60 Ambient Temp.: 83.32 F
14:42:43.445 -> Average Cal sensor Temp.: 93.30 Ambient Temp.: 83.35 F
14:42:45.004 -> Average Cal sensor Temp.: 93.53 Ambient Temp.: 83.32 F
14:42:46.558 -> Average Cal sensor Temp.: 93.50 Ambient Temp.: 83.28 F
14:42:48.104 -> Average Cal sensor Temp.: 93.39 Ambient Temp.: 83.32 F
14:42:49.660 -> Average Cal sensor Temp.: 93.41 Ambient Temp.: 83.32 F
14:42:51.216 -> Average Cal sensor Temp.: 93.49 Ambient Temp.: 83.32 F
14:42:52.771 -> Average Cal sensor Temp.: 93.03 Ambient Temp.: 83.28 F
14:42:54.328 -> Average Cal sensor Temp.: 93.34 Ambient Temp.: 83.25 F
14:42:55.885 -> Average Cal sensor Temp.: 93.37 Ambient Temp.: 83.25 F
14:42:57.435 -> Average Cal sensor Temp.: 93.52 Ambient Temp.: 83.17 F
14:42:58.991 -> Average Cal sensor Temp.: 93.29 Ambient Temp.: 83.25 F
14:43:00.541 -> Average Cal sensor Temp.: 92.97 Ambient Temp.: 83.14 F
14:43:02.125 -> Average Cal sensor Temp.: 92.95 Ambient Temp.: 83.14 F
14:43:03.675 -> Average Cal sensor Temp.: 93.52 Ambient Temp.: 83.14 F
14:43:05.229 -> Average Cal sensor Temp.: 93.07 Ambient Temp.: 83.07 F
14:43:06.782 -> Average Cal sensor Temp.: 93.19 Ambient Temp.: 83.07 F
14:43:08.341 -> Average Cal sensor Temp.: 93.31 Ambient Temp.: 83.07 F
14:43:09.893 -> Average Cal sensor Temp.: 93.52 Ambient Temp.: 82.99 F
14:43:11.439 -> Average Cal sensor Temp.: 93.75 Ambient Temp.: 82.99 F
14:43:12.991 -> Average Cal sensor Temp.: 93.71 Ambient Temp.: 83.03 F
14:43:14.537 -> Average Cal sensor Temp.: 93.18 Ambient Temp.: 83.03 F
14:43:16.124 -> Average Cal sensor Temp.: 93.15 Ambient Temp.: 82.99 F
14:43:17.647 -> Average Cal sensor Temp.: 93.59 Ambient Temp.: 82.96 F
14:43:19.233 -> Average Cal sensor Temp.: 93.47 Ambient Temp.: 82.96 F
14:43:20.784 -> Average Cal sensor Temp.: 93.69 Ambient Temp.: 82.96 F
14:43:22.338 -> Average Cal sensor Temp.: 93.94 Ambient Temp.: 82.96 F
14:43:23.899 -> Average Cal sensor Temp.: 93.66 Ambient Temp.: 82.92 F
14:43:25.457 -> Average Cal sensor Temp.: 93.69 Ambient Temp.: 82.92 F
14:43:27.012 -> Average Cal sensor Temp.: 93.18 Ambient Temp.: 82.96 F
14:43:28.564 -> Average Cal sensor Temp.: 93.16 Ambient Temp.: 82.92 F
14:43:29.157 -> Ambient Temp.: 82.92 F
14:43:30.711 -> Average Cal sensor Temp.: 93.22 Ambient Temp.: 82.92 F
14:43:32.267 -> Average Cal sensor Temp.: 93.65 Ambient Temp.: 82.92 F
14:43:33.821 -> Average Cal sensor Temp.: 93.81 Ambient Temp.: 82.96 F
14:43:35.364 -> Average Cal sensor Temp.: 93.98 Ambient Temp.: 82.96 F
14:43:36.919 -> Average Cal sensor Temp.: 93.79 Ambient Temp.: 82.89 F
14:43:38.476 -> Average Cal sensor Temp.: 93.86 Ambient Temp.: 82.92 F
14:43:40.031 -> Average Cal sensor Temp.: 93.93 Ambient Temp.: 82.99 F
14:43:41.579 -> Average Cal sensor Temp.: 93.86 Ambient Temp.: 82.92 F
14:43:43.166 -> Average Cal sensor Temp.: 93.48 Ambient Temp.: 82.99 F
14:43:44.724 -> Average Cal sensor Temp.: 93.95 Ambient Temp.: 83.03 F
14:43:46.280 -> Average Cal sensor Temp.: 93.95 Ambient Temp.: 82.99 F
14:43:47.832 -> Average Cal sensor Temp.: 93.94 Ambient Temp.: 82.99 F
14:43:49.381 -> Average Cal sensor Temp.: 94.16 Ambient Temp.: 83.03 F
14:43:50.937 -> Average Cal sensor Temp.: 93.81 Ambient Temp.: 83.07 F
14:43:52.495 -> Average Cal sensor Temp.: 93.53 Ambient Temp.: 83.07 F
14:43:54.051 -> Average Cal sensor Temp.: 93.42 Ambient Temp.: 83.10 F
14:43:55.608 -> Average Cal sensor Temp.: 93.72 Ambient Temp.: 83.03 F
14:43:57.163 -> Average Cal sensor Temp.: 93.80 Ambient Temp.: 83.07 F
14:43:58.720 -> Average Cal sensor Temp.: 93.67 Ambient Temp.: 83.10 F
14:44:00.272 -> Average Cal sensor Temp.: 93.83 Ambient Temp.: 83.14 F
14:44:01.822 -> Average Cal sensor Temp.: 93.97 Ambient Temp.: 83.10 F
14:44:03.368 -> Average Cal sensor Temp.: 93.83 Ambient Temp.: 83.10 F
14:44:04.924 -> Average Cal sensor Temp.: 93.54 Ambient Temp.: 83.10 F
14:44:06.482 -> Average Cal sensor Temp.: 93.57 Ambient Temp.: 83.10 F
14:44:08.037 -> Average Cal sensor Temp.: 93.70 Ambient Temp.: 83.14 F
14:44:09.588 -> Average Cal sensor Temp.: 93.89 Ambient Temp.: 83.25 F
14:44:11.145 -> Average Cal sensor Temp.: 93.52 Ambient Temp.: 83.28 F
14:44:12.694 -> Average Cal sensor Temp.: 93.46 Ambient Temp.: 83.25 F
14:44:14.282 -> Average Cal sensor Temp.: 93.74 Ambient Temp.: 83.25 F
14:44:15.839 -> Average Cal sensor Temp.: 93.93 Ambient Temp.: 83.25 F
14:44:17.388 -> Average Cal sensor Temp.: 93.97 Ambient Temp.: 83.28 F
14:44:18.943 -> Average Cal sensor Temp.: 93.97 Ambient Temp.: 83.28 F
14:44:20.496 -> Average Cal sensor Temp.: 93.63 Ambient Temp.: 83.32 F
14:44:22.052 -> Average Cal sensor Temp.: 93.79 Ambient Temp.: 83.28 F
14:44:23.602 -> Average Cal sensor Temp.: 93.85 Ambient Temp.: 83.32 F
14:44:25.156 -> Average Cal sensor Temp.: 93.80 Ambient Temp.: 83.39 F
14:44:26.711 -> Average Cal sensor Temp.: 94.09 Ambient Temp.: 83.39 F
14:44:28.265 -> Average Cal sensor Temp.: 93.99 Ambient Temp.: 83.35 F
14:44:29.822 -> Average Cal sensor Temp.: 94.10 Ambient Temp.: 83.39 F
14:44:31.370 -> Average Cal sensor Temp.: 93.74 Ambient Temp.: 83.43 F
14:44:32.922 -> Average Cal sensor Temp.: 93.69 Ambient Temp.: 83.46 F
14:44:34.505 -> Average Cal sensor Temp.: 93.47 Ambient Temp.: 83.43 F
14:44:36.061 -> Average Cal sensor Temp.: 93.61 Ambient Temp.: 83.46 F
14:44:37.615 -> Average Cal sensor Temp.: 93.57 Ambient Temp.: 83.43 F
14:44:39.166 -> Average Cal sensor Temp.: 93.90 Ambient Temp.: 83.46 F
14:44:40.722 -> Average Cal sensor Temp.: 94.10 Ambient Temp.: 83.50 F
14:44:42.272 -> Average Cal sensor Temp.: 94.01 Ambient Temp.: 83.50 F
14:44:43.829 -> Average Cal sensor Temp.: 93.68 Ambient Temp.: 83.53 F
14:44:45.385 -> Average Cal sensor Temp.: 94.07 Ambient Temp.: 83.53 F
14:44:46.944 -> Average Cal sensor Temp.: 94.22 Ambient Temp.: 83.57 F
14:44:48.496 -> Average Cal sensor Temp.: 93.79 Ambient Temp.: 83.57 F
14:44:50.052 -> Average Cal sensor Temp.: 93.60 Ambient Temp.: 83.57 F
14:44:51.609 -> Average Cal sensor Temp.: 93.43 Ambient Temp.: 83.68 F
14:44:53.162 -> Average Cal sensor Temp.: 93.57 Ambient Temp.: 83.68 F
14:44:54.714 -> Average Cal sensor Temp.: 93.88 Ambient Temp.: 83.68 F
14:44:56.269 -> Average Cal sensor Temp.: 93.74 Ambient Temp.: 83.68 F
14:44:57.826 -> Average Cal sensor Temp.: 93.53 Ambient Temp.: 83.71 F
14:44:59.378 -> Average Cal sensor Temp.: 93.58 Ambient Temp.: 83.71 F
14:45:00.930 -> Average Cal sensor Temp.: 93.58 Ambient Temp.: 83.68 F
14:45:02.514 -> Average Cal sensor Temp.: 93.61 Ambient Temp.: 83.71 F
14:45:04.063 -> Average Cal sensor Temp.: 93.52 Ambient Temp.: 83.79 F
14:45:05.620 -> Average Cal sensor Temp.: 93.33 Ambient Temp.: 83.75 F
14:45:07.176 -> Average Cal sensor Temp.: 93.38 Ambient Temp.: 83.75 F
14:45:08.730 -> Average Cal sensor Temp.: 93.56 Ambient Temp.: 83.86 F
14:45:10.262 -> Average Cal sensor Temp.: 93.58 Ambient Temp.: 83.79 F
14:45:11.824 -> Average Cal sensor Temp.: 93.71 Ambient Temp.: 83.79 F
14:45:13.377 -> Average Cal sensor Temp.: 93.69 Ambient Temp.: 83.79 F
14:45:14.927 -> Average Cal sensor Temp.: 93.71 Ambient Temp.: 83.86 F
14:45:16.516 -> Average Cal sensor Temp.: 93.85 Ambient Temp.: 83.86 F
14:45:18.065 -> Average Cal sensor Temp.: 93.83 Ambient Temp.: 83.89 F
14:45:19.618 -> Average Cal sensor Temp.: 94.14 Ambient Temp.: 83.89 F
14:45:21.179 -> Average Cal sensor Temp.: 93.74 Ambient Temp.: 83.89 F
14:45:22.727 -> Average Cal sensor Temp.: 93.72 Ambient Temp.: 83.93 F
14:45:24.279 -> Average Cal sensor Temp.: 93.74 Ambient Temp.: 83.89 F
14:45:25.833 -> Average Cal sensor Temp.: 94.08 Ambient Temp.: 83.93 F
14:45:27.384 -> Average Cal sensor Temp.: 93.76 Ambient Temp.: 83.93 F
14:45:28.939 -> Average Cal sensor Temp.: 93.80 Ambient Temp.: 83.97 F
14:45:30.492 -> Average Cal sensor Temp.: 93.73 Ambient Temp.: 84.00 F
14:45:32.079 -> Average Cal sensor Temp.: 93.88 Ambient Temp.: 84.00 F
14:45:33.630 -> Average Cal sensor Temp.: 93.82 Ambient Temp.: 84.11 F
14:45:35.180 -> Average Cal sensor Temp.: 93.79 Ambient Temp.: 84.07 F
14:45:36.734 -> Average Cal sensor Temp.: 93.88 Ambient Temp.: 84.15 F
14:45:38.284 -> Average Cal sensor Temp.: 94.08 Ambient Temp.: 84.15 F
14:45:39.837 -> Average Cal sensor Temp.: 93.82 Ambient Temp.: 84.18 F
14:45:41.388 -> Average Cal sensor Temp.: 93.65 Ambient Temp.: 84.18 F
14:45:42.975 -> Average Cal sensor Temp.: 93.92 Ambient Temp.: 84.18 F
14:45:44.531 -> Average Cal sensor Temp.: 93.91 Ambient Temp.: 84.22 F
14:45:46.087 -> Average Cal sensor Temp.: 93.96 Ambient Temp.: 84.18 F
14:45:47.644 -> Average Cal sensor Temp.: 93.79 Ambient Temp.: 84.25 F
14:45:49.196 -> Average Cal sensor Temp.: 93.80 Ambient Temp.: 84.33 F
14:45:50.754 -> Average Cal sensor Temp.: 93.71 Ambient Temp.: 84.29 F
14:45:52.304 -> Average Cal sensor Temp.: 93.69 Ambient Temp.: 84.33 F
14:45:53.856 -> Average Cal sensor Temp.: 93.87 Ambient Temp.: 84.40 F
14:45:55.406 -> Average Cal sensor Temp.: 93.59 Ambient Temp.: 84.40 F
14:45:56.961 -> Average Cal sensor Temp.: 93.49 Ambient Temp.: 84.40 F
14:45:58.511 -> Average Cal sensor Temp.: 93.73 Ambient Temp.: 84.47 F
14:46:00.064 -> Average Cal sensor Temp.: 93.98 Ambient Temp.: 84.47 F
14:46:01.623 -> Average Cal sensor Temp.: 93.81 Ambient Temp.: 84.51 F
14:46:03.177 -> Average Cal sensor Temp.: 93.47 Ambient Temp.: 84.54 F
14:46:04.762 -> Average Cal sensor Temp.: 93.87 Ambient Temp.: 84.51 F
14:46:06.286 -> Average Cal sensor Temp.: 93.62 Ambient Temp.: 84.58 F
14:46:07.876 -> Average Cal sensor Temp.: 93.46 Ambient Temp.: 84.61 F
14:46:09.402 -> Average Cal sensor Temp.: 93.73 Ambient Temp.: 84.54 F
14:46:10.959 -> Average Cal sensor Temp.: 93.30 Ambient Temp.: 84.58 F
14:46:12.544 -> Average Cal sensor Temp.: 93.15 Ambient Temp.: 84.61 F
14:46:13.096 -> Ambient Temp.: 84.61 F
14:46:14.543 -> Ambient Temp.: 84.58 F
14:46:15.129 -> Ambient Temp.: 84.58 F
14:46:15.716 -> Ambient Temp.: 84.58 F
14:46:16.269 -> Ambient Temp.: 84.58 F
14:46:17.469 -> Ambient Temp.: 84.69 F
14:46:18.056 -> Ambient Temp.: 84.69 F
14:46:19.614 -> Average Cal sensor Temp.: 94.10 Ambient Temp.: 84.69 F
14:46:21.173 -> Average Cal sensor Temp.: 94.14 Ambient Temp.: 84.76 F
14:46:22.721 -> Average Cal sensor Temp.: 93.88 Ambient Temp.: 84.72 F
14:46:23.685 -> Ambient Temp.: 84.72 F
14:46:24.237 -> Ambient Temp.: 84.72 F
14:46:24.960 -> Ambient Temp.: 84.72 F
14:46:26.513 -> Average Cal sensor Temp.: 93.83 Ambient Temp.: 84.69 F
14:46:28.070 -> Average Cal sensor Temp.: 94.14 Ambient Temp.: 84.79 F
14:46:29.624 -> Average Cal sensor Temp.: 93.66 Ambient Temp.: 84.79 F
14:46:31.185 -> Average Cal sensor Temp.: 93.62 Ambient Temp.: 84.87 F
14:46:32.734 -> Average Cal sensor Temp.: 93.43 Ambient Temp.: 84.79 F
14:46:34.291 -> Average Cal sensor Temp.: 93.39 Ambient Temp.: 84.87 F
14:46:35.852 -> Average Cal sensor Temp.: 93.57 Ambient Temp.: 84.87 F
14:46:37.401 -> Average Cal sensor Temp.: 93.39 Ambient Temp.: 84.90 F
14:46:38.952 -> Average Cal sensor Temp.: 93.57 Ambient Temp.: 84.90 F
14:46:40.510 -> Average Cal sensor Temp.: 93.50 Ambient Temp.: 84.87 F
14:46:42.065 -> Average Cal sensor Temp.: 93.57 Ambient Temp.: 84.94 F
14:46:43.620 -> Average Cal sensor Temp.: 93.52 Ambient Temp.: 84.94 F
14:46:45.176 -> Average Cal sensor Temp.: 93.55 Ambient Temp.: 84.94 F
14:46:46.731 -> Average Cal sensor Temp.: 93.34 Ambient Temp.: 84.90 F
14:46:48.284 -> Average Cal sensor Temp.: 93.35 Ambient Temp.: 84.97 F
14:46:49.832 -> Average Cal sensor Temp.: 93.63 Ambient Temp.: 84.94 F
14:46:51.389 -> Average Cal sensor Temp.: 93.52 Ambient Temp.: 84.87 F
14:46:52.944 -> Average Cal sensor Temp.: 93.72 Ambient Temp.: 84.94 F
14:46:54.534 -> Average Cal sensor Temp.: 93.78 Ambient Temp.: 84.94 F
14:46:56.058 -> Average Cal sensor Temp.: 93.92 Ambient Temp.: 84.94 F
14:46:57.643 -> Average Cal sensor Temp.: 93.80 Ambient Temp.: 84.97 F
14:46:59.198 -> Average Cal sensor Temp.: 93.69 Ambient Temp.: 85.01 F
14:47:00.753 -> Average Cal sensor Temp.: 93.82 Ambient Temp.: 85.05 F
14:47:02.299 -> Average Cal sensor Temp.: 93.72 Ambient Temp.: 85.01 F
14:47:03.849 -> Average Cal sensor Temp.: 93.45 Ambient Temp.: 84.97 F
14:47:05.404 -> Average Cal sensor Temp.: 93.38 Ambient Temp.: 84.97 F
14:47:06.955 -> Average Cal sensor Temp.: 93.69 Ambient Temp.: 84.97 F
14:47:08.515 -> Average Cal sensor Temp.: 93.54 Ambient Temp.: 84.97 F
14:47:10.074 -> Average Cal sensor Temp.: 93.82 Ambient Temp.: 85.01 F
14:47:11.631 -> Average Cal sensor Temp.: 93.63 Ambient Temp.: 85.05 F
14:47:13.183 -> Average Cal sensor Temp.: 93.65 Ambient Temp.: 84.94 F
14:47:14.736 -> Average Cal sensor Temp.: 93.93 Ambient Temp.: 85.01 F
14:47:16.296 -> Average Cal sensor Temp.: 93.99 Ambient Temp.: 85.01 F
14:47:17.845 -> Average Cal sensor Temp.: 93.50 Ambient Temp.: 84.97 F
14:47:19.431 -> Average Cal sensor Temp.: 93.77 Ambient Temp.: 84.94 F
14:47:20.987 -> Average Cal sensor Temp.: 93.59 Ambient Temp.: 84.90 F
14:47:22.547 -> Average Cal sensor Temp.: 93.70 Ambient Temp.: 84.90 F
14:47:24.102 -> Average Cal sensor Temp.: 93.98 Ambient Temp.: 84.90 F
14:47:25.654 -> Average Cal sensor Temp.: 93.94 Ambient Temp.: 84.97 F
14:47:27.207 -> Average Cal sensor Temp.: 93.73 Ambient Temp.: 84.94 F
14:47:28.758 -> Average Cal sensor Temp.: 93.73 Ambient Temp.: 84.94 F
14:47:30.302 -> Average Cal sensor Temp.: 93.91 Ambient Temp.: 84.94 F
14:47:31.873 -> Average Cal sensor Temp.: 93.97 Ambient Temp.: 84.97 F
14:47:33.441 -> Average Cal sensor Temp.: 93.95 Ambient Temp.: 84.94 F
14:47:34.996 -> Average Cal sensor Temp.: 93.88 Ambient Temp.: 84.94 F
14:47:36.548 -> Average Cal sensor Temp.: 93.63 Ambient Temp.: 84.97 F
14:47:38.105 -> Average Cal sensor Temp.: 93.72 Ambient Temp.: 84.94 F
14:47:39.660 -> Average Cal sensor Temp.: 93.56 Ambient Temp.: 84.90 F
14:47:41.212 -> Average Cal sensor Temp.: 93.22 Ambient Temp.: 84.87 F
14:47:42.760 -> Average Cal sensor Temp.: 93.16 Ambient Temp.: 84.87 F
14:47:44.316 -> Average Cal sensor Temp.: 93.10 Ambient Temp.: 84.79 F
14:47:45.868 -> Average Cal sensor Temp.: 93.14 Ambient Temp.: 84.76 F
14:47:47.420 -> Average Cal sensor Temp.: 92.92 Ambient Temp.: 84.79 F
14:47:49.003 -> Average Cal sensor Temp.: 92.93 Ambient Temp.: 84.76 F
14:47:50.559 -> Average Cal sensor Temp.: 92.64 Ambient Temp.: 84.72 F
14:47:52.110 -> Average Cal sensor Temp.: 93.22 Ambient Temp.: 84.76 F
14:47:53.665 -> Average Cal sensor Temp.: 93.43 Ambient Temp.: 84.76 F
14:47:55.219 -> Average Cal sensor Temp.: 93.10 Ambient Temp.: 84.72 F
14:47:56.775 -> Average Cal sensor Temp.: 92.99 Ambient Temp.: 84.69 F
14:47:58.323 -> Average Cal sensor Temp.: 93.06 Ambient Temp.: 84.69 F
14:47:59.902 -> Average Cal sensor Temp.: 92.95 Ambient Temp.: 84.69 F
14:48:01.454 -> Average Cal sensor Temp.: 93.12 Ambient Temp.: 84.69 F
14:48:03.006 -> Average Cal sensor Temp.: 93.33 Ambient Temp.: 84.69 F
14:48:04.560 -> Average Cal sensor Temp.: 93.47 Ambient Temp.: 84.69 F
14:48:06.111 -> Average Cal sensor Temp.: 93.25 Ambient Temp.: 84.69 F
14:48:07.660 -> Average Cal sensor Temp.: 92.99 Ambient Temp.: 84.65 F
14:48:09.212 -> Average Cal sensor Temp.: 93.40 Ambient Temp.: 84.61 F
14:48:10.793 -> Average Cal sensor Temp.: 93.27 Ambient Temp.: 84.65 F
14:48:12.345 -> Average Cal sensor Temp.: 93.41 Ambient Temp.: 84.61 F
14:48:13.895 -> Average Cal sensor Temp.: 93.61 Ambient Temp.: 84.58 F
14:48:15.452 -> Average Cal sensor Temp.: 93.43 Ambient Temp.: 84.58 F
14:48:17.000 -> Average Cal sensor Temp.: 93.28 Ambient Temp.: 84.54 F
14:48:18.553 -> Average Cal sensor Temp.: 93.23 Ambient Temp.: 84.58 F
14:48:20.133 -> Average Cal sensor Temp.: 93.79 Ambient Temp.: 84.51 F
14:48:21.683 -> Average Cal sensor Temp.: 93.52 Ambient Temp.: 84.54 F
14:48:23.236 -> Average Cal sensor Temp.: 93.48 Ambient Temp.: 84.47 F
14:48:24.784 -> Average Cal sensor Temp.: 93.54 Ambient Temp.: 84.51 F
14:48:26.345 -> Average Cal sensor Temp.: 93.35 Ambient Temp.: 84.54 F
14:48:27.898 -> Average Cal sensor Temp.: 93.35 Ambient Temp.: 84.40 F
14:48:29.452 -> Average Cal sensor Temp.: 93.74 Ambient Temp.: 84.40 F
14:48:31.002 -> Average Cal sensor Temp.: 93.31 Ambient Temp.: 84.47 F
14:48:32.548 -> Average Cal sensor Temp.: 93.21 Ambient Temp.: 84.47 F
14:48:34.134 -> Average Cal sensor Temp.: 93.07 Ambient Temp.: 84.40 F
14:48:35.687 -> Average Cal sensor Temp.: 93.45 Ambient Temp.: 84.40 F
14:48:37.239 -> Average Cal sensor Temp.: 93.17 Ambient Temp.: 84.36 F
14:48:38.793 -> Average Cal sensor Temp.: 92.77 Ambient Temp.: 84.36 F
14:48:40.346 -> Average Cal sensor Temp.: 93.06 Ambient Temp.: 84.29 F
14:48:41.901 -> Average Cal sensor Temp.: 93.22 Ambient Temp.: 84.33 F
14:48:43.451 -> Average Cal sensor Temp.: 93.25 Ambient Temp.: 84.29 F
14:48:45.002 -> Average Cal sensor Temp.: 93.23 Ambient Temp.: 84.29 F
14:48:46.555 -> Average Cal sensor Temp.: 93.07 Ambient Temp.: 84.25 F
14:48:48.111 -> Average Cal sensor Temp.: 93.24 Ambient Temp.: 84.22 F
14:48:49.666 -> Average Cal sensor Temp.: 93.36 Ambient Temp.: 84.22 F
14:48:51.219 -> Average Cal sensor Temp.: 93.07 Ambient Temp.: 84.22 F
14:48:52.772 -> Average Cal sensor Temp.: 92.64 Ambient Temp.: 84.18 F
14:48:54.357 -> Average Cal sensor Temp.: 93.00 Ambient Temp.: 84.18 F
14:48:55.919 -> Average Cal sensor Temp.: 93.08 Ambient Temp.: 84.11 F
14:48:57.472 -> Average Cal sensor Temp.: 92.94 Ambient Temp.: 84.15 F
14:48:59.024 -> Average Cal sensor Temp.: 93.36 Ambient Temp.: 84.11 F
14:49:00.574 -> Average Cal sensor Temp.: 93.25 Ambient Temp.: 84.11 F
14:49:02.134 -> Average Cal sensor Temp.: 92.92 Ambient Temp.: 84.11 F
14:49:03.691 -> Average Cal sensor Temp.: 92.47 Ambient Temp.: 84.07 F
14:49:05.248 -> Average Cal sensor Temp.: 92.96 Ambient Temp.: 84.11 F
14:49:06.805 -> Average Cal sensor Temp.: 92.85 Ambient Temp.: 84.00 F
14:49:08.331 -> Average Cal sensor Temp.: 93.33 Ambient Temp.: 84.00 F
14:49:09.894 -> Average Cal sensor Temp.: 93.61 Ambient Temp.: 84.07 F
14:49:11.450 -> Average Cal sensor Temp.: 93.11 Ambient Temp.: 83.97 F
14:49:13.005 -> Average Cal sensor Temp.: 93.36 Ambient Temp.: 83.93 F
14:49:14.557 -> Average Cal sensor Temp.: 93.46 Ambient Temp.: 83.93 F
14:49:16.143 -> Average Cal sensor Temp.: 93.52 Ambient Temp.: 83.93 F
14:49:17.698 -> Average Cal sensor Temp.: 93.66 Ambient Temp.: 83.86 F
14:49:19.250 -> Average Cal sensor Temp.: 93.67 Ambient Temp.: 83.89 F
14:49:20.804 -> Average Cal sensor Temp.: 93.70 Ambient Temp.: 83.82 F
14:49:22.359 -> Average Cal sensor Temp.: 93.46 Ambient Temp.: 83.82 F
14:49:23.909 -> Average Cal sensor Temp.: 93.43 Ambient Temp.: 83.86 F
14:49:25.467 -> Average Cal sensor Temp.: 93.29 Ambient Temp.: 83.79 F
14:49:27.023 -> Average Cal sensor Temp.: 93.53 Ambient Temp.: 83.79 F
14:49:28.583 -> Average Cal sensor Temp.: 93.46 Ambient Temp.: 83.79 F
14:49:30.141 -> Average Cal sensor Temp.: 92.72 Ambient Temp.: 83.71 F
14:49:31.695 -> Average Cal sensor Temp.: 92.91 Ambient Temp.: 83.71 F
14:49:33.254 -> Average Cal sensor Temp.: 93.37 Ambient Temp.: 83.75 F
14:49:34.811 -> Average Cal sensor Temp.: 92.97 Ambient Temp.: 83.71 F
14:49:36.366 -> Average Cal sensor Temp.: 93.02 Ambient Temp.: 83.71 F
14:49:37.889 -> Average Cal sensor Temp.: 93.09 Ambient Temp.: 83.71 F
14:49:39.476 -> Average Cal sensor Temp.: 93.11 Ambient Temp.: 83.68 F
14:49:40.305 -> Ambient Temp.: 83.68 F
14:49:40.857 -> Ambient Temp.: 83.68 F
14:49:42.311 -> Ambient Temp.: 83.57 F
14:49:42.864 -> Ambient Temp.: 83.57 F
14:49:43.588 -> Ambient Temp.: 83.68 F
14:49:44.178 -> Ambient Temp.: 83.68 F
14:49:45.700 -> Average Cal sensor Temp.: 93.38 Ambient Temp.: 83.64 F
14:49:47.256 -> Average Cal sensor Temp.: 93.47 Ambient Temp.: 83.57 F
14:49:48.811 -> Average Cal sensor Temp.: 93.43 Ambient Temp.: 83.57 F
14:49:50.365 -> Average Cal sensor Temp.: 93.38 Ambient Temp.: 83.53 F
14:49:51.919 -> Average Cal sensor Temp.: 93.26 Ambient Temp.: 83.50 F
14:49:53.507 -> Average Cal sensor Temp.: 92.88 Ambient Temp.: 83.50 F
14:49:55.058 -> Average Cal sensor Temp.: 92.69 Ambient Temp.: 83.46 F
14:49:56.613 -> Average Cal sensor Temp.: 92.97 Ambient Temp.: 83.46 F
14:49:58.157 -> Average Cal sensor Temp.: 93.25 Ambient Temp.: 83.39 F
14:49:59.709 -> Average Cal sensor Temp.: 93.47 Ambient Temp.: 83.39 F
14:50:01.257 -> Average Cal sensor Temp.: 93.99 Ambient Temp.: 83.32 F
14:50:02.844 -> Average Cal sensor Temp.: 93.93 Ambient Temp.: 83.39 F
14:50:04.390 -> Average Cal sensor Temp.: 93.96 Ambient Temp.: 83.43 F
14:50:05.945 -> Average Cal sensor Temp.: 93.91 Ambient Temp.: 83.39 F
14:50:07.502 -> Average Cal sensor Temp.: 93.56 Ambient Temp.: 83.39 F
14:50:09.062 -> Average Cal sensor Temp.: 93.09 Ambient Temp.: 83.35 F
14:50:10.615 -> Average Cal sensor Temp.: 93.20 Ambient Temp.: 83.35 F
14:50:12.167 -> Average Cal sensor Temp.: 93.29 Ambient Temp.: 83.39 F
14:50:13.720 -> Average Cal sensor Temp.: 93.33 Ambient Temp.: 83.35 F
14:50:15.037 -> Ambient Temp.: 83.35 F
14:50:15.624 -> Ambient Temp.: 83.35 F
14:50:16.212 -> Ambient Temp.: 83.35 F
14:50:16.765 -> Ambient Temp.: 83.35 F
14:50:18.317 -> Average Cal sensor Temp.: 93.06 Ambient Temp.: 83.28 F
14:50:19.875 -> Average Cal sensor Temp.: 93.30 Ambient Temp.: 83.32 F
14:50:21.429 -> Average Cal sensor Temp.: 93.68 Ambient Temp.: 83.35 F
14:50:22.983 -> Average Cal sensor Temp.: 93.80 Ambient Temp.: 83.35 F
14:50:24.541 -> Average Cal sensor Temp.: 93.54 Ambient Temp.: 83.35 F
14:50:26.099 -> Average Cal sensor Temp.: 93.96 Ambient Temp.: 83.32 F
14:50:27.657 -> Average Cal sensor Temp.: 93.59 Ambient Temp.: 83.35 F
14:50:29.210 -> Average Cal sensor Temp.: 93.58 Ambient Temp.: 83.35 F
14:50:30.767 -> Average Cal sensor Temp.: 93.75 Ambient Temp.: 83.39 F
14:50:32.324 -> Average Cal sensor Temp.: 93.61 Ambient Temp.: 83.35 F
14:50:33.877 -> Average Cal sensor Temp.: 93.52 Ambient Temp.: 83.39 F
14:50:35.431 -> Average Cal sensor Temp.: 93.32 Ambient Temp.: 83.35 F
14:50:36.981 -> Average Cal sensor Temp.: 93.63 Ambient Temp.: 83.39 F
14:50:38.531 -> Average Cal sensor Temp.: 93.80 Ambient Temp.: 83.35 F
14:50:40.117 -> Average Cal sensor Temp.: 94.00 Ambient Temp.: 83.35 F
14:50:41.670 -> Average Cal sensor Temp.: 94.05 Ambient Temp.: 83.35 F
14:50:43.223 -> Average Cal sensor Temp.: 93.98 Ambient Temp.: 83.39 F
14:50:44.774 -> Average Cal sensor Temp.: 93.47 Ambient Temp.: 83.39 F
14:50:46.327 -> Average Cal sensor Temp.: 93.37 Ambient Temp.: 83.35 F
14:50:47.880 -> Average Cal sensor Temp.: 93.87 Ambient Temp.: 83.43 F
14:50:49.434 -> Average Cal sensor Temp.: 93.83 Ambient Temp.: 83.35 F
14:50:50.985 -> Average Cal sensor Temp.: 93.47 Ambient Temp.: 83.39 F
14:50:52.539 -> Average Cal sensor Temp.: 93.77 Ambient Temp.: 83.39 F
14:50:54.091 -> Average Cal sensor Temp.: 93.73 Ambient Temp.: 83.32 F
14:50:55.644 -> Average Cal sensor Temp.: 93.83 Ambient Temp.: 83.39 F
14:50:57.198 -> Average Cal sensor Temp.: 93.82 Ambient Temp.: 83.43 F
14:50:58.782 -> Average Cal sensor Temp.: 94.04 Ambient Temp.: 83.43 F
14:51:00.337 -> Average Cal sensor Temp.: 93.66 Ambient Temp.: 83.39 F
14:51:01.886 -> Average Cal sensor Temp.: 93.76 Ambient Temp.: 83.35 F
14:51:03.434 -> Average Cal sensor Temp.: 93.93 Ambient Temp.: 83.50 F
14:51:04.990 -> Average Cal sensor Temp.: 93.67 Ambient Temp.: 83.43 F
14:51:06.540 -> Average Cal sensor Temp.: 93.69 Ambient Temp.: 83.46 F
14:51:08.100 -> Average Cal sensor Temp.: 93.74 Ambient Temp.: 83.46 F
14:51:09.652 -> Average Cal sensor Temp.: 93.72 Ambient Temp.: 83.50 F
14:51:11.236 -> Average Cal sensor Temp.: 94.38 Ambient Temp.: 83.53 F
14:51:12.760 -> Average Cal sensor Temp.: 93.75 Ambient Temp.: 83.50 F
14:51:14.347 -> Average Cal sensor Temp.: 93.73 Ambient Temp.: 83.50 F
14:51:15.892 -> Average Cal sensor Temp.: 93.79 Ambient Temp.: 83.64 F
14:51:17.444 -> Average Cal sensor Temp.: 93.51 Ambient Temp.: 83.64 F
14:51:18.989 -> Average Cal sensor Temp.: 93.30 Ambient Temp.: 83.57 F
14:51:20.544 -> Average Cal sensor Temp.: 93.08 Ambient Temp.: 83.53 F
14:51:22.099 -> Average Cal sensor Temp.: 93.42 Ambient Temp.: 83.57 F
14:51:23.654 -> Average Cal sensor Temp.: 93.55 Ambient Temp.: 83.57 F
14:51:25.214 -> Average Cal sensor Temp.: 93.90 Ambient Temp.: 83.57 F
14:51:26.775 -> Average Cal sensor Temp.: 93.77 Ambient Temp.: 83.64 F
14:51:28.330 -> Average Cal sensor Temp.: 93.88 Ambient Temp.: 83.68 F
14:51:29.875 -> Average Cal sensor Temp.: 93.89 Ambient Temp.: 83.64 F
14:51:31.462 -> Average Cal sensor Temp.: 93.81 Ambient Temp.: 83.64 F
14:51:33.016 -> Average Cal sensor Temp.: 93.89 Ambient Temp.: 83.64 F
14:51:34.573 -> Average Cal sensor Temp.: 93.92 Ambient Temp.: 83.68 F
14:51:36.130 -> Average Cal sensor Temp.: 93.95 Ambient Temp.: 83.71 F
14:51:37.679 -> Average Cal sensor Temp.: 94.10 Ambient Temp.: 83.68 F
14:51:39.230 -> Average Cal sensor Temp.: 93.99 Ambient Temp.: 83.71 F
14:51:40.780 -> Average Cal sensor Temp.: 94.00 Ambient Temp.: 83.75 F
14:51:42.337 -> Average Cal sensor Temp.: 93.96 Ambient Temp.: 83.75 F
14:51:43.889 -> Average Cal sensor Temp.: 93.66 Ambient Temp.: 83.79 F
14:51:45.438 -> Average Cal sensor Temp.: 93.83 Ambient Temp.: 83.79 F
14:51:46.992 -> Average Cal sensor Temp.: 93.83 Ambient Temp.: 83.75 F
14:51:48.549 -> Average Cal sensor Temp.: 93.76 Ambient Temp.: 83.79 F
14:51:50.105 -> Average Cal sensor Temp.: 93.92 Ambient Temp.: 83.86 F
14:51:51.661 -> Average Cal sensor Temp.: 93.77 Ambient Temp.: 83.82 F
14:51:53.248 -> Average Cal sensor Temp.: 93.93 Ambient Temp.: 83.82 F
14:51:54.804 -> Average Cal sensor Temp.: 93.70 Ambient Temp.: 83.86 F
14:51:56.357 -> Average Cal sensor Temp.: 93.98 Ambient Temp.: 83.86 F
14:51:57.913 -> Average Cal sensor Temp.: 94.16 Ambient Temp.: 83.89 F
14:51:59.467 -> Average Cal sensor Temp.: 93.98 Ambient Temp.: 83.89 F
14:52:01.027 -> Average Cal sensor Temp.: 93.88 Ambient Temp.: 83.93 F
14:52:02.587 -> Average Cal sensor Temp.: 93.72 Ambient Temp.: 83.93 F
14:52:04.140 -> Average Cal sensor Temp.: 93.92 Ambient Temp.: 83.97 F
14:52:05.695 -> Average Cal sensor Temp.: 93.92 Ambient Temp.: 83.97 F
14:52:07.245 -> Average Cal sensor Temp.: 93.81 Ambient Temp.: 84.07 F
14:52:08.794 -> Average Cal sensor Temp.: 93.75 Ambient Temp.: 84.00 F
14:52:10.354 -> Average Cal sensor Temp.: 93.51 Ambient Temp.: 84.11 F
14:52:11.900 -> Average Cal sensor Temp.: 93.79 Ambient Temp.: 84.07 F
14:52:13.453 -> Average Cal sensor Temp.: 93.65 Ambient Temp.: 84.07 F
14:52:15.006 -> Average Cal sensor Temp.: 93.60 Ambient Temp.: 84.11 F
14:52:16.558 -> Average Cal sensor Temp.: 93.70 Ambient Temp.: 84.18 F
14:52:18.117 -> Average Cal sensor Temp.: 94.06 Ambient Temp.: 84.15 F
14:52:19.674 -> Average Cal sensor Temp.: 93.78 Ambient Temp.: 84.18 F
14:52:21.228 -> Average Cal sensor Temp.: 93.70 Ambient Temp.: 84.25 F
14:52:22.782 -> Average Cal sensor Temp.: 93.83 Ambient Temp.: 84.18 F
14:52:24.338 -> Average Cal sensor Temp.: 93.49 Ambient Temp.: 84.29 F
14:52:25.928 -> Average Cal sensor Temp.: 93.15 Ambient Temp.: 84.22 F
14:52:27.456 -> Average Cal sensor Temp.: 93.44 Ambient Temp.: 84.25 F
14:52:29.012 -> Average Cal sensor Temp.: 93.87 Ambient Temp.: 84.25 F
14:52:30.568 -> Average Cal sensor Temp.: 93.85 Ambient Temp.: 84.25 F
14:52:32.121 -> Average Cal sensor Temp.: 94.01 Ambient Temp.: 84.25 F
14:52:33.705 -> Average Cal sensor Temp.: 93.96 Ambient Temp.: 84.36 F
14:52:35.257 -> Average Cal sensor Temp.: 93.72 Ambient Temp.: 84.40 F
14:52:36.816 -> Average Cal sensor Temp.: 93.96 Ambient Temp.: 84.40 F
14:52:38.367 -> Average Cal sensor Temp.: 94.05 Ambient Temp.: 84.40 F
14:52:39.919 -> Average Cal sensor Temp.: 93.51 Ambient Temp.: 84.47 F
14:52:41.468 -> Average Cal sensor Temp.: 93.73 Ambient Temp.: 84.54 F
14:52:43.022 -> Average Cal sensor Temp.: 93.48 Ambient Temp.: 84.51 F
14:52:44.577 -> Average Cal sensor Temp.: 93.42 Ambient Temp.: 84.54 F
14:52:46.131 -> Average Cal sensor Temp.: 93.71 Ambient Temp.: 84.51 F
14:52:47.689 -> Average Cal sensor Temp.: 93.91 Ambient Temp.: 84.58 F
14:52:49.240 -> Average Cal sensor Temp.: 93.85 Ambient Temp.: 84.61 F
14:52:50.792 -> Average Cal sensor Temp.: 94.14 Ambient Temp.: 84.61 F
14:52:52.347 -> Average Cal sensor Temp.: 94.01 Ambient Temp.: 84.69 F
14:52:53.937 -> Average Cal sensor Temp.: 93.91 Ambient Temp.: 84.65 F
14:52:55.490 -> Average Cal sensor Temp.: 93.77 Ambient Temp.: 84.72 F
14:52:57.017 -> Average Cal sensor Temp.: 93.97 Ambient Temp.: 84.72 F
14:52:58.569 -> Average Cal sensor Temp.: 94.01 Ambient Temp.: 84.76 F
14:53:00.127 -> Average Cal sensor Temp.: 94.33 Ambient Temp.: 84.79 F
14:53:01.685 -> Average Cal sensor Temp.: 94.27 Ambient Temp.: 84.79 F
14:53:03.238 -> Average Cal sensor Temp.: 94.12 Ambient Temp.: 84.87 F
14:53:04.826 -> Average Cal sensor Temp.: 94.23 Ambient Temp.: 84.90 F
14:53:06.383 -> Average Cal sensor Temp.: 94.35 Ambient Temp.: 84.90 F
14:53:07.940 -> Average Cal sensor Temp.: 94.50 Ambient Temp.: 84.97 F
14:53:09.496 -> Average Cal sensor Temp.: 94.12 Ambient Temp.: 84.97 F
14:53:11.050 -> Average Cal sensor Temp.: 93.79 Ambient Temp.: 84.90 F
14:53:12.598 -> Average Cal sensor Temp.: 93.62 Ambient Temp.: 85.01 F
14:53:14.156 -> Average Cal sensor Temp.: 93.78 Ambient Temp.: 85.01 F
14:53:15.707 -> Average Cal sensor Temp.: 93.81 Ambient Temp.: 85.05 F
14:53:17.258 -> Average Cal sensor Temp.: 94.00 Ambient Temp.: 85.05 F
14:53:18.807 -> Average Cal sensor Temp.: 94.35 Ambient Temp.: 85.08 F
14:53:20.364 -> Average Cal sensor Temp.: 94.31 Ambient Temp.: 85.08 F
14:53:21.915 -> Average Cal sensor Temp.: 93.90 Ambient Temp.: 85.08 F
14:53:23.500 -> Average Cal sensor Temp.: 93.76 Ambient Temp.: 85.12 F
14:53:25.052 -> Average Cal sensor Temp.: 94.04 Ambient Temp.: 85.12 F
14:53:26.607 -> Average Cal sensor Temp.: 94.27 Ambient Temp.: 85.19 F
14:53:28.159 -> Average Cal sensor Temp.: 94.03 Ambient Temp.: 85.12 F
14:53:29.718 -> Average Cal sensor Temp.: 93.93 Ambient Temp.: 85.19 F
14:53:31.272 -> Average Cal sensor Temp.: 93.66 Ambient Temp.: 85.19 F
14:53:32.825 -> Average Cal sensor Temp.: 94.08 Ambient Temp.: 85.12 F
14:53:34.381 -> Average Cal sensor Temp.: 93.83 Ambient Temp.: 85.12 F
14:53:35.932 -> Average Cal sensor Temp.: 94.20 Ambient Temp.: 85.19 F
14:53:37.489 -> Average Cal sensor Temp.: 94.15 Ambient Temp.: 85.19 F
14:53:39.050 -> Average Cal sensor Temp.: 93.93 Ambient Temp.: 85.19 F
14:53:40.607 -> Average Cal sensor Temp.: 94.30 Ambient Temp.: 85.19 F
14:53:42.160 -> Average Cal sensor Temp.: 93.95 Ambient Temp.: 85.19 F
14:53:43.714 -> Average Cal sensor Temp.: 94.18 Ambient Temp.: 85.30 F
14:53:45.272 -> Average Cal sensor Temp.: 94.02 Ambient Temp.: 85.30 F
14:53:46.833 -> Average Cal sensor Temp.: 94.13 Ambient Temp.: 85.26 F
14:53:48.383 -> Average Cal sensor Temp.: 94.04 Ambient Temp.: 85.15 F
14:53:49.924 -> Average Cal sensor Temp.: 94.19 Ambient Temp.: 85.26 F
14:53:51.477 -> Average Cal sensor Temp.: 93.93 Ambient Temp.: 85.30 F
14:53:53.032 -> Average Cal sensor Temp.: 93.92 Ambient Temp.: 85.26 F
14:53:54.615 -> Average Cal sensor Temp.: 93.79 Ambient Temp.: 85.26 F
14:53:56.165 -> Average Cal sensor Temp.: 94.04 Ambient Temp.: 85.19 F
14:53:57.711 -> Average Cal sensor Temp.: 93.92 Ambient Temp.: 85.26 F
14:53:59.261 -> Average Cal sensor Temp.: 93.79 Ambient Temp.: 85.26 F
14:54:00.809 -> Average Cal sensor Temp.: 93.49 Ambient Temp.: 85.19 F
14:54:02.396 -> Average Cal sensor Temp.: 93.76 Ambient Temp.: 85.19 F
14:54:03.949 -> Average Cal sensor Temp.: 93.93 Ambient Temp.: 85.26 F
14:54:05.503 -> Average Cal sensor Temp.: 93.47 Ambient Temp.: 85.15 F
14:54:07.058 -> Average Cal sensor Temp.: 93.68 Ambient Temp.: 85.15 F
14:54:08.607 -> Average Cal sensor Temp.: 93.75 Ambient Temp.: 85.12 F
14:54:10.159 -> Average Cal sensor Temp.: 93.55 Ambient Temp.: 85.12 F
14:54:11.712 -> Average Cal sensor Temp.: 93.87 Ambient Temp.: 85.12 F
14:54:13.270 -> Average Cal sensor Temp.: 93.46 Ambient Temp.: 85.08 F
14:54:14.817 -> Average Cal sensor Temp.: 93.84 Ambient Temp.: 85.08 F
14:54:16.369 -> Average Cal sensor Temp.: 93.93 Ambient Temp.: 85.05 F
14:54:17.926 -> Average Cal sensor Temp.: 94.15 Ambient Temp.: 85.08 F
14:54:19.479 -> Average Cal sensor Temp.: 93.94 Ambient Temp.: 85.05 F
14:54:21.063 -> Average Cal sensor Temp.: 94.43 Ambient Temp.: 85.08 F
14:54:22.624 -> Average Cal sensor Temp.: 94.18 Ambient Temp.: 85.01 F
14:54:24.148 -> Average Cal sensor Temp.: 94.06 Ambient Temp.: 85.08 F
14:54:25.703 -> Average Cal sensor Temp.: 94.18 Ambient Temp.: 85.05 F
14:54:27.289 -> Average Cal sensor Temp.: 94.37 Ambient Temp.: 85.05 F
14:54:28.845 -> Average Cal sensor Temp.: 94.05 Ambient Temp.: 85.08 F
14:54:30.397 -> Average Cal sensor Temp.: 94.30 Ambient Temp.: 85.01 F
14:54:31.957 -> Average Cal sensor Temp.: 94.31 Ambient Temp.: 85.08 F
14:54:33.505 -> Average Cal sensor Temp.: 94.17 Ambient Temp.: 85.08 F
14:54:35.058 -> Average Cal sensor Temp.: 93.82 Ambient Temp.: 85.05 F
14:54:36.613 -> Average Cal sensor Temp.: 93.92 Ambient Temp.: 85.08 F
14:54:38.169 -> Average Cal sensor Temp.: 93.95 Ambient Temp.: 85.08 F
14:54:39.718 -> Average Cal sensor Temp.: 94.16 Ambient Temp.: 85.08 F
14:54:41.267 -> Average Cal sensor Temp.: 94.10 Ambient Temp.: 85.05 F
14:54:42.847 -> Average Cal sensor Temp.: 94.15 Ambient Temp.: 85.05 F
14:54:44.405 -> Average Cal sensor Temp.: 94.20 Ambient Temp.: 85.08 F
14:54:45.965 -> Average Cal sensor Temp.: 94.37 Ambient Temp.: 85.05 F
14:54:47.516 -> Average Cal sensor Temp.: 94.20 Ambient Temp.: 85.08 F
14:54:49.073 -> Average Cal sensor Temp.: 93.90 Ambient Temp.: 85.05 F
14:54:50.620 -> Average Cal sensor Temp.: 93.95 Ambient Temp.: 85.08 F
14:54:52.173 -> Average Cal sensor Temp.: 93.73 Ambient Temp.: 85.01 F
14:54:53.723 -> Average Cal sensor Temp.: 93.76 Ambient Temp.: 85.05 F
14:54:55.276 -> Average Cal sensor Temp.: 93.66 Ambient Temp.: 85.01 F
14:54:56.832 -> Average Cal sensor Temp.: 93.84 Ambient Temp.: 85.01 F
14:54:58.381 -> Average Cal sensor Temp.: 93.73 Ambient Temp.: 84.94 F
14:54:59.967 -> Average Cal sensor Temp.: 93.80 Ambient Temp.: 84.97 F
14:55:01.518 -> Average Cal sensor Temp.: 94.13 Ambient Temp.: 84.97 F
14:55:03.078 -> Average Cal sensor Temp.: 94.22 Ambient Temp.: 84.97 F
14:55:04.633 -> Average Cal sensor Temp.: 94.28 Ambient Temp.: 84.97 F
14:55:06.190 -> Average Cal sensor Temp.: 93.81 Ambient Temp.: 84.97 F
14:55:07.742 -> Average Cal sensor Temp.: 93.95 Ambient Temp.: 84.90 F
14:55:09.295 -> Average Cal sensor Temp.: 93.83 Ambient Temp.: 84.94 F
14:55:10.852 -> Average Cal sensor Temp.: 94.29 Ambient Temp.: 84.90 F
14:55:12.400 -> Average Cal sensor Temp.: 94.22 Ambient Temp.: 84.90 F
14:55:13.950 -> Average Cal sensor Temp.: 94.44 Ambient Temp.: 84.87 F
14:55:15.540 -> Average Cal sensor Temp.: 94.46 Ambient Temp.: 84.90 F
14:55:17.094 -> Average Cal sensor Temp.: 94.56 Ambient Temp.: 84.94 F
14:55:18.640 -> Average Cal sensor Temp.: 94.62 Ambient Temp.: 84.87 F
14:55:20.198 -> Average Cal sensor Temp.: 94.73 Ambient Temp.: 84.87 F
14:55:21.750 -> Average Cal sensor Temp.: 94.50 Ambient Temp.: 84.90 F
14:55:23.299 -> Average Cal sensor Temp.: 93.96 Ambient Temp.: 84.87 F
14:55:24.862 -> Average Cal sensor Temp.: 93.87 Ambient Temp.: 84.87 F
14:55:26.415 -> Average Cal sensor Temp.: 94.45 Ambient Temp.: 84.87 F
14:55:27.971 -> Average Cal sensor Temp.: 94.00 Ambient Temp.: 84.79 F
14:55:29.522 -> Average Cal sensor Temp.: 93.57 Ambient Temp.: 84.87 F
14:55:31.108 -> Average Cal sensor Temp.: 93.99 Ambient Temp.: 84.72 F
14:55:32.665 -> Average Cal sensor Temp.: 93.95 Ambient Temp.: 84.79 F
14:55:34.213 -> Average Cal sensor Temp.: 94.41 Ambient Temp.: 84.79 F
14:55:35.769 -> Average Cal sensor Temp.: 94.33 Ambient Temp.: 84.79 F
14:55:37.321 -> Average Cal sensor Temp.: 93.96 Ambient Temp.: 84.76 F
14:55:38.875 -> Average Cal sensor Temp.: 94.17 Ambient Temp.: 84.76 F
14:55:40.429 -> Average Cal sensor Temp.: 93.81 Ambient Temp.: 84.69 F
14:55:41.983 -> Average Cal sensor Temp.: 93.91 Ambient Temp.: 84.65 F
14:55:43.532 -> Average Cal sensor Temp.: 93.81 Ambient Temp.: 84.65 F
14:55:45.088 -> Average Cal sensor Temp.: 93.38 Ambient Temp.: 84.65 F
14:55:46.643 -> Average Cal sensor Temp.: 94.06 Ambient Temp.: 84.58 F
14:55:48.229 -> Average Cal sensor Temp.: 93.98 Ambient Temp.: 84.58 F
14:55:49.781 -> Average Cal sensor Temp.: 93.35 Ambient Temp.: 84.58 F
14:55:51.339 -> Average Cal sensor Temp.: 94.20 Ambient Temp.: 84.51 F
14:55:52.889 -> Average Cal sensor Temp.: 94.17 Ambient Temp.: 84.51 F
14:55:54.443 -> Average Cal sensor Temp.: 93.66 Ambient Temp.: 84.40 F
14:55:55.998 -> Average Cal sensor Temp.: 93.90 Ambient Temp.: 84.47 F
14:55:57.551 -> Average Cal sensor Temp.: 94.22 Ambient Temp.: 84.47 F
14:55:59.101 -> Average Cal sensor Temp.: 93.89 Ambient Temp.: 84.47 F
14:56:00.684 -> Average Cal sensor Temp.: 93.95 Ambient Temp.: 84.40 F
14:56:02.236 -> Average Cal sensor Temp.: 94.06 Ambient Temp.: 84.40 F
14:56:03.783 -> Average Cal sensor Temp.: 94.17 Ambient Temp.: 84.40 F
14:56:05.342 -> Average Cal sensor Temp.: 93.96 Ambient Temp.: 84.33 F
14:56:06.896 -> Average Cal sensor Temp.: 94.09 Ambient Temp.: 84.36 F
14:56:08.438 -> Average Cal sensor Temp.: 93.90 Ambient Temp.: 84.29 F
14:56:10.025 -> Average Cal sensor Temp.: 93.79 Ambient Temp.: 84.36 F
14:56:11.578 -> Average Cal sensor Temp.: 94.12 Ambient Temp.: 84.33 F
14:56:13.128 -> Average Cal sensor Temp.: 94.22 Ambient Temp.: 84.33 F
14:56:14.686 -> Average Cal sensor Temp.: 94.38 Ambient Temp.: 84.33 F
14:56:16.243 -> Average Cal sensor Temp.: 94.18 Ambient Temp.: 84.25 F
14:56:17.802 -> Average Cal sensor Temp.: 93.78 Ambient Temp.: 84.25 F
14:56:19.353 -> Average Cal sensor Temp.: 93.89 Ambient Temp.: 84.25 F
14:56:20.904 -> Average Cal sensor Temp.: 93.95 Ambient Temp.: 84.22 F
14:56:22.459 -> Average Cal sensor Temp.: 93.79 Ambient Temp.: 84.22 F
14:56:24.008 -> Average Cal sensor Temp.: 94.11 Ambient Temp.: 84.22 F
14:56:25.570 -> Average Cal sensor Temp.: 93.70 Ambient Temp.: 84.18 F
14:56:27.131 -> Average Cal sensor Temp.: 93.87 Ambient Temp.: 84.15 F
14:56:28.676 -> Average Cal sensor Temp.: 94.04 Ambient Temp.: 84.15 F
14:56:30.229 -> Average Cal sensor Temp.: 94.24 Ambient Temp.: 84.11 F
14:56:31.779 -> Average Cal sensor Temp.: 93.93 Ambient Temp.: 84.07 F
14:56:33.333 -> Average Cal sensor Temp.: 93.90 Ambient Temp.: 84.07 F
14:56:34.916 -> Average Cal sensor Temp.: 94.13 Ambient Temp.: 84.00 F
14:56:36.472 -> Average Cal sensor Temp.: 94.04 Ambient Temp.: 84.00 F
14:56:38.025 -> Average Cal sensor Temp.: 94.50 Ambient Temp.: 84.07 F
14:56:39.582 -> Average Cal sensor Temp.: 94.14 Ambient Temp.: 84.07 F
14:56:41.136 -> Average Cal sensor Temp.: 94.20 Ambient Temp.: 84.00 F
14:56:42.688 -> Average Cal sensor Temp.: 94.17 Ambient Temp.: 83.97 F
14:56:44.251 -> Average Cal sensor Temp.: 94.00 Ambient Temp.: 83.97 F
14:56:45.809 -> Average Cal sensor Temp.: 93.84 Ambient Temp.: 83.89 F
14:56:47.361 -> Average Cal sensor Temp.: 93.59 Ambient Temp.: 83.86 F
14:56:48.913 -> Average Cal sensor Temp.: 93.43 Ambient Temp.: 83.86 F
14:56:50.461 -> Average Cal sensor Temp.: 93.84 Ambient Temp.: 83.86 F
14:56:52.044 -> Average Cal sensor Temp.: 94.06 Ambient Temp.: 83.82 F
14:56:53.590 -> Average Cal sensor Temp.: 94.23 Ambient Temp.: 83.79 F
14:56:55.137 -> Average Cal sensor Temp.: 94.09 Ambient Temp.: 83.75 F
14:56:56.691 -> Average Cal sensor Temp.: 94.42 Ambient Temp.: 83.75 F
14:56:58.240 -> Average Cal sensor Temp.: 94.28 Ambient Temp.: 83.75 F
14:56:59.825 -> Average Cal sensor Temp.: 94.29 Ambient Temp.: 83.75 F
14:57:01.378 -> Average Cal sensor Temp.: 94.20 Ambient Temp.: 83.71 F
14:57:02.932 -> Average Cal sensor Temp.: 93.99 Ambient Temp.: 83.68 F
14:57:04.489 -> Average Cal sensor Temp.: 93.92 Ambient Temp.: 83.68 F
14:57:06.041 -> Average Cal sensor Temp.: 94.38 Ambient Temp.: 83.71 F
14:57:07.595 -> Average Cal sensor Temp.: 94.51 Ambient Temp.: 83.75 F
14:57:09.146 -> Average Cal sensor Temp.: 94.40 Ambient Temp.: 83.71 F
14:57:10.701 -> Average Cal sensor Temp.: 94.19 Ambient Temp.: 83.75 F
14:57:12.253 -> Average Cal sensor Temp.: 94.16 Ambient Temp.: 83.71 F
14:57:13.804 -> Average Cal sensor Temp.: 94.43 Ambient Temp.: 83.75 F
14:57:15.358 -> Average Cal sensor Temp.: 94.10 Ambient Temp.: 83.68 F
14:57:16.911 -> Average Cal sensor Temp.: 94.29 Ambient Temp.: 83.68 F
14:57:18.500 -> Average Cal sensor Temp.: 94.40 Ambient Temp.: 83.71 F
14:57:20.025 -> Average Cal sensor Temp.: 94.23 Ambient Temp.: 83.68 F
14:57:21.611 -> Average Cal sensor Temp.: 94.04 Ambient Temp.: 83.71 F
14:57:23.162 -> Average Cal sensor Temp.: 94.28 Ambient Temp.: 83.75 F
14:57:24.707 -> Average Cal sensor Temp.: 94.48 Ambient Temp.: 83.71 F
14:57:26.260 -> Average Cal sensor Temp.: 94.33 Ambient Temp.: 83.75 F
14:57:27.814 -> Average Cal sensor Temp.: 93.42 Ambient Temp.: 83.71 F
14:57:29.368 -> Average Cal sensor Temp.: 93.79 Ambient Temp.: 83.68 F
14:57:30.921 -> Average Cal sensor Temp.: 94.44 Ambient Temp.: 83.71 F
14:57:32.480 -> Average Cal sensor Temp.: 94.33 Ambient Temp.: 83.68 F
14:57:34.035 -> Average Cal sensor Temp.: 94.22 Ambient Temp.: 83.64 F
14:57:35.590 -> Average Cal sensor Temp.: 94.06 Ambient Temp.: 83.64 F
14:57:37.175 -> Average Cal sensor Temp.: 93.93 Ambient Temp.: 83.68 F
14:57:38.727 -> Average Cal sensor Temp.: 94.73 Ambient Temp.: 83.68 F
14:57:40.281 -> Average Cal sensor Temp.: 94.47 Ambient Temp.: 83.68 F
14:57:41.837 -> Average Cal sensor Temp.: 94.59 Ambient Temp.: 83.71 F
14:57:43.392 -> Average Cal sensor Temp.: 94.57 Ambient Temp.: 83.68 F
14:57:44.944 -> Average Cal sensor Temp.: 94.56 Ambient Temp.: 83.68 F
14:57:46.500 -> Average Cal sensor Temp.: 94.42 Ambient Temp.: 83.68 F
14:57:48.049 -> Average Cal sensor Temp.: 94.26 Ambient Temp.: 83.71 F
14:57:49.603 -> Average Cal sensor Temp.: 94.65 Ambient Temp.: 83.75 F
14:57:51.161 -> Average Cal sensor Temp.: 94.79 Ambient Temp.: 83.75 F
14:57:52.713 -> Average Cal sensor Temp.: 94.94 Ambient Temp.: 83.75 F
14:57:54.264 -> Average Cal sensor Temp.: 95.28 Ambient Temp.: 83.79 F
14:57:55.819 -> Average Cal sensor Temp.: 95.10 Ambient Temp.: 83.79 F
14:57:57.374 -> Average Cal sensor Temp.: 94.94 Ambient Temp.: 83.75 F
14:57:58.962 -> Average Cal sensor Temp.: 94.86 Ambient Temp.: 83.79 F
14:58:00.515 -> Average Cal sensor Temp.: 94.90 Ambient Temp.: 83.82 F
14:58:02.064 -> Average Cal sensor Temp.: 94.51 Ambient Temp.: 83.82 F
14:58:03.617 -> Average Cal sensor Temp.: 94.55 Ambient Temp.: 83.79 F
14:58:05.170 -> Average Cal sensor Temp.: 94.21 Ambient Temp.: 83.82 F
14:58:06.722 -> Average Cal sensor Temp.: 94.34 Ambient Temp.: 83.82 F
14:58:08.276 -> Average Cal sensor Temp.: 94.47 Ambient Temp.: 83.89 F
14:58:09.826 -> Average Cal sensor Temp.: 94.60 Ambient Temp.: 83.89 F
14:58:11.410 -> Average Cal sensor Temp.: 94.22 Ambient Temp.: 83.82 F
14:58:12.966 -> Average Cal sensor Temp.: 94.17 Ambient Temp.: 83.86 F
14:58:14.514 -> Average Cal sensor Temp.: 94.44 Ambient Temp.: 83.89 F
14:58:16.067 -> Average Cal sensor Temp.: 94.69 Ambient Temp.: 83.93 F
14:58:17.622 -> Average Cal sensor Temp.: 94.91 Ambient Temp.: 83.86 F
14:58:19.171 -> Average Cal sensor Temp.: 94.60 Ambient Temp.: 83.89 F
14:58:20.730 -> Average Cal sensor Temp.: 94.74 Ambient Temp.: 83.93 F
14:58:22.282 -> Average Cal sensor Temp.: 94.63 Ambient Temp.: 83.97 F
14:58:23.839 -> Average Cal sensor Temp.: 94.32 Ambient Temp.: 84.00 F
14:58:25.420 -> Average Cal sensor Temp.: 94.41 Ambient Temp.: 83.97 F
14:58:26.976 -> Average Cal sensor Temp.: 93.94 Ambient Temp.: 84.00 F
14:58:28.530 -> Average Cal sensor Temp.: 93.86 Ambient Temp.: 83.97 F
14:58:30.082 -> Average Cal sensor Temp.: 93.99 Ambient Temp.: 83.97 F
14:58:31.638 -> Average Cal sensor Temp.: 94.10 Ambient Temp.: 84.07 F
14:58:33.199 -> Average Cal sensor Temp.: 94.13 Ambient Temp.: 84.00 F
14:58:34.757 -> Average Cal sensor Temp.: 94.17 Ambient Temp.: 84.07 F
14:58:36.321 -> Average Cal sensor Temp.: 94.72 Ambient Temp.: 84.11 F
14:58:37.876 -> Average Cal sensor Temp.: 94.56 Ambient Temp.: 84.15 F
14:58:39.432 -> Average Cal sensor Temp.: 94.82 Ambient Temp.: 84.18 F
14:58:40.986 -> Average Cal sensor Temp.: 94.60 Ambient Temp.: 84.11 F
14:58:42.543 -> Average Cal sensor Temp.: 94.28 Ambient Temp.: 84.15 F
14:58:44.097 -> Average Cal sensor Temp.: 93.89 Ambient Temp.: 84.18 F
14:58:45.655 -> Average Cal sensor Temp.: 94.46 Ambient Temp.: 84.18 F
14:58:47.205 -> Average Cal sensor Temp.: 94.58 Ambient Temp.: 84.22 F
14:58:48.754 -> Average Cal sensor Temp.: 94.52 Ambient Temp.: 84.18 F
14:58:50.301 -> Average Cal sensor Temp.: 94.76 Ambient Temp.: 84.25 F
14:58:51.856 -> Average Cal sensor Temp.: 94.28 Ambient Temp.: 84.22 F
14:58:53.414 -> Average Cal sensor Temp.: 94.41 Ambient Temp.: 84.25 F
14:58:54.974 -> Average Cal sensor Temp.: 94.41 Ambient Temp.: 84.25 F
14:58:56.530 -> Average Cal sensor Temp.: 94.40 Ambient Temp.: 84.22 F
14:58:58.079 -> Average Cal sensor Temp.: 94.89 Ambient Temp.: 84.25 F
14:58:59.658 -> Average Cal sensor Temp.: 94.47 Ambient Temp.: 84.36 F
14:59:01.214 -> Average Cal sensor Temp.: 94.38 Ambient Temp.: 84.33 F
14:59:02.765 -> Average Cal sensor Temp.: 94.43 Ambient Temp.: 84.33 F
14:59:04.316 -> Average Cal sensor Temp.: 94.50 Ambient Temp.: 84.40 F
14:59:05.866 -> Average Cal sensor Temp.: 94.20 Ambient Temp.: 84.40 F
14:59:07.448 -> Average Cal sensor Temp.: 94.28 Ambient Temp.: 84.40 F
14:59:09.009 -> Average Cal sensor Temp.: 94.25 Ambient Temp.: 84.47 F
14:59:10.559 -> Average Cal sensor Temp.: 94.52 Ambient Temp.: 84.36 F
14:59:12.114 -> Average Cal sensor Temp.: 94.26 Ambient Temp.: 84.47 F
14:59:13.672 -> Average Cal sensor Temp.: 94.35 Ambient Temp.: 84.47 F
14:59:15.225 -> Average Cal sensor Temp.: 94.63 Ambient Temp.: 84.51 F
14:59:16.780 -> Average Cal sensor Temp.: 94.29 Ambient Temp.: 84.51 F
14:59:18.340 -> Average Cal sensor Temp.: 94.23 Ambient Temp.: 84.54 F
14:59:19.894 -> Average Cal sensor Temp.: 94.34 Ambient Temp.: 84.54 F
14:59:21.449 -> Average Cal sensor Temp.: 94.49 Ambient Temp.: 84.61 F
14:59:23.009 -> Average Cal sensor Temp.: 94.25 Ambient Temp.: 84.61 F
14:59:24.568 -> Average Cal sensor Temp.: 94.37 Ambient Temp.: 84.61 F
14:59:26.124 -> Average Cal sensor Temp.: 94.32 Ambient Temp.: 84.61 F
14:59:27.673 -> Average Cal sensor Temp.: 94.59 Ambient Temp.: 84.65 F
14:59:29.224 -> Average Cal sensor Temp.: 94.55 Ambient Temp.: 84.72 F
14:59:30.773 -> Average Cal sensor Temp.: 94.33 Ambient Temp.: 84.72 F
14:59:32.333 -> Average Cal sensor Temp.: 94.28 Ambient Temp.: 84.79 F
14:59:33.887 -> Average Cal sensor Temp.: 94.40 Ambient Temp.: 84.76 F
14:59:35.439 -> Average Cal sensor Temp.: 94.41 Ambient Temp.: 84.87 F
14:59:37.025 -> Average Cal sensor Temp.: 94.73 Ambient Temp.: 84.79 F
14:59:38.582 -> Average Cal sensor Temp.: 94.80 Ambient Temp.: 84.87 F
14:59:40.136 -> Average Cal sensor Temp.: 94.34 Ambient Temp.: 84.79 F
14:59:41.693 -> Average Cal sensor Temp.: 94.61 Ambient Temp.: 84.90 F
14:59:43.248 -> Average Cal sensor Temp.: 94.47 Ambient Temp.: 84.90 F
14:59:44.801 -> Average Cal sensor Temp.: 94.46 Ambient Temp.: 84.94 F
14:59:46.352 -> Average Cal sensor Temp.: 94.54 Ambient Temp.: 84.97 F
14:59:47.903 -> Average Cal sensor Temp.: 94.49 Ambient Temp.: 85.01 F
14:59:49.459 -> Average Cal sensor Temp.: 94.47 Ambient Temp.: 85.01 F
14:59:51.015 -> Average Cal sensor Temp.: 94.45 Ambient Temp.: 85.05 F
14:59:52.571 -> Average Cal sensor Temp.: 94.47 Ambient Temp.: 85.05 F
14:59:54.129 -> Average Cal sensor Temp.: 94.44 Ambient Temp.: 85.12 F
14:59:55.685 -> Average Cal sensor Temp.: 94.63 Ambient Temp.: 85.12 F
14:59:57.239 -> Average Cal sensor Temp.: 94.37 Ambient Temp.: 85.12 F
14:59:58.793 -> Average Cal sensor Temp.: 94.55 Ambient Temp.: 85.19 F
15:00:00.345 -> Average Cal sensor Temp.: 94.02 Ambient Temp.: 85.26 F
15:00:01.903 -> Average Cal sensor Temp.: 94.10 Ambient Temp.: 85.26 F
15:00:03.458 -> Average Cal sensor Temp.: 94.21 Ambient Temp.: 85.26 F
15:00:05.011 -> Average Cal sensor Temp.: 94.49 Ambient Temp.: 85.33 F
15:00:06.567 -> Average Cal sensor Temp.: 95.03 Ambient Temp.: 85.30 F
15:00:08.119 -> Average Cal sensor Temp.: 95.27 Ambient Temp.: 85.33 F
15:00:09.677 -> Average Cal sensor Temp.: 95.48 Ambient Temp.: 85.33 F
15:00:11.233 -> Average Cal sensor Temp.: 95.79 Ambient Temp.: 85.41 F
15:00:12.785 -> Average Cal sensor Temp.: 95.51 Ambient Temp.: 85.37 F
15:00:14.340 -> Average Cal sensor Temp.: 94.95 Ambient Temp.: 85.37 F
15:00:15.926 -> Average Cal sensor Temp.: 95.25 Ambient Temp.: 85.37 F
15:00:17.481 -> Average Cal sensor Temp.: 95.61 Ambient Temp.: 85.37 F
15:00:19.036 -> Average Cal sensor Temp.: 95.53 Ambient Temp.: 85.41 F
15:00:20.589 -> Average Cal sensor Temp.: 95.08 Ambient Temp.: 85.41 F
15:00:22.137 -> Average Cal sensor Temp.: 94.43 Ambient Temp.: 85.41 F
15:00:23.693 -> Average Cal sensor Temp.: 94.37 Ambient Temp.: 85.41 F
15:00:25.245 -> Average Cal sensor Temp.: 94.93 Ambient Temp.: 85.41 F
15:00:26.802 -> Average Cal sensor Temp.: 94.57 Ambient Temp.: 85.41 F
15:00:28.355 -> Average Cal sensor Temp.: 94.74 Ambient Temp.: 85.41 F
15:00:29.909 -> Average Cal sensor Temp.: 94.61 Ambient Temp.: 85.44 F
15:00:31.461 -> Average Cal sensor Temp.: 94.82 Ambient Temp.: 85.44 F
15:00:33.013 -> Average Cal sensor Temp.: 94.81 Ambient Temp.: 85.51 F
15:00:34.570 -> Average Cal sensor Temp.: 94.97 Ambient Temp.: 85.44 F
15:00:36.127 -> Average Cal sensor Temp.: 94.77 Ambient Temp.: 85.44 F
15:00:37.682 -> Average Cal sensor Temp.: 95.07 Ambient Temp.: 85.41 F
15:00:39.235 -> Average Cal sensor Temp.: 95.00 Ambient Temp.: 85.44 F
15:00:40.787 -> Average Cal sensor Temp.: 95.06 Ambient Temp.: 85.48 F
15:00:42.370 -> Average Cal sensor Temp.: 94.92 Ambient Temp.: 85.48 F
15:00:43.924 -> Average Cal sensor Temp.: 94.98 Ambient Temp.: 85.48 F
15:00:45.475 -> Average Cal sensor Temp.: 94.68 Ambient Temp.: 85.44 F
15:00:47.028 -> Average Cal sensor Temp.: 94.68 Ambient Temp.: 85.48 F
15:00:48.585 -> Average Cal sensor Temp.: 94.59 Ambient Temp.: 85.48 F
15:00:50.143 -> Average Cal sensor Temp.: 94.26 Ambient Temp.: 85.51 F
15:00:51.696 -> Average Cal sensor Temp.: 94.22 Ambient Temp.: 85.44 F
15:00:53.256 -> Average Cal sensor Temp.: 94.64 Ambient Temp.: 85.44 F
15:00:54.706 -> Ambient Temp.: 85.48 F
15:00:55.293 -> Ambient Temp.: 85.48 F
15:00:55.847 -> Ambient Temp.: 85.48 F
15:00:57.399 -> Average Cal sensor Temp.: 94.55 Ambient Temp.: 85.48 F
15:00:58.954 -> Average Cal sensor Temp.: 94.46 Ambient Temp.: 85.44 F
15:01:00.505 -> Average Cal sensor Temp.: 94.38 Ambient Temp.: 85.48 F
15:01:02.063 -> Average Cal sensor Temp.: 94.43 Ambient Temp.: 85.51 F
15:01:03.649 -> Average Cal sensor Temp.: 94.53 Ambient Temp.: 85.44 F
15:01:05.199 -> Average Cal sensor Temp.: 94.60 Ambient Temp.: 85.48 F
15:01:06.755 -> Average Cal sensor Temp.: 94.93 Ambient Temp.: 85.48 F
15:01:08.310 -> Average Cal sensor Temp.: 94.64 Ambient Temp.: 85.44 F
15:01:09.857 -> Average Cal sensor Temp.: 94.77 Ambient Temp.: 85.48 F
15:01:11.417 -> Average Cal sensor Temp.: 94.61 Ambient Temp.: 85.48 F
15:01:12.974 -> Average Cal sensor Temp.: 94.74 Ambient Temp.: 85.51 F
15:01:14.524 -> Average Cal sensor Temp.: 94.57 Ambient Temp.: 85.44 F
15:01:16.079 -> Average Cal sensor Temp.: 94.10 Ambient Temp.: 85.48 F
15:01:17.636 -> Average Cal sensor Temp.: 94.35 Ambient Temp.: 85.44 F
15:01:19.185 -> Average Cal sensor Temp.: 94.08 Ambient Temp.: 85.41 F
15:01:20.772 -> Average Cal sensor Temp.: 94.32 Ambient Temp.: 85.48 F
15:01:22.324 -> Average Cal sensor Temp.: 94.15 Ambient Temp.: 85.44 F
15:01:23.875 -> Average Cal sensor Temp.: 94.22 Ambient Temp.: 85.37 F
15:01:25.426 -> Average Cal sensor Temp.: 94.10 Ambient Temp.: 85.41 F
15:01:26.976 -> Average Cal sensor Temp.: 94.16 Ambient Temp.: 85.33 F
15:01:28.533 -> Average Cal sensor Temp.: 94.03 Ambient Temp.: 85.33 F
15:01:30.079 -> Average Cal sensor Temp.: 94.44 Ambient Temp.: 85.30 F
15:01:31.633 -> Average Cal sensor Temp.: 94.72 Ambient Temp.: 85.30 F
15:01:33.188 -> Average Cal sensor Temp.: 94.52 Ambient Temp.: 85.37 F
15:01:34.748 -> Average Cal sensor Temp.: 94.25 Ambient Temp.: 85.30 F
15:01:36.310 -> Average Cal sensor Temp.: 94.46 Ambient Temp.: 85.26 F
15:01:37.864 -> Average Cal sensor Temp.: 94.83 Ambient Temp.: 85.26 F
15:01:39.413 -> Average Cal sensor Temp.: 94.49 Ambient Temp.: 85.26 F
15:01:40.997 -> Average Cal sensor Temp.: 94.84 Ambient Temp.: 85.26 F
15:01:42.545 -> Average Cal sensor Temp.: 94.47 Ambient Temp.: 85.15 F
15:01:44.095 -> Average Cal sensor Temp.: 94.28 Ambient Temp.: 85.15 F
15:01:45.647 -> Average Cal sensor Temp.: 94.48 Ambient Temp.: 85.26 F
15:01:47.197 -> Average Cal sensor Temp.: 94.51 Ambient Temp.: 85.19 F
15:01:48.748 -> Average Cal sensor Temp.: 94.31 Ambient Temp.: 85.15 F
15:01:50.303 -> Average Cal sensor Temp.: 94.65 Ambient Temp.: 85.15 F
15:01:51.857 -> Average Cal sensor Temp.: 94.32 Ambient Temp.: 85.15 F
15:01:53.443 -> Average Cal sensor Temp.: 94.48 Ambient Temp.: 85.15 F
15:01:54.997 -> Average Cal sensor Temp.: 94.74 Ambient Temp.: 85.15 F
15:01:56.551 -> Average Cal sensor Temp.: 94.38 Ambient Temp.: 85.15 F
15:01:58.102 -> Average Cal sensor Temp.: 94.59 Ambient Temp.: 85.12 F
15:01:59.656 -> Average Cal sensor Temp.: 94.24 Ambient Temp.: 85.08 F
15:02:01.213 -> Average Cal sensor Temp.: 94.23 Ambient Temp.: 85.05 F
15:02:02.768 -> Average Cal sensor Temp.: 94.29 Ambient Temp.: 85.05 F
15:02:04.321 -> Average Cal sensor Temp.: 94.18 Ambient Temp.: 85.08 F
15:02:05.878 -> Average Cal sensor Temp.: 94.68 Ambient Temp.: 85.08 F
15:02:07.434 -> Average Cal sensor Temp.: 94.51 Ambient Temp.: 85.12 F
15:02:08.993 -> Average Cal sensor Temp.: 94.61 Ambient Temp.: 85.05 F
15:02:10.548 -> Average Cal sensor Temp.: 94.47 Ambient Temp.: 85.05 F
15:02:12.100 -> Average Cal sensor Temp.: 94.52 Ambient Temp.: 85.05 F
15:02:13.653 -> Average Cal sensor Temp.: 94.46 Ambient Temp.: 85.05 F
15:02:15.202 -> Average Cal sensor Temp.: 94.52 Ambient Temp.: 85.05 F
15:02:16.755 -> Average Cal sensor Temp.: 94.34 Ambient Temp.: 85.01 F
15:02:18.343 -> Average Cal sensor Temp.: 94.27 Ambient Temp.: 85.01 F
15:02:19.895 -> Average Cal sensor Temp.: 94.24 Ambient Temp.: 85.01 F
15:02:21.450 -> Average Cal sensor Temp.: 94.15 Ambient Temp.: 85.01 F
15:02:23.002 -> Average Cal sensor Temp.: 94.17 Ambient Temp.: 84.97 F
15:02:24.557 -> Average Cal sensor Temp.: 94.58 Ambient Temp.: 84.97 F
15:02:26.110 -> Average Cal sensor Temp.: 94.14 Ambient Temp.: 84.94 F
15:02:27.662 -> Average Cal sensor Temp.: 94.07 Ambient Temp.: 84.97 F
15:02:29.221 -> Average Cal sensor Temp.: 94.55 Ambient Temp.: 84.94 F
15:02:30.776 -> Average Cal sensor Temp.: 94.35 Ambient Temp.: 84.90 F
15:02:32.330 -> Average Cal sensor Temp.: 94.50 Ambient Temp.: 84.94 F
15:02:33.885 -> Average Cal sensor Temp.: 94.44 Ambient Temp.: 84.87 F
15:02:35.439 -> Average Cal sensor Temp.: 94.82 Ambient Temp.: 84.97 F
15:02:36.989 -> Average Cal sensor Temp.: 94.74 Ambient Temp.: 84.87 F
15:02:38.544 -> Average Cal sensor Temp.: 94.50 Ambient Temp.: 84.87 F
15:02:40.125 -> Average Cal sensor Temp.: 94.39 Ambient Temp.: 84.87 F
15:02:41.681 -> Average Cal sensor Temp.: 93.86 Ambient Temp.: 84.79 F
15:02:43.233 -> Average Cal sensor Temp.: 94.22 Ambient Temp.: 84.76 F
15:02:44.783 -> Average Cal sensor Temp.: 94.42 Ambient Temp.: 84.76 F
15:02:46.335 -> Average Cal sensor Temp.: 94.31 Ambient Temp.: 84.76 F
15:02:47.884 -> Average Cal sensor Temp.: 94.44 Ambient Temp.: 84.79 F
15:02:49.466 -> Average Cal sensor Temp.: 94.35 Ambient Temp.: 84.72 F
15:02:51.012 -> Average Cal sensor Temp.: 94.45 Ambient Temp.: 84.72 F
15:02:52.565 -> Average Cal sensor Temp.: 94.37 Ambient Temp.: 84.76 F
15:02:54.123 -> Average Cal sensor Temp.: 94.33 Ambient Temp.: 84.72 F
15:02:55.671 -> Average Cal sensor Temp.: 94.50 Ambient Temp.: 84.76 F
15:02:57.222 -> Average Cal sensor Temp.: 94.50 Ambient Temp.: 84.72 F
15:02:58.778 -> Average Cal sensor Temp.: 94.35 Ambient Temp.: 84.72 F
15:03:00.334 -> Average Cal sensor Temp.: 94.41 Ambient Temp.: 84.65 F
15:03:01.920 -> Average Cal sensor Temp.: 94.61 Ambient Temp.: 84.72 F
15:03:03.478 -> Average Cal sensor Temp.: 94.55 Ambient Temp.: 84.69 F
15:03:05.030 -> Average Cal sensor Temp.: 94.66 Ambient Temp.: 84.72 F
15:03:06.588 -> Average Cal sensor Temp.: 94.96 Ambient Temp.: 84.69 F
15:03:08.140 -> Average Cal sensor Temp.: 95.01 Ambient Temp.: 84.69 F
15:03:09.693 -> Average Cal sensor Temp.: 94.80 Ambient Temp.: 84.76 F
15:03:11.245 -> Average Cal sensor Temp.: 94.91 Ambient Temp.: 84.72 F
15:03:12.800 -> Average Cal sensor Temp.: 94.65 Ambient Temp.: 84.65 F
15:03:14.352 -> Average Cal sensor Temp.: 94.74 Ambient Temp.: 84.61 F
15:03:15.902 -> Average Cal sensor Temp.: 94.54 Ambient Temp.: 84.58 F
15:03:17.490 -> Average Cal sensor Temp.: 94.49 Ambient Temp.: 84.51 F
15:03:19.047 -> Average Cal sensor Temp.: 94.55 Ambient Temp.: 84.51 F
15:03:20.569 -> Average Cal sensor Temp.: 94.55 Ambient Temp.: 84.54 F
15:03:22.156 -> Average Cal sensor Temp.: 94.45 Ambient Temp.: 84.51 F
15:03:23.713 -> Average Cal sensor Temp.: 94.27 Ambient Temp.: 84.54 F
15:03:25.264 -> Average Cal sensor Temp.: 94.23 Ambient Temp.: 84.47 F
15:03:26.822 -> Average Cal sensor Temp.: 94.24 Ambient Temp.: 84.47 F
15:03:28.377 -> Average Cal sensor Temp.: 94.82 Ambient Temp.: 84.51 F
15:03:29.927 -> Average Cal sensor Temp.: 94.41 Ambient Temp.: 84.47 F
15:03:31.477 -> Average Cal sensor Temp.: 94.46 Ambient Temp.: 84.47 F
15:03:33.025 -> Average Cal sensor Temp.: 94.24 Ambient Temp.: 84.36 F
15:03:34.583 -> Average Cal sensor Temp.: 93.97 Ambient Temp.: 84.40 F
15:03:36.134 -> Average Cal sensor Temp.: 94.18 Ambient Temp.: 84.36 F
15:03:37.721 -> Average Cal sensor Temp.: 94.20 Ambient Temp.: 84.33 F
15:03:39.271 -> Average Cal sensor Temp.: 94.29 Ambient Temp.: 84.33 F
15:03:40.821 -> Average Cal sensor Temp.: 94.28 Ambient Temp.: 84.25 F
15:03:42.368 -> Average Cal sensor Temp.: 94.40 Ambient Temp.: 84.29 F
15:03:43.924 -> Average Cal sensor Temp.: 94.92 Ambient Temp.: 84.25 F
15:03:45.507 -> Average Cal sensor Temp.: 94.23 Ambient Temp.: 84.29 F
15:03:47.058 -> Average Cal sensor Temp.: 94.05 Ambient Temp.: 84.29 F
15:03:48.605 -> Average Cal sensor Temp.: 94.39 Ambient Temp.: 84.22 F
15:03:50.155 -> Average Cal sensor Temp.: 94.52 Ambient Temp.: 84.22 F
15:03:51.708 -> Average Cal sensor Temp.: 94.56 Ambient Temp.: 84.22 F
15:03:53.288 -> Average Cal sensor Temp.: 94.35 Ambient Temp.: 84.22 F
15:03:54.836 -> Average Cal sensor Temp.: 94.38 Ambient Temp.: 84.18 F
15:03:56.387 -> Average Cal sensor Temp.: 94.73 Ambient Temp.: 84.22 F
15:03:57.944 -> Average Cal sensor Temp.: 94.74 Ambient Temp.: 84.18 F
15:03:59.492 -> Average Cal sensor Temp.: 94.65 Ambient Temp.: 84.22 F
15:04:01.042 -> Average Cal sensor Temp.: 94.61 Ambient Temp.: 84.18 F
15:04:02.594 -> Average Cal sensor Temp.: 94.45 Ambient Temp.: 84.18 F
15:04:04.150 -> Average Cal sensor Temp.: 94.07 Ambient Temp.: 84.11 F
15:04:05.705 -> Average Cal sensor Temp.: 94.14 Ambient Temp.: 84.15 F
15:04:07.285 -> Average Cal sensor Temp.: 94.39 Ambient Temp.: 84.11 F
15:04:08.835 -> Average Cal sensor Temp.: 94.32 Ambient Temp.: 84.11 F
15:04:10.387 -> Average Cal sensor Temp.: 94.45 Ambient Temp.: 84.07 F
15:04:11.944 -> Average Cal sensor Temp.: 94.51 Ambient Temp.: 84.07 F
15:04:13.495 -> Average Cal sensor Temp.: 94.18 Ambient Temp.: 84.11 F
15:04:15.046 -> Average Cal sensor Temp.: 94.31 Ambient Temp.: 83.97 F
15:04:16.605 -> Average Cal sensor Temp.: 94.73 Ambient Temp.: 84.07 F
15:04:18.153 -> Average Cal sensor Temp.: 94.76 Ambient Temp.: 84.00 F
15:04:19.732 -> Average Cal sensor Temp.: 94.65 Ambient Temp.: 84.07 F
15:04:21.285 -> Average Cal sensor Temp.: 94.70 Ambient Temp.: 84.00 F
15:04:22.835 -> Average Cal sensor Temp.: 94.66 Ambient Temp.: 83.97 F
15:04:24.387 -> Average Cal sensor Temp.: 94.44 Ambient Temp.: 83.97 F
15:04:25.935 -> Average Cal sensor Temp.: 94.37 Ambient Temp.: 84.00 F
15:04:27.524 -> Average Cal sensor Temp.: 94.23 Ambient Temp.: 83.97 F
15:04:29.077 -> Average Cal sensor Temp.: 94.38 Ambient Temp.: 83.97 F
15:04:30.635 -> Average Cal sensor Temp.: 94.36 Ambient Temp.: 83.93 F
15:04:32.190 -> Average Cal sensor Temp.: 94.76 Ambient Temp.: 83.97 F
15:04:33.747 -> Average Cal sensor Temp.: 94.55 Ambient Temp.: 83.93 F
15:04:35.299 -> Average Cal sensor Temp.: 94.61 Ambient Temp.: 83.93 F
15:04:36.854 -> Average Cal sensor Temp.: 94.52 Ambient Temp.: 83.89 F
15:04:38.407 -> Average Cal sensor Temp.: 94.41 Ambient Temp.: 83.89 F
15:04:39.965 -> Average Cal sensor Temp.: 94.65 Ambient Temp.: 83.89 F
15:04:41.519 -> Average Cal sensor Temp.: 94.61 Ambient Temp.: 83.93 F
15:04:43.074 -> Average Cal sensor Temp.: 94.26 Ambient Temp.: 83.86 F
15:04:44.622 -> Average Cal sensor Temp.: 94.68 Ambient Temp.: 83.89 F
15:04:46.173 -> Average Cal sensor Temp.: 93.97 Ambient Temp.: 83.89 F
15:04:47.728 -> Average Cal sensor Temp.: 93.81 Ambient Temp.: 83.93 F
15:04:49.314 -> Average Cal sensor Temp.: 94.17 Ambient Temp.: 83.86 F
15:04:50.868 -> Average Cal sensor Temp.: 94.13 Ambient Temp.: 83.82 F
15:04:52.422 -> Average Cal sensor Temp.: 94.17 Ambient Temp.: 83.82 F
15:04:53.975 -> Average Cal sensor Temp.: 94.45 Ambient Temp.: 83.82 F
15:04:55.521 -> Average Cal sensor Temp.: 94.35 Ambient Temp.: 83.79 F
15:04:57.078 -> Average Cal sensor Temp.: 94.85 Ambient Temp.: 83.82 F
15:04:58.629 -> Average Cal sensor Temp.: 94.40 Ambient Temp.: 83.79 F
15:05:00.180 -> Average Cal sensor Temp.: 94.61 Ambient Temp.: 83.79 F
15:05:01.764 -> Average Cal sensor Temp.: 94.47 Ambient Temp.: 83.79 F
15:05:03.321 -> Average Cal sensor Temp.: 94.32 Ambient Temp.: 83.79 F
15:05:04.873 -> Average Cal sensor Temp.: 94.50 Ambient Temp.: 83.79 F
15:05:06.432 -> Average Cal sensor Temp.: 94.67 Ambient Temp.: 83.75 F
15:05:07.989 -> Average Cal sensor Temp.: 94.63 Ambient Temp.: 83.71 F
15:05:09.543 -> Average Cal sensor Temp.: 94.73 Ambient Temp.: 83.79 F
15:05:11.098 -> Average Cal sensor Temp.: 94.41 Ambient Temp.: 83.79 F
15:05:12.652 -> Average Cal sensor Temp.: 94.30 Ambient Temp.: 83.75 F
15:05:14.204 -> Average Cal sensor Temp.: 94.34 Ambient Temp.: 83.79 F
15:05:15.756 -> Average Cal sensor Temp.: 94.58 Ambient Temp.: 83.79 F
15:05:17.313 -> Average Cal sensor Temp.: 94.81 Ambient Temp.: 83.79 F
15:05:18.868 -> Average Cal sensor Temp.: 94.84 Ambient Temp.: 83.82 F
15:05:20.424 -> Average Cal sensor Temp.: 94.81 Ambient Temp.: 83.82 F
15:05:21.978 -> Average Cal sensor Temp.: 94.68 Ambient Temp.: 83.86 F
15:05:23.527 -> Average Cal sensor Temp.: 95.14 Ambient Temp.: 83.86 F
15:05:25.078 -> Average Cal sensor Temp.: 95.15 Ambient Temp.: 83.89 F
15:05:26.635 -> Average Cal sensor Temp.: 94.69 Ambient Temp.: 83.86 F
15:05:28.221 -> Average Cal sensor Temp.: 94.59 Ambient Temp.: 83.86 F
15:05:29.771 -> Average Cal sensor Temp.: 94.71 Ambient Temp.: 83.86 F
15:05:31.323 -> Average Cal sensor Temp.: 94.66 Ambient Temp.: 83.89 F
15:05:32.875 -> Average Cal sensor Temp.: 94.33 Ambient Temp.: 83.89 F
15:05:34.432 -> Average Cal sensor Temp.: 94.44 Ambient Temp.: 83.86 F
15:05:35.985 -> Average Cal sensor Temp.: 94.47 Ambient Temp.: 83.89 F
15:05:37.542 -> Average Cal sensor Temp.: 94.30 Ambient Temp.: 83.93 F
15:05:39.123 -> Average Cal sensor Temp.: 94.70 Ambient Temp.: 83.93 F
15:05:40.676 -> Average Cal sensor Temp.: 94.76 Ambient Temp.: 83.89 F
15:05:42.229 -> Average Cal sensor Temp.: 94.72 Ambient Temp.: 83.97 F
15:05:43.783 -> Average Cal sensor Temp.: 94.63 Ambient Temp.: 83.89 F
15:05:45.331 -> Average Cal sensor Temp.: 94.77 Ambient Temp.: 83.93 F
15:05:46.892 -> Average Cal sensor Temp.: 94.64 Ambient Temp.: 84.00 F
15:05:48.440 -> Average Cal sensor Temp.: 94.25 Ambient Temp.: 83.97 F
15:05:49.990 -> Average Cal sensor Temp.: 94.56 Ambient Temp.: 84.00 F
15:05:51.544 -> Average Cal sensor Temp.: 94.69 Ambient Temp.: 83.93 F
15:05:53.131 -> Average Cal sensor Temp.: 94.75 Ambient Temp.: 83.97 F
15:05:54.687 -> Average Cal sensor Temp.: 94.42 Ambient Temp.: 84.00 F
15:05:56.243 -> Average Cal sensor Temp.: 94.59 Ambient Temp.: 84.00 F
15:05:57.790 -> Average Cal sensor Temp.: 94.92 Ambient Temp.: 83.93 F
15:05:59.352 -> Average Cal sensor Temp.: 94.58 Ambient Temp.: 84.00 F
15:06:00.901 -> Average Cal sensor Temp.: 94.28 Ambient Temp.: 84.00 F
15:06:02.460 -> Average Cal sensor Temp.: 94.22 Ambient Temp.: 84.00 F
15:06:04.015 -> Average Cal sensor Temp.: 94.41 Ambient Temp.: 84.00 F
15:06:05.566 -> Average Cal sensor Temp.: 94.58 Ambient Temp.: 84.11 F
15:06:07.123 -> Average Cal sensor Temp.: 94.91 Ambient Temp.: 84.11 F
15:06:08.678 -> Average Cal sensor Temp.: 94.58 Ambient Temp.: 84.11 F
15:06:10.232 -> Average Cal sensor Temp.: 94.60 Ambient Temp.: 84.07 F
15:06:11.782 -> Average Cal sensor Temp.: 94.64 Ambient Temp.: 84.00 F
15:06:13.339 -> Average Cal sensor Temp.: 94.42 Ambient Temp.: 84.11 F
15:06:14.927 -> Average Cal sensor Temp.: 94.85 Ambient Temp.: 84.00 F
15:06:16.480 -> Average Cal sensor Temp.: 94.60 Ambient Temp.: 84.07 F
15:06:18.031 -> Average Cal sensor Temp.: 94.66 Ambient Temp.: 84.07 F
15:06:19.589 -> Average Cal sensor Temp.: 94.70 Ambient Temp.: 84.11 F
15:06:21.137 -> Average Cal sensor Temp.: 94.82 Ambient Temp.: 84.11 F
15:06:22.690 -> Average Cal sensor Temp.: 94.83 Ambient Temp.: 84.15 F
15:06:24.243 -> Average Cal sensor Temp.: 94.94 Ambient Temp.: 84.18 F
15:06:25.793 -> Average Cal sensor Temp.: 94.96 Ambient Temp.: 84.15 F
15:06:27.379 -> Average Cal sensor Temp.: 94.89 Ambient Temp.: 84.15 F
15:06:28.930 -> Average Cal sensor Temp.: 94.69 Ambient Temp.: 84.18 F
15:06:30.481 -> Average Cal sensor Temp.: 94.79 Ambient Temp.: 84.15 F
15:06:32.036 -> Average Cal sensor Temp.: 94.74 Ambient Temp.: 84.15 F
15:06:33.592 -> Average Cal sensor Temp.: 94.75 Ambient Temp.: 84.15 F
15:06:35.153 -> Average Cal sensor Temp.: 94.73 Ambient Temp.: 84.15 F
15:06:36.705 -> Average Cal sensor Temp.: 94.76 Ambient Temp.: 84.18 F
15:06:37.293 -> Ambient Temp.: 84.18 F
15:06:37.880 -> Ambient Temp.: 84.18 F
15:06:38.468 -> Ambient Temp.: 84.18 F
15:06:39.747 -> Ambient Temp.: 84.22 F
15:06:40.336 -> Ambient Temp.: 84.22 F
15:06:41.890 -> Average Cal sensor Temp.: 95.01 Ambient Temp.: 84.22 F
15:06:43.444 -> Average Cal sensor Temp.: 94.39 Ambient Temp.: 84.25 F
15:06:44.990 -> Average Cal sensor Temp.: 94.51 Ambient Temp.: 84.18 F
15:06:46.550 -> Average Cal sensor Temp.: 94.63 Ambient Temp.: 84.22 F
15:06:48.109 -> Average Cal sensor Temp.: 94.74 Ambient Temp.: 84.25 F
15:06:49.664 -> Average Cal sensor Temp.: 94.47 Ambient Temp.: 84.25 F
15:06:51.215 -> Average Cal sensor Temp.: 94.56 Ambient Temp.: 84.29 F
15:06:52.767 -> Average Cal sensor Temp.: 94.46 Ambient Temp.: 84.29 F
15:06:54.325 -> Average Cal sensor Temp.: 94.66 Ambient Temp.: 84.29 F
15:06:55.910 -> Average Cal sensor Temp.: 94.64 Ambient Temp.: 84.33 F
15:06:57.459 -> Average Cal sensor Temp.: 94.38 Ambient Temp.: 84.33 F
15:06:59.012 -> Average Cal sensor Temp.: 94.43 Ambient Temp.: 84.33 F
15:07:00.570 -> Average Cal sensor Temp.: 94.96 Ambient Temp.: 84.33 F
15:07:02.124 -> Average Cal sensor Temp.: 94.80 Ambient Temp.: 84.40 F
15:07:03.677 -> Average Cal sensor Temp.: 94.72 Ambient Temp.: 84.47 F
15:07:05.225 -> Average Cal sensor Temp.: 94.71 Ambient Temp.: 84.47 F
15:07:06.780 -> Average Cal sensor Temp.: 94.92 Ambient Temp.: 84.51 F
15:07:08.337 -> Average Cal sensor Temp.: 94.90 Ambient Temp.: 84.47 F
15:07:09.922 -> Average Cal sensor Temp.: 94.79 Ambient Temp.: 84.47 F
15:07:11.481 -> Average Cal sensor Temp.: 94.65 Ambient Temp.: 84.51 F
15:07:13.036 -> Average Cal sensor Temp.: 94.82 Ambient Temp.: 84.51 F
15:07:14.587 -> Average Cal sensor Temp.: 94.70 Ambient Temp.: 84.61 F
15:07:16.140 -> Average Cal sensor Temp.: 94.75 Ambient Temp.: 84.58 F
15:07:17.690 -> Average Cal sensor Temp.: 94.86 Ambient Temp.: 84.58 F
15:07:19.243 -> Average Cal sensor Temp.: 95.08 Ambient Temp.: 84.69 F
15:07:20.803 -> Average Cal sensor Temp.: 94.75 Ambient Temp.: 84.61 F
15:07:22.354 -> Average Cal sensor Temp.: 94.87 Ambient Temp.: 84.69 F
15:07:23.910 -> Average Cal sensor Temp.: 94.82 Ambient Temp.: 84.76 F
15:07:25.466 -> Average Cal sensor Temp.: 95.09 Ambient Temp.: 84.79 F
15:07:27.025 -> Average Cal sensor Temp.: 94.87 Ambient Temp.: 84.76 F
15:07:28.583 -> Average Cal sensor Temp.: 94.97 Ambient Temp.: 84.76 F
15:07:30.138 -> Average Cal sensor Temp.: 94.75 Ambient Temp.: 84.79 F
15:07:31.696 -> Average Cal sensor Temp.: 94.67 Ambient Temp.: 84.76 F
15:07:33.252 -> Average Cal sensor Temp.: 94.78 Ambient Temp.: 84.87 F
15:07:34.810 -> Average Cal sensor Temp.: 94.86 Ambient Temp.: 84.90 F
15:07:36.365 -> Average Cal sensor Temp.: 94.99 Ambient Temp.: 84.97 F
15:07:37.920 -> Average Cal sensor Temp.: 94.89 Ambient Temp.: 84.97 F
15:07:39.508 -> Average Cal sensor Temp.: 94.84 Ambient Temp.: 84.97 F
15:07:41.065 -> Average Cal sensor Temp.: 95.04 Ambient Temp.: 85.01 F
15:07:42.620 -> Average Cal sensor Temp.: 94.90 Ambient Temp.: 84.97 F
15:07:44.175 -> Average Cal sensor Temp.: 94.42 Ambient Temp.: 85.01 F
15:07:45.732 -> Average Cal sensor Temp.: 94.15 Ambient Temp.: 85.05 F
15:07:47.285 -> Average Cal sensor Temp.: 94.52 Ambient Temp.: 85.08 F
15:07:48.837 -> Average Cal sensor Temp.: 94.63 Ambient Temp.: 85.12 F
15:07:50.389 -> Average Cal sensor Temp.: 94.68 Ambient Temp.: 85.08 F
15:07:51.941 -> Average Cal sensor Temp.: 95.36 Ambient Temp.: 85.15 F
15:07:53.492 -> Average Cal sensor Temp.: 94.68 Ambient Temp.: 85.15 F
15:07:55.045 -> Average Cal sensor Temp.: 94.80 Ambient Temp.: 85.19 F
15:07:56.603 -> Average Cal sensor Temp.: 94.76 Ambient Temp.: 85.19 F
15:07:58.184 -> Average Cal sensor Temp.: 94.91 Ambient Temp.: 85.19 F
15:07:59.740 -> Average Cal sensor Temp.: 95.01 Ambient Temp.: 85.19 F
15:08:01.297 -> Average Cal sensor Temp.: 94.96 Ambient Temp.: 85.30 F
15:08:02.850 -> Average Cal sensor Temp.: 94.97 Ambient Temp.: 85.30 F
15:08:04.405 -> Average Cal sensor Temp.: 95.01 Ambient Temp.: 85.33 F
15:08:05.960 -> Average Cal sensor Temp.: 94.92 Ambient Temp.: 85.30 F
15:08:07.507 -> Average Cal sensor Temp.: 94.95 Ambient Temp.: 85.33 F
15:08:09.055 -> Average Cal sensor Temp.: 94.73 Ambient Temp.: 85.33 F
15:08:10.612 -> Average Cal sensor Temp.: 94.50 Ambient Temp.: 85.37 F
15:08:12.167 -> Average Cal sensor Temp.: 94.65 Ambient Temp.: 85.30 F
15:08:13.719 -> Average Cal sensor Temp.: 94.45 Ambient Temp.: 85.33 F
15:08:15.301 -> Average Cal sensor Temp.: 94.95 Ambient Temp.: 85.37 F
15:08:16.855 -> Average Cal sensor Temp.: 94.81 Ambient Temp.: 85.33 F
15:08:18.411 -> Average Cal sensor Temp.: 94.94 Ambient Temp.: 85.37 F
15:08:19.936 -> Average Cal sensor Temp.: 94.37 Ambient Temp.: 85.41 F
15:08:21.523 -> Average Cal sensor Temp.: 94.57 Ambient Temp.: 85.37 F
15:08:23.047 -> Average Cal sensor Temp.: 95.13 Ambient Temp.: 85.37 F
15:08:24.630 -> Average Cal sensor Temp.: 94.94 Ambient Temp.: 85.41 F
15:08:26.181 -> Average Cal sensor Temp.: 94.95 Ambient Temp.: 85.44 F
15:08:27.736 -> Average Cal sensor Temp.: 95.02 Ambient Temp.: 85.37 F
15:08:29.293 -> Average Cal sensor Temp.: 94.70 Ambient Temp.: 85.41 F
15:08:30.850 -> Average Cal sensor Temp.: 94.77 Ambient Temp.: 85.33 F
15:08:32.408 -> Average Cal sensor Temp.: 94.83 Ambient Temp.: 85.41 F
15:08:33.964 -> Average Cal sensor Temp.: 94.57 Ambient Temp.: 85.41 F
15:08:35.523 -> Average Cal sensor Temp.: 94.80 Ambient Temp.: 85.41 F
15:08:37.078 -> Average Cal sensor Temp.: 94.95 Ambient Temp.: 85.41 F
15:08:38.632 -> Average Cal sensor Temp.: 94.71 Ambient Temp.: 85.44 F
15:08:40.183 -> Average Cal sensor Temp.: 94.91 Ambient Temp.: 85.37 F
15:08:41.738 -> Average Cal sensor Temp.: 94.85 Ambient Temp.: 85.44 F
15:08:43.294 -> Average Cal sensor Temp.: 94.59 Ambient Temp.: 85.44 F
15:08:44.845 -> Average Cal sensor Temp.: 94.79 Ambient Temp.: 85.44 F
15:08:46.398 -> Average Cal sensor Temp.: 95.03 Ambient Temp.: 85.44 F
15:08:47.954 -> Average Cal sensor Temp.: 94.84 Ambient Temp.: 85.44 F
15:08:49.506 -> Average Cal sensor Temp.: 94.66 Ambient Temp.: 85.44 F
15:08:51.057 -> Average Cal sensor Temp.: 94.88 Ambient Temp.: 85.44 F
15:08:52.644 -> Average Cal sensor Temp.: 94.86 Ambient Temp.: 85.44 F
15:08:54.199 -> Average Cal sensor Temp.: 94.70 Ambient Temp.: 85.44 F
15:08:55.750 -> Average Cal sensor Temp.: 94.93 Ambient Temp.: 85.44 F
15:08:57.305 -> Average Cal sensor Temp.: 94.89 Ambient Temp.: 85.44 F
15:08:58.858 -> Average Cal sensor Temp.: 95.17 Ambient Temp.: 85.51 F
15:09:00.404 -> Average Cal sensor Temp.: 94.56 Ambient Temp.: 85.44 F
15:09:01.955 -> Average Cal sensor Temp.: 94.50 Ambient Temp.: 85.44 F
15:09:03.513 -> Average Cal sensor Temp.: 94.32 Ambient Temp.: 85.41 F
15:09:05.069 -> Average Cal sensor Temp.: 94.64 Ambient Temp.: 85.41 F
15:09:06.656 -> Average Cal sensor Temp.: 94.48 Ambient Temp.: 85.41 F
15:09:08.204 -> Average Cal sensor Temp.: 94.99 Ambient Temp.: 85.44 F
15:09:09.756 -> Average Cal sensor Temp.: 94.87 Ambient Temp.: 85.44 F
15:09:11.308 -> Average Cal sensor Temp.: 94.96 Ambient Temp.: 85.48 F
15:09:12.861 -> Average Cal sensor Temp.: 94.99 Ambient Temp.: 85.48 F
15:09:14.443 -> Average Cal sensor Temp.: 94.98 Ambient Temp.: 85.44 F
15:09:16.000 -> Average Cal sensor Temp.: 94.96 Ambient Temp.: 85.41 F
15:09:17.550 -> Average Cal sensor Temp.: 94.59 Ambient Temp.: 85.41 F
15:09:19.104 -> Average Cal sensor Temp.: 95.01 Ambient Temp.: 85.37 F
15:09:20.657 -> Average Cal sensor Temp.: 95.17 Ambient Temp.: 85.37 F
15:09:22.204 -> Average Cal sensor Temp.: 94.45 Ambient Temp.: 85.37 F
15:09:23.760 -> Average Cal sensor Temp.: 94.68 Ambient Temp.: 85.30 F
15:09:25.319 -> Average Cal sensor Temp.: 94.67 Ambient Temp.: 85.26 F
15:09:26.876 -> Average Cal sensor Temp.: 94.87 Ambient Temp.: 85.30 F
15:09:28.431 -> Average Cal sensor Temp.: 94.93 Ambient Temp.: 85.26 F
15:09:29.984 -> Average Cal sensor Temp.: 95.08 Ambient Temp.: 85.26 F
15:09:31.542 -> Average Cal sensor Temp.: 95.44 Ambient Temp.: 85.26 F
15:09:33.092 -> Average Cal sensor Temp.: 95.00 Ambient Temp.: 85.30 F
15:09:34.645 -> Average Cal sensor Temp.: 94.93 Ambient Temp.: 85.30 F
15:09:36.201 -> Average Cal sensor Temp.: 94.86 Ambient Temp.: 85.26 F
15:09:37.757 -> Average Cal sensor Temp.: 94.63 Ambient Temp.: 85.26 F
15:09:39.342 -> Average Cal sensor Temp.: 94.94 Ambient Temp.: 85.26 F
15:09:40.896 -> Average Cal sensor Temp.: 94.80 Ambient Temp.: 85.26 F
15:09:42.443 -> Average Cal sensor Temp.: 94.89 Ambient Temp.: 85.26 F
15:09:44.000 -> Average Cal sensor Temp.: 94.85 Ambient Temp.: 85.19 F
15:09:45.556 -> Average Cal sensor Temp.: 95.04 Ambient Temp.: 85.19 F
15:09:47.106 -> Average Cal sensor Temp.: 94.69 Ambient Temp.: 85.15 F
15:09:48.656 -> Average Cal sensor Temp.: 95.13 Ambient Temp.: 85.15 F
15:09:50.215 -> Average Cal sensor Temp.: 94.96 Ambient Temp.: 85.12 F
15:09:51.771 -> Average Cal sensor Temp.: 94.91 Ambient Temp.: 85.08 F
15:09:53.324 -> Average Cal sensor Temp.: 94.37 Ambient Temp.: 85.08 F
15:09:54.879 -> Average Cal sensor Temp.: 94.21 Ambient Temp.: 85.08 F
15:09:56.434 -> Average Cal sensor Temp.: 94.46 Ambient Temp.: 85.05 F
15:09:58.021 -> Average Cal sensor Temp.: 94.11 Ambient Temp.: 84.97 F
15:09:59.572 -> Average Cal sensor Temp.: 94.62 Ambient Temp.: 84.94 F
15:10:01.127 -> Average Cal sensor Temp.: 94.95 Ambient Temp.: 84.97 F
15:10:02.677 -> Average Cal sensor Temp.: 94.86 Ambient Temp.: 84.97 F
15:10:04.231 -> Average Cal sensor Temp.: 94.84 Ambient Temp.: 84.94 F
15:10:05.790 -> Average Cal sensor Temp.: 94.96 Ambient Temp.: 84.90 F
15:10:07.347 -> Average Cal sensor Temp.: 94.90 Ambient Temp.: 84.87 F
15:10:08.910 -> Average Cal sensor Temp.: 94.28 Ambient Temp.: 84.79 F
15:10:10.469 -> Average Cal sensor Temp.: 94.60 Ambient Temp.: 84.79 F
15:10:12.021 -> Average Cal sensor Temp.: 94.85 Ambient Temp.: 84.76 F
15:10:13.090 -> Ambient Temp.: 84.76 F
15:10:13.674 -> Ambient Temp.: 84.76 F
15:10:15.231 -> Average Cal sensor Temp.: 95.13 Ambient Temp.: 84.76 F
15:10:16.786 -> Average Cal sensor Temp.: 94.73 Ambient Temp.: 84.76 F
15:10:18.341 -> Average Cal sensor Temp.: 94.90 Ambient Temp.: 84.72 F
15:10:19.901 -> Average Cal sensor Temp.: 94.89 Ambient Temp.: 84.72 F
15:10:21.457 -> Average Cal sensor Temp.: 94.55 Ambient Temp.: 84.65 F
15:10:23.007 -> Average Cal sensor Temp.: 94.50 Ambient Temp.: 84.72 F
15:10:24.560 -> Average Cal sensor Temp.: 94.79 Ambient Temp.: 84.65 F
15:10:26.115 -> Average Cal sensor Temp.: 94.86 Ambient Temp.: 84.65 F
15:10:27.670 -> Average Cal sensor Temp.: 94.85 Ambient Temp.: 84.61 F
15:10:28.639 -> Ambient Temp.: 84.65 F
15:10:29.225 -> Ambient Temp.: 84.65 F
15:10:29.810 -> Ambient Temp.: 84.65 F
15:10:30.364 -> Ambient Temp.: 84.65 F
15:10:30.950 -> Ambient Temp.: 84.65 F
15:10:32.503 -> Average Cal sensor Temp.: 94.71 Ambient Temp.: 84.65 F
15:10:34.058 -> Average Cal sensor Temp.: 94.77 Ambient Temp.: 84.69 F
15:10:35.607 -> Average Cal sensor Temp.: 94.44 Ambient Temp.: 84.58 F
15:10:37.162 -> Average Cal sensor Temp.: 94.70 Ambient Temp.: 84.61 F
15:10:38.720 -> Average Cal sensor Temp.: 94.69 Ambient Temp.: 84.58 F
15:10:40.275 -> Average Cal sensor Temp.: 94.38 Ambient Temp.: 84.61 F
15:10:41.862 -> Average Cal sensor Temp.: 94.41 Ambient Temp.: 84.51 F
15:10:43.415 -> Average Cal sensor Temp.: 94.48 Ambient Temp.: 84.51 F
15:10:44.971 -> Average Cal sensor Temp.: 94.86 Ambient Temp.: 84.47 F
15:10:46.525 -> Average Cal sensor Temp.: 94.73 Ambient Temp.: 84.51 F
15:10:48.075 -> Average Cal sensor Temp.: 94.81 Ambient Temp.: 84.47 F
15:10:49.629 -> Average Cal sensor Temp.: 94.59 Ambient Temp.: 84.47 F
15:10:51.179 -> Average Cal sensor Temp.: 94.23 Ambient Temp.: 84.36 F
15:10:52.731 -> Average Cal sensor Temp.: 94.60 Ambient Temp.: 84.40 F
15:10:54.317 -> Average Cal sensor Temp.: 94.71 Ambient Temp.: 84.47 F
15:10:55.839 -> Average Cal sensor Temp.: 94.55 Ambient Temp.: 84.33 F
15:10:57.394 -> Average Cal sensor Temp.: 94.38 Ambient Temp.: 84.36 F
15:10:58.976 -> Average Cal sensor Temp.: 94.59 Ambient Temp.: 84.36 F
15:11:00.532 -> Average Cal sensor Temp.: 94.95 Ambient Temp.: 84.33 F
15:11:02.085 -> Average Cal sensor Temp.: 94.99 Ambient Temp.: 84.33 F
15:11:03.643 -> Average Cal sensor Temp.: 94.74 Ambient Temp.: 84.29 F
15:11:05.190 -> Average Cal sensor Temp.: 94.77 Ambient Temp.: 84.29 F
15:11:06.743 -> Average Cal sensor Temp.: 95.02 Ambient Temp.: 84.33 F
15:11:08.296 -> Average Cal sensor Temp.: 95.26 Ambient Temp.: 84.33 F
15:11:09.851 -> Average Cal sensor Temp.: 95.36 Ambient Temp.: 84.33 F
15:11:11.413 -> Average Cal sensor Temp.: 95.12 Ambient Temp.: 84.29 F
15:11:12.970 -> Average Cal sensor Temp.: 95.47 Ambient Temp.: 84.25 F
15:11:14.521 -> Average Cal sensor Temp.: 94.94 Ambient Temp.: 84.25 F
15:11:16.079 -> Average Cal sensor Temp.: 95.26 Ambient Temp.: 84.22 F
15:11:17.629 -> Average Cal sensor Temp.: 95.28 Ambient Temp.: 84.29 F
15:11:19.215 -> Average Cal sensor Temp.: 94.62 Ambient Temp.: 84.22 F
15:11:20.771 -> Average Cal sensor Temp.: 94.72 Ambient Temp.: 84.18 F
15:11:22.326 -> Average Cal sensor Temp.: 95.19 Ambient Temp.: 84.18 F
15:11:23.853 -> Average Cal sensor Temp.: 95.41 Ambient Temp.: 84.15 F
15:11:25.410 -> Average Cal sensor Temp.: 95.45 Ambient Temp.: 84.11 F
15:11:26.972 -> Average Cal sensor Temp.: 95.17 Ambient Temp.: 84.11 F
15:11:28.524 -> Average Cal sensor Temp.: 95.32 Ambient Temp.: 84.11 F
15:11:30.079 -> Average Cal sensor Temp.: 95.48 Ambient Temp.: 84.15 F
15:11:31.632 -> Average Cal sensor Temp.: 95.64 Ambient Temp.: 84.11 F
15:11:33.185 -> Average Cal sensor Temp.: 94.87 Ambient Temp.: 84.00 F
15:11:34.740 -> Average Cal sensor Temp.: 94.87 Ambient Temp.: 84.00 F
15:11:36.296 -> Average Cal sensor Temp.: 94.61 Ambient Temp.: 84.07 F
15:11:37.851 -> Average Cal sensor Temp.: 94.69 Ambient Temp.: 83.89 F
15:11:39.440 -> Average Cal sensor Temp.: 94.76 Ambient Temp.: 83.82 F
15:11:40.966 -> Average Cal sensor Temp.: 94.96 Ambient Temp.: 83.86 F
15:11:42.551 -> Average Cal sensor Temp.: 94.90 Ambient Temp.: 83.89 F
15:11:44.105 -> Average Cal sensor Temp.: 94.71 Ambient Temp.: 83.82 F
15:11:45.655 -> Average Cal sensor Temp.: 94.59 Ambient Temp.: 83.89 F
15:11:47.213 -> Average Cal sensor Temp.: 94.67 Ambient Temp.: 83.82 F
15:11:48.765 -> Average Cal sensor Temp.: 94.91 Ambient Temp.: 83.82 F
15:11:50.317 -> Average Cal sensor Temp.: 94.78 Ambient Temp.: 83.86 F
15:11:51.872 -> Average Cal sensor Temp.: 94.99 Ambient Temp.: 83.82 F
15:11:53.425 -> Average Cal sensor Temp.: 94.71 Ambient Temp.: 83.79 F
15:11:54.980 -> Average Cal sensor Temp.: 94.99 Ambient Temp.: 83.79 F
15:11:56.539 -> Average Cal sensor Temp.: 94.69 Ambient Temp.: 83.75 F
15:11:58.126 -> Average Cal sensor Temp.: 94.55 Ambient Temp.: 83.71 F
15:11:59.674 -> Average Cal sensor Temp.: 94.76 Ambient Temp.: 83.68 F
15:12:01.233 -> Average Cal sensor Temp.: 94.78 Ambient Temp.: 83.68 F
15:12:02.789 -> Average Cal sensor Temp.: 94.77 Ambient Temp.: 83.68 F
15:12:04.344 -> Average Cal sensor Temp.: 94.85 Ambient Temp.: 83.64 F
15:12:05.897 -> Average Cal sensor Temp.: 94.82 Ambient Temp.: 83.64 F
15:12:07.455 -> Average Cal sensor Temp.: 94.79 Ambient Temp.: 83.68 F
15:12:09.013 -> Average Cal sensor Temp.: 94.67 Ambient Temp.: 83.57 F
15:12:10.571 -> Average Cal sensor Temp.: 94.89 Ambient Temp.: 83.50 F
15:12:12.129 -> Average Cal sensor Temp.: 95.26 Ambient Temp.: 83.57 F
15:12:13.679 -> Average Cal sensor Temp.: 95.28 Ambient Temp.: 83.64 F
15:12:15.227 -> Average Cal sensor Temp.: 95.01 Ambient Temp.: 83.64 F
15:12:16.780 -> Average Cal sensor Temp.: 94.80 Ambient Temp.: 83.64 F
15:12:18.330 -> Average Cal sensor Temp.: 94.95 Ambient Temp.: 83.68 F
15:12:19.919 -> Average Cal sensor Temp.: 95.22 Ambient Temp.: 83.71 F
15:12:21.443 -> Average Cal sensor Temp.: 95.02 Ambient Temp.: 83.68 F
15:12:23.001 -> Average Cal sensor Temp.: 94.56 Ambient Temp.: 83.71 F
15:12:24.558 -> Average Cal sensor Temp.: 94.76 Ambient Temp.: 83.71 F
15:12:26.112 -> Average Cal sensor Temp.: 94.70 Ambient Temp.: 83.68 F
15:12:27.665 -> Average Cal sensor Temp.: 94.84 Ambient Temp.: 83.68 F
15:12:29.251 -> Average Cal sensor Temp.: 95.20 Ambient Temp.: 83.71 F
15:12:30.805 -> Average Cal sensor Temp.: 95.07 Ambient Temp.: 83.71 F
15:12:32.362 -> Average Cal sensor Temp.: 95.43 Ambient Temp.: 83.71 F
15:12:33.915 -> Average Cal sensor Temp.: 95.31 Ambient Temp.: 83.71 F
15:12:35.466 -> Average Cal sensor Temp.: 94.89 Ambient Temp.: 83.68 F
15:12:37.019 -> Average Cal sensor Temp.: 94.95 Ambient Temp.: 83.68 F
15:12:38.575 -> Average Cal sensor Temp.: 94.95 Ambient Temp.: 83.71 F
15:12:40.130 -> Average Cal sensor Temp.: 94.62 Ambient Temp.: 83.75 F
15:12:41.689 -> Average Cal sensor Temp.: 94.80 Ambient Temp.: 83.71 F
15:12:43.243 -> Average Cal sensor Temp.: 94.93 Ambient Temp.: 83.68 F
15:12:44.798 -> Average Cal sensor Temp.: 95.21 Ambient Temp.: 83.71 F
15:12:46.356 -> Average Cal sensor Temp.: 94.82 Ambient Temp.: 83.75 F
15:12:47.906 -> Average Cal sensor Temp.: 94.45 Ambient Temp.: 83.79 F
15:12:49.458 -> Average Cal sensor Temp.: 94.35 Ambient Temp.: 83.71 F
15:12:51.006 -> Average Cal sensor Temp.: 94.62 Ambient Temp.: 83.75 F
15:12:52.561 -> Average Cal sensor Temp.: 94.62 Ambient Temp.: 83.75 F
15:12:54.118 -> Average Cal sensor Temp.: 94.63 Ambient Temp.: 83.71 F
15:12:55.675 -> Average Cal sensor Temp.: 94.96 Ambient Temp.: 83.71 F
15:12:57.228 -> Average Cal sensor Temp.: 95.07 Ambient Temp.: 83.79 F
15:12:58.811 -> Average Cal sensor Temp.: 94.80 Ambient Temp.: 83.75 F
15:13:00.364 -> Average Cal sensor Temp.: 94.74 Ambient Temp.: 83.82 F
15:13:01.919 -> Average Cal sensor Temp.: 94.65 Ambient Temp.: 83.75 F
15:13:03.463 -> Average Cal sensor Temp.: 95.07 Ambient Temp.: 83.79 F
15:13:05.015 -> Average Cal sensor Temp.: 94.78 Ambient Temp.: 83.79 F
15:13:06.564 -> Average Cal sensor Temp.: 94.79 Ambient Temp.: 83.82 F
15:13:08.151 -> Average Cal sensor Temp.: 94.65 Ambient Temp.: 83.86 F
15:13:09.707 -> Average Cal sensor Temp.: 94.69 Ambient Temp.: 83.82 F
15:13:11.266 -> Average Cal sensor Temp.: 94.78 Ambient Temp.: 83.86 F
15:13:12.815 -> Average Cal sensor Temp.: 94.79 Ambient Temp.: 83.93 F
15:13:14.365 -> Average Cal sensor Temp.: 94.55 Ambient Temp.: 83.89 F
15:13:15.912 -> Average Cal sensor Temp.: 94.92 Ambient Temp.: 83.86 F
15:13:17.469 -> Average Cal sensor Temp.: 94.87 Ambient Temp.: 83.93 F
15:13:19.027 -> Average Cal sensor Temp.: 95.07 Ambient Temp.: 83.93 F
15:13:20.580 -> Average Cal sensor Temp.: 94.80 Ambient Temp.: 83.97 F
15:13:22.164 -> Average Cal sensor Temp.: 95.15 Ambient Temp.: 83.93 F
15:13:23.689 -> Average Cal sensor Temp.: 94.86 Ambient Temp.: 83.93 F
15:13:25.276 -> Average Cal sensor Temp.: 94.91 Ambient Temp.: 84.07 F
15:13:26.830 -> Average Cal sensor Temp.: 95.09 Ambient Temp.: 83.97 F
15:13:27.416 -> Ambient Temp.: 83.97 F
15:13:27.968 -> Ambient Temp.: 83.97 F
15:13:29.283 -> Ambient Temp.: 83.97 F
15:13:30.837 -> Average Cal sensor Temp.: 94.85 Ambient Temp.: 84.07 F
15:13:32.396 -> Average Cal sensor Temp.: 94.90 Ambient Temp.: 84.11 F
15:13:33.949 -> Average Cal sensor Temp.: 95.32 Ambient Temp.: 84.11 F
15:13:35.502 -> Average Cal sensor Temp.: 94.89 Ambient Temp.: 84.11 F
15:13:37.055 -> Average Cal sensor Temp.: 94.85 Ambient Temp.: 84.07 F
15:13:38.611 -> Average Cal sensor Temp.: 94.98 Ambient Temp.: 84.15 F
15:13:40.193 -> Average Cal sensor Temp.: 95.52 Ambient Temp.: 84.15 F
15:13:41.747 -> Average Cal sensor Temp.: 95.05 Ambient Temp.: 84.18 F
15:13:43.299 -> Average Cal sensor Temp.: 95.15 Ambient Temp.: 84.22 F
15:13:44.853 -> Average Cal sensor Temp.: 95.13 Ambient Temp.: 84.22 F
15:13:46.407 -> Average Cal sensor Temp.: 94.91 Ambient Temp.: 84.22 F
15:13:47.967 -> Average Cal sensor Temp.: 94.87 Ambient Temp.: 84.25 F
15:13:49.522 -> Average Cal sensor Temp.: 94.78 Ambient Temp.: 84.25 F
15:13:51.081 -> Average Cal sensor Temp.: 94.79 Ambient Temp.: 84.29 F
15:13:52.631 -> Average Cal sensor Temp.: 94.72 Ambient Temp.: 84.25 F
15:13:54.185 -> Average Cal sensor Temp.: 94.58 Ambient Temp.: 84.29 F
15:13:55.745 -> Average Cal sensor Temp.: 94.70 Ambient Temp.: 84.25 F
15:13:57.300 -> Average Cal sensor Temp.: 94.91 Ambient Temp.: 84.33 F
15:13:58.859 -> Average Cal sensor Temp.: 94.87 Ambient Temp.: 84.33 F
15:14:00.413 -> Average Cal sensor Temp.: 95.07 Ambient Temp.: 84.29 F
15:14:01.972 -> Average Cal sensor Temp.: 95.57 Ambient Temp.: 84.29 F
15:14:03.524 -> Average Cal sensor Temp.: 95.58 Ambient Temp.: 84.33 F
15:14:05.082 -> Average Cal sensor Temp.: 95.52 Ambient Temp.: 84.36 F
15:14:06.632 -> Average Cal sensor Temp.: 95.45 Ambient Temp.: 84.36 F
15:14:08.185 -> Average Cal sensor Temp.: 95.51 Ambient Temp.: 84.47 F
15:14:09.740 -> Average Cal sensor Temp.: 95.29 Ambient Temp.: 84.47 F
15:14:11.290 -> Average Cal sensor Temp.: 95.35 Ambient Temp.: 84.36 F
15:14:12.844 -> Average Cal sensor Temp.: 95.28 Ambient Temp.: 84.51 F
15:14:14.401 -> Average Cal sensor Temp.: 95.14 Ambient Temp.: 84.47 F
15:14:15.954 -> Average Cal sensor Temp.: 94.86 Ambient Temp.: 84.47 F
15:14:17.508 -> Average Cal sensor Temp.: 95.18 Ambient Temp.: 84.51 F
15:14:19.065 -> Average Cal sensor Temp.: 94.78 Ambient Temp.: 84.51 F
15:14:20.627 -> Average Cal sensor Temp.: 94.75 Ambient Temp.: 84.51 F
15:14:22.178 -> Average Cal sensor Temp.: 94.89 Ambient Temp.: 84.51 F
15:14:23.762 -> Average Cal sensor Temp.: 95.04 Ambient Temp.: 84.54 F
15:14:25.314 -> Average Cal sensor Temp.: 94.92 Ambient Temp.: 84.54 F
15:14:26.870 -> Average Cal sensor Temp.: 94.68 Ambient Temp.: 84.58 F
15:14:28.425 -> Average Cal sensor Temp.: 94.94 Ambient Temp.: 84.54 F
15:14:29.976 -> Average Cal sensor Temp.: 95.05 Ambient Temp.: 84.58 F
15:14:30.912 -> Ambient Temp.: 84.61 F
15:14:31.500 -> Ambient Temp.: 84.61 F
15:14:33.060 -> Average Cal sensor Temp.: 95.00 Ambient Temp.: 84.61 F
15:14:34.617 -> Average Cal sensor Temp.: 95.16 Ambient Temp.: 84.61 F
15:14:36.172 -> Average Cal sensor Temp.: 95.00 Ambient Temp.: 84.72 F
15:14:37.732 -> Average Cal sensor Temp.: 95.02 Ambient Temp.: 84.69 F
15:14:39.287 -> Average Cal sensor Temp.: 95.33 Ambient Temp.: 84.72 F
15:14:40.836 -> Average Cal sensor Temp.: 95.83 Ambient Temp.: 84.76 F
15:14:42.392 -> Average Cal sensor Temp.: 95.45 Ambient Temp.: 84.72 F
15:14:43.944 -> Average Cal sensor Temp.: 95.47 Ambient Temp.: 84.79 F
15:14:45.502 -> Average Cal sensor Temp.: 95.77 Ambient Temp.: 84.87 F
15:14:47.061 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 84.87 F
15:14:48.618 -> Average Cal sensor Temp.: 95.52 Ambient Temp.: 84.87 F
15:14:50.175 -> Average Cal sensor Temp.: 95.25 Ambient Temp.: 84.90 F
15:14:51.726 -> Average Cal sensor Temp.: 95.06 Ambient Temp.: 84.90 F
15:14:53.277 -> Average Cal sensor Temp.: 95.21 Ambient Temp.: 84.94 F
15:14:54.836 -> Average Cal sensor Temp.: 94.95 Ambient Temp.: 84.97 F
15:14:56.387 -> Average Cal sensor Temp.: 95.04 Ambient Temp.: 84.94 F
15:14:57.942 -> Average Cal sensor Temp.: 95.12 Ambient Temp.: 84.97 F
15:14:59.496 -> Average Cal sensor Temp.: 95.58 Ambient Temp.: 84.97 F
15:15:00.325 -> Ambient Temp.: 85.01 F
15:15:00.912 -> Ambient Temp.: 85.01 F
15:15:01.498 -> Ambient Temp.: 85.01 F
15:15:03.052 -> Average Cal sensor Temp.: 95.38 Ambient Temp.: 85.08 F
15:15:04.606 -> Average Cal sensor Temp.: 95.23 Ambient Temp.: 85.08 F
15:15:06.160 -> Average Cal sensor Temp.: 95.21 Ambient Temp.: 85.15 F
15:15:06.745 -> Ambient Temp.: 85.15 F
15:15:07.332 -> Ambient Temp.: 85.15 F
15:15:08.889 -> Average Cal sensor Temp.: 95.30 Ambient Temp.: 85.19 F
15:15:10.412 -> Average Cal sensor Temp.: 95.05 Ambient Temp.: 85.26 F
15:15:11.987 -> Average Cal sensor Temp.: 95.08 Ambient Temp.: 85.30 F
15:15:13.528 -> Average Cal sensor Temp.: 94.96 Ambient Temp.: 85.26 F
15:15:15.105 -> Average Cal sensor Temp.: 94.97 Ambient Temp.: 85.30 F
15:15:16.662 -> Average Cal sensor Temp.: 94.96 Ambient Temp.: 85.33 F
15:15:18.213 -> Average Cal sensor Temp.: 95.03 Ambient Temp.: 85.37 F
15:15:19.770 -> Average Cal sensor Temp.: 95.05 Ambient Temp.: 85.37 F
15:15:21.325 -> Average Cal sensor Temp.: 94.97 Ambient Temp.: 85.37 F
15:15:22.880 -> Average Cal sensor Temp.: 95.09 Ambient Temp.: 85.37 F
15:15:24.437 -> Average Cal sensor Temp.: 95.11 Ambient Temp.: 85.41 F
15:15:25.990 -> Average Cal sensor Temp.: 95.08 Ambient Temp.: 85.41 F
15:15:27.548 -> Average Cal sensor Temp.: 95.04 Ambient Temp.: 85.44 F
15:15:29.103 -> Average Cal sensor Temp.: 94.95 Ambient Temp.: 85.51 F
15:15:30.662 -> Average Cal sensor Temp.: 95.13 Ambient Temp.: 85.51 F
15:15:32.217 -> Average Cal sensor Temp.: 95.06 Ambient Temp.: 85.48 F
15:15:33.768 -> Average Cal sensor Temp.: 95.01 Ambient Temp.: 85.44 F
15:15:35.314 -> Average Cal sensor Temp.: 95.27 Ambient Temp.: 85.51 F
15:15:36.869 -> Average Cal sensor Temp.: 95.35 Ambient Temp.: 85.48 F
15:15:38.424 -> Average Cal sensor Temp.: 94.95 Ambient Temp.: 85.48 F
15:15:39.972 -> Average Cal sensor Temp.: 95.21 Ambient Temp.: 85.55 F
15:15:41.528 -> Average Cal sensor Temp.: 95.26 Ambient Temp.: 85.48 F
15:15:43.085 -> Average Cal sensor Temp.: 95.53 Ambient Temp.: 85.55 F
15:15:44.636 -> Average Cal sensor Temp.: 95.39 Ambient Temp.: 85.55 F
15:15:46.221 -> Average Cal sensor Temp.: 95.23 Ambient Temp.: 85.55 F
15:15:47.776 -> Average Cal sensor Temp.: 95.10 Ambient Temp.: 85.55 F
15:15:49.323 -> Average Cal sensor Temp.: 95.05 Ambient Temp.: 85.59 F
15:15:50.870 -> Average Cal sensor Temp.: 94.80 Ambient Temp.: 85.59 F
15:15:52.419 -> Average Cal sensor Temp.: 94.69 Ambient Temp.: 85.59 F
15:15:54.003 -> Average Cal sensor Temp.: 94.87 Ambient Temp.: 85.55 F
15:15:55.559 -> Average Cal sensor Temp.: 94.59 Ambient Temp.: 85.55 F
15:15:57.113 -> Average Cal sensor Temp.: 94.52 Ambient Temp.: 85.55 F
15:15:58.667 -> Average Cal sensor Temp.: 94.60 Ambient Temp.: 85.48 F
15:16:00.221 -> Average Cal sensor Temp.: 94.78 Ambient Temp.: 85.51 F
15:16:01.777 -> Average Cal sensor Temp.: 95.04 Ambient Temp.: 85.51 F
15:16:03.331 -> Average Cal sensor Temp.: 95.11 Ambient Temp.: 85.51 F
15:16:04.877 -> Average Cal sensor Temp.: 95.00 Ambient Temp.: 85.48 F
15:16:06.431 -> Average Cal sensor Temp.: 94.96 Ambient Temp.: 85.48 F
15:16:07.983 -> Average Cal sensor Temp.: 95.09 Ambient Temp.: 85.48 F
15:16:09.535 -> Average Cal sensor Temp.: 95.12 Ambient Temp.: 85.48 F
15:16:11.090 -> Average Cal sensor Temp.: 95.18 Ambient Temp.: 85.55 F
15:16:12.641 -> Average Cal sensor Temp.: 95.28 Ambient Temp.: 85.51 F
15:16:14.229 -> Average Cal sensor Temp.: 95.09 Ambient Temp.: 85.48 F
15:16:15.780 -> Average Cal sensor Temp.: 95.03 Ambient Temp.: 85.48 F
15:16:17.333 -> Average Cal sensor Temp.: 95.12 Ambient Temp.: 85.51 F
15:16:18.886 -> Average Cal sensor Temp.: 94.75 Ambient Temp.: 85.48 F
15:16:20.438 -> Average Cal sensor Temp.: 94.94 Ambient Temp.: 85.48 F
15:16:21.987 -> Average Cal sensor Temp.: 95.08 Ambient Temp.: 85.48 F
15:16:23.562 -> Average Cal sensor Temp.: 95.04 Ambient Temp.: 85.44 F
15:16:25.114 -> Average Cal sensor Temp.: 95.12 Ambient Temp.: 85.44 F
15:16:26.657 -> Average Cal sensor Temp.: 94.99 Ambient Temp.: 85.41 F
15:16:28.235 -> Average Cal sensor Temp.: 95.17 Ambient Temp.: 85.41 F
15:16:29.778 -> Average Cal sensor Temp.: 95.25 Ambient Temp.: 85.48 F
15:16:31.322 -> Average Cal sensor Temp.: 95.09 Ambient Temp.: 85.41 F
15:16:32.901 -> Average Cal sensor Temp.: 95.44 Ambient Temp.: 85.41 F
15:16:34.453 -> Average Cal sensor Temp.: 95.10 Ambient Temp.: 85.41 F
15:16:36.003 -> Average Cal sensor Temp.: 95.38 Ambient Temp.: 85.44 F
15:16:37.548 -> Average Cal sensor Temp.: 95.04 Ambient Temp.: 85.41 F
15:16:39.100 -> Average Cal sensor Temp.: 95.24 Ambient Temp.: 85.44 F
15:16:40.654 -> Average Cal sensor Temp.: 95.11 Ambient Temp.: 85.44 F
15:16:42.234 -> Average Cal sensor Temp.: 95.39 Ambient Temp.: 85.41 F
15:16:43.778 -> Average Cal sensor Temp.: 95.32 Ambient Temp.: 85.41 F
15:16:45.322 -> Average Cal sensor Temp.: 95.17 Ambient Temp.: 85.44 F
15:16:46.874 -> Average Cal sensor Temp.: 95.46 Ambient Temp.: 85.44 F
15:16:48.451 -> Average Cal sensor Temp.: 94.79 Ambient Temp.: 85.41 F
15:16:49.995 -> Average Cal sensor Temp.: 94.67 Ambient Temp.: 85.37 F
15:16:51.544 -> Average Cal sensor Temp.: 94.86 Ambient Temp.: 85.37 F
15:16:53.118 -> Average Cal sensor Temp.: 95.04 Ambient Temp.: 85.33 F
15:16:54.662 -> Average Cal sensor Temp.: 95.53 Ambient Temp.: 85.41 F
15:16:56.215 -> Average Cal sensor Temp.: 95.02 Ambient Temp.: 85.33 F
15:16:57.758 -> Average Cal sensor Temp.: 94.54 Ambient Temp.: 85.37 F
15:16:59.331 -> Average Cal sensor Temp.: 95.31 Ambient Temp.: 85.37 F
15:17:00.876 -> Average Cal sensor Temp.: 95.32 Ambient Temp.: 85.33 F
15:17:02.448 -> Average Cal sensor Temp.: 95.18 Ambient Temp.: 85.30 F
15:17:03.995 -> Average Cal sensor Temp.: 95.18 Ambient Temp.: 85.26 F
15:17:05.545 -> Average Cal sensor Temp.: 95.14 Ambient Temp.: 85.33 F
15:17:07.096 -> Average Cal sensor Temp.: 94.83 Ambient Temp.: 85.30 F
15:17:08.651 -> Average Cal sensor Temp.: 94.64 Ambient Temp.: 85.30 F
15:17:10.205 -> Average Cal sensor Temp.: 94.94 Ambient Temp.: 85.26 F
15:17:11.757 -> Average Cal sensor Temp.: 95.36 Ambient Temp.: 85.19 F
15:17:13.343 -> Average Cal sensor Temp.: 95.28 Ambient Temp.: 85.30 F
15:17:14.897 -> Average Cal sensor Temp.: 94.92 Ambient Temp.: 85.19 F
15:17:16.440 -> Average Cal sensor Temp.: 94.91 Ambient Temp.: 85.26 F
15:17:17.996 -> Average Cal sensor Temp.: 95.31 Ambient Temp.: 85.19 F
15:17:19.554 -> Average Cal sensor Temp.: 95.77 Ambient Temp.: 85.26 F
15:17:21.106 -> Average Cal sensor Temp.: 95.83 Ambient Temp.: 85.30 F
15:17:22.659 -> Average Cal sensor Temp.: 95.00 Ambient Temp.: 85.26 F
15:17:23.970 -> Ambient Temp.: 85.19 F
15:17:24.557 -> Ambient Temp.: 85.19 F
15:17:26.112 -> Average Cal sensor Temp.: 94.76 Ambient Temp.: 85.08 F
15:17:27.669 -> Average Cal sensor Temp.: 94.63 Ambient Temp.: 85.12 F
15:17:29.231 -> Average Cal sensor Temp.: 95.13 Ambient Temp.: 85.08 F
15:17:30.784 -> Average Cal sensor Temp.: 94.77 Ambient Temp.: 85.05 F
15:17:32.345 -> Average Cal sensor Temp.: 95.18 Ambient Temp.: 85.08 F
15:17:33.901 -> Average Cal sensor Temp.: 95.13 Ambient Temp.: 85.08 F
15:17:35.449 -> Average Cal sensor Temp.: 94.90 Ambient Temp.: 85.08 F
15:17:37.008 -> Average Cal sensor Temp.: 94.89 Ambient Temp.: 85.12 F
15:17:38.561 -> Average Cal sensor Temp.: 95.28 Ambient Temp.: 85.08 F
15:17:39.736 -> Ambient Temp.: 85.08 F
15:17:40.324 -> Ambient Temp.: 85.08 F
15:17:40.913 -> Ambient Temp.: 85.08 F
15:17:42.473 -> Average Cal sensor Temp.: 95.18 Ambient Temp.: 85.05 F
15:17:44.033 -> Average Cal sensor Temp.: 95.07 Ambient Temp.: 85.12 F
15:17:45.583 -> Average Cal sensor Temp.: 95.12 Ambient Temp.: 85.05 F
15:17:47.135 -> Average Cal sensor Temp.: 95.15 Ambient Temp.: 85.01 F
15:17:48.688 -> Average Cal sensor Temp.: 95.18 Ambient Temp.: 85.01 F
15:17:50.246 -> Average Cal sensor Temp.: 95.19 Ambient Temp.: 85.05 F
15:17:51.802 -> Average Cal sensor Temp.: 95.06 Ambient Temp.: 84.97 F
15:17:53.359 -> Average Cal sensor Temp.: 95.04 Ambient Temp.: 84.97 F
15:17:54.914 -> Average Cal sensor Temp.: 95.18 Ambient Temp.: 84.97 F
15:17:56.467 -> Average Cal sensor Temp.: 95.18 Ambient Temp.: 84.97 F
15:17:58.020 -> Average Cal sensor Temp.: 94.87 Ambient Temp.: 85.01 F
15:17:59.573 -> Average Cal sensor Temp.: 95.40 Ambient Temp.: 84.97 F
15:18:01.131 -> Average Cal sensor Temp.: 95.44 Ambient Temp.: 84.97 F
15:18:02.690 -> Average Cal sensor Temp.: 95.23 Ambient Temp.: 84.94 F
15:18:04.244 -> Average Cal sensor Temp.: 94.91 Ambient Temp.: 84.94 F
15:18:05.798 -> Average Cal sensor Temp.: 95.15 Ambient Temp.: 84.94 F
15:18:07.344 -> Average Cal sensor Temp.: 95.09 Ambient Temp.: 84.94 F
15:18:08.929 -> Average Cal sensor Temp.: 94.96 Ambient Temp.: 84.94 F
15:18:10.484 -> Average Cal sensor Temp.: 95.16 Ambient Temp.: 84.94 F
15:18:12.039 -> Average Cal sensor Temp.: 95.48 Ambient Temp.: 85.01 F
15:18:13.583 -> Average Cal sensor Temp.: 95.53 Ambient Temp.: 84.97 F
15:18:15.135 -> Average Cal sensor Temp.: 95.14 Ambient Temp.: 84.94 F
15:18:16.691 -> Average Cal sensor Temp.: 95.08 Ambient Temp.: 84.97 F
15:18:18.248 -> Average Cal sensor Temp.: 94.72 Ambient Temp.: 84.94 F
15:18:19.805 -> Average Cal sensor Temp.: 94.96 Ambient Temp.: 84.87 F
15:18:21.362 -> Average Cal sensor Temp.: 95.06 Ambient Temp.: 84.87 F
15:18:22.922 -> Average Cal sensor Temp.: 95.11 Ambient Temp.: 84.87 F
15:18:24.474 -> Average Cal sensor Temp.: 95.29 Ambient Temp.: 84.79 F
15:18:26.024 -> Average Cal sensor Temp.: 95.51 Ambient Temp.: 84.79 F
15:18:27.580 -> Average Cal sensor Temp.: 95.39 Ambient Temp.: 84.79 F
15:18:29.139 -> Average Cal sensor Temp.: 95.43 Ambient Temp.: 84.79 F
15:18:30.696 -> Average Cal sensor Temp.: 95.42 Ambient Temp.: 84.87 F
15:18:32.254 -> Average Cal sensor Temp.: 95.12 Ambient Temp.: 84.79 F
15:18:33.805 -> Average Cal sensor Temp.: 95.10 Ambient Temp.: 84.79 F
15:18:35.360 -> Average Cal sensor Temp.: 95.47 Ambient Temp.: 84.76 F
15:18:36.918 -> Average Cal sensor Temp.: 95.64 Ambient Temp.: 84.72 F
15:18:38.476 -> Average Cal sensor Temp.: 95.73 Ambient Temp.: 84.79 F
15:18:40.027 -> Average Cal sensor Temp.: 95.35 Ambient Temp.: 84.79 F
15:18:41.583 -> Average Cal sensor Temp.: 95.07 Ambient Temp.: 84.76 F
15:18:43.139 -> Average Cal sensor Temp.: 95.50 Ambient Temp.: 84.72 F
15:18:44.697 -> Average Cal sensor Temp.: 95.26 Ambient Temp.: 84.76 F
15:18:46.254 -> Average Cal sensor Temp.: 94.96 Ambient Temp.: 84.72 F
15:18:47.804 -> Average Cal sensor Temp.: 95.10 Ambient Temp.: 84.69 F
15:18:49.358 -> Average Cal sensor Temp.: 95.20 Ambient Temp.: 84.65 F
15:18:50.914 -> Average Cal sensor Temp.: 95.08 Ambient Temp.: 84.65 F
15:18:52.504 -> Average Cal sensor Temp.: 95.13 Ambient Temp.: 84.65 F
15:18:54.050 -> Average Cal sensor Temp.: 95.27 Ambient Temp.: 84.61 F
15:18:55.603 -> Average Cal sensor Temp.: 95.32 Ambient Temp.: 84.58 F
15:18:57.161 -> Average Cal sensor Temp.: 95.05 Ambient Temp.: 84.58 F
15:18:58.714 -> Average Cal sensor Temp.: 95.15 Ambient Temp.: 84.54 F
15:19:00.266 -> Average Cal sensor Temp.: 94.95 Ambient Temp.: 84.51 F
15:19:01.821 -> Average Cal sensor Temp.: 95.44 Ambient Temp.: 84.58 F
15:19:03.382 -> Average Cal sensor Temp.: 95.34 Ambient Temp.: 84.47 F
15:19:04.933 -> Average Cal sensor Temp.: 95.22 Ambient Temp.: 84.51 F
15:19:06.486 -> Average Cal sensor Temp.: 95.18 Ambient Temp.: 84.47 F
15:19:08.042 -> Average Cal sensor Temp.: 94.88 Ambient Temp.: 84.36 F
15:19:09.595 -> Average Cal sensor Temp.: 95.18 Ambient Temp.: 84.36 F
15:19:11.152 -> Average Cal sensor Temp.: 95.33 Ambient Temp.: 84.36 F
15:19:12.706 -> Average Cal sensor Temp.: 95.33 Ambient Temp.: 84.36 F
15:19:13.297 -> Ambient Temp.: 84.36 F
15:19:13.884 -> Ambient Temp.: 84.36 F
15:19:15.440 -> Average Cal sensor Temp.: 94.86 Ambient Temp.: 84.33 F
15:19:16.993 -> Average Cal sensor Temp.: 95.19 Ambient Temp.: 84.33 F
15:19:18.541 -> Average Cal sensor Temp.: 95.22 Ambient Temp.: 84.29 F
15:19:20.094 -> Average Cal sensor Temp.: 95.02 Ambient Temp.: 84.29 F
15:19:21.650 -> Average Cal sensor Temp.: 94.98 Ambient Temp.: 84.22 F
15:19:23.209 -> Average Cal sensor Temp.: 95.05 Ambient Temp.: 84.22 F
15:19:24.762 -> Average Cal sensor Temp.: 95.51 Ambient Temp.: 84.25 F
15:19:26.319 -> Average Cal sensor Temp.: 95.34 Ambient Temp.: 84.18 F
15:19:27.870 -> Average Cal sensor Temp.: 94.86 Ambient Temp.: 84.22 F
15:19:29.423 -> Average Cal sensor Temp.: 94.99 Ambient Temp.: 84.15 F
15:19:30.977 -> Average Cal sensor Temp.: 95.06 Ambient Temp.: 84.15 F
15:19:32.563 -> Average Cal sensor Temp.: 94.91 Ambient Temp.: 84.18 F
15:19:34.113 -> Average Cal sensor Temp.: 95.13 Ambient Temp.: 84.18 F
15:19:35.661 -> Average Cal sensor Temp.: 95.15 Ambient Temp.: 84.18 F
15:19:37.217 -> Average Cal sensor Temp.: 95.12 Ambient Temp.: 84.18 F
15:19:38.757 -> Average Cal sensor Temp.: 95.17 Ambient Temp.: 84.18 F
15:19:40.333 -> Average Cal sensor Temp.: 95.35 Ambient Temp.: 84.18 F
15:19:41.878 -> Average Cal sensor Temp.: 95.47 Ambient Temp.: 84.11 F
15:19:43.433 -> Average Cal sensor Temp.: 95.20 Ambient Temp.: 84.15 F
15:19:44.976 -> Average Cal sensor Temp.: 95.23 Ambient Temp.: 84.15 F
15:19:46.529 -> Average Cal sensor Temp.: 95.54 Ambient Temp.: 84.15 F
15:19:48.110 -> Average Cal sensor Temp.: 95.31 Ambient Temp.: 84.15 F
15:19:49.657 -> Average Cal sensor Temp.: 95.43 Ambient Temp.: 84.00 F
15:19:51.201 -> Average Cal sensor Temp.: 95.50 Ambient Temp.: 84.07 F
15:19:52.750 -> Average Cal sensor Temp.: 94.91 Ambient Temp.: 84.11 F
15:19:54.331 -> Average Cal sensor Temp.: 94.89 Ambient Temp.: 84.15 F
15:19:55.880 -> Average Cal sensor Temp.: 95.27 Ambient Temp.: 84.07 F
15:19:57.436 -> Average Cal sensor Temp.: 95.70 Ambient Temp.: 84.15 F
15:19:58.988 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 84.15 F
15:20:00.538 -> Average Cal sensor Temp.: 95.53 Ambient Temp.: 84.11 F
15:20:02.113 -> Average Cal sensor Temp.: 95.54 Ambient Temp.: 84.11 F
15:20:03.658 -> Average Cal sensor Temp.: 95.26 Ambient Temp.: 84.07 F
15:20:05.208 -> Average Cal sensor Temp.: 95.16 Ambient Temp.: 84.07 F
15:20:06.767 -> Average Cal sensor Temp.: 94.94 Ambient Temp.: 84.00 F
15:20:08.316 -> Average Cal sensor Temp.: 95.52 Ambient Temp.: 84.11 F
15:20:09.874 -> Average Cal sensor Temp.: 95.64 Ambient Temp.: 84.07 F
15:20:11.431 -> Average Cal sensor Temp.: 95.39 Ambient Temp.: 84.07 F
15:20:12.995 -> Average Cal sensor Temp.: 95.27 Ambient Temp.: 84.07 F
15:20:14.549 -> Average Cal sensor Temp.: 95.35 Ambient Temp.: 84.07 F
15:20:16.106 -> Average Cal sensor Temp.: 95.27 Ambient Temp.: 84.07 F
15:20:17.657 -> Average Cal sensor Temp.: 95.27 Ambient Temp.: 84.07 F
15:20:19.214 -> Average Cal sensor Temp.: 95.14 Ambient Temp.: 84.15 F
15:20:20.767 -> Average Cal sensor Temp.: 95.29 Ambient Temp.: 84.11 F
15:20:22.323 -> Average Cal sensor Temp.: 95.67 Ambient Temp.: 84.15 F
15:20:23.876 -> Average Cal sensor Temp.: 95.78 Ambient Temp.: 84.15 F
15:20:25.429 -> Average Cal sensor Temp.: 95.61 Ambient Temp.: 84.15 F
15:20:26.987 -> Average Cal sensor Temp.: 95.26 Ambient Temp.: 84.18 F
15:20:28.540 -> Average Cal sensor Temp.: 95.11 Ambient Temp.: 84.18 F
15:20:30.093 -> Average Cal sensor Temp.: 95.22 Ambient Temp.: 84.15 F
15:20:31.649 -> Average Cal sensor Temp.: 95.23 Ambient Temp.: 84.18 F
15:20:33.205 -> Average Cal sensor Temp.: 94.73 Ambient Temp.: 84.18 F
15:20:34.761 -> Average Cal sensor Temp.: 94.92 Ambient Temp.: 84.18 F
15:20:36.315 -> Average Cal sensor Temp.: 94.95 Ambient Temp.: 84.22 F
15:20:37.869 -> Average Cal sensor Temp.: 95.00 Ambient Temp.: 84.18 F
15:20:39.421 -> Average Cal sensor Temp.: 95.09 Ambient Temp.: 84.18 F
15:20:40.977 -> Average Cal sensor Temp.: 95.09 Ambient Temp.: 84.22 F
15:20:42.529 -> Average Cal sensor Temp.: 95.21 Ambient Temp.: 84.22 F
15:20:44.081 -> Average Cal sensor Temp.: 95.12 Ambient Temp.: 84.22 F
15:20:45.637 -> Average Cal sensor Temp.: 95.39 Ambient Temp.: 84.18 F
15:20:47.191 -> Average Cal sensor Temp.: 95.51 Ambient Temp.: 84.25 F
15:20:48.746 -> Average Cal sensor Temp.: 95.26 Ambient Temp.: 84.22 F
15:20:50.329 -> Average Cal sensor Temp.: 95.17 Ambient Temp.: 84.18 F
15:20:51.878 -> Average Cal sensor Temp.: 95.25 Ambient Temp.: 84.22 F
15:20:53.434 -> Average Cal sensor Temp.: 95.33 Ambient Temp.: 84.18 F
15:20:54.987 -> Average Cal sensor Temp.: 94.97 Ambient Temp.: 84.22 F
15:20:56.536 -> Average Cal sensor Temp.: 95.11 Ambient Temp.: 84.25 F
15:20:58.092 -> Average Cal sensor Temp.: 95.32 Ambient Temp.: 84.22 F
15:20:59.636 -> Average Cal sensor Temp.: 95.23 Ambient Temp.: 84.25 F
15:21:01.187 -> Average Cal sensor Temp.: 95.31 Ambient Temp.: 84.29 F
15:21:02.741 -> Average Cal sensor Temp.: 95.29 Ambient Temp.: 84.25 F
15:21:04.327 -> Average Cal sensor Temp.: 94.97 Ambient Temp.: 84.22 F
15:21:05.884 -> Average Cal sensor Temp.: 94.99 Ambient Temp.: 84.25 F
15:21:07.436 -> Average Cal sensor Temp.: 94.94 Ambient Temp.: 84.22 F
15:21:08.985 -> Average Cal sensor Temp.: 95.05 Ambient Temp.: 84.29 F
15:21:10.538 -> Average Cal sensor Temp.: 95.05 Ambient Temp.: 84.25 F
15:21:12.090 -> Average Cal sensor Temp.: 95.26 Ambient Temp.: 84.29 F
15:21:13.643 -> Average Cal sensor Temp.: 95.32 Ambient Temp.: 84.25 F
15:21:15.200 -> Average Cal sensor Temp.: 94.85 Ambient Temp.: 84.25 F
15:21:16.759 -> Average Cal sensor Temp.: 95.06 Ambient Temp.: 84.25 F
15:21:18.313 -> Average Cal sensor Temp.: 95.03 Ambient Temp.: 84.33 F
15:21:19.866 -> Average Cal sensor Temp.: 95.38 Ambient Temp.: 84.33 F
15:21:21.419 -> Average Cal sensor Temp.: 95.29 Ambient Temp.: 84.36 F
15:21:23.004 -> Average Cal sensor Temp.: 95.58 Ambient Temp.: 84.33 F
15:21:24.561 -> Average Cal sensor Temp.: 95.36 Ambient Temp.: 84.36 F
15:21:26.108 -> Average Cal sensor Temp.: 95.13 Ambient Temp.: 84.36 F
15:21:27.655 -> Average Cal sensor Temp.: 94.97 Ambient Temp.: 84.40 F
15:21:29.198 -> Average Cal sensor Temp.: 94.72 Ambient Temp.: 84.40 F
15:21:30.787 -> Average Cal sensor Temp.: 95.21 Ambient Temp.: 84.40 F
15:21:32.338 -> Average Cal sensor Temp.: 95.07 Ambient Temp.: 84.40 F
15:21:33.886 -> Average Cal sensor Temp.: 95.04 Ambient Temp.: 84.40 F
15:21:35.437 -> Average Cal sensor Temp.: 95.25 Ambient Temp.: 84.47 F
15:21:36.989 -> Average Cal sensor Temp.: 95.65 Ambient Temp.: 84.47 F
15:21:38.543 -> Average Cal sensor Temp.: 95.77 Ambient Temp.: 84.51 F
15:21:40.102 -> Average Cal sensor Temp.: 95.65 Ambient Temp.: 84.47 F
15:21:41.659 -> Average Cal sensor Temp.: 95.69 Ambient Temp.: 84.54 F
15:21:43.211 -> Average Cal sensor Temp.: 95.72 Ambient Temp.: 84.51 F
15:21:44.766 -> Average Cal sensor Temp.: 95.69 Ambient Temp.: 84.54 F
15:21:46.320 -> Average Cal sensor Temp.: 95.46 Ambient Temp.: 84.58 F
15:21:47.869 -> Average Cal sensor Temp.: 95.67 Ambient Temp.: 84.58 F
15:21:49.426 -> Average Cal sensor Temp.: 95.67 Ambient Temp.: 84.65 F
15:21:51.005 -> Average Cal sensor Temp.: 95.27 Ambient Temp.: 84.65 F
15:21:52.552 -> Average Cal sensor Temp.: 95.31 Ambient Temp.: 84.61 F
15:21:54.104 -> Average Cal sensor Temp.: 95.33 Ambient Temp.: 84.65 F
15:21:55.660 -> Average Cal sensor Temp.: 95.53 Ambient Temp.: 84.69 F
15:21:57.217 -> Average Cal sensor Temp.: 95.35 Ambient Temp.: 84.65 F
15:21:58.769 -> Average Cal sensor Temp.: 95.77 Ambient Temp.: 84.72 F
15:22:00.322 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 84.65 F
15:22:01.903 -> Average Cal sensor Temp.: 95.90 Ambient Temp.: 84.72 F
15:22:03.450 -> Average Cal sensor Temp.: 95.86 Ambient Temp.: 84.76 F
15:22:05.004 -> Average Cal sensor Temp.: 95.87 Ambient Temp.: 84.69 F
15:22:06.556 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 84.79 F
15:22:08.108 -> Average Cal sensor Temp.: 95.71 Ambient Temp.: 84.87 F
15:22:09.660 -> Average Cal sensor Temp.: 95.57 Ambient Temp.: 84.87 F
15:22:11.216 -> Average Cal sensor Temp.: 95.61 Ambient Temp.: 84.87 F
15:22:12.768 -> Average Cal sensor Temp.: 95.78 Ambient Temp.: 84.79 F
15:22:14.352 -> Average Cal sensor Temp.: 95.99 Ambient Temp.: 84.87 F
15:22:15.905 -> Average Cal sensor Temp.: 96.12 Ambient Temp.: 84.94 F
15:22:17.464 -> Average Cal sensor Temp.: 95.78 Ambient Temp.: 84.90 F
15:22:19.017 -> Average Cal sensor Temp.: 95.74 Ambient Temp.: 84.94 F
15:22:20.571 -> Average Cal sensor Temp.: 95.67 Ambient Temp.: 85.01 F
15:22:22.122 -> Average Cal sensor Temp.: 95.61 Ambient Temp.: 85.01 F
15:22:23.676 -> Average Cal sensor Temp.: 95.58 Ambient Temp.: 85.01 F
15:22:25.225 -> Average Cal sensor Temp.: 95.50 Ambient Temp.: 85.05 F
15:22:26.778 -> Average Cal sensor Temp.: 95.82 Ambient Temp.: 85.05 F
15:22:28.330 -> Average Cal sensor Temp.: 95.85 Ambient Temp.: 85.08 F
15:22:29.883 -> Average Cal sensor Temp.: 95.72 Ambient Temp.: 85.15 F
15:22:31.465 -> Average Cal sensor Temp.: 95.69 Ambient Temp.: 85.08 F
15:22:33.016 -> Average Cal sensor Temp.: 95.54 Ambient Temp.: 85.15 F
15:22:34.572 -> Average Cal sensor Temp.: 95.29 Ambient Temp.: 85.19 F
15:22:36.130 -> Average Cal sensor Temp.: 95.30 Ambient Temp.: 85.15 F
15:22:37.686 -> Average Cal sensor Temp.: 95.54 Ambient Temp.: 85.19 F
15:22:39.237 -> Average Cal sensor Temp.: 95.94 Ambient Temp.: 85.26 F
15:22:40.791 -> Average Cal sensor Temp.: 95.30 Ambient Temp.: 85.30 F
15:22:42.346 -> Average Cal sensor Temp.: 95.38 Ambient Temp.: 85.30 F
15:22:43.894 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 85.30 F
15:22:45.447 -> Average Cal sensor Temp.: 95.66 Ambient Temp.: 85.33 F
15:22:47.002 -> Average Cal sensor Temp.: 95.54 Ambient Temp.: 85.33 F
15:22:48.553 -> Average Cal sensor Temp.: 95.60 Ambient Temp.: 85.30 F
15:22:50.109 -> Average Cal sensor Temp.: 95.51 Ambient Temp.: 85.33 F
15:22:51.661 -> Average Cal sensor Temp.: 95.49 Ambient Temp.: 85.41 F
15:22:53.215 -> Average Cal sensor Temp.: 95.49 Ambient Temp.: 85.41 F
15:22:54.770 -> Average Cal sensor Temp.: 95.32 Ambient Temp.: 85.37 F
15:22:56.325 -> Average Cal sensor Temp.: 95.45 Ambient Temp.: 85.41 F
15:22:57.910 -> Average Cal sensor Temp.: 95.52 Ambient Temp.: 85.44 F
15:22:59.464 -> Average Cal sensor Temp.: 95.54 Ambient Temp.: 85.41 F
15:23:01.021 -> Average Cal sensor Temp.: 95.58 Ambient Temp.: 85.37 F
15:23:02.543 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 85.44 F
15:23:04.126 -> Average Cal sensor Temp.: 95.76 Ambient Temp.: 85.48 F
15:23:05.684 -> Average Cal sensor Temp.: 96.03 Ambient Temp.: 85.51 F
15:23:07.235 -> Average Cal sensor Temp.: 95.72 Ambient Temp.: 85.51 F
15:23:08.786 -> Average Cal sensor Temp.: 95.75 Ambient Temp.: 85.48 F
15:23:10.337 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 85.55 F
15:23:11.888 -> Average Cal sensor Temp.: 96.21 Ambient Temp.: 85.59 F
15:23:13.440 -> Average Cal sensor Temp.: 95.94 Ambient Temp.: 85.59 F
15:23:14.991 -> Average Cal sensor Temp.: 96.09 Ambient Temp.: 85.59 F
15:23:16.576 -> Average Cal sensor Temp.: 96.02 Ambient Temp.: 85.66 F
15:23:18.129 -> Average Cal sensor Temp.: 96.02 Ambient Temp.: 85.66 F
15:23:19.671 -> Average Cal sensor Temp.: 95.80 Ambient Temp.: 85.66 F
15:23:21.228 -> Average Cal sensor Temp.: 95.86 Ambient Temp.: 85.66 F
15:23:22.783 -> Average Cal sensor Temp.: 95.76 Ambient Temp.: 85.59 F
15:23:24.332 -> Average Cal sensor Temp.: 95.86 Ambient Temp.: 85.66 F
15:23:25.917 -> Average Cal sensor Temp.: 95.82 Ambient Temp.: 85.69 F
15:23:27.473 -> Average Cal sensor Temp.: 95.78 Ambient Temp.: 85.66 F
15:23:29.033 -> Average Cal sensor Temp.: 96.02 Ambient Temp.: 85.66 F
15:23:30.587 -> Average Cal sensor Temp.: 95.78 Ambient Temp.: 85.59 F
15:23:32.144 -> Average Cal sensor Temp.: 95.48 Ambient Temp.: 85.66 F
15:23:33.699 -> Average Cal sensor Temp.: 95.68 Ambient Temp.: 85.66 F
15:23:35.248 -> Average Cal sensor Temp.: 95.89 Ambient Temp.: 85.66 F
15:23:36.807 -> Average Cal sensor Temp.: 95.84 Ambient Temp.: 85.66 F
15:23:38.365 -> Average Cal sensor Temp.: 96.29 Ambient Temp.: 85.66 F
15:23:39.916 -> Average Cal sensor Temp.: 95.72 Ambient Temp.: 85.69 F
15:23:41.469 -> Average Cal sensor Temp.: 95.86 Ambient Temp.: 85.66 F
15:23:43.024 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 85.69 F
15:23:44.577 -> Average Cal sensor Temp.: 95.94 Ambient Temp.: 85.66 F
15:23:46.128 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 85.69 F
15:23:47.716 -> Average Cal sensor Temp.: 96.05 Ambient Temp.: 85.66 F
15:23:49.263 -> Average Cal sensor Temp.: 96.19 Ambient Temp.: 85.66 F
15:23:50.819 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 85.66 F
15:23:52.376 -> Average Cal sensor Temp.: 96.11 Ambient Temp.: 85.69 F
15:23:53.934 -> Average Cal sensor Temp.: 95.82 Ambient Temp.: 85.66 F
15:23:55.487 -> Average Cal sensor Temp.: 95.62 Ambient Temp.: 85.66 F
15:23:57.042 -> Average Cal sensor Temp.: 95.54 Ambient Temp.: 85.66 F
15:23:58.589 -> Average Cal sensor Temp.: 95.86 Ambient Temp.: 85.59 F
15:24:00.137 -> Average Cal sensor Temp.: 95.95 Ambient Temp.: 85.66 F
15:24:01.721 -> Average Cal sensor Temp.: 96.18 Ambient Temp.: 85.66 F
15:24:03.269 -> Average Cal sensor Temp.: 96.36 Ambient Temp.: 85.66 F
15:24:04.821 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 85.66 F
15:24:06.378 -> Average Cal sensor Temp.: 95.90 Ambient Temp.: 85.59 F
15:24:07.928 -> Average Cal sensor Temp.: 95.68 Ambient Temp.: 85.59 F
15:24:09.485 -> Average Cal sensor Temp.: 95.80 Ambient Temp.: 85.59 F
15:24:11.041 -> Average Cal sensor Temp.: 95.58 Ambient Temp.: 85.55 F
15:24:12.598 -> Average Cal sensor Temp.: 95.61 Ambient Temp.: 85.55 F
15:24:14.152 -> Average Cal sensor Temp.: 95.84 Ambient Temp.: 85.66 F
15:24:15.706 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 85.55 F
15:24:17.291 -> Average Cal sensor Temp.: 95.85 Ambient Temp.: 85.55 F
15:24:18.841 -> Average Cal sensor Temp.: 95.90 Ambient Temp.: 85.59 F
15:24:20.393 -> Average Cal sensor Temp.: 95.73 Ambient Temp.: 85.59 F
15:24:21.946 -> Average Cal sensor Temp.: 95.87 Ambient Temp.: 85.59 F
15:24:23.498 -> Average Cal sensor Temp.: 96.31 Ambient Temp.: 85.51 F
15:24:25.049 -> Average Cal sensor Temp.: 96.10 Ambient Temp.: 85.51 F
15:24:26.595 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 85.51 F
15:24:28.152 -> Average Cal sensor Temp.: 95.59 Ambient Temp.: 85.55 F
15:24:29.703 -> Average Cal sensor Temp.: 95.83 Ambient Temp.: 85.51 F
15:24:31.258 -> Average Cal sensor Temp.: 95.80 Ambient Temp.: 85.48 F
15:24:32.837 -> Average Cal sensor Temp.: 95.90 Ambient Temp.: 85.48 F
15:24:34.395 -> Average Cal sensor Temp.: 95.85 Ambient Temp.: 85.48 F
15:24:35.952 -> Average Cal sensor Temp.: 96.03 Ambient Temp.: 85.55 F
15:24:37.505 -> Average Cal sensor Temp.: 96.01 Ambient Temp.: 85.51 F
15:24:39.051 -> Average Cal sensor Temp.: 95.93 Ambient Temp.: 85.55 F
15:24:40.604 -> Average Cal sensor Temp.: 95.98 Ambient Temp.: 85.48 F
15:24:42.158 -> Average Cal sensor Temp.: 95.74 Ambient Temp.: 85.51 F
15:24:43.704 -> Average Cal sensor Temp.: 95.49 Ambient Temp.: 85.51 F
15:24:45.288 -> Average Cal sensor Temp.: 94.91 Ambient Temp.: 85.48 F
15:24:46.840 -> Average Cal sensor Temp.: 95.54 Ambient Temp.: 85.48 F
15:24:48.388 -> Average Cal sensor Temp.: 95.24 Ambient Temp.: 85.48 F
15:24:49.948 -> Average Cal sensor Temp.: 95.90 Ambient Temp.: 85.48 F
15:24:51.499 -> Average Cal sensor Temp.: 95.72 Ambient Temp.: 85.51 F
15:24:53.047 -> Average Cal sensor Temp.: 96.01 Ambient Temp.: 85.48 F
15:24:54.597 -> Average Cal sensor Temp.: 95.92 Ambient Temp.: 85.44 F
15:24:56.185 -> Average Cal sensor Temp.: 95.74 Ambient Temp.: 85.48 F
15:24:57.741 -> Average Cal sensor Temp.: 95.93 Ambient Temp.: 85.44 F
15:24:59.298 -> Average Cal sensor Temp.: 95.57 Ambient Temp.: 85.41 F
15:25:00.846 -> Average Cal sensor Temp.: 95.31 Ambient Temp.: 85.37 F
15:25:02.397 -> Average Cal sensor Temp.: 95.59 Ambient Temp.: 85.33 F
15:25:03.951 -> Average Cal sensor Temp.: 95.62 Ambient Temp.: 85.33 F
15:25:05.503 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 85.30 F
15:25:07.066 -> Average Cal sensor Temp.: 95.76 Ambient Temp.: 85.26 F
15:25:08.623 -> Average Cal sensor Temp.: 96.06 Ambient Temp.: 85.19 F
15:25:10.176 -> Average Cal sensor Temp.: 96.13 Ambient Temp.: 85.12 F
15:25:11.729 -> Average Cal sensor Temp.: 95.79 Ambient Temp.: 85.15 F
15:25:13.278 -> Average Cal sensor Temp.: 95.83 Ambient Temp.: 85.19 F
15:25:14.829 -> Average Cal sensor Temp.: 95.94 Ambient Temp.: 85.15 F
15:25:16.379 -> Average Cal sensor Temp.: 95.92 Ambient Temp.: 85.08 F
15:25:17.958 -> Average Cal sensor Temp.: 95.42 Ambient Temp.: 85.12 F
15:25:19.504 -> Average Cal sensor Temp.: 95.77 Ambient Temp.: 85.05 F
15:25:21.053 -> Average Cal sensor Temp.: 95.70 Ambient Temp.: 85.01 F
15:25:22.598 -> Average Cal sensor Temp.: 95.91 Ambient Temp.: 85.05 F
15:25:24.148 -> Average Cal sensor Temp.: 96.16 Ambient Temp.: 85.05 F
15:25:25.730 -> Average Cal sensor Temp.: 96.09 Ambient Temp.: 85.05 F
15:25:27.277 -> Average Cal sensor Temp.: 96.10 Ambient Temp.: 85.01 F
15:25:28.830 -> Average Cal sensor Temp.: 96.11 Ambient Temp.: 84.97 F
15:25:30.381 -> Average Cal sensor Temp.: 96.34 Ambient Temp.: 84.97 F
15:25:31.933 -> Average Cal sensor Temp.: 96.24 Ambient Temp.: 85.05 F
15:25:33.519 -> Average Cal sensor Temp.: 95.85 Ambient Temp.: 85.01 F
15:25:35.076 -> Average Cal sensor Temp.: 96.18 Ambient Temp.: 85.05 F
15:25:36.624 -> Average Cal sensor Temp.: 96.28 Ambient Temp.: 85.05 F
15:25:38.172 -> Average Cal sensor Temp.: 96.14 Ambient Temp.: 85.01 F
15:25:39.721 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 85.01 F
15:25:41.277 -> Average Cal sensor Temp.: 95.99 Ambient Temp.: 85.05 F
15:25:42.828 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 85.01 F
15:25:44.389 -> Average Cal sensor Temp.: 96.27 Ambient Temp.: 85.05 F
15:25:45.951 -> Average Cal sensor Temp.: 96.13 Ambient Temp.: 85.01 F
15:25:47.506 -> Average Cal sensor Temp.: 95.31 Ambient Temp.: 85.05 F
15:25:49.062 -> Average Cal sensor Temp.: 96.03 Ambient Temp.: 84.97 F
15:25:50.616 -> Average Cal sensor Temp.: 95.99 Ambient Temp.: 84.97 F
15:25:52.172 -> Average Cal sensor Temp.: 95.93 Ambient Temp.: 85.01 F
15:25:53.724 -> Average Cal sensor Temp.: 95.87 Ambient Temp.: 84.94 F
15:25:55.282 -> Average Cal sensor Temp.: 95.78 Ambient Temp.: 84.90 F
15:25:56.835 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 84.87 F
15:25:58.387 -> Average Cal sensor Temp.: 95.96 Ambient Temp.: 84.87 F
15:25:59.944 -> Average Cal sensor Temp.: 96.21 Ambient Temp.: 84.90 F
15:26:01.531 -> Average Cal sensor Temp.: 96.15 Ambient Temp.: 84.90 F
15:26:03.088 -> Average Cal sensor Temp.: 96.38 Ambient Temp.: 84.87 F
15:26:04.644 -> Average Cal sensor Temp.: 96.73 Ambient Temp.: 84.90 F
15:26:06.200 -> Average Cal sensor Temp.: 96.35 Ambient Temp.: 84.90 F
15:26:07.755 -> Average Cal sensor Temp.: 96.38 Ambient Temp.: 84.87 F
15:26:09.307 -> Average Cal sensor Temp.: 96.31 Ambient Temp.: 84.90 F
15:26:10.856 -> Average Cal sensor Temp.: 96.31 Ambient Temp.: 84.79 F
15:26:12.413 -> Average Cal sensor Temp.: 96.31 Ambient Temp.: 84.90 F
15:26:13.963 -> Average Cal sensor Temp.: 95.88 Ambient Temp.: 84.90 F
15:26:15.513 -> Average Cal sensor Temp.: 96.04 Ambient Temp.: 84.87 F
15:26:17.098 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 84.90 F
15:26:18.652 -> Average Cal sensor Temp.: 95.94 Ambient Temp.: 84.87 F
15:26:20.206 -> Average Cal sensor Temp.: 95.80 Ambient Temp.: 84.87 F
15:26:21.757 -> Average Cal sensor Temp.: 95.86 Ambient Temp.: 84.76 F
15:26:23.309 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 84.79 F
15:26:24.869 -> Average Cal sensor Temp.: 95.50 Ambient Temp.: 84.72 F
15:26:26.422 -> Average Cal sensor Temp.: 95.92 Ambient Temp.: 84.65 F
15:26:27.970 -> Average Cal sensor Temp.: 96.21 Ambient Temp.: 84.69 F
15:26:29.524 -> Average Cal sensor Temp.: 96.15 Ambient Temp.: 84.65 F
15:26:31.080 -> Average Cal sensor Temp.: 96.06 Ambient Temp.: 84.61 F
15:26:32.666 -> Average Cal sensor Temp.: 96.05 Ambient Temp.: 84.65 F
15:26:34.222 -> Average Cal sensor Temp.: 96.11 Ambient Temp.: 84.69 F
15:26:35.776 -> Average Cal sensor Temp.: 96.03 Ambient Temp.: 84.65 F
15:26:37.328 -> Average Cal sensor Temp.: 95.93 Ambient Temp.: 84.61 F
15:26:38.880 -> Average Cal sensor Temp.: 96.06 Ambient Temp.: 84.65 F
15:26:40.430 -> Average Cal sensor Temp.: 96.52 Ambient Temp.: 84.65 F
15:26:42.012 -> Average Cal sensor Temp.: 95.61 Ambient Temp.: 84.61 F
15:26:43.564 -> Average Cal sensor Temp.: 96.05 Ambient Temp.: 84.65 F
15:26:45.113 -> Average Cal sensor Temp.: 96.40 Ambient Temp.: 84.65 F
15:26:46.665 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 84.54 F
15:26:48.217 -> Average Cal sensor Temp.: 96.35 Ambient Temp.: 84.58 F
15:26:49.773 -> Average Cal sensor Temp.: 95.80 Ambient Temp.: 84.58 F
15:26:51.330 -> Average Cal sensor Temp.: 96.07 Ambient Temp.: 84.58 F
15:26:52.889 -> Average Cal sensor Temp.: 96.10 Ambient Temp.: 84.54 F
15:26:54.442 -> Average Cal sensor Temp.: 96.00 Ambient Temp.: 84.51 F
15:26:55.998 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 84.54 F
15:26:57.560 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 84.58 F
15:26:59.114 -> Average Cal sensor Temp.: 95.93 Ambient Temp.: 84.54 F
15:27:00.671 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 84.51 F
15:27:02.220 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 84.47 F
15:27:03.774 -> Average Cal sensor Temp.: 96.14 Ambient Temp.: 84.40 F
15:27:05.356 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 84.47 F
15:27:06.911 -> Average Cal sensor Temp.: 95.88 Ambient Temp.: 84.47 F
15:27:08.466 -> Average Cal sensor Temp.: 95.80 Ambient Temp.: 84.47 F
15:27:10.022 -> Average Cal sensor Temp.: 95.75 Ambient Temp.: 84.40 F
15:27:11.572 -> Average Cal sensor Temp.: 96.01 Ambient Temp.: 84.40 F
15:27:13.126 -> Average Cal sensor Temp.: 96.00 Ambient Temp.: 84.33 F
15:27:14.678 -> Average Cal sensor Temp.: 96.25 Ambient Temp.: 84.36 F
15:27:16.230 -> Average Cal sensor Temp.: 96.21 Ambient Temp.: 84.33 F
15:27:17.787 -> Average Cal sensor Temp.: 96.20 Ambient Temp.: 84.36 F
15:27:19.342 -> Average Cal sensor Temp.: 96.16 Ambient Temp.: 84.36 F
15:27:20.897 -> Average Cal sensor Temp.: 96.16 Ambient Temp.: 84.33 F
15:27:22.452 -> Average Cal sensor Temp.: 96.04 Ambient Temp.: 84.36 F
15:27:24.035 -> Average Cal sensor Temp.: 95.90 Ambient Temp.: 84.40 F
15:27:25.591 -> Average Cal sensor Temp.: 95.95 Ambient Temp.: 84.33 F
15:27:27.149 -> Average Cal sensor Temp.: 95.68 Ambient Temp.: 84.36 F
15:27:28.704 -> Average Cal sensor Temp.: 95.72 Ambient Temp.: 84.29 F
15:27:30.265 -> Average Cal sensor Temp.: 95.98 Ambient Temp.: 84.33 F
15:27:31.822 -> Average Cal sensor Temp.: 95.84 Ambient Temp.: 84.33 F
15:27:33.373 -> Average Cal sensor Temp.: 95.84 Ambient Temp.: 84.29 F
15:27:34.928 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 84.25 F
15:27:36.487 -> Average Cal sensor Temp.: 96.65 Ambient Temp.: 84.29 F
15:27:38.040 -> Average Cal sensor Temp.: 96.86 Ambient Temp.: 84.33 F
15:27:39.591 -> Average Cal sensor Temp.: 96.62 Ambient Temp.: 84.33 F
15:27:41.149 -> Average Cal sensor Temp.: 96.20 Ambient Temp.: 84.36 F
15:27:42.703 -> Average Cal sensor Temp.: 96.02 Ambient Temp.: 84.29 F
15:27:44.257 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 84.33 F
15:27:45.811 -> Average Cal sensor Temp.: 96.15 Ambient Temp.: 84.33 F
15:27:47.364 -> Average Cal sensor Temp.: 96.22 Ambient Temp.: 84.33 F
15:27:48.922 -> Average Cal sensor Temp.: 96.25 Ambient Temp.: 84.36 F
15:27:50.473 -> Average Cal sensor Temp.: 96.58 Ambient Temp.: 84.36 F
15:27:52.059 -> Average Cal sensor Temp.: 96.40 Ambient Temp.: 84.36 F
15:27:53.610 -> Average Cal sensor Temp.: 96.31 Ambient Temp.: 84.40 F
15:27:55.165 -> Average Cal sensor Temp.: 96.12 Ambient Temp.: 84.40 F
15:27:56.720 -> Average Cal sensor Temp.: 96.27 Ambient Temp.: 84.40 F
15:27:58.271 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 84.40 F
15:27:59.824 -> Average Cal sensor Temp.: 96.20 Ambient Temp.: 84.40 F
15:28:01.380 -> Average Cal sensor Temp.: 96.28 Ambient Temp.: 84.40 F
15:28:02.930 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 84.51 F
15:28:04.511 -> Average Cal sensor Temp.: 96.07 Ambient Temp.: 84.47 F
15:28:06.069 -> Average Cal sensor Temp.: 96.05 Ambient Temp.: 84.40 F
15:28:07.624 -> Average Cal sensor Temp.: 96.31 Ambient Temp.: 84.40 F
15:28:09.175 -> Average Cal sensor Temp.: 96.48 Ambient Temp.: 84.47 F
15:28:10.732 -> Average Cal sensor Temp.: 96.48 Ambient Temp.: 84.40 F
15:28:12.288 -> Average Cal sensor Temp.: 96.43 Ambient Temp.: 84.54 F
15:28:13.838 -> Average Cal sensor Temp.: 96.26 Ambient Temp.: 84.47 F
15:28:15.385 -> Average Cal sensor Temp.: 96.23 Ambient Temp.: 84.40 F
15:28:16.940 -> Average Cal sensor Temp.: 96.34 Ambient Temp.: 84.54 F
15:28:18.494 -> Average Cal sensor Temp.: 96.00 Ambient Temp.: 84.51 F
15:28:20.082 -> Average Cal sensor Temp.: 96.14 Ambient Temp.: 84.51 F
15:28:21.608 -> Average Cal sensor Temp.: 95.94 Ambient Temp.: 84.51 F
15:28:23.195 -> Average Cal sensor Temp.: 96.18 Ambient Temp.: 84.51 F
15:28:24.749 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 84.54 F
15:28:26.306 -> Average Cal sensor Temp.: 96.37 Ambient Temp.: 84.51 F
15:28:27.863 -> Average Cal sensor Temp.: 96.34 Ambient Temp.: 84.54 F
15:28:29.421 -> Average Cal sensor Temp.: 96.45 Ambient Temp.: 84.54 F
15:28:30.975 -> Average Cal sensor Temp.: 96.50 Ambient Temp.: 84.58 F
15:28:32.534 -> Average Cal sensor Temp.: 96.49 Ambient Temp.: 84.61 F
15:28:34.078 -> Average Cal sensor Temp.: 96.34 Ambient Temp.: 84.58 F
15:28:35.631 -> Average Cal sensor Temp.: 96.48 Ambient Temp.: 84.61 F
15:28:37.186 -> Average Cal sensor Temp.: 96.47 Ambient Temp.: 84.58 F
15:28:38.736 -> Average Cal sensor Temp.: 96.34 Ambient Temp.: 84.65 F
15:28:40.291 -> Average Cal sensor Temp.: 96.39 Ambient Temp.: 84.61 F
15:28:41.876 -> Average Cal sensor Temp.: 96.21 Ambient Temp.: 84.61 F
15:28:43.428 -> Average Cal sensor Temp.: 96.03 Ambient Temp.: 84.61 F
15:28:44.980 -> Average Cal sensor Temp.: 96.06 Ambient Temp.: 84.65 F
15:28:46.534 -> Average Cal sensor Temp.: 96.40 Ambient Temp.: 84.69 F
15:28:48.084 -> Average Cal sensor Temp.: 96.42 Ambient Temp.: 84.72 F
15:28:49.630 -> Average Cal sensor Temp.: 96.38 Ambient Temp.: 84.72 F
15:28:51.217 -> Average Cal sensor Temp.: 96.38 Ambient Temp.: 84.72 F
15:28:52.763 -> Average Cal sensor Temp.: 95.97 Ambient Temp.: 84.72 F
15:28:54.314 -> Average Cal sensor Temp.: 96.26 Ambient Temp.: 84.76 F
15:28:55.860 -> Average Cal sensor Temp.: 96.19 Ambient Temp.: 84.79 F
15:28:57.413 -> Average Cal sensor Temp.: 96.54 Ambient Temp.: 84.76 F
15:28:59.003 -> Average Cal sensor Temp.: 96.60 Ambient Temp.: 84.79 F
15:29:00.557 -> Average Cal sensor Temp.: 96.56 Ambient Temp.: 84.87 F
15:29:02.116 -> Average Cal sensor Temp.: 96.51 Ambient Temp.: 84.87 F
15:29:03.669 -> Average Cal sensor Temp.: 96.58 Ambient Temp.: 84.90 F
15:29:05.220 -> Average Cal sensor Temp.: 96.25 Ambient Temp.: 84.90 F
15:29:06.772 -> Average Cal sensor Temp.: 96.26 Ambient Temp.: 84.90 F
15:29:08.322 -> Average Cal sensor Temp.: 96.13 Ambient Temp.: 84.94 F
15:29:09.878 -> Average Cal sensor Temp.: 96.48 Ambient Temp.: 84.97 F
15:29:11.434 -> Average Cal sensor Temp.: 96.12 Ambient Temp.: 85.01 F
15:29:12.986 -> Average Cal sensor Temp.: 96.15 Ambient Temp.: 85.01 F
15:29:14.538 -> Average Cal sensor Temp.: 96.26 Ambient Temp.: 85.01 F
15:29:16.097 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 85.05 F
15:29:17.682 -> Average Cal sensor Temp.: 96.14 Ambient Temp.: 85.01 F
15:29:19.234 -> Average Cal sensor Temp.: 96.18 Ambient Temp.: 85.08 F
15:29:20.789 -> Average Cal sensor Temp.: 96.29 Ambient Temp.: 85.08 F
15:29:22.346 -> Average Cal sensor Temp.: 96.21 Ambient Temp.: 85.12 F
15:29:23.899 -> Average Cal sensor Temp.: 96.42 Ambient Temp.: 85.12 F
15:29:25.452 -> Average Cal sensor Temp.: 96.18 Ambient Temp.: 85.15 F
15:29:27.007 -> Average Cal sensor Temp.: 96.43 Ambient Temp.: 85.15 F
15:29:28.556 -> Average Cal sensor Temp.: 96.26 Ambient Temp.: 85.19 F
15:29:30.105 -> Average Cal sensor Temp.: 96.38 Ambient Temp.: 85.26 F
15:29:31.659 -> Average Cal sensor Temp.: 96.27 Ambient Temp.: 85.26 F
15:29:33.211 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 85.33 F
15:29:34.766 -> Average Cal sensor Temp.: 96.14 Ambient Temp.: 85.30 F
15:29:36.324 -> Average Cal sensor Temp.: 96.04 Ambient Temp.: 85.30 F
15:29:37.910 -> Average Cal sensor Temp.: 96.29 Ambient Temp.: 85.37 F
15:29:39.468 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 85.37 F
15:29:41.020 -> Average Cal sensor Temp.: 96.59 Ambient Temp.: 85.41 F
15:29:42.569 -> Average Cal sensor Temp.: 96.51 Ambient Temp.: 85.41 F
15:29:44.125 -> Average Cal sensor Temp.: 96.38 Ambient Temp.: 85.41 F
15:29:45.681 -> Average Cal sensor Temp.: 96.13 Ambient Temp.: 85.44 F
15:29:47.233 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 85.41 F
15:29:48.789 -> Average Cal sensor Temp.: 96.39 Ambient Temp.: 85.48 F
15:29:50.343 -> Average Cal sensor Temp.: 96.42 Ambient Temp.: 85.44 F
15:29:51.898 -> Average Cal sensor Temp.: 96.60 Ambient Temp.: 85.48 F
15:29:53.456 -> Average Cal sensor Temp.: 96.58 Ambient Temp.: 85.51 F
15:29:55.015 -> Average Cal sensor Temp.: 96.36 Ambient Temp.: 85.51 F
15:29:56.569 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 85.55 F
15:29:58.122 -> Average Cal sensor Temp.: 96.35 Ambient Temp.: 85.55 F
15:29:59.668 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 85.55 F
15:30:01.251 -> Average Cal sensor Temp.: 96.42 Ambient Temp.: 85.59 F
15:30:01.836 -> Ambient Temp.: 85.59 F
15:30:02.420 -> Ambient Temp.: 85.59 F
15:30:02.967 -> Ambient Temp.: 85.59 F
15:30:04.523 -> Average Cal sensor Temp.: 96.37 Ambient Temp.: 85.59 F
15:30:06.079 -> Average Cal sensor Temp.: 96.62 Ambient Temp.: 85.69 F
15:30:07.636 -> Average Cal sensor Temp.: 96.52 Ambient Temp.: 85.73 F
15:30:09.187 -> Average Cal sensor Temp.: 96.45 Ambient Temp.: 85.69 F
15:30:10.771 -> Average Cal sensor Temp.: 96.41 Ambient Temp.: 85.77 F
15:30:12.318 -> Average Cal sensor Temp.: 96.38 Ambient Temp.: 85.77 F
15:30:13.878 -> Average Cal sensor Temp.: 96.23 Ambient Temp.: 85.69 F
15:30:15.429 -> Average Cal sensor Temp.: 96.39 Ambient Temp.: 85.73 F
15:30:16.980 -> Average Cal sensor Temp.: 96.58 Ambient Temp.: 85.77 F
15:30:18.530 -> Average Cal sensor Temp.: 96.41 Ambient Temp.: 85.80 F
15:30:20.080 -> Average Cal sensor Temp.: 96.32 Ambient Temp.: 85.84 F
15:30:21.665 -> Average Cal sensor Temp.: 96.14 Ambient Temp.: 85.80 F
15:30:23.213 -> Average Cal sensor Temp.: 96.23 Ambient Temp.: 85.77 F
15:30:24.759 -> Average Cal sensor Temp.: 95.77 Ambient Temp.: 85.77 F
15:30:26.318 -> Average Cal sensor Temp.: 95.87 Ambient Temp.: 85.84 F
15:30:27.872 -> Average Cal sensor Temp.: 96.10 Ambient Temp.: 85.77 F
15:30:29.423 -> Average Cal sensor Temp.: 96.12 Ambient Temp.: 85.77 F
15:30:30.974 -> Average Cal sensor Temp.: 96.13 Ambient Temp.: 85.77 F
15:30:32.531 -> Average Cal sensor Temp.: 96.21 Ambient Temp.: 85.84 F
15:30:34.081 -> Average Cal sensor Temp.: 96.58 Ambient Temp.: 85.80 F
15:30:35.663 -> Average Cal sensor Temp.: 96.60 Ambient Temp.: 85.80 F
15:30:37.217 -> Average Cal sensor Temp.: 96.86 Ambient Temp.: 85.84 F
15:30:38.774 -> Average Cal sensor Temp.: 96.87 Ambient Temp.: 85.80 F
15:30:40.330 -> Average Cal sensor Temp.: 96.35 Ambient Temp.: 85.80 F
15:30:41.883 -> Average Cal sensor Temp.: 96.65 Ambient Temp.: 85.77 F
15:30:43.435 -> Average Cal sensor Temp.: 96.90 Ambient Temp.: 85.87 F
15:30:44.985 -> Average Cal sensor Temp.: 96.60 Ambient Temp.: 85.84 F
15:30:46.545 -> Average Cal sensor Temp.: 96.51 Ambient Temp.: 85.87 F
15:30:48.110 -> Average Cal sensor Temp.: 96.72 Ambient Temp.: 85.87 F
15:30:49.664 -> Average Cal sensor Temp.: 96.49 Ambient Temp.: 85.87 F
15:30:51.221 -> Average Cal sensor Temp.: 96.50 Ambient Temp.: 85.87 F
15:30:52.775 -> Average Cal sensor Temp.: 96.10 Ambient Temp.: 85.87 F
15:30:54.332 -> Average Cal sensor Temp.: 96.53 Ambient Temp.: 85.87 F
15:30:55.887 -> Average Cal sensor Temp.: 96.70 Ambient Temp.: 85.84 F
15:30:57.436 -> Average Cal sensor Temp.: 96.43 Ambient Temp.: 85.87 F
15:30:58.989 -> Average Cal sensor Temp.: 96.40 Ambient Temp.: 85.87 F
15:31:00.577 -> Average Cal sensor Temp.: 96.57 Ambient Temp.: 85.87 F
15:31:02.123 -> Average Cal sensor Temp.: 96.80 Ambient Temp.: 85.91 F
15:31:03.675 -> Average Cal sensor Temp.: 96.65 Ambient Temp.: 85.91 F
15:31:05.229 -> Average Cal sensor Temp.: 96.80 Ambient Temp.: 85.87 F
15:31:06.784 -> Average Cal sensor Temp.: 96.72 Ambient Temp.: 85.87 F
15:31:08.338 -> Average Cal sensor Temp.: 96.58 Ambient Temp.: 85.91 F
15:31:09.892 -> Average Cal sensor Temp.: 96.28 Ambient Temp.: 85.84 F
15:31:11.444 -> Average Cal sensor Temp.: 96.38 Ambient Temp.: 85.84 F
15:31:13.033 -> Average Cal sensor Temp.: 96.55 Ambient Temp.: 85.87 F
15:31:14.588 -> Average Cal sensor Temp.: 96.78 Ambient Temp.: 85.84 F
15:31:16.144 -> Average Cal sensor Temp.: 96.71 Ambient Temp.: 85.84 F
15:31:17.702 -> Average Cal sensor Temp.: 96.48 Ambient Temp.: 85.84 F
15:31:19.257 -> Average Cal sensor Temp.: 96.37 Ambient Temp.: 85.84 F
15:31:20.814 -> Average Cal sensor Temp.: 95.94 Ambient Temp.: 85.87 F
15:31:22.362 -> Average Cal sensor Temp.: 96.37 Ambient Temp.: 85.87 F
15:31:23.910 -> Average Cal sensor Temp.: 96.21 Ambient Temp.: 85.87 F
15:31:25.459 -> Average Cal sensor Temp.: 95.86 Ambient Temp.: 85.91 F
15:31:27.009 -> Average Cal sensor Temp.: 95.86 Ambient Temp.: 85.84 F
15:31:28.566 -> Average Cal sensor Temp.: 96.21 Ambient Temp.: 85.80 F
15:31:30.123 -> Average Cal sensor Temp.: 96.09 Ambient Temp.: 85.80 F
15:31:31.683 -> Average Cal sensor Temp.: 96.44 Ambient Temp.: 85.80 F
15:31:33.238 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 85.80 F
15:31:34.828 -> Average Cal sensor Temp.: 96.25 Ambient Temp.: 85.87 F
15:31:36.384 -> Average Cal sensor Temp.: 96.42 Ambient Temp.: 85.87 F
15:31:37.933 -> Average Cal sensor Temp.: 96.46 Ambient Temp.: 85.87 F
15:31:39.487 -> Average Cal sensor Temp.: 96.61 Ambient Temp.: 85.84 F
15:31:41.041 -> Average Cal sensor Temp.: 96.16 Ambient Temp.: 85.80 F
15:31:42.595 -> Average Cal sensor Temp.: 95.99 Ambient Temp.: 85.77 F
15:31:44.154 -> Average Cal sensor Temp.: 95.72 Ambient Temp.: 85.73 F
15:31:45.710 -> Average Cal sensor Temp.: 96.24 Ambient Temp.: 85.69 F
15:31:47.269 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 85.69 F
15:31:48.826 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 85.66 F
15:31:50.384 -> Average Cal sensor Temp.: 96.10 Ambient Temp.: 85.66 F
15:31:51.932 -> Average Cal sensor Temp.: 96.06 Ambient Temp.: 85.66 F
15:31:53.488 -> Average Cal sensor Temp.: 95.88 Ambient Temp.: 85.59 F
15:31:55.043 -> Average Cal sensor Temp.: 96.15 Ambient Temp.: 85.51 F
15:31:56.598 -> Average Cal sensor Temp.: 96.41 Ambient Temp.: 85.51 F
15:31:58.153 -> Average Cal sensor Temp.: 96.51 Ambient Temp.: 85.55 F
15:31:59.710 -> Average Cal sensor Temp.: 96.38 Ambient Temp.: 85.55 F
15:32:01.266 -> Average Cal sensor Temp.: 96.37 Ambient Temp.: 85.51 F
15:32:02.814 -> Average Cal sensor Temp.: 96.48 Ambient Temp.: 85.51 F
15:32:04.368 -> Average Cal sensor Temp.: 96.13 Ambient Temp.: 85.51 F
15:32:05.924 -> Average Cal sensor Temp.: 96.61 Ambient Temp.: 85.51 F
15:32:07.479 -> Average Cal sensor Temp.: 96.60 Ambient Temp.: 85.48 F
15:32:09.036 -> Average Cal sensor Temp.: 96.54 Ambient Temp.: 85.44 F
15:32:10.589 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 85.41 F
15:32:12.170 -> Average Cal sensor Temp.: 96.45 Ambient Temp.: 85.44 F
15:32:13.729 -> Average Cal sensor Temp.: 96.68 Ambient Temp.: 85.41 F
15:32:15.280 -> Average Cal sensor Temp.: 96.14 Ambient Temp.: 85.41 F
15:32:16.831 -> Average Cal sensor Temp.: 96.47 Ambient Temp.: 85.37 F
15:32:18.382 -> Average Cal sensor Temp.: 96.51 Ambient Temp.: 85.33 F
15:32:19.936 -> Average Cal sensor Temp.: 96.26 Ambient Temp.: 85.30 F
15:32:21.495 -> Average Cal sensor Temp.: 96.19 Ambient Temp.: 85.30 F
15:32:23.045 -> Average Cal sensor Temp.: 96.12 Ambient Temp.: 85.30 F
15:32:24.601 -> Average Cal sensor Temp.: 96.72 Ambient Temp.: 85.26 F
15:32:26.155 -> Average Cal sensor Temp.: 96.54 Ambient Temp.: 85.19 F
15:32:26.743 -> Ambient Temp.: 85.19 F
15:32:27.330 -> Ambient Temp.: 85.19 F
15:32:27.913 -> Ambient Temp.: 85.19 F
15:32:29.471 -> Average Cal sensor Temp.: 96.57 Ambient Temp.: 85.30 F
15:32:31.025 -> Average Cal sensor Temp.: 96.22 Ambient Temp.: 85.19 F
15:32:32.574 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 85.19 F
15:32:34.130 -> Average Cal sensor Temp.: 96.45 Ambient Temp.: 85.12 F
15:32:35.683 -> Average Cal sensor Temp.: 96.52 Ambient Temp.: 85.05 F
15:32:37.233 -> Average Cal sensor Temp.: 96.53 Ambient Temp.: 85.08 F
15:32:38.792 -> Average Cal sensor Temp.: 96.28 Ambient Temp.: 85.05 F
15:32:40.347 -> Average Cal sensor Temp.: 96.64 Ambient Temp.: 85.12 F
15:32:41.898 -> Average Cal sensor Temp.: 96.43 Ambient Temp.: 85.12 F
15:32:43.452 -> Average Cal sensor Temp.: 96.15 Ambient Temp.: 85.08 F
15:32:45.035 -> Average Cal sensor Temp.: 95.89 Ambient Temp.: 85.08 F
15:32:46.561 -> Average Cal sensor Temp.: 96.03 Ambient Temp.: 85.01 F
15:32:48.148 -> Average Cal sensor Temp.: 96.14 Ambient Temp.: 85.01 F
15:32:49.704 -> Average Cal sensor Temp.: 96.35 Ambient Temp.: 85.01 F
15:32:51.259 -> Average Cal sensor Temp.: 96.62 Ambient Temp.: 84.97 F
15:32:52.813 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 84.94 F
15:32:54.370 -> Average Cal sensor Temp.: 96.68 Ambient Temp.: 84.97 F
15:32:55.924 -> Average Cal sensor Temp.: 96.48 Ambient Temp.: 84.94 F
15:32:57.476 -> Average Cal sensor Temp.: 96.61 Ambient Temp.: 84.90 F
15:32:59.031 -> Average Cal sensor Temp.: 96.80 Ambient Temp.: 84.94 F
15:33:00.584 -> Average Cal sensor Temp.: 96.80 Ambient Temp.: 84.97 F
15:33:02.135 -> Average Cal sensor Temp.: 96.58 Ambient Temp.: 84.90 F
15:33:03.693 -> Average Cal sensor Temp.: 96.95 Ambient Temp.: 84.94 F
15:33:05.282 -> Average Cal sensor Temp.: 96.54 Ambient Temp.: 84.87 F
15:33:06.836 -> Average Cal sensor Temp.: 96.43 Ambient Temp.: 84.87 F
15:33:08.387 -> Average Cal sensor Temp.: 96.12 Ambient Temp.: 84.79 F
15:33:09.943 -> Average Cal sensor Temp.: 96.46 Ambient Temp.: 84.87 F
15:33:11.494 -> Average Cal sensor Temp.: 96.37 Ambient Temp.: 84.87 F
15:33:13.040 -> Average Cal sensor Temp.: 96.43 Ambient Temp.: 84.87 F
15:33:14.595 -> Average Cal sensor Temp.: 96.31 Ambient Temp.: 84.76 F
15:33:16.150 -> Average Cal sensor Temp.: 96.29 Ambient Temp.: 84.87 F
15:33:17.736 -> Average Cal sensor Temp.: 96.47 Ambient Temp.: 84.79 F
15:33:19.288 -> Average Cal sensor Temp.: 96.80 Ambient Temp.: 84.79 F
15:33:20.842 -> Average Cal sensor Temp.: 96.52 Ambient Temp.: 84.76 F
15:33:22.391 -> Average Cal sensor Temp.: 96.51 Ambient Temp.: 84.76 F
15:33:23.944 -> Average Cal sensor Temp.: 96.68 Ambient Temp.: 84.79 F
15:33:25.493 -> Average Cal sensor Temp.: 96.83 Ambient Temp.: 84.79 F
15:33:27.073 -> Average Cal sensor Temp.: 96.46 Ambient Temp.: 84.76 F
15:33:28.628 -> Average Cal sensor Temp.: 96.20 Ambient Temp.: 84.76 F
15:33:30.185 -> Average Cal sensor Temp.: 96.03 Ambient Temp.: 84.72 F
15:33:31.740 -> Average Cal sensor Temp.: 96.35 Ambient Temp.: 84.61 F
15:33:33.296 -> Average Cal sensor Temp.: 96.67 Ambient Temp.: 84.65 F
15:33:34.855 -> Average Cal sensor Temp.: 96.37 Ambient Temp.: 84.65 F
15:33:36.413 -> Average Cal sensor Temp.: 96.46 Ambient Temp.: 84.58 F
15:33:37.972 -> Average Cal sensor Temp.: 96.31 Ambient Temp.: 84.61 F
15:33:39.530 -> Average Cal sensor Temp.: 96.64 Ambient Temp.: 84.61 F
15:33:41.082 -> Average Cal sensor Temp.: 96.01 Ambient Temp.: 84.47 F
15:33:42.638 -> Average Cal sensor Temp.: 96.13 Ambient Temp.: 84.51 F
15:33:44.188 -> Average Cal sensor Temp.: 96.09 Ambient Temp.: 84.47 F
15:33:45.745 -> Average Cal sensor Temp.: 96.10 Ambient Temp.: 84.47 F
15:33:47.299 -> Average Cal sensor Temp.: 95.98 Ambient Temp.: 84.36 F
15:33:48.853 -> Average Cal sensor Temp.: 96.33 Ambient Temp.: 84.33 F
15:33:50.407 -> Average Cal sensor Temp.: 96.68 Ambient Temp.: 84.36 F
15:33:51.960 -> Average Cal sensor Temp.: 95.97 Ambient Temp.: 84.33 F
15:33:53.525 -> Average Cal sensor Temp.: 96.41 Ambient Temp.: 84.36 F
15:33:55.081 -> Average Cal sensor Temp.: 96.53 Ambient Temp.: 84.29 F
15:33:56.639 -> Average Cal sensor Temp.: 96.66 Ambient Temp.: 84.25 F
15:33:58.193 -> Average Cal sensor Temp.: 96.62 Ambient Temp.: 84.25 F
15:33:59.750 -> Average Cal sensor Temp.: 96.77 Ambient Temp.: 84.29 F
15:34:01.297 -> Average Cal sensor Temp.: 96.52 Ambient Temp.: 84.25 F
15:34:02.885 -> Average Cal sensor Temp.: 96.68 Ambient Temp.: 84.25 F
15:34:04.410 -> Average Cal sensor Temp.: 96.31 Ambient Temp.: 84.25 F
15:34:05.998 -> Average Cal sensor Temp.: 95.91 Ambient Temp.: 84.18 F
15:34:07.551 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 84.22 F
15:34:09.107 -> Average Cal sensor Temp.: 96.33 Ambient Temp.: 84.15 F
15:34:10.661 -> Average Cal sensor Temp.: 96.44 Ambient Temp.: 84.22 F
15:34:12.215 -> Average Cal sensor Temp.: 96.37 Ambient Temp.: 84.15 F
15:34:13.768 -> Average Cal sensor Temp.: 96.42 Ambient Temp.: 84.15 F
15:34:15.321 -> Average Cal sensor Temp.: 96.33 Ambient Temp.: 84.22 F
15:34:16.880 -> Average Cal sensor Temp.: 96.22 Ambient Temp.: 84.22 F
15:34:18.432 -> Average Cal sensor Temp.: 96.58 Ambient Temp.: 84.25 F
15:34:19.984 -> Average Cal sensor Temp.: 96.22 Ambient Temp.: 84.22 F
15:34:21.535 -> Average Cal sensor Temp.: 96.77 Ambient Temp.: 84.25 F
15:34:23.091 -> Average Cal sensor Temp.: 96.44 Ambient Temp.: 84.22 F
15:34:24.641 -> Average Cal sensor Temp.: 96.46 Ambient Temp.: 84.18 F
15:34:26.229 -> Average Cal sensor Temp.: 96.27 Ambient Temp.: 84.22 F
15:34:27.782 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 84.15 F
15:34:29.332 -> Average Cal sensor Temp.: 96.16 Ambient Temp.: 84.15 F
15:34:30.890 -> Average Cal sensor Temp.: 96.46 Ambient Temp.: 84.18 F
15:34:32.445 -> Average Cal sensor Temp.: 96.06 Ambient Temp.: 84.15 F
15:34:33.999 -> Average Cal sensor Temp.: 96.26 Ambient Temp.: 84.15 F
15:34:35.555 -> Average Cal sensor Temp.: 96.36 Ambient Temp.: 84.18 F
15:34:37.109 -> Average Cal sensor Temp.: 96.42 Ambient Temp.: 84.22 F
15:34:38.669 -> Average Cal sensor Temp.: 96.49 Ambient Temp.: 84.29 F
15:34:40.221 -> Average Cal sensor Temp.: 96.15 Ambient Temp.: 84.22 F
15:34:41.774 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 84.22 F
15:34:43.322 -> Average Cal sensor Temp.: 96.20 Ambient Temp.: 84.22 F
15:34:44.876 -> Average Cal sensor Temp.: 96.54 Ambient Temp.: 84.22 F
15:34:46.463 -> Average Cal sensor Temp.: 96.55 Ambient Temp.: 84.18 F
15:34:48.019 -> Average Cal sensor Temp.: 96.56 Ambient Temp.: 84.18 F
15:34:49.576 -> Average Cal sensor Temp.: 96.38 Ambient Temp.: 84.22 F
15:34:51.130 -> Average Cal sensor Temp.: 96.52 Ambient Temp.: 84.22 F
15:34:52.685 -> Average Cal sensor Temp.: 96.62 Ambient Temp.: 84.18 F
15:34:54.243 -> Average Cal sensor Temp.: 96.63 Ambient Temp.: 84.25 F
15:34:55.800 -> Average Cal sensor Temp.: 96.65 Ambient Temp.: 84.25 F
15:34:57.352 -> Average Cal sensor Temp.: 96.54 Ambient Temp.: 84.22 F
15:34:58.907 -> Average Cal sensor Temp.: 96.85 Ambient Temp.: 84.22 F
15:35:00.465 -> Average Cal sensor Temp.: 96.60 Ambient Temp.: 84.25 F
15:35:02.015 -> Average Cal sensor Temp.: 96.70 Ambient Temp.: 84.29 F
15:35:03.565 -> Average Cal sensor Temp.: 96.49 Ambient Temp.: 84.29 F
15:35:05.121 -> Average Cal sensor Temp.: 96.31 Ambient Temp.: 84.29 F
15:35:06.675 -> Average Cal sensor Temp.: 96.50 Ambient Temp.: 84.33 F
15:35:08.229 -> Average Cal sensor Temp.: 96.52 Ambient Temp.: 84.33 F
15:35:09.813 -> Average Cal sensor Temp.: 96.65 Ambient Temp.: 84.40 F
15:35:11.370 -> Average Cal sensor Temp.: 96.38 Ambient Temp.: 84.36 F
15:35:12.920 -> Average Cal sensor Temp.: 96.66 Ambient Temp.: 84.36 F
15:35:14.469 -> Average Cal sensor Temp.: 96.56 Ambient Temp.: 84.36 F
15:35:16.025 -> Average Cal sensor Temp.: 96.19 Ambient Temp.: 84.51 F
15:35:17.576 -> Average Cal sensor Temp.: 96.38 Ambient Temp.: 84.40 F
15:35:19.132 -> Average Cal sensor Temp.: 96.49 Ambient Temp.: 84.47 F
15:35:20.685 -> Average Cal sensor Temp.: 96.37 Ambient Temp.: 84.47 F
15:35:22.271 -> Average Cal sensor Temp.: 96.29 Ambient Temp.: 84.51 F
15:35:23.825 -> Average Cal sensor Temp.: 96.39 Ambient Temp.: 84.54 F
15:35:25.374 -> Average Cal sensor Temp.: 96.43 Ambient Temp.: 84.51 F
15:35:26.930 -> Average Cal sensor Temp.: 96.22 Ambient Temp.: 84.51 F
15:35:28.486 -> Average Cal sensor Temp.: 96.36 Ambient Temp.: 84.58 F
15:35:30.042 -> Average Cal sensor Temp.: 96.22 Ambient Temp.: 84.54 F
15:35:31.590 -> Average Cal sensor Temp.: 96.45 Ambient Temp.: 84.58 F
15:35:33.140 -> Average Cal sensor Temp.: 96.61 Ambient Temp.: 84.58 F
15:35:34.694 -> Average Cal sensor Temp.: 96.72 Ambient Temp.: 84.58 F
15:35:36.280 -> Average Cal sensor Temp.: 96.55 Ambient Temp.: 84.58 F
15:35:37.830 -> Average Cal sensor Temp.: 96.29 Ambient Temp.: 84.61 F
15:35:39.382 -> Average Cal sensor Temp.: 96.27 Ambient Temp.: 84.54 F
15:35:40.931 -> Average Cal sensor Temp.: 96.31 Ambient Temp.: 84.58 F
15:35:42.485 -> Average Cal sensor Temp.: 96.20 Ambient Temp.: 84.61 F
15:35:44.036 -> Average Cal sensor Temp.: 96.26 Ambient Temp.: 84.61 F
15:35:45.592 -> Average Cal sensor Temp.: 96.85 Ambient Temp.: 84.65 F
15:35:47.146 -> Average Cal sensor Temp.: 96.42 Ambient Temp.: 84.58 F
15:35:48.735 -> Average Cal sensor Temp.: 96.35 Ambient Temp.: 84.61 F
15:35:50.290 -> Average Cal sensor Temp.: 96.53 Ambient Temp.: 84.69 F
15:35:51.846 -> Average Cal sensor Temp.: 96.58 Ambient Temp.: 84.65 F
15:35:53.398 -> Average Cal sensor Temp.: 96.06 Ambient Temp.: 84.69 F
15:35:54.950 -> Average Cal sensor Temp.: 96.21 Ambient Temp.: 84.69 F
15:35:56.504 -> Average Cal sensor Temp.: 96.65 Ambient Temp.: 84.65 F
15:35:58.049 -> Average Cal sensor Temp.: 96.35 Ambient Temp.: 84.76 F
15:35:59.604 -> Average Cal sensor Temp.: 96.15 Ambient Temp.: 84.72 F
15:36:01.163 -> Average Cal sensor Temp.: 96.39 Ambient Temp.: 84.87 F
15:36:02.717 -> Average Cal sensor Temp.: 96.65 Ambient Temp.: 84.76 F
15:36:04.301 -> Average Cal sensor Temp.: 96.53 Ambient Temp.: 84.76 F
15:36:05.855 -> Average Cal sensor Temp.: 96.42 Ambient Temp.: 84.76 F
15:36:07.409 -> Average Cal sensor Temp.: 96.57 Ambient Temp.: 84.87 F
15:36:08.967 -> Average Cal sensor Temp.: 96.72 Ambient Temp.: 84.87 F
15:36:10.524 -> Average Cal sensor Temp.: 96.79 Ambient Temp.: 84.87 F
15:36:12.073 -> Average Cal sensor Temp.: 96.80 Ambient Temp.: 84.90 F
15:36:13.630 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 84.97 F
15:36:15.183 -> Average Cal sensor Temp.: 96.76 Ambient Temp.: 84.97 F
15:36:16.740 -> Average Cal sensor Temp.: 96.80 Ambient Temp.: 84.94 F
15:36:18.297 -> Average Cal sensor Temp.: 96.63 Ambient Temp.: 84.94 F
15:36:19.856 -> Average Cal sensor Temp.: 96.77 Ambient Temp.: 85.05 F
15:36:21.411 -> Average Cal sensor Temp.: 96.90 Ambient Temp.: 85.08 F
15:36:22.962 -> Average Cal sensor Temp.: 97.03 Ambient Temp.: 85.08 F
15:36:24.516 -> Average Cal sensor Temp.: 96.90 Ambient Temp.: 85.08 F
15:36:26.072 -> Average Cal sensor Temp.: 96.90 Ambient Temp.: 85.08 F
15:36:27.662 -> Average Cal sensor Temp.: 96.82 Ambient Temp.: 85.12 F
15:36:29.219 -> Average Cal sensor Temp.: 96.72 Ambient Temp.: 85.15 F
15:36:30.741 -> Average Cal sensor Temp.: 96.61 Ambient Temp.: 85.19 F
15:36:32.330 -> Average Cal sensor Temp.: 96.62 Ambient Temp.: 85.19 F
15:36:33.882 -> Average Cal sensor Temp.: 96.66 Ambient Temp.: 85.30 F
15:36:35.435 -> Average Cal sensor Temp.: 96.86 Ambient Temp.: 85.33 F
15:36:36.987 -> Average Cal sensor Temp.: 96.66 Ambient Temp.: 85.30 F
15:36:38.546 -> Average Cal sensor Temp.: 96.74 Ambient Temp.: 85.37 F
15:36:40.093 -> Average Cal sensor Temp.: 96.84 Ambient Temp.: 85.37 F
15:36:41.647 -> Average Cal sensor Temp.: 96.71 Ambient Temp.: 85.41 F
15:36:43.197 -> Average Cal sensor Temp.: 96.65 Ambient Temp.: 85.41 F
15:36:44.779 -> Average Cal sensor Temp.: 96.80 Ambient Temp.: 85.41 F
15:36:46.330 -> Average Cal sensor Temp.: 96.86 Ambient Temp.: 85.51 F
15:36:47.886 -> Average Cal sensor Temp.: 96.96 Ambient Temp.: 85.44 F
15:36:49.444 -> Average Cal sensor Temp.: 96.89 Ambient Temp.: 85.51 F
15:36:50.997 -> Average Cal sensor Temp.: 96.88 Ambient Temp.: 85.55 F
15:36:52.548 -> Average Cal sensor Temp.: 96.85 Ambient Temp.: 85.59 F
15:36:54.108 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 85.66 F
15:36:55.658 -> Average Cal sensor Temp.: 96.84 Ambient Temp.: 85.66 F
15:36:57.207 -> Average Cal sensor Temp.: 97.30 Ambient Temp.: 85.69 F
15:36:58.762 -> Average Cal sensor Temp.: 97.05 Ambient Temp.: 85.69 F
15:37:00.351 -> Average Cal sensor Temp.: 96.86 Ambient Temp.: 85.77 F
15:37:01.904 -> Average Cal sensor Temp.: 96.83 Ambient Temp.: 85.77 F
15:37:03.453 -> Average Cal sensor Temp.: 96.92 Ambient Temp.: 85.80 F
15:37:05.000 -> Average Cal sensor Temp.: 96.94 Ambient Temp.: 85.84 F
15:37:06.552 -> Average Cal sensor Temp.: 96.94 Ambient Temp.: 85.84 F
15:37:08.137 -> Average Cal sensor Temp.: 96.92 Ambient Temp.: 85.91 F
15:37:09.693 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 85.91 F
15:37:11.249 -> Average Cal sensor Temp.: 96.51 Ambient Temp.: 85.87 F
15:37:12.805 -> Average Cal sensor Temp.: 96.68 Ambient Temp.: 85.91 F
15:37:14.359 -> Average Cal sensor Temp.: 96.74 Ambient Temp.: 85.87 F
15:37:15.913 -> Average Cal sensor Temp.: 96.39 Ambient Temp.: 85.95 F
15:37:17.467 -> Average Cal sensor Temp.: 96.65 Ambient Temp.: 85.87 F
15:37:19.025 -> Average Cal sensor Temp.: 96.55 Ambient Temp.: 85.87 F
15:37:20.582 -> Average Cal sensor Temp.: 96.82 Ambient Temp.: 85.87 F
15:37:22.143 -> Average Cal sensor Temp.: 96.91 Ambient Temp.: 85.98 F
15:37:23.697 -> Average Cal sensor Temp.: 96.96 Ambient Temp.: 85.98 F
15:37:25.252 -> Average Cal sensor Temp.: 97.18 Ambient Temp.: 85.98 F
15:37:26.805 -> Average Cal sensor Temp.: 97.12 Ambient Temp.: 85.95 F
15:37:28.358 -> Average Cal sensor Temp.: 96.70 Ambient Temp.: 85.98 F
15:37:29.917 -> Average Cal sensor Temp.: 97.10 Ambient Temp.: 85.98 F
15:37:31.478 -> Average Cal sensor Temp.: 96.66 Ambient Temp.: 86.05 F
15:37:33.033 -> Average Cal sensor Temp.: 96.84 Ambient Temp.: 86.05 F
15:37:34.588 -> Average Cal sensor Temp.: 96.82 Ambient Temp.: 86.09 F
15:37:36.141 -> Average Cal sensor Temp.: 97.16 Ambient Temp.: 86.05 F
15:37:37.701 -> Average Cal sensor Temp.: 97.21 Ambient Temp.: 86.13 F
15:37:39.257 -> Average Cal sensor Temp.: 97.00 Ambient Temp.: 86.16 F
15:37:40.812 -> Average Cal sensor Temp.: 96.80 Ambient Temp.: 86.13 F
15:37:42.364 -> Average Cal sensor Temp.: 96.68 Ambient Temp.: 86.13 F
15:37:43.920 -> Average Cal sensor Temp.: 96.95 Ambient Temp.: 86.16 F
15:37:45.476 -> Average Cal sensor Temp.: 97.09 Ambient Temp.: 86.20 F
15:37:47.032 -> Average Cal sensor Temp.: 96.94 Ambient Temp.: 86.20 F
15:37:48.587 -> Average Cal sensor Temp.: 97.00 Ambient Temp.: 86.20 F
15:37:50.170 -> Average Cal sensor Temp.: 97.18 Ambient Temp.: 86.27 F
15:37:51.693 -> Average Cal sensor Temp.: 96.74 Ambient Temp.: 86.23 F
15:37:53.248 -> Average Cal sensor Temp.: 96.07 Ambient Temp.: 86.23 F
15:37:54.834 -> Average Cal sensor Temp.: 96.00 Ambient Temp.: 86.16 F
15:37:56.387 -> Average Cal sensor Temp.: 96.10 Ambient Temp.: 86.16 F
15:37:57.929 -> Average Cal sensor Temp.: 96.12 Ambient Temp.: 86.13 F
15:37:59.506 -> Average Cal sensor Temp.: 96.59 Ambient Temp.: 86.16 F
15:38:01.054 -> Average Cal sensor Temp.: 96.74 Ambient Temp.: 86.20 F
15:38:02.598 -> Average Cal sensor Temp.: 96.53 Ambient Temp.: 86.16 F
15:38:04.144 -> Average Cal sensor Temp.: 96.64 Ambient Temp.: 86.16 F
15:38:05.731 -> Average Cal sensor Temp.: 96.82 Ambient Temp.: 86.16 F
15:38:07.285 -> Average Cal sensor Temp.: 96.54 Ambient Temp.: 86.16 F
15:38:08.838 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 86.16 F
15:38:10.394 -> Average Cal sensor Temp.: 96.52 Ambient Temp.: 86.16 F
15:38:11.946 -> Average Cal sensor Temp.: 96.43 Ambient Temp.: 86.20 F
15:38:13.492 -> Average Cal sensor Temp.: 96.41 Ambient Temp.: 86.13 F
15:38:15.046 -> Average Cal sensor Temp.: 96.55 Ambient Temp.: 86.13 F
15:38:16.597 -> Average Cal sensor Temp.: 96.76 Ambient Temp.: 86.13 F
15:38:18.157 -> Average Cal sensor Temp.: 97.19 Ambient Temp.: 86.13 F
15:38:19.716 -> Average Cal sensor Temp.: 96.91 Ambient Temp.: 86.13 F
15:38:21.274 -> Average Cal sensor Temp.: 96.84 Ambient Temp.: 86.16 F
15:38:22.826 -> Average Cal sensor Temp.: 97.57 Ambient Temp.: 86.20 F
15:38:24.383 -> Average Cal sensor Temp.: 96.69 Ambient Temp.: 86.20 F
15:38:25.936 -> Average Cal sensor Temp.: 96.67 Ambient Temp.: 86.13 F
15:38:27.491 -> Average Cal sensor Temp.: 97.09 Ambient Temp.: 86.16 F
15:38:29.078 -> Average Cal sensor Temp.: 97.33 Ambient Temp.: 86.16 F
15:38:30.629 -> Average Cal sensor Temp.: 96.71 Ambient Temp.: 86.09 F
15:38:32.178 -> Average Cal sensor Temp.: 97.10 Ambient Temp.: 86.13 F
15:38:33.728 -> Average Cal sensor Temp.: 96.58 Ambient Temp.: 86.13 F
15:38:35.284 -> Average Cal sensor Temp.: 96.70 Ambient Temp.: 86.13 F
15:38:36.840 -> Average Cal sensor Temp.: 96.66 Ambient Temp.: 86.13 F
15:38:38.395 -> Average Cal sensor Temp.: 96.71 Ambient Temp.: 86.09 F
15:38:39.947 -> Average Cal sensor Temp.: 96.13 Ambient Temp.: 86.09 F
15:38:41.501 -> Average Cal sensor Temp.: 96.13 Ambient Temp.: 85.98 F
15:38:43.084 -> Average Cal sensor Temp.: 96.09 Ambient Temp.: 86.05 F
15:38:44.640 -> Average Cal sensor Temp.: 96.52 Ambient Temp.: 85.98 F
15:38:46.195 -> Average Cal sensor Temp.: 96.85 Ambient Temp.: 86.05 F
15:38:47.753 -> Average Cal sensor Temp.: 97.01 Ambient Temp.: 85.98 F
15:38:49.298 -> Average Cal sensor Temp.: 97.23 Ambient Temp.: 86.05 F
15:38:50.849 -> Average Cal sensor Temp.: 96.69 Ambient Temp.: 85.98 F
15:38:52.429 -> Average Cal sensor Temp.: 96.79 Ambient Temp.: 85.98 F
15:38:53.985 -> Average Cal sensor Temp.: 96.59 Ambient Temp.: 85.98 F
15:38:55.537 -> Average Cal sensor Temp.: 96.14 Ambient Temp.: 85.95 F
15:38:57.088 -> Average Cal sensor Temp.: 96.82 Ambient Temp.: 85.95 F
15:38:58.635 -> Average Cal sensor Temp.: 96.96 Ambient Temp.: 85.91 F
15:39:00.211 -> Average Cal sensor Temp.: 96.96 Ambient Temp.: 85.98 F
15:39:01.763 -> Average Cal sensor Temp.: 96.91 Ambient Temp.: 85.91 F
15:39:03.318 -> Average Cal sensor Temp.: 96.93 Ambient Temp.: 85.95 F
15:39:04.875 -> Average Cal sensor Temp.: 97.14 Ambient Temp.: 85.95 F
15:39:06.430 -> Average Cal sensor Temp.: 96.50 Ambient Temp.: 85.87 F
15:39:07.980 -> Average Cal sensor Temp.: 96.67 Ambient Temp.: 85.84 F
15:39:09.525 -> Average Cal sensor Temp.: 96.50 Ambient Temp.: 85.91 F
15:39:11.077 -> Average Cal sensor Temp.: 96.49 Ambient Temp.: 85.84 F
15:39:12.636 -> Average Cal sensor Temp.: 96.52 Ambient Temp.: 85.80 F
15:39:14.223 -> Average Cal sensor Temp.: 96.44 Ambient Temp.: 85.87 F
15:39:15.770 -> Average Cal sensor Temp.: 95.98 Ambient Temp.: 85.84 F
15:39:17.324 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 85.80 F
15:39:18.876 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 85.77 F
15:39:20.430 -> Average Cal sensor Temp.: 96.32 Ambient Temp.: 85.73 F
15:39:21.982 -> Average Cal sensor Temp.: 96.76 Ambient Temp.: 85.73 F
15:39:23.531 -> Average Cal sensor Temp.: 96.86 Ambient Temp.: 85.73 F
15:39:25.114 -> Average Cal sensor Temp.: 96.91 Ambient Temp.: 85.69 F
15:39:26.669 -> Average Cal sensor Temp.: 96.51 Ambient Temp.: 85.69 F
15:39:28.215 -> Average Cal sensor Temp.: 96.45 Ambient Temp.: 85.69 F
15:39:29.766 -> Average Cal sensor Temp.: 96.61 Ambient Temp.: 85.69 F
15:39:31.319 -> Average Cal sensor Temp.: 96.21 Ambient Temp.: 85.66 F
15:39:32.872 -> Average Cal sensor Temp.: 96.77 Ambient Temp.: 85.59 F
15:39:34.426 -> Average Cal sensor Temp.: 96.82 Ambient Temp.: 85.59 F
15:39:36.013 -> Average Cal sensor Temp.: 96.32 Ambient Temp.: 85.51 F
15:39:37.569 -> Average Cal sensor Temp.: 96.09 Ambient Temp.: 85.55 F
15:39:39.122 -> Average Cal sensor Temp.: 96.03 Ambient Temp.: 85.48 F
15:39:40.680 -> Average Cal sensor Temp.: 96.58 Ambient Temp.: 85.41 F
15:39:42.228 -> Average Cal sensor Temp.: 97.20 Ambient Temp.: 85.44 F
15:39:43.782 -> Average Cal sensor Temp.: 97.48 Ambient Temp.: 85.48 F
15:39:45.335 -> Average Cal sensor Temp.: 97.23 Ambient Temp.: 85.41 F
15:39:46.887 -> Average Cal sensor Temp.: 96.66 Ambient Temp.: 85.44 F
15:39:48.436 -> Average Cal sensor Temp.: 96.70 Ambient Temp.: 85.41 F
15:39:49.990 -> Average Cal sensor Temp.: 96.57 Ambient Temp.: 85.41 F
15:39:51.576 -> Average Cal sensor Temp.: 96.84 Ambient Temp.: 85.41 F
15:39:53.128 -> Average Cal sensor Temp.: 96.79 Ambient Temp.: 85.41 F
15:39:54.675 -> Average Cal sensor Temp.: 96.84 Ambient Temp.: 85.33 F
15:39:56.225 -> Average Cal sensor Temp.: 96.79 Ambient Temp.: 85.37 F
15:39:57.776 -> Average Cal sensor Temp.: 96.87 Ambient Temp.: 85.37 F
15:39:59.352 -> Average Cal sensor Temp.: 96.99 Ambient Temp.: 85.41 F
15:40:00.903 -> Average Cal sensor Temp.: 96.47 Ambient Temp.: 85.33 F
15:40:02.449 -> Average Cal sensor Temp.: 97.11 Ambient Temp.: 85.33 F
15:40:04.007 -> Average Cal sensor Temp.: 97.23 Ambient Temp.: 85.37 F
15:40:05.561 -> Average Cal sensor Temp.: 96.80 Ambient Temp.: 85.33 F
15:40:07.114 -> Average Cal sensor Temp.: 96.67 Ambient Temp.: 85.33 F
15:40:08.693 -> Average Cal sensor Temp.: 96.93 Ambient Temp.: 85.37 F
15:40:10.246 -> Average Cal sensor Temp.: 97.18 Ambient Temp.: 85.37 F
15:40:11.789 -> Average Cal sensor Temp.: 96.88 Ambient Temp.: 85.33 F
15:40:13.332 -> Average Cal sensor Temp.: 96.91 Ambient Temp.: 85.33 F
15:40:14.918 -> Average Cal sensor Temp.: 96.65 Ambient Temp.: 85.30 F
15:40:16.471 -> Average Cal sensor Temp.: 96.94 Ambient Temp.: 85.33 F
15:40:18.018 -> Average Cal sensor Temp.: 97.25 Ambient Temp.: 85.33 F
15:40:19.571 -> Average Cal sensor Temp.: 97.26 Ambient Temp.: 85.37 F
15:40:21.119 -> Average Cal sensor Temp.: 97.20 Ambient Temp.: 85.37 F
15:40:22.669 -> Average Cal sensor Temp.: 96.80 Ambient Temp.: 85.33 F
15:40:24.220 -> Average Cal sensor Temp.: 96.77 Ambient Temp.: 85.33 F
15:40:25.801 -> Average Cal sensor Temp.: 96.58 Ambient Temp.: 85.26 F
15:40:27.354 -> Average Cal sensor Temp.: 96.94 Ambient Temp.: 85.37 F
15:40:28.908 -> Average Cal sensor Temp.: 97.22 Ambient Temp.: 85.30 F
15:40:30.461 -> Average Cal sensor Temp.: 97.24 Ambient Temp.: 85.30 F
15:40:32.015 -> Average Cal sensor Temp.: 97.03 Ambient Temp.: 85.33 F
15:40:33.565 -> Average Cal sensor Temp.: 96.94 Ambient Temp.: 85.30 F
15:40:35.118 -> Average Cal sensor Temp.: 97.35 Ambient Temp.: 85.33 F
15:40:36.671 -> Average Cal sensor Temp.: 97.07 Ambient Temp.: 85.30 F
15:40:38.225 -> Average Cal sensor Temp.: 97.28 Ambient Temp.: 85.33 F
15:40:39.805 -> Average Cal sensor Temp.: 97.45 Ambient Temp.: 85.37 F
15:40:41.356 -> Average Cal sensor Temp.: 97.07 Ambient Temp.: 85.37 F
15:40:42.910 -> Average Cal sensor Temp.: 97.06 Ambient Temp.: 85.37 F
15:40:44.460 -> Average Cal sensor Temp.: 97.02 Ambient Temp.: 85.37 F
15:40:46.013 -> Average Cal sensor Temp.: 97.00 Ambient Temp.: 85.33 F
15:40:47.558 -> Average Cal sensor Temp.: 97.08 Ambient Temp.: 85.33 F
15:40:49.113 -> Average Cal sensor Temp.: 96.94 Ambient Temp.: 85.37 F
15:40:50.696 -> Average Cal sensor Temp.: 97.45 Ambient Temp.: 85.37 F
15:40:52.244 -> Average Cal sensor Temp.: 97.45 Ambient Temp.: 85.44 F
15:40:53.796 -> Average Cal sensor Temp.: 97.32 Ambient Temp.: 85.44 F
15:40:55.342 -> Average Cal sensor Temp.: 97.19 Ambient Temp.: 85.44 F
15:40:56.894 -> Average Cal sensor Temp.: 97.32 Ambient Temp.: 85.48 F
15:40:58.474 -> Average Cal sensor Temp.: 97.06 Ambient Temp.: 85.44 F
15:41:00.022 -> Average Cal sensor Temp.: 96.85 Ambient Temp.: 85.41 F
15:41:01.569 -> Average Cal sensor Temp.: 97.06 Ambient Temp.: 85.44 F
15:41:03.114 -> Average Cal sensor Temp.: 97.25 Ambient Temp.: 85.44 F
15:41:04.671 -> Average Cal sensor Temp.: 97.26 Ambient Temp.: 85.44 F
15:41:06.231 -> Average Cal sensor Temp.: 97.47 Ambient Temp.: 85.51 F
15:41:07.779 -> Average Cal sensor Temp.: 97.54 Ambient Temp.: 85.51 F
15:41:09.362 -> Average Cal sensor Temp.: 97.34 Ambient Temp.: 85.44 F
15:41:10.917 -> Average Cal sensor Temp.: 97.37 Ambient Temp.: 85.51 F
15:41:12.468 -> Average Cal sensor Temp.: 97.48 Ambient Temp.: 85.44 F
15:41:14.025 -> Average Cal sensor Temp.: 97.44 Ambient Temp.: 85.51 F
15:41:15.583 -> Average Cal sensor Temp.: 97.09 Ambient Temp.: 85.55 F
15:41:17.139 -> Average Cal sensor Temp.: 97.18 Ambient Temp.: 85.51 F
15:41:18.696 -> Average Cal sensor Temp.: 97.18 Ambient Temp.: 85.55 F
15:41:20.250 -> Average Cal sensor Temp.: 97.39 Ambient Temp.: 85.59 F
15:41:21.802 -> Average Cal sensor Temp.: 97.44 Ambient Temp.: 85.59 F
15:41:23.358 -> Average Cal sensor Temp.: 97.20 Ambient Temp.: 85.55 F
15:41:24.910 -> Average Cal sensor Temp.: 97.54 Ambient Temp.: 85.66 F
15:41:26.465 -> Average Cal sensor Temp.: 97.34 Ambient Temp.: 85.69 F
15:41:28.019 -> Average Cal sensor Temp.: 97.40 Ambient Temp.: 85.69 F
15:41:29.574 -> Average Cal sensor Temp.: 97.28 Ambient Temp.: 85.69 F
15:41:31.126 -> Average Cal sensor Temp.: 97.50 Ambient Temp.: 85.73 F
15:41:32.682 -> Average Cal sensor Temp.: 97.57 Ambient Temp.: 85.77 F
15:41:34.237 -> Average Cal sensor Temp.: 97.65 Ambient Temp.: 85.77 F
15:41:35.793 -> Average Cal sensor Temp.: 97.68 Ambient Temp.: 85.84 F
15:41:37.342 -> Average Cal sensor Temp.: 97.32 Ambient Temp.: 85.84 F
15:41:38.893 -> Average Cal sensor Temp.: 97.35 Ambient Temp.: 85.80 F
15:41:40.446 -> Average Cal sensor Temp.: 97.46 Ambient Temp.: 85.87 F
15:41:42.031 -> Average Cal sensor Temp.: 97.34 Ambient Temp.: 85.87 F
15:41:43.586 -> Average Cal sensor Temp.: 97.15 Ambient Temp.: 85.80 F
15:41:45.137 -> Average Cal sensor Temp.: 97.44 Ambient Temp.: 85.87 F
15:41:46.687 -> Average Cal sensor Temp.: 97.49 Ambient Temp.: 85.91 F
15:41:48.242 -> Average Cal sensor Temp.: 97.54 Ambient Temp.: 85.84 F
15:41:49.793 -> Average Cal sensor Temp.: 97.18 Ambient Temp.: 85.87 F
15:41:51.346 -> Average Cal sensor Temp.: 97.70 Ambient Temp.: 85.91 F
15:41:52.897 -> Average Cal sensor Temp.: 97.47 Ambient Temp.: 85.87 F
15:41:54.450 -> Average Cal sensor Temp.: 97.58 Ambient Temp.: 85.87 F
15:41:56.008 -> Average Cal sensor Temp.: 97.42 Ambient Temp.: 85.91 F
15:41:57.558 -> Average Cal sensor Temp.: 97.49 Ambient Temp.: 85.95 F
15:41:59.143 -> Average Cal sensor Temp.: 97.53 Ambient Temp.: 85.95 F
15:42:00.696 -> Average Cal sensor Temp.: 97.45 Ambient Temp.: 85.98 F
15:42:02.248 -> Average Cal sensor Temp.: 97.58 Ambient Temp.: 85.95 F
15:42:03.806 -> Average Cal sensor Temp.: 97.62 Ambient Temp.: 85.98 F
15:42:05.355 -> Average Cal sensor Temp.: 97.64 Ambient Temp.: 86.05 F
15:42:06.907 -> Average Cal sensor Temp.: 97.50 Ambient Temp.: 86.09 F
15:42:08.458 -> Average Cal sensor Temp.: 97.47 Ambient Temp.: 86.09 F
15:42:10.041 -> Average Cal sensor Temp.: 97.58 Ambient Temp.: 86.13 F
15:42:11.564 -> Average Cal sensor Temp.: 97.22 Ambient Temp.: 86.09 F
15:42:13.150 -> Average Cal sensor Temp.: 97.83 Ambient Temp.: 86.13 F
15:42:14.699 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 86.20 F
15:42:16.254 -> Average Cal sensor Temp.: 97.37 Ambient Temp.: 86.23 F
15:42:17.806 -> Average Cal sensor Temp.: 97.53 Ambient Temp.: 86.13 F
15:42:19.359 -> Average Cal sensor Temp.: 97.63 Ambient Temp.: 86.20 F
15:42:20.912 -> Average Cal sensor Temp.: 97.85 Ambient Temp.: 86.27 F
15:42:22.467 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 86.31 F
15:42:24.017 -> Average Cal sensor Temp.: 97.57 Ambient Temp.: 86.31 F
15:42:25.575 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 86.27 F
15:42:27.127 -> Average Cal sensor Temp.: 97.72 Ambient Temp.: 86.27 F
15:42:28.678 -> Average Cal sensor Temp.: 97.72 Ambient Temp.: 86.27 F
15:42:30.229 -> Average Cal sensor Temp.: 97.74 Ambient Temp.: 86.31 F
15:42:31.783 -> Average Cal sensor Temp.: 97.67 Ambient Temp.: 86.34 F
15:42:33.368 -> Average Cal sensor Temp.: 97.65 Ambient Temp.: 86.34 F
15:42:34.925 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 86.38 F
15:42:36.480 -> Average Cal sensor Temp.: 97.72 Ambient Temp.: 86.38 F
15:42:38.025 -> Average Cal sensor Temp.: 97.56 Ambient Temp.: 86.38 F
15:42:39.569 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 86.49 F
15:42:41.123 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 86.52 F
15:42:42.712 -> Average Cal sensor Temp.: 97.70 Ambient Temp.: 86.49 F
15:42:44.268 -> Average Cal sensor Temp.: 97.68 Ambient Temp.: 86.49 F
15:42:45.820 -> Average Cal sensor Temp.: 98.04 Ambient Temp.: 86.56 F
15:42:47.367 -> Average Cal sensor Temp.: 97.68 Ambient Temp.: 86.52 F
15:42:48.916 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 86.59 F
15:42:50.467 -> Average Cal sensor Temp.: 97.96 Ambient Temp.: 86.56 F
15:42:52.025 -> Average Cal sensor Temp.: 97.84 Ambient Temp.: 86.59 F
15:42:53.610 -> Average Cal sensor Temp.: 97.52 Ambient Temp.: 86.63 F
15:42:55.161 -> Average Cal sensor Temp.: 97.42 Ambient Temp.: 86.63 F
15:42:56.717 -> Average Cal sensor Temp.: 97.55 Ambient Temp.: 86.67 F
15:42:58.275 -> Average Cal sensor Temp.: 97.76 Ambient Temp.: 86.63 F
15:42:59.828 -> Average Cal sensor Temp.: 97.85 Ambient Temp.: 86.70 F
15:43:01.381 -> Average Cal sensor Temp.: 97.76 Ambient Temp.: 86.70 F
15:43:02.934 -> Average Cal sensor Temp.: 97.98 Ambient Temp.: 86.74 F
15:43:04.489 -> Average Cal sensor Temp.: 97.96 Ambient Temp.: 86.74 F
15:43:06.044 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 86.74 F
15:43:07.592 -> Average Cal sensor Temp.: 97.86 Ambient Temp.: 86.77 F
15:43:09.173 -> Average Cal sensor Temp.: 97.87 Ambient Temp.: 86.88 F
15:43:10.723 -> Average Cal sensor Temp.: 98.05 Ambient Temp.: 86.88 F
15:43:12.275 -> Average Cal sensor Temp.: 98.50 Ambient Temp.: 86.88 F
15:43:13.826 -> Average Cal sensor Temp.: 98.12 Ambient Temp.: 86.92 F
15:43:15.374 -> Average Cal sensor Temp.: 98.00 Ambient Temp.: 86.92 F
15:43:16.964 -> Average Cal sensor Temp.: 97.76 Ambient Temp.: 86.95 F
15:43:18.519 -> Average Cal sensor Temp.: 97.83 Ambient Temp.: 86.95 F
15:43:20.071 -> Average Cal sensor Temp.: 98.00 Ambient Temp.: 86.99 F
15:43:21.624 -> Average Cal sensor Temp.: 97.97 Ambient Temp.: 86.99 F
15:43:23.174 -> Average Cal sensor Temp.: 98.12 Ambient Temp.: 87.06 F
15:43:24.721 -> Average Cal sensor Temp.: 97.78 Ambient Temp.: 87.06 F
15:43:26.306 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 87.10 F
15:43:27.862 -> Average Cal sensor Temp.: 97.78 Ambient Temp.: 87.13 F
15:43:29.412 -> Average Cal sensor Temp.: 97.81 Ambient Temp.: 87.10 F
15:43:30.965 -> Average Cal sensor Temp.: 97.76 Ambient Temp.: 87.13 F
15:43:32.519 -> Average Cal sensor Temp.: 97.52 Ambient Temp.: 87.03 F
15:43:34.073 -> Average Cal sensor Temp.: 97.67 Ambient Temp.: 87.10 F
15:43:35.626 -> Average Cal sensor Temp.: 97.71 Ambient Temp.: 87.10 F
15:43:37.182 -> Average Cal sensor Temp.: 97.44 Ambient Temp.: 87.06 F
15:43:38.733 -> Average Cal sensor Temp.: 97.68 Ambient Temp.: 87.03 F
15:43:40.287 -> Average Cal sensor Temp.: 97.49 Ambient Temp.: 87.06 F
15:43:41.839 -> Average Cal sensor Temp.: 97.98 Ambient Temp.: 86.99 F
15:43:43.388 -> Average Cal sensor Temp.: 98.05 Ambient Temp.: 87.06 F
15:43:44.974 -> Average Cal sensor Temp.: 97.62 Ambient Temp.: 87.06 F
15:43:46.527 -> Average Cal sensor Temp.: 97.40 Ambient Temp.: 87.03 F
15:43:48.077 -> Average Cal sensor Temp.: 97.43 Ambient Temp.: 86.99 F
15:43:49.628 -> Average Cal sensor Temp.: 97.89 Ambient Temp.: 87.03 F
15:43:51.175 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 87.03 F
15:43:52.727 -> Average Cal sensor Temp.: 97.67 Ambient Temp.: 86.99 F
15:43:54.308 -> Average Cal sensor Temp.: 98.08 Ambient Temp.: 86.99 F
15:43:55.863 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 87.03 F
15:43:57.409 -> Average Cal sensor Temp.: 97.53 Ambient Temp.: 86.99 F
15:43:58.955 -> Average Cal sensor Temp.: 97.19 Ambient Temp.: 87.03 F
15:44:00.539 -> Average Cal sensor Temp.: 97.31 Ambient Temp.: 86.99 F
15:44:02.084 -> Average Cal sensor Temp.: 97.35 Ambient Temp.: 86.99 F
15:44:03.639 -> Average Cal sensor Temp.: 97.36 Ambient Temp.: 86.95 F
15:44:05.190 -> Average Cal sensor Temp.: 97.58 Ambient Temp.: 86.95 F
15:44:06.743 -> Average Cal sensor Temp.: 97.49 Ambient Temp.: 86.95 F
15:44:08.292 -> Average Cal sensor Temp.: 97.60 Ambient Temp.: 86.92 F
15:44:09.881 -> Average Cal sensor Temp.: 97.83 Ambient Temp.: 86.88 F
15:44:11.429 -> Average Cal sensor Temp.: 97.46 Ambient Temp.: 86.92 F
15:44:12.982 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 86.95 F
15:44:14.528 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 86.95 F
15:44:16.078 -> Average Cal sensor Temp.: 97.60 Ambient Temp.: 86.92 F
15:44:17.665 -> Average Cal sensor Temp.: 97.34 Ambient Temp.: 86.92 F
15:44:19.217 -> Average Cal sensor Temp.: 97.59 Ambient Temp.: 86.88 F
15:44:20.769 -> Average Cal sensor Temp.: 97.67 Ambient Temp.: 86.92 F
15:44:22.322 -> Average Cal sensor Temp.: 97.04 Ambient Temp.: 86.88 F
15:44:23.876 -> Average Cal sensor Temp.: 97.51 Ambient Temp.: 86.77 F
15:44:25.428 -> Average Cal sensor Temp.: 97.75 Ambient Temp.: 86.88 F
15:44:26.978 -> Average Cal sensor Temp.: 97.77 Ambient Temp.: 86.88 F
15:44:28.562 -> Average Cal sensor Temp.: 97.38 Ambient Temp.: 86.81 F
15:44:30.118 -> Average Cal sensor Temp.: 97.32 Ambient Temp.: 86.77 F
15:44:31.639 -> Average Cal sensor Temp.: 97.58 Ambient Temp.: 86.77 F
15:44:33.228 -> Average Cal sensor Temp.: 97.77 Ambient Temp.: 86.81 F
15:44:34.777 -> Average Cal sensor Temp.: 98.03 Ambient Temp.: 86.74 F
15:44:36.333 -> Average Cal sensor Temp.: 98.09 Ambient Temp.: 86.81 F
15:44:37.890 -> Average Cal sensor Temp.: 97.25 Ambient Temp.: 86.77 F
15:44:39.443 -> Average Cal sensor Temp.: 97.52 Ambient Temp.: 86.74 F
15:44:40.995 -> Average Cal sensor Temp.: 97.65 Ambient Temp.: 86.74 F
15:44:42.551 -> Average Cal sensor Temp.: 97.60 Ambient Temp.: 86.74 F
15:44:44.108 -> Average Cal sensor Temp.: 97.39 Ambient Temp.: 86.74 F
15:44:45.664 -> Average Cal sensor Temp.: 97.28 Ambient Temp.: 86.74 F
15:44:47.225 -> Average Cal sensor Temp.: 97.14 Ambient Temp.: 86.70 F
15:44:48.778 -> Average Cal sensor Temp.: 96.79 Ambient Temp.: 86.70 F
15:44:50.330 -> Average Cal sensor Temp.: 97.56 Ambient Temp.: 86.70 F
15:44:51.887 -> Average Cal sensor Temp.: 97.15 Ambient Temp.: 86.70 F
15:44:53.441 -> Average Cal sensor Temp.: 97.17 Ambient Temp.: 86.70 F
15:44:54.994 -> Average Cal sensor Temp.: 97.49 Ambient Temp.: 86.59 F
15:44:56.549 -> Average Cal sensor Temp.: 97.58 Ambient Temp.: 86.63 F
15:44:58.104 -> Average Cal sensor Temp.: 97.20 Ambient Temp.: 86.56 F
15:44:59.661 -> Average Cal sensor Temp.: 97.11 Ambient Temp.: 86.59 F
15:45:01.219 -> Average Cal sensor Temp.: 96.83 Ambient Temp.: 86.56 F
15:45:02.776 -> Average Cal sensor Temp.: 97.26 Ambient Temp.: 86.52 F
15:45:04.329 -> Average Cal sensor Temp.: 97.50 Ambient Temp.: 86.52 F
15:45:05.891 -> Average Cal sensor Temp.: 97.37 Ambient Temp.: 86.49 F
15:45:07.446 -> Average Cal sensor Temp.: 97.50 Ambient Temp.: 86.38 F
15:45:08.997 -> Average Cal sensor Temp.: 97.26 Ambient Temp.: 86.38 F
15:45:10.550 -> Average Cal sensor Temp.: 96.89 Ambient Temp.: 86.31 F
15:45:12.132 -> Average Cal sensor Temp.: 97.36 Ambient Temp.: 86.31 F
15:45:13.683 -> Average Cal sensor Temp.: 96.97 Ambient Temp.: 86.31 F
15:45:15.237 -> Average Cal sensor Temp.: 97.03 Ambient Temp.: 86.27 F
15:45:16.791 -> Average Cal sensor Temp.: 97.51 Ambient Temp.: 86.27 F
15:45:18.342 -> Average Cal sensor Temp.: 97.44 Ambient Temp.: 86.20 F
15:45:19.899 -> Average Cal sensor Temp.: 97.76 Ambient Temp.: 86.23 F
15:45:21.452 -> Average Cal sensor Temp.: 97.78 Ambient Temp.: 86.31 F
15:45:22.995 -> Average Cal sensor Temp.: 97.65 Ambient Temp.: 86.31 F
15:45:24.577 -> Average Cal sensor Temp.: 97.22 Ambient Temp.: 86.23 F
15:45:26.126 -> Average Cal sensor Temp.: 97.40 Ambient Temp.: 86.23 F
15:45:27.674 -> Average Cal sensor Temp.: 97.33 Ambient Temp.: 86.23 F
15:45:29.229 -> Average Cal sensor Temp.: 97.52 Ambient Temp.: 86.23 F
15:45:30.787 -> Average Cal sensor Temp.: 97.48 Ambient Temp.: 86.16 F
15:45:32.334 -> Average Cal sensor Temp.: 97.14 Ambient Temp.: 86.13 F
15:45:33.892 -> Average Cal sensor Temp.: 97.13 Ambient Temp.: 86.13 F
15:45:35.439 -> Average Cal sensor Temp.: 97.56 Ambient Temp.: 86.16 F
15:45:37.003 -> Average Cal sensor Temp.: 97.51 Ambient Temp.: 86.09 F
15:45:38.560 -> Average Cal sensor Temp.: 97.25 Ambient Temp.: 86.05 F
15:45:40.113 -> Average Cal sensor Temp.: 97.58 Ambient Temp.: 86.05 F
15:45:41.663 -> Average Cal sensor Temp.: 97.55 Ambient Temp.: 86.09 F
15:45:43.213 -> Average Cal sensor Temp.: 97.87 Ambient Temp.: 86.13 F
15:45:44.768 -> Average Cal sensor Temp.: 97.34 Ambient Temp.: 85.98 F
15:45:46.326 -> Average Cal sensor Temp.: 97.71 Ambient Temp.: 85.98 F
15:45:47.883 -> Average Cal sensor Temp.: 97.52 Ambient Temp.: 85.98 F
15:45:49.435 -> Average Cal sensor Temp.: 97.69 Ambient Temp.: 85.95 F
15:45:51.018 -> Average Cal sensor Temp.: 97.94 Ambient Temp.: 85.98 F
15:45:52.577 -> Average Cal sensor Temp.: 97.71 Ambient Temp.: 85.98 F
15:45:54.132 -> Average Cal sensor Temp.: 97.77 Ambient Temp.: 85.95 F
15:45:55.685 -> Average Cal sensor Temp.: 97.75 Ambient Temp.: 85.98 F
15:45:57.238 -> Average Cal sensor Temp.: 97.66 Ambient Temp.: 85.98 F
15:45:58.797 -> Average Cal sensor Temp.: 97.97 Ambient Temp.: 85.98 F
15:46:00.344 -> Average Cal sensor Temp.: 97.74 Ambient Temp.: 85.95 F
15:46:01.897 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 85.98 F
15:46:03.452 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 85.91 F
15:46:05.004 -> Average Cal sensor Temp.: 97.73 Ambient Temp.: 85.91 F
15:46:06.564 -> Average Cal sensor Temp.: 97.52 Ambient Temp.: 85.87 F
15:46:08.122 -> Average Cal sensor Temp.: 97.66 Ambient Temp.: 85.95 F
15:46:09.672 -> Average Cal sensor Temp.: 97.37 Ambient Temp.: 85.87 F
15:46:10.261 -> Ambient Temp.: 85.87 F
15:46:11.816 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 85.87 F
15:46:13.366 -> Average Cal sensor Temp.: 98.29 Ambient Temp.: 85.91 F
15:46:14.914 -> Average Cal sensor Temp.: 98.42 Ambient Temp.: 85.84 F
15:46:16.469 -> Average Cal sensor Temp.: 98.10 Ambient Temp.: 85.87 F
15:46:18.057 -> Average Cal sensor Temp.: 98.48 Ambient Temp.: 85.84 F
15:46:19.233 -> Ambient Temp.: 85.87 F
15:46:20.788 -> Average Cal sensor Temp.: 98.40 Ambient Temp.: 85.91 F
15:46:22.340 -> Average Cal sensor Temp.: 98.50 Ambient Temp.: 85.95 F
15:46:23.889 -> Average Cal sensor Temp.: 98.94 Ambient Temp.: 85.91 F
15:46:25.474 -> Average Cal sensor Temp.: 98.67 Ambient Temp.: 85.91 F
15:46:27.031 -> Average Cal sensor Temp.: 98.53 Ambient Temp.: 86.05 F
15:46:28.588 -> Average Cal sensor Temp.: 98.97 Ambient Temp.: 85.98 F
15:46:30.143 -> Average Cal sensor Temp.: 98.50 Ambient Temp.: 85.98 F
15:46:31.701 -> Average Cal sensor Temp.: 98.42 Ambient Temp.: 86.05 F
15:46:33.256 -> Average Cal sensor Temp.: 97.99 Ambient Temp.: 85.95 F
15:46:34.807 -> Average Cal sensor Temp.: 97.83 Ambient Temp.: 86.05 F
15:46:36.360 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 86.09 F
15:46:37.912 -> Average Cal sensor Temp.: 97.69 Ambient Temp.: 86.05 F
15:46:39.464 -> Average Cal sensor Temp.: 97.46 Ambient Temp.: 85.98 F
15:46:41.020 -> Average Cal sensor Temp.: 97.30 Ambient Temp.: 85.91 F
15:46:42.572 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 85.98 F
15:46:44.128 -> Average Cal sensor Temp.: 97.48 Ambient Temp.: 85.95 F
15:46:45.715 -> Average Cal sensor Temp.: 97.48 Ambient Temp.: 85.95 F
15:46:47.267 -> Average Cal sensor Temp.: 97.71 Ambient Temp.: 85.84 F
15:46:48.819 -> Average Cal sensor Temp.: 97.82 Ambient Temp.: 85.91 F
15:46:50.374 -> Average Cal sensor Temp.: 97.91 Ambient Temp.: 85.87 F
15:46:51.932 -> Average Cal sensor Temp.: 97.57 Ambient Temp.: 85.87 F
15:46:53.483 -> Average Cal sensor Temp.: 97.49 Ambient Temp.: 85.87 F
15:46:55.038 -> Average Cal sensor Temp.: 97.52 Ambient Temp.: 85.80 F
15:46:56.593 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 85.80 F
15:46:58.148 -> Average Cal sensor Temp.: 98.30 Ambient Temp.: 85.84 F
15:46:59.703 -> Average Cal sensor Temp.: 98.13 Ambient Temp.: 85.84 F
15:47:01.255 -> Average Cal sensor Temp.: 97.97 Ambient Temp.: 85.80 F
15:47:02.810 -> Average Cal sensor Temp.: 98.17 Ambient Temp.: 85.87 F
15:47:04.361 -> Average Cal sensor Temp.: 97.93 Ambient Temp.: 85.84 F
15:47:05.914 -> Average Cal sensor Temp.: 98.19 Ambient Temp.: 85.84 F
15:47:07.467 -> Average Cal sensor Temp.: 97.94 Ambient Temp.: 85.80 F
15:47:09.056 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 85.80 F
15:47:10.611 -> Average Cal sensor Temp.: 98.08 Ambient Temp.: 85.87 F
15:47:12.168 -> Average Cal sensor Temp.: 97.98 Ambient Temp.: 85.84 F
15:47:13.719 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 85.91 F
15:47:15.265 -> Average Cal sensor Temp.: 97.56 Ambient Temp.: 85.84 F
15:47:16.815 -> Average Cal sensor Temp.: 97.41 Ambient Temp.: 85.87 F
15:47:18.369 -> Average Cal sensor Temp.: 97.70 Ambient Temp.: 85.84 F
15:47:19.922 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 85.80 F
15:47:21.511 -> Average Cal sensor Temp.: 98.34 Ambient Temp.: 85.87 F
15:47:23.070 -> Average Cal sensor Temp.: 97.69 Ambient Temp.: 85.87 F
15:47:24.627 -> Average Cal sensor Temp.: 97.87 Ambient Temp.: 85.87 F
15:47:26.153 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 85.91 F
15:47:27.737 -> Average Cal sensor Temp.: 97.96 Ambient Temp.: 85.87 F
15:47:29.292 -> Average Cal sensor Temp.: 98.04 Ambient Temp.: 85.91 F
15:47:30.847 -> Average Cal sensor Temp.: 98.31 Ambient Temp.: 85.87 F
15:47:32.398 -> Average Cal sensor Temp.: 97.77 Ambient Temp.: 85.84 F
15:47:33.954 -> Average Cal sensor Temp.: 97.73 Ambient Temp.: 85.84 F
15:47:35.508 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 85.84 F
15:47:37.065 -> Average Cal sensor Temp.: 97.93 Ambient Temp.: 85.95 F
15:47:38.625 -> Average Cal sensor Temp.: 97.83 Ambient Temp.: 85.87 F
15:47:40.173 -> Average Cal sensor Temp.: 98.04 Ambient Temp.: 85.98 F
15:47:41.724 -> Average Cal sensor Temp.: 97.98 Ambient Temp.: 85.98 F
15:47:43.276 -> Average Cal sensor Temp.: 97.57 Ambient Temp.: 85.95 F
15:47:44.827 -> Average Cal sensor Temp.: 97.52 Ambient Temp.: 85.95 F
15:47:46.411 -> Average Cal sensor Temp.: 97.80 Ambient Temp.: 85.98 F
15:47:47.962 -> Average Cal sensor Temp.: 97.83 Ambient Temp.: 85.98 F
15:47:49.523 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 85.95 F
15:47:51.073 -> Average Cal sensor Temp.: 97.80 Ambient Temp.: 86.05 F
15:47:52.621 -> Average Cal sensor Temp.: 97.75 Ambient Temp.: 86.05 F
15:47:54.175 -> Average Cal sensor Temp.: 97.52 Ambient Temp.: 85.98 F
15:47:55.734 -> Average Cal sensor Temp.: 97.77 Ambient Temp.: 85.98 F
15:47:57.282 -> Average Cal sensor Temp.: 97.60 Ambient Temp.: 85.98 F
15:47:58.837 -> Average Cal sensor Temp.: 97.81 Ambient Temp.: 85.95 F
15:48:00.392 -> Average Cal sensor Temp.: 97.83 Ambient Temp.: 85.98 F
15:48:01.943 -> Average Cal sensor Temp.: 97.74 Ambient Temp.: 85.95 F
15:48:03.526 -> Average Cal sensor Temp.: 97.87 Ambient Temp.: 85.95 F
15:48:05.083 -> Average Cal sensor Temp.: 98.02 Ambient Temp.: 85.95 F
15:48:06.638 -> Average Cal sensor Temp.: 97.86 Ambient Temp.: 85.91 F
15:48:08.193 -> Average Cal sensor Temp.: 97.94 Ambient Temp.: 85.91 F
15:48:09.744 -> Average Cal sensor Temp.: 98.04 Ambient Temp.: 85.98 F
15:48:11.303 -> Average Cal sensor Temp.: 97.99 Ambient Temp.: 85.95 F
15:48:12.862 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 86.05 F
15:48:14.416 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 86.09 F
15:48:15.968 -> Average Cal sensor Temp.: 97.83 Ambient Temp.: 86.05 F
15:48:17.522 -> Average Cal sensor Temp.: 97.93 Ambient Temp.: 85.95 F
15:48:19.077 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 86.09 F
15:48:20.632 -> Average Cal sensor Temp.: 98.00 Ambient Temp.: 86.05 F
15:48:22.186 -> Average Cal sensor Temp.: 97.89 Ambient Temp.: 86.13 F
15:48:23.740 -> Average Cal sensor Temp.: 98.05 Ambient Temp.: 86.09 F
15:48:25.322 -> Average Cal sensor Temp.: 98.19 Ambient Temp.: 86.13 F
15:48:26.875 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 86.13 F
15:48:28.430 -> Average Cal sensor Temp.: 97.96 Ambient Temp.: 86.13 F
15:48:29.983 -> Average Cal sensor Temp.: 98.10 Ambient Temp.: 86.13 F
15:48:31.537 -> Average Cal sensor Temp.: 97.81 Ambient Temp.: 86.20 F
15:48:33.093 -> Average Cal sensor Temp.: 97.80 Ambient Temp.: 86.20 F
15:48:34.645 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 86.23 F
15:48:36.197 -> Average Cal sensor Temp.: 97.94 Ambient Temp.: 86.20 F
15:48:37.749 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 86.23 F
15:48:39.305 -> Average Cal sensor Temp.: 97.82 Ambient Temp.: 86.23 F
15:48:40.861 -> Average Cal sensor Temp.: 97.93 Ambient Temp.: 86.31 F
15:48:42.440 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 86.34 F
15:48:43.989 -> Average Cal sensor Temp.: 97.95 Ambient Temp.: 86.27 F
15:48:45.538 -> Average Cal sensor Temp.: 97.73 Ambient Temp.: 86.27 F
15:48:47.090 -> Average Cal sensor Temp.: 97.83 Ambient Temp.: 86.34 F
15:48:48.670 -> Average Cal sensor Temp.: 97.60 Ambient Temp.: 86.27 F
15:48:50.218 -> Average Cal sensor Temp.: 97.72 Ambient Temp.: 86.31 F
15:48:51.770 -> Average Cal sensor Temp.: 97.75 Ambient Temp.: 86.38 F
15:48:53.330 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 86.38 F
15:48:54.885 -> Average Cal sensor Temp.: 97.97 Ambient Temp.: 86.45 F
15:48:56.440 -> Average Cal sensor Temp.: 98.16 Ambient Temp.: 86.49 F
15:48:57.997 -> Average Cal sensor Temp.: 98.09 Ambient Temp.: 86.45 F
15:48:59.554 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 86.49 F
15:49:01.107 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 86.56 F
15:49:02.657 -> Average Cal sensor Temp.: 98.13 Ambient Temp.: 86.56 F
15:49:04.215 -> Average Cal sensor Temp.: 97.98 Ambient Temp.: 86.52 F
15:49:05.772 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 86.56 F
15:49:07.326 -> Average Cal sensor Temp.: 98.34 Ambient Temp.: 86.59 F
15:49:08.892 -> Average Cal sensor Temp.: 98.12 Ambient Temp.: 86.63 F
15:49:10.452 -> Average Cal sensor Temp.: 97.83 Ambient Temp.: 86.59 F
15:49:12.010 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 86.67 F
15:49:13.565 -> Average Cal sensor Temp.: 98.04 Ambient Temp.: 86.63 F
15:49:15.120 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 86.67 F
15:49:16.675 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 86.67 F
15:49:18.226 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 86.77 F
15:49:19.780 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 86.67 F
15:49:21.364 -> Average Cal sensor Temp.: 98.00 Ambient Temp.: 86.74 F
15:49:22.917 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 86.77 F
15:49:24.475 -> Average Cal sensor Temp.: 98.05 Ambient Temp.: 86.77 F
15:49:26.029 -> Average Cal sensor Temp.: 98.13 Ambient Temp.: 86.81 F
15:49:27.587 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 86.81 F
15:49:29.142 -> Average Cal sensor Temp.: 98.17 Ambient Temp.: 86.88 F
15:49:30.696 -> Average Cal sensor Temp.: 98.17 Ambient Temp.: 86.88 F
15:49:32.248 -> Average Cal sensor Temp.: 97.97 Ambient Temp.: 86.92 F
15:49:33.801 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 86.95 F
15:49:35.357 -> Average Cal sensor Temp.: 97.80 Ambient Temp.: 86.95 F
15:49:36.907 -> Average Cal sensor Temp.: 98.09 Ambient Temp.: 86.95 F
15:49:38.459 -> Average Cal sensor Temp.: 98.02 Ambient Temp.: 86.99 F
15:49:40.046 -> Average Cal sensor Temp.: 97.97 Ambient Temp.: 87.03 F
15:49:41.570 -> Average Cal sensor Temp.: 98.08 Ambient Temp.: 87.03 F
15:49:43.158 -> Average Cal sensor Temp.: 98.42 Ambient Temp.: 87.03 F
15:49:44.713 -> Average Cal sensor Temp.: 98.34 Ambient Temp.: 87.06 F
15:49:46.267 -> Average Cal sensor Temp.: 98.13 Ambient Temp.: 87.13 F
15:49:47.820 -> Average Cal sensor Temp.: 97.82 Ambient Temp.: 87.17 F
15:49:49.372 -> Average Cal sensor Temp.: 98.13 Ambient Temp.: 87.17 F
15:49:50.925 -> Average Cal sensor Temp.: 98.23 Ambient Temp.: 87.13 F
15:49:52.480 -> Average Cal sensor Temp.: 98.16 Ambient Temp.: 87.17 F
15:49:54.030 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 87.13 F
15:49:55.589 -> Average Cal sensor Temp.: 98.08 Ambient Temp.: 87.10 F
15:49:57.141 -> Average Cal sensor Temp.: 97.75 Ambient Temp.: 87.21 F
15:49:58.699 -> Average Cal sensor Temp.: 97.71 Ambient Temp.: 87.17 F
15:50:00.254 -> Average Cal sensor Temp.: 97.68 Ambient Temp.: 87.17 F
15:50:01.805 -> Average Cal sensor Temp.: 97.71 Ambient Temp.: 87.21 F
15:50:03.360 -> Average Cal sensor Temp.: 97.42 Ambient Temp.: 87.17 F
15:50:04.911 -> Average Cal sensor Temp.: 97.53 Ambient Temp.: 87.17 F
15:50:06.467 -> Average Cal sensor Temp.: 97.28 Ambient Temp.: 87.13 F
15:50:08.025 -> Average Cal sensor Temp.: 97.51 Ambient Temp.: 87.17 F
15:50:09.580 -> Average Cal sensor Temp.: 97.31 Ambient Temp.: 87.17 F
15:50:11.137 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 87.17 F
15:50:12.692 -> Average Cal sensor Temp.: 97.35 Ambient Temp.: 87.13 F
15:50:14.247 -> Average Cal sensor Temp.: 97.71 Ambient Temp.: 87.10 F
15:50:15.803 -> Average Cal sensor Temp.: 97.62 Ambient Temp.: 87.13 F
15:50:17.354 -> Average Cal sensor Temp.: 97.46 Ambient Temp.: 87.10 F
15:50:18.911 -> Average Cal sensor Temp.: 97.65 Ambient Temp.: 87.17 F
15:50:20.473 -> Average Cal sensor Temp.: 97.69 Ambient Temp.: 87.13 F
15:50:22.025 -> Average Cal sensor Temp.: 98.02 Ambient Temp.: 87.13 F
15:50:23.579 -> Average Cal sensor Temp.: 97.85 Ambient Temp.: 87.10 F
15:50:25.128 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 87.17 F
15:50:26.683 -> Average Cal sensor Temp.: 97.64 Ambient Temp.: 87.13 F
15:50:28.239 -> Average Cal sensor Temp.: 97.98 Ambient Temp.: 87.06 F
15:50:29.788 -> Average Cal sensor Temp.: 97.66 Ambient Temp.: 87.10 F
15:50:31.346 -> Average Cal sensor Temp.: 97.75 Ambient Temp.: 87.10 F
15:50:32.903 -> Average Cal sensor Temp.: 97.59 Ambient Temp.: 87.06 F
15:50:34.452 -> Average Cal sensor Temp.: 97.37 Ambient Temp.: 87.10 F
15:50:36.007 -> Average Cal sensor Temp.: 97.59 Ambient Temp.: 87.06 F
15:50:37.558 -> Average Cal sensor Temp.: 97.71 Ambient Temp.: 87.06 F
15:50:39.112 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 87.10 F
15:50:40.668 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 87.10 F
15:50:42.227 -> Average Cal sensor Temp.: 97.98 Ambient Temp.: 87.06 F
15:50:43.788 -> Average Cal sensor Temp.: 97.58 Ambient Temp.: 87.10 F
15:50:45.345 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 87.10 F
15:50:46.904 -> Average Cal sensor Temp.: 97.33 Ambient Temp.: 87.03 F
15:50:48.463 -> Average Cal sensor Temp.: 97.43 Ambient Temp.: 86.99 F
15:50:50.021 -> Average Cal sensor Temp.: 97.05 Ambient Temp.: 87.03 F
15:50:51.573 -> Average Cal sensor Temp.: 97.17 Ambient Temp.: 87.03 F
15:50:53.131 -> Average Cal sensor Temp.: 97.11 Ambient Temp.: 86.99 F
15:50:54.679 -> Average Cal sensor Temp.: 96.77 Ambient Temp.: 87.03 F
15:50:56.231 -> Average Cal sensor Temp.: 97.64 Ambient Temp.: 86.99 F
15:50:57.785 -> Average Cal sensor Temp.: 98.07 Ambient Temp.: 86.99 F
15:50:59.338 -> Average Cal sensor Temp.: 97.98 Ambient Temp.: 86.95 F
15:51:00.895 -> Average Cal sensor Temp.: 97.74 Ambient Temp.: 86.99 F
15:51:02.448 -> Average Cal sensor Temp.: 98.17 Ambient Temp.: 86.99 F
15:51:04.004 -> Average Cal sensor Temp.: 97.69 Ambient Temp.: 86.95 F
15:51:05.588 -> Average Cal sensor Temp.: 97.66 Ambient Temp.: 86.92 F
15:51:07.144 -> Average Cal sensor Temp.: 97.57 Ambient Temp.: 86.92 F
15:51:08.692 -> Average Cal sensor Temp.: 97.36 Ambient Temp.: 86.88 F
15:51:10.245 -> Average Cal sensor Temp.: 97.52 Ambient Temp.: 86.81 F
15:51:11.798 -> Average Cal sensor Temp.: 97.69 Ambient Temp.: 86.81 F
15:51:13.346 -> Average Cal sensor Temp.: 97.74 Ambient Temp.: 86.74 F
15:51:14.900 -> Average Cal sensor Temp.: 97.54 Ambient Temp.: 86.74 F
15:51:16.457 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 86.74 F
15:51:18.011 -> Average Cal sensor Temp.: 98.19 Ambient Temp.: 86.67 F
15:51:19.566 -> Average Cal sensor Temp.: 98.52 Ambient Temp.: 86.77 F
15:51:21.155 -> Average Cal sensor Temp.: 98.05 Ambient Temp.: 86.77 F
15:51:22.711 -> Average Cal sensor Temp.: 98.05 Ambient Temp.: 86.67 F
15:51:24.264 -> Average Cal sensor Temp.: 97.93 Ambient Temp.: 86.63 F
15:51:25.818 -> Average Cal sensor Temp.: 97.82 Ambient Temp.: 86.67 F
15:51:27.378 -> Average Cal sensor Temp.: 97.78 Ambient Temp.: 86.67 F
15:51:28.929 -> Average Cal sensor Temp.: 97.54 Ambient Temp.: 86.59 F
15:51:30.489 -> Average Cal sensor Temp.: 97.82 Ambient Temp.: 86.59 F
15:51:32.044 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 86.67 F
15:51:33.603 -> Average Cal sensor Temp.: 98.04 Ambient Temp.: 86.63 F
15:51:35.158 -> Average Cal sensor Temp.: 97.69 Ambient Temp.: 86.59 F
15:51:36.716 -> Average Cal sensor Temp.: 97.58 Ambient Temp.: 86.56 F
15:51:38.272 -> Average Cal sensor Temp.: 97.33 Ambient Temp.: 86.59 F
15:51:39.824 -> Average Cal sensor Temp.: 97.02 Ambient Temp.: 86.52 F
15:51:41.380 -> Average Cal sensor Temp.: 97.49 Ambient Temp.: 86.52 F
15:51:42.929 -> Average Cal sensor Temp.: 97.83 Ambient Temp.: 86.49 F
15:51:44.484 -> Average Cal sensor Temp.: 97.37 Ambient Temp.: 86.45 F
15:51:46.038 -> Average Cal sensor Temp.: 97.25 Ambient Temp.: 86.49 F
15:51:47.592 -> Average Cal sensor Temp.: 97.31 Ambient Temp.: 86.49 F
15:51:49.152 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 86.45 F
15:51:50.699 -> Average Cal sensor Temp.: 97.58 Ambient Temp.: 86.38 F
15:51:52.254 -> Average Cal sensor Temp.: 98.00 Ambient Temp.: 86.45 F
15:51:53.843 -> Average Cal sensor Temp.: 97.68 Ambient Temp.: 86.34 F
15:51:55.398 -> Average Cal sensor Temp.: 97.11 Ambient Temp.: 86.31 F
15:51:56.921 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 86.31 F
15:51:58.508 -> Average Cal sensor Temp.: 97.78 Ambient Temp.: 86.31 F
15:52:00.060 -> Average Cal sensor Temp.: 97.93 Ambient Temp.: 86.31 F
15:52:01.616 -> Average Cal sensor Temp.: 97.76 Ambient Temp.: 86.31 F
15:52:03.168 -> Average Cal sensor Temp.: 97.71 Ambient Temp.: 86.34 F
15:52:04.718 -> Average Cal sensor Temp.: 97.63 Ambient Temp.: 86.31 F
15:52:06.274 -> Average Cal sensor Temp.: 97.37 Ambient Temp.: 86.27 F
15:52:07.835 -> Average Cal sensor Temp.: 97.46 Ambient Temp.: 86.23 F
15:52:09.395 -> Average Cal sensor Temp.: 97.20 Ambient Temp.: 86.16 F
15:52:10.922 -> Average Cal sensor Temp.: 97.00 Ambient Temp.: 86.13 F
15:52:12.509 -> Average Cal sensor Temp.: 97.59 Ambient Temp.: 86.16 F
15:52:14.062 -> Average Cal sensor Temp.: 97.73 Ambient Temp.: 86.13 F
15:52:15.615 -> Average Cal sensor Temp.: 97.64 Ambient Temp.: 86.13 F
15:52:17.175 -> Average Cal sensor Temp.: 97.47 Ambient Temp.: 86.13 F
15:52:18.729 -> Average Cal sensor Temp.: 97.77 Ambient Temp.: 86.13 F
15:52:20.281 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 86.13 F
15:52:21.833 -> Average Cal sensor Temp.: 97.49 Ambient Temp.: 86.09 F
15:52:23.389 -> Average Cal sensor Temp.: 96.79 Ambient Temp.: 85.98 F
15:52:24.942 -> Average Cal sensor Temp.: 97.34 Ambient Temp.: 85.95 F
15:52:26.497 -> Average Cal sensor Temp.: 97.68 Ambient Temp.: 85.98 F
15:52:28.054 -> Average Cal sensor Temp.: 97.82 Ambient Temp.: 85.95 F
15:52:29.608 -> Average Cal sensor Temp.: 97.73 Ambient Temp.: 85.95 F
15:52:31.164 -> Average Cal sensor Temp.: 97.26 Ambient Temp.: 85.91 F
15:52:32.715 -> Average Cal sensor Temp.: 97.48 Ambient Temp.: 85.91 F
15:52:34.270 -> Average Cal sensor Temp.: 97.69 Ambient Temp.: 85.91 F
15:52:35.827 -> Average Cal sensor Temp.: 97.80 Ambient Temp.: 85.87 F
15:52:37.385 -> Average Cal sensor Temp.: 97.97 Ambient Temp.: 85.91 F
15:52:38.942 -> Average Cal sensor Temp.: 97.98 Ambient Temp.: 85.87 F
15:52:40.493 -> Average Cal sensor Temp.: 97.77 Ambient Temp.: 85.95 F
15:52:42.045 -> Average Cal sensor Temp.: 97.10 Ambient Temp.: 85.87 F
15:52:43.598 -> Average Cal sensor Temp.: 97.36 Ambient Temp.: 85.84 F
15:52:45.154 -> Average Cal sensor Temp.: 97.39 Ambient Temp.: 85.80 F
15:52:46.708 -> Average Cal sensor Temp.: 97.67 Ambient Temp.: 85.80 F
15:52:48.291 -> Average Cal sensor Temp.: 97.28 Ambient Temp.: 85.77 F
15:52:49.818 -> Average Cal sensor Temp.: 97.27 Ambient Temp.: 85.73 F
15:52:51.373 -> Average Cal sensor Temp.: 97.35 Ambient Temp.: 85.66 F
15:52:52.929 -> Average Cal sensor Temp.: 97.53 Ambient Temp.: 85.69 F
15:52:54.484 -> Average Cal sensor Temp.: 97.77 Ambient Temp.: 85.59 F
15:52:56.039 -> Average Cal sensor Temp.: 97.76 Ambient Temp.: 85.59 F
15:52:57.595 -> Average Cal sensor Temp.: 97.72 Ambient Temp.: 85.66 F
15:52:59.151 -> Average Cal sensor Temp.: 97.69 Ambient Temp.: 85.59 F
15:53:00.707 -> Average Cal sensor Temp.: 97.82 Ambient Temp.: 85.59 F
15:53:02.295 -> Average Cal sensor Temp.: 97.56 Ambient Temp.: 85.55 F
15:53:03.844 -> Average Cal sensor Temp.: 97.32 Ambient Temp.: 85.55 F
15:53:05.388 -> Average Cal sensor Temp.: 97.08 Ambient Temp.: 85.55 F
15:53:06.945 -> Average Cal sensor Temp.: 97.06 Ambient Temp.: 85.48 F
15:53:08.503 -> Average Cal sensor Temp.: 96.80 Ambient Temp.: 85.48 F
15:53:10.051 -> Average Cal sensor Temp.: 97.23 Ambient Temp.: 85.51 F
15:53:11.603 -> Average Cal sensor Temp.: 97.20 Ambient Temp.: 85.44 F
15:53:13.161 -> Average Cal sensor Temp.: 97.43 Ambient Temp.: 85.44 F
15:53:14.711 -> Average Cal sensor Temp.: 97.56 Ambient Temp.: 85.41 F
15:53:16.267 -> Average Cal sensor Temp.: 97.93 Ambient Temp.: 85.41 F
15:53:17.819 -> Average Cal sensor Temp.: 97.99 Ambient Temp.: 85.41 F
15:53:19.376 -> Average Cal sensor Temp.: 97.34 Ambient Temp.: 85.41 F
15:53:20.963 -> Average Cal sensor Temp.: 97.64 Ambient Temp.: 85.37 F
15:53:22.489 -> Average Cal sensor Temp.: 97.47 Ambient Temp.: 85.37 F
15:53:24.041 -> Average Cal sensor Temp.: 97.74 Ambient Temp.: 85.33 F
15:53:25.599 -> Average Cal sensor Temp.: 97.93 Ambient Temp.: 85.33 F
15:53:27.155 -> Average Cal sensor Temp.: 97.91 Ambient Temp.: 85.33 F
15:53:28.712 -> Average Cal sensor Temp.: 97.76 Ambient Temp.: 85.33 F
15:53:30.267 -> Average Cal sensor Temp.: 97.33 Ambient Temp.: 85.33 F
15:53:31.818 -> Average Cal sensor Temp.: 97.12 Ambient Temp.: 85.26 F
15:53:33.375 -> Average Cal sensor Temp.: 97.64 Ambient Temp.: 85.26 F
15:53:34.933 -> Average Cal sensor Temp.: 97.89 Ambient Temp.: 85.26 F
15:53:36.486 -> Average Cal sensor Temp.: 97.73 Ambient Temp.: 85.19 F
15:53:37.455 -> Ambient Temp.: 85.26 F
15:53:38.043 -> Ambient Temp.: 85.26 F
15:53:38.598 -> Ambient Temp.: 85.26 F
15:53:39.808 -> Ambient Temp.: 85.26 F
15:53:40.364 -> Ambient Temp.: 85.26 F
15:53:41.920 -> Average Cal sensor Temp.: 97.89 Ambient Temp.: 85.19 F
15:53:43.475 -> Average Cal sensor Temp.: 97.97 Ambient Temp.: 85.26 F
15:53:45.032 -> Average Cal sensor Temp.: 98.10 Ambient Temp.: 85.33 F
15:53:46.587 -> Average Cal sensor Temp.: 98.13 Ambient Temp.: 85.33 F
15:53:48.137 -> Average Cal sensor Temp.: 98.37 Ambient Temp.: 85.33 F
15:53:49.690 -> Average Cal sensor Temp.: 97.83 Ambient Temp.: 85.30 F
15:53:51.248 -> Average Cal sensor Temp.: 97.84 Ambient Temp.: 85.33 F
15:53:52.804 -> Average Cal sensor Temp.: 97.93 Ambient Temp.: 85.30 F
15:53:54.359 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 85.33 F
15:53:55.916 -> Average Cal sensor Temp.: 98.03 Ambient Temp.: 85.26 F
15:53:57.478 -> Average Cal sensor Temp.: 97.73 Ambient Temp.: 85.30 F
15:53:59.035 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 85.26 F
15:54:00.594 -> Average Cal sensor Temp.: 97.89 Ambient Temp.: 85.26 F
15:54:02.149 -> Average Cal sensor Temp.: 97.86 Ambient Temp.: 85.26 F
15:54:03.704 -> Average Cal sensor Temp.: 97.76 Ambient Temp.: 85.30 F
15:54:05.252 -> Average Cal sensor Temp.: 97.70 Ambient Temp.: 85.26 F
15:54:06.811 -> Average Cal sensor Temp.: 97.81 Ambient Temp.: 85.19 F
15:54:08.363 -> Average Cal sensor Temp.: 97.78 Ambient Temp.: 85.19 F
15:54:09.920 -> Average Cal sensor Temp.: 97.80 Ambient Temp.: 85.19 F
15:54:11.477 -> Average Cal sensor Temp.: 97.98 Ambient Temp.: 85.26 F
15:54:13.028 -> Average Cal sensor Temp.: 97.77 Ambient Temp.: 85.19 F
15:54:14.612 -> Average Cal sensor Temp.: 98.08 Ambient Temp.: 85.26 F
15:54:16.164 -> Average Cal sensor Temp.: 97.77 Ambient Temp.: 85.30 F
15:54:17.234 -> Ambient Temp.: 85.30 F
15:54:17.821 -> Ambient Temp.: 85.30 F
15:54:19.371 -> Average Cal sensor Temp.: 97.75 Ambient Temp.: 85.19 F
15:54:20.922 -> Average Cal sensor Temp.: 98.12 Ambient Temp.: 85.19 F
15:54:22.483 -> Average Cal sensor Temp.: 98.10 Ambient Temp.: 85.26 F
15:54:24.006 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 85.15 F
15:54:25.563 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 85.30 F
15:54:27.117 -> Average Cal sensor Temp.: 97.67 Ambient Temp.: 85.26 F
15:54:28.672 -> Average Cal sensor Temp.: 97.76 Ambient Temp.: 85.26 F
15:54:30.229 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 85.19 F
15:54:31.787 -> Average Cal sensor Temp.: 97.96 Ambient Temp.: 85.19 F
15:54:33.346 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 85.26 F
15:54:34.904 -> Average Cal sensor Temp.: 97.94 Ambient Temp.: 85.26 F
15:54:36.454 -> Average Cal sensor Temp.: 98.12 Ambient Temp.: 85.26 F
15:54:38.042 -> Average Cal sensor Temp.: 97.97 Ambient Temp.: 85.37 F
15:54:39.598 -> Average Cal sensor Temp.: 98.28 Ambient Temp.: 85.30 F
15:54:41.152 -> Average Cal sensor Temp.: 98.27 Ambient Temp.: 85.37 F
15:54:42.706 -> Average Cal sensor Temp.: 98.27 Ambient Temp.: 85.41 F
15:54:44.264 -> Average Cal sensor Temp.: 98.23 Ambient Temp.: 85.41 F
15:54:45.817 -> Average Cal sensor Temp.: 98.08 Ambient Temp.: 85.37 F
15:54:47.375 -> Average Cal sensor Temp.: 98.10 Ambient Temp.: 85.41 F
15:54:48.927 -> Average Cal sensor Temp.: 98.07 Ambient Temp.: 85.41 F
15:54:50.479 -> Average Cal sensor Temp.: 97.81 Ambient Temp.: 85.44 F
15:54:52.034 -> Average Cal sensor Temp.: 98.27 Ambient Temp.: 85.44 F
15:54:53.592 -> Average Cal sensor Temp.: 98.05 Ambient Temp.: 85.44 F
15:54:55.143 -> Average Cal sensor Temp.: 97.91 Ambient Temp.: 85.51 F
15:54:56.700 -> Average Cal sensor Temp.: 97.52 Ambient Temp.: 85.44 F
15:54:58.256 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 85.48 F
15:54:59.807 -> Average Cal sensor Temp.: 97.67 Ambient Temp.: 85.51 F
15:55:01.366 -> Average Cal sensor Temp.: 97.68 Ambient Temp.: 85.48 F
15:55:02.921 -> Average Cal sensor Temp.: 98.03 Ambient Temp.: 85.48 F
15:55:04.477 -> Average Cal sensor Temp.: 98.41 Ambient Temp.: 85.44 F
15:55:06.030 -> Average Cal sensor Temp.: 98.40 Ambient Temp.: 85.55 F
15:55:07.582 -> Average Cal sensor Temp.: 98.47 Ambient Temp.: 85.59 F
15:55:09.132 -> Average Cal sensor Temp.: 98.57 Ambient Temp.: 85.69 F
15:55:10.718 -> Average Cal sensor Temp.: 98.31 Ambient Temp.: 85.59 F
15:55:12.270 -> Average Cal sensor Temp.: 98.35 Ambient Temp.: 85.59 F
15:55:13.818 -> Average Cal sensor Temp.: 98.04 Ambient Temp.: 85.66 F
15:55:15.369 -> Average Cal sensor Temp.: 98.03 Ambient Temp.: 85.66 F
15:55:16.925 -> Average Cal sensor Temp.: 98.04 Ambient Temp.: 85.69 F
15:55:18.480 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 85.69 F
15:55:20.035 -> Average Cal sensor Temp.: 97.75 Ambient Temp.: 85.73 F
15:55:21.591 -> Average Cal sensor Temp.: 97.96 Ambient Temp.: 85.69 F
15:55:23.143 -> Average Cal sensor Temp.: 98.27 Ambient Temp.: 85.77 F
15:55:24.699 -> Average Cal sensor Temp.: 98.29 Ambient Temp.: 85.73 F
15:55:25.770 -> Ambient Temp.: 85.77 F
15:55:26.359 -> Ambient Temp.: 85.77 F
15:55:26.950 -> Ambient Temp.: 85.77 F
15:55:27.537 -> Ambient Temp.: 85.77 F
15:55:28.713 -> Ambient Temp.: 85.77 F
15:55:29.297 -> Ambient Temp.: 85.77 F
15:55:30.851 -> Average Cal sensor Temp.: 98.45 Ambient Temp.: 85.84 F
15:55:32.407 -> Average Cal sensor Temp.: 98.45 Ambient Temp.: 85.87 F
15:55:33.959 -> Average Cal sensor Temp.: 98.69 Ambient Temp.: 85.91 F
15:55:35.507 -> Average Cal sensor Temp.: 98.52 Ambient Temp.: 85.91 F
15:55:37.065 -> Average Cal sensor Temp.: 98.47 Ambient Temp.: 85.95 F
15:55:38.278 -> Ambient Temp.: 85.98 F
15:55:38.831 -> Ambient Temp.: 85.98 F
15:55:39.423 -> Ambient Temp.: 85.98 F
15:55:40.009 -> Ambient Temp.: 85.98 F
15:55:40.598 -> Ambient Temp.: 85.98 F
15:55:41.151 -> Ambient Temp.: 85.98 F
15:55:41.736 -> Ambient Temp.: 85.98 F
15:55:42.325 -> Ambient Temp.: 85.98 F
15:55:42.911 -> Ambient Temp.: 85.98 F
15:55:43.498 -> Ambient Temp.: 85.98 F
15:55:44.085 -> Ambient Temp.: 85.98 F
15:55:44.673 -> Ambient Temp.: 85.98 F
15:55:45.226 -> Ambient Temp.: 85.98 F
15:55:45.810 -> Ambient Temp.: 85.98 F
15:55:46.395 -> Ambient Temp.: 85.98 F
15:55:46.983 -> Ambient Temp.: 85.98 F
15:55:47.572 -> Ambient Temp.: 85.98 F
15:55:48.159 -> Ambient Temp.: 85.98 F
15:55:48.711 -> Ambient Temp.: 85.98 F
15:55:50.268 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 86.59 F
15:55:51.823 -> Average Cal sensor Temp.: 98.35 Ambient Temp.: 86.67 F
15:55:53.382 -> Average Cal sensor Temp.: 98.26 Ambient Temp.: 86.70 F
15:55:54.936 -> Average Cal sensor Temp.: 98.50 Ambient Temp.: 86.81 F
15:55:56.488 -> Average Cal sensor Temp.: 98.25 Ambient Temp.: 86.88 F
15:55:58.076 -> Average Cal sensor Temp.: 98.40 Ambient Temp.: 86.95 F
15:55:59.628 -> Average Cal sensor Temp.: 98.46 Ambient Temp.: 86.99 F
15:56:01.181 -> Average Cal sensor Temp.: 98.32 Ambient Temp.: 87.03 F
15:56:02.738 -> Average Cal sensor Temp.: 98.43 Ambient Temp.: 87.06 F
15:56:04.292 -> Average Cal sensor Temp.: 98.24 Ambient Temp.: 87.03 F
15:56:05.850 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 87.06 F
15:56:07.387 -> Average Cal sensor Temp.: 98.18 Ambient Temp.: 87.06 F
15:56:08.961 -> Average Cal sensor Temp.: 98.38 Ambient Temp.: 87.03 F
15:56:10.509 -> Average Cal sensor Temp.: 98.31 Ambient Temp.: 87.10 F
15:56:12.062 -> Average Cal sensor Temp.: 98.33 Ambient Temp.: 87.13 F
15:56:13.617 -> Average Cal sensor Temp.: 98.40 Ambient Temp.: 87.13 F
15:56:15.201 -> Average Cal sensor Temp.: 98.46 Ambient Temp.: 87.17 F
15:56:16.725 -> Average Cal sensor Temp.: 98.33 Ambient Temp.: 87.17 F
15:56:18.281 -> Average Cal sensor Temp.: 98.43 Ambient Temp.: 87.21 F
15:56:19.837 -> Average Cal sensor Temp.: 98.32 Ambient Temp.: 87.28 F
15:56:21.393 -> Average Cal sensor Temp.: 98.53 Ambient Temp.: 87.21 F
15:56:22.946 -> Average Cal sensor Temp.: 98.60 Ambient Temp.: 87.21 F
15:56:24.506 -> Average Cal sensor Temp.: 98.45 Ambient Temp.: 87.31 F
15:56:26.063 -> Average Cal sensor Temp.: 98.48 Ambient Temp.: 87.35 F
15:56:27.647 -> Average Cal sensor Temp.: 98.44 Ambient Temp.: 87.35 F
15:56:29.204 -> Average Cal sensor Temp.: 98.26 Ambient Temp.: 87.35 F
15:56:30.732 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 87.35 F
15:56:32.285 -> Average Cal sensor Temp.: 98.24 Ambient Temp.: 87.39 F
15:56:33.841 -> Average Cal sensor Temp.: 98.38 Ambient Temp.: 87.39 F
15:56:35.425 -> Average Cal sensor Temp.: 98.12 Ambient Temp.: 87.39 F
15:56:36.980 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 87.39 F
15:56:38.536 -> Average Cal sensor Temp.: 98.10 Ambient Temp.: 87.35 F
15:56:40.091 -> Average Cal sensor Temp.: 98.02 Ambient Temp.: 87.39 F
15:56:41.649 -> Average Cal sensor Temp.: 97.66 Ambient Temp.: 87.39 F
15:56:43.200 -> Average Cal sensor Temp.: 97.56 Ambient Temp.: 87.35 F
15:56:44.753 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 87.31 F
15:56:46.306 -> Average Cal sensor Temp.: 97.97 Ambient Temp.: 87.35 F
15:56:47.863 -> Average Cal sensor Temp.: 98.49 Ambient Temp.: 87.35 F
15:56:49.418 -> Average Cal sensor Temp.: 98.19 Ambient Temp.: 87.39 F
15:56:50.972 -> Average Cal sensor Temp.: 98.27 Ambient Temp.: 87.39 F
15:56:52.522 -> Average Cal sensor Temp.: 97.85 Ambient Temp.: 87.31 F
15:56:54.072 -> Average Cal sensor Temp.: 98.02 Ambient Temp.: 87.28 F
15:56:55.656 -> Average Cal sensor Temp.: 98.40 Ambient Temp.: 87.28 F
15:56:57.207 -> Average Cal sensor Temp.: 98.02 Ambient Temp.: 87.21 F
15:56:58.761 -> Average Cal sensor Temp.: 98.14 Ambient Temp.: 87.28 F
15:57:00.316 -> Average Cal sensor Temp.: 98.49 Ambient Temp.: 87.28 F
15:57:01.872 -> Average Cal sensor Temp.: 98.36 Ambient Temp.: 87.31 F
15:57:03.422 -> Average Cal sensor Temp.: 98.09 Ambient Temp.: 87.28 F
15:57:04.977 -> Average Cal sensor Temp.: 98.26 Ambient Temp.: 87.28 F
15:57:06.529 -> Average Cal sensor Temp.: 97.96 Ambient Temp.: 87.21 F
15:57:08.082 -> Average Cal sensor Temp.: 97.98 Ambient Temp.: 87.21 F
15:57:09.638 -> Average Cal sensor Temp.: 97.89 Ambient Temp.: 87.21 F
15:57:11.197 -> Average Cal sensor Temp.: 97.74 Ambient Temp.: 87.17 F
15:57:12.748 -> Average Cal sensor Temp.: 97.87 Ambient Temp.: 87.28 F
15:57:14.333 -> Average Cal sensor Temp.: 98.48 Ambient Temp.: 87.21 F
15:57:15.886 -> Average Cal sensor Temp.: 98.25 Ambient Temp.: 87.17 F
15:57:17.437 -> Average Cal sensor Temp.: 97.84 Ambient Temp.: 87.17 F
15:57:18.997 -> Average Cal sensor Temp.: 98.19 Ambient Temp.: 87.17 F
15:57:20.556 -> Average Cal sensor Temp.: 98.17 Ambient Temp.: 87.13 F
15:57:22.082 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 87.13 F
15:57:23.639 -> Average Cal sensor Temp.: 98.04 Ambient Temp.: 87.10 F
15:57:25.196 -> Average Cal sensor Temp.: 98.08 Ambient Temp.: 87.10 F
15:57:26.753 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 87.06 F
15:57:28.311 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 87.13 F
15:57:29.869 -> Average Cal sensor Temp.: 98.08 Ambient Temp.: 87.10 F
15:57:31.422 -> Average Cal sensor Temp.: 98.48 Ambient Temp.: 87.10 F
15:57:33.007 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 87.06 F
15:57:34.558 -> Average Cal sensor Temp.: 98.18 Ambient Temp.: 87.10 F
15:57:36.108 -> Average Cal sensor Temp.: 97.75 Ambient Temp.: 87.06 F
15:57:37.664 -> Average Cal sensor Temp.: 98.29 Ambient Temp.: 87.06 F
15:57:39.223 -> Average Cal sensor Temp.: 98.56 Ambient Temp.: 87.06 F
15:57:40.780 -> Average Cal sensor Temp.: 98.46 Ambient Temp.: 87.13 F
15:57:42.341 -> Average Cal sensor Temp.: 98.03 Ambient Temp.: 87.10 F
15:57:43.886 -> Average Cal sensor Temp.: 98.45 Ambient Temp.: 87.06 F
15:57:45.440 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 87.03 F
15:57:46.997 -> Average Cal sensor Temp.: 98.19 Ambient Temp.: 87.06 F
15:57:48.551 -> Average Cal sensor Temp.: 97.83 Ambient Temp.: 87.03 F
15:57:50.109 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 86.99 F
15:57:51.662 -> Average Cal sensor Temp.: 98.23 Ambient Temp.: 86.99 F
15:57:53.243 -> Average Cal sensor Temp.: 98.18 Ambient Temp.: 86.99 F
15:57:54.793 -> Average Cal sensor Temp.: 97.73 Ambient Temp.: 86.95 F
15:57:56.350 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 86.95 F
15:57:57.897 -> Average Cal sensor Temp.: 98.23 Ambient Temp.: 86.92 F
15:57:59.446 -> Average Cal sensor Temp.: 98.29 Ambient Temp.: 86.92 F
15:58:01.031 -> Average Cal sensor Temp.: 98.28 Ambient Temp.: 86.92 F
15:58:02.585 -> Average Cal sensor Temp.: 98.46 Ambient Temp.: 86.92 F
15:58:04.143 -> Average Cal sensor Temp.: 98.13 Ambient Temp.: 86.95 F
15:58:05.699 -> Average Cal sensor Temp.: 98.29 Ambient Temp.: 86.92 F
15:58:07.231 -> Average Cal sensor Temp.: 98.14 Ambient Temp.: 86.88 F
15:58:08.813 -> Average Cal sensor Temp.: 98.08 Ambient Temp.: 86.92 F
15:58:10.366 -> Average Cal sensor Temp.: 98.28 Ambient Temp.: 86.92 F
15:58:11.923 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 86.88 F
15:58:13.476 -> Average Cal sensor Temp.: 98.28 Ambient Temp.: 86.88 F
15:58:15.030 -> Average Cal sensor Temp.: 97.95 Ambient Temp.: 86.81 F
15:58:16.582 -> Average Cal sensor Temp.: 98.68 Ambient Temp.: 86.88 F
15:58:18.164 -> Average Cal sensor Temp.: 98.28 Ambient Temp.: 86.81 F
15:58:19.715 -> Average Cal sensor Temp.: 98.40 Ambient Temp.: 86.74 F
15:58:21.266 -> Average Cal sensor Temp.: 98.23 Ambient Temp.: 86.81 F
15:58:22.817 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 86.70 F
15:58:24.373 -> Average Cal sensor Temp.: 98.10 Ambient Temp.: 86.70 F
15:58:25.926 -> Average Cal sensor Temp.: 98.16 Ambient Temp.: 86.70 F
15:58:27.480 -> Average Cal sensor Temp.: 98.62 Ambient Temp.: 86.67 F
15:58:29.067 -> Average Cal sensor Temp.: 98.22 Ambient Temp.: 86.67 F
15:58:30.620 -> Average Cal sensor Temp.: 98.27 Ambient Temp.: 86.67 F
15:58:32.177 -> Average Cal sensor Temp.: 97.97 Ambient Temp.: 86.67 F
15:58:33.730 -> Average Cal sensor Temp.: 98.04 Ambient Temp.: 86.59 F
15:58:35.284 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 86.56 F
15:58:36.839 -> Average Cal sensor Temp.: 97.89 Ambient Temp.: 86.52 F
15:58:38.392 -> Average Cal sensor Temp.: 97.87 Ambient Temp.: 86.49 F
15:58:39.950 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 86.52 F
15:58:41.504 -> Average Cal sensor Temp.: 98.35 Ambient Temp.: 86.45 F
15:58:43.061 -> Average Cal sensor Temp.: 98.19 Ambient Temp.: 86.45 F
15:58:44.618 -> Average Cal sensor Temp.: 98.35 Ambient Temp.: 86.45 F
15:58:46.167 -> Average Cal sensor Temp.: 97.83 Ambient Temp.: 86.45 F
15:58:47.727 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 86.31 F
15:58:49.282 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 86.34 F
15:58:50.835 -> Average Cal sensor Temp.: 97.63 Ambient Temp.: 86.31 F
15:58:52.394 -> Average Cal sensor Temp.: 98.02 Ambient Temp.: 86.31 F
15:58:53.953 -> Average Cal sensor Temp.: 97.99 Ambient Temp.: 86.31 F
15:58:55.503 -> Average Cal sensor Temp.: 98.43 Ambient Temp.: 86.23 F
15:58:57.090 -> Average Cal sensor Temp.: 98.45 Ambient Temp.: 86.20 F
15:58:58.614 -> Average Cal sensor Temp.: 98.44 Ambient Temp.: 86.20 F
15:59:00.203 -> Average Cal sensor Temp.: 98.29 Ambient Temp.: 86.23 F
15:59:01.762 -> Average Cal sensor Temp.: 98.18 Ambient Temp.: 86.16 F
15:59:03.287 -> Average Cal sensor Temp.: 98.19 Ambient Temp.: 86.13 F
15:59:04.874 -> Average Cal sensor Temp.: 98.49 Ambient Temp.: 86.13 F
15:59:06.433 -> Average Cal sensor Temp.: 98.62 Ambient Temp.: 86.16 F
15:59:07.957 -> Average Cal sensor Temp.: 98.24 Ambient Temp.: 86.09 F
15:59:09.547 -> Average Cal sensor Temp.: 98.47 Ambient Temp.: 86.09 F
15:59:11.100 -> Average Cal sensor Temp.: 98.24 Ambient Temp.: 86.05 F
15:59:12.649 -> Average Cal sensor Temp.: 97.85 Ambient Temp.: 86.09 F
15:59:14.202 -> Average Cal sensor Temp.: 97.84 Ambient Temp.: 86.05 F
15:59:15.756 -> Average Cal sensor Temp.: 98.36 Ambient Temp.: 85.98 F
15:59:17.318 -> Average Cal sensor Temp.: 98.23 Ambient Temp.: 85.98 F
15:59:18.875 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 85.91 F
15:59:20.431 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 85.95 F
15:59:21.988 -> Average Cal sensor Temp.: 97.68 Ambient Temp.: 85.87 F
15:59:23.547 -> Average Cal sensor Temp.: 97.47 Ambient Temp.: 85.87 F
15:59:25.102 -> Average Cal sensor Temp.: 97.52 Ambient Temp.: 85.80 F
15:59:26.660 -> Average Cal sensor Temp.: 97.94 Ambient Temp.: 85.84 F
15:59:28.211 -> Average Cal sensor Temp.: 97.50 Ambient Temp.: 85.80 F
15:59:29.758 -> Average Cal sensor Temp.: 97.78 Ambient Temp.: 85.80 F
15:59:31.310 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 85.73 F
15:59:32.897 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 85.66 F
15:59:34.452 -> Average Cal sensor Temp.: 97.87 Ambient Temp.: 85.66 F
15:59:36.006 -> Average Cal sensor Temp.: 98.04 Ambient Temp.: 85.66 F
15:59:37.562 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 85.55 F
15:59:39.123 -> Average Cal sensor Temp.: 98.16 Ambient Temp.: 85.51 F
15:59:40.678 -> Average Cal sensor Temp.: 97.72 Ambient Temp.: 85.55 F
15:59:42.227 -> Average Cal sensor Temp.: 97.70 Ambient Temp.: 85.44 F
15:59:43.780 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 85.48 F
15:59:45.332 -> Average Cal sensor Temp.: 98.25 Ambient Temp.: 85.44 F
15:59:46.896 -> Average Cal sensor Temp.: 98.12 Ambient Temp.: 85.44 F
15:59:48.455 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 85.41 F
15:59:50.005 -> Average Cal sensor Temp.: 98.47 Ambient Temp.: 85.41 F
15:59:51.564 -> Average Cal sensor Temp.: 98.46 Ambient Temp.: 85.41 F
15:59:53.116 -> Average Cal sensor Temp.: 98.29 Ambient Temp.: 85.41 F
15:59:54.667 -> Average Cal sensor Temp.: 98.12 Ambient Temp.: 85.41 F
15:59:56.224 -> Average Cal sensor Temp.: 98.22 Ambient Temp.: 85.33 F
15:59:57.781 -> Average Cal sensor Temp.: 98.44 Ambient Temp.: 85.41 F
15:59:59.332 -> Average Cal sensor Temp.: 97.82 Ambient Temp.: 85.33 F
16:00:00.885 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 85.37 F
16:00:02.442 -> Average Cal sensor Temp.: 98.43 Ambient Temp.: 85.33 F
16:00:03.999 -> Average Cal sensor Temp.: 98.41 Ambient Temp.: 85.41 F
16:00:05.561 -> Average Cal sensor Temp.: 98.13 Ambient Temp.: 85.41 F
16:00:07.113 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 85.41 F
16:00:08.664 -> Average Cal sensor Temp.: 97.91 Ambient Temp.: 85.41 F
16:00:10.222 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 85.37 F
16:00:11.771 -> Average Cal sensor Temp.: 97.78 Ambient Temp.: 85.33 F
16:00:13.326 -> Average Cal sensor Temp.: 97.81 Ambient Temp.: 85.30 F
16:00:14.880 -> Average Cal sensor Temp.: 97.85 Ambient Temp.: 85.37 F
16:00:16.438 -> Average Cal sensor Temp.: 98.08 Ambient Temp.: 85.37 F
16:00:17.996 -> Average Cal sensor Temp.: 98.09 Ambient Temp.: 85.33 F
16:00:19.551 -> Average Cal sensor Temp.: 98.26 Ambient Temp.: 85.30 F
16:00:21.102 -> Average Cal sensor Temp.: 98.51 Ambient Temp.: 85.33 F
16:00:22.656 -> Average Cal sensor Temp.: 98.43 Ambient Temp.: 85.41 F
16:00:24.213 -> Average Cal sensor Temp.: 98.12 Ambient Temp.: 85.37 F
16:00:25.765 -> Average Cal sensor Temp.: 98.27 Ambient Temp.: 85.33 F
16:00:27.322 -> Average Cal sensor Temp.: 98.22 Ambient Temp.: 85.33 F
16:00:28.874 -> Average Cal sensor Temp.: 98.38 Ambient Temp.: 85.37 F
16:00:30.462 -> Average Cal sensor Temp.: 98.54 Ambient Temp.: 85.41 F
16:00:32.014 -> Average Cal sensor Temp.: 98.84 Ambient Temp.: 85.37 F
16:00:33.569 -> Average Cal sensor Temp.: 98.44 Ambient Temp.: 85.37 F
16:00:35.126 -> Average Cal sensor Temp.: 98.46 Ambient Temp.: 85.37 F
16:00:36.679 -> Average Cal sensor Temp.: 98.22 Ambient Temp.: 85.33 F
16:00:38.234 -> Average Cal sensor Temp.: 98.47 Ambient Temp.: 85.41 F
16:00:39.785 -> Average Cal sensor Temp.: 99.19 Ambient Temp.: 85.48 F
16:00:41.342 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 85.41 F
16:00:42.896 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 85.41 F
16:00:44.447 -> Average Cal sensor Temp.: 98.60 Ambient Temp.: 85.44 F
16:00:45.999 -> Average Cal sensor Temp.: 98.64 Ambient Temp.: 85.44 F
16:00:47.555 -> Average Cal sensor Temp.: 98.52 Ambient Temp.: 85.41 F
16:00:49.110 -> Average Cal sensor Temp.: 98.88 Ambient Temp.: 85.44 F
16:00:50.666 -> Average Cal sensor Temp.: 98.72 Ambient Temp.: 85.48 F
16:00:52.254 -> Average Cal sensor Temp.: 98.82 Ambient Temp.: 85.48 F
16:00:53.805 -> Average Cal sensor Temp.: 98.69 Ambient Temp.: 85.51 F
16:00:55.357 -> Average Cal sensor Temp.: 98.48 Ambient Temp.: 85.48 F
16:00:56.913 -> Average Cal sensor Temp.: 98.27 Ambient Temp.: 85.44 F
16:00:58.464 -> Average Cal sensor Temp.: 98.62 Ambient Temp.: 85.51 F
16:01:00.018 -> Average Cal sensor Temp.: 98.68 Ambient Temp.: 85.51 F
16:01:01.572 -> Average Cal sensor Temp.: 98.66 Ambient Temp.: 85.51 F
16:01:03.122 -> Average Cal sensor Temp.: 98.85 Ambient Temp.: 85.59 F
16:01:04.679 -> Average Cal sensor Temp.: 98.74 Ambient Temp.: 85.59 F
16:01:06.237 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 85.59 F
16:01:07.823 -> Average Cal sensor Temp.: 98.32 Ambient Temp.: 85.59 F
16:01:09.348 -> Average Cal sensor Temp.: 98.52 Ambient Temp.: 85.66 F
16:01:10.911 -> Average Cal sensor Temp.: 98.47 Ambient Temp.: 85.66 F
16:01:12.467 -> Average Cal sensor Temp.: 98.51 Ambient Temp.: 85.69 F
16:01:14.018 -> Average Cal sensor Temp.: 98.67 Ambient Temp.: 85.66 F
16:01:15.606 -> Average Cal sensor Temp.: 98.85 Ambient Temp.: 85.73 F
16:01:17.165 -> Average Cal sensor Temp.: 98.43 Ambient Temp.: 85.69 F
16:01:18.688 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 85.77 F
16:01:20.274 -> Average Cal sensor Temp.: 98.25 Ambient Temp.: 85.69 F
16:01:21.830 -> Average Cal sensor Temp.: 98.51 Ambient Temp.: 85.73 F
16:01:23.385 -> Average Cal sensor Temp.: 98.84 Ambient Temp.: 85.73 F
16:01:24.938 -> Average Cal sensor Temp.: 98.42 Ambient Temp.: 85.77 F
16:01:26.495 -> Average Cal sensor Temp.: 98.65 Ambient Temp.: 85.73 F
16:01:28.052 -> Average Cal sensor Temp.: 98.78 Ambient Temp.: 85.77 F
16:01:29.605 -> Average Cal sensor Temp.: 98.93 Ambient Temp.: 85.77 F
16:01:31.158 -> Average Cal sensor Temp.: 98.75 Ambient Temp.: 85.80 F
16:01:32.718 -> Average Cal sensor Temp.: 98.48 Ambient Temp.: 85.80 F
16:01:34.279 -> Average Cal sensor Temp.: 98.57 Ambient Temp.: 85.84 F
16:01:35.834 -> Average Cal sensor Temp.: 98.42 Ambient Temp.: 85.84 F
16:01:37.395 -> Average Cal sensor Temp.: 98.46 Ambient Temp.: 85.84 F
16:01:38.950 -> Average Cal sensor Temp.: 98.43 Ambient Temp.: 85.80 F
16:01:40.504 -> Average Cal sensor Temp.: 98.45 Ambient Temp.: 85.91 F
16:01:42.059 -> Average Cal sensor Temp.: 98.19 Ambient Temp.: 85.87 F
16:01:43.616 -> Average Cal sensor Temp.: 98.31 Ambient Temp.: 85.91 F
16:01:45.172 -> Average Cal sensor Temp.: 98.42 Ambient Temp.: 85.87 F
16:01:46.726 -> Average Cal sensor Temp.: 99.19 Ambient Temp.: 85.91 F
16:01:48.278 -> Average Cal sensor Temp.: 98.37 Ambient Temp.: 85.95 F
16:01:49.829 -> Average Cal sensor Temp.: 98.60 Ambient Temp.: 85.95 F
16:01:51.385 -> Average Cal sensor Temp.: 98.93 Ambient Temp.: 85.98 F
16:01:52.940 -> Average Cal sensor Temp.: 98.59 Ambient Temp.: 85.95 F
16:01:54.501 -> Average Cal sensor Temp.: 98.48 Ambient Temp.: 86.05 F
16:01:56.052 -> Average Cal sensor Temp.: 98.51 Ambient Temp.: 85.95 F
16:01:57.608 -> Average Cal sensor Temp.: 98.45 Ambient Temp.: 85.98 F
16:01:59.164 -> Average Cal sensor Temp.: 98.47 Ambient Temp.: 86.05 F
16:02:00.751 -> Average Cal sensor Temp.: 98.39 Ambient Temp.: 86.09 F
16:02:02.306 -> Average Cal sensor Temp.: 98.57 Ambient Temp.: 86.13 F
16:02:03.862 -> Average Cal sensor Temp.: 98.62 Ambient Temp.: 86.13 F
16:02:05.418 -> Average Cal sensor Temp.: 98.59 Ambient Temp.: 86.16 F
16:02:06.971 -> Average Cal sensor Temp.: 98.60 Ambient Temp.: 86.16 F
16:02:08.528 -> Average Cal sensor Temp.: 98.64 Ambient Temp.: 86.20 F
16:02:10.078 -> Average Cal sensor Temp.: 98.65 Ambient Temp.: 86.27 F
16:02:11.632 -> Average Cal sensor Temp.: 98.67 Ambient Temp.: 86.27 F
16:02:13.180 -> Average Cal sensor Temp.: 98.56 Ambient Temp.: 86.27 F
16:02:14.763 -> Average Cal sensor Temp.: 98.41 Ambient Temp.: 86.27 F
16:02:16.322 -> Average Cal sensor Temp.: 98.50 Ambient Temp.: 86.34 F
16:02:17.871 -> Average Cal sensor Temp.: 98.55 Ambient Temp.: 86.31 F
16:02:19.430 -> Average Cal sensor Temp.: 99.05 Ambient Temp.: 86.34 F
16:02:20.986 -> Average Cal sensor Temp.: 98.64 Ambient Temp.: 86.38 F
16:02:22.540 -> Average Cal sensor Temp.: 98.53 Ambient Temp.: 86.38 F
16:02:24.088 -> Average Cal sensor Temp.: 98.69 Ambient Temp.: 86.45 F
16:02:25.642 -> Average Cal sensor Temp.: 98.72 Ambient Temp.: 86.49 F
16:02:27.197 -> Average Cal sensor Temp.: 98.64 Ambient Temp.: 86.49 F
16:02:28.752 -> Average Cal sensor Temp.: 98.69 Ambient Temp.: 86.52 F
16:02:30.312 -> Average Cal sensor Temp.: 98.64 Ambient Temp.: 86.56 F
16:02:31.867 -> Average Cal sensor Temp.: 98.86 Ambient Temp.: 86.59 F
16:02:33.449 -> Average Cal sensor Temp.: 98.83 Ambient Temp.: 86.59 F
16:02:35.000 -> Average Cal sensor Temp.: 98.89 Ambient Temp.: 86.59 F
16:02:36.558 -> Average Cal sensor Temp.: 98.74 Ambient Temp.: 86.67 F
16:02:38.110 -> Average Cal sensor Temp.: 98.76 Ambient Temp.: 86.70 F
16:02:39.663 -> Average Cal sensor Temp.: 98.69 Ambient Temp.: 86.70 F
16:02:41.222 -> Average Cal sensor Temp.: 98.68 Ambient Temp.: 86.74 F
16:02:42.776 -> Average Cal sensor Temp.: 98.68 Ambient Temp.: 86.74 F
16:02:44.332 -> Average Cal sensor Temp.: 98.33 Ambient Temp.: 86.70 F
16:02:45.889 -> Average Cal sensor Temp.: 98.75 Ambient Temp.: 86.77 F
16:02:47.442 -> Average Cal sensor Temp.: 98.36 Ambient Temp.: 86.81 F
16:02:49.028 -> Average Cal sensor Temp.: 98.65 Ambient Temp.: 86.81 F
16:02:50.585 -> Average Cal sensor Temp.: 98.73 Ambient Temp.: 86.81 F
16:02:52.136 -> Average Cal sensor Temp.: 98.46 Ambient Temp.: 86.81 F
16:02:53.692 -> Average Cal sensor Temp.: 98.58 Ambient Temp.: 86.88 F
16:02:55.243 -> Average Cal sensor Temp.: 98.45 Ambient Temp.: 86.95 F
16:02:56.803 -> Average Cal sensor Temp.: 98.37 Ambient Temp.: 86.95 F
16:02:58.360 -> Average Cal sensor Temp.: 98.64 Ambient Temp.: 86.92 F
16:02:59.923 -> Average Cal sensor Temp.: 98.64 Ambient Temp.: 86.99 F
16:03:01.473 -> Average Cal sensor Temp.: 98.59 Ambient Temp.: 86.95 F
16:03:03.051 -> Average Cal sensor Temp.: 98.60 Ambient Temp.: 87.03 F
16:03:04.596 -> Average Cal sensor Temp.: 98.50 Ambient Temp.: 87.03 F
16:03:06.149 -> Average Cal sensor Temp.: 98.42 Ambient Temp.: 87.03 F
16:03:07.699 -> Average Cal sensor Temp.: 98.27 Ambient Temp.: 87.06 F
16:03:09.250 -> Average Cal sensor Temp.: 98.33 Ambient Temp.: 87.06 F
16:03:10.835 -> Average Cal sensor Temp.: 98.46 Ambient Temp.: 87.03 F
16:03:12.389 -> Average Cal sensor Temp.: 98.78 Ambient Temp.: 87.06 F
16:03:13.937 -> Average Cal sensor Temp.: 98.44 Ambient Temp.: 87.06 F
16:03:15.485 -> Average Cal sensor Temp.: 98.48 Ambient Temp.: 87.10 F
16:03:17.061 -> Average Cal sensor Temp.: 98.31 Ambient Temp.: 87.03 F
16:03:18.608 -> Average Cal sensor Temp.: 98.60 Ambient Temp.: 87.10 F
16:03:20.149 -> Average Cal sensor Temp.: 98.67 Ambient Temp.: 87.06 F
16:03:21.734 -> Average Cal sensor Temp.: 98.65 Ambient Temp.: 87.10 F
16:03:23.272 -> Average Cal sensor Temp.: 98.53 Ambient Temp.: 87.13 F
16:03:24.848 -> Average Cal sensor Temp.: 98.38 Ambient Temp.: 87.13 F
16:03:26.398 -> Average Cal sensor Temp.: 98.29 Ambient Temp.: 87.21 F
16:03:27.947 -> Average Cal sensor Temp.: 98.51 Ambient Temp.: 87.13 F
16:03:29.493 -> Average Cal sensor Temp.: 98.39 Ambient Temp.: 87.10 F
16:03:31.038 -> Average Cal sensor Temp.: 98.28 Ambient Temp.: 87.10 F
16:03:32.617 -> Average Cal sensor Temp.: 98.50 Ambient Temp.: 87.13 F
16:03:34.162 -> Average Cal sensor Temp.: 98.61 Ambient Temp.: 87.13 F
16:03:35.709 -> Average Cal sensor Temp.: 97.81 Ambient Temp.: 87.10 F
16:03:37.293 -> Average Cal sensor Temp.: 97.58 Ambient Temp.: 87.03 F
16:03:38.838 -> Average Cal sensor Temp.: 97.51 Ambient Temp.: 86.99 F
16:03:40.382 -> Average Cal sensor Temp.: 97.54 Ambient Temp.: 87.06 F
16:03:41.957 -> Average Cal sensor Temp.: 97.94 Ambient Temp.: 86.95 F
16:03:43.508 -> Average Cal sensor Temp.: 98.07 Ambient Temp.: 86.95 F
16:03:45.059 -> Average Cal sensor Temp.: 97.50 Ambient Temp.: 86.95 F
16:03:46.603 -> Average Cal sensor Temp.: 97.87 Ambient Temp.: 86.92 F
16:03:48.184 -> Average Cal sensor Temp.: 98.27 Ambient Temp.: 86.92 F
16:03:49.742 -> Average Cal sensor Temp.: 98.27 Ambient Temp.: 86.88 F
16:03:51.290 -> Average Cal sensor Temp.: 98.19 Ambient Temp.: 86.81 F
16:03:52.839 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 86.81 F
16:03:54.382 -> Average Cal sensor Temp.: 97.76 Ambient Temp.: 86.77 F
16:03:55.967 -> Average Cal sensor Temp.: 98.00 Ambient Temp.: 86.77 F
16:03:57.520 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 86.77 F
16:03:59.070 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 86.74 F
16:04:00.621 -> Average Cal sensor Temp.: 98.28 Ambient Temp.: 86.74 F
16:04:02.175 -> Average Cal sensor Temp.: 98.31 Ambient Temp.: 86.74 F
16:04:03.732 -> Average Cal sensor Temp.: 98.65 Ambient Temp.: 86.74 F
16:04:05.280 -> Average Cal sensor Temp.: 98.55 Ambient Temp.: 86.70 F
16:04:06.834 -> Average Cal sensor Temp.: 98.13 Ambient Temp.: 86.70 F
16:04:07.422 -> Ambient Temp.: 86.70 F
16:04:08.004 -> Ambient Temp.: 86.70 F
16:04:08.591 -> Ambient Temp.: 86.70 F
16:04:09.179 -> Ambient Temp.: 86.70 F
16:04:10.732 -> Average Cal sensor Temp.: 98.49 Ambient Temp.: 86.67 F
16:04:12.287 -> Average Cal sensor Temp.: 98.33 Ambient Temp.: 86.63 F
16:04:13.842 -> Average Cal sensor Temp.: 98.24 Ambient Temp.: 86.59 F
16:04:15.395 -> Average Cal sensor Temp.: 98.31 Ambient Temp.: 86.63 F
16:04:16.953 -> Average Cal sensor Temp.: 98.42 Ambient Temp.: 86.63 F
16:04:18.502 -> Average Cal sensor Temp.: 98.07 Ambient Temp.: 86.59 F
16:04:20.088 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 86.67 F
16:04:21.637 -> Average Cal sensor Temp.: 98.13 Ambient Temp.: 86.63 F
16:04:23.193 -> Average Cal sensor Temp.: 98.45 Ambient Temp.: 86.56 F
16:04:24.748 -> Average Cal sensor Temp.: 97.97 Ambient Temp.: 86.59 F
16:04:26.303 -> Average Cal sensor Temp.: 98.29 Ambient Temp.: 86.63 F
16:04:27.861 -> Average Cal sensor Temp.: 98.22 Ambient Temp.: 86.63 F
16:04:29.418 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 86.59 F
16:04:30.977 -> Average Cal sensor Temp.: 98.16 Ambient Temp.: 86.59 F
16:04:32.535 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 86.56 F
16:04:34.085 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 86.63 F
16:04:35.641 -> Average Cal sensor Temp.: 97.93 Ambient Temp.: 86.59 F
16:04:37.197 -> Average Cal sensor Temp.: 98.13 Ambient Temp.: 86.52 F
16:04:38.749 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 86.56 F
16:04:40.305 -> Average Cal sensor Temp.: 98.19 Ambient Temp.: 86.52 F
16:04:41.860 -> Average Cal sensor Temp.: 98.27 Ambient Temp.: 86.52 F
16:04:43.413 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 86.52 F
16:04:44.969 -> Average Cal sensor Temp.: 97.83 Ambient Temp.: 86.52 F
16:04:46.533 -> Average Cal sensor Temp.: 98.31 Ambient Temp.: 86.45 F
16:04:48.090 -> Average Cal sensor Temp.: 98.46 Ambient Temp.: 86.49 F
16:04:49.650 -> Average Cal sensor Temp.: 98.48 Ambient Temp.: 86.45 F
16:04:51.199 -> Average Cal sensor Temp.: 98.60 Ambient Temp.: 86.45 F
16:04:52.757 -> Average Cal sensor Temp.: 98.65 Ambient Temp.: 86.49 F
16:04:54.341 -> Average Cal sensor Temp.: 98.27 Ambient Temp.: 86.49 F
16:04:55.897 -> Average Cal sensor Temp.: 98.72 Ambient Temp.: 86.45 F
16:04:57.454 -> Average Cal sensor Temp.: 98.45 Ambient Temp.: 86.49 F
16:04:59.003 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 86.34 F
16:05:00.558 -> Average Cal sensor Temp.: 98.37 Ambient Temp.: 86.38 F
16:05:02.109 -> Average Cal sensor Temp.: 98.78 Ambient Temp.: 86.38 F
16:05:03.664 -> Average Cal sensor Temp.: 98.71 Ambient Temp.: 86.45 F
16:05:05.217 -> Average Cal sensor Temp.: 98.46 Ambient Temp.: 86.45 F
16:05:06.774 -> Average Cal sensor Temp.: 98.38 Ambient Temp.: 86.45 F
16:05:08.329 -> Average Cal sensor Temp.: 97.95 Ambient Temp.: 86.45 F
16:05:09.885 -> Average Cal sensor Temp.: 97.94 Ambient Temp.: 86.34 F
16:05:11.440 -> Average Cal sensor Temp.: 98.39 Ambient Temp.: 86.38 F
16:05:13.027 -> Average Cal sensor Temp.: 98.32 Ambient Temp.: 86.34 F
16:05:14.578 -> Average Cal sensor Temp.: 98.37 Ambient Temp.: 86.34 F
16:05:16.135 -> Average Cal sensor Temp.: 98.31 Ambient Temp.: 86.34 F
16:05:16.721 -> Ambient Temp.: 86.34 F
16:05:17.276 -> Ambient Temp.: 86.34 F
16:05:17.863 -> Ambient Temp.: 86.34 F
16:05:18.448 -> Ambient Temp.: 86.34 F
16:05:19.036 -> Ambient Temp.: 86.34 F
16:05:19.622 -> Ambient Temp.: 86.34 F
16:05:20.209 -> Ambient Temp.: 86.34 F
16:05:20.763 -> Ambient Temp.: 86.34 F
16:05:21.352 -> Ambient Temp.: 86.34 F
16:05:21.934 -> Ambient Temp.: 86.34 F
16:05:22.519 -> Ambient Temp.: 86.34 F
16:05:23.107 -> Ambient Temp.: 86.34 F
16:05:23.695 -> Ambient Temp.: 86.34 F
16:05:24.280 -> Ambient Temp.: 86.34 F
16:05:24.834 -> Ambient Temp.: 86.34 F
16:05:25.421 -> Ambient Temp.: 86.34 F
16:05:26.011 -> Ambient Temp.: 86.34 F
16:05:26.600 -> Ambient Temp.: 86.34 F
16:05:28.155 -> Average Cal sensor Temp.: 98.53 Ambient Temp.: 86.27 F
16:05:29.709 -> Average Cal sensor Temp.: 98.48 Ambient Temp.: 86.31 F
16:05:31.259 -> Average Cal sensor Temp.: 98.46 Ambient Temp.: 86.20 F
16:05:32.806 -> Average Cal sensor Temp.: 98.86 Ambient Temp.: 86.31 F
16:05:33.395 -> Ambient Temp.: 86.31 F
16:05:33.984 -> Ambient Temp.: 86.31 F
16:05:34.573 -> Ambient Temp.: 86.31 F
16:05:35.124 -> Ambient Temp.: 86.31 F
16:05:35.714 -> Ambient Temp.: 86.31 F
16:05:36.303 -> Ambient Temp.: 86.31 F
16:05:36.888 -> Ambient Temp.: 86.31 F
16:05:37.472 -> Ambient Temp.: 86.31 F
16:05:38.057 -> Ambient Temp.: 86.31 F
16:05:38.610 -> Ambient Temp.: 86.31 F
16:05:39.197 -> Ambient Temp.: 86.31 F
16:05:39.787 -> Ambient Temp.: 86.31 F
16:05:40.376 -> Ambient Temp.: 86.31 F
16:05:40.925 -> Ambient Temp.: 86.31 F
16:05:41.517 -> Ambient Temp.: 86.31 F
16:05:42.103 -> Ambient Temp.: 86.31 F
16:05:42.688 -> Ambient Temp.: 86.31 F
16:05:43.277 -> Ambient Temp.: 86.31 F
16:05:43.833 -> Ambient Temp.: 86.31 F
16:05:44.420 -> Ambient Temp.: 86.31 F
16:05:45.006 -> Ambient Temp.: 86.31 F
16:05:45.594 -> Ambient Temp.: 86.31 F
16:05:46.148 -> Ambient Temp.: 86.31 F
16:05:46.737 -> Ambient Temp.: 86.31 F
16:05:47.325 -> Ambient Temp.: 86.31 F
16:05:48.883 -> Average Cal sensor Temp.: 98.37 Ambient Temp.: 86.23 F
16:05:50.438 -> Average Cal sensor Temp.: 98.33 Ambient Temp.: 86.20 F
16:05:51.751 -> Ambient Temp.: 86.16 F
16:05:52.336 -> Ambient Temp.: 86.16 F
16:05:52.891 -> Ambient Temp.: 86.16 F
16:05:53.480 -> Ambient Temp.: 86.16 F
16:05:54.066 -> Ambient Temp.: 86.16 F
16:05:54.656 -> Ambient Temp.: 86.16 F
16:05:55.239 -> Ambient Temp.: 86.16 F
16:05:55.827 -> Ambient Temp.: 86.16 F
16:05:56.523 -> Ambient Temp.: 86.13 F
16:05:57.108 -> Ambient Temp.: 86.13 F
16:05:57.664 -> Ambient Temp.: 86.13 F
16:05:59.221 -> Average Cal sensor Temp.: 98.62 Ambient Temp.: 86.20 F
16:06:00.777 -> Average Cal sensor Temp.: 98.79 Ambient Temp.: 86.16 F
16:06:02.328 -> Average Cal sensor Temp.: 98.70 Ambient Temp.: 86.13 F
16:06:03.887 -> Average Cal sensor Temp.: 98.48 Ambient Temp.: 86.16 F
16:06:05.440 -> Average Cal sensor Temp.: 98.66 Ambient Temp.: 86.20 F
16:06:06.995 -> Average Cal sensor Temp.: 98.51 Ambient Temp.: 86.13 F
16:06:08.554 -> Average Cal sensor Temp.: 98.61 Ambient Temp.: 86.09 F
16:06:10.110 -> Average Cal sensor Temp.: 98.86 Ambient Temp.: 86.13 F
16:06:11.669 -> Average Cal sensor Temp.: 98.90 Ambient Temp.: 86.16 F
16:06:13.254 -> Average Cal sensor Temp.: 98.72 Ambient Temp.: 86.13 F
16:06:14.808 -> Average Cal sensor Temp.: 98.34 Ambient Temp.: 86.09 F
16:06:16.362 -> Average Cal sensor Temp.: 98.73 Ambient Temp.: 86.13 F
16:06:17.919 -> Average Cal sensor Temp.: 98.82 Ambient Temp.: 86.09 F
16:06:19.479 -> Average Cal sensor Temp.: 98.82 Ambient Temp.: 86.05 F
16:06:21.033 -> Average Cal sensor Temp.: 98.96 Ambient Temp.: 86.05 F
16:06:22.594 -> Average Cal sensor Temp.: 98.88 Ambient Temp.: 86.05 F
16:06:24.147 -> Average Cal sensor Temp.: 98.75 Ambient Temp.: 86.09 F
16:06:25.700 -> Average Cal sensor Temp.: 98.76 Ambient Temp.: 86.09 F
16:06:27.254 -> Average Cal sensor Temp.: 98.96 Ambient Temp.: 86.05 F
16:06:28.808 -> Average Cal sensor Temp.: 98.91 Ambient Temp.: 86.09 F
16:06:30.362 -> Average Cal sensor Temp.: 98.57 Ambient Temp.: 86.05 F
16:06:31.913 -> Average Cal sensor Temp.: 98.62 Ambient Temp.: 85.98 F
16:06:33.471 -> Average Cal sensor Temp.: 98.80 Ambient Temp.: 86.05 F
16:06:35.062 -> Average Cal sensor Temp.: 98.86 Ambient Temp.: 85.98 F
16:06:36.618 -> Average Cal sensor Temp.: 98.90 Ambient Temp.: 86.05 F
16:06:38.170 -> Average Cal sensor Temp.: 98.68 Ambient Temp.: 85.98 F
16:06:39.725 -> Average Cal sensor Temp.: 98.37 Ambient Temp.: 85.98 F
16:06:41.279 -> Average Cal sensor Temp.: 98.57 Ambient Temp.: 85.98 F
16:06:42.837 -> Average Cal sensor Temp.: 99.00 Ambient Temp.: 86.05 F
16:06:44.387 -> Average Cal sensor Temp.: 99.19 Ambient Temp.: 85.95 F
16:06:45.937 -> Average Cal sensor Temp.: 98.66 Ambient Temp.: 86.05 F
16:06:47.523 -> Average Cal sensor Temp.: 98.54 Ambient Temp.: 85.98 F
16:06:49.073 -> Average Cal sensor Temp.: 98.56 Ambient Temp.: 85.98 F
16:06:50.623 -> Average Cal sensor Temp.: 98.59 Ambient Temp.: 86.05 F
16:06:52.170 -> Average Cal sensor Temp.: 98.59 Ambient Temp.: 86.09 F
16:06:53.722 -> Average Cal sensor Temp.: 98.58 Ambient Temp.: 86.05 F
16:06:55.309 -> Average Cal sensor Temp.: 98.69 Ambient Temp.: 86.05 F
16:06:56.857 -> Average Cal sensor Temp.: 98.65 Ambient Temp.: 86.09 F
16:06:58.411 -> Average Cal sensor Temp.: 98.70 Ambient Temp.: 86.09 F
16:06:59.961 -> Average Cal sensor Temp.: 98.89 Ambient Temp.: 86.13 F
16:07:01.543 -> Average Cal sensor Temp.: 98.58 Ambient Temp.: 86.16 F
16:07:03.091 -> Average Cal sensor Temp.: 98.58 Ambient Temp.: 86.13 F
16:07:04.646 -> Average Cal sensor Temp.: 98.70 Ambient Temp.: 86.13 F
16:07:06.201 -> Average Cal sensor Temp.: 98.84 Ambient Temp.: 86.16 F
16:07:07.752 -> Average Cal sensor Temp.: 99.12 Ambient Temp.: 86.20 F
16:07:09.303 -> Average Cal sensor Temp.: 99.22 Ambient Temp.: 86.16 F
16:07:10.858 -> Average Cal sensor Temp.: 99.01 Ambient Temp.: 86.23 F
16:07:12.408 -> Average Cal sensor Temp.: 98.66 Ambient Temp.: 86.20 F
16:07:13.990 -> Average Cal sensor Temp.: 98.76 Ambient Temp.: 86.20 F
16:07:15.546 -> Average Cal sensor Temp.: 98.74 Ambient Temp.: 86.20 F
16:07:17.104 -> Average Cal sensor Temp.: 98.63 Ambient Temp.: 86.27 F
16:07:18.660 -> Average Cal sensor Temp.: 99.02 Ambient Temp.: 86.23 F
16:07:20.213 -> Average Cal sensor Temp.: 99.01 Ambient Temp.: 86.27 F
16:07:21.768 -> Average Cal sensor Temp.: 99.27 Ambient Temp.: 86.23 F
16:07:23.325 -> Average Cal sensor Temp.: 99.18 Ambient Temp.: 86.31 F
16:07:24.875 -> Average Cal sensor Temp.: 99.23 Ambient Temp.: 86.27 F
16:07:26.430 -> Average Cal sensor Temp.: 99.14 Ambient Temp.: 86.27 F
16:07:27.979 -> Average Cal sensor Temp.: 99.03 Ambient Temp.: 86.31 F
16:07:29.535 -> Average Cal sensor Temp.: 99.11 Ambient Temp.: 86.27 F
16:07:31.116 -> Average Cal sensor Temp.: 99.40 Ambient Temp.: 86.31 F
16:07:32.665 -> Average Cal sensor Temp.: 99.14 Ambient Temp.: 86.31 F
16:07:34.213 -> Average Cal sensor Temp.: 98.82 Ambient Temp.: 86.31 F
16:07:35.770 -> Average Cal sensor Temp.: 98.92 Ambient Temp.: 86.31 F
16:07:37.328 -> Average Cal sensor Temp.: 98.81 Ambient Temp.: 86.31 F
16:07:38.883 -> Average Cal sensor Temp.: 99.24 Ambient Temp.: 86.31 F
16:07:40.437 -> Average Cal sensor Temp.: 99.51 Ambient Temp.: 86.31 F
16:07:41.993 -> Average Cal sensor Temp.: 99.03 Ambient Temp.: 86.27 F
16:07:43.546 -> Average Cal sensor Temp.: 98.78 Ambient Temp.: 86.34 F
16:07:45.100 -> Average Cal sensor Temp.: 99.06 Ambient Temp.: 86.34 F
16:07:46.691 -> Average Cal sensor Temp.: 99.04 Ambient Temp.: 86.31 F
16:07:48.240 -> Average Cal sensor Temp.: 99.15 Ambient Temp.: 86.38 F
16:07:49.788 -> Average Cal sensor Temp.: 99.14 Ambient Temp.: 86.38 F
16:07:51.335 -> Average Cal sensor Temp.: 99.05 Ambient Temp.: 86.38 F
16:07:52.919 -> Average Cal sensor Temp.: 99.26 Ambient Temp.: 86.34 F
16:07:54.443 -> Average Cal sensor Temp.: 98.96 Ambient Temp.: 86.38 F
16:07:56.028 -> Average Cal sensor Temp.: 99.28 Ambient Temp.: 86.34 F
16:07:57.591 -> Average Cal sensor Temp.: 99.52 Ambient Temp.: 86.38 F
16:07:59.143 -> Average Cal sensor Temp.: 99.51 Ambient Temp.: 86.45 F
16:08:00.701 -> Average Cal sensor Temp.: 99.28 Ambient Temp.: 86.45 F
16:08:02.252 -> Average Cal sensor Temp.: 98.84 Ambient Temp.: 86.38 F
16:08:03.800 -> Average Cal sensor Temp.: 99.18 Ambient Temp.: 86.38 F
16:08:05.356 -> Average Cal sensor Temp.: 98.99 Ambient Temp.: 86.45 F
16:08:06.947 -> Average Cal sensor Temp.: 99.18 Ambient Temp.: 86.45 F
16:08:08.470 -> Average Cal sensor Temp.: 99.43 Ambient Temp.: 86.45 F
16:08:10.055 -> Average Cal sensor Temp.: 99.11 Ambient Temp.: 86.49 F
16:08:11.611 -> Average Cal sensor Temp.: 98.91 Ambient Temp.: 86.45 F
16:08:13.165 -> Average Cal sensor Temp.: 99.08 Ambient Temp.: 86.45 F
16:08:14.718 -> Average Cal sensor Temp.: 99.42 Ambient Temp.: 86.49 F
16:08:16.273 -> Average Cal sensor Temp.: 99.13 Ambient Temp.: 86.52 F
16:08:17.831 -> Average Cal sensor Temp.: 99.15 Ambient Temp.: 86.49 F
16:08:19.394 -> Average Cal sensor Temp.: 99.08 Ambient Temp.: 86.56 F
16:08:20.946 -> Average Cal sensor Temp.: 99.55 Ambient Temp.: 86.52 F
16:08:22.503 -> Average Cal sensor Temp.: 99.13 Ambient Temp.: 86.52 F
16:08:24.053 -> Average Cal sensor Temp.: 99.16 Ambient Temp.: 86.59 F
16:08:25.603 -> Average Cal sensor Temp.: 99.00 Ambient Temp.: 86.56 F
16:08:27.191 -> Average Cal sensor Temp.: 99.27 Ambient Temp.: 86.67 F
16:08:28.745 -> Average Cal sensor Temp.: 99.22 Ambient Temp.: 86.63 F
16:08:30.299 -> Average Cal sensor Temp.: 99.16 Ambient Temp.: 86.67 F
16:08:31.854 -> Average Cal sensor Temp.: 99.14 Ambient Temp.: 86.59 F
16:08:33.409 -> Average Cal sensor Temp.: 99.00 Ambient Temp.: 86.59 F
16:08:34.965 -> Average Cal sensor Temp.: 99.16 Ambient Temp.: 86.67 F
16:08:36.519 -> Average Cal sensor Temp.: 99.11 Ambient Temp.: 86.63 F
16:08:38.078 -> Average Cal sensor Temp.: 99.67 Ambient Temp.: 86.67 F
16:08:39.639 -> Average Cal sensor Temp.: 99.81 Ambient Temp.: 86.67 F
16:08:41.194 -> Average Cal sensor Temp.: 99.71 Ambient Temp.: 86.70 F
16:08:42.748 -> Average Cal sensor Temp.: 99.92 Ambient Temp.: 86.77 F
16:08:44.338 -> Average Cal sensor Temp.: 99.38 Ambient Temp.: 86.74 F
16:08:45.886 -> Average Cal sensor Temp.: 99.56 Ambient Temp.: 86.77 F
16:08:47.434 -> Average Cal sensor Temp.: 99.34 Ambient Temp.: 86.77 F
16:08:48.992 -> Average Cal sensor Temp.: 99.38 Ambient Temp.: 86.77 F
16:08:50.545 -> Average Cal sensor Temp.: 99.23 Ambient Temp.: 86.81 F
16:08:52.098 -> Average Cal sensor Temp.: 99.51 Ambient Temp.: 86.88 F
16:08:53.689 -> Average Cal sensor Temp.: 99.40 Ambient Temp.: 86.88 F
16:08:55.238 -> Average Cal sensor Temp.: 99.44 Ambient Temp.: 86.92 F
16:08:56.793 -> Average Cal sensor Temp.: 99.56 Ambient Temp.: 86.92 F
16:08:58.340 -> Average Cal sensor Temp.: 99.40 Ambient Temp.: 86.99 F
16:08:59.923 -> Average Cal sensor Temp.: 99.64 Ambient Temp.: 87.03 F
16:09:01.472 -> Average Cal sensor Temp.: 99.43 Ambient Temp.: 87.03 F
16:09:03.028 -> Average Cal sensor Temp.: 99.18 Ambient Temp.: 87.06 F
16:09:04.575 -> Average Cal sensor Temp.: 99.44 Ambient Temp.: 87.03 F
16:09:06.129 -> Average Cal sensor Temp.: 99.21 Ambient Temp.: 87.10 F
16:09:07.710 -> Average Cal sensor Temp.: 99.24 Ambient Temp.: 87.10 F
16:09:09.256 -> Average Cal sensor Temp.: 99.22 Ambient Temp.: 87.13 F
16:09:10.810 -> Average Cal sensor Temp.: 99.27 Ambient Temp.: 87.10 F
16:09:12.365 -> Average Cal sensor Temp.: 99.46 Ambient Temp.: 87.13 F
16:09:13.911 -> Average Cal sensor Temp.: 99.38 Ambient Temp.: 87.17 F
16:09:15.471 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 87.17 F
16:09:17.029 -> Average Cal sensor Temp.: 99.49 Ambient Temp.: 87.17 F
16:09:18.585 -> Average Cal sensor Temp.: 99.62 Ambient Temp.: 87.21 F
16:09:20.170 -> Average Cal sensor Temp.: 99.53 Ambient Temp.: 87.21 F
16:09:21.726 -> Average Cal sensor Temp.: 99.42 Ambient Temp.: 87.21 F
16:09:23.281 -> Average Cal sensor Temp.: 99.48 Ambient Temp.: 87.28 F
16:09:24.838 -> Average Cal sensor Temp.: 99.20 Ambient Temp.: 87.28 F
16:09:26.396 -> Average Cal sensor Temp.: 99.50 Ambient Temp.: 87.28 F
16:09:27.946 -> Average Cal sensor Temp.: 99.45 Ambient Temp.: 87.31 F
16:09:29.501 -> Average Cal sensor Temp.: 99.33 Ambient Temp.: 87.31 F
16:09:31.052 -> Average Cal sensor Temp.: 99.24 Ambient Temp.: 87.28 F
16:09:32.613 -> Average Cal sensor Temp.: 99.07 Ambient Temp.: 87.31 F
16:09:34.171 -> Average Cal sensor Temp.: 99.17 Ambient Temp.: 87.35 F
16:09:35.726 -> Average Cal sensor Temp.: 99.00 Ambient Temp.: 87.28 F
16:09:37.279 -> Average Cal sensor Temp.: 98.92 Ambient Temp.: 87.31 F
16:09:38.834 -> Average Cal sensor Temp.: 98.87 Ambient Temp.: 87.31 F
16:09:40.389 -> Average Cal sensor Temp.: 98.87 Ambient Temp.: 87.35 F
16:09:41.940 -> Average Cal sensor Temp.: 98.57 Ambient Temp.: 87.31 F
16:09:43.497 -> Average Cal sensor Temp.: 98.78 Ambient Temp.: 87.35 F
16:09:45.073 -> Average Cal sensor Temp.: 99.98 Ambient Temp.: 87.31 F
16:09:46.631 -> Average Cal sensor Temp.: 99.15 Ambient Temp.: 87.35 F
16:09:48.184 -> Average Cal sensor Temp.: 99.29 Ambient Temp.: 87.35 F
16:09:49.733 -> Average Cal sensor Temp.: 99.08 Ambient Temp.: 87.35 F
16:09:51.282 -> Average Cal sensor Temp.: 99.29 Ambient Temp.: 87.35 F
16:09:52.836 -> Average Cal sensor Temp.: 98.96 Ambient Temp.: 87.31 F
16:09:54.393 -> Average Cal sensor Temp.: 99.27 Ambient Temp.: 87.35 F
16:09:55.948 -> Average Cal sensor Temp.: 99.31 Ambient Temp.: 87.39 F
16:09:57.505 -> Average Cal sensor Temp.: 98.74 Ambient Temp.: 87.31 F
16:09:59.061 -> Average Cal sensor Temp.: 98.83 Ambient Temp.: 87.35 F
16:10:00.611 -> Average Cal sensor Temp.: 99.03 Ambient Temp.: 87.35 F
16:10:02.165 -> Average Cal sensor Temp.: 98.83 Ambient Temp.: 87.31 F
16:10:03.753 -> Average Cal sensor Temp.: 97.78 Ambient Temp.: 87.28 F
16:10:05.301 -> Average Cal sensor Temp.: 98.39 Ambient Temp.: 87.21 F
16:10:06.850 -> Average Cal sensor Temp.: 98.55 Ambient Temp.: 87.28 F
16:10:08.404 -> Average Cal sensor Temp.: 98.60 Ambient Temp.: 87.28 F
16:10:09.964 -> Average Cal sensor Temp.: 99.05 Ambient Temp.: 87.21 F
16:10:11.515 -> Average Cal sensor Temp.: 99.43 Ambient Temp.: 87.17 F
16:10:13.071 -> Average Cal sensor Temp.: 99.29 Ambient Temp.: 87.21 F
16:10:14.627 -> Average Cal sensor Temp.: 98.75 Ambient Temp.: 87.17 F
16:10:16.181 -> Average Cal sensor Temp.: 98.63 Ambient Temp.: 87.13 F
16:10:17.735 -> Average Cal sensor Temp.: 98.54 Ambient Temp.: 87.13 F
16:10:19.288 -> Average Cal sensor Temp.: 98.66 Ambient Temp.: 87.10 F
16:10:20.841 -> Average Cal sensor Temp.: 98.83 Ambient Temp.: 87.17 F
16:10:22.398 -> Average Cal sensor Temp.: 98.87 Ambient Temp.: 87.10 F
16:10:23.950 -> Average Cal sensor Temp.: 98.79 Ambient Temp.: 87.06 F
16:10:25.533 -> Average Cal sensor Temp.: 99.16 Ambient Temp.: 87.10 F
16:10:27.079 -> Average Cal sensor Temp.: 99.33 Ambient Temp.: 86.99 F
16:10:28.633 -> Average Cal sensor Temp.: 99.37 Ambient Temp.: 87.03 F
16:10:30.182 -> Average Cal sensor Temp.: 98.87 Ambient Temp.: 87.03 F
16:10:31.736 -> Average Cal sensor Temp.: 98.97 Ambient Temp.: 87.03 F
16:10:33.317 -> Average Cal sensor Temp.: 99.03 Ambient Temp.: 87.03 F
16:10:34.870 -> Average Cal sensor Temp.: 99.12 Ambient Temp.: 86.99 F
16:10:36.422 -> Average Cal sensor Temp.: 99.00 Ambient Temp.: 86.99 F
16:10:37.971 -> Average Cal sensor Temp.: 98.90 Ambient Temp.: 86.95 F
16:10:39.531 -> Average Cal sensor Temp.: 98.93 Ambient Temp.: 86.95 F
16:10:41.087 -> Average Cal sensor Temp.: 98.96 Ambient Temp.: 86.92 F
16:10:42.642 -> Average Cal sensor Temp.: 98.89 Ambient Temp.: 86.95 F
16:10:44.195 -> Average Cal sensor Temp.: 99.06 Ambient Temp.: 86.92 F
16:10:45.755 -> Average Cal sensor Temp.: 99.49 Ambient Temp.: 86.92 F
16:10:47.338 -> Average Cal sensor Temp.: 99.75 Ambient Temp.: 86.99 F
16:10:48.890 -> Average Cal sensor Temp.: 99.50 Ambient Temp.: 86.99 F
16:10:50.443 -> Average Cal sensor Temp.: 99.23 Ambient Temp.: 86.95 F
16:10:52.023 -> Average Cal sensor Temp.: 99.33 Ambient Temp.: 86.92 F
16:10:53.564 -> Average Cal sensor Temp.: 99.40 Ambient Temp.: 86.92 F
16:10:55.112 -> Average Cal sensor Temp.: 98.81 Ambient Temp.: 86.88 F
16:10:56.673 -> Average Cal sensor Temp.: 99.08 Ambient Temp.: 86.92 F
16:10:58.228 -> Average Cal sensor Temp.: 99.41 Ambient Temp.: 86.88 F
16:10:59.779 -> Average Cal sensor Temp.: 98.92 Ambient Temp.: 86.92 F
16:11:01.336 -> Average Cal sensor Temp.: 98.82 Ambient Temp.: 86.77 F
16:11:02.916 -> Average Cal sensor Temp.: 99.23 Ambient Temp.: 86.81 F
16:11:04.467 -> Average Cal sensor Temp.: 99.09 Ambient Temp.: 86.88 F
16:11:06.029 -> Average Cal sensor Temp.: 98.70 Ambient Temp.: 86.77 F
16:11:07.582 -> Average Cal sensor Temp.: 99.10 Ambient Temp.: 86.77 F
16:11:09.138 -> Average Cal sensor Temp.: 99.04 Ambient Temp.: 86.74 F
16:11:10.695 -> Average Cal sensor Temp.: 98.85 Ambient Temp.: 86.74 F
16:11:12.251 -> Average Cal sensor Temp.: 98.96 Ambient Temp.: 86.70 F
16:11:13.809 -> Average Cal sensor Temp.: 99.16 Ambient Temp.: 86.74 F
16:11:15.360 -> Average Cal sensor Temp.: 99.08 Ambient Temp.: 86.70 F
16:11:16.914 -> Average Cal sensor Temp.: 98.61 Ambient Temp.: 86.70 F
16:11:18.499 -> Average Cal sensor Temp.: 98.53 Ambient Temp.: 86.63 F
16:11:20.050 -> Average Cal sensor Temp.: 99.08 Ambient Temp.: 86.67 F
16:11:21.605 -> Average Cal sensor Temp.: 99.20 Ambient Temp.: 86.67 F
16:11:23.159 -> Average Cal sensor Temp.: 99.18 Ambient Temp.: 86.67 F
16:11:24.712 -> Average Cal sensor Temp.: 99.00 Ambient Temp.: 86.67 F
16:11:26.265 -> Average Cal sensor Temp.: 99.32 Ambient Temp.: 86.63 F
16:11:27.820 -> Average Cal sensor Temp.: 99.14 Ambient Temp.: 86.63 F
16:11:29.371 -> Average Cal sensor Temp.: 99.13 Ambient Temp.: 86.63 F
16:11:30.925 -> Average Cal sensor Temp.: 98.97 Ambient Temp.: 86.63 F
16:11:32.515 -> Average Cal sensor Temp.: 98.86 Ambient Temp.: 86.63 F
16:11:34.067 -> Average Cal sensor Temp.: 99.13 Ambient Temp.: 86.63 F
16:11:35.619 -> Average Cal sensor Temp.: 99.12 Ambient Temp.: 86.67 F
16:11:37.172 -> Average Cal sensor Temp.: 98.65 Ambient Temp.: 86.59 F
16:11:38.726 -> Average Cal sensor Temp.: 98.78 Ambient Temp.: 86.56 F
16:11:40.278 -> Average Cal sensor Temp.: 99.08 Ambient Temp.: 86.63 F
16:11:41.835 -> Average Cal sensor Temp.: 98.96 Ambient Temp.: 86.52 F
16:11:43.393 -> Average Cal sensor Temp.: 98.53 Ambient Temp.: 86.56 F
16:11:44.948 -> Average Cal sensor Temp.: 98.91 Ambient Temp.: 86.52 F
16:11:46.531 -> Average Cal sensor Temp.: 98.68 Ambient Temp.: 86.56 F
16:11:48.083 -> Average Cal sensor Temp.: 98.99 Ambient Temp.: 86.45 F
16:11:49.635 -> Average Cal sensor Temp.: 99.19 Ambient Temp.: 86.52 F
16:11:51.192 -> Average Cal sensor Temp.: 99.04 Ambient Temp.: 86.49 F
16:11:52.747 -> Average Cal sensor Temp.: 98.69 Ambient Temp.: 86.38 F
16:11:54.303 -> Average Cal sensor Temp.: 98.97 Ambient Temp.: 86.45 F
16:11:55.863 -> Average Cal sensor Temp.: 98.90 Ambient Temp.: 86.49 F
16:11:57.419 -> Average Cal sensor Temp.: 98.22 Ambient Temp.: 86.45 F
16:11:58.970 -> Average Cal sensor Temp.: 97.93 Ambient Temp.: 86.34 F
16:12:00.523 -> Average Cal sensor Temp.: 98.25 Ambient Temp.: 86.31 F
16:12:02.075 -> Average Cal sensor Temp.: 98.38 Ambient Temp.: 86.23 F
16:12:03.636 -> Average Cal sensor Temp.: 98.68 Ambient Temp.: 86.23 F
16:12:05.192 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 86.23 F
16:12:06.749 -> Average Cal sensor Temp.: 98.22 Ambient Temp.: 86.16 F
16:12:08.305 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 86.13 F
16:12:09.890 -> Average Cal sensor Temp.: 98.48 Ambient Temp.: 86.16 F
16:12:11.438 -> Average Cal sensor Temp.: 97.62 Ambient Temp.: 86.09 F
16:12:12.990 -> Average Cal sensor Temp.: 98.05 Ambient Temp.: 86.05 F
16:12:14.542 -> Average Cal sensor Temp.: 98.58 Ambient Temp.: 85.95 F
16:12:16.100 -> Average Cal sensor Temp.: 98.85 Ambient Temp.: 85.98 F
16:12:17.655 -> Average Cal sensor Temp.: 99.02 Ambient Temp.: 85.98 F
16:12:19.218 -> Average Cal sensor Temp.: 98.83 Ambient Temp.: 85.91 F
16:12:20.772 -> Average Cal sensor Temp.: 98.84 Ambient Temp.: 85.95 F
16:12:22.348 -> Average Cal sensor Temp.: 98.91 Ambient Temp.: 85.95 F
16:12:23.899 -> Average Cal sensor Temp.: 99.02 Ambient Temp.: 85.91 F
16:12:25.450 -> Average Cal sensor Temp.: 98.99 Ambient Temp.: 85.87 F
16:12:27.004 -> Average Cal sensor Temp.: 98.69 Ambient Temp.: 85.91 F
16:12:28.557 -> Average Cal sensor Temp.: 98.82 Ambient Temp.: 85.87 F
16:12:30.112 -> Average Cal sensor Temp.: 99.19 Ambient Temp.: 85.91 F
16:12:31.669 -> Average Cal sensor Temp.: 99.32 Ambient Temp.: 85.84 F
16:12:33.218 -> Average Cal sensor Temp.: 99.50 Ambient Temp.: 85.87 F
16:12:34.800 -> Average Cal sensor Temp.: 98.72 Ambient Temp.: 85.87 F
16:12:36.356 -> Average Cal sensor Temp.: 98.76 Ambient Temp.: 85.91 F
16:12:37.902 -> Average Cal sensor Temp.: 99.15 Ambient Temp.: 85.84 F
16:12:39.449 -> Average Cal sensor Temp.: 98.69 Ambient Temp.: 85.84 F
16:12:41.005 -> Average Cal sensor Temp.: 98.77 Ambient Temp.: 85.84 F
16:12:42.563 -> Average Cal sensor Temp.: 99.07 Ambient Temp.: 85.77 F
16:12:44.115 -> Average Cal sensor Temp.: 99.10 Ambient Temp.: 85.80 F
16:12:45.669 -> Average Cal sensor Temp.: 99.24 Ambient Temp.: 85.80 F
16:12:47.225 -> Average Cal sensor Temp.: 99.11 Ambient Temp.: 85.80 F
16:12:48.781 -> Average Cal sensor Temp.: 99.09 Ambient Temp.: 85.84 F
16:12:50.364 -> Average Cal sensor Temp.: 99.15 Ambient Temp.: 85.80 F
16:12:51.918 -> Average Cal sensor Temp.: 99.29 Ambient Temp.: 85.84 F
16:12:53.470 -> Average Cal sensor Temp.: 99.62 Ambient Temp.: 85.87 F
16:12:55.023 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 85.87 F
16:12:56.580 -> Average Cal sensor Temp.: 100.31 Ambient Temp.: 85.91 F
16:12:58.169 -> Average Cal sensor Temp.: 99.81 Ambient Temp.: 85.95 F
16:12:59.721 -> Average Cal sensor Temp.: 99.45 Ambient Temp.: 85.84 F
16:13:01.273 -> Average Cal sensor Temp.: 99.58 Ambient Temp.: 85.91 F
16:13:02.825 -> Average Cal sensor Temp.: 99.29 Ambient Temp.: 85.91 F
16:13:04.410 -> Average Cal sensor Temp.: 99.02 Ambient Temp.: 85.91 F
16:13:05.963 -> Average Cal sensor Temp.: 99.40 Ambient Temp.: 85.87 F
16:13:07.520 -> Average Cal sensor Temp.: 99.12 Ambient Temp.: 85.95 F
16:13:09.076 -> Average Cal sensor Temp.: 99.18 Ambient Temp.: 85.91 F
16:13:10.626 -> Average Cal sensor Temp.: 99.32 Ambient Temp.: 85.91 F
16:13:12.179 -> Average Cal sensor Temp.: 99.03 Ambient Temp.: 85.95 F
16:13:13.728 -> Average Cal sensor Temp.: 99.39 Ambient Temp.: 85.91 F
16:13:15.284 -> Average Cal sensor Temp.: 99.06 Ambient Temp.: 85.91 F
16:13:16.837 -> Average Cal sensor Temp.: 99.26 Ambient Temp.: 85.91 F
16:13:18.389 -> Average Cal sensor Temp.: 99.49 Ambient Temp.: 85.95 F
16:13:19.944 -> Average Cal sensor Temp.: 99.45 Ambient Temp.: 85.95 F
16:13:21.504 -> Average Cal sensor Temp.: 99.45 Ambient Temp.: 86.05 F
16:13:23.058 -> Average Cal sensor Temp.: 99.33 Ambient Temp.: 85.98 F
16:13:24.642 -> Average Cal sensor Temp.: 99.59 Ambient Temp.: 86.05 F
16:13:26.198 -> Average Cal sensor Temp.: 99.51 Ambient Temp.: 86.09 F
16:13:27.755 -> Average Cal sensor Temp.: 99.38 Ambient Temp.: 86.09 F
16:13:29.313 -> Average Cal sensor Temp.: 99.40 Ambient Temp.: 86.13 F
16:13:30.872 -> Average Cal sensor Temp.: 99.30 Ambient Temp.: 86.13 F
16:13:32.423 -> Average Cal sensor Temp.: 99.53 Ambient Temp.: 86.16 F
16:13:33.976 -> Average Cal sensor Temp.: 99.15 Ambient Temp.: 86.16 F
16:13:35.532 -> Average Cal sensor Temp.: 99.46 Ambient Temp.: 86.13 F
16:13:37.089 -> Average Cal sensor Temp.: 99.37 Ambient Temp.: 86.13 F
16:13:38.671 -> Average Cal sensor Temp.: 99.24 Ambient Temp.: 86.16 F
16:13:40.230 -> Average Cal sensor Temp.: 99.38 Ambient Temp.: 86.20 F
16:13:41.785 -> Average Cal sensor Temp.: 99.23 Ambient Temp.: 86.20 F
16:13:43.341 -> Average Cal sensor Temp.: 99.22 Ambient Temp.: 86.16 F
16:13:44.890 -> Average Cal sensor Temp.: 99.60 Ambient Temp.: 86.23 F
16:13:46.446 -> Average Cal sensor Temp.: 99.65 Ambient Temp.: 86.27 F
16:13:47.999 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 86.27 F
16:13:49.557 -> Average Cal sensor Temp.: 99.57 Ambient Temp.: 86.27 F
16:13:51.145 -> Average Cal sensor Temp.: 99.66 Ambient Temp.: 86.31 F
16:13:52.705 -> Average Cal sensor Temp.: 99.43 Ambient Temp.: 86.27 F
16:13:54.258 -> Average Cal sensor Temp.: 99.36 Ambient Temp.: 86.31 F
16:13:55.807 -> Average Cal sensor Temp.: 99.79 Ambient Temp.: 86.27 F
16:13:57.370 -> Average Cal sensor Temp.: 99.68 Ambient Temp.: 86.34 F
16:13:58.925 -> Average Cal sensor Temp.: 99.69 Ambient Temp.: 86.34 F
16:14:00.510 -> Average Cal sensor Temp.: 99.55 Ambient Temp.: 86.38 F
16:14:02.061 -> Average Cal sensor Temp.: 99.49 Ambient Temp.: 86.34 F
16:14:03.615 -> Average Cal sensor Temp.: 99.69 Ambient Temp.: 86.34 F
16:14:05.195 -> Average Cal sensor Temp.: 99.29 Ambient Temp.: 86.38 F
16:14:06.744 -> Average Cal sensor Temp.: 99.32 Ambient Temp.: 86.45 F
16:14:08.300 -> Average Cal sensor Temp.: 99.39 Ambient Temp.: 86.45 F
16:14:09.852 -> Average Cal sensor Temp.: 99.67 Ambient Temp.: 86.45 F
16:14:11.404 -> Average Cal sensor Temp.: 99.56 Ambient Temp.: 86.49 F
16:14:12.993 -> Average Cal sensor Temp.: 99.48 Ambient Temp.: 86.56 F
16:14:14.539 -> Average Cal sensor Temp.: 99.73 Ambient Temp.: 86.49 F
16:14:16.090 -> Average Cal sensor Temp.: 99.77 Ambient Temp.: 86.52 F
16:14:17.677 -> Average Cal sensor Temp.: 99.56 Ambient Temp.: 86.52 F
16:14:19.233 -> Average Cal sensor Temp.: 99.55 Ambient Temp.: 86.56 F
16:14:20.785 -> Average Cal sensor Temp.: 99.85 Ambient Temp.: 86.52 F
16:14:22.338 -> Average Cal sensor Temp.: 99.60 Ambient Temp.: 86.59 F
16:14:23.924 -> Average Cal sensor Temp.: 99.58 Ambient Temp.: 86.59 F
16:14:25.479 -> Average Cal sensor Temp.: 99.62 Ambient Temp.: 86.59 F
16:14:27.035 -> Average Cal sensor Temp.: 99.16 Ambient Temp.: 86.59 F
16:14:28.580 -> Average Cal sensor Temp.: 99.27 Ambient Temp.: 86.63 F
16:14:30.134 -> Average Cal sensor Temp.: 99.19 Ambient Temp.: 86.63 F
16:14:31.694 -> Average Cal sensor Temp.: 99.38 Ambient Temp.: 86.67 F
16:14:33.280 -> Average Cal sensor Temp.: 99.49 Ambient Temp.: 86.67 F
16:14:34.806 -> Average Cal sensor Temp.: 99.54 Ambient Temp.: 86.63 F
16:14:36.390 -> Average Cal sensor Temp.: 99.59 Ambient Temp.: 86.70 F
16:14:37.946 -> Average Cal sensor Temp.: 99.59 Ambient Temp.: 86.70 F
16:14:39.504 -> Average Cal sensor Temp.: 99.66 Ambient Temp.: 86.70 F
16:14:41.058 -> Average Cal sensor Temp.: 99.77 Ambient Temp.: 86.74 F
16:14:42.650 -> Average Cal sensor Temp.: 99.71 Ambient Temp.: 86.77 F
16:14:44.198 -> Average Cal sensor Temp.: 99.68 Ambient Temp.: 86.77 F
16:14:45.755 -> Average Cal sensor Temp.: 99.82 Ambient Temp.: 86.77 F
16:14:47.307 -> Average Cal sensor Temp.: 99.80 Ambient Temp.: 86.92 F
16:14:48.895 -> Average Cal sensor Temp.: 99.34 Ambient Temp.: 86.92 F
16:14:50.452 -> Average Cal sensor Temp.: 99.20 Ambient Temp.: 86.92 F
16:14:52.003 -> Average Cal sensor Temp.: 99.41 Ambient Temp.: 86.92 F
16:14:53.562 -> Average Cal sensor Temp.: 99.71 Ambient Temp.: 86.95 F
16:14:55.113 -> Average Cal sensor Temp.: 99.56 Ambient Temp.: 86.95 F
16:14:56.677 -> Average Cal sensor Temp.: 99.61 Ambient Temp.: 87.06 F
16:14:58.236 -> Average Cal sensor Temp.: 99.55 Ambient Temp.: 87.03 F
16:14:59.791 -> Average Cal sensor Temp.: 99.72 Ambient Temp.: 87.06 F
16:15:01.349 -> Average Cal sensor Temp.: 99.80 Ambient Temp.: 87.10 F
16:15:02.936 -> Average Cal sensor Temp.: 99.96 Ambient Temp.: 87.10 F
16:15:04.491 -> Average Cal sensor Temp.: 99.97 Ambient Temp.: 87.21 F
16:15:06.047 -> Average Cal sensor Temp.: 99.95 Ambient Temp.: 87.17 F
16:15:07.607 -> Average Cal sensor Temp.: 99.94 Ambient Temp.: 87.28 F
16:15:09.189 -> Average Cal sensor Temp.: 99.69 Ambient Temp.: 87.31 F
16:15:10.743 -> Average Cal sensor Temp.: 100.04 Ambient Temp.: 87.31 F
16:15:12.297 -> Average Cal sensor Temp.: 100.07 Ambient Temp.: 87.31 F
16:15:13.850 -> Average Cal sensor Temp.: 99.78 Ambient Temp.: 87.35 F
16:15:15.437 -> Average Cal sensor Temp.: 99.95 Ambient Temp.: 87.39 F
16:15:16.994 -> Average Cal sensor Temp.: 99.87 Ambient Temp.: 87.39 F
16:15:18.548 -> Average Cal sensor Temp.: 99.98 Ambient Temp.: 87.46 F
16:15:20.132 -> Average Cal sensor Temp.: 99.72 Ambient Temp.: 87.42 F
16:15:21.688 -> Average Cal sensor Temp.: 99.72 Ambient Temp.: 87.49 F
16:15:23.242 -> Average Cal sensor Temp.: 99.41 Ambient Temp.: 87.46 F
16:15:24.793 -> Average Cal sensor Temp.: 99.68 Ambient Temp.: 87.49 F
16:15:26.346 -> Average Cal sensor Temp.: 99.72 Ambient Temp.: 87.57 F
16:15:27.935 -> Average Cal sensor Temp.: 99.67 Ambient Temp.: 87.49 F
16:15:29.491 -> Average Cal sensor Temp.: 99.65 Ambient Temp.: 87.57 F
16:15:31.040 -> Average Cal sensor Temp.: 99.59 Ambient Temp.: 87.53 F
16:15:32.589 -> Average Cal sensor Temp.: 99.18 Ambient Temp.: 87.57 F
16:15:34.143 -> Average Cal sensor Temp.: 99.46 Ambient Temp.: 87.60 F
16:15:35.733 -> Average Cal sensor Temp.: 99.55 Ambient Temp.: 87.60 F
16:15:37.285 -> Average Cal sensor Temp.: 99.76 Ambient Temp.: 87.60 F
16:15:38.836 -> Average Cal sensor Temp.: 99.60 Ambient Temp.: 87.67 F
16:15:40.421 -> Average Cal sensor Temp.: 99.34 Ambient Temp.: 87.60 F
16:15:41.973 -> Average Cal sensor Temp.: 99.46 Ambient Temp.: 87.67 F
16:15:43.524 -> Average Cal sensor Temp.: 99.38 Ambient Temp.: 87.60 F
16:15:45.079 -> Average Cal sensor Temp.: 99.53 Ambient Temp.: 87.67 F
16:15:46.631 -> Average Cal sensor Temp.: 99.64 Ambient Temp.: 87.71 F
16:15:48.185 -> Average Cal sensor Temp.: 99.82 Ambient Temp.: 87.67 F
16:15:49.771 -> Average Cal sensor Temp.: 99.49 Ambient Temp.: 87.60 F
16:15:51.327 -> Average Cal sensor Temp.: 99.87 Ambient Temp.: 87.71 F
16:15:52.884 -> Average Cal sensor Temp.: 99.65 Ambient Temp.: 87.67 F
16:15:54.438 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 87.60 F
16:15:55.994 -> Average Cal sensor Temp.: 99.54 Ambient Temp.: 87.71 F
16:15:57.582 -> Average Cal sensor Temp.: 99.99 Ambient Temp.: 87.71 F
16:15:59.132 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 87.71 F
16:16:00.689 -> Average Cal sensor Temp.: 100.14 Ambient Temp.: 87.75 F
16:16:02.242 -> Average Cal sensor Temp.: 99.86 Ambient Temp.: 87.78 F
16:16:03.821 -> Average Cal sensor Temp.: 99.68 Ambient Temp.: 87.75 F
16:16:05.374 -> Average Cal sensor Temp.: 99.54 Ambient Temp.: 87.78 F
16:16:06.925 -> Average Cal sensor Temp.: 99.44 Ambient Temp.: 87.78 F
16:16:08.515 -> Average Cal sensor Temp.: 99.67 Ambient Temp.: 87.82 F
16:16:10.069 -> Average Cal sensor Temp.: 99.50 Ambient Temp.: 87.75 F
16:16:11.623 -> Average Cal sensor Temp.: 99.84 Ambient Temp.: 87.78 F
16:16:13.173 -> Average Cal sensor Temp.: 99.80 Ambient Temp.: 87.75 F
16:16:14.762 -> Average Cal sensor Temp.: 99.77 Ambient Temp.: 87.78 F
16:16:16.319 -> Average Cal sensor Temp.: 100.03 Ambient Temp.: 87.85 F
16:16:17.869 -> Average Cal sensor Temp.: 99.94 Ambient Temp.: 87.85 F
16:16:19.422 -> Average Cal sensor Temp.: 99.59 Ambient Temp.: 87.82 F
16:16:21.010 -> Average Cal sensor Temp.: 99.45 Ambient Temp.: 87.82 F
16:16:22.565 -> Average Cal sensor Temp.: 99.64 Ambient Temp.: 87.85 F
16:16:24.120 -> Average Cal sensor Temp.: 99.20 Ambient Temp.: 87.78 F
16:16:25.675 -> Average Cal sensor Temp.: 99.30 Ambient Temp.: 87.75 F
16:16:27.228 -> Average Cal sensor Temp.: 99.59 Ambient Temp.: 87.75 F
16:16:28.785 -> Average Cal sensor Temp.: 99.73 Ambient Temp.: 87.78 F
16:16:30.337 -> Average Cal sensor Temp.: 99.72 Ambient Temp.: 87.78 F
16:16:31.927 -> Average Cal sensor Temp.: 99.66 Ambient Temp.: 87.82 F
16:16:33.474 -> Average Cal sensor Temp.: 99.49 Ambient Temp.: 87.75 F
16:16:35.025 -> Average Cal sensor Temp.: 99.71 Ambient Temp.: 87.82 F
16:16:36.610 -> Average Cal sensor Temp.: 99.60 Ambient Temp.: 87.78 F
16:16:38.159 -> Average Cal sensor Temp.: 99.39 Ambient Temp.: 87.75 F
16:16:39.712 -> Average Cal sensor Temp.: 99.99 Ambient Temp.: 87.78 F
16:16:41.264 -> Average Cal sensor Temp.: 99.80 Ambient Temp.: 87.75 F
16:16:42.851 -> Average Cal sensor Temp.: 99.67 Ambient Temp.: 87.75 F
16:16:44.405 -> Average Cal sensor Temp.: 99.55 Ambient Temp.: 87.71 F
16:16:45.950 -> Average Cal sensor Temp.: 99.84 Ambient Temp.: 87.75 F
16:16:47.537 -> Average Cal sensor Temp.: 99.74 Ambient Temp.: 87.75 F
16:16:49.084 -> Average Cal sensor Temp.: 99.58 Ambient Temp.: 87.75 F
16:16:50.631 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 87.75 F
16:16:52.216 -> Average Cal sensor Temp.: 99.43 Ambient Temp.: 87.71 F
16:16:53.773 -> Average Cal sensor Temp.: 99.13 Ambient Temp.: 87.71 F
16:16:55.323 -> Average Cal sensor Temp.: 99.56 Ambient Temp.: 87.67 F
16:16:56.877 -> Average Cal sensor Temp.: 99.46 Ambient Temp.: 87.67 F
16:16:58.420 -> Average Cal sensor Temp.: 99.41 Ambient Temp.: 87.67 F
16:17:00.002 -> Average Cal sensor Temp.: 99.45 Ambient Temp.: 87.67 F
16:17:01.556 -> Average Cal sensor Temp.: 98.85 Ambient Temp.: 87.57 F
16:17:03.114 -> Average Cal sensor Temp.: 98.48 Ambient Temp.: 87.53 F
16:17:04.669 -> Average Cal sensor Temp.: 98.54 Ambient Temp.: 87.49 F
16:17:06.225 -> Average Cal sensor Temp.: 99.01 Ambient Temp.: 87.53 F
16:17:07.776 -> Average Cal sensor Temp.: 99.25 Ambient Temp.: 87.46 F
16:17:09.333 -> Average Cal sensor Temp.: 99.15 Ambient Temp.: 87.46 F
16:17:10.888 -> Average Cal sensor Temp.: 99.61 Ambient Temp.: 87.42 F
16:17:12.437 -> Average Cal sensor Temp.: 99.29 Ambient Temp.: 87.39 F
16:17:13.994 -> Average Cal sensor Temp.: 99.71 Ambient Temp.: 87.39 F
16:17:15.583 -> Average Cal sensor Temp.: 99.74 Ambient Temp.: 87.42 F
16:17:17.137 -> Average Cal sensor Temp.: 99.21 Ambient Temp.: 87.31 F
16:17:18.691 -> Average Cal sensor Temp.: 99.51 Ambient Temp.: 87.31 F
16:17:20.244 -> Average Cal sensor Temp.: 99.54 Ambient Temp.: 87.28 F
16:17:21.803 -> Average Cal sensor Temp.: 99.61 Ambient Temp.: 87.28 F
16:17:23.350 -> Average Cal sensor Temp.: 99.60 Ambient Temp.: 87.21 F
16:17:24.941 -> Average Cal sensor Temp.: 99.73 Ambient Temp.: 87.28 F
16:17:26.492 -> Average Cal sensor Temp.: 99.94 Ambient Temp.: 87.21 F
16:17:28.040 -> Average Cal sensor Temp.: 100.09 Ambient Temp.: 87.28 F
16:17:29.631 -> Average Cal sensor Temp.: 99.54 Ambient Temp.: 87.28 F
16:17:31.183 -> Average Cal sensor Temp.: 99.68 Ambient Temp.: 87.21 F
16:17:32.733 -> Average Cal sensor Temp.: 99.54 Ambient Temp.: 87.28 F
16:17:34.319 -> Average Cal sensor Temp.: 99.18 Ambient Temp.: 87.21 F
16:17:35.873 -> Average Cal sensor Temp.: 99.31 Ambient Temp.: 87.17 F
16:17:37.426 -> Average Cal sensor Temp.: 99.68 Ambient Temp.: 87.17 F
16:17:38.977 -> Average Cal sensor Temp.: 99.65 Ambient Temp.: 87.17 F
16:17:40.530 -> Average Cal sensor Temp.: 99.77 Ambient Temp.: 87.17 F
16:17:42.086 -> Average Cal sensor Temp.: 99.10 Ambient Temp.: 87.17 F
16:17:43.644 -> Average Cal sensor Temp.: 99.41 Ambient Temp.: 87.13 F
16:17:45.195 -> Average Cal sensor Temp.: 99.74 Ambient Temp.: 87.10 F
16:17:46.752 -> Average Cal sensor Temp.: 99.91 Ambient Temp.: 87.10 F
16:17:48.344 -> Average Cal sensor Temp.: 99.54 Ambient Temp.: 87.06 F
16:17:49.897 -> Average Cal sensor Temp.: 99.46 Ambient Temp.: 87.10 F
16:17:51.453 -> Average Cal sensor Temp.: 99.66 Ambient Temp.: 87.10 F
16:17:53.012 -> Average Cal sensor Temp.: 99.39 Ambient Temp.: 87.06 F
16:17:54.571 -> Average Cal sensor Temp.: 99.61 Ambient Temp.: 87.13 F
16:17:56.124 -> Average Cal sensor Temp.: 99.53 Ambient Temp.: 87.10 F
16:17:57.679 -> Average Cal sensor Temp.: 99.72 Ambient Temp.: 87.13 F
16:17:59.239 -> Average Cal sensor Temp.: 99.90 Ambient Temp.: 87.06 F
16:18:00.827 -> Average Cal sensor Temp.: 99.88 Ambient Temp.: 87.06 F
16:18:02.384 -> Average Cal sensor Temp.: 99.46 Ambient Temp.: 87.03 F
16:18:03.933 -> Average Cal sensor Temp.: 99.75 Ambient Temp.: 87.03 F
16:18:05.487 -> Average Cal sensor Temp.: 99.67 Ambient Temp.: 87.06 F
16:18:07.046 -> Average Cal sensor Temp.: 99.02 Ambient Temp.: 86.99 F
16:18:08.602 -> Average Cal sensor Temp.: 98.71 Ambient Temp.: 86.99 F
16:18:10.158 -> Average Cal sensor Temp.: 99.06 Ambient Temp.: 86.92 F
16:18:11.748 -> Average Cal sensor Temp.: 99.07 Ambient Temp.: 86.99 F
16:18:13.298 -> Average Cal sensor Temp.: 99.42 Ambient Temp.: 86.88 F
16:18:14.847 -> Average Cal sensor Temp.: 99.90 Ambient Temp.: 86.88 F
16:18:16.399 -> Average Cal sensor Temp.: 99.72 Ambient Temp.: 86.88 F
16:18:17.978 -> Average Cal sensor Temp.: 99.93 Ambient Temp.: 86.92 F
16:18:19.535 -> Average Cal sensor Temp.: 100.03 Ambient Temp.: 86.88 F
16:18:21.120 -> Average Cal sensor Temp.: 99.57 Ambient Temp.: 86.77 F
16:18:22.674 -> Average Cal sensor Temp.: 99.79 Ambient Temp.: 86.77 F
16:18:24.232 -> Average Cal sensor Temp.: 99.82 Ambient Temp.: 86.77 F
16:18:25.785 -> Average Cal sensor Temp.: 99.31 Ambient Temp.: 86.77 F
16:18:27.334 -> Average Cal sensor Temp.: 99.35 Ambient Temp.: 86.67 F
16:18:28.888 -> Average Cal sensor Temp.: 99.17 Ambient Temp.: 86.74 F
16:18:30.445 -> Average Cal sensor Temp.: 99.81 Ambient Temp.: 86.67 F
16:18:32.036 -> Average Cal sensor Temp.: 99.80 Ambient Temp.: 86.67 F
16:18:33.589 -> Average Cal sensor Temp.: 99.97 Ambient Temp.: 86.67 F
16:18:35.143 -> Average Cal sensor Temp.: 99.23 Ambient Temp.: 86.70 F
16:18:36.702 -> Average Cal sensor Temp.: 98.55 Ambient Temp.: 86.59 F
16:18:38.261 -> Average Cal sensor Temp.: 98.37 Ambient Temp.: 86.49 F
16:18:39.821 -> Average Cal sensor Temp.: 98.43 Ambient Temp.: 86.52 F
16:18:41.383 -> Average Cal sensor Temp.: 99.00 Ambient Temp.: 86.52 F
16:18:42.942 -> Average Cal sensor Temp.: 98.91 Ambient Temp.: 86.38 F
16:18:44.503 -> Average Cal sensor Temp.: 99.41 Ambient Temp.: 86.34 F
16:18:46.051 -> Average Cal sensor Temp.: 99.60 Ambient Temp.: 86.31 F
16:18:47.607 -> Average Cal sensor Temp.: 99.37 Ambient Temp.: 86.31 F
16:18:49.159 -> Average Cal sensor Temp.: 99.39 Ambient Temp.: 86.31 F
16:18:50.714 -> Average Cal sensor Temp.: 99.65 Ambient Temp.: 86.34 F
16:18:52.267 -> Average Cal sensor Temp.: 99.62 Ambient Temp.: 86.31 F
16:18:53.849 -> Average Cal sensor Temp.: 99.50 Ambient Temp.: 86.27 F
16:18:55.406 -> Average Cal sensor Temp.: 99.55 Ambient Temp.: 86.23 F
16:18:56.966 -> Average Cal sensor Temp.: 99.68 Ambient Temp.: 86.31 F
16:18:58.520 -> Average Cal sensor Temp.: 99.42 Ambient Temp.: 86.27 F
16:19:00.079 -> Average Cal sensor Temp.: 99.51 Ambient Temp.: 86.20 F
16:19:01.636 -> Average Cal sensor Temp.: 99.09 Ambient Temp.: 86.20 F
16:19:03.192 -> Average Cal sensor Temp.: 99.22 Ambient Temp.: 86.13 F
16:19:04.745 -> Average Cal sensor Temp.: 99.34 Ambient Temp.: 86.13 F
16:19:06.330 -> Average Cal sensor Temp.: 99.81 Ambient Temp.: 86.13 F
16:19:07.887 -> Average Cal sensor Temp.: 100.31 Ambient Temp.: 86.16 F
16:19:09.442 -> Average Cal sensor Temp.: 99.84 Ambient Temp.: 86.13 F
16:19:11.002 -> Average Cal sensor Temp.: 100.07 Ambient Temp.: 86.16 F
16:19:12.554 -> Average Cal sensor Temp.: 99.94 Ambient Temp.: 86.13 F
16:19:14.141 -> Average Cal sensor Temp.: 99.59 Ambient Temp.: 86.16 F
16:19:15.698 -> Average Cal sensor Temp.: 99.64 Ambient Temp.: 86.13 F
16:19:17.247 -> Average Cal sensor Temp.: 99.68 Ambient Temp.: 86.09 F
16:19:18.832 -> Average Cal sensor Temp.: 99.94 Ambient Temp.: 86.13 F
16:19:20.385 -> Average Cal sensor Temp.: 99.70 Ambient Temp.: 86.13 F
16:19:21.937 -> Average Cal sensor Temp.: 99.83 Ambient Temp.: 86.16 F
16:19:23.494 -> Average Cal sensor Temp.: 99.87 Ambient Temp.: 86.16 F
16:19:25.078 -> Average Cal sensor Temp.: 99.71 Ambient Temp.: 86.16 F
16:19:26.630 -> Average Cal sensor Temp.: 100.01 Ambient Temp.: 86.20 F
16:19:28.188 -> Average Cal sensor Temp.: 99.95 Ambient Temp.: 86.16 F
16:19:29.773 -> Average Cal sensor Temp.: 99.48 Ambient Temp.: 86.20 F
16:19:31.329 -> Average Cal sensor Temp.: 99.92 Ambient Temp.: 86.20 F
16:19:32.887 -> Average Cal sensor Temp.: 99.96 Ambient Temp.: 86.23 F
16:19:34.442 -> Average Cal sensor Temp.: 99.55 Ambient Temp.: 86.23 F
16:19:36.000 -> Average Cal sensor Temp.: 99.85 Ambient Temp.: 86.31 F
16:19:37.554 -> Average Cal sensor Temp.: 100.01 Ambient Temp.: 86.31 F
16:19:39.140 -> Average Cal sensor Temp.: 99.64 Ambient Temp.: 86.31 F
16:19:40.703 -> Average Cal sensor Temp.: 99.54 Ambient Temp.: 86.31 F
16:19:42.254 -> Average Cal sensor Temp.: 99.65 Ambient Temp.: 86.27 F
16:19:43.809 -> Average Cal sensor Temp.: 99.77 Ambient Temp.: 86.34 F
16:19:45.364 -> Average Cal sensor Temp.: 99.61 Ambient Temp.: 86.31 F
16:19:46.921 -> Average Cal sensor Temp.: 99.26 Ambient Temp.: 86.27 F
16:19:48.511 -> Average Cal sensor Temp.: 99.40 Ambient Temp.: 86.27 F
16:19:50.067 -> Average Cal sensor Temp.: 99.20 Ambient Temp.: 86.31 F
16:19:51.614 -> Average Cal sensor Temp.: 99.51 Ambient Temp.: 86.34 F
16:19:53.169 -> Average Cal sensor Temp.: 99.81 Ambient Temp.: 86.34 F
16:19:54.724 -> Average Cal sensor Temp.: 99.92 Ambient Temp.: 86.34 F
16:19:56.279 -> Average Cal sensor Temp.: 100.02 Ambient Temp.: 86.34 F
16:19:57.865 -> Average Cal sensor Temp.: 99.90 Ambient Temp.: 86.38 F
16:19:59.422 -> Average Cal sensor Temp.: 99.90 Ambient Temp.: 86.38 F
16:20:00.978 -> Average Cal sensor Temp.: 100.14 Ambient Temp.: 86.38 F
16:20:02.559 -> Average Cal sensor Temp.: 99.85 Ambient Temp.: 86.38 F
16:20:04.112 -> Average Cal sensor Temp.: 100.13 Ambient Temp.: 86.45 F
16:20:05.694 -> Average Cal sensor Temp.: 100.03 Ambient Temp.: 86.49 F
16:20:07.245 -> Average Cal sensor Temp.: 99.79 Ambient Temp.: 86.52 F
16:20:08.800 -> Average Cal sensor Temp.: 99.51 Ambient Temp.: 86.49 F
16:20:10.385 -> Average Cal sensor Temp.: 99.27 Ambient Temp.: 86.52 F
16:20:11.937 -> Average Cal sensor Temp.: 99.08 Ambient Temp.: 86.45 F
16:20:13.489 -> Average Cal sensor Temp.: 99.56 Ambient Temp.: 86.45 F
16:20:15.044 -> Average Cal sensor Temp.: 99.72 Ambient Temp.: 86.52 F
16:20:16.596 -> Average Cal sensor Temp.: 99.78 Ambient Temp.: 86.49 F
16:20:18.181 -> Average Cal sensor Temp.: 100.02 Ambient Temp.: 86.52 F
16:20:19.738 -> Average Cal sensor Temp.: 100.01 Ambient Temp.: 86.56 F
16:20:21.290 -> Average Cal sensor Temp.: 99.89 Ambient Temp.: 86.52 F
16:20:22.871 -> Average Cal sensor Temp.: 100.06 Ambient Temp.: 86.59 F
16:20:24.422 -> Average Cal sensor Temp.: 100.11 Ambient Temp.: 86.59 F
16:20:26.005 -> Average Cal sensor Temp.: 99.92 Ambient Temp.: 86.56 F
16:20:27.556 -> Average Cal sensor Temp.: 100.10 Ambient Temp.: 86.63 F
16:20:29.111 -> Average Cal sensor Temp.: 100.10 Ambient Temp.: 86.67 F
16:20:30.699 -> Average Cal sensor Temp.: 100.25 Ambient Temp.: 86.70 F
16:20:32.250 -> Average Cal sensor Temp.: 100.20 Ambient Temp.: 86.70 F
16:20:33.839 -> Average Cal sensor Temp.: 100.00 Ambient Temp.: 86.74 F
16:20:35.392 -> Average Cal sensor Temp.: 100.04 Ambient Temp.: 86.70 F
16:20:36.948 -> Average Cal sensor Temp.: 100.04 Ambient Temp.: 86.77 F
16:20:38.533 -> Average Cal sensor Temp.: 100.22 Ambient Temp.: 86.81 F
16:20:40.081 -> Average Cal sensor Temp.: 99.91 Ambient Temp.: 86.88 F
16:20:41.635 -> Average Cal sensor Temp.: 99.68 Ambient Temp.: 86.92 F
16:20:43.219 -> Average Cal sensor Temp.: 99.76 Ambient Temp.: 86.88 F
16:20:44.772 -> Average Cal sensor Temp.: 99.95 Ambient Temp.: 86.95 F
16:20:46.329 -> Average Cal sensor Temp.: 99.80 Ambient Temp.: 86.92 F
16:20:47.917 -> Average Cal sensor Temp.: 100.02 Ambient Temp.: 86.95 F
16:20:49.468 -> Average Cal sensor Temp.: 100.10 Ambient Temp.: 86.99 F
16:20:51.053 -> Average Cal sensor Temp.: 100.20 Ambient Temp.: 87.03 F
16:20:52.608 -> Average Cal sensor Temp.: 99.90 Ambient Temp.: 87.06 F
16:20:54.163 -> Average Cal sensor Temp.: 99.84 Ambient Temp.: 87.06 F
16:20:55.717 -> Average Cal sensor Temp.: 99.90 Ambient Temp.: 87.06 F
16:20:57.277 -> Average Cal sensor Temp.: 100.02 Ambient Temp.: 87.06 F
16:20:58.867 -> Average Cal sensor Temp.: 99.85 Ambient Temp.: 87.13 F
16:21:00.423 -> Average Cal sensor Temp.: 99.72 Ambient Temp.: 87.13 F
16:21:01.977 -> Average Cal sensor Temp.: 99.99 Ambient Temp.: 87.13 F
16:21:03.534 -> Average Cal sensor Temp.: 100.17 Ambient Temp.: 87.13 F
16:21:05.122 -> Average Cal sensor Temp.: 100.27 Ambient Temp.: 87.17 F
16:21:06.675 -> Average Cal sensor Temp.: 100.38 Ambient Temp.: 87.21 F
16:21:08.226 -> Average Cal sensor Temp.: 100.22 Ambient Temp.: 87.21 F
16:21:09.783 -> Average Cal sensor Temp.: 100.11 Ambient Temp.: 87.21 F
16:21:11.373 -> Average Cal sensor Temp.: 100.18 Ambient Temp.: 87.17 F
16:21:12.929 -> Average Cal sensor Temp.: 100.18 Ambient Temp.: 87.28 F
16:21:14.486 -> Average Cal sensor Temp.: 100.14 Ambient Temp.: 87.28 F
16:21:16.043 -> Average Cal sensor Temp.: 100.05 Ambient Temp.: 87.28 F
16:21:17.631 -> Average Cal sensor Temp.: 100.08 Ambient Temp.: 87.31 F
16:21:19.182 -> Average Cal sensor Temp.: 100.32 Ambient Temp.: 87.35 F
16:21:20.738 -> Average Cal sensor Temp.: 100.45 Ambient Temp.: 87.39 F
16:21:22.325 -> Average Cal sensor Temp.: 100.33 Ambient Temp.: 87.42 F
16:21:23.881 -> Average Cal sensor Temp.: 100.18 Ambient Temp.: 87.42 F
16:21:25.434 -> Average Cal sensor Temp.: 100.13 Ambient Temp.: 87.42 F
16:21:27.024 -> Average Cal sensor Temp.: 100.32 Ambient Temp.: 87.46 F
16:21:28.576 -> Average Cal sensor Temp.: 100.19 Ambient Temp.: 87.53 F
16:21:30.123 -> Average Cal sensor Temp.: 100.19 Ambient Temp.: 87.53 F
16:21:31.707 -> Average Cal sensor Temp.: 100.26 Ambient Temp.: 87.57 F
16:21:33.261 -> Average Cal sensor Temp.: 99.99 Ambient Temp.: 87.60 F
16:21:34.813 -> Average Cal sensor Temp.: 100.35 Ambient Temp.: 87.57 F
16:21:36.401 -> Average Cal sensor Temp.: 100.23 Ambient Temp.: 87.67 F
16:21:37.947 -> Average Cal sensor Temp.: 100.08 Ambient Temp.: 87.57 F
16:21:39.528 -> Average Cal sensor Temp.: 100.37 Ambient Temp.: 87.71 F
16:21:41.083 -> Average Cal sensor Temp.: 100.27 Ambient Temp.: 87.75 F
16:21:42.666 -> Average Cal sensor Temp.: 100.32 Ambient Temp.: 87.78 F
16:21:44.218 -> Average Cal sensor Temp.: 100.17 Ambient Temp.: 87.78 F
16:21:45.776 -> Average Cal sensor Temp.: 100.14 Ambient Temp.: 87.75 F
16:21:47.359 -> Average Cal sensor Temp.: 100.25 Ambient Temp.: 87.78 F
16:21:48.904 -> Average Cal sensor Temp.: 100.18 Ambient Temp.: 87.85 F
16:21:50.459 -> Average Cal sensor Temp.: 100.37 Ambient Temp.: 87.82 F
16:21:52.043 -> Average Cal sensor Temp.: 100.69 Ambient Temp.: 87.85 F
16:21:53.593 -> Average Cal sensor Temp.: 100.69 Ambient Temp.: 87.89 F
16:21:55.174 -> Average Cal sensor Temp.: 100.34 Ambient Temp.: 87.96 F
16:21:56.728 -> Average Cal sensor Temp.: 100.50 Ambient Temp.: 87.93 F
16:21:58.317 -> Average Cal sensor Temp.: 100.36 Ambient Temp.: 87.96 F
16:21:59.873 -> Average Cal sensor Temp.: 100.66 Ambient Temp.: 87.96 F
16:22:01.430 -> Average Cal sensor Temp.: 100.53 Ambient Temp.: 88.07 F
16:22:03.018 -> Average Cal sensor Temp.: 100.56 Ambient Temp.: 88.07 F
16:22:04.567 -> Average Cal sensor Temp.: 100.36 Ambient Temp.: 88.00 F
16:22:06.119 -> Average Cal sensor Temp.: 100.18 Ambient Temp.: 88.07 F
16:22:07.704 -> Average Cal sensor Temp.: 100.73 Ambient Temp.: 88.11 F
16:22:09.252 -> Average Cal sensor Temp.: 100.02 Ambient Temp.: 88.14 F
16:22:10.840 -> Average Cal sensor Temp.: 99.95 Ambient Temp.: 88.07 F
16:22:12.396 -> Average Cal sensor Temp.: 100.33 Ambient Temp.: 88.11 F
16:22:13.948 -> Average Cal sensor Temp.: 100.32 Ambient Temp.: 88.18 F
16:22:15.536 -> Average Cal sensor Temp.: 99.87 Ambient Temp.: 88.14 F
16:22:17.093 -> Average Cal sensor Temp.: 100.22 Ambient Temp.: 88.11 F
16:22:18.645 -> Average Cal sensor Temp.: 100.28 Ambient Temp.: 88.18 F
16:22:20.230 -> Average Cal sensor Temp.: 99.94 Ambient Temp.: 88.18 F
16:22:21.787 -> Average Cal sensor Temp.: 99.97 Ambient Temp.: 88.11 F
16:22:23.343 -> Average Cal sensor Temp.: 99.88 Ambient Temp.: 88.21 F
16:22:24.894 -> Average Cal sensor Temp.: 99.74 Ambient Temp.: 88.14 F
16:22:26.449 -> Average Cal sensor Temp.: 99.96 Ambient Temp.: 88.14 F
16:22:28.033 -> Average Cal sensor Temp.: 99.74 Ambient Temp.: 88.18 F
16:22:29.588 -> Average Cal sensor Temp.: 99.86 Ambient Temp.: 88.14 F
16:22:31.172 -> Average Cal sensor Temp.: 100.19 Ambient Temp.: 88.11 F
16:22:32.717 -> Average Cal sensor Temp.: 99.87 Ambient Temp.: 88.14 F
16:22:34.265 -> Average Cal sensor Temp.: 99.58 Ambient Temp.: 88.11 F
16:22:35.857 -> Average Cal sensor Temp.: 99.50 Ambient Temp.: 88.14 F
16:22:37.406 -> Average Cal sensor Temp.: 99.76 Ambient Temp.: 88.11 F
16:22:38.963 -> Average Cal sensor Temp.: 100.33 Ambient Temp.: 88.00 F
16:22:40.514 -> Average Cal sensor Temp.: 99.82 Ambient Temp.: 88.11 F
16:22:42.101 -> Average Cal sensor Temp.: 99.56 Ambient Temp.: 88.00 F
16:22:43.658 -> Average Cal sensor Temp.: 99.81 Ambient Temp.: 87.96 F
16:22:45.214 -> Average Cal sensor Temp.: 99.50 Ambient Temp.: 88.00 F
16:22:46.772 -> Average Cal sensor Temp.: 99.96 Ambient Temp.: 87.96 F
16:22:48.324 -> Average Cal sensor Temp.: 99.47 Ambient Temp.: 87.93 F
16:22:49.882 -> Average Cal sensor Temp.: 99.52 Ambient Temp.: 87.93 F
16:22:51.468 -> Average Cal sensor Temp.: 99.91 Ambient Temp.: 87.93 F
16:22:53.023 -> Average Cal sensor Temp.: 100.21 Ambient Temp.: 87.93 F
16:22:54.567 -> Average Cal sensor Temp.: 100.07 Ambient Temp.: 87.93 F
16:22:56.158 -> Average Cal sensor Temp.: 99.99 Ambient Temp.: 87.89 F
16:22:57.708 -> Average Cal sensor Temp.: 99.41 Ambient Temp.: 87.89 F
16:22:59.263 -> Average Cal sensor Temp.: 99.26 Ambient Temp.: 87.85 F
16:23:00.820 -> Average Cal sensor Temp.: 99.44 Ambient Temp.: 87.89 F
16:23:02.375 -> Average Cal sensor Temp.: 99.97 Ambient Temp.: 87.85 F
16:23:03.926 -> Average Cal sensor Temp.: 100.05 Ambient Temp.: 87.85 F
16:23:05.512 -> Average Cal sensor Temp.: 100.12 Ambient Temp.: 87.85 F
16:23:07.069 -> Average Cal sensor Temp.: 99.98 Ambient Temp.: 87.85 F
16:23:08.619 -> Average Cal sensor Temp.: 99.46 Ambient Temp.: 87.85 F
16:23:10.206 -> Average Cal sensor Temp.: 99.30 Ambient Temp.: 87.82 F
16:23:11.762 -> Average Cal sensor Temp.: 99.46 Ambient Temp.: 87.89 F
16:23:13.315 -> Average Cal sensor Temp.: 99.67 Ambient Temp.: 87.78 F
16:23:14.870 -> Average Cal sensor Temp.: 100.01 Ambient Temp.: 87.82 F
16:23:16.425 -> Average Cal sensor Temp.: 99.47 Ambient Temp.: 87.78 F
16:23:17.975 -> Average Cal sensor Temp.: 99.29 Ambient Temp.: 87.75 F
16:23:19.553 -> Average Cal sensor Temp.: 99.73 Ambient Temp.: 87.78 F
16:23:21.111 -> Average Cal sensor Temp.: 99.27 Ambient Temp.: 87.75 F
16:23:22.657 -> Average Cal sensor Temp.: 99.95 Ambient Temp.: 87.75 F
16:23:24.244 -> Average Cal sensor Temp.: 100.35 Ambient Temp.: 87.67 F
16:23:25.792 -> Average Cal sensor Temp.: 99.72 Ambient Temp.: 87.67 F
16:23:27.342 -> Average Cal sensor Temp.: 99.81 Ambient Temp.: 87.67 F
16:23:28.925 -> Average Cal sensor Temp.: 99.59 Ambient Temp.: 87.60 F
16:23:30.480 -> Average Cal sensor Temp.: 99.25 Ambient Temp.: 87.60 F
16:23:32.035 -> Average Cal sensor Temp.: 99.45 Ambient Temp.: 87.57 F
16:23:33.589 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 87.57 F
16:23:35.147 -> Average Cal sensor Temp.: 99.95 Ambient Temp.: 87.53 F
16:23:36.704 -> Average Cal sensor Temp.: 99.52 Ambient Temp.: 87.53 F
16:23:38.259 -> Average Cal sensor Temp.: 99.88 Ambient Temp.: 87.53 F
16:23:39.852 -> Average Cal sensor Temp.: 99.90 Ambient Temp.: 87.46 F
16:23:41.412 -> Average Cal sensor Temp.: 100.07 Ambient Temp.: 87.49 F
16:23:42.968 -> Average Cal sensor Temp.: 100.23 Ambient Temp.: 87.46 F
16:23:44.518 -> Average Cal sensor Temp.: 99.90 Ambient Temp.: 87.42 F
16:23:46.079 -> Average Cal sensor Temp.: 99.66 Ambient Temp.: 87.42 F
16:23:47.670 -> Average Cal sensor Temp.: 99.96 Ambient Temp.: 87.42 F
16:23:49.224 -> Average Cal sensor Temp.: 100.19 Ambient Temp.: 87.46 F
16:23:50.770 -> Average Cal sensor Temp.: 100.04 Ambient Temp.: 87.39 F
16:23:52.347 -> Average Cal sensor Temp.: 100.16 Ambient Temp.: 87.42 F
16:23:53.898 -> Average Cal sensor Temp.: 99.30 Ambient Temp.: 87.42 F
16:23:55.478 -> Average Cal sensor Temp.: 99.57 Ambient Temp.: 87.35 F
16:23:57.037 -> Average Cal sensor Temp.: 100.19 Ambient Temp.: 87.35 F
16:23:58.591 -> Average Cal sensor Temp.: 100.26 Ambient Temp.: 87.39 F
16:24:00.177 -> Average Cal sensor Temp.: 100.09 Ambient Temp.: 87.35 F
16:24:01.729 -> Average Cal sensor Temp.: 99.98 Ambient Temp.: 87.39 F
16:24:03.274 -> Average Cal sensor Temp.: 100.09 Ambient Temp.: 87.31 F
16:24:04.860 -> Average Cal sensor Temp.: 99.73 Ambient Temp.: 87.31 F
16:24:06.417 -> Average Cal sensor Temp.: 99.84 Ambient Temp.: 87.28 F
16:24:07.973 -> Average Cal sensor Temp.: 99.77 Ambient Temp.: 87.21 F
16:24:09.532 -> Average Cal sensor Temp.: 100.16 Ambient Temp.: 87.31 F
16:24:11.090 -> Average Cal sensor Temp.: 100.47 Ambient Temp.: 87.21 F
16:24:12.682 -> Average Cal sensor Temp.: 100.22 Ambient Temp.: 87.28 F
16:24:14.239 -> Average Cal sensor Temp.: 100.22 Ambient Temp.: 87.31 F
16:24:15.790 -> Average Cal sensor Temp.: 99.88 Ambient Temp.: 87.28 F
16:24:17.346 -> Average Cal sensor Temp.: 99.84 Ambient Temp.: 87.21 F
16:24:18.933 -> Average Cal sensor Temp.: 99.51 Ambient Temp.: 87.21 F
16:24:20.497 -> Average Cal sensor Temp.: 99.40 Ambient Temp.: 87.17 F
16:24:22.051 -> Average Cal sensor Temp.: 99.00 Ambient Temp.: 87.17 F
16:24:23.601 -> Average Cal sensor Temp.: 99.72 Ambient Temp.: 87.13 F
16:24:25.154 -> Average Cal sensor Temp.: 99.68 Ambient Temp.: 87.17 F
16:24:26.711 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 87.06 F
16:24:28.267 -> Average Cal sensor Temp.: 99.62 Ambient Temp.: 87.10 F
16:24:29.823 -> Average Cal sensor Temp.: 99.65 Ambient Temp.: 87.06 F
16:24:31.414 -> Average Cal sensor Temp.: 99.73 Ambient Temp.: 87.06 F
16:24:32.972 -> Average Cal sensor Temp.: 100.19 Ambient Temp.: 87.06 F
16:24:34.523 -> Average Cal sensor Temp.: 100.13 Ambient Temp.: 87.06 F
16:24:36.078 -> Average Cal sensor Temp.: 100.40 Ambient Temp.: 87.03 F
16:24:37.672 -> Average Cal sensor Temp.: 100.45 Ambient Temp.: 87.06 F
16:24:39.226 -> Average Cal sensor Temp.: 100.43 Ambient Temp.: 87.06 F
16:24:40.781 -> Average Cal sensor Temp.: 100.38 Ambient Temp.: 87.03 F
16:24:42.372 -> Average Cal sensor Temp.: 99.84 Ambient Temp.: 87.03 F
16:24:43.926 -> Average Cal sensor Temp.: 100.33 Ambient Temp.: 87.03 F
16:24:45.475 -> Average Cal sensor Temp.: 99.86 Ambient Temp.: 86.95 F
16:24:47.035 -> Average Cal sensor Temp.: 99.33 Ambient Temp.: 86.95 F
16:24:48.621 -> Average Cal sensor Temp.: 99.76 Ambient Temp.: 86.92 F
16:24:50.174 -> Average Cal sensor Temp.: 100.19 Ambient Temp.: 86.92 F
16:24:51.729 -> Average Cal sensor Temp.: 100.36 Ambient Temp.: 86.95 F
16:24:53.316 -> Average Cal sensor Temp.: 100.46 Ambient Temp.: 86.92 F
16:24:54.868 -> Average Cal sensor Temp.: 99.69 Ambient Temp.: 86.92 F
16:24:56.426 -> Average Cal sensor Temp.: 99.79 Ambient Temp.: 86.92 F
16:24:57.985 -> Average Cal sensor Temp.: 100.95 Ambient Temp.: 86.92 F
16:24:59.570 -> Average Cal sensor Temp.: 100.80 Ambient Temp.: 86.88 F
16:25:01.123 -> Average Cal sensor Temp.: 100.29 Ambient Temp.: 86.88 F
16:25:02.682 -> Average Cal sensor Temp.: 99.79 Ambient Temp.: 86.77 F
16:25:04.242 -> Average Cal sensor Temp.: 100.49 Ambient Temp.: 86.81 F
16:25:05.833 -> Average Cal sensor Temp.: 100.67 Ambient Temp.: 86.74 F
16:25:07.388 -> Average Cal sensor Temp.: 100.22 Ambient Temp.: 86.70 F
16:25:08.970 -> Average Cal sensor Temp.: 99.94 Ambient Temp.: 86.74 F
16:25:10.522 -> Average Cal sensor Temp.: 99.77 Ambient Temp.: 86.70 F
16:25:12.076 -> Average Cal sensor Temp.: 100.29 Ambient Temp.: 86.67 F
16:25:13.631 -> Average Cal sensor Temp.: 100.45 Ambient Temp.: 86.67 F
16:25:15.223 -> Average Cal sensor Temp.: 100.12 Ambient Temp.: 86.67 F
16:25:16.776 -> Average Cal sensor Temp.: 100.35 Ambient Temp.: 86.67 F
16:25:18.348 -> Average Cal sensor Temp.: 100.12 Ambient Temp.: 86.59 F
16:25:19.896 -> Average Cal sensor Temp.: 100.40 Ambient Temp.: 86.63 F
16:25:21.484 -> Average Cal sensor Temp.: 100.00 Ambient Temp.: 86.56 F
16:25:23.036 -> Average Cal sensor Temp.: 99.96 Ambient Temp.: 86.56 F
16:25:24.593 -> Average Cal sensor Temp.: 99.73 Ambient Temp.: 86.56 F
16:25:26.149 -> Average Cal sensor Temp.: 100.20 Ambient Temp.: 86.59 F
16:25:27.736 -> Average Cal sensor Temp.: 100.38 Ambient Temp.: 86.63 F
16:25:29.296 -> Average Cal sensor Temp.: 99.74 Ambient Temp.: 86.52 F
16:25:30.847 -> Average Cal sensor Temp.: 99.90 Ambient Temp.: 86.52 F
16:25:32.435 -> Average Cal sensor Temp.: 99.73 Ambient Temp.: 86.49 F
16:25:33.990 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 86.45 F
16:25:35.547 -> Average Cal sensor Temp.: 99.83 Ambient Temp.: 86.45 F
16:25:37.105 -> Average Cal sensor Temp.: 100.41 Ambient Temp.: 86.38 F
16:25:38.689 -> Average Cal sensor Temp.: 100.53 Ambient Temp.: 86.34 F
16:25:40.247 -> Average Cal sensor Temp.: 99.79 Ambient Temp.: 86.38 F
16:25:41.802 -> Average Cal sensor Temp.: 100.14 Ambient Temp.: 86.31 F
16:25:43.384 -> Average Cal sensor Temp.: 99.90 Ambient Temp.: 86.31 F
16:25:44.931 -> Average Cal sensor Temp.: 100.46 Ambient Temp.: 86.27 F
16:25:46.480 -> Average Cal sensor Temp.: 100.49 Ambient Temp.: 86.27 F
16:25:48.070 -> Average Cal sensor Temp.: 100.22 Ambient Temp.: 86.27 F
16:25:49.617 -> Average Cal sensor Temp.: 99.95 Ambient Temp.: 86.27 F
16:25:51.195 -> Average Cal sensor Temp.: 100.84 Ambient Temp.: 86.31 F
16:25:52.753 -> Average Cal sensor Temp.: 100.04 Ambient Temp.: 86.23 F
16:25:54.338 -> Average Cal sensor Temp.: 100.14 Ambient Temp.: 86.27 F
16:25:55.891 -> Average Cal sensor Temp.: 100.27 Ambient Temp.: 86.23 F
16:25:57.439 -> Average Cal sensor Temp.: 100.48 Ambient Temp.: 86.34 F
16:25:59.027 -> Average Cal sensor Temp.: 100.54 Ambient Temp.: 86.27 F
16:26:00.580 -> Average Cal sensor Temp.: 100.31 Ambient Temp.: 86.23 F
16:26:02.162 -> Average Cal sensor Temp.: 100.60 Ambient Temp.: 86.27 F
16:26:03.715 -> Average Cal sensor Temp.: 100.68 Ambient Temp.: 86.31 F
16:26:05.273 -> Average Cal sensor Temp.: 100.62 Ambient Temp.: 86.31 F
16:26:06.828 -> Average Cal sensor Temp.: 100.51 Ambient Temp.: 86.34 F
16:26:08.415 -> Average Cal sensor Temp.: 100.29 Ambient Temp.: 86.27 F
16:26:09.969 -> Average Cal sensor Temp.: 100.27 Ambient Temp.: 86.34 F
16:26:11.525 -> Average Cal sensor Temp.: 100.55 Ambient Temp.: 86.38 F
16:26:13.109 -> Average Cal sensor Temp.: 100.15 Ambient Temp.: 86.45 F
16:26:14.662 -> Average Cal sensor Temp.: 100.18 Ambient Temp.: 86.38 F
16:26:16.219 -> Average Cal sensor Temp.: 99.92 Ambient Temp.: 86.38 F
16:26:17.800 -> Average Cal sensor Temp.: 100.07 Ambient Temp.: 86.45 F
16:26:19.360 -> Average Cal sensor Temp.: 100.19 Ambient Temp.: 86.45 F
16:26:20.942 -> Average Cal sensor Temp.: 99.96 Ambient Temp.: 86.45 F
16:26:22.499 -> Average Cal sensor Temp.: 99.88 Ambient Temp.: 86.45 F
16:26:24.054 -> Average Cal sensor Temp.: 100.26 Ambient Temp.: 86.49 F
16:26:25.640 -> Average Cal sensor Temp.: 100.19 Ambient Temp.: 86.45 F
16:26:27.196 -> Average Cal sensor Temp.: 100.31 Ambient Temp.: 86.45 F
16:26:28.750 -> Average Cal sensor Temp.: 100.41 Ambient Temp.: 86.52 F
16:26:30.308 -> Average Cal sensor Temp.: 100.12 Ambient Temp.: 86.52 F
16:26:31.897 -> Average Cal sensor Temp.: 100.52 Ambient Temp.: 86.52 F
16:26:33.447 -> Average Cal sensor Temp.: 100.66 Ambient Temp.: 86.52 F
16:26:35.004 -> Average Cal sensor Temp.: 100.72 Ambient Temp.: 86.49 F
16:26:36.590 -> Average Cal sensor Temp.: 100.42 Ambient Temp.: 86.56 F
16:26:38.148 -> Average Cal sensor Temp.: 100.26 Ambient Temp.: 86.56 F
16:26:39.704 -> Average Cal sensor Temp.: 100.29 Ambient Temp.: 86.59 F
16:26:41.296 -> Average Cal sensor Temp.: 100.57 Ambient Temp.: 86.59 F
16:26:42.846 -> Average Cal sensor Temp.: 100.11 Ambient Temp.: 86.63 F
16:26:44.400 -> Average Cal sensor Temp.: 99.62 Ambient Temp.: 86.67 F
16:26:45.984 -> Average Cal sensor Temp.: 99.75 Ambient Temp.: 86.63 F
16:26:47.533 -> Average Cal sensor Temp.: 99.70 Ambient Temp.: 86.63 F
16:26:49.088 -> Average Cal sensor Temp.: 99.90 Ambient Temp.: 86.63 F
16:26:50.674 -> Average Cal sensor Temp.: 99.44 Ambient Temp.: 86.63 F
16:26:52.231 -> Average Cal sensor Temp.: 99.84 Ambient Temp.: 86.52 F
16:26:53.783 -> Average Cal sensor Temp.: 99.92 Ambient Temp.: 86.63 F
16:26:55.336 -> Average Cal sensor Temp.: 99.81 Ambient Temp.: 86.63 F
16:26:56.923 -> Average Cal sensor Temp.: 99.81 Ambient Temp.: 86.56 F
16:26:58.479 -> Average Cal sensor Temp.: 100.13 Ambient Temp.: 86.63 F
16:27:00.031 -> Average Cal sensor Temp.: 100.17 Ambient Temp.: 86.67 F
16:27:01.591 -> Average Cal sensor Temp.: 100.39 Ambient Temp.: 86.63 F
16:27:03.180 -> Average Cal sensor Temp.: 100.34 Ambient Temp.: 86.67 F
16:27:04.732 -> Average Cal sensor Temp.: 100.26 Ambient Temp.: 86.63 F
16:27:06.288 -> Average Cal sensor Temp.: 100.58 Ambient Temp.: 86.67 F
16:27:07.845 -> Average Cal sensor Temp.: 100.68 Ambient Temp.: 86.70 F
16:27:09.430 -> Average Cal sensor Temp.: 100.42 Ambient Temp.: 86.74 F
16:27:10.980 -> Average Cal sensor Temp.: 100.62 Ambient Temp.: 86.67 F
16:27:12.571 -> Average Cal sensor Temp.: 100.79 Ambient Temp.: 86.70 F
16:27:14.120 -> Average Cal sensor Temp.: 100.64 Ambient Temp.: 86.77 F
16:27:15.708 -> Average Cal sensor Temp.: 100.43 Ambient Temp.: 86.77 F
16:27:17.253 -> Average Cal sensor Temp.: 100.17 Ambient Temp.: 86.77 F
16:27:18.806 -> Average Cal sensor Temp.: 100.48 Ambient Temp.: 86.81 F
16:27:20.398 -> Average Cal sensor Temp.: 100.16 Ambient Temp.: 86.81 F
16:27:21.951 -> Average Cal sensor Temp.: 100.11 Ambient Temp.: 86.88 F
16:27:23.503 -> Average Cal sensor Temp.: 100.07 Ambient Temp.: 86.88 F
16:27:25.094 -> Average Cal sensor Temp.: 100.03 Ambient Temp.: 86.92 F
16:27:26.650 -> Average Cal sensor Temp.: 100.30 Ambient Temp.: 86.88 F
16:27:28.232 -> Average Cal sensor Temp.: 100.43 Ambient Temp.: 86.95 F
16:27:29.787 -> Average Cal sensor Temp.: 100.34 Ambient Temp.: 86.99 F
16:27:31.338 -> Average Cal sensor Temp.: 100.51 Ambient Temp.: 87.03 F
16:27:32.924 -> Average Cal sensor Temp.: 100.69 Ambient Temp.: 86.99 F
16:27:34.469 -> Average Cal sensor Temp.: 100.30 Ambient Temp.: 87.03 F
16:27:36.052 -> Average Cal sensor Temp.: 99.90 Ambient Temp.: 87.03 F
16:27:37.602 -> Average Cal sensor Temp.: 100.50 Ambient Temp.: 87.03 F
16:27:39.181 -> Average Cal sensor Temp.: 100.56 Ambient Temp.: 87.10 F
16:27:40.737 -> Average Cal sensor Temp.: 100.40 Ambient Temp.: 87.03 F
16:27:42.321 -> Average Cal sensor Temp.: 100.31 Ambient Temp.: 87.10 F
16:27:43.873 -> Average Cal sensor Temp.: 100.38 Ambient Temp.: 87.13 F
16:27:45.448 -> Average Cal sensor Temp.: 100.45 Ambient Temp.: 87.10 F
16:27:47.001 -> Average Cal sensor Temp.: 100.38 Ambient Temp.: 87.13 F
16:27:48.548 -> Average Cal sensor Temp.: 100.40 Ambient Temp.: 87.17 F
16:27:50.134 -> Average Cal sensor Temp.: 100.29 Ambient Temp.: 87.17 F
16:27:51.687 -> Average Cal sensor Temp.: 100.22 Ambient Temp.: 87.21 F
16:27:53.272 -> Average Cal sensor Temp.: 100.22 Ambient Temp.: 87.21 F
16:27:54.821 -> Average Cal sensor Temp.: 100.08 Ambient Temp.: 87.28 F
16:27:56.405 -> Average Cal sensor Temp.: 100.08 Ambient Temp.: 87.28 F
16:27:57.959 -> Average Cal sensor Temp.: 100.27 Ambient Temp.: 87.21 F
16:27:59.514 -> Average Cal sensor Temp.: 100.77 Ambient Temp.: 87.28 F
16:28:01.071 -> Average Cal sensor Temp.: 100.99 Ambient Temp.: 87.35 F
16:28:02.661 -> Average Cal sensor Temp.: 100.77 Ambient Temp.: 87.35 F
16:28:04.218 -> Average Cal sensor Temp.: 101.29 Ambient Temp.: 87.39 F
16:28:05.777 -> Average Cal sensor Temp.: 100.82 Ambient Temp.: 87.46 F
16:28:07.330 -> Average Cal sensor Temp.: 100.97 Ambient Temp.: 87.46 F
16:28:08.922 -> Average Cal sensor Temp.: 100.67 Ambient Temp.: 87.49 F
16:28:10.471 -> Average Cal sensor Temp.: 100.84 Ambient Temp.: 87.49 F
16:28:12.059 -> Average Cal sensor Temp.: 100.39 Ambient Temp.: 87.49 F
16:28:13.611 -> Average Cal sensor Temp.: 100.78 Ambient Temp.: 87.53 F
16:28:15.163 -> Average Cal sensor Temp.: 100.52 Ambient Temp.: 87.49 F
16:28:16.750 -> Average Cal sensor Temp.: 100.47 Ambient Temp.: 87.53 F
16:28:18.303 -> Average Cal sensor Temp.: 100.85 Ambient Temp.: 87.60 F
16:28:19.852 -> Average Cal sensor Temp.: 100.73 Ambient Temp.: 87.57 F
16:28:21.444 -> Average Cal sensor Temp.: 100.63 Ambient Temp.: 87.60 F
16:28:22.994 -> Average Cal sensor Temp.: 100.50 Ambient Temp.: 87.71 F
16:28:24.549 -> Average Cal sensor Temp.: 100.31 Ambient Temp.: 87.67 F
16:28:26.135 -> Average Cal sensor Temp.: 100.32 Ambient Temp.: 87.71 F
16:28:27.683 -> Average Cal sensor Temp.: 100.58 Ambient Temp.: 87.75 F
16:28:29.271 -> Average Cal sensor Temp.: 100.48 Ambient Temp.: 87.75 F
16:28:30.827 -> Average Cal sensor Temp.: 100.78 Ambient Temp.: 87.71 F
16:28:32.385 -> Average Cal sensor Temp.: 100.94 Ambient Temp.: 87.78 F
16:28:33.943 -> Average Cal sensor Temp.: 100.57 Ambient Temp.: 87.75 F
16:28:35.530 -> Average Cal sensor Temp.: 101.06 Ambient Temp.: 87.78 F
16:28:37.091 -> Average Cal sensor Temp.: 100.75 Ambient Temp.: 87.85 F
16:28:38.650 -> Average Cal sensor Temp.: 100.94 Ambient Temp.: 87.82 F
16:28:40.207 -> Average Cal sensor Temp.: 100.65 Ambient Temp.: 87.89 F
16:28:41.767 -> Average Cal sensor Temp.: 100.68 Ambient Temp.: 87.89 F
16:28:43.349 -> Average Cal sensor Temp.: 100.44 Ambient Temp.: 87.89 F
16:28:44.910 -> Average Cal sensor Temp.: 100.63 Ambient Temp.: 87.89 F
16:28:46.469 -> Average Cal sensor Temp.: 100.43 Ambient Temp.: 87.93 F
16:28:48.025 -> Average Cal sensor Temp.: 100.36 Ambient Temp.: 87.93 F
16:28:49.608 -> Average Cal sensor Temp.: 100.53 Ambient Temp.: 87.93 F
16:28:51.164 -> Average Cal sensor Temp.: 100.79 Ambient Temp.: 87.96 F
16:28:52.753 -> Average Cal sensor Temp.: 100.46 Ambient Temp.: 88.00 F
16:28:54.303 -> Average Cal sensor Temp.: 100.36 Ambient Temp.: 88.00 F
16:28:55.854 -> Average Cal sensor Temp.: 100.25 Ambient Temp.: 87.96 F
16:28:57.445 -> Average Cal sensor Temp.: 100.25 Ambient Temp.: 88.07 F
16:28:59.003 -> Average Cal sensor Temp.: 99.93 Ambient Temp.: 88.07 F
16:29:00.559 -> Average Cal sensor Temp.: 100.26 Ambient Temp.: 88.00 F
16:29:02.145 -> Average Cal sensor Temp.: 100.13 Ambient Temp.: 88.07 F
16:29:03.698 -> Average Cal sensor Temp.: 100.13 Ambient Temp.: 88.00 F
16:29:05.256 -> Average Cal sensor Temp.: 99.74 Ambient Temp.: 88.00 F
16:29:06.811 -> Average Cal sensor Temp.: 100.01 Ambient Temp.: 88.07 F
16:29:08.406 -> Average Cal sensor Temp.: 100.02 Ambient Temp.: 87.96 F
16:29:09.963 -> Average Cal sensor Temp.: 100.13 Ambient Temp.: 87.96 F
16:29:11.520 -> Average Cal sensor Temp.: 100.05 Ambient Temp.: 87.96 F
16:29:13.073 -> Average Cal sensor Temp.: 100.15 Ambient Temp.: 88.07 F
16:29:14.656 -> Average Cal sensor Temp.: 100.31 Ambient Temp.: 88.11 F
16:29:16.208 -> Average Cal sensor Temp.: 100.02 Ambient Temp.: 87.96 F
16:29:17.793 -> Average Cal sensor Temp.: 99.81 Ambient Temp.: 87.93 F
16:29:19.348 -> Average Cal sensor Temp.: 100.12 Ambient Temp.: 88.07 F
16:29:20.906 -> Average Cal sensor Temp.: 99.82 Ambient Temp.: 87.89 F
16:29:22.462 -> Average Cal sensor Temp.: 99.88 Ambient Temp.: 87.93 F
16:29:24.049 -> Average Cal sensor Temp.: 99.88 Ambient Temp.: 87.89 F
16:29:25.607 -> Average Cal sensor Temp.: 100.17 Ambient Temp.: 87.89 F
16:29:27.166 -> Average Cal sensor Temp.: 100.48 Ambient Temp.: 87.89 F
16:29:28.721 -> Average Cal sensor Temp.: 100.68 Ambient Temp.: 87.89 F
16:29:30.283 -> Average Cal sensor Temp.: 100.41 Ambient Temp.: 87.85 F
16:29:31.875 -> Average Cal sensor Temp.: 100.19 Ambient Temp.: 87.85 F
16:29:33.433 -> Average Cal sensor Temp.: 100.08 Ambient Temp.: 87.89 F
16:29:34.985 -> Average Cal sensor Temp.: 100.07 Ambient Temp.: 87.85 F
16:29:36.572 -> Average Cal sensor Temp.: 100.03 Ambient Temp.: 87.82 F
16:29:38.119 -> Average Cal sensor Temp.: 100.18 Ambient Temp.: 87.82 F
16:29:39.675 -> Average Cal sensor Temp.: 99.48 Ambient Temp.: 87.89 F
16:29:41.260 -> Average Cal sensor Temp.: 99.97 Ambient Temp.: 87.85 F
16:29:42.812 -> Average Cal sensor Temp.: 100.03 Ambient Temp.: 87.82 F
16:29:44.369 -> Average Cal sensor Temp.: 100.07 Ambient Temp.: 87.78 F
16:29:45.926 -> Average Cal sensor Temp.: 100.54 Ambient Temp.: 87.82 F
16:29:47.510 -> Average Cal sensor Temp.: 100.48 Ambient Temp.: 87.75 F
16:29:49.061 -> Average Cal sensor Temp.: 100.04 Ambient Temp.: 87.82 F
16:29:50.652 -> Average Cal sensor Temp.: 99.82 Ambient Temp.: 87.78 F
16:29:52.201 -> Average Cal sensor Temp.: 100.29 Ambient Temp.: 87.82 F
16:29:53.752 -> Average Cal sensor Temp.: 100.01 Ambient Temp.: 87.82 F
16:29:55.342 -> Average Cal sensor Temp.: 100.12 Ambient Temp.: 87.75 F
16:29:56.900 -> Average Cal sensor Temp.: 100.32 Ambient Temp.: 87.78 F
16:29:58.450 -> Average Cal sensor Temp.: 100.19 Ambient Temp.: 87.75 F
16:30:00.041 -> Average Cal sensor Temp.: 99.99 Ambient Temp.: 87.78 F
16:30:01.593 -> Average Cal sensor Temp.: 100.06 Ambient Temp.: 87.78 F
16:30:03.145 -> Average Cal sensor Temp.: 100.31 Ambient Temp.: 87.78 F
16:30:04.729 -> Average Cal sensor Temp.: 100.53 Ambient Temp.: 87.78 F
16:30:06.284 -> Average Cal sensor Temp.: 100.31 Ambient Temp.: 87.78 F
16:30:07.835 -> Average Cal sensor Temp.: 100.15 Ambient Temp.: 87.71 F
16:30:09.421 -> Average Cal sensor Temp.: 100.08 Ambient Temp.: 87.78 F
16:30:10.971 -> Average Cal sensor Temp.: 100.13 Ambient Temp.: 87.75 F
16:30:12.527 -> Average Cal sensor Temp.: 100.09 Ambient Temp.: 87.71 F
16:30:14.113 -> Average Cal sensor Temp.: 100.02 Ambient Temp.: 87.71 F
16:30:15.663 -> Average Cal sensor Temp.: 100.22 Ambient Temp.: 87.71 F
16:30:17.220 -> Average Cal sensor Temp.: 99.95 Ambient Temp.: 87.67 F
16:30:18.778 -> Average Cal sensor Temp.: 99.88 Ambient Temp.: 87.60 F
16:30:20.367 -> Average Cal sensor Temp.: 100.04 Ambient Temp.: 87.57 F
16:30:21.919 -> Average Cal sensor Temp.: 100.30 Ambient Temp.: 87.60 F
16:30:23.473 -> Average Cal sensor Temp.: 100.44 Ambient Temp.: 87.67 F
16:30:25.062 -> Average Cal sensor Temp.: 100.06 Ambient Temp.: 87.57 F
16:30:26.617 -> Average Cal sensor Temp.: 100.49 Ambient Temp.: 87.57 F
16:30:28.170 -> Average Cal sensor Temp.: 100.16 Ambient Temp.: 87.53 F
16:30:29.759 -> Average Cal sensor Temp.: 100.51 Ambient Temp.: 87.53 F
16:30:31.321 -> Average Cal sensor Temp.: 100.49 Ambient Temp.: 87.53 F
16:30:32.873 -> Average Cal sensor Temp.: 100.18 Ambient Temp.: 87.57 F
16:30:34.431 -> Average Cal sensor Temp.: 100.08 Ambient Temp.: 87.53 F
16:30:36.014 -> Average Cal sensor Temp.: 99.99 Ambient Temp.: 87.57 F
16:30:37.572 -> Average Cal sensor Temp.: 99.78 Ambient Temp.: 87.57 F
16:30:39.128 -> Average Cal sensor Temp.: 99.77 Ambient Temp.: 87.49 F
16:30:40.689 -> Average Cal sensor Temp.: 100.02 Ambient Temp.: 87.53 F
16:30:42.248 -> Average Cal sensor Temp.: 100.08 Ambient Temp.: 87.53 F
16:30:43.833 -> Average Cal sensor Temp.: 99.64 Ambient Temp.: 87.49 F
16:30:45.387 -> Average Cal sensor Temp.: 100.27 Ambient Temp.: 87.49 F
16:30:46.948 -> Average Cal sensor Temp.: 100.31 Ambient Temp.: 87.46 F
16:30:48.532 -> Average Cal sensor Temp.: 100.09 Ambient Temp.: 87.46 F
16:30:50.089 -> Average Cal sensor Temp.: 99.90 Ambient Temp.: 87.39 F
16:30:51.649 -> Average Cal sensor Temp.: 99.82 Ambient Temp.: 87.42 F
16:30:53.207 -> Average Cal sensor Temp.: 99.87 Ambient Temp.: 87.39 F
16:30:54.760 -> Average Cal sensor Temp.: 100.06 Ambient Temp.: 87.39 F
16:30:56.322 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 87.35 F
16:30:57.913 -> Average Cal sensor Temp.: 99.71 Ambient Temp.: 87.35 F
16:30:59.474 -> Average Cal sensor Temp.: 100.09 Ambient Temp.: 87.31 F
16:31:01.034 -> Average Cal sensor Temp.: 100.26 Ambient Temp.: 87.21 F
16:31:02.588 -> Average Cal sensor Temp.: 100.42 Ambient Temp.: 87.28 F
16:31:04.144 -> Average Cal sensor Temp.: 100.01 Ambient Temp.: 87.21 F
16:31:05.730 -> Average Cal sensor Temp.: 99.70 Ambient Temp.: 87.28 F
16:31:07.281 -> Average Cal sensor Temp.: 100.11 Ambient Temp.: 87.21 F
16:31:08.835 -> Average Cal sensor Temp.: 100.29 Ambient Temp.: 87.21 F
16:31:10.424 -> Average Cal sensor Temp.: 100.03 Ambient Temp.: 87.17 F
16:31:11.978 -> Average Cal sensor Temp.: 99.79 Ambient Temp.: 87.17 F
16:31:13.536 -> Average Cal sensor Temp.: 99.69 Ambient Temp.: 87.13 F
16:31:15.089 -> Average Cal sensor Temp.: 100.21 Ambient Temp.: 87.13 F
16:31:16.645 -> Average Cal sensor Temp.: 100.49 Ambient Temp.: 87.17 F
16:31:18.233 -> Average Cal sensor Temp.: 100.12 Ambient Temp.: 87.17 F
16:31:19.791 -> Average Cal sensor Temp.: 99.86 Ambient Temp.: 87.17 F
16:31:21.338 -> Average Cal sensor Temp.: 99.96 Ambient Temp.: 87.13 F
16:31:22.919 -> Average Cal sensor Temp.: 100.13 Ambient Temp.: 87.17 F
16:31:24.474 -> Average Cal sensor Temp.: 99.94 Ambient Temp.: 87.13 F
16:31:26.028 -> Average Cal sensor Temp.: 99.60 Ambient Temp.: 87.10 F
16:31:27.584 -> Average Cal sensor Temp.: 100.11 Ambient Temp.: 87.10 F
16:31:29.170 -> Average Cal sensor Temp.: 100.47 Ambient Temp.: 87.10 F
16:31:30.724 -> Average Cal sensor Temp.: 100.56 Ambient Temp.: 87.10 F
16:31:32.274 -> Average Cal sensor Temp.: 100.44 Ambient Temp.: 87.03 F
16:31:33.862 -> Average Cal sensor Temp.: 100.20 Ambient Temp.: 87.10 F
16:31:35.409 -> Average Cal sensor Temp.: 100.38 Ambient Temp.: 87.10 F
16:31:36.994 -> Average Cal sensor Temp.: 100.46 Ambient Temp.: 87.13 F
16:31:38.540 -> Average Cal sensor Temp.: 100.87 Ambient Temp.: 87.17 F
16:31:40.121 -> Average Cal sensor Temp.: 100.03 Ambient Temp.: 87.17 F
16:31:41.675 -> Average Cal sensor Temp.: 100.50 Ambient Temp.: 87.28 F
16:31:43.253 -> Average Cal sensor Temp.: 100.14 Ambient Temp.: 87.31 F
16:31:44.799 -> Average Cal sensor Temp.: 100.49 Ambient Temp.: 87.31 F
16:31:46.388 -> Average Cal sensor Temp.: 100.76 Ambient Temp.: 87.28 F
16:31:47.937 -> Average Cal sensor Temp.: 100.76 Ambient Temp.: 87.35 F
16:31:49.489 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 87.42 F
16:31:51.077 -> Average Cal sensor Temp.: 100.72 Ambient Temp.: 87.42 F
16:31:52.625 -> Average Cal sensor Temp.: 100.00 Ambient Temp.: 87.46 F
16:31:54.217 -> Average Cal sensor Temp.: 99.73 Ambient Temp.: 87.39 F
16:31:55.771 -> Average Cal sensor Temp.: 99.89 Ambient Temp.: 87.42 F
16:31:57.321 -> Average Cal sensor Temp.: 100.43 Ambient Temp.: 87.46 F
16:31:58.875 -> Average Cal sensor Temp.: 100.54 Ambient Temp.: 87.42 F
16:32:00.461 -> Average Cal sensor Temp.: 100.42 Ambient Temp.: 87.42 F
16:32:02.013 -> Average Cal sensor Temp.: 100.70 Ambient Temp.: 87.42 F
16:32:03.599 -> Average Cal sensor Temp.: 100.61 Ambient Temp.: 87.46 F
16:32:05.145 -> Average Cal sensor Temp.: 100.93 Ambient Temp.: 87.46 F
16:32:06.699 -> Average Cal sensor Temp.: 100.64 Ambient Temp.: 87.42 F
16:32:08.278 -> Average Cal sensor Temp.: 100.57 Ambient Temp.: 87.46 F
16:32:09.837 -> Average Cal sensor Temp.: 100.31 Ambient Temp.: 87.42 F
16:32:11.426 -> Average Cal sensor Temp.: 100.47 Ambient Temp.: 87.42 F
16:32:12.985 -> Average Cal sensor Temp.: 100.41 Ambient Temp.: 87.39 F
16:32:14.532 -> Average Cal sensor Temp.: 100.46 Ambient Temp.: 87.39 F
16:32:16.119 -> Average Cal sensor Temp.: 100.67 Ambient Temp.: 87.42 F
16:32:17.675 -> Average Cal sensor Temp.: 100.76 Ambient Temp.: 87.42 F
16:32:19.230 -> Average Cal sensor Temp.: 100.73 Ambient Temp.: 87.42 F
16:32:20.818 -> Average Cal sensor Temp.: 100.75 Ambient Temp.: 87.49 F
16:32:22.369 -> Average Cal sensor Temp.: 100.64 Ambient Temp.: 87.39 F
16:32:23.924 -> Average Cal sensor Temp.: 100.18 Ambient Temp.: 87.39 F
16:32:25.503 -> Average Cal sensor Temp.: 100.52 Ambient Temp.: 87.42 F
16:32:27.055 -> Average Cal sensor Temp.: 101.32 Ambient Temp.: 87.42 F
16:32:28.610 -> Average Cal sensor Temp.: 100.65 Ambient Temp.: 87.46 F
16:32:30.187 -> Average Cal sensor Temp.: 100.69 Ambient Temp.: 87.46 F
16:32:31.745 -> Average Cal sensor Temp.: 101.06 Ambient Temp.: 87.39 F
16:32:33.337 -> Average Cal sensor Temp.: 101.14 Ambient Temp.: 87.39 F
16:32:34.890 -> Average Cal sensor Temp.: 101.20 Ambient Temp.: 87.46 F
16:32:36.447 -> Average Cal sensor Temp.: 100.52 Ambient Temp.: 87.42 F
16:32:38.001 -> Average Cal sensor Temp.: 100.81 Ambient Temp.: 87.46 F
16:32:39.589 -> Average Cal sensor Temp.: 100.96 Ambient Temp.: 87.46 F
16:32:41.145 -> Average Cal sensor Temp.: 100.82 Ambient Temp.: 87.46 F
16:32:42.699 -> Average Cal sensor Temp.: 101.03 Ambient Temp.: 87.46 F
16:32:44.284 -> Average Cal sensor Temp.: 100.96 Ambient Temp.: 87.46 F
16:32:45.835 -> Average Cal sensor Temp.: 100.68 Ambient Temp.: 87.42 F
16:32:47.424 -> Average Cal sensor Temp.: 100.71 Ambient Temp.: 87.46 F
16:32:48.983 -> Average Cal sensor Temp.: 100.82 Ambient Temp.: 87.46 F
16:32:50.533 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 87.53 F
16:32:52.118 -> Average Cal sensor Temp.: 100.64 Ambient Temp.: 87.46 F
16:32:53.667 -> Average Cal sensor Temp.: 100.65 Ambient Temp.: 87.49 F
16:32:55.222 -> Average Cal sensor Temp.: 101.06 Ambient Temp.: 87.49 F
16:32:56.807 -> Average Cal sensor Temp.: 101.12 Ambient Temp.: 87.49 F
16:32:58.364 -> Average Cal sensor Temp.: 101.06 Ambient Temp.: 87.53 F
16:32:59.919 -> Average Cal sensor Temp.: 101.02 Ambient Temp.: 87.57 F
16:33:01.476 -> Average Cal sensor Temp.: 101.25 Ambient Temp.: 87.53 F
16:33:03.058 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 87.60 F
16:33:04.613 -> Average Cal sensor Temp.: 101.18 Ambient Temp.: 87.67 F
16:33:06.169 -> Average Cal sensor Temp.: 101.37 Ambient Temp.: 87.60 F
16:33:07.755 -> Average Cal sensor Temp.: 101.19 Ambient Temp.: 87.75 F
16:33:09.313 -> Average Cal sensor Temp.: 100.87 Ambient Temp.: 87.71 F
16:33:10.864 -> Average Cal sensor Temp.: 100.83 Ambient Temp.: 87.71 F
16:33:12.419 -> Average Cal sensor Temp.: 100.96 Ambient Temp.: 87.75 F
16:33:14.007 -> Average Cal sensor Temp.: 100.95 Ambient Temp.: 87.75 F
16:33:15.558 -> Average Cal sensor Temp.: 101.33 Ambient Temp.: 87.78 F
16:33:17.141 -> Average Cal sensor Temp.: 101.18 Ambient Temp.: 87.71 F
16:33:18.688 -> Average Cal sensor Temp.: 101.33 Ambient Temp.: 87.85 F
16:33:20.244 -> Average Cal sensor Temp.: 101.48 Ambient Temp.: 87.85 F
16:33:21.831 -> Average Cal sensor Temp.: 101.28 Ambient Temp.: 87.78 F
16:33:23.385 -> Average Cal sensor Temp.: 101.31 Ambient Temp.: 87.85 F
16:33:24.937 -> Average Cal sensor Temp.: 101.22 Ambient Temp.: 87.85 F
16:33:26.496 -> Average Cal sensor Temp.: 101.00 Ambient Temp.: 87.89 F
16:33:28.082 -> Average Cal sensor Temp.: 100.98 Ambient Temp.: 87.89 F
16:33:29.636 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 87.89 F
16:33:31.193 -> Average Cal sensor Temp.: 101.14 Ambient Temp.: 87.82 F
16:33:32.780 -> Average Cal sensor Temp.: 101.31 Ambient Temp.: 87.85 F
16:33:34.328 -> Average Cal sensor Temp.: 101.35 Ambient Temp.: 87.89 F
16:33:35.886 -> Average Cal sensor Temp.: 101.18 Ambient Temp.: 87.96 F
16:33:37.443 -> Average Cal sensor Temp.: 101.20 Ambient Temp.: 87.93 F
16:33:39.026 -> Average Cal sensor Temp.: 100.79 Ambient Temp.: 87.89 F
16:33:40.584 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 87.89 F
16:33:42.133 -> Average Cal sensor Temp.: 101.16 Ambient Temp.: 87.96 F
16:33:43.724 -> Average Cal sensor Temp.: 100.98 Ambient Temp.: 87.93 F
16:33:45.279 -> Average Cal sensor Temp.: 101.27 Ambient Temp.: 87.93 F
16:33:46.834 -> Average Cal sensor Temp.: 101.20 Ambient Temp.: 87.93 F
16:33:48.392 -> Average Cal sensor Temp.: 100.92 Ambient Temp.: 88.00 F
16:33:49.979 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 88.00 F
16:33:51.528 -> Average Cal sensor Temp.: 101.04 Ambient Temp.: 87.96 F
16:33:53.111 -> Average Cal sensor Temp.: 101.15 Ambient Temp.: 87.96 F
16:33:54.664 -> Average Cal sensor Temp.: 101.22 Ambient Temp.: 88.07 F
16:33:56.218 -> Average Cal sensor Temp.: 101.29 Ambient Temp.: 88.00 F
16:33:57.808 -> Average Cal sensor Temp.: 101.40 Ambient Temp.: 88.11 F
16:33:59.359 -> Average Cal sensor Temp.: 101.13 Ambient Temp.: 88.07 F
16:34:00.910 -> Average Cal sensor Temp.: 101.18 Ambient Temp.: 88.07 F
16:34:02.496 -> Average Cal sensor Temp.: 101.41 Ambient Temp.: 88.14 F
16:34:04.048 -> Average Cal sensor Temp.: 101.48 Ambient Temp.: 88.14 F
16:34:05.600 -> Average Cal sensor Temp.: 101.34 Ambient Temp.: 88.14 F
16:34:07.158 -> Average Cal sensor Temp.: 101.33 Ambient Temp.: 88.11 F
16:34:08.751 -> Average Cal sensor Temp.: 101.47 Ambient Temp.: 88.14 F
16:34:10.304 -> Average Cal sensor Temp.: 101.17 Ambient Temp.: 88.18 F
16:34:11.849 -> Average Cal sensor Temp.: 101.18 Ambient Temp.: 88.25 F
16:34:13.430 -> Average Cal sensor Temp.: 101.12 Ambient Temp.: 88.18 F
16:34:14.977 -> Average Cal sensor Temp.: 101.32 Ambient Temp.: 88.21 F
16:34:16.565 -> Average Cal sensor Temp.: 101.44 Ambient Temp.: 88.25 F
16:34:18.117 -> Average Cal sensor Temp.: 101.59 Ambient Temp.: 88.25 F
16:34:19.673 -> Average Cal sensor Temp.: 101.34 Ambient Temp.: 88.29 F
16:34:21.227 -> Average Cal sensor Temp.: 101.24 Ambient Temp.: 88.32 F
16:34:22.810 -> Average Cal sensor Temp.: 101.57 Ambient Temp.: 88.29 F
16:34:24.362 -> Average Cal sensor Temp.: 101.48 Ambient Temp.: 88.32 F
16:34:25.950 -> Average Cal sensor Temp.: 101.60 Ambient Temp.: 88.29 F
16:34:27.500 -> Average Cal sensor Temp.: 101.55 Ambient Temp.: 88.36 F
16:34:29.053 -> Average Cal sensor Temp.: 101.40 Ambient Temp.: 88.32 F
16:34:30.638 -> Average Cal sensor Temp.: 101.44 Ambient Temp.: 88.36 F
16:34:32.186 -> Average Cal sensor Temp.: 101.66 Ambient Temp.: 88.36 F
16:34:33.743 -> Average Cal sensor Temp.: 101.67 Ambient Temp.: 88.36 F
16:34:35.321 -> Average Cal sensor Temp.: 101.74 Ambient Temp.: 88.39 F
16:34:36.871 -> Average Cal sensor Temp.: 101.62 Ambient Temp.: 88.39 F
16:34:38.453 -> Average Cal sensor Temp.: 101.45 Ambient Temp.: 88.47 F
16:34:40.005 -> Average Cal sensor Temp.: 101.76 Ambient Temp.: 88.50 F
16:34:41.594 -> Average Cal sensor Temp.: 101.74 Ambient Temp.: 88.50 F
16:34:43.137 -> Average Cal sensor Temp.: 101.75 Ambient Temp.: 88.47 F
16:34:44.719 -> Average Cal sensor Temp.: 101.44 Ambient Temp.: 88.54 F
16:34:46.267 -> Average Cal sensor Temp.: 101.55 Ambient Temp.: 88.50 F
16:34:47.849 -> Average Cal sensor Temp.: 101.52 Ambient Temp.: 88.54 F
16:34:49.399 -> Average Cal sensor Temp.: 101.51 Ambient Temp.: 88.50 F
16:34:50.946 -> Average Cal sensor Temp.: 101.71 Ambient Temp.: 88.54 F
16:34:52.525 -> Average Cal sensor Temp.: 101.46 Ambient Temp.: 88.61 F
16:34:54.080 -> Average Cal sensor Temp.: 101.72 Ambient Temp.: 88.57 F
16:34:55.638 -> Average Cal sensor Temp.: 101.47 Ambient Temp.: 88.57 F
16:34:57.200 -> Average Cal sensor Temp.: 101.53 Ambient Temp.: 88.57 F
16:34:58.789 -> Average Cal sensor Temp.: 101.59 Ambient Temp.: 88.61 F
16:35:00.344 -> Average Cal sensor Temp.: 101.89 Ambient Temp.: 88.65 F
16:35:01.900 -> Average Cal sensor Temp.: 101.82 Ambient Temp.: 88.65 F
16:35:03.488 -> Average Cal sensor Temp.: 101.83 Ambient Temp.: 88.68 F
16:35:05.037 -> Average Cal sensor Temp.: 101.83 Ambient Temp.: 88.61 F
16:35:06.593 -> Average Cal sensor Temp.: 101.57 Ambient Temp.: 88.72 F
16:35:08.177 -> Average Cal sensor Temp.: 101.51 Ambient Temp.: 88.68 F
16:35:09.731 -> Average Cal sensor Temp.: 101.50 Ambient Temp.: 88.72 F
16:35:11.283 -> Average Cal sensor Temp.: 101.65 Ambient Temp.: 88.65 F
16:35:12.865 -> Average Cal sensor Temp.: 101.69 Ambient Temp.: 88.72 F
16:35:14.419 -> Average Cal sensor Temp.: 101.82 Ambient Temp.: 88.86 F
16:35:15.980 -> Average Cal sensor Temp.: 101.15 Ambient Temp.: 88.75 F
16:35:17.537 -> Average Cal sensor Temp.: 101.11 Ambient Temp.: 88.79 F
16:35:19.125 -> Average Cal sensor Temp.: 101.18 Ambient Temp.: 88.75 F
16:35:20.681 -> Average Cal sensor Temp.: 101.28 Ambient Temp.: 88.75 F
16:35:22.239 -> Average Cal sensor Temp.: 101.11 Ambient Temp.: 88.75 F
16:35:23.794 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 88.79 F
16:35:25.379 -> Average Cal sensor Temp.: 101.14 Ambient Temp.: 88.79 F
16:35:26.929 -> Average Cal sensor Temp.: 100.73 Ambient Temp.: 88.79 F
16:35:28.484 -> Average Cal sensor Temp.: 100.98 Ambient Temp.: 88.86 F
16:35:30.069 -> Average Cal sensor Temp.: 100.96 Ambient Temp.: 88.79 F
16:35:31.626 -> Average Cal sensor Temp.: 101.09 Ambient Temp.: 88.75 F
16:35:33.178 -> Average Cal sensor Temp.: 101.16 Ambient Temp.: 88.79 F
16:35:34.755 -> Average Cal sensor Temp.: 100.78 Ambient Temp.: 88.75 F
16:35:36.304 -> Average Cal sensor Temp.: 100.99 Ambient Temp.: 88.75 F
16:35:37.889 -> Average Cal sensor Temp.: 100.88 Ambient Temp.: 88.65 F
16:35:39.444 -> Average Cal sensor Temp.: 101.01 Ambient Temp.: 88.72 F
16:35:41.030 -> Average Cal sensor Temp.: 101.03 Ambient Temp.: 88.68 F
16:35:42.585 -> Average Cal sensor Temp.: 101.27 Ambient Temp.: 88.65 F
16:35:44.136 -> Average Cal sensor Temp.: 101.25 Ambient Temp.: 88.65 F
16:35:45.725 -> Average Cal sensor Temp.: 101.29 Ambient Temp.: 88.65 F
16:35:47.285 -> Average Cal sensor Temp.: 101.19 Ambient Temp.: 88.65 F
16:35:48.829 -> Average Cal sensor Temp.: 101.06 Ambient Temp.: 88.61 F
16:35:50.416 -> Average Cal sensor Temp.: 101.29 Ambient Temp.: 88.61 F
16:35:51.967 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 88.68 F
16:35:53.518 -> Average Cal sensor Temp.: 101.33 Ambient Temp.: 88.65 F
16:35:55.104 -> Average Cal sensor Temp.: 101.29 Ambient Temp.: 88.54 F
16:35:56.657 -> Average Cal sensor Temp.: 100.86 Ambient Temp.: 88.54 F
16:35:58.242 -> Average Cal sensor Temp.: 101.14 Ambient Temp.: 88.57 F
16:35:59.791 -> Average Cal sensor Temp.: 100.78 Ambient Temp.: 88.57 F
16:36:01.342 -> Average Cal sensor Temp.: 100.31 Ambient Temp.: 88.54 F
16:36:02.925 -> Average Cal sensor Temp.: 100.76 Ambient Temp.: 88.50 F
16:36:04.474 -> Average Cal sensor Temp.: 101.22 Ambient Temp.: 88.50 F
16:36:06.063 -> Average Cal sensor Temp.: 101.23 Ambient Temp.: 88.50 F
16:36:07.621 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 88.47 F
16:36:09.169 -> Average Cal sensor Temp.: 101.15 Ambient Temp.: 88.50 F
16:36:10.727 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 88.50 F
16:36:12.313 -> Average Cal sensor Temp.: 101.23 Ambient Temp.: 88.47 F
16:36:13.863 -> Average Cal sensor Temp.: 100.88 Ambient Temp.: 88.47 F
16:36:15.450 -> Average Cal sensor Temp.: 100.85 Ambient Temp.: 88.39 F
16:36:16.997 -> Average Cal sensor Temp.: 100.31 Ambient Temp.: 88.32 F
16:36:18.579 -> Average Cal sensor Temp.: 100.26 Ambient Temp.: 88.32 F
16:36:20.134 -> Average Cal sensor Temp.: 100.47 Ambient Temp.: 88.29 F
16:36:21.692 -> Average Cal sensor Temp.: 100.94 Ambient Temp.: 88.29 F
16:36:23.247 -> Average Cal sensor Temp.: 101.02 Ambient Temp.: 88.25 F
16:36:24.829 -> Average Cal sensor Temp.: 101.28 Ambient Temp.: 88.29 F
16:36:26.384 -> Average Cal sensor Temp.: 101.35 Ambient Temp.: 88.29 F
16:36:27.969 -> Average Cal sensor Temp.: 100.85 Ambient Temp.: 88.25 F
16:36:29.527 -> Average Cal sensor Temp.: 100.98 Ambient Temp.: 88.29 F
16:36:31.083 -> Average Cal sensor Temp.: 101.50 Ambient Temp.: 88.21 F
16:36:32.641 -> Average Cal sensor Temp.: 101.58 Ambient Temp.: 88.21 F
16:36:34.230 -> Average Cal sensor Temp.: 100.94 Ambient Temp.: 88.21 F
16:36:35.787 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 88.18 F
16:36:37.340 -> Average Cal sensor Temp.: 101.12 Ambient Temp.: 88.14 F
16:36:38.895 -> Average Cal sensor Temp.: 101.47 Ambient Temp.: 88.18 F
16:36:40.483 -> Average Cal sensor Temp.: 101.46 Ambient Temp.: 88.14 F
16:36:42.039 -> Average Cal sensor Temp.: 101.37 Ambient Temp.: 88.18 F
16:36:43.590 -> Average Cal sensor Temp.: 101.34 Ambient Temp.: 88.14 F
16:36:45.173 -> Average Cal sensor Temp.: 101.70 Ambient Temp.: 88.18 F
16:36:46.725 -> Average Cal sensor Temp.: 100.88 Ambient Temp.: 88.14 F
16:36:48.278 -> Average Cal sensor Temp.: 101.24 Ambient Temp.: 88.14 F
16:36:49.864 -> Average Cal sensor Temp.: 101.64 Ambient Temp.: 88.00 F
16:36:51.417 -> Average Cal sensor Temp.: 101.81 Ambient Temp.: 88.07 F
16:36:52.973 -> Average Cal sensor Temp.: 101.27 Ambient Temp.: 88.11 F
16:36:54.552 -> Average Cal sensor Temp.: 101.26 Ambient Temp.: 88.00 F
16:36:56.109 -> Average Cal sensor Temp.: 101.48 Ambient Temp.: 88.07 F
16:36:57.663 -> Average Cal sensor Temp.: 101.18 Ambient Temp.: 88.14 F
16:36:59.257 -> Average Cal sensor Temp.: 101.17 Ambient Temp.: 88.07 F
16:37:00.811 -> Average Cal sensor Temp.: 100.99 Ambient Temp.: 88.00 F
16:37:02.360 -> Average Cal sensor Temp.: 100.72 Ambient Temp.: 87.96 F
16:37:03.948 -> Average Cal sensor Temp.: 101.18 Ambient Temp.: 87.96 F
16:37:05.507 -> Average Cal sensor Temp.: 101.17 Ambient Temp.: 88.00 F
16:37:07.063 -> Average Cal sensor Temp.: 100.57 Ambient Temp.: 87.89 F
16:37:08.615 -> Average Cal sensor Temp.: 100.43 Ambient Temp.: 87.85 F
16:37:10.199 -> Average Cal sensor Temp.: 100.99 Ambient Temp.: 87.85 F
16:37:11.751 -> Average Cal sensor Temp.: 101.04 Ambient Temp.: 87.82 F
16:37:13.342 -> Average Cal sensor Temp.: 100.71 Ambient Temp.: 87.78 F
16:37:14.892 -> Average Cal sensor Temp.: 101.03 Ambient Temp.: 87.78 F
16:37:16.452 -> Average Cal sensor Temp.: 101.37 Ambient Temp.: 87.82 F
16:37:18.002 -> Average Cal sensor Temp.: 101.42 Ambient Temp.: 87.78 F
16:37:19.594 -> Average Cal sensor Temp.: 101.50 Ambient Temp.: 87.82 F
16:37:21.147 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 87.78 F
16:37:22.703 -> Average Cal sensor Temp.: 101.43 Ambient Temp.: 87.75 F
16:37:24.261 -> Average Cal sensor Temp.: 100.71 Ambient Temp.: 87.78 F
16:37:25.851 -> Average Cal sensor Temp.: 101.34 Ambient Temp.: 87.78 F
16:37:27.414 -> Average Cal sensor Temp.: 101.42 Ambient Temp.: 87.71 F
16:37:28.967 -> Average Cal sensor Temp.: 101.15 Ambient Temp.: 87.78 F
16:37:30.521 -> Average Cal sensor Temp.: 100.95 Ambient Temp.: 87.71 F
16:37:32.077 -> Average Cal sensor Temp.: 101.16 Ambient Temp.: 87.78 F
16:37:33.661 -> Average Cal sensor Temp.: 101.19 Ambient Temp.: 87.75 F
16:37:35.209 -> Average Cal sensor Temp.: 101.55 Ambient Temp.: 87.75 F
16:37:36.758 -> Average Cal sensor Temp.: 101.35 Ambient Temp.: 87.75 F
16:37:38.339 -> Average Cal sensor Temp.: 101.23 Ambient Temp.: 87.75 F
16:37:39.891 -> Average Cal sensor Temp.: 101.00 Ambient Temp.: 87.71 F
16:37:41.450 -> Average Cal sensor Temp.: 101.12 Ambient Temp.: 87.75 F
16:37:43.036 -> Average Cal sensor Temp.: 100.97 Ambient Temp.: 87.75 F
16:37:44.589 -> Average Cal sensor Temp.: 101.32 Ambient Temp.: 87.75 F
16:37:46.139 -> Average Cal sensor Temp.: 101.33 Ambient Temp.: 87.71 F
16:37:47.725 -> Average Cal sensor Temp.: 101.15 Ambient Temp.: 87.67 F
16:37:49.278 -> Average Cal sensor Temp.: 100.85 Ambient Temp.: 87.71 F
16:37:50.834 -> Average Cal sensor Temp.: 101.15 Ambient Temp.: 87.67 F
16:37:52.390 -> Average Cal sensor Temp.: 100.56 Ambient Temp.: 87.67 F
16:37:53.980 -> Average Cal sensor Temp.: 100.70 Ambient Temp.: 87.67 F
16:37:55.523 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 87.60 F
16:37:57.078 -> Average Cal sensor Temp.: 101.08 Ambient Temp.: 87.67 F
16:37:58.665 -> Average Cal sensor Temp.: 101.21 Ambient Temp.: 87.57 F
16:38:00.214 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 87.60 F
16:38:01.798 -> Average Cal sensor Temp.: 101.43 Ambient Temp.: 87.67 F
16:38:03.347 -> Average Cal sensor Temp.: 101.39 Ambient Temp.: 87.60 F
16:38:04.895 -> Average Cal sensor Temp.: 100.95 Ambient Temp.: 87.53 F
16:38:06.478 -> Average Cal sensor Temp.: 100.76 Ambient Temp.: 87.49 F
16:38:08.033 -> Average Cal sensor Temp.: 101.13 Ambient Temp.: 87.53 F
16:38:09.593 -> Average Cal sensor Temp.: 101.22 Ambient Temp.: 87.53 F
16:38:11.152 -> Average Cal sensor Temp.: 101.13 Ambient Temp.: 87.49 F
16:38:12.746 -> Average Cal sensor Temp.: 101.15 Ambient Temp.: 87.49 F
16:38:14.301 -> Average Cal sensor Temp.: 101.13 Ambient Temp.: 87.53 F
16:38:15.855 -> Average Cal sensor Temp.: 101.39 Ambient Temp.: 87.46 F
16:38:17.409 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 87.46 F
16:38:18.968 -> Average Cal sensor Temp.: 101.15 Ambient Temp.: 87.49 F
16:38:20.553 -> Average Cal sensor Temp.: 101.24 Ambient Temp.: 87.46 F
16:38:22.104 -> Average Cal sensor Temp.: 101.19 Ambient Temp.: 87.46 F
16:38:23.655 -> Average Cal sensor Temp.: 101.24 Ambient Temp.: 87.42 F
16:38:25.242 -> Average Cal sensor Temp.: 101.25 Ambient Temp.: 87.46 F
16:38:26.798 -> Average Cal sensor Temp.: 100.90 Ambient Temp.: 87.46 F
16:38:28.354 -> Average Cal sensor Temp.: 100.59 Ambient Temp.: 87.42 F
16:38:29.945 -> Average Cal sensor Temp.: 100.96 Ambient Temp.: 87.46 F
16:38:31.499 -> Average Cal sensor Temp.: 100.88 Ambient Temp.: 87.42 F
16:38:33.055 -> Average Cal sensor Temp.: 101.23 Ambient Temp.: 87.39 F
16:38:34.642 -> Average Cal sensor Temp.: 101.55 Ambient Temp.: 87.46 F
16:38:36.190 -> Average Cal sensor Temp.: 101.27 Ambient Temp.: 87.42 F
16:38:37.741 -> Average Cal sensor Temp.: 101.39 Ambient Temp.: 87.46 F
16:38:39.329 -> Average Cal sensor Temp.: 101.79 Ambient Temp.: 87.46 F
16:38:40.880 -> Average Cal sensor Temp.: 102.03 Ambient Temp.: 87.49 F
16:38:42.430 -> Average Cal sensor Temp.: 102.02 Ambient Temp.: 87.46 F
16:38:44.019 -> Average Cal sensor Temp.: 101.66 Ambient Temp.: 87.46 F
16:38:45.568 -> Average Cal sensor Temp.: 102.04 Ambient Temp.: 87.49 F
16:38:47.129 -> Average Cal sensor Temp.: 102.04 Ambient Temp.: 87.49 F
16:38:48.689 -> Average Cal sensor Temp.: 101.87 Ambient Temp.: 87.57 F
16:38:50.276 -> Average Cal sensor Temp.: 101.93 Ambient Temp.: 87.53 F
16:38:51.834 -> Average Cal sensor Temp.: 101.96 Ambient Temp.: 87.60 F
16:38:53.391 -> Average Cal sensor Temp.: 101.77 Ambient Temp.: 87.67 F
16:38:54.944 -> Average Cal sensor Temp.: 101.63 Ambient Temp.: 87.67 F
16:38:56.532 -> Average Cal sensor Temp.: 101.52 Ambient Temp.: 87.71 F
16:38:58.088 -> Average Cal sensor Temp.: 101.14 Ambient Temp.: 87.67 F
16:38:59.643 -> Average Cal sensor Temp.: 101.05 Ambient Temp.: 87.71 F
16:39:01.199 -> Average Cal sensor Temp.: 101.57 Ambient Temp.: 87.71 F
16:39:02.756 -> Average Cal sensor Temp.: 101.57 Ambient Temp.: 87.75 F
16:39:04.346 -> Average Cal sensor Temp.: 101.52 Ambient Temp.: 87.78 F
16:39:05.900 -> Average Cal sensor Temp.: 101.62 Ambient Temp.: 87.78 F
16:39:07.454 -> Average Cal sensor Temp.: 101.19 Ambient Temp.: 87.82 F
16:39:09.008 -> Average Cal sensor Temp.: 101.39 Ambient Temp.: 87.82 F
16:39:10.596 -> Average Cal sensor Temp.: 101.40 Ambient Temp.: 87.75 F
16:39:12.152 -> Average Cal sensor Temp.: 101.20 Ambient Temp.: 87.78 F
16:39:13.702 -> Average Cal sensor Temp.: 101.23 Ambient Temp.: 87.75 F
16:39:15.288 -> Average Cal sensor Temp.: 101.21 Ambient Temp.: 87.71 F
16:39:16.843 -> Average Cal sensor Temp.: 101.69 Ambient Temp.: 87.78 F
16:39:18.400 -> Average Cal sensor Temp.: 101.81 Ambient Temp.: 87.82 F
16:39:19.953 -> Average Cal sensor Temp.: 101.84 Ambient Temp.: 87.82 F
16:39:21.517 -> Average Cal sensor Temp.: 101.66 Ambient Temp.: 87.85 F
16:39:23.107 -> Average Cal sensor Temp.: 101.80 Ambient Temp.: 87.85 F
16:39:24.660 -> Average Cal sensor Temp.: 101.84 Ambient Temp.: 87.85 F
16:39:26.216 -> Average Cal sensor Temp.: 101.88 Ambient Temp.: 87.89 F
16:39:27.800 -> Average Cal sensor Temp.: 101.88 Ambient Temp.: 87.89 F
16:39:29.353 -> Average Cal sensor Temp.: 102.12 Ambient Temp.: 87.89 F
16:39:30.905 -> Average Cal sensor Temp.: 101.97 Ambient Temp.: 87.89 F
16:39:32.465 -> Average Cal sensor Temp.: 101.35 Ambient Temp.: 87.89 F
16:39:34.055 -> Average Cal sensor Temp.: 101.41 Ambient Temp.: 87.96 F
16:39:35.607 -> Average Cal sensor Temp.: 101.58 Ambient Temp.: 87.96 F
16:39:37.167 -> Average Cal sensor Temp.: 101.58 Ambient Temp.: 87.96 F
16:39:38.719 -> Average Cal sensor Temp.: 102.11 Ambient Temp.: 87.93 F
16:39:40.304 -> Average Cal sensor Temp.: 102.16 Ambient Temp.: 88.07 F
16:39:41.862 -> Average Cal sensor Temp.: 101.85 Ambient Temp.: 88.07 F
16:39:43.419 -> Average Cal sensor Temp.: 101.98 Ambient Temp.: 88.07 F
16:39:45.004 -> Average Cal sensor Temp.: 102.16 Ambient Temp.: 88.07 F
16:39:46.557 -> Average Cal sensor Temp.: 102.10 Ambient Temp.: 88.11 F
16:39:48.130 -> Average Cal sensor Temp.: 102.02 Ambient Temp.: 88.14 F
16:39:49.678 -> Average Cal sensor Temp.: 102.14 Ambient Temp.: 88.14 F
16:39:51.230 -> Average Cal sensor Temp.: 102.25 Ambient Temp.: 88.14 F
16:39:52.817 -> Average Cal sensor Temp.: 102.32 Ambient Temp.: 88.21 F
16:39:54.368 -> Average Cal sensor Temp.: 102.15 Ambient Temp.: 88.18 F
16:39:55.921 -> Average Cal sensor Temp.: 101.91 Ambient Temp.: 88.18 F
16:39:57.506 -> Average Cal sensor Temp.: 102.07 Ambient Temp.: 88.25 F
16:39:59.052 -> Average Cal sensor Temp.: 101.94 Ambient Temp.: 88.21 F
16:40:00.639 -> Average Cal sensor Temp.: 102.18 Ambient Temp.: 88.21 F
16:40:02.192 -> Average Cal sensor Temp.: 102.46 Ambient Temp.: 88.25 F
16:40:03.742 -> Average Cal sensor Temp.: 102.34 Ambient Temp.: 88.25 F
16:40:05.325 -> Average Cal sensor Temp.: 102.29 Ambient Temp.: 88.25 F
16:40:06.877 -> Average Cal sensor Temp.: 102.08 Ambient Temp.: 88.25 F
16:40:08.451 -> Average Cal sensor Temp.: 102.45 Ambient Temp.: 88.32 F
16:40:10.003 -> Average Cal sensor Temp.: 102.10 Ambient Temp.: 88.32 F
16:40:11.582 -> Average Cal sensor Temp.: 102.11 Ambient Temp.: 88.32 F
16:40:13.131 -> Average Cal sensor Temp.: 102.15 Ambient Temp.: 88.32 F
16:40:14.714 -> Average Cal sensor Temp.: 102.41 Ambient Temp.: 88.29 F
16:40:16.268 -> Average Cal sensor Temp.: 102.28 Ambient Temp.: 88.32 F
16:40:17.823 -> Average Cal sensor Temp.: 102.61 Ambient Temp.: 88.36 F
16:40:19.415 -> Average Cal sensor Temp.: 102.15 Ambient Temp.: 88.36 F
16:40:20.965 -> Average Cal sensor Temp.: 102.22 Ambient Temp.: 88.39 F
16:40:22.525 -> Average Cal sensor Temp.: 102.06 Ambient Temp.: 88.47 F
16:40:24.077 -> Average Cal sensor Temp.: 102.17 Ambient Temp.: 88.47 F
16:40:25.668 -> Average Cal sensor Temp.: 102.12 Ambient Temp.: 88.39 F
16:40:27.227 -> Average Cal sensor Temp.: 101.88 Ambient Temp.: 88.50 F
16:40:28.782 -> Average Cal sensor Temp.: 102.01 Ambient Temp.: 88.54 F
16:40:30.337 -> Average Cal sensor Temp.: 101.84 Ambient Temp.: 88.50 F
16:40:31.930 -> Average Cal sensor Temp.: 101.86 Ambient Temp.: 88.47 F
16:40:33.482 -> Average Cal sensor Temp.: 101.60 Ambient Temp.: 88.47 F
16:40:35.031 -> Average Cal sensor Temp.: 102.14 Ambient Temp.: 88.50 F
16:40:36.617 -> Average Cal sensor Temp.: 102.71 Ambient Temp.: 88.50 F
16:40:38.177 -> Average Cal sensor Temp.: 102.24 Ambient Temp.: 88.47 F
16:40:39.730 -> Average Cal sensor Temp.: 102.36 Ambient Temp.: 88.50 F
16:40:41.312 -> Average Cal sensor Temp.: 102.14 Ambient Temp.: 88.47 F
16:40:42.869 -> Average Cal sensor Temp.: 102.42 Ambient Temp.: 88.47 F
16:40:44.425 -> Average Cal sensor Temp.: 102.08 Ambient Temp.: 88.50 F
16:40:45.981 -> Average Cal sensor Temp.: 102.27 Ambient Temp.: 88.50 F
16:40:47.562 -> Average Cal sensor Temp.: 102.33 Ambient Temp.: 88.57 F
16:40:49.113 -> Average Cal sensor Temp.: 102.51 Ambient Temp.: 88.57 F
16:40:50.701 -> Average Cal sensor Temp.: 102.23 Ambient Temp.: 88.57 F
16:40:52.255 -> Average Cal sensor Temp.: 102.29 Ambient Temp.: 88.61 F
16:40:53.805 -> Average Cal sensor Temp.: 102.62 Ambient Temp.: 88.61 F
16:40:55.390 -> Average Cal sensor Temp.: 102.33 Ambient Temp.: 88.61 F
16:40:56.949 -> Average Cal sensor Temp.: 102.16 Ambient Temp.: 88.57 F
16:40:58.497 -> Average Cal sensor Temp.: 102.35 Ambient Temp.: 88.61 F
16:41:00.077 -> Average Cal sensor Temp.: 102.60 Ambient Temp.: 88.68 F
16:41:01.635 -> Average Cal sensor Temp.: 102.75 Ambient Temp.: 88.68 F
16:41:03.192 -> Average Cal sensor Temp.: 102.51 Ambient Temp.: 88.68 F
16:41:04.777 -> Average Cal sensor Temp.: 102.42 Ambient Temp.: 88.72 F
16:41:06.336 -> Average Cal sensor Temp.: 102.34 Ambient Temp.: 88.72 F
16:41:07.881 -> Average Cal sensor Temp.: 102.37 Ambient Temp.: 88.72 F
16:41:09.470 -> Average Cal sensor Temp.: 102.56 Ambient Temp.: 88.75 F
16:41:11.029 -> Average Cal sensor Temp.: 102.54 Ambient Temp.: 88.75 F
16:41:12.582 -> Average Cal sensor Temp.: 102.50 Ambient Temp.: 88.86 F
16:41:14.164 -> Average Cal sensor Temp.: 102.70 Ambient Temp.: 88.79 F
16:41:15.719 -> Average Cal sensor Temp.: 102.86 Ambient Temp.: 88.90 F
16:41:17.299 -> Average Cal sensor Temp.: 102.70 Ambient Temp.: 88.86 F
16:41:18.850 -> Average Cal sensor Temp.: 102.76 Ambient Temp.: 88.86 F
16:41:20.403 -> Average Cal sensor Temp.: 102.93 Ambient Temp.: 88.97 F
16:41:21.995 -> Average Cal sensor Temp.: 102.66 Ambient Temp.: 88.86 F
16:41:23.547 -> Average Cal sensor Temp.: 102.68 Ambient Temp.: 88.93 F
16:41:25.101 -> Average Cal sensor Temp.: 102.54 Ambient Temp.: 88.90 F
16:41:26.693 -> Average Cal sensor Temp.: 102.83 Ambient Temp.: 89.01 F
16:41:28.241 -> Average Cal sensor Temp.: 102.40 Ambient Temp.: 88.97 F
16:41:29.796 -> Average Cal sensor Temp.: 102.38 Ambient Temp.: 89.04 F
16:41:31.382 -> Average Cal sensor Temp.: 102.13 Ambient Temp.: 88.97 F
16:41:32.936 -> Average Cal sensor Temp.: 102.51 Ambient Temp.: 89.01 F
16:41:34.490 -> Average Cal sensor Temp.: 102.11 Ambient Temp.: 89.04 F
16:41:36.075 -> Average Cal sensor Temp.: 102.23 Ambient Temp.: 89.04 F
16:41:37.631 -> Average Cal sensor Temp.: 102.17 Ambient Temp.: 88.97 F
16:41:39.185 -> Average Cal sensor Temp.: 102.33 Ambient Temp.: 89.01 F
16:41:40.741 -> Average Cal sensor Temp.: 102.28 Ambient Temp.: 89.04 F
16:41:42.325 -> Average Cal sensor Temp.: 102.38 Ambient Temp.: 89.04 F
16:41:43.880 -> Average Cal sensor Temp.: 102.53 Ambient Temp.: 89.04 F
16:41:45.464 -> Average Cal sensor Temp.: 102.44 Ambient Temp.: 89.08 F
16:41:47.013 -> Average Cal sensor Temp.: 102.33 Ambient Temp.: 89.08 F
16:41:48.569 -> Average Cal sensor Temp.: 102.22 Ambient Temp.: 89.08 F
16:41:50.154 -> Average Cal sensor Temp.: 102.02 Ambient Temp.: 89.08 F
16:41:51.711 -> Average Cal sensor Temp.: 102.03 Ambient Temp.: 89.08 F
16:41:53.294 -> Average Cal sensor Temp.: 102.15 Ambient Temp.: 89.08 F
16:41:54.849 -> Average Cal sensor Temp.: 101.84 Ambient Temp.: 89.04 F
16:41:56.404 -> Average Cal sensor Temp.: 101.70 Ambient Temp.: 89.04 F
16:41:57.987 -> Average Cal sensor Temp.: 102.09 Ambient Temp.: 89.08 F
16:41:59.541 -> Average Cal sensor Temp.: 102.29 Ambient Temp.: 89.01 F
16:42:01.120 -> Average Cal sensor Temp.: 102.59 Ambient Temp.: 89.08 F
16:42:02.668 -> Average Cal sensor Temp.: 102.29 Ambient Temp.: 89.04 F
16:42:04.255 -> Average Cal sensor Temp.: 102.44 Ambient Temp.: 89.04 F
16:42:05.802 -> Average Cal sensor Temp.: 102.75 Ambient Temp.: 89.04 F
16:42:07.382 -> Average Cal sensor Temp.: 102.53 Ambient Temp.: 89.08 F
16:42:08.931 -> Average Cal sensor Temp.: 102.48 Ambient Temp.: 89.11 F
16:42:10.514 -> Average Cal sensor Temp.: 102.23 Ambient Temp.: 89.08 F
16:42:12.066 -> Average Cal sensor Temp.: 101.88 Ambient Temp.: 89.04 F
16:42:13.621 -> Average Cal sensor Temp.: 102.01 Ambient Temp.: 89.08 F
16:42:15.207 -> Average Cal sensor Temp.: 102.21 Ambient Temp.: 89.11 F
16:42:16.767 -> Average Cal sensor Temp.: 101.94 Ambient Temp.: 89.04 F
16:42:18.321 -> Average Cal sensor Temp.: 101.47 Ambient Temp.: 89.08 F
16:42:19.879 -> Average Cal sensor Temp.: 101.80 Ambient Temp.: 89.01 F
16:42:21.466 -> Average Cal sensor Temp.: 101.77 Ambient Temp.: 89.01 F
16:42:23.018 -> Average Cal sensor Temp.: 101.49 Ambient Temp.: 89.04 F
16:42:24.574 -> Average Cal sensor Temp.: 101.71 Ambient Temp.: 88.97 F
16:42:26.127 -> Average Cal sensor Temp.: 101.41 Ambient Temp.: 89.01 F
16:42:27.714 -> Average Cal sensor Temp.: 101.84 Ambient Temp.: 88.93 F
16:42:29.269 -> Average Cal sensor Temp.: 101.76 Ambient Temp.: 88.90 F
16:42:30.820 -> Average Cal sensor Temp.: 101.53 Ambient Temp.: 88.93 F
16:42:32.403 -> Average Cal sensor Temp.: 102.12 Ambient Temp.: 88.90 F
16:42:33.954 -> Average Cal sensor Temp.: 102.22 Ambient Temp.: 88.93 F
16:42:35.538 -> Average Cal sensor Temp.: 101.86 Ambient Temp.: 88.93 F
16:42:37.091 -> Average Cal sensor Temp.: 101.50 Ambient Temp.: 88.90 F
16:42:38.643 -> Average Cal sensor Temp.: 101.31 Ambient Temp.: 88.79 F
16:42:40.223 -> Average Cal sensor Temp.: 101.50 Ambient Temp.: 88.79 F
16:42:41.768 -> Average Cal sensor Temp.: 101.65 Ambient Temp.: 88.72 F
16:42:43.356 -> Average Cal sensor Temp.: 102.00 Ambient Temp.: 88.68 F
16:42:44.906 -> Average Cal sensor Temp.: 102.17 Ambient Temp.: 88.72 F
16:42:46.465 -> Average Cal sensor Temp.: 102.20 Ambient Temp.: 88.65 F
16:42:48.037 -> Average Cal sensor Temp.: 102.25 Ambient Temp.: 88.65 F
16:42:49.607 -> Average Cal sensor Temp.: 101.36 Ambient Temp.: 88.61 F
16:42:51.159 -> Average Cal sensor Temp.: 101.65 Ambient Temp.: 88.57 F
16:42:52.748 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 88.54 F
16:42:54.301 -> Average Cal sensor Temp.: 101.26 Ambient Temp.: 88.54 F
16:42:55.856 -> Average Cal sensor Temp.: 101.61 Ambient Temp.: 88.47 F
16:42:57.443 -> Average Cal sensor Temp.: 101.87 Ambient Temp.: 88.47 F
16:42:58.996 -> Average Cal sensor Temp.: 102.18 Ambient Temp.: 88.50 F
16:43:00.549 -> Average Cal sensor Temp.: 101.74 Ambient Temp.: 88.47 F
16:43:02.130 -> Average Cal sensor Temp.: 101.33 Ambient Temp.: 88.36 F
16:43:03.681 -> Average Cal sensor Temp.: 101.79 Ambient Temp.: 88.36 F
16:43:05.230 -> Average Cal sensor Temp.: 101.61 Ambient Temp.: 88.32 F
16:43:06.819 -> Average Cal sensor Temp.: 102.36 Ambient Temp.: 88.32 F
16:43:08.371 -> Average Cal sensor Temp.: 101.96 Ambient Temp.: 88.36 F
16:43:09.952 -> Average Cal sensor Temp.: 101.81 Ambient Temp.: 88.32 F
16:43:11.503 -> Average Cal sensor Temp.: 102.18 Ambient Temp.: 88.32 F
16:43:13.053 -> Average Cal sensor Temp.: 102.60 Ambient Temp.: 88.25 F
16:43:14.643 -> Average Cal sensor Temp.: 101.87 Ambient Temp.: 88.29 F
16:43:16.194 -> Average Cal sensor Temp.: 101.35 Ambient Temp.: 88.25 F
16:43:17.744 -> Average Cal sensor Temp.: 101.85 Ambient Temp.: 88.25 F
16:43:19.333 -> Average Cal sensor Temp.: 101.52 Ambient Temp.: 88.25 F
16:43:20.893 -> Average Cal sensor Temp.: 102.20 Ambient Temp.: 88.21 F
16:43:22.440 -> Average Cal sensor Temp.: 102.45 Ambient Temp.: 88.29 F
16:43:23.999 -> Average Cal sensor Temp.: 102.92 Ambient Temp.: 88.29 F
16:43:25.589 -> Average Cal sensor Temp.: 102.35 Ambient Temp.: 88.32 F
16:43:27.144 -> Average Cal sensor Temp.: 101.28 Ambient Temp.: 88.29 F
16:43:28.699 -> Average Cal sensor Temp.: 101.32 Ambient Temp.: 88.25 F
16:43:30.286 -> Average Cal sensor Temp.: 101.35 Ambient Temp.: 88.21 F
16:43:31.839 -> Average Cal sensor Temp.: 101.42 Ambient Temp.: 88.21 F
16:43:33.392 -> Average Cal sensor Temp.: 101.37 Ambient Temp.: 88.14 F
16:43:34.946 -> Average Cal sensor Temp.: 101.45 Ambient Temp.: 88.11 F
16:43:36.535 -> Average Cal sensor Temp.: 101.78 Ambient Temp.: 88.07 F
16:43:38.095 -> Average Cal sensor Temp.: 101.94 Ambient Temp.: 88.11 F
16:43:39.650 -> Average Cal sensor Temp.: 101.40 Ambient Temp.: 88.07 F
16:43:41.208 -> Average Cal sensor Temp.: 101.63 Ambient Temp.: 88.07 F
16:43:42.794 -> Average Cal sensor Temp.: 101.83 Ambient Temp.: 88.07 F
16:43:44.343 -> Average Cal sensor Temp.: 101.64 Ambient Temp.: 87.96 F
16:43:45.897 -> Average Cal sensor Temp.: 101.40 Ambient Temp.: 87.96 F
16:43:47.488 -> Average Cal sensor Temp.: 101.29 Ambient Temp.: 87.93 F
16:43:49.045 -> Average Cal sensor Temp.: 101.49 Ambient Temp.: 87.93 F
16:43:50.597 -> Average Cal sensor Temp.: 102.32 Ambient Temp.: 87.89 F
16:43:52.182 -> Average Cal sensor Temp.: 101.67 Ambient Temp.: 87.89 F
16:43:53.739 -> Average Cal sensor Temp.: 101.40 Ambient Temp.: 87.85 F
16:43:55.288 -> Average Cal sensor Temp.: 101.48 Ambient Temp.: 87.85 F
16:43:56.873 -> Average Cal sensor Temp.: 101.71 Ambient Temp.: 87.82 F
16:43:58.425 -> Average Cal sensor Temp.: 102.09 Ambient Temp.: 87.82 F
16:43:59.983 -> Average Cal sensor Temp.: 102.23 Ambient Temp.: 87.82 F
16:44:01.539 -> Average Cal sensor Temp.: 102.00 Ambient Temp.: 87.82 F
16:44:03.132 -> Average Cal sensor Temp.: 102.09 Ambient Temp.: 87.85 F
16:44:04.692 -> Average Cal sensor Temp.: 101.40 Ambient Temp.: 87.82 F
16:44:06.247 -> Average Cal sensor Temp.: 101.60 Ambient Temp.: 87.82 F
16:44:07.801 -> Average Cal sensor Temp.: 101.90 Ambient Temp.: 87.82 F
16:44:09.390 -> Average Cal sensor Temp.: 101.48 Ambient Temp.: 87.75 F
16:44:10.941 -> Average Cal sensor Temp.: 101.21 Ambient Temp.: 87.78 F
16:44:12.499 -> Average Cal sensor Temp.: 101.66 Ambient Temp.: 87.78 F
16:44:14.056 -> Average Cal sensor Temp.: 102.04 Ambient Temp.: 87.75 F
16:44:15.638 -> Average Cal sensor Temp.: 101.31 Ambient Temp.: 87.82 F
16:44:17.192 -> Average Cal sensor Temp.: 101.38 Ambient Temp.: 87.75 F
16:44:18.747 -> Average Cal sensor Temp.: 101.77 Ambient Temp.: 87.75 F
16:44:20.333 -> Average Cal sensor Temp.: 101.47 Ambient Temp.: 87.71 F
16:44:21.892 -> Average Cal sensor Temp.: 101.44 Ambient Temp.: 87.67 F
16:44:23.442 -> Average Cal sensor Temp.: 101.17 Ambient Temp.: 87.71 F
16:44:24.990 -> Average Cal sensor Temp.: 101.00 Ambient Temp.: 87.60 F
16:44:26.552 -> Average Cal sensor Temp.: 101.41 Ambient Temp.: 87.60 F
16:44:28.138 -> Average Cal sensor Temp.: 101.50 Ambient Temp.: 87.57 F
16:44:29.687 -> Average Cal sensor Temp.: 101.92 Ambient Temp.: 87.60 F
16:44:31.269 -> Average Cal sensor Temp.: 101.98 Ambient Temp.: 87.57 F
16:44:32.824 -> Average Cal sensor Temp.: 101.34 Ambient Temp.: 87.60 F
16:44:34.375 -> Average Cal sensor Temp.: 101.49 Ambient Temp.: 87.53 F
16:44:35.962 -> Average Cal sensor Temp.: 100.79 Ambient Temp.: 87.49 F
16:44:37.508 -> Average Cal sensor Temp.: 101.17 Ambient Temp.: 87.46 F
16:44:39.095 -> Average Cal sensor Temp.: 101.43 Ambient Temp.: 87.46 F
16:44:40.645 -> Average Cal sensor Temp.: 101.63 Ambient Temp.: 87.46 F
16:44:42.198 -> Average Cal sensor Temp.: 101.42 Ambient Temp.: 87.39 F
16:44:43.757 -> Average Cal sensor Temp.: 102.13 Ambient Temp.: 87.42 F
16:44:45.345 -> Average Cal sensor Temp.: 102.51 Ambient Temp.: 87.35 F
16:44:46.902 -> Average Cal sensor Temp.: 101.10 Ambient Temp.: 87.39 F
16:44:48.454 -> Average Cal sensor Temp.: 101.10 Ambient Temp.: 87.31 F
16:44:50.010 -> Average Cal sensor Temp.: 101.93 Ambient Temp.: 87.31 F
16:44:51.597 -> Average Cal sensor Temp.: 101.95 Ambient Temp.: 87.31 F
16:44:53.152 -> Average Cal sensor Temp.: 101.03 Ambient Temp.: 87.17 F
16:44:54.706 -> Average Cal sensor Temp.: 101.89 Ambient Temp.: 87.21 F
16:44:56.263 -> Average Cal sensor Temp.: 101.80 Ambient Temp.: 87.21 F
16:44:57.821 -> Average Cal sensor Temp.: 101.96 Ambient Temp.: 87.17 F
16:44:59.413 -> Average Cal sensor Temp.: 101.72 Ambient Temp.: 87.13 F
16:45:00.968 -> Average Cal sensor Temp.: 101.38 Ambient Temp.: 87.13 F
16:45:02.526 -> Average Cal sensor Temp.: 101.61 Ambient Temp.: 87.10 F
16:45:04.082 -> Average Cal sensor Temp.: 101.94 Ambient Temp.: 87.13 F
16:45:05.664 -> Average Cal sensor Temp.: 102.09 Ambient Temp.: 87.13 F
16:45:07.217 -> Average Cal sensor Temp.: 102.14 Ambient Temp.: 87.17 F
16:45:08.773 -> Average Cal sensor Temp.: 102.03 Ambient Temp.: 87.21 F
16:45:10.323 -> Average Cal sensor Temp.: 101.92 Ambient Temp.: 87.17 F
16:45:11.913 -> Average Cal sensor Temp.: 102.26 Ambient Temp.: 87.17 F
16:45:13.468 -> Average Cal sensor Temp.: 102.38 Ambient Temp.: 87.21 F
16:45:15.023 -> Average Cal sensor Temp.: 102.19 Ambient Temp.: 87.28 F
16:45:16.574 -> Average Cal sensor Temp.: 102.26 Ambient Temp.: 87.31 F
16:45:18.166 -> Average Cal sensor Temp.: 102.37 Ambient Temp.: 87.35 F
16:45:19.722 -> Average Cal sensor Temp.: 102.47 Ambient Temp.: 87.31 F
16:45:21.273 -> Average Cal sensor Temp.: 102.47 Ambient Temp.: 87.35 F
16:45:22.830 -> Average Cal sensor Temp.: 102.56 Ambient Temp.: 87.42 F
16:45:24.420 -> Average Cal sensor Temp.: 102.33 Ambient Temp.: 87.42 F
16:45:25.977 -> Average Cal sensor Temp.: 102.04 Ambient Temp.: 87.39 F
16:45:27.523 -> Average Cal sensor Temp.: 102.11 Ambient Temp.: 87.46 F
16:45:29.110 -> Average Cal sensor Temp.: 102.27 Ambient Temp.: 87.49 F
16:45:30.663 -> Average Cal sensor Temp.: 102.14 Ambient Temp.: 87.49 F
16:45:32.216 -> Average Cal sensor Temp.: 102.13 Ambient Temp.: 87.53 F
16:45:33.805 -> Average Cal sensor Temp.: 102.51 Ambient Temp.: 87.49 F
16:45:35.361 -> Average Cal sensor Temp.: 102.69 Ambient Temp.: 87.57 F
16:45:36.912 -> Average Cal sensor Temp.: 102.58 Ambient Temp.: 87.60 F
16:45:38.467 -> Average Cal sensor Temp.: 102.34 Ambient Temp.: 87.60 F
16:45:40.061 -> Average Cal sensor Temp.: 102.10 Ambient Temp.: 87.75 F
16:45:41.615 -> Average Cal sensor Temp.: 102.38 Ambient Temp.: 87.75 F
16:45:43.171 -> Average Cal sensor Temp.: 102.21 Ambient Temp.: 87.75 F
16:45:44.729 -> Average Cal sensor Temp.: 101.74 Ambient Temp.: 87.78 F
16:45:46.290 -> Average Cal sensor Temp.: 102.24 Ambient Temp.: 87.75 F
16:45:47.874 -> Average Cal sensor Temp.: 102.73 Ambient Temp.: 87.78 F
16:45:49.427 -> Average Cal sensor Temp.: 102.52 Ambient Temp.: 87.78 F
16:45:50.982 -> Average Cal sensor Temp.: 102.66 Ambient Temp.: 87.82 F
16:45:52.567 -> Average Cal sensor Temp.: 102.37 Ambient Temp.: 87.85 F
16:45:54.125 -> Average Cal sensor Temp.: 102.55 Ambient Temp.: 87.85 F
16:45:55.679 -> Average Cal sensor Temp.: 102.54 Ambient Temp.: 87.85 F
16:45:57.238 -> Average Cal sensor Temp.: 102.31 Ambient Temp.: 87.89 F
16:45:58.794 -> Average Cal sensor Temp.: 102.61 Ambient Temp.: 88.00 F
16:46:00.384 -> Average Cal sensor Temp.: 102.82 Ambient Temp.: 87.96 F
16:46:01.935 -> Average Cal sensor Temp.: 102.46 Ambient Temp.: 87.96 F
16:46:03.518 -> Average Cal sensor Temp.: 102.42 Ambient Temp.: 87.96 F
16:46:05.067 -> Average Cal sensor Temp.: 102.61 Ambient Temp.: 87.93 F
16:46:06.622 -> Average Cal sensor Temp.: 102.59 Ambient Temp.: 88.00 F
16:46:08.206 -> Average Cal sensor Temp.: 102.34 Ambient Temp.: 88.00 F
16:46:09.764 -> Average Cal sensor Temp.: 102.95 Ambient Temp.: 88.07 F
16:46:11.322 -> Average Cal sensor Temp.: 103.00 Ambient Temp.: 88.11 F
16:46:12.905 -> Average Cal sensor Temp.: 102.82 Ambient Temp.: 88.11 F
16:46:14.455 -> Average Cal sensor Temp.: 103.20 Ambient Temp.: 88.14 F
16:46:16.011 -> Average Cal sensor Temp.: 102.45 Ambient Temp.: 88.14 F
16:46:17.569 -> Average Cal sensor Temp.: 102.74 Ambient Temp.: 88.18 F
16:46:19.157 -> Average Cal sensor Temp.: 102.66 Ambient Temp.: 88.14 F
16:46:20.715 -> Average Cal sensor Temp.: 102.77 Ambient Temp.: 88.18 F
16:46:22.268 -> Average Cal sensor Temp.: 102.65 Ambient Temp.: 88.21 F
16:46:23.856 -> Average Cal sensor Temp.: 102.60 Ambient Temp.: 88.25 F
16:46:25.417 -> Average Cal sensor Temp.: 102.63 Ambient Temp.: 88.25 F
16:46:26.972 -> Average Cal sensor Temp.: 102.37 Ambient Temp.: 88.25 F
16:46:28.521 -> Average Cal sensor Temp.: 102.44 Ambient Temp.: 88.25 F
16:46:30.107 -> Average Cal sensor Temp.: 102.72 Ambient Temp.: 88.29 F
16:46:31.662 -> Average Cal sensor Temp.: 102.84 Ambient Temp.: 88.25 F
16:46:33.219 -> Average Cal sensor Temp.: 102.59 Ambient Temp.: 88.32 F
16:46:34.809 -> Average Cal sensor Temp.: 102.72 Ambient Temp.: 88.32 F
16:46:36.357 -> Average Cal sensor Temp.: 102.62 Ambient Temp.: 88.32 F
16:46:37.910 -> Average Cal sensor Temp.: 102.69 Ambient Temp.: 88.32 F
16:46:39.502 -> Average Cal sensor Temp.: 102.92 Ambient Temp.: 88.39 F
16:46:41.060 -> Average Cal sensor Temp.: 102.56 Ambient Temp.: 88.32 F
16:46:42.619 -> Average Cal sensor Temp.: 102.61 Ambient Temp.: 88.29 F
16:46:44.175 -> Average Cal sensor Temp.: 102.72 Ambient Temp.: 88.32 F
16:46:45.729 -> Average Cal sensor Temp.: 102.53 Ambient Temp.: 88.39 F
16:46:47.317 -> Average Cal sensor Temp.: 102.77 Ambient Temp.: 88.36 F
16:46:48.879 -> Average Cal sensor Temp.: 102.60 Ambient Temp.: 88.36 F
16:46:50.431 -> Average Cal sensor Temp.: 102.88 Ambient Temp.: 88.50 F
16:46:52.015 -> Average Cal sensor Temp.: 102.74 Ambient Temp.: 88.50 F
16:46:53.573 -> Average Cal sensor Temp.: 102.42 Ambient Temp.: 88.50 F
16:46:55.126 -> Average Cal sensor Temp.: 102.74 Ambient Temp.: 88.47 F
16:46:56.687 -> Average Cal sensor Temp.: 102.54 Ambient Temp.: 88.50 F
16:46:58.274 -> Average Cal sensor Temp.: 102.64 Ambient Temp.: 88.50 F
16:46:59.833 -> Average Cal sensor Temp.: 102.71 Ambient Temp.: 88.54 F
16:47:01.384 -> Average Cal sensor Temp.: 102.59 Ambient Temp.: 88.54 F
16:47:02.972 -> Average Cal sensor Temp.: 102.62 Ambient Temp.: 88.54 F
16:47:04.522 -> Average Cal sensor Temp.: 102.71 Ambient Temp.: 88.57 F
16:47:06.079 -> Average Cal sensor Temp.: 102.66 Ambient Temp.: 88.57 F
16:47:07.664 -> Average Cal sensor Temp.: 102.84 Ambient Temp.: 88.61 F
16:47:09.218 -> Average Cal sensor Temp.: 103.17 Ambient Temp.: 88.61 F
16:47:09.807 -> Ambient Temp.: 88.61 F
16:47:10.395 -> Ambient Temp.: 88.61 F
16:47:10.986 -> Ambient Temp.: 88.61 F
16:47:11.540 -> Ambient Temp.: 88.61 F
16:47:12.126 -> Ambient Temp.: 88.61 F
16:47:12.708 -> Ambient Temp.: 88.61 F
16:47:13.299 -> Ambient Temp.: 88.61 F
16:47:13.882 -> Ambient Temp.: 88.61 F
16:47:14.473 -> Ambient Temp.: 88.61 F
16:47:15.022 -> Ambient Temp.: 88.61 F
16:47:15.610 -> Ambient Temp.: 88.61 F
16:47:16.199 -> Ambient Temp.: 88.61 F
16:47:16.782 -> Ambient Temp.: 88.61 F
16:47:17.365 -> Ambient Temp.: 88.61 F
16:47:17.948 -> Ambient Temp.: 88.61 F
16:47:18.535 -> Ambient Temp.: 88.61 F
16:47:19.088 -> Ambient Temp.: 88.61 F
16:47:19.675 -> Ambient Temp.: 88.61 F
16:47:20.264 -> Ambient Temp.: 88.61 F
16:47:20.851 -> Ambient Temp.: 88.61 F
16:47:21.439 -> Ambient Temp.: 88.61 F
16:47:22.025 -> Ambient Temp.: 88.61 F
16:47:22.575 -> Ambient Temp.: 88.61 F
16:47:23.164 -> Ambient Temp.: 88.61 F
16:47:23.751 -> Ambient Temp.: 88.61 F
16:47:24.338 -> Ambient Temp.: 88.61 F
16:47:25.897 -> Average Cal sensor Temp.: 102.91 Ambient Temp.: 88.72 F
16:47:27.449 -> Average Cal sensor Temp.: 102.95 Ambient Temp.: 88.86 F
16:47:29.005 -> Average Cal sensor Temp.: 103.04 Ambient Temp.: 88.86 F
16:47:30.589 -> Average Cal sensor Temp.: 102.76 Ambient Temp.: 88.86 F
16:47:32.142 -> Average Cal sensor Temp.: 102.57 Ambient Temp.: 88.90 F
16:47:33.697 -> Average Cal sensor Temp.: 102.58 Ambient Temp.: 88.86 F
16:47:35.288 -> Average Cal sensor Temp.: 102.67 Ambient Temp.: 88.86 F
16:47:36.845 -> Average Cal sensor Temp.: 102.96 Ambient Temp.: 88.90 F
16:47:38.407 -> Average Cal sensor Temp.: 102.90 Ambient Temp.: 88.93 F
16:47:39.967 -> Average Cal sensor Temp.: 102.88 Ambient Temp.: 88.97 F
16:47:41.522 -> Average Cal sensor Temp.: 102.65 Ambient Temp.: 88.97 F
16:47:43.114 -> Average Cal sensor Temp.: 102.91 Ambient Temp.: 89.01 F
16:47:44.669 -> Average Cal sensor Temp.: 102.94 Ambient Temp.: 88.97 F
16:47:46.224 -> Average Cal sensor Temp.: 102.87 Ambient Temp.: 89.01 F
16:47:47.813 -> Average Cal sensor Temp.: 103.10 Ambient Temp.: 89.04 F
16:47:49.371 -> Average Cal sensor Temp.: 103.03 Ambient Temp.: 89.08 F
16:47:50.925 -> Average Cal sensor Temp.: 103.44 Ambient Temp.: 89.08 F
16:47:52.481 -> Average Cal sensor Temp.: 103.37 Ambient Temp.: 89.08 F
16:47:54.071 -> Average Cal sensor Temp.: 102.70 Ambient Temp.: 89.11 F
16:47:55.625 -> Average Cal sensor Temp.: 103.01 Ambient Temp.: 89.11 F
16:47:57.181 -> Average Cal sensor Temp.: 103.24 Ambient Temp.: 89.11 F
16:47:58.732 -> Average Cal sensor Temp.: 103.19 Ambient Temp.: 89.15 F
16:48:00.320 -> Average Cal sensor Temp.: 102.53 Ambient Temp.: 89.15 F
16:48:01.871 -> Average Cal sensor Temp.: 102.79 Ambient Temp.: 89.11 F
16:48:03.455 -> Average Cal sensor Temp.: 103.03 Ambient Temp.: 89.15 F
16:48:05.001 -> Average Cal sensor Temp.: 102.68 Ambient Temp.: 89.15 F
16:48:06.556 -> Average Cal sensor Temp.: 102.39 Ambient Temp.: 89.19 F
16:48:08.141 -> Average Cal sensor Temp.: 102.33 Ambient Temp.: 89.19 F
16:48:09.689 -> Average Cal sensor Temp.: 102.63 Ambient Temp.: 89.26 F
16:48:11.272 -> Average Cal sensor Temp.: 102.74 Ambient Temp.: 89.19 F
16:48:12.828 -> Average Cal sensor Temp.: 102.79 Ambient Temp.: 89.26 F
16:48:14.383 -> Average Cal sensor Temp.: 102.15 Ambient Temp.: 89.19 F
16:48:15.943 -> Average Cal sensor Temp.: 101.97 Ambient Temp.: 89.19 F
16:48:17.532 -> Average Cal sensor Temp.: 102.14 Ambient Temp.: 89.19 F
16:48:19.078 -> Average Cal sensor Temp.: 102.10 Ambient Temp.: 89.11 F
16:48:20.660 -> Average Cal sensor Temp.: 102.16 Ambient Temp.: 89.11 F
16:48:22.212 -> Average Cal sensor Temp.: 102.56 Ambient Temp.: 89.15 F
16:48:23.768 -> Average Cal sensor Temp.: 102.78 Ambient Temp.: 89.08 F
16:48:25.319 -> Average Cal sensor Temp.: 102.62 Ambient Temp.: 89.15 F
16:48:26.915 -> Average Cal sensor Temp.: 102.77 Ambient Temp.: 89.11 F
16:48:28.472 -> Average Cal sensor Temp.: 102.53 Ambient Temp.: 89.11 F
16:48:30.021 -> Average Cal sensor Temp.: 102.89 Ambient Temp.: 89.11 F
16:48:31.577 -> Average Cal sensor Temp.: 102.93 Ambient Temp.: 89.19 F
16:48:33.160 -> Average Cal sensor Temp.: 102.62 Ambient Temp.: 89.19 F
16:48:34.709 -> Average Cal sensor Temp.: 103.16 Ambient Temp.: 89.15 F
16:48:36.265 -> Average Cal sensor Temp.: 102.85 Ambient Temp.: 89.15 F
16:48:37.850 -> Average Cal sensor Temp.: 102.46 Ambient Temp.: 89.19 F
16:48:39.405 -> Average Cal sensor Temp.: 102.55 Ambient Temp.: 89.19 F
16:48:40.990 -> Average Cal sensor Temp.: 102.66 Ambient Temp.: 89.15 F
16:48:42.545 -> Average Cal sensor Temp.: 102.79 Ambient Temp.: 89.19 F
16:48:44.096 -> Average Cal sensor Temp.: 102.86 Ambient Temp.: 89.19 F
16:48:45.687 -> Average Cal sensor Temp.: 102.44 Ambient Temp.: 89.19 F
16:48:47.242 -> Average Cal sensor Temp.: 102.43 Ambient Temp.: 89.19 F
16:48:48.797 -> Average Cal sensor Temp.: 102.45 Ambient Temp.: 89.19 F
16:48:50.347 -> Average Cal sensor Temp.: 102.71 Ambient Temp.: 89.29 F
16:48:51.936 -> Average Cal sensor Temp.: 102.32 Ambient Temp.: 89.19 F
16:48:53.492 -> Average Cal sensor Temp.: 102.39 Ambient Temp.: 89.29 F
16:48:55.044 -> Average Cal sensor Temp.: 102.31 Ambient Temp.: 89.26 F
16:48:56.634 -> Average Cal sensor Temp.: 102.30 Ambient Temp.: 89.19 F
16:48:58.191 -> Average Cal sensor Temp.: 102.32 Ambient Temp.: 89.19 F
16:48:59.742 -> Average Cal sensor Temp.: 101.84 Ambient Temp.: 89.11 F
16:49:01.327 -> Average Cal sensor Temp.: 102.29 Ambient Temp.: 89.15 F
16:49:02.878 -> Average Cal sensor Temp.: 102.02 Ambient Temp.: 89.04 F
16:49:04.430 -> Average Cal sensor Temp.: 102.37 Ambient Temp.: 89.04 F
16:49:06.021 -> Average Cal sensor Temp.: 102.38 Ambient Temp.: 89.04 F
16:49:07.570 -> Average Cal sensor Temp.: 102.12 Ambient Temp.: 88.97 F
16:49:09.139 -> Average Cal sensor Temp.: 102.29 Ambient Temp.: 88.97 F
16:49:10.691 -> Average Cal sensor Temp.: 102.34 Ambient Temp.: 88.97 F
16:49:12.279 -> Average Cal sensor Temp.: 102.43 Ambient Temp.: 88.97 F
16:49:13.828 -> Average Cal sensor Temp.: 101.73 Ambient Temp.: 88.93 F
16:49:15.379 -> Average Cal sensor Temp.: 101.95 Ambient Temp.: 88.90 F
16:49:16.974 -> Average Cal sensor Temp.: 102.10 Ambient Temp.: 88.86 F
16:49:18.528 -> Average Cal sensor Temp.: 102.62 Ambient Temp.: 88.86 F
16:49:20.082 -> Average Cal sensor Temp.: 102.39 Ambient Temp.: 88.75 F
16:49:21.667 -> Average Cal sensor Temp.: 102.10 Ambient Temp.: 88.75 F
16:49:23.222 -> Average Cal sensor Temp.: 101.64 Ambient Temp.: 88.72 F
16:49:24.778 -> Average Cal sensor Temp.: 102.43 Ambient Temp.: 88.68 F
16:49:26.336 -> Average Cal sensor Temp.: 102.45 Ambient Temp.: 88.72 F
16:49:27.923 -> Average Cal sensor Temp.: 102.38 Ambient Temp.: 88.65 F
16:49:29.480 -> Average Cal sensor Temp.: 102.12 Ambient Temp.: 88.68 F
16:49:31.033 -> Average Cal sensor Temp.: 102.43 Ambient Temp.: 88.68 F
16:49:32.617 -> Average Cal sensor Temp.: 102.61 Ambient Temp.: 88.65 F
16:49:34.168 -> Average Cal sensor Temp.: 102.50 Ambient Temp.: 88.65 F
16:49:35.755 -> Average Cal sensor Temp.: 102.78 Ambient Temp.: 88.65 F
16:49:37.306 -> Average Cal sensor Temp.: 102.56 Ambient Temp.: 88.65 F
16:49:38.858 -> Average Cal sensor Temp.: 102.02 Ambient Temp.: 88.61 F
16:49:40.444 -> Average Cal sensor Temp.: 101.70 Ambient Temp.: 88.54 F
16:49:42.003 -> Average Cal sensor Temp.: 102.28 Ambient Temp.: 88.61 F
16:49:43.560 -> Average Cal sensor Temp.: 102.45 Ambient Temp.: 88.57 F
16:49:45.120 -> Average Cal sensor Temp.: 102.32 Ambient Temp.: 88.54 F
16:49:46.681 -> Average Cal sensor Temp.: 101.81 Ambient Temp.: 88.47 F
16:49:48.267 -> Average Cal sensor Temp.: 102.42 Ambient Temp.: 88.50 F
16:49:49.821 -> Average Cal sensor Temp.: 102.44 Ambient Temp.: 88.50 F
16:49:51.371 -> Average Cal sensor Temp.: 102.47 Ambient Temp.: 88.47 F
16:49:52.955 -> Average Cal sensor Temp.: 102.99 Ambient Temp.: 88.50 F
16:49:54.510 -> Average Cal sensor Temp.: 102.75 Ambient Temp.: 88.50 F
16:49:56.061 -> Average Cal sensor Temp.: 102.53 Ambient Temp.: 88.54 F
16:49:57.645 -> Average Cal sensor Temp.: 102.56 Ambient Temp.: 88.50 F
16:49:59.199 -> Average Cal sensor Temp.: 102.74 Ambient Temp.: 88.47 F
16:50:00.788 -> Average Cal sensor Temp.: 102.41 Ambient Temp.: 88.50 F
16:50:02.336 -> Average Cal sensor Temp.: 102.14 Ambient Temp.: 88.47 F
16:50:03.890 -> Average Cal sensor Temp.: 102.43 Ambient Temp.: 88.50 F
16:50:05.477 -> Average Cal sensor Temp.: 101.64 Ambient Temp.: 88.39 F
16:50:07.039 -> Average Cal sensor Temp.: 101.70 Ambient Temp.: 88.39 F
16:50:08.595 -> Average Cal sensor Temp.: 101.75 Ambient Temp.: 88.36 F
16:50:10.149 -> Average Cal sensor Temp.: 102.28 Ambient Temp.: 88.47 F
16:50:11.733 -> Average Cal sensor Temp.: 101.72 Ambient Temp.: 88.39 F
16:50:13.286 -> Average Cal sensor Temp.: 102.01 Ambient Temp.: 88.36 F
16:50:14.845 -> Average Cal sensor Temp.: 102.34 Ambient Temp.: 88.39 F
16:50:16.398 -> Average Cal sensor Temp.: 102.28 Ambient Temp.: 88.29 F
16:50:17.987 -> Average Cal sensor Temp.: 101.66 Ambient Temp.: 88.36 F
16:50:19.541 -> Average Cal sensor Temp.: 101.64 Ambient Temp.: 88.29 F
16:50:21.098 -> Average Cal sensor Temp.: 102.45 Ambient Temp.: 88.25 F
16:50:22.681 -> Average Cal sensor Temp.: 102.15 Ambient Temp.: 88.25 F
16:50:24.232 -> Average Cal sensor Temp.: 101.84 Ambient Temp.: 88.25 F
16:50:25.788 -> Average Cal sensor Temp.: 102.14 Ambient Temp.: 88.18 F
16:50:27.375 -> Average Cal sensor Temp.: 102.50 Ambient Temp.: 88.18 F
16:50:28.935 -> Average Cal sensor Temp.: 102.41 Ambient Temp.: 88.18 F
16:50:30.493 -> Average Cal sensor Temp.: 102.60 Ambient Temp.: 88.25 F
16:50:32.050 -> Average Cal sensor Temp.: 102.55 Ambient Temp.: 88.21 F
16:50:33.636 -> Average Cal sensor Temp.: 102.64 Ambient Temp.: 88.18 F
16:50:35.188 -> Average Cal sensor Temp.: 102.65 Ambient Temp.: 88.21 F
16:50:36.742 -> Average Cal sensor Temp.: 102.51 Ambient Temp.: 88.21 F
16:50:38.330 -> Average Cal sensor Temp.: 102.41 Ambient Temp.: 88.18 F
16:50:39.891 -> Average Cal sensor Temp.: 101.83 Ambient Temp.: 88.14 F
16:50:41.448 -> Average Cal sensor Temp.: 102.15 Ambient Temp.: 88.14 F
16:50:43.006 -> Average Cal sensor Temp.: 101.66 Ambient Temp.: 88.07 F
16:50:44.558 -> Average Cal sensor Temp.: 101.63 Ambient Temp.: 88.07 F
16:50:46.148 -> Average Cal sensor Temp.: 102.02 Ambient Temp.: 88.07 F
16:50:47.703 -> Average Cal sensor Temp.: 102.40 Ambient Temp.: 88.11 F
16:50:49.257 -> Average Cal sensor Temp.: 101.77 Ambient Temp.: 88.00 F
16:50:50.845 -> Average Cal sensor Temp.: 102.29 Ambient Temp.: 88.07 F
16:50:52.395 -> Average Cal sensor Temp.: 101.90 Ambient Temp.: 88.07 F
16:50:53.955 -> Average Cal sensor Temp.: 102.11 Ambient Temp.: 88.00 F
16:50:55.536 -> Average Cal sensor Temp.: 101.57 Ambient Temp.: 87.93 F
16:50:57.099 -> Average Cal sensor Temp.: 101.71 Ambient Temp.: 87.89 F
16:50:58.652 -> Average Cal sensor Temp.: 102.47 Ambient Temp.: 87.89 F
16:51:00.203 -> Average Cal sensor Temp.: 102.59 Ambient Temp.: 87.85 F
16:51:01.790 -> Average Cal sensor Temp.: 102.49 Ambient Temp.: 87.85 F
16:51:03.347 -> Average Cal sensor Temp.: 102.54 Ambient Temp.: 87.85 F
16:51:04.896 -> Average Cal sensor Temp.: 102.34 Ambient Temp.: 87.78 F
16:51:06.486 -> Average Cal sensor Temp.: 102.80 Ambient Temp.: 87.85 F
16:51:08.042 -> Average Cal sensor Temp.: 103.40 Ambient Temp.: 87.89 F
16:51:09.602 -> Average Cal sensor Temp.: 102.27 Ambient Temp.: 87.85 F
16:51:11.159 -> Average Cal sensor Temp.: 102.32 Ambient Temp.: 87.85 F
16:51:12.712 -> Average Cal sensor Temp.: 102.98 Ambient Temp.: 87.85 F
16:51:14.306 -> Average Cal sensor Temp.: 103.16 Ambient Temp.: 87.89 F
16:51:15.861 -> Average Cal sensor Temp.: 102.76 Ambient Temp.: 87.89 F
16:51:17.411 -> Average Cal sensor Temp.: 102.66 Ambient Temp.: 87.89 F
16:51:18.996 -> Average Cal sensor Temp.: 102.64 Ambient Temp.: 87.89 F
16:51:20.555 -> Average Cal sensor Temp.: 102.81 Ambient Temp.: 87.93 F
16:51:22.103 -> Average Cal sensor Temp.: 102.84 Ambient Temp.: 87.93 F
16:51:23.684 -> Average Cal sensor Temp.: 102.42 Ambient Temp.: 87.96 F
16:51:25.239 -> Average Cal sensor Temp.: 102.89 Ambient Temp.: 87.96 F
16:51:26.830 -> Average Cal sensor Temp.: 102.35 Ambient Temp.: 88.00 F
16:51:28.385 -> Average Cal sensor Temp.: 102.69 Ambient Temp.: 87.93 F
16:51:29.937 -> Average Cal sensor Temp.: 103.05 Ambient Temp.: 87.93 F
16:51:31.492 -> Average Cal sensor Temp.: 103.09 Ambient Temp.: 87.96 F
16:51:33.083 -> Average Cal sensor Temp.: 102.94 Ambient Temp.: 87.89 F
16:51:34.639 -> Average Cal sensor Temp.: 102.72 Ambient Temp.: 87.96 F
16:51:36.192 -> Average Cal sensor Temp.: 102.69 Ambient Temp.: 87.93 F
16:51:37.776 -> Average Cal sensor Temp.: 102.51 Ambient Temp.: 87.96 F
16:51:39.332 -> Average Cal sensor Temp.: 102.56 Ambient Temp.: 87.96 F
16:51:40.889 -> Average Cal sensor Temp.: 102.78 Ambient Temp.: 88.00 F
16:51:42.479 -> Average Cal sensor Temp.: 103.18 Ambient Temp.: 87.96 F
16:51:44.035 -> Average Cal sensor Temp.: 102.41 Ambient Temp.: 87.93 F
16:51:45.589 -> Average Cal sensor Temp.: 102.79 Ambient Temp.: 87.93 F
16:51:47.141 -> Average Cal sensor Temp.: 102.85 Ambient Temp.: 87.93 F
16:51:48.734 -> Average Cal sensor Temp.: 102.89 Ambient Temp.: 87.96 F
16:51:50.286 -> Average Cal sensor Temp.: 102.87 Ambient Temp.: 88.00 F
16:51:51.848 -> Average Cal sensor Temp.: 102.84 Ambient Temp.: 87.89 F
16:51:53.407 -> Average Cal sensor Temp.: 102.72 Ambient Temp.: 87.96 F
16:51:54.966 -> Average Cal sensor Temp.: 102.94 Ambient Temp.: 87.96 F
16:51:56.557 -> Average Cal sensor Temp.: 103.37 Ambient Temp.: 87.96 F
16:51:58.115 -> Average Cal sensor Temp.: 102.90 Ambient Temp.: 88.00 F
16:51:59.668 -> Average Cal sensor Temp.: 102.29 Ambient Temp.: 88.07 F
16:52:01.224 -> Average Cal sensor Temp.: 102.79 Ambient Temp.: 88.11 F
16:52:02.810 -> Average Cal sensor Temp.: 102.74 Ambient Temp.: 87.96 F
16:52:04.366 -> Average Cal sensor Temp.: 102.70 Ambient Temp.: 88.00 F
16:52:05.921 -> Average Cal sensor Temp.: 102.86 Ambient Temp.: 88.00 F
16:52:07.514 -> Average Cal sensor Temp.: 102.99 Ambient Temp.: 87.96 F
16:52:09.069 -> Average Cal sensor Temp.: 102.68 Ambient Temp.: 88.00 F
16:52:10.621 -> Average Cal sensor Temp.: 102.61 Ambient Temp.: 88.00 F
16:52:12.211 -> Average Cal sensor Temp.: 102.50 Ambient Temp.: 88.00 F
16:52:13.771 -> Average Cal sensor Temp.: 102.50 Ambient Temp.: 87.93 F
16:52:15.324 -> Average Cal sensor Temp.: 102.92 Ambient Temp.: 87.96 F
16:52:16.883 -> Average Cal sensor Temp.: 103.13 Ambient Temp.: 88.00 F
16:52:18.470 -> Average Cal sensor Temp.: 102.94 Ambient Temp.: 88.00 F
16:52:20.024 -> Average Cal sensor Temp.: 103.13 Ambient Temp.: 88.07 F
16:52:21.575 -> Average Cal sensor Temp.: 102.95 Ambient Temp.: 88.00 F
16:52:23.159 -> Average Cal sensor Temp.: 103.03 Ambient Temp.: 88.00 F
16:52:24.709 -> Average Cal sensor Temp.: 103.23 Ambient Temp.: 87.96 F
16:52:26.266 -> Average Cal sensor Temp.: 102.59 Ambient Temp.: 88.07 F
16:52:27.860 -> Average Cal sensor Temp.: 103.00 Ambient Temp.: 88.00 F
16:52:29.412 -> Average Cal sensor Temp.: 102.84 Ambient Temp.: 88.07 F
16:52:30.976 -> Average Cal sensor Temp.: 103.11 Ambient Temp.: 88.07 F
16:52:32.527 -> Average Cal sensor Temp.: 103.01 Ambient Temp.: 88.11 F
16:52:34.116 -> Average Cal sensor Temp.: 103.33 Ambient Temp.: 88.11 F
16:52:35.665 -> Average Cal sensor Temp.: 103.19 Ambient Temp.: 88.11 F
16:52:37.220 -> Average Cal sensor Temp.: 103.26 Ambient Temp.: 88.14 F
16:52:38.811 -> Average Cal sensor Temp.: 103.01 Ambient Temp.: 88.18 F
16:52:40.366 -> Average Cal sensor Temp.: 103.22 Ambient Temp.: 88.14 F
16:52:41.926 -> Average Cal sensor Temp.: 103.40 Ambient Temp.: 88.14 F
16:52:43.486 -> Average Cal sensor Temp.: 102.88 Ambient Temp.: 88.11 F
16:52:45.040 -> Average Cal sensor Temp.: 102.86 Ambient Temp.: 88.18 F
16:52:46.633 -> Average Cal sensor Temp.: 103.13 Ambient Temp.: 88.14 F
16:52:48.190 -> Average Cal sensor Temp.: 103.08 Ambient Temp.: 88.18 F
16:52:49.741 -> Average Cal sensor Temp.: 103.28 Ambient Temp.: 88.21 F
16:52:51.296 -> Average Cal sensor Temp.: 103.10 Ambient Temp.: 88.18 F
16:52:52.884 -> Average Cal sensor Temp.: 102.80 Ambient Temp.: 88.21 F
16:52:54.437 -> Average Cal sensor Temp.: 102.91 Ambient Temp.: 88.21 F
16:52:55.995 -> Average Cal sensor Temp.: 102.87 Ambient Temp.: 88.21 F
16:52:57.581 -> Average Cal sensor Temp.: 102.81 Ambient Temp.: 88.25 F
16:52:59.132 -> Average Cal sensor Temp.: 103.18 Ambient Temp.: 88.21 F
16:53:00.688 -> Average Cal sensor Temp.: 102.89 Ambient Temp.: 88.21 F
16:53:02.275 -> Average Cal sensor Temp.: 102.76 Ambient Temp.: 88.21 F
16:53:03.831 -> Average Cal sensor Temp.: 103.10 Ambient Temp.: 88.21 F
16:53:05.383 -> Average Cal sensor Temp.: 102.99 Ambient Temp.: 88.25 F
16:53:06.970 -> Average Cal sensor Temp.: 103.02 Ambient Temp.: 88.25 F
16:53:08.530 -> Average Cal sensor Temp.: 102.82 Ambient Temp.: 88.29 F
16:53:10.085 -> Average Cal sensor Temp.: 102.71 Ambient Temp.: 88.25 F
16:53:11.646 -> Average Cal sensor Temp.: 102.81 Ambient Temp.: 88.25 F
16:53:13.200 -> Average Cal sensor Temp.: 102.85 Ambient Temp.: 88.32 F
16:53:14.789 -> Average Cal sensor Temp.: 103.23 Ambient Temp.: 88.29 F
16:53:16.337 -> Average Cal sensor Temp.: 103.70 Ambient Temp.: 88.29 F
16:53:17.892 -> Average Cal sensor Temp.: 103.41 Ambient Temp.: 88.32 F
16:53:19.476 -> Average Cal sensor Temp.: 103.22 Ambient Temp.: 88.32 F
16:53:21.029 -> Average Cal sensor Temp.: 103.20 Ambient Temp.: 88.32 F
16:53:22.592 -> Average Cal sensor Temp.: 103.25 Ambient Temp.: 88.39 F
16:53:24.148 -> Average Cal sensor Temp.: 103.15 Ambient Temp.: 88.36 F
16:53:25.735 -> Average Cal sensor Temp.: 103.54 Ambient Temp.: 88.47 F
16:53:27.287 -> Average Cal sensor Temp.: 103.51 Ambient Temp.: 88.54 F
16:53:28.841 -> Average Cal sensor Temp.: 103.16 Ambient Temp.: 88.50 F
16:53:30.430 -> Average Cal sensor Temp.: 103.52 Ambient Temp.: 88.50 F
16:53:31.986 -> Average Cal sensor Temp.: 103.56 Ambient Temp.: 88.54 F
16:53:33.535 -> Average Cal sensor Temp.: 103.05 Ambient Temp.: 88.57 F
16:53:35.121 -> Average Cal sensor Temp.: 103.45 Ambient Temp.: 88.57 F
16:53:36.674 -> Average Cal sensor Temp.: 103.46 Ambient Temp.: 88.61 F
16:53:38.225 -> Average Cal sensor Temp.: 103.72 Ambient Temp.: 88.65 F
16:53:39.813 -> Average Cal sensor Temp.: 103.38 Ambient Temp.: 88.65 F
16:53:41.366 -> Average Cal sensor Temp.: 103.58 Ambient Temp.: 88.65 F
16:53:42.949 -> Average Cal sensor Temp.: 103.56 Ambient Temp.: 88.65 F
16:53:44.504 -> Average Cal sensor Temp.: 103.55 Ambient Temp.: 88.75 F
16:53:46.054 -> Average Cal sensor Temp.: 102.96 Ambient Temp.: 88.75 F
16:53:47.610 -> Average Cal sensor Temp.: 103.20 Ambient Temp.: 88.75 F
16:53:49.194 -> Average Cal sensor Temp.: 103.07 Ambient Temp.: 88.75 F
16:53:50.748 -> Average Cal sensor Temp.: 103.52 Ambient Temp.: 88.75 F
16:53:52.333 -> Average Cal sensor Temp.: 103.33 Ambient Temp.: 88.86 F
16:53:53.891 -> Average Cal sensor Temp.: 103.55 Ambient Temp.: 88.86 F
16:53:55.440 -> Average Cal sensor Temp.: 103.49 Ambient Temp.: 88.90 F
16:53:56.998 -> Average Cal sensor Temp.: 103.40 Ambient Temp.: 88.93 F
16:53:58.580 -> Average Cal sensor Temp.: 103.39 Ambient Temp.: 88.97 F
16:54:00.128 -> Average Cal sensor Temp.: 103.77 Ambient Temp.: 89.01 F
16:54:00.749 -> Ambient Temp.: 89.01 F
16:54:02.300 -> Average Cal sensor Temp.: 103.47 Ambient Temp.: 89.04 F
16:54:03.854 -> Average Cal sensor Temp.: 103.53 Ambient Temp.: 89.04 F
16:54:05.406 -> Average Cal sensor Temp.: 103.53 Ambient Temp.: 89.11 F
16:54:06.965 -> Average Cal sensor Temp.: 103.61 Ambient Temp.: 89.11 F
16:54:08.551 -> Average Cal sensor Temp.: 103.44 Ambient Temp.: 89.15 F
16:54:10.099 -> Average Cal sensor Temp.: 103.46 Ambient Temp.: 89.19 F
16:54:11.685 -> Average Cal sensor Temp.: 103.27 Ambient Temp.: 89.19 F
16:54:13.238 -> Average Cal sensor Temp.: 103.56 Ambient Temp.: 89.26 F
16:54:14.790 -> Average Cal sensor Temp.: 103.74 Ambient Temp.: 89.33 F
16:54:16.382 -> Average Cal sensor Temp.: 103.35 Ambient Temp.: 89.33 F
16:54:17.931 -> Average Cal sensor Temp.: 103.47 Ambient Temp.: 89.37 F
16:54:19.490 -> Average Cal sensor Temp.: 103.57 Ambient Temp.: 89.33 F
16:54:21.045 -> Average Cal sensor Temp.: 103.53 Ambient Temp.: 89.40 F
16:54:22.631 -> Average Cal sensor Temp.: 103.27 Ambient Temp.: 89.44 F
16:54:24.185 -> Average Cal sensor Temp.: 103.08 Ambient Temp.: 89.47 F
16:54:25.737 -> Average Cal sensor Temp.: 103.17 Ambient Temp.: 89.44 F
16:54:27.325 -> Average Cal sensor Temp.: 103.08 Ambient Temp.: 89.47 F
16:54:28.883 -> Average Cal sensor Temp.: 102.73 Ambient Temp.: 89.51 F
16:54:30.440 -> Average Cal sensor Temp.: 102.78 Ambient Temp.: 89.51 F
16:54:31.994 -> Average Cal sensor Temp.: 102.82 Ambient Temp.: 89.51 F
16:54:33.567 -> Average Cal sensor Temp.: 102.55 Ambient Temp.: 89.47 F
16:54:35.122 -> Average Cal sensor Temp.: 102.91 Ambient Temp.: 89.55 F
16:54:36.712 -> Average Cal sensor Temp.: 102.80 Ambient Temp.: 89.47 F
16:54:38.262 -> Average Cal sensor Temp.: 102.67 Ambient Temp.: 89.47 F
16:54:39.815 -> Average Cal sensor Temp.: 102.71 Ambient Temp.: 89.47 F
16:54:41.402 -> Average Cal sensor Temp.: 102.99 Ambient Temp.: 89.44 F
16:54:42.953 -> Average Cal sensor Temp.: 103.42 Ambient Temp.: 89.47 F
16:54:44.528 -> Average Cal sensor Temp.: 103.30 Ambient Temp.: 89.47 F
16:54:46.079 -> Average Cal sensor Temp.: 103.38 Ambient Temp.: 89.47 F
16:54:47.666 -> Average Cal sensor Temp.: 102.91 Ambient Temp.: 89.51 F
16:54:49.220 -> Average Cal sensor Temp.: 102.75 Ambient Temp.: 89.51 F
16:54:50.771 -> Average Cal sensor Temp.: 103.14 Ambient Temp.: 89.47 F
16:54:52.356 -> Average Cal sensor Temp.: 103.17 Ambient Temp.: 89.51 F
16:54:53.906 -> Average Cal sensor Temp.: 102.97 Ambient Temp.: 89.47 F
16:54:55.456 -> Average Cal sensor Temp.: 102.78 Ambient Temp.: 89.47 F
16:54:57.049 -> Average Cal sensor Temp.: 102.78 Ambient Temp.: 89.44 F
16:54:58.602 -> Average Cal sensor Temp.: 103.20 Ambient Temp.: 89.51 F
16:55:00.153 -> Average Cal sensor Temp.: 103.28 Ambient Temp.: 89.47 F
16:55:01.740 -> Average Cal sensor Temp.: 103.44 Ambient Temp.: 89.51 F
16:55:03.294 -> Average Cal sensor Temp.: 102.87 Ambient Temp.: 89.51 F
16:55:04.844 -> Average Cal sensor Temp.: 103.12 Ambient Temp.: 89.44 F
16:55:06.433 -> Average Cal sensor Temp.: 103.52 Ambient Temp.: 89.47 F
16:55:07.986 -> Average Cal sensor Temp.: 103.77 Ambient Temp.: 89.51 F
16:55:09.538 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 89.47 F
16:55:11.097 -> Average Cal sensor Temp.: 103.47 Ambient Temp.: 89.47 F
16:55:12.686 -> Average Cal sensor Temp.: 103.02 Ambient Temp.: 89.47 F
16:55:14.236 -> Average Cal sensor Temp.: 103.10 Ambient Temp.: 89.44 F
16:55:15.821 -> Average Cal sensor Temp.: 103.18 Ambient Temp.: 89.47 F
16:55:17.378 -> Average Cal sensor Temp.: 102.91 Ambient Temp.: 89.47 F
16:55:18.930 -> Average Cal sensor Temp.: 102.09 Ambient Temp.: 89.44 F
16:55:20.486 -> Average Cal sensor Temp.: 101.46 Ambient Temp.: 89.40 F
16:55:22.078 -> Average Cal sensor Temp.: 101.91 Ambient Temp.: 89.33 F
16:55:23.634 -> Average Cal sensor Temp.: 102.20 Ambient Temp.: 89.26 F
16:55:25.187 -> Average Cal sensor Temp.: 102.61 Ambient Temp.: 89.26 F
16:55:26.744 -> Average Cal sensor Temp.: 103.14 Ambient Temp.: 89.29 F
16:55:28.334 -> Average Cal sensor Temp.: 102.80 Ambient Temp.: 89.29 F
16:55:29.885 -> Average Cal sensor Temp.: 102.73 Ambient Temp.: 89.29 F
16:55:31.440 -> Average Cal sensor Temp.: 102.87 Ambient Temp.: 89.29 F
16:55:33.025 -> Average Cal sensor Temp.: 103.29 Ambient Temp.: 89.19 F
16:55:34.581 -> Average Cal sensor Temp.: 103.32 Ambient Temp.: 89.19 F
16:55:36.132 -> Average Cal sensor Temp.: 103.55 Ambient Temp.: 89.29 F
16:55:37.713 -> Average Cal sensor Temp.: 102.99 Ambient Temp.: 89.15 F
16:55:39.266 -> Average Cal sensor Temp.: 103.48 Ambient Temp.: 89.26 F
16:55:40.825 -> Average Cal sensor Temp.: 103.78 Ambient Temp.: 89.26 F
16:55:42.413 -> Average Cal sensor Temp.: 103.57 Ambient Temp.: 89.26 F
16:55:43.962 -> Average Cal sensor Temp.: 103.13 Ambient Temp.: 89.15 F
16:55:45.549 -> Average Cal sensor Temp.: 103.22 Ambient Temp.: 89.19 F
16:55:47.109 -> Average Cal sensor Temp.: 102.65 Ambient Temp.: 89.11 F
16:55:48.664 -> Average Cal sensor Temp.: 102.53 Ambient Temp.: 89.15 F
16:55:50.221 -> Average Cal sensor Temp.: 102.92 Ambient Temp.: 89.15 F
16:55:51.775 -> Average Cal sensor Temp.: 103.29 Ambient Temp.: 89.11 F
16:55:53.365 -> Average Cal sensor Temp.: 103.40 Ambient Temp.: 89.15 F
16:55:54.919 -> Average Cal sensor Temp.: 102.92 Ambient Temp.: 89.19 F
16:55:56.476 -> Average Cal sensor Temp.: 102.98 Ambient Temp.: 89.15 F
16:55:58.032 -> Average Cal sensor Temp.: 102.77 Ambient Temp.: 89.11 F
16:55:59.620 -> Average Cal sensor Temp.: 102.39 Ambient Temp.: 89.11 F
16:56:01.173 -> Average Cal sensor Temp.: 102.38 Ambient Temp.: 89.08 F
16:56:02.725 -> Average Cal sensor Temp.: 102.29 Ambient Temp.: 88.97 F
16:56:04.305 -> Average Cal sensor Temp.: 102.34 Ambient Temp.: 88.97 F
16:56:05.853 -> Average Cal sensor Temp.: 102.82 Ambient Temp.: 88.90 F
16:56:07.437 -> Average Cal sensor Temp.: 103.29 Ambient Temp.: 88.93 F
16:56:08.997 -> Average Cal sensor Temp.: 102.96 Ambient Temp.: 88.93 F
16:56:10.548 -> Average Cal sensor Temp.: 102.94 Ambient Temp.: 88.93 F
16:56:12.135 -> Average Cal sensor Temp.: 102.73 Ambient Temp.: 88.86 F
16:56:13.711 -> Average Cal sensor Temp.: 102.89 Ambient Temp.: 88.90 F
16:56:15.262 -> Average Cal sensor Temp.: 103.19 Ambient Temp.: 88.90 F
16:56:16.817 -> Average Cal sensor Temp.: 103.16 Ambient Temp.: 88.86 F
16:56:18.399 -> Average Cal sensor Temp.: 102.43 Ambient Temp.: 88.75 F
16:56:19.953 -> Average Cal sensor Temp.: 102.41 Ambient Temp.: 88.72 F
16:56:21.536 -> Average Cal sensor Temp.: 102.52 Ambient Temp.: 88.72 F
16:56:23.083 -> Average Cal sensor Temp.: 102.25 Ambient Temp.: 88.72 F
16:56:24.633 -> Average Cal sensor Temp.: 102.68 Ambient Temp.: 88.68 F
16:56:26.221 -> Average Cal sensor Temp.: 102.55 Ambient Temp.: 88.68 F
16:56:27.768 -> Average Cal sensor Temp.: 102.89 Ambient Temp.: 88.68 F
16:56:29.359 -> Average Cal sensor Temp.: 102.20 Ambient Temp.: 88.65 F
16:56:30.915 -> Average Cal sensor Temp.: 102.20 Ambient Temp.: 88.65 F
16:56:32.490 -> Average Cal sensor Temp.: 102.46 Ambient Temp.: 88.57 F
16:56:34.044 -> Average Cal sensor Temp.: 102.70 Ambient Temp.: 88.54 F
16:56:35.595 -> Average Cal sensor Temp.: 102.92 Ambient Temp.: 88.50 F
16:56:37.186 -> Average Cal sensor Temp.: 103.09 Ambient Temp.: 88.50 F
16:56:38.739 -> Average Cal sensor Temp.: 102.25 Ambient Temp.: 88.50 F
16:56:40.289 -> Average Cal sensor Temp.: 101.91 Ambient Temp.: 88.47 F
16:56:41.876 -> Average Cal sensor Temp.: 102.68 Ambient Temp.: 88.39 F
16:56:43.430 -> Average Cal sensor Temp.: 102.71 Ambient Temp.: 88.39 F
16:56:44.989 -> Average Cal sensor Temp.: 102.70 Ambient Temp.: 88.39 F
16:56:46.544 -> Average Cal sensor Temp.: 102.71 Ambient Temp.: 88.32 F
16:56:48.130 -> Average Cal sensor Temp.: 102.51 Ambient Temp.: 88.29 F
16:56:49.687 -> Average Cal sensor Temp.: 102.89 Ambient Temp.: 88.36 F
16:56:51.247 -> Average Cal sensor Temp.: 102.79 Ambient Temp.: 88.25 F
16:56:52.801 -> Average Cal sensor Temp.: 102.81 Ambient Temp.: 88.29 F
16:56:54.387 -> Average Cal sensor Temp.: 103.16 Ambient Temp.: 88.25 F
16:56:55.939 -> Average Cal sensor Temp.: 102.65 Ambient Temp.: 88.25 F
16:56:57.492 -> Average Cal sensor Temp.: 102.67 Ambient Temp.: 88.14 F
16:56:59.072 -> Average Cal sensor Temp.: 102.78 Ambient Temp.: 88.21 F
16:57:00.619 -> Average Cal sensor Temp.: 102.73 Ambient Temp.: 88.21 F
16:57:02.203 -> Average Cal sensor Temp.: 102.76 Ambient Temp.: 88.21 F
16:57:03.750 -> Average Cal sensor Temp.: 102.43 Ambient Temp.: 88.21 F
16:57:05.335 -> Average Cal sensor Temp.: 103.22 Ambient Temp.: 88.21 F
16:57:06.883 -> Average Cal sensor Temp.: 102.71 Ambient Temp.: 88.25 F
16:57:08.438 -> Average Cal sensor Temp.: 102.55 Ambient Temp.: 88.18 F
16:57:10.027 -> Average Cal sensor Temp.: 101.89 Ambient Temp.: 88.18 F
16:57:11.574 -> Average Cal sensor Temp.: 102.22 Ambient Temp.: 88.18 F
16:57:13.163 -> Average Cal sensor Temp.: 102.22 Ambient Temp.: 88.14 F
16:57:14.717 -> Average Cal sensor Temp.: 102.56 Ambient Temp.: 88.14 F
16:57:16.269 -> Average Cal sensor Temp.: 102.49 Ambient Temp.: 88.11 F
16:57:17.824 -> Average Cal sensor Temp.: 103.14 Ambient Temp.: 88.11 F
16:57:19.410 -> Average Cal sensor Temp.: 103.14 Ambient Temp.: 88.11 F
16:57:20.964 -> Average Cal sensor Temp.: 102.92 Ambient Temp.: 88.14 F
16:57:22.545 -> Average Cal sensor Temp.: 102.85 Ambient Temp.: 88.14 F
16:57:24.094 -> Average Cal sensor Temp.: 102.86 Ambient Temp.: 88.11 F
16:57:25.643 -> Average Cal sensor Temp.: 102.82 Ambient Temp.: 88.11 F
16:57:27.231 -> Average Cal sensor Temp.: 102.93 Ambient Temp.: 88.11 F
16:57:28.780 -> Average Cal sensor Temp.: 102.74 Ambient Temp.: 88.11 F
16:57:30.362 -> Average Cal sensor Temp.: 103.17 Ambient Temp.: 88.00 F
16:57:31.920 -> Average Cal sensor Temp.: 103.17 Ambient Temp.: 88.11 F
16:57:33.480 -> Average Cal sensor Temp.: 103.19 Ambient Temp.: 88.11 F
16:57:35.027 -> Average Cal sensor Temp.: 103.07 Ambient Temp.: 88.18 F
16:57:36.612 -> Average Cal sensor Temp.: 102.85 Ambient Temp.: 88.07 F
16:57:38.167 -> Average Cal sensor Temp.: 102.65 Ambient Temp.: 88.11 F
16:57:39.720 -> Average Cal sensor Temp.: 103.16 Ambient Temp.: 88.00 F
16:57:41.311 -> Average Cal sensor Temp.: 103.21 Ambient Temp.: 88.07 F
16:57:42.859 -> Average Cal sensor Temp.: 102.68 Ambient Temp.: 88.11 F
16:57:44.417 -> Average Cal sensor Temp.: 102.38 Ambient Temp.: 88.11 F
16:57:45.972 -> Average Cal sensor Temp.: 102.79 Ambient Temp.: 88.00 F
16:57:47.563 -> Average Cal sensor Temp.: 103.05 Ambient Temp.: 88.07 F
16:57:49.120 -> Average Cal sensor Temp.: 102.56 Ambient Temp.: 88.07 F
16:57:50.677 -> Average Cal sensor Temp.: 102.87 Ambient Temp.: 88.14 F
16:57:52.265 -> Average Cal sensor Temp.: 103.19 Ambient Temp.: 88.14 F
16:57:53.818 -> Average Cal sensor Temp.: 103.14 Ambient Temp.: 88.14 F
16:57:55.369 -> Average Cal sensor Temp.: 103.18 Ambient Temp.: 88.14 F
16:57:56.953 -> Average Cal sensor Temp.: 103.44 Ambient Temp.: 88.14 F
16:57:58.504 -> Average Cal sensor Temp.: 103.02 Ambient Temp.: 88.14 F
16:58:00.057 -> Average Cal sensor Temp.: 102.83 Ambient Temp.: 88.14 F
16:58:01.650 -> Average Cal sensor Temp.: 103.18 Ambient Temp.: 88.14 F
16:58:03.209 -> Average Cal sensor Temp.: 103.51 Ambient Temp.: 88.21 F
16:58:04.763 -> Average Cal sensor Temp.: 103.32 Ambient Temp.: 88.21 F
16:58:06.323 -> Average Cal sensor Temp.: 103.42 Ambient Temp.: 88.21 F
16:58:07.878 -> Average Cal sensor Temp.: 103.20 Ambient Temp.: 88.21 F
16:58:09.439 -> Average Cal sensor Temp.: 103.04 Ambient Temp.: 88.25 F
16:58:11.027 -> Average Cal sensor Temp.: 103.36 Ambient Temp.: 88.25 F
16:58:12.582 -> Average Cal sensor Temp.: 103.76 Ambient Temp.: 88.29 F
16:58:14.136 -> Average Cal sensor Temp.: 103.59 Ambient Temp.: 88.29 F
16:58:15.725 -> Average Cal sensor Temp.: 103.54 Ambient Temp.: 88.32 F
16:58:17.286 -> Average Cal sensor Temp.: 103.25 Ambient Temp.: 88.36 F
16:58:18.841 -> Average Cal sensor Temp.: 103.16 Ambient Temp.: 88.36 F
16:58:20.392 -> Average Cal sensor Temp.: 103.80 Ambient Temp.: 88.36 F
16:58:21.981 -> Average Cal sensor Temp.: 103.56 Ambient Temp.: 88.36 F
16:58:23.534 -> Average Cal sensor Temp.: 102.91 Ambient Temp.: 88.39 F
16:58:25.092 -> Average Cal sensor Temp.: 103.21 Ambient Temp.: 88.36 F
16:58:26.680 -> Average Cal sensor Temp.: 103.47 Ambient Temp.: 88.36 F
16:58:28.231 -> Average Cal sensor Temp.: 103.32 Ambient Temp.: 88.39 F
16:58:29.788 -> Average Cal sensor Temp.: 103.27 Ambient Temp.: 88.36 F
16:58:31.374 -> Average Cal sensor Temp.: 103.46 Ambient Temp.: 88.39 F
16:58:32.930 -> Average Cal sensor Temp.: 103.00 Ambient Temp.: 88.39 F
16:58:34.483 -> Average Cal sensor Temp.: 103.48 Ambient Temp.: 88.47 F
16:58:36.038 -> Average Cal sensor Temp.: 103.61 Ambient Temp.: 88.50 F
16:58:37.624 -> Average Cal sensor Temp.: 103.02 Ambient Temp.: 88.47 F
16:58:39.181 -> Average Cal sensor Temp.: 103.08 Ambient Temp.: 88.39 F
16:58:40.739 -> Average Cal sensor Temp.: 103.13 Ambient Temp.: 88.39 F
16:58:42.329 -> Average Cal sensor Temp.: 103.44 Ambient Temp.: 88.39 F
16:58:43.878 -> Average Cal sensor Temp.: 103.33 Ambient Temp.: 88.47 F
16:58:45.434 -> Average Cal sensor Temp.: 103.63 Ambient Temp.: 88.47 F
16:58:46.990 -> Average Cal sensor Temp.: 103.36 Ambient Temp.: 88.39 F
16:58:48.578 -> Average Cal sensor Temp.: 103.60 Ambient Temp.: 88.47 F
16:58:50.137 -> Average Cal sensor Temp.: 103.42 Ambient Temp.: 88.47 F
16:58:51.692 -> Average Cal sensor Temp.: 103.58 Ambient Temp.: 88.50 F
16:58:53.284 -> Average Cal sensor Temp.: 103.21 Ambient Temp.: 88.47 F
16:58:54.841 -> Average Cal sensor Temp.: 103.31 Ambient Temp.: 88.39 F
16:58:56.400 -> Average Cal sensor Temp.: 103.12 Ambient Temp.: 88.47 F
16:58:57.953 -> Average Cal sensor Temp.: 103.00 Ambient Temp.: 88.36 F
16:58:59.539 -> Average Cal sensor Temp.: 103.37 Ambient Temp.: 88.47 F
16:59:01.088 -> Average Cal sensor Temp.: 103.16 Ambient Temp.: 88.50 F
16:59:02.653 -> Average Cal sensor Temp.: 102.80 Ambient Temp.: 88.39 F
16:59:04.223 -> Average Cal sensor Temp.: 103.13 Ambient Temp.: 88.47 F
16:59:05.776 -> Average Cal sensor Temp.: 103.87 Ambient Temp.: 88.50 F
16:59:07.336 -> Average Cal sensor Temp.: 103.58 Ambient Temp.: 88.47 F
16:59:08.925 -> Average Cal sensor Temp.: 103.10 Ambient Temp.: 88.50 F
16:59:10.482 -> Average Cal sensor Temp.: 103.28 Ambient Temp.: 88.54 F
16:59:12.035 -> Average Cal sensor Temp.: 103.57 Ambient Temp.: 88.50 F
16:59:13.592 -> Average Cal sensor Temp.: 103.80 Ambient Temp.: 88.50 F
16:59:15.185 -> Average Cal sensor Temp.: 103.44 Ambient Temp.: 88.54 F
16:59:16.742 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 88.54 F
16:59:18.296 -> Average Cal sensor Temp.: 103.44 Ambient Temp.: 88.54 F
16:59:19.853 -> Average Cal sensor Temp.: 103.27 Ambient Temp.: 88.57 F
16:59:21.442 -> Average Cal sensor Temp.: 103.84 Ambient Temp.: 88.57 F
16:59:22.999 -> Average Cal sensor Temp.: 103.70 Ambient Temp.: 88.54 F
16:59:24.552 -> Average Cal sensor Temp.: 104.18 Ambient Temp.: 88.65 F
16:59:26.107 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 88.57 F
16:59:27.697 -> Average Cal sensor Temp.: 103.55 Ambient Temp.: 88.65 F
16:59:29.254 -> Average Cal sensor Temp.: 103.40 Ambient Temp.: 88.68 F
16:59:30.802 -> Average Cal sensor Temp.: 103.58 Ambient Temp.: 88.68 F
16:59:32.382 -> Average Cal sensor Temp.: 103.52 Ambient Temp.: 88.65 F
16:59:33.943 -> Average Cal sensor Temp.: 103.71 Ambient Temp.: 88.68 F
16:59:35.528 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 88.68 F
16:59:37.079 -> Average Cal sensor Temp.: 103.48 Ambient Temp.: 88.72 F
16:59:38.634 -> Average Cal sensor Temp.: 103.52 Ambient Temp.: 88.72 F
16:59:40.223 -> Average Cal sensor Temp.: 103.60 Ambient Temp.: 88.68 F
16:59:41.778 -> Average Cal sensor Temp.: 103.47 Ambient Temp.: 88.86 F
16:59:43.334 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 88.75 F
16:59:44.887 -> Average Cal sensor Temp.: 103.70 Ambient Temp.: 88.79 F
16:59:46.471 -> Average Cal sensor Temp.: 103.99 Ambient Temp.: 88.79 F
16:59:48.030 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 88.90 F
16:59:49.584 -> Average Cal sensor Temp.: 103.80 Ambient Temp.: 88.90 F
16:59:51.175 -> Average Cal sensor Temp.: 104.07 Ambient Temp.: 88.97 F
16:59:52.725 -> Average Cal sensor Temp.: 104.10 Ambient Temp.: 88.90 F
16:59:54.278 -> Average Cal sensor Temp.: 103.92 Ambient Temp.: 88.97 F
16:59:55.868 -> Average Cal sensor Temp.: 103.90 Ambient Temp.: 88.93 F
16:59:57.421 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 89.01 F
16:59:58.975 -> Average Cal sensor Temp.: 104.04 Ambient Temp.: 89.04 F
17:00:00.565 -> Average Cal sensor Temp.: 103.95 Ambient Temp.: 89.08 F
17:00:02.118 -> Average Cal sensor Temp.: 103.70 Ambient Temp.: 89.08 F
17:00:03.700 -> Average Cal sensor Temp.: 103.92 Ambient Temp.: 89.11 F
17:00:05.250 -> Average Cal sensor Temp.: 104.19 Ambient Temp.: 89.08 F
17:00:06.801 -> Average Cal sensor Temp.: 103.96 Ambient Temp.: 89.11 F
17:00:08.384 -> Average Cal sensor Temp.: 104.02 Ambient Temp.: 89.11 F
17:00:09.958 -> Average Cal sensor Temp.: 103.91 Ambient Temp.: 89.15 F
17:00:11.511 -> Average Cal sensor Temp.: 103.89 Ambient Temp.: 89.15 F
17:00:13.062 -> Average Cal sensor Temp.: 104.18 Ambient Temp.: 89.29 F
17:00:14.644 -> Average Cal sensor Temp.: 104.16 Ambient Temp.: 89.26 F
17:00:16.199 -> Average Cal sensor Temp.: 103.77 Ambient Temp.: 89.26 F
17:00:17.784 -> Average Cal sensor Temp.: 104.04 Ambient Temp.: 89.29 F
17:00:19.336 -> Average Cal sensor Temp.: 104.02 Ambient Temp.: 89.29 F
17:00:20.893 -> Average Cal sensor Temp.: 103.66 Ambient Temp.: 89.29 F
17:00:22.451 -> Average Cal sensor Temp.: 103.95 Ambient Temp.: 89.40 F
17:00:24.040 -> Average Cal sensor Temp.: 104.11 Ambient Temp.: 89.29 F
17:00:25.593 -> Average Cal sensor Temp.: 104.12 Ambient Temp.: 89.37 F
17:00:27.147 -> Average Cal sensor Temp.: 104.17 Ambient Temp.: 89.44 F
17:00:28.733 -> Average Cal sensor Temp.: 104.08 Ambient Temp.: 89.40 F
17:00:30.279 -> Average Cal sensor Temp.: 103.96 Ambient Temp.: 89.40 F
17:00:31.841 -> Average Cal sensor Temp.: 104.24 Ambient Temp.: 89.47 F
17:00:33.429 -> Average Cal sensor Temp.: 104.05 Ambient Temp.: 89.47 F
17:00:34.982 -> Average Cal sensor Temp.: 103.85 Ambient Temp.: 89.44 F
17:00:36.567 -> Average Cal sensor Temp.: 103.84 Ambient Temp.: 89.51 F
17:00:38.115 -> Average Cal sensor Temp.: 103.49 Ambient Temp.: 89.55 F
17:00:39.698 -> Average Cal sensor Temp.: 103.54 Ambient Temp.: 89.55 F
17:00:41.248 -> Average Cal sensor Temp.: 103.57 Ambient Temp.: 89.55 F
17:00:42.823 -> Average Cal sensor Temp.: 103.64 Ambient Temp.: 89.55 F
17:00:44.378 -> Average Cal sensor Temp.: 103.19 Ambient Temp.: 89.58 F
17:00:45.926 -> Average Cal sensor Temp.: 103.48 Ambient Temp.: 89.65 F
17:00:47.513 -> Average Cal sensor Temp.: 103.56 Ambient Temp.: 89.58 F
17:00:49.065 -> Average Cal sensor Temp.: 103.69 Ambient Temp.: 89.69 F
17:00:50.619 -> Average Cal sensor Temp.: 103.94 Ambient Temp.: 89.69 F
17:00:52.206 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 89.65 F
17:00:53.756 -> Average Cal sensor Temp.: 104.22 Ambient Temp.: 89.65 F
17:00:55.345 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 89.73 F
17:00:56.899 -> Average Cal sensor Temp.: 103.90 Ambient Temp.: 89.69 F
17:00:58.447 -> Average Cal sensor Temp.: 103.88 Ambient Temp.: 89.73 F
17:01:00.036 -> Average Cal sensor Temp.: 103.96 Ambient Temp.: 89.73 F
17:01:01.592 -> Average Cal sensor Temp.: 104.51 Ambient Temp.: 89.76 F
17:01:03.146 -> Average Cal sensor Temp.: 103.60 Ambient Temp.: 89.73 F
17:01:04.729 -> Average Cal sensor Temp.: 102.63 Ambient Temp.: 89.76 F
17:01:06.288 -> Average Cal sensor Temp.: 103.05 Ambient Temp.: 89.69 F
17:01:07.838 -> Average Cal sensor Temp.: 103.58 Ambient Temp.: 89.69 F
17:01:09.434 -> Average Cal sensor Temp.: 103.47 Ambient Temp.: 89.76 F
17:01:10.987 -> Average Cal sensor Temp.: 103.46 Ambient Temp.: 89.69 F
17:01:12.542 -> Average Cal sensor Temp.: 103.33 Ambient Temp.: 89.69 F
17:01:14.104 -> Average Cal sensor Temp.: 103.63 Ambient Temp.: 89.73 F
17:01:15.689 -> Average Cal sensor Temp.: 103.30 Ambient Temp.: 89.69 F
17:01:17.247 -> Average Cal sensor Temp.: 103.34 Ambient Temp.: 89.69 F
17:01:18.800 -> Average Cal sensor Temp.: 103.52 Ambient Temp.: 89.69 F
17:01:20.389 -> Average Cal sensor Temp.: 103.88 Ambient Temp.: 89.69 F
17:01:21.946 -> Average Cal sensor Temp.: 103.47 Ambient Temp.: 89.69 F
17:01:23.501 -> Average Cal sensor Temp.: 103.31 Ambient Temp.: 89.73 F
17:01:25.056 -> Average Cal sensor Temp.: 103.59 Ambient Temp.: 89.69 F
17:01:26.645 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 89.69 F
17:01:28.200 -> Average Cal sensor Temp.: 103.64 Ambient Temp.: 89.69 F
17:01:29.752 -> Average Cal sensor Temp.: 103.30 Ambient Temp.: 89.76 F
17:01:31.340 -> Average Cal sensor Temp.: 103.46 Ambient Temp.: 89.69 F
17:01:32.896 -> Average Cal sensor Temp.: 103.77 Ambient Temp.: 89.69 F
17:01:34.477 -> Average Cal sensor Temp.: 103.41 Ambient Temp.: 89.73 F
17:01:36.032 -> Average Cal sensor Temp.: 103.81 Ambient Temp.: 89.73 F
17:01:37.581 -> Average Cal sensor Temp.: 103.59 Ambient Temp.: 89.73 F
17:01:39.169 -> Average Cal sensor Temp.: 103.82 Ambient Temp.: 89.69 F
17:01:40.727 -> Average Cal sensor Temp.: 103.27 Ambient Temp.: 89.69 F
17:01:42.278 -> Average Cal sensor Temp.: 102.84 Ambient Temp.: 89.58 F
17:01:43.864 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 89.65 F
17:01:45.414 -> Average Cal sensor Temp.: 102.94 Ambient Temp.: 89.58 F
17:01:46.969 -> Average Cal sensor Temp.: 103.35 Ambient Temp.: 89.47 F
17:01:48.557 -> Average Cal sensor Temp.: 103.78 Ambient Temp.: 89.51 F
17:01:50.110 -> Average Cal sensor Temp.: 103.22 Ambient Temp.: 89.51 F
17:01:51.663 -> Average Cal sensor Temp.: 102.58 Ambient Temp.: 89.47 F
17:01:53.253 -> Average Cal sensor Temp.: 102.91 Ambient Temp.: 89.40 F
17:01:54.805 -> Average Cal sensor Temp.: 102.74 Ambient Temp.: 89.37 F
17:01:56.356 -> Average Cal sensor Temp.: 103.23 Ambient Temp.: 89.37 F
17:01:57.941 -> Average Cal sensor Temp.: 103.02 Ambient Temp.: 89.37 F
17:01:59.495 -> Average Cal sensor Temp.: 103.70 Ambient Temp.: 89.40 F
17:02:01.056 -> Average Cal sensor Temp.: 104.06 Ambient Temp.: 89.33 F
17:02:02.613 -> Average Cal sensor Temp.: 103.53 Ambient Temp.: 89.33 F
17:02:04.173 -> Average Cal sensor Temp.: 103.40 Ambient Temp.: 89.33 F
17:02:05.764 -> Average Cal sensor Temp.: 102.99 Ambient Temp.: 89.29 F
17:02:07.324 -> Average Cal sensor Temp.: 103.36 Ambient Temp.: 89.26 F
17:02:08.875 -> Average Cal sensor Temp.: 103.19 Ambient Temp.: 89.26 F
17:02:10.461 -> Average Cal sensor Temp.: 102.84 Ambient Temp.: 89.19 F
17:02:12.006 -> Average Cal sensor Temp.: 103.28 Ambient Temp.: 89.15 F
17:02:13.592 -> Average Cal sensor Temp.: 103.13 Ambient Temp.: 89.11 F
17:02:15.148 -> Average Cal sensor Temp.: 102.89 Ambient Temp.: 89.11 F
17:02:16.693 -> Average Cal sensor Temp.: 102.40 Ambient Temp.: 89.08 F
17:02:18.276 -> Average Cal sensor Temp.: 102.93 Ambient Temp.: 89.04 F
17:02:19.832 -> Average Cal sensor Temp.: 103.25 Ambient Temp.: 89.01 F
17:02:21.384 -> Average Cal sensor Temp.: 103.49 Ambient Temp.: 89.04 F
17:02:22.969 -> Average Cal sensor Temp.: 103.25 Ambient Temp.: 89.01 F
17:02:24.527 -> Average Cal sensor Temp.: 103.13 Ambient Temp.: 88.97 F
17:02:26.079 -> Average Cal sensor Temp.: 103.37 Ambient Temp.: 88.93 F
17:02:27.658 -> Average Cal sensor Temp.: 102.93 Ambient Temp.: 88.90 F
17:02:29.215 -> Average Cal sensor Temp.: 103.35 Ambient Temp.: 88.90 F
17:02:30.800 -> Average Cal sensor Temp.: 103.35 Ambient Temp.: 88.86 F
17:02:32.353 -> Average Cal sensor Temp.: 103.59 Ambient Temp.: 88.86 F
17:02:33.911 -> Average Cal sensor Temp.: 103.91 Ambient Temp.: 88.90 F
17:02:35.495 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 88.90 F
17:02:37.037 -> Average Cal sensor Temp.: 103.77 Ambient Temp.: 88.90 F
17:02:38.603 -> Average Cal sensor Temp.: 103.66 Ambient Temp.: 88.93 F
17:02:40.165 -> Average Cal sensor Temp.: 103.87 Ambient Temp.: 88.90 F
17:02:41.752 -> Average Cal sensor Temp.: 103.89 Ambient Temp.: 88.90 F
17:02:43.304 -> Average Cal sensor Temp.: 103.87 Ambient Temp.: 88.93 F
17:02:44.890 -> Average Cal sensor Temp.: 103.19 Ambient Temp.: 88.90 F
17:02:46.439 -> Average Cal sensor Temp.: 102.73 Ambient Temp.: 88.90 F
17:02:47.995 -> Average Cal sensor Temp.: 102.71 Ambient Temp.: 88.79 F
17:02:49.583 -> Average Cal sensor Temp.: 102.53 Ambient Temp.: 88.86 F
17:02:51.138 -> Average Cal sensor Temp.: 102.73 Ambient Temp.: 88.79 F
17:02:52.687 -> Average Cal sensor Temp.: 103.09 Ambient Temp.: 88.72 F
17:02:54.263 -> Average Cal sensor Temp.: 102.63 Ambient Temp.: 88.68 F
17:02:55.848 -> Average Cal sensor Temp.: 102.67 Ambient Temp.: 88.72 F
17:02:57.399 -> Average Cal sensor Temp.: 103.21 Ambient Temp.: 88.65 F
17:02:58.950 -> Average Cal sensor Temp.: 103.07 Ambient Temp.: 88.68 F
17:03:00.537 -> Average Cal sensor Temp.: 103.46 Ambient Temp.: 88.68 F
17:03:02.086 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 88.68 F
17:03:03.670 -> Average Cal sensor Temp.: 103.48 Ambient Temp.: 88.65 F
17:03:05.221 -> Average Cal sensor Temp.: 102.59 Ambient Temp.: 88.61 F
17:03:06.770 -> Average Cal sensor Temp.: 102.52 Ambient Temp.: 88.61 F
17:03:08.353 -> Average Cal sensor Temp.: 102.67 Ambient Temp.: 88.61 F
17:03:09.907 -> Average Cal sensor Temp.: 103.08 Ambient Temp.: 88.57 F
17:03:11.495 -> Average Cal sensor Temp.: 102.78 Ambient Temp.: 88.57 F
17:03:13.046 -> Average Cal sensor Temp.: 103.03 Ambient Temp.: 88.57 F
17:03:14.602 -> Average Cal sensor Temp.: 103.46 Ambient Temp.: 88.54 F
17:03:16.192 -> Average Cal sensor Temp.: 103.17 Ambient Temp.: 88.57 F
17:03:17.740 -> Average Cal sensor Temp.: 103.06 Ambient Temp.: 88.57 F
17:03:19.286 -> Average Cal sensor Temp.: 103.01 Ambient Temp.: 88.61 F
17:03:20.873 -> Average Cal sensor Temp.: 103.41 Ambient Temp.: 88.57 F
17:03:22.428 -> Average Cal sensor Temp.: 103.35 Ambient Temp.: 88.61 F
17:03:23.981 -> Average Cal sensor Temp.: 103.37 Ambient Temp.: 88.65 F
17:03:25.570 -> Average Cal sensor Temp.: 103.37 Ambient Temp.: 88.65 F
17:03:27.131 -> Average Cal sensor Temp.: 103.04 Ambient Temp.: 88.65 F
17:03:28.682 -> Average Cal sensor Temp.: 102.99 Ambient Temp.: 88.57 F
17:03:30.237 -> Average Cal sensor Temp.: 102.51 Ambient Temp.: 88.57 F
17:03:31.823 -> Average Cal sensor Temp.: 102.44 Ambient Temp.: 88.57 F
17:03:33.374 -> Average Cal sensor Temp.: 102.11 Ambient Temp.: 88.50 F
17:03:34.929 -> Average Cal sensor Temp.: 102.02 Ambient Temp.: 88.50 F
17:03:36.517 -> Average Cal sensor Temp.: 102.51 Ambient Temp.: 88.36 F
17:03:38.072 -> Average Cal sensor Temp.: 102.92 Ambient Temp.: 88.39 F
17:03:39.627 -> Average Cal sensor Temp.: 102.73 Ambient Temp.: 88.39 F
17:03:41.214 -> Average Cal sensor Temp.: 102.93 Ambient Temp.: 88.36 F
17:03:42.766 -> Average Cal sensor Temp.: 103.07 Ambient Temp.: 88.32 F
17:03:44.317 -> Average Cal sensor Temp.: 103.58 Ambient Temp.: 88.36 F
17:03:45.911 -> Average Cal sensor Temp.: 103.10 Ambient Temp.: 88.32 F
17:03:47.465 -> Average Cal sensor Temp.: 103.16 Ambient Temp.: 88.32 F
17:03:49.014 -> Average Cal sensor Temp.: 103.13 Ambient Temp.: 88.29 F
17:03:50.570 -> Average Cal sensor Temp.: 103.06 Ambient Temp.: 88.29 F
17:03:52.158 -> Average Cal sensor Temp.: 103.41 Ambient Temp.: 88.29 F
17:03:53.712 -> Average Cal sensor Temp.: 103.71 Ambient Temp.: 88.32 F
17:03:55.271 -> Average Cal sensor Temp.: 103.25 Ambient Temp.: 88.32 F
17:03:56.857 -> Average Cal sensor Temp.: 103.32 Ambient Temp.: 88.36 F
17:03:58.411 -> Average Cal sensor Temp.: 103.54 Ambient Temp.: 88.32 F
17:03:59.966 -> Average Cal sensor Temp.: 103.45 Ambient Temp.: 88.32 F
17:04:01.523 -> Average Cal sensor Temp.: 103.69 Ambient Temp.: 88.36 F
17:04:03.104 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 88.36 F
17:04:04.659 -> Average Cal sensor Temp.: 103.63 Ambient Temp.: 88.39 F
17:04:06.245 -> Average Cal sensor Temp.: 103.39 Ambient Temp.: 88.36 F
17:04:07.793 -> Average Cal sensor Temp.: 103.28 Ambient Temp.: 88.39 F
17:04:09.346 -> Average Cal sensor Temp.: 103.46 Ambient Temp.: 88.39 F
17:04:10.934 -> Average Cal sensor Temp.: 103.14 Ambient Temp.: 88.36 F
17:04:12.491 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 88.32 F
17:04:14.042 -> Average Cal sensor Temp.: 103.47 Ambient Temp.: 88.36 F
17:04:15.637 -> Average Cal sensor Temp.: 103.51 Ambient Temp.: 88.32 F
17:04:17.193 -> Average Cal sensor Temp.: 103.60 Ambient Temp.: 88.39 F
17:04:18.749 -> Average Cal sensor Temp.: 103.06 Ambient Temp.: 88.29 F
17:04:20.305 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 88.25 F
17:04:21.862 -> Average Cal sensor Temp.: 103.88 Ambient Temp.: 88.32 F
17:04:23.449 -> Average Cal sensor Temp.: 103.32 Ambient Temp.: 88.36 F
17:04:25.008 -> Average Cal sensor Temp.: 103.15 Ambient Temp.: 88.29 F
17:04:26.564 -> Average Cal sensor Temp.: 103.66 Ambient Temp.: 88.32 F
17:04:28.152 -> Average Cal sensor Temp.: 103.55 Ambient Temp.: 88.25 F
17:04:29.707 -> Average Cal sensor Temp.: 103.65 Ambient Temp.: 88.25 F
17:04:31.259 -> Average Cal sensor Temp.: 104.16 Ambient Temp.: 88.25 F
17:04:32.817 -> Average Cal sensor Temp.: 103.69 Ambient Temp.: 88.32 F
17:04:34.408 -> Average Cal sensor Temp.: 103.77 Ambient Temp.: 88.25 F
17:04:35.965 -> Average Cal sensor Temp.: 103.94 Ambient Temp.: 88.29 F
17:04:37.520 -> Average Cal sensor Temp.: 104.08 Ambient Temp.: 88.25 F
17:04:39.075 -> Average Cal sensor Temp.: 104.46 Ambient Temp.: 88.29 F
17:04:40.664 -> Average Cal sensor Temp.: 103.83 Ambient Temp.: 88.29 F
17:04:42.219 -> Average Cal sensor Temp.: 103.91 Ambient Temp.: 88.29 F
17:04:43.772 -> Average Cal sensor Temp.: 104.02 Ambient Temp.: 88.25 F
17:04:45.364 -> Average Cal sensor Temp.: 104.12 Ambient Temp.: 88.29 F
17:04:46.916 -> Average Cal sensor Temp.: 103.94 Ambient Temp.: 88.29 F
17:04:48.498 -> Average Cal sensor Temp.: 103.80 Ambient Temp.: 88.29 F
17:04:50.047 -> Average Cal sensor Temp.: 104.14 Ambient Temp.: 88.32 F
17:04:51.605 -> Average Cal sensor Temp.: 104.26 Ambient Temp.: 88.32 F
17:04:53.160 -> Average Cal sensor Temp.: 104.32 Ambient Temp.: 88.39 F
17:04:54.745 -> Average Cal sensor Temp.: 104.17 Ambient Temp.: 88.32 F
17:04:56.297 -> Average Cal sensor Temp.: 103.95 Ambient Temp.: 88.39 F
17:04:57.885 -> Average Cal sensor Temp.: 103.74 Ambient Temp.: 88.36 F
17:04:59.438 -> Average Cal sensor Temp.: 103.52 Ambient Temp.: 88.39 F
17:05:00.992 -> Average Cal sensor Temp.: 104.09 Ambient Temp.: 88.39 F
17:05:02.574 -> Average Cal sensor Temp.: 103.90 Ambient Temp.: 88.47 F
17:05:04.122 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 88.39 F
17:05:05.708 -> Average Cal sensor Temp.: 104.11 Ambient Temp.: 88.39 F
17:05:07.267 -> Average Cal sensor Temp.: 103.42 Ambient Temp.: 88.39 F
17:05:08.818 -> Average Cal sensor Temp.: 103.46 Ambient Temp.: 88.47 F
17:05:10.402 -> Average Cal sensor Temp.: 103.46 Ambient Temp.: 88.32 F
17:05:11.958 -> Average Cal sensor Temp.: 103.78 Ambient Temp.: 88.47 F
17:05:13.515 -> Average Cal sensor Temp.: 104.44 Ambient Temp.: 88.39 F
17:05:15.072 -> Average Cal sensor Temp.: 105.23 Ambient Temp.: 88.47 F
17:05:16.658 -> Average Cal sensor Temp.: 104.92 Ambient Temp.: 88.47 F
17:05:18.215 -> Average Cal sensor Temp.: 104.53 Ambient Temp.: 88.39 F
17:05:19.773 -> Average Cal sensor Temp.: 104.48 Ambient Temp.: 88.39 F
17:05:21.363 -> Average Cal sensor Temp.: 104.10 Ambient Temp.: 88.50 F
17:05:22.919 -> Average Cal sensor Temp.: 104.18 Ambient Temp.: 88.47 F
17:05:24.476 -> Average Cal sensor Temp.: 104.35 Ambient Temp.: 88.47 F
17:05:26.035 -> Average Cal sensor Temp.: 104.37 Ambient Temp.: 88.50 F
17:05:27.594 -> Average Cal sensor Temp.: 103.81 Ambient Temp.: 88.54 F
17:05:29.181 -> Average Cal sensor Temp.: 103.98 Ambient Temp.: 88.47 F
17:05:30.734 -> Average Cal sensor Temp.: 103.70 Ambient Temp.: 88.50 F
17:05:32.286 -> Average Cal sensor Temp.: 103.94 Ambient Temp.: 88.50 F
17:05:33.883 -> Average Cal sensor Temp.: 104.11 Ambient Temp.: 88.50 F
17:05:35.435 -> Average Cal sensor Temp.: 104.23 Ambient Temp.: 88.50 F
17:05:36.994 -> Average Cal sensor Temp.: 104.15 Ambient Temp.: 88.50 F
17:05:38.548 -> Average Cal sensor Temp.: 104.11 Ambient Temp.: 88.54 F
17:05:40.139 -> Average Cal sensor Temp.: 103.88 Ambient Temp.: 88.57 F
17:05:41.699 -> Average Cal sensor Temp.: 104.25 Ambient Temp.: 88.61 F
17:05:43.254 -> Average Cal sensor Temp.: 103.90 Ambient Temp.: 88.65 F
17:05:44.813 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 88.61 F
17:05:46.373 -> Average Cal sensor Temp.: 103.87 Ambient Temp.: 88.61 F
17:05:47.964 -> Average Cal sensor Temp.: 104.16 Ambient Temp.: 88.65 F
17:05:49.525 -> Average Cal sensor Temp.: 104.03 Ambient Temp.: 88.61 F
17:05:51.084 -> Average Cal sensor Temp.: 104.11 Ambient Temp.: 88.68 F
17:05:52.631 -> Average Cal sensor Temp.: 104.01 Ambient Temp.: 88.72 F
17:05:54.216 -> Average Cal sensor Temp.: 104.30 Ambient Temp.: 88.72 F
17:05:55.769 -> Average Cal sensor Temp.: 104.16 Ambient Temp.: 88.68 F
17:05:57.348 -> Average Cal sensor Temp.: 104.40 Ambient Temp.: 88.75 F
17:05:58.918 -> Average Cal sensor Temp.: 104.07 Ambient Temp.: 88.75 F
17:06:00.463 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 88.79 F
17:06:02.044 -> Average Cal sensor Temp.: 103.93 Ambient Temp.: 88.75 F
17:06:03.582 -> Average Cal sensor Temp.: 104.55 Ambient Temp.: 88.86 F
17:06:05.162 -> Average Cal sensor Temp.: 103.96 Ambient Temp.: 88.90 F
17:06:06.738 -> Average Cal sensor Temp.: 103.87 Ambient Temp.: 88.86 F
17:06:08.286 -> Average Cal sensor Temp.: 103.91 Ambient Temp.: 88.93 F
17:06:09.863 -> Average Cal sensor Temp.: 104.06 Ambient Temp.: 88.93 F
17:06:11.412 -> Average Cal sensor Temp.: 104.18 Ambient Temp.: 88.86 F
17:06:12.998 -> Average Cal sensor Temp.: 104.34 Ambient Temp.: 88.93 F
17:06:14.552 -> Average Cal sensor Temp.: 104.14 Ambient Temp.: 88.90 F
17:06:16.135 -> Average Cal sensor Temp.: 104.31 Ambient Temp.: 88.93 F
17:06:17.686 -> Average Cal sensor Temp.: 104.48 Ambient Temp.: 88.97 F
17:06:19.237 -> Average Cal sensor Temp.: 104.55 Ambient Temp.: 89.01 F
17:06:20.823 -> Average Cal sensor Temp.: 104.25 Ambient Temp.: 89.04 F
17:06:22.374 -> Average Cal sensor Temp.: 104.36 Ambient Temp.: 89.04 F
17:06:23.964 -> Average Cal sensor Temp.: 103.95 Ambient Temp.: 89.04 F
17:06:25.516 -> Average Cal sensor Temp.: 104.47 Ambient Temp.: 89.08 F
17:06:27.069 -> Average Cal sensor Temp.: 104.27 Ambient Temp.: 89.04 F
17:06:28.658 -> Average Cal sensor Temp.: 104.02 Ambient Temp.: 89.11 F
17:06:30.216 -> Average Cal sensor Temp.: 104.14 Ambient Temp.: 89.11 F
17:06:31.766 -> Average Cal sensor Temp.: 103.90 Ambient Temp.: 89.04 F
17:06:33.355 -> Average Cal sensor Temp.: 104.01 Ambient Temp.: 89.08 F
17:06:34.908 -> Average Cal sensor Temp.: 104.24 Ambient Temp.: 89.11 F
17:06:36.457 -> Average Cal sensor Temp.: 104.54 Ambient Temp.: 89.15 F
17:06:38.045 -> Average Cal sensor Temp.: 104.36 Ambient Temp.: 89.15 F
17:06:39.595 -> Average Cal sensor Temp.: 104.17 Ambient Temp.: 89.19 F
17:06:41.148 -> Average Cal sensor Temp.: 104.27 Ambient Temp.: 89.26 F
17:06:42.733 -> Average Cal sensor Temp.: 104.33 Ambient Temp.: 89.26 F
17:06:44.290 -> Average Cal sensor Temp.: 104.13 Ambient Temp.: 89.26 F
17:06:45.842 -> Average Cal sensor Temp.: 104.16 Ambient Temp.: 89.26 F
17:06:47.430 -> Average Cal sensor Temp.: 104.29 Ambient Temp.: 89.19 F
17:06:48.975 -> Average Cal sensor Temp.: 104.35 Ambient Temp.: 89.29 F
17:06:50.565 -> Average Cal sensor Temp.: 104.27 Ambient Temp.: 89.33 F
17:06:52.126 -> Average Cal sensor Temp.: 104.32 Ambient Temp.: 89.29 F
17:06:53.677 -> Average Cal sensor Temp.: 104.18 Ambient Temp.: 89.33 F
17:06:55.236 -> Average Cal sensor Temp.: 104.35 Ambient Temp.: 89.37 F
17:06:56.824 -> Average Cal sensor Temp.: 104.43 Ambient Temp.: 89.37 F
17:06:58.384 -> Average Cal sensor Temp.: 104.21 Ambient Temp.: 89.40 F
17:06:59.939 -> Average Cal sensor Temp.: 104.02 Ambient Temp.: 89.44 F
17:07:01.492 -> Average Cal sensor Temp.: 103.77 Ambient Temp.: 89.44 F
17:07:03.083 -> Average Cal sensor Temp.: 104.02 Ambient Temp.: 89.44 F
17:07:04.630 -> Average Cal sensor Temp.: 104.09 Ambient Temp.: 89.44 F
17:07:06.185 -> Average Cal sensor Temp.: 103.90 Ambient Temp.: 89.44 F
17:07:07.770 -> Average Cal sensor Temp.: 104.22 Ambient Temp.: 89.51 F
17:07:09.325 -> Average Cal sensor Temp.: 104.04 Ambient Temp.: 89.47 F
17:07:10.882 -> Average Cal sensor Temp.: 103.91 Ambient Temp.: 89.47 F
17:07:12.472 -> Average Cal sensor Temp.: 103.91 Ambient Temp.: 89.51 F
17:07:14.026 -> Average Cal sensor Temp.: 104.47 Ambient Temp.: 89.55 F
17:07:15.582 -> Average Cal sensor Temp.: 104.10 Ambient Temp.: 89.51 F
17:07:17.139 -> Average Cal sensor Temp.: 103.93 Ambient Temp.: 89.55 F
17:07:18.728 -> Average Cal sensor Temp.: 104.09 Ambient Temp.: 89.65 F
17:07:20.284 -> Average Cal sensor Temp.: 103.94 Ambient Temp.: 89.51 F
17:07:21.841 -> Average Cal sensor Temp.: 103.98 Ambient Temp.: 89.65 F
17:07:23.400 -> Average Cal sensor Temp.: 103.65 Ambient Temp.: 89.55 F
17:07:24.987 -> Average Cal sensor Temp.: 104.18 Ambient Temp.: 89.69 F
17:07:26.543 -> Average Cal sensor Temp.: 104.14 Ambient Temp.: 89.65 F
17:07:28.096 -> Average Cal sensor Temp.: 103.98 Ambient Temp.: 89.69 F
17:07:29.687 -> Average Cal sensor Temp.: 104.15 Ambient Temp.: 89.65 F
17:07:31.238 -> Average Cal sensor Temp.: 104.45 Ambient Temp.: 89.65 F
17:07:32.792 -> Average Cal sensor Temp.: 103.98 Ambient Temp.: 89.65 F
17:07:34.383 -> Average Cal sensor Temp.: 103.88 Ambient Temp.: 89.65 F
17:07:35.937 -> Average Cal sensor Temp.: 103.85 Ambient Temp.: 89.69 F
17:07:37.491 -> Average Cal sensor Temp.: 103.95 Ambient Temp.: 89.65 F
17:07:39.081 -> Average Cal sensor Temp.: 103.83 Ambient Temp.: 89.58 F
17:07:40.640 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 89.58 F
17:07:42.200 -> Average Cal sensor Temp.: 103.91 Ambient Temp.: 89.58 F
17:07:43.754 -> Average Cal sensor Temp.: 104.10 Ambient Temp.: 89.58 F
17:07:45.339 -> Average Cal sensor Temp.: 103.97 Ambient Temp.: 89.55 F
17:07:46.895 -> Average Cal sensor Temp.: 104.09 Ambient Temp.: 89.65 F
17:07:48.450 -> Average Cal sensor Temp.: 104.09 Ambient Temp.: 89.58 F
17:07:50.039 -> Average Cal sensor Temp.: 103.76 Ambient Temp.: 89.58 F
17:07:51.591 -> Average Cal sensor Temp.: 103.28 Ambient Temp.: 89.55 F
17:07:53.146 -> Average Cal sensor Temp.: 103.36 Ambient Temp.: 89.47 F
17:07:54.732 -> Average Cal sensor Temp.: 103.60 Ambient Temp.: 89.51 F
17:07:56.287 -> Average Cal sensor Temp.: 103.55 Ambient Temp.: 89.51 F
17:07:57.837 -> Average Cal sensor Temp.: 103.68 Ambient Temp.: 89.44 F
17:07:58.459 -> Ambient Temp.: 89.44 F
17:07:59.010 -> Ambient Temp.: 89.44 F
17:07:59.598 -> Ambient Temp.: 89.44 F
17:08:00.182 -> Ambient Temp.: 89.44 F
17:08:01.118 -> Ambient Temp.: 89.47 F
17:08:02.672 -> Average Cal sensor Temp.: 104.97 Ambient Temp.: 89.51 F
17:08:04.261 -> Average Cal sensor Temp.: 104.47 Ambient Temp.: 89.47 F
17:08:05.815 -> Average Cal sensor Temp.: 104.56 Ambient Temp.: 89.44 F
17:08:07.361 -> Average Cal sensor Temp.: 104.20 Ambient Temp.: 89.44 F
17:08:08.947 -> Average Cal sensor Temp.: 104.15 Ambient Temp.: 89.40 F
17:08:10.498 -> Average Cal sensor Temp.: 104.04 Ambient Temp.: 89.44 F
17:08:12.090 -> Average Cal sensor Temp.: 104.21 Ambient Temp.: 89.37 F
17:08:13.652 -> Average Cal sensor Temp.: 104.23 Ambient Temp.: 89.40 F
17:08:15.205 -> Average Cal sensor Temp.: 104.18 Ambient Temp.: 89.37 F
17:08:16.758 -> Average Cal sensor Temp.: 104.11 Ambient Temp.: 89.40 F
17:08:18.316 -> Average Cal sensor Temp.: 104.44 Ambient Temp.: 89.37 F
17:08:19.906 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 89.33 F
17:08:21.457 -> Average Cal sensor Temp.: 103.92 Ambient Temp.: 89.29 F
17:08:23.044 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 89.33 F
17:08:24.601 -> Average Cal sensor Temp.: 103.67 Ambient Temp.: 89.37 F
17:08:26.157 -> Average Cal sensor Temp.: 103.49 Ambient Temp.: 89.26 F
17:08:27.710 -> Average Cal sensor Temp.: 103.81 Ambient Temp.: 89.19 F
17:08:29.297 -> Average Cal sensor Temp.: 103.76 Ambient Temp.: 89.26 F
17:08:30.850 -> Average Cal sensor Temp.: 103.54 Ambient Temp.: 89.15 F
17:08:32.408 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 89.15 F
17:08:33.964 -> Average Cal sensor Temp.: 104.18 Ambient Temp.: 89.15 F
17:08:35.548 -> Average Cal sensor Temp.: 103.47 Ambient Temp.: 89.15 F
17:08:37.101 -> Average Cal sensor Temp.: 103.77 Ambient Temp.: 89.08 F
17:08:38.658 -> Average Cal sensor Temp.: 103.58 Ambient Temp.: 89.08 F
17:08:40.246 -> Average Cal sensor Temp.: 104.09 Ambient Temp.: 89.08 F
17:08:41.802 -> Average Cal sensor Temp.: 104.04 Ambient Temp.: 89.08 F
17:08:43.358 -> Average Cal sensor Temp.: 103.46 Ambient Temp.: 89.01 F
17:08:44.948 -> Average Cal sensor Temp.: 102.96 Ambient Temp.: 88.97 F
17:08:46.504 -> Average Cal sensor Temp.: 103.53 Ambient Temp.: 88.93 F
17:08:48.058 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 88.93 F
17:08:49.648 -> Average Cal sensor Temp.: 103.94 Ambient Temp.: 88.93 F
17:08:51.200 -> Average Cal sensor Temp.: 103.63 Ambient Temp.: 88.93 F
17:08:52.751 -> Average Cal sensor Temp.: 103.52 Ambient Temp.: 88.90 F
17:08:54.339 -> Average Cal sensor Temp.: 103.70 Ambient Temp.: 88.79 F
17:08:55.886 -> Average Cal sensor Temp.: 103.96 Ambient Temp.: 88.86 F
17:08:57.469 -> Average Cal sensor Temp.: 103.71 Ambient Temp.: 88.90 F
17:08:59.022 -> Average Cal sensor Temp.: 103.94 Ambient Temp.: 88.75 F
17:09:00.575 -> Average Cal sensor Temp.: 103.85 Ambient Temp.: 88.72 F
17:09:02.168 -> Average Cal sensor Temp.: 104.10 Ambient Temp.: 88.75 F
17:09:03.724 -> Average Cal sensor Temp.: 104.29 Ambient Temp.: 88.79 F
17:09:05.275 -> Average Cal sensor Temp.: 104.01 Ambient Temp.: 88.75 F
17:09:06.861 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 88.68 F
17:09:08.412 -> Average Cal sensor Temp.: 103.10 Ambient Temp.: 88.75 F
17:09:09.966 -> Average Cal sensor Temp.: 103.81 Ambient Temp.: 88.68 F
17:09:11.525 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 88.61 F
17:09:13.111 -> Average Cal sensor Temp.: 103.32 Ambient Temp.: 88.65 F
17:09:14.659 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 88.65 F
17:09:16.245 -> Average Cal sensor Temp.: 103.89 Ambient Temp.: 88.57 F
17:09:17.800 -> Average Cal sensor Temp.: 103.09 Ambient Temp.: 88.57 F
17:09:19.357 -> Average Cal sensor Temp.: 103.11 Ambient Temp.: 88.50 F
17:09:20.916 -> Average Cal sensor Temp.: 103.48 Ambient Temp.: 88.50 F
17:09:22.476 -> Average Cal sensor Temp.: 103.72 Ambient Temp.: 88.39 F
17:09:24.064 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 88.36 F
17:09:25.622 -> Average Cal sensor Temp.: 103.04 Ambient Temp.: 88.47 F
17:09:27.179 -> Average Cal sensor Temp.: 103.18 Ambient Temp.: 88.29 F
17:09:28.762 -> Average Cal sensor Temp.: 103.39 Ambient Temp.: 88.32 F
17:09:30.314 -> Average Cal sensor Temp.: 103.53 Ambient Temp.: 88.25 F
17:09:31.867 -> Average Cal sensor Temp.: 103.59 Ambient Temp.: 88.29 F
17:09:33.456 -> Average Cal sensor Temp.: 103.94 Ambient Temp.: 88.29 F
17:09:35.010 -> Average Cal sensor Temp.: 104.18 Ambient Temp.: 88.25 F
17:09:36.563 -> Average Cal sensor Temp.: 103.94 Ambient Temp.: 88.25 F
17:09:38.152 -> Average Cal sensor Temp.: 104.11 Ambient Temp.: 88.29 F
17:09:39.707 -> Average Cal sensor Temp.: 103.35 Ambient Temp.: 88.25 F
17:09:41.264 -> Average Cal sensor Temp.: 103.83 Ambient Temp.: 88.21 F
17:09:42.816 -> Average Cal sensor Temp.: 103.93 Ambient Temp.: 88.21 F
17:09:44.405 -> Average Cal sensor Temp.: 103.51 Ambient Temp.: 88.21 F
17:09:45.962 -> Average Cal sensor Temp.: 103.89 Ambient Temp.: 88.25 F
17:09:47.520 -> Average Cal sensor Temp.: 103.66 Ambient Temp.: 88.21 F
17:09:49.072 -> Average Cal sensor Temp.: 103.14 Ambient Temp.: 88.14 F
17:09:50.662 -> Average Cal sensor Temp.: 103.72 Ambient Temp.: 88.14 F
17:09:52.218 -> Average Cal sensor Temp.: 103.64 Ambient Temp.: 88.14 F
17:09:53.775 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 88.00 F
17:09:55.363 -> Average Cal sensor Temp.: 103.70 Ambient Temp.: 88.11 F
17:09:56.922 -> Average Cal sensor Temp.: 103.17 Ambient Temp.: 88.00 F
17:09:58.478 -> Average Cal sensor Temp.: 103.18 Ambient Temp.: 87.96 F
17:10:00.034 -> Average Cal sensor Temp.: 103.85 Ambient Temp.: 88.00 F
17:10:01.593 -> Average Cal sensor Temp.: 103.77 Ambient Temp.: 87.93 F
17:10:03.179 -> Average Cal sensor Temp.: 103.43 Ambient Temp.: 87.93 F
17:10:04.731 -> Average Cal sensor Temp.: 104.06 Ambient Temp.: 87.93 F
17:10:06.289 -> Average Cal sensor Temp.: 104.77 Ambient Temp.: 87.96 F
17:10:07.878 -> Average Cal sensor Temp.: 104.16 Ambient Temp.: 87.96 F
17:10:09.435 -> Average Cal sensor Temp.: 103.90 Ambient Temp.: 88.00 F
17:10:10.991 -> Average Cal sensor Temp.: 104.20 Ambient Temp.: 87.96 F
17:10:12.574 -> Average Cal sensor Temp.: 104.23 Ambient Temp.: 87.96 F
17:10:14.135 -> Average Cal sensor Temp.: 104.16 Ambient Temp.: 87.93 F
17:10:15.687 -> Average Cal sensor Temp.: 104.23 Ambient Temp.: 87.93 F
17:10:17.238 -> Average Cal sensor Temp.: 103.83 Ambient Temp.: 87.96 F
17:10:18.829 -> Average Cal sensor Temp.: 103.24 Ambient Temp.: 87.93 F
17:10:20.384 -> Average Cal sensor Temp.: 103.15 Ambient Temp.: 87.89 F
17:10:21.941 -> Average Cal sensor Temp.: 103.35 Ambient Temp.: 87.89 F
17:10:23.497 -> Average Cal sensor Temp.: 103.49 Ambient Temp.: 87.82 F
17:10:25.078 -> Average Cal sensor Temp.: 103.80 Ambient Temp.: 87.85 F
17:10:26.637 -> Average Cal sensor Temp.: 103.90 Ambient Temp.: 87.82 F
17:10:28.192 -> Average Cal sensor Temp.: 103.46 Ambient Temp.: 87.78 F
17:10:29.777 -> Average Cal sensor Temp.: 103.41 Ambient Temp.: 87.75 F
17:10:31.333 -> Average Cal sensor Temp.: 103.87 Ambient Temp.: 87.75 F
17:10:32.882 -> Average Cal sensor Temp.: 103.95 Ambient Temp.: 87.78 F
17:10:34.471 -> Average Cal sensor Temp.: 103.63 Ambient Temp.: 87.75 F
17:10:36.027 -> Average Cal sensor Temp.: 104.06 Ambient Temp.: 87.71 F
17:10:37.583 -> Average Cal sensor Temp.: 104.05 Ambient Temp.: 87.75 F
17:10:39.139 -> Average Cal sensor Temp.: 104.06 Ambient Temp.: 87.75 F
17:10:40.729 -> Average Cal sensor Temp.: 104.18 Ambient Temp.: 87.75 F
17:10:42.282 -> Average Cal sensor Temp.: 103.88 Ambient Temp.: 87.71 F
17:10:43.831 -> Average Cal sensor Temp.: 103.93 Ambient Temp.: 87.75 F
17:10:45.414 -> Average Cal sensor Temp.: 104.19 Ambient Temp.: 87.75 F
17:10:46.973 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 87.71 F
17:10:48.519 -> Average Cal sensor Temp.: 104.22 Ambient Temp.: 87.75 F
17:10:50.111 -> Average Cal sensor Temp.: 104.98 Ambient Temp.: 87.75 F
17:10:51.669 -> Average Cal sensor Temp.: 104.79 Ambient Temp.: 87.82 F
17:10:53.227 -> Average Cal sensor Temp.: 104.56 Ambient Temp.: 87.78 F
17:10:54.781 -> Average Cal sensor Temp.: 104.34 Ambient Temp.: 87.85 F
17:10:56.372 -> Average Cal sensor Temp.: 104.19 Ambient Temp.: 87.85 F
17:10:57.929 -> Average Cal sensor Temp.: 104.24 Ambient Temp.: 87.85 F
17:10:59.479 -> Average Cal sensor Temp.: 104.46 Ambient Temp.: 87.89 F
17:11:01.061 -> Average Cal sensor Temp.: 104.24 Ambient Temp.: 87.96 F
17:11:02.619 -> Average Cal sensor Temp.: 104.65 Ambient Temp.: 87.96 F
17:11:04.174 -> Average Cal sensor Temp.: 104.44 Ambient Temp.: 87.96 F
17:11:05.762 -> Average Cal sensor Temp.: 104.34 Ambient Temp.: 88.07 F
17:11:07.321 -> Average Cal sensor Temp.: 103.96 Ambient Temp.: 88.07 F
17:11:08.877 -> Average Cal sensor Temp.: 104.19 Ambient Temp.: 88.00 F
17:11:10.429 -> Average Cal sensor Temp.: 104.16 Ambient Temp.: 88.07 F
17:11:12.021 -> Average Cal sensor Temp.: 104.10 Ambient Temp.: 88.00 F
17:11:13.575 -> Average Cal sensor Temp.: 104.19 Ambient Temp.: 88.07 F
17:11:15.130 -> Average Cal sensor Temp.: 104.55 Ambient Temp.: 88.11 F
17:11:16.719 -> Average Cal sensor Temp.: 104.40 Ambient Temp.: 88.14 F
17:11:18.274 -> Average Cal sensor Temp.: 104.36 Ambient Temp.: 88.11 F
17:11:19.836 -> Average Cal sensor Temp.: 104.52 Ambient Temp.: 88.14 F
17:11:21.386 -> Average Cal sensor Temp.: 104.63 Ambient Temp.: 88.11 F
17:11:22.973 -> Average Cal sensor Temp.: 104.91 Ambient Temp.: 88.14 F
17:11:24.525 -> Average Cal sensor Temp.: 104.69 Ambient Temp.: 88.14 F
17:11:26.076 -> Average Cal sensor Temp.: 104.52 Ambient Temp.: 88.14 F
17:11:27.664 -> Average Cal sensor Temp.: 104.63 Ambient Temp.: 88.18 F
17:11:29.219 -> Average Cal sensor Temp.: 104.79 Ambient Temp.: 88.18 F
17:11:30.772 -> Average Cal sensor Temp.: 104.79 Ambient Temp.: 88.18 F
17:11:32.363 -> Average Cal sensor Temp.: 104.49 Ambient Temp.: 88.25 F
17:11:33.915 -> Average Cal sensor Temp.: 104.14 Ambient Temp.: 88.18 F
17:11:35.499 -> Average Cal sensor Temp.: 104.15 Ambient Temp.: 88.21 F
17:11:37.052 -> Average Cal sensor Temp.: 104.37 Ambient Temp.: 88.18 F
17:11:38.610 -> Average Cal sensor Temp.: 104.24 Ambient Temp.: 88.18 F
17:11:40.164 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 88.18 F
17:11:41.759 -> Average Cal sensor Temp.: 104.32 Ambient Temp.: 88.14 F
17:11:43.314 -> Average Cal sensor Temp.: 104.18 Ambient Temp.: 88.14 F
17:11:44.866 -> Average Cal sensor Temp.: 104.52 Ambient Temp.: 88.18 F
17:11:46.453 -> Average Cal sensor Temp.: 104.54 Ambient Temp.: 88.21 F
17:11:48.007 -> Average Cal sensor Temp.: 104.80 Ambient Temp.: 88.21 F
17:11:49.564 -> Average Cal sensor Temp.: 104.37 Ambient Temp.: 88.18 F
17:11:51.118 -> Average Cal sensor Temp.: 104.53 Ambient Temp.: 88.25 F
17:11:52.710 -> Average Cal sensor Temp.: 104.59 Ambient Temp.: 88.21 F
17:11:54.264 -> Average Cal sensor Temp.: 104.45 Ambient Temp.: 88.21 F
17:11:55.814 -> Average Cal sensor Temp.: 104.56 Ambient Temp.: 88.21 F
17:11:57.372 -> Average Cal sensor Temp.: 104.48 Ambient Temp.: 88.21 F
17:11:58.960 -> Average Cal sensor Temp.: 104.39 Ambient Temp.: 88.29 F
17:12:00.516 -> Average Cal sensor Temp.: 104.55 Ambient Temp.: 88.21 F
17:12:02.071 -> Average Cal sensor Temp.: 104.51 Ambient Temp.: 88.18 F
17:12:03.664 -> Average Cal sensor Temp.: 105.11 Ambient Temp.: 88.25 F
17:12:05.214 -> Average Cal sensor Temp.: 105.13 Ambient Temp.: 88.29 F
17:12:06.768 -> Average Cal sensor Temp.: 104.91 Ambient Temp.: 88.25 F
17:12:08.354 -> Average Cal sensor Temp.: 104.90 Ambient Temp.: 88.29 F
17:12:09.909 -> Average Cal sensor Temp.: 105.25 Ambient Temp.: 88.32 F
17:12:11.465 -> Average Cal sensor Temp.: 104.89 Ambient Temp.: 88.36 F
17:12:13.020 -> Average Cal sensor Temp.: 105.08 Ambient Temp.: 88.32 F
17:12:14.615 -> Average Cal sensor Temp.: 104.97 Ambient Temp.: 88.47 F
17:12:16.172 -> Average Cal sensor Temp.: 104.77 Ambient Temp.: 88.47 F
17:12:17.729 -> Average Cal sensor Temp.: 104.72 Ambient Temp.: 88.39 F
17:12:19.309 -> Average Cal sensor Temp.: 104.81 Ambient Temp.: 88.39 F
17:12:20.867 -> Average Cal sensor Temp.: 104.61 Ambient Temp.: 88.50 F
17:12:22.425 -> Average Cal sensor Temp.: 104.58 Ambient Temp.: 88.50 F
17:12:23.010 -> Ambient Temp.: 88.50 F
17:12:24.570 -> Average Cal sensor Temp.: 105.18 Ambient Temp.: 88.50 F
17:12:26.121 -> Average Cal sensor Temp.: 105.12 Ambient Temp.: 88.50 F
17:12:27.707 -> Average Cal sensor Temp.: 105.21 Ambient Temp.: 88.57 F
17:12:29.263 -> Average Cal sensor Temp.: 104.98 Ambient Temp.: 88.57 F
17:12:30.822 -> Average Cal sensor Temp.: 104.74 Ambient Temp.: 88.57 F
17:12:32.409 -> Average Cal sensor Temp.: 104.90 Ambient Temp.: 88.61 F
17:12:33.965 -> Average Cal sensor Temp.: 104.78 Ambient Temp.: 88.61 F
17:12:35.528 -> Average Cal sensor Temp.: 104.77 Ambient Temp.: 88.61 F
17:12:37.088 -> Average Cal sensor Temp.: 104.72 Ambient Temp.: 88.65 F
17:12:38.648 -> Average Cal sensor Temp.: 105.01 Ambient Temp.: 88.65 F
17:12:40.232 -> Average Cal sensor Temp.: 104.91 Ambient Temp.: 88.68 F
17:12:40.822 -> Ambient Temp.: 88.68 F
17:12:41.378 -> Ambient Temp.: 88.68 F
17:12:41.962 -> Ambient Temp.: 88.68 F
17:12:42.552 -> Ambient Temp.: 88.68 F
17:12:43.139 -> Ambient Temp.: 88.68 F
17:12:44.693 -> Average Cal sensor Temp.: 101.69 Ambient Temp.: 88.79 F
17:12:46.249 -> Average Cal sensor Temp.: 105.57 Ambient Temp.: 88.90 F
17:12:47.805 -> Average Cal sensor Temp.: 105.64 Ambient Temp.: 88.97 F
17:12:49.394 -> Average Cal sensor Temp.: 105.14 Ambient Temp.: 88.93 F
17:12:50.947 -> Average Cal sensor Temp.: 105.34 Ambient Temp.: 88.97 F
17:12:52.500 -> Average Cal sensor Temp.: 105.49 Ambient Temp.: 88.97 F
17:12:54.090 -> Average Cal sensor Temp.: 105.30 Ambient Temp.: 89.04 F
17:12:55.643 -> Average Cal sensor Temp.: 104.85 Ambient Temp.: 89.04 F
17:12:57.197 -> Average Cal sensor Temp.: 105.05 Ambient Temp.: 89.04 F
17:12:58.773 -> Average Cal sensor Temp.: 105.28 Ambient Temp.: 89.11 F
17:13:00.357 -> Average Cal sensor Temp.: 104.82 Ambient Temp.: 89.08 F
17:13:01.917 -> Average Cal sensor Temp.: 105.09 Ambient Temp.: 89.19 F
17:13:03.468 -> Average Cal sensor Temp.: 104.90 Ambient Temp.: 89.15 F
17:13:05.054 -> Average Cal sensor Temp.: 105.13 Ambient Temp.: 89.19 F
17:13:06.608 -> Average Cal sensor Temp.: 105.35 Ambient Temp.: 89.19 F
17:13:08.162 -> Average Cal sensor Temp.: 105.13 Ambient Temp.: 89.19 F
17:13:09.713 -> Average Cal sensor Temp.: 105.20 Ambient Temp.: 89.29 F
17:13:11.307 -> Average Cal sensor Temp.: 105.20 Ambient Temp.: 89.33 F
17:13:12.859 -> Average Cal sensor Temp.: 105.14 Ambient Temp.: 89.37 F
17:13:14.407 -> Average Cal sensor Temp.: 105.16 Ambient Temp.: 89.37 F
17:13:15.993 -> Average Cal sensor Temp.: 105.29 Ambient Temp.: 89.37 F
17:13:17.547 -> Average Cal sensor Temp.: 104.97 Ambient Temp.: 89.47 F
17:13:19.101 -> Average Cal sensor Temp.: 104.86 Ambient Temp.: 89.47 F
17:13:20.686 -> Average Cal sensor Temp.: 105.05 Ambient Temp.: 89.44 F
17:13:22.238 -> Average Cal sensor Temp.: 105.04 Ambient Temp.: 89.47 F
17:13:23.795 -> Average Cal sensor Temp.: 105.04 Ambient Temp.: 89.47 F
17:13:25.380 -> Average Cal sensor Temp.: 105.22 Ambient Temp.: 89.58 F
17:13:26.935 -> Average Cal sensor Temp.: 104.74 Ambient Temp.: 89.55 F
17:13:28.517 -> Average Cal sensor Temp.: 104.67 Ambient Temp.: 89.58 F
17:13:30.063 -> Average Cal sensor Temp.: 104.98 Ambient Temp.: 89.55 F
17:13:31.622 -> Average Cal sensor Temp.: 104.64 Ambient Temp.: 89.65 F
17:13:33.205 -> Average Cal sensor Temp.: 104.80 Ambient Temp.: 89.58 F
17:13:34.764 -> Average Cal sensor Temp.: 104.39 Ambient Temp.: 89.58 F
17:13:36.321 -> Average Cal sensor Temp.: 104.73 Ambient Temp.: 89.65 F
17:13:37.912 -> Average Cal sensor Temp.: 104.71 Ambient Temp.: 89.65 F
17:13:39.466 -> Average Cal sensor Temp.: 104.53 Ambient Temp.: 89.69 F
17:13:41.023 -> Average Cal sensor Temp.: 104.63 Ambient Temp.: 89.65 F
17:13:42.610 -> Average Cal sensor Temp.: 105.49 Ambient Temp.: 89.73 F
17:13:44.167 -> Average Cal sensor Temp.: 104.86 Ambient Temp.: 89.58 F
17:13:45.721 -> Average Cal sensor Temp.: 104.87 Ambient Temp.: 89.69 F
17:13:47.279 -> Average Cal sensor Temp.: 105.31 Ambient Temp.: 89.69 F
17:13:48.871 -> Average Cal sensor Temp.: 104.91 Ambient Temp.: 89.73 F
17:13:50.423 -> Average Cal sensor Temp.: 104.87 Ambient Temp.: 89.76 F
17:13:51.982 -> Average Cal sensor Temp.: 104.82 Ambient Temp.: 89.73 F
17:13:53.536 -> Average Cal sensor Temp.: 104.68 Ambient Temp.: 89.73 F
17:13:55.125 -> Average Cal sensor Temp.: 104.18 Ambient Temp.: 89.73 F
17:13:56.683 -> Average Cal sensor Temp.: 104.32 Ambient Temp.: 89.73 F
17:13:58.239 -> Average Cal sensor Temp.: 104.16 Ambient Temp.: 89.69 F
17:13:59.827 -> Average Cal sensor Temp.: 104.72 Ambient Temp.: 89.69 F
17:14:01.383 -> Average Cal sensor Temp.: 104.26 Ambient Temp.: 89.65 F
17:14:02.940 -> Average Cal sensor Temp.: 103.85 Ambient Temp.: 89.55 F
17:14:04.493 -> Average Cal sensor Temp.: 104.36 Ambient Temp.: 89.55 F
17:14:06.078 -> Average Cal sensor Temp.: 104.87 Ambient Temp.: 89.58 F
17:14:07.635 -> Average Cal sensor Temp.: 104.96 Ambient Temp.: 89.69 F
17:14:09.190 -> Average Cal sensor Temp.: 104.63 Ambient Temp.: 89.58 F
17:14:10.780 -> Average Cal sensor Temp.: 104.16 Ambient Temp.: 89.58 F
17:14:12.335 -> Average Cal sensor Temp.: 104.50 Ambient Temp.: 89.55 F
17:14:13.885 -> Average Cal sensor Temp.: 104.39 Ambient Temp.: 89.58 F
17:14:15.471 -> Average Cal sensor Temp.: 104.69 Ambient Temp.: 89.55 F
17:14:17.025 -> Average Cal sensor Temp.: 105.10 Ambient Temp.: 89.65 F
17:14:18.611 -> Average Cal sensor Temp.: 105.57 Ambient Temp.: 89.69 F
17:14:20.169 -> Average Cal sensor Temp.: 105.74 Ambient Temp.: 89.69 F
17:14:21.721 -> Average Cal sensor Temp.: 105.29 Ambient Temp.: 89.65 F
17:14:23.274 -> Average Cal sensor Temp.: 105.31 Ambient Temp.: 89.73 F
17:14:24.869 -> Average Cal sensor Temp.: 104.97 Ambient Temp.: 89.76 F
17:14:26.427 -> Average Cal sensor Temp.: 105.48 Ambient Temp.: 89.76 F
17:14:27.980 -> Average Cal sensor Temp.: 105.06 Ambient Temp.: 89.73 F
17:14:29.531 -> Average Cal sensor Temp.: 105.31 Ambient Temp.: 89.76 F
17:14:31.117 -> Average Cal sensor Temp.: 104.96 Ambient Temp.: 89.80 F
17:14:32.667 -> Average Cal sensor Temp.: 105.62 Ambient Temp.: 89.80 F
17:14:34.247 -> Average Cal sensor Temp.: 105.58 Ambient Temp.: 89.87 F
17:14:35.800 -> Average Cal sensor Temp.: 105.81 Ambient Temp.: 89.91 F
17:14:37.384 -> Average Cal sensor Temp.: 105.84 Ambient Temp.: 89.91 F
17:14:38.943 -> Average Cal sensor Temp.: 106.07 Ambient Temp.: 90.01 F
17:14:40.491 -> Average Cal sensor Temp.: 105.35 Ambient Temp.: 89.98 F
17:14:42.075 -> Average Cal sensor Temp.: 105.17 Ambient Temp.: 89.98 F
17:14:43.631 -> Average Cal sensor Temp.: 106.05 Ambient Temp.: 90.12 F
17:14:45.186 -> Average Cal sensor Temp.: 105.86 Ambient Temp.: 90.09 F
17:14:46.750 -> Average Cal sensor Temp.: 106.55 Ambient Temp.: 90.09 F
17:14:48.344 -> Average Cal sensor Temp.: 105.77 Ambient Temp.: 90.12 F
17:14:49.901 -> Average Cal sensor Temp.: 106.00 Ambient Temp.: 90.16 F
17:14:51.463 -> Average Cal sensor Temp.: 105.87 Ambient Temp.: 90.19 F
17:14:53.016 -> Average Cal sensor Temp.: 105.57 Ambient Temp.: 90.12 F
17:14:54.571 -> Average Cal sensor Temp.: 105.29 Ambient Temp.: 90.12 F
17:14:56.157 -> Average Cal sensor Temp.: 105.16 Ambient Temp.: 90.16 F
17:14:57.709 -> Average Cal sensor Temp.: 105.39 Ambient Temp.: 90.12 F
17:14:59.285 -> Average Cal sensor Temp.: 105.69 Ambient Temp.: 90.16 F
17:15:00.835 -> Average Cal sensor Temp.: 106.23 Ambient Temp.: 90.23 F
17:15:02.421 -> Average Cal sensor Temp.: 105.84 Ambient Temp.: 90.23 F
17:15:03.974 -> Average Cal sensor Temp.: 105.78 Ambient Temp.: 90.27 F
17:15:05.528 -> Average Cal sensor Temp.: 105.73 Ambient Temp.: 90.30 F
17:15:07.083 -> Average Cal sensor Temp.: 105.53 Ambient Temp.: 90.30 F
17:15:08.669 -> Average Cal sensor Temp.: 105.98 Ambient Temp.: 90.30 F
17:15:10.226 -> Average Cal sensor Temp.: 105.56 Ambient Temp.: 90.41 F
17:15:11.778 -> Average Cal sensor Temp.: 105.82 Ambient Temp.: 90.41 F
17:15:13.368 -> Average Cal sensor Temp.: 105.77 Ambient Temp.: 90.48 F
17:15:14.919 -> Average Cal sensor Temp.: 105.56 Ambient Temp.: 90.55 F
17:15:16.476 -> Average Cal sensor Temp.: 105.12 Ambient Temp.: 90.55 F
17:15:18.061 -> Average Cal sensor Temp.: 105.02 Ambient Temp.: 90.59 F
17:15:19.609 -> Average Cal sensor Temp.: 105.47 Ambient Temp.: 90.59 F
17:15:21.188 -> Average Cal sensor Temp.: 105.12 Ambient Temp.: 90.59 F
17:15:22.740 -> Average Cal sensor Temp.: 105.31 Ambient Temp.: 90.63 F
17:15:24.291 -> Average Cal sensor Temp.: 105.31 Ambient Temp.: 90.70 F
17:15:25.879 -> Average Cal sensor Temp.: 105.56 Ambient Temp.: 90.66 F
17:15:27.427 -> Average Cal sensor Temp.: 105.22 Ambient Temp.: 90.73 F
17:15:29.006 -> Average Cal sensor Temp.: 105.42 Ambient Temp.: 90.70 F
17:15:30.555 -> Average Cal sensor Temp.: 105.39 Ambient Temp.: 90.77 F
17:15:32.135 -> Average Cal sensor Temp.: 105.11 Ambient Temp.: 90.88 F
17:15:33.688 -> Average Cal sensor Temp.: 104.88 Ambient Temp.: 90.81 F
17:15:35.241 -> Average Cal sensor Temp.: 105.51 Ambient Temp.: 90.88 F
17:15:36.829 -> Average Cal sensor Temp.: 105.71 Ambient Temp.: 90.95 F
17:15:38.375 -> Average Cal sensor Temp.: 105.27 Ambient Temp.: 90.99 F
17:15:39.928 -> Average Cal sensor Temp.: 105.39 Ambient Temp.: 91.09 F
17:15:41.517 -> Average Cal sensor Temp.: 105.26 Ambient Temp.: 91.06 F
17:15:43.070 -> Average Cal sensor Temp.: 105.44 Ambient Temp.: 91.09 F
17:15:44.620 -> Average Cal sensor Temp.: 105.62 Ambient Temp.: 91.17 F
17:15:46.206 -> Average Cal sensor Temp.: 105.33 Ambient Temp.: 91.20 F
17:15:47.752 -> Average Cal sensor Temp.: 105.29 Ambient Temp.: 91.20 F
17:15:49.339 -> Average Cal sensor Temp.: 105.22 Ambient Temp.: 91.20 F
17:15:50.904 -> Average Cal sensor Temp.: 105.41 Ambient Temp.: 91.17 F
17:15:52.437 -> Average Cal sensor Temp.: 105.26 Ambient Temp.: 91.27 F