## SET-3
Reference Clinical Thermometer Readout | Reference DIGI Thermometer Readout | Reference Water Temperature (Sensor) | Calibrating contactless sensor Readout (2-3 cm) | Calibrating contactless sensor AMBIENT Readout | DIFF bw clinical and sensor | DIFF bw DIGI and sensor | DIFF BETWEEN CLINICAL AND DIGI
Read at (deg F) |  | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F)
94 | 93.9 | 92.41 | 91.35 | 83.46 | 2.65 | 2.55 | 0.1
94.4 | 94.4 | 93.65 | 92.95 |  | 1.45 | 1.45 | 0
94.8 | 94.8 | 93.87 | 93.42 |  | 1.38 | 1.38 | 0
95 | 94.9 | 94.1 | 93.51 |  | 1.49 | 1.39 | 0.1
95.4 | 95.4 | 94.55 | 93.65 |  | 1.75 | 1.75 | 0
95.8 | 95.7 | 94.89 | 93.96 |  | 1.84 | 1.74 | 0.1
96 | 95.9 | 95 | 94.11 |  | 1.89 | 1.79 | 0.1
96.4 | 96.4 | 95.45 | 94.9 |  | 1.5 | 1.5 | 0
96.8 | 96.7 | 95.9 | 95.15 |  | 1.65 | 1.55 | 0.1
97 | 96.9 | 96.01 | 95 |  | 2 | 1.9 | 0.1
97.4 | 97.3 | 96.35 | 95.45 |  | 1.95 | 1.85 | 0.1
97.8 | 97.7 | 96.91 | 95.81 |  | 1.99 | 1.89 | 0.1
98 | 97.9 | 96.91 | 96.15 |  | 1.85 | 1.75 | 0.1
98.2 | 98.1 | 97.25 | 96.3 |  | 1.9 | 1.8 | 0.1
98.4 | 98.3 | 97.47 | 96.47 |  | 1.93 | 1.83 | 0.1
98.6 | 98.5 | 97.7 | 96.6 |  | 2 | 1.9 | 0.1
98.8 | 98.7 | 97.81 | 96.75 | 85.77 | 2.05 | 1.95 | 0.1
99 | 99 | 98.04 | 96.99 |  | 2.01 | 2.01 | 0
99.2 | 99.1 | 98.3 | 97.15 |  | 2.05 | 1.95 | 0.1
99.4 | 99.3 | 98.49 | 97.45 |  | 1.95 | 1.85 | 0.1
99.6 | 99.5 | 98.64 | 97.64 | 84.51 | 1.96 | 1.86 | 0.1
99.8 | 99.7 | 98.81 | 97.78 |  | 2.02 | 1.92 | 0.1
100 | 100 | 99.05 | 97.96 |  | 2.04 | 2.04 | 0
100.2 | 100.2 | 99.16 | 98.1 |  | 2.1 | 2.1 | 0
100.4 | 100.4 | 99.27 | 98.22 |  | 2.18 | 2.18 | 0
100.6 | 100.6 | 99.35 | 98.31 |  | 2.29 | 2.29 | 0
100.8 | 100.8 | 99.5 | 98.35 |  | 2.45 | 2.45 | 0
101 | 100.9 | 99.72 | 98.65 |  | 2.35 | 2.25 | 0.1
101.2 | 101.2 | 99.95 | 98.72 |  | 2.48 | 2.48 | 0
101.4 | 101.4 | 100.17 | 98.85 |  | 2.55 | 2.55 | 0
101.6 | 101.6 | 100.29 | 99.03 |  | 2.57 | 2.57 | 0
101.8 | 101.9 | 100.51 | 99.23 |  | 2.57 | 2.67 | -0.1
102 | 102 | 100.62 | 99.45 |  | 2.55 | 2.55 | 0
102.2 | 102.2 | 100.74 | 99.67 |  | 2.53 | 2.53 | 0
102.4 | 102.4 | 101.41 | 99.9 |  | 2.5 | 2.5 | 0
102.6 | 102.7 | 101.52 | 100.3 |  | 2.3 | 2.4 | -0.1
102.8 | 102.9 | 101.64 | 100.45 |  | 2.35 | 2.45 | -0.1
103 | 103 | 101.75 | 100.53 |  | 2.47 | 2.47 | 0
103.4 | 103.5 | 101.89 | 100.85 |  | 2.55 | 2.65 | -0.1
103.8 | 103.7 | 102.42 | 101.13 |  | 2.67 | 2.57 | 0.1
104 | 104 | 102.6 | 101.49 |  | 2.51 | 2.51 | 0
104.4 | 104.5 | 103.21 | 101.97 |  | 2.43 | 2.53 | -0.1
104.8 | 104.8 | 103.32 | 102.2 |  | 2.6 | 2.6 | 0
105 | 105.2 | 103.44 | 102.45 |  | 2.55 | 2.75 | -0.2
105.4 | 105.4 | 103.89 | 102.75 |  | 2.65 | 2.65 | 0
105.8 | 105.8 | 104.34 | 102.94 |  | 2.86 | 2.86 | 0
106 | 106.2 | 104.56 | 103.15 |  | 2.85 | 3.05 | -0.2
106.4 | 106.6 | 105.01 | 103.75 |  | 2.65 | 2.85 | -0.2
106.8 | 107 | 105.35 | 104.15 |  | 2.65 | 2.85 | -0.2
107 | 107.1 | 105.69 | 104.6 |  | 2.4 | 2.5 | -0.1
107.4 | 107.5 | 106.14 | 105.09 |  | 2.31 | 2.41 | -0.1
107.8 | 107.9 | 106.25 | 105.66 | 88.25 | 2.14 | 2.24 | -0.1
108 | 108.1 | 106.29 | 105.69 |  | 2.31 | 2.41 | -0.1

## RAW_HEATER-3
22:40:00.630 -> Water_Temp.:83.07, PWM:0
22:40:01.425 -> Water_Temp.:83.07, PWM:0
22:40:02.189 -> Water_Temp.:83.07, PWM:0
22:40:02.986 -> Water_Temp.:83.07, PWM:0
22:40:03.744 -> Water_Temp.:83.07, PWM:0
22:40:04.539 -> Water_Temp.:83.07, PWM:0
22:40:05.301 -> Water_Temp.:83.07, PWM:0
22:40:06.089 -> Water_Temp.:83.07, PWM:0
22:40:06.885 -> Water_Temp.:83.07, PWM:0
22:40:07.643 -> Water_Temp.:83.07, PWM:0
22:40:08.437 -> Water_Temp.:83.07, PWM:0
22:40:09.200 -> Water_Temp.:83.07, PWM:0
22:40:09.994 -> Water_Temp.:83.07, PWM:0
22:40:10.757 -> Water_Temp.:83.07, PWM:0
22:40:11.550 -> Water_Temp.:83.07, PWM:0
22:40:12.345 -> Water_Temp.:83.07, PWM:0
22:40:13.106 -> Water_Temp.:83.07, PWM:0
22:40:13.900 -> Water_Temp.:83.07, PWM:0
22:40:14.659 -> Water_Temp.:82.96, PWM:0
22:40:15.453 -> Water_Temp.:82.96, PWM:0
22:40:16.214 -> Water_Temp.:82.96, PWM:0
22:40:17.007 -> Water_Temp.:82.96, PWM:0
22:40:17.804 -> Water_Temp.:82.96, PWM:0
22:40:18.568 -> Water_Temp.:82.96, PWM:0
22:40:19.332 -> Water_Temp.:82.96, PWM:0
22:40:20.132 -> Water_Temp.:82.96, PWM:0
22:40:20.894 -> Water_Temp.:82.96, PWM:0
22:40:21.692 -> Water_Temp.:82.96, PWM:0
22:40:22.484 -> Water_Temp.:82.96, PWM:0
22:40:23.247 -> Water_Temp.:82.96, PWM:0
22:40:24.045 -> Water_Temp.:82.96, PWM:0
22:40:24.804 -> Water_Temp.:82.96, PWM:0
22:40:25.598 -> Water_Temp.:82.96, PWM:0
22:40:26.358 -> Water_Temp.:82.96, PWM:0
22:40:27.146 -> Water_Temp.:82.96, PWM:0
22:40:27.943 -> Water_Temp.:82.96, PWM:0
22:40:28.706 -> Water_Temp.:82.96, PWM:0
22:40:29.500 -> Water_Temp.:82.96, PWM:0
22:40:30.259 -> Water_Temp.:82.96, PWM:0
22:40:31.059 -> Water_Temp.:82.85, PWM:0
22:40:31.816 -> Water_Temp.:82.96, PWM:0
22:40:32.610 -> Water_Temp.:82.85, PWM:0
22:40:33.371 -> Water_Temp.:82.85, PWM:0
22:40:34.170 -> Water_Temp.:82.85, PWM:0
22:40:34.930 -> Water_Temp.:82.85, PWM:0
22:40:35.721 -> Water_Temp.:82.85, PWM:0
22:40:36.518 -> Water_Temp.:82.85, PWM:0
22:40:37.282 -> Water_Temp.:82.85, PWM:0
22:40:38.078 -> Water_Temp.:82.85, PWM:0
22:40:38.843 -> Water_Temp.:82.85, PWM:0
22:40:39.639 -> Water_Temp.:82.85, PWM:0
22:40:40.403 -> Water_Temp.:82.85, PWM:0
22:40:41.196 -> Water_Temp.:82.85, PWM:0
22:40:41.957 -> Water_Temp.:82.85, PWM:0
22:40:42.757 -> Water_Temp.:82.85, PWM:0
22:40:43.521 -> Water_Temp.:82.85, PWM:0
22:40:44.286 -> Water_Temp.:82.85, PWM:0
22:40:45.084 -> Water_Temp.:82.85, PWM:0
22:40:45.846 -> Water_Temp.:82.85, PWM:0
22:40:46.643 -> Water_Temp.:82.85, PWM:0
22:40:47.434 -> Water_Temp.:82.85, PWM:0
22:40:48.196 -> Water_Temp.:82.85, PWM:0
22:40:48.961 -> Water_Temp.:82.85, PWM:0
22:40:49.758 -> Water_Temp.:82.85, PWM:0
22:40:50.552 -> Water_Temp.:82.85, PWM:0
22:40:51.312 -> Water_Temp.:82.85, PWM:0
22:40:52.110 -> Water_Temp.:82.85, PWM:0
22:40:52.870 -> Water_Temp.:82.85, PWM:0
22:40:53.660 -> Water_Temp.:82.74, PWM:0
22:40:54.423 -> Water_Temp.:82.74, PWM:0
22:40:55.220 -> Water_Temp.:82.74, PWM:0
22:40:56.015 -> Water_Temp.:82.74, PWM:0
22:40:56.773 -> Water_Temp.:82.74, PWM:0
22:40:57.570 -> Water_Temp.:82.74, PWM:0
22:40:58.333 -> Water_Temp.:82.74, PWM:0
22:40:59.130 -> Water_Temp.:82.74, PWM:0
22:40:59.889 -> Water_Temp.:82.74, PWM:0
22:41:00.685 -> Water_Temp.:82.74, PWM:0
22:41:01.447 -> Water_Temp.:82.74, PWM:0
22:41:02.243 -> Water_Temp.:82.74, PWM:0
22:41:03.003 -> Water_Temp.:82.74, PWM:0
22:41:03.800 -> Water_Temp.:82.74, PWM:0
22:41:04.564 -> Water_Temp.:82.74, PWM:0
22:41:05.364 -> Water_Temp.:82.74, PWM:0
22:41:06.123 -> Water_Temp.:82.74, PWM:0
22:41:06.915 -> Water_Temp.:82.74, PWM:0
22:41:07.713 -> Water_Temp.:82.74, PWM:0
22:41:08.472 -> Water_Temp.:82.74, PWM:0
22:41:09.270 -> Water_Temp.:82.74, PWM:4
22:41:10.035 -> Water_Temp.:82.74, PWM:49
22:41:10.830 -> Water_Temp.:82.74, PWM:49
22:41:12.356 -> Water_Temp.:82.74, PWM:189
22:41:13.877 -> Water_Temp.:82.62, PWM:254
22:41:14.676 -> Water_Temp.:82.62, PWM:255
22:41:15.435 -> Water_Temp.:82.62, PWM:255
22:41:16.237 -> Water_Temp.:82.62, PWM:255
22:41:16.997 -> Water_Temp.:82.62, PWM:255
22:41:17.793 -> Water_Temp.:82.62, PWM:255
22:41:18.557 -> Water_Temp.:82.62, PWM:255
22:41:19.356 -> Water_Temp.:82.62, PWM:255
22:41:20.116 -> Water_Temp.:82.62, PWM:255
22:41:20.909 -> Water_Temp.:82.62, PWM:254
22:41:21.671 -> Water_Temp.:82.62, PWM:254
22:41:22.471 -> Water_Temp.:82.62, PWM:255
22:41:23.237 -> Water_Temp.:82.62, PWM:254
22:41:24.035 -> Water_Temp.:82.62, PWM:255
22:41:24.800 -> Water_Temp.:82.62, PWM:255
22:41:25.597 -> Water_Temp.:82.62, PWM:254
22:41:26.359 -> Water_Temp.:82.62, PWM:254
22:41:27.919 -> Water_Temp.:82.62, PWM:254
22:41:28.719 -> Water_Temp.:82.51, PWM:255
22:41:29.483 -> Water_Temp.:82.51, PWM:254
22:41:30.249 -> Water_Temp.:82.51, PWM:254
22:41:31.039 -> Water_Temp.:82.51, PWM:255
22:41:31.837 -> Water_Temp.:82.51, PWM:254
22:41:32.595 -> Water_Temp.:82.51, PWM:255
22:41:33.392 -> Water_Temp.:82.51, PWM:254
22:41:34.156 -> Water_Temp.:82.51, PWM:255
22:41:35.713 -> Water_Temp.:82.51, PWM:254
22:41:38.031 -> Water_Temp.:82.51, PWM:254
22:41:38.826 -> Water_Temp.:82.51, PWM:254
22:41:40.387 -> Water_Temp.:82.51, PWM:254
22:41:41.151 -> Water_Temp.:82.51, PWM:255
22:41:41.943 -> Water_Temp.:82.51, PWM:255
22:41:42.708 -> Water_Temp.:82.62, PWM:255
22:41:43.506 -> Water_Temp.:82.51, PWM:255
22:41:44.273 -> Water_Temp.:82.62, PWM:254
22:41:45.066 -> Water_Temp.:82.62, PWM:254
22:41:45.827 -> Water_Temp.:82.62, PWM:255
22:41:46.622 -> Water_Temp.:82.62, PWM:255
22:41:47.380 -> Water_Temp.:82.62, PWM:255
22:41:48.933 -> Water_Temp.:82.74, PWM:254
22:41:49.729 -> Water_Temp.:82.74, PWM:255
22:41:52.011 -> Water_Temp.:82.85, PWM:254
22:41:52.810 -> Water_Temp.:82.85, PWM:254
22:41:53.573 -> Water_Temp.:82.85, PWM:255
22:41:54.371 -> Water_Temp.:82.96, PWM:255
22:41:55.927 -> Water_Temp.:82.96, PWM:254
22:41:56.684 -> Water_Temp.:83.07, PWM:255
22:41:57.480 -> Water_Temp.:83.07, PWM:255
22:41:58.244 -> Water_Temp.:83.19, PWM:255
22:41:59.042 -> Water_Temp.:83.30, PWM:254
22:41:59.803 -> Water_Temp.:83.30, PWM:254
22:42:00.599 -> Water_Temp.:83.41, PWM:255
22:42:01.361 -> Water_Temp.:83.52, PWM:255
22:42:03.679 -> Water_Temp.:83.64, PWM:254
22:42:04.477 -> Water_Temp.:83.75, PWM:255
22:42:05.271 -> Water_Temp.:83.75, PWM:255
22:42:06.037 -> Water_Temp.:83.75, PWM:255
22:42:07.557 -> Water_Temp.:83.86, PWM:254
22:42:08.353 -> Water_Temp.:83.97, PWM:255
22:42:09.115 -> Water_Temp.:83.97, PWM:255
22:42:09.914 -> Water_Temp.:84.09, PWM:254
22:42:10.678 -> Water_Temp.:84.09, PWM:255
22:42:11.475 -> Water_Temp.:84.09, PWM:255
22:42:12.239 -> Water_Temp.:84.09, PWM:254
22:42:13.032 -> Water_Temp.:84.20, PWM:255
22:42:13.795 -> Water_Temp.:84.20, PWM:254
22:42:15.352 -> Water_Temp.:84.20, PWM:254
22:42:16.142 -> Water_Temp.:84.31, PWM:255
22:42:16.937 -> Water_Temp.:84.43, PWM:255
22:42:17.699 -> Water_Temp.:84.43, PWM:254
22:42:19.255 -> Water_Temp.:84.65, PWM:254
22:42:20.014 -> Water_Temp.:84.65, PWM:255
22:42:20.814 -> Water_Temp.:84.76, PWM:255
22:42:21.579 -> Water_Temp.:84.76, PWM:254
22:42:22.375 -> Water_Temp.:84.87, PWM:255
22:42:23.137 -> Water_Temp.:84.87, PWM:255
22:42:23.933 -> Water_Temp.:84.87, PWM:255
22:42:24.695 -> Water_Temp.:84.87, PWM:255
22:42:25.490 -> Water_Temp.:84.99, PWM:255
22:42:26.283 -> Water_Temp.:84.99, PWM:255
22:42:27.047 -> Water_Temp.:84.99, PWM:254
22:42:27.810 -> Water_Temp.:85.10, PWM:255
22:42:28.607 -> Water_Temp.:85.10, PWM:255
22:42:29.403 -> Water_Temp.:85.21, PWM:255
22:42:30.162 -> Water_Temp.:85.21, PWM:255
22:42:31.717 -> Water_Temp.:85.32, PWM:254
22:42:32.479 -> Water_Temp.:85.44, PWM:255
22:42:34.037 -> Water_Temp.:85.44, PWM:254
22:42:34.832 -> Water_Temp.:85.55, PWM:255
22:42:35.592 -> Water_Temp.:85.55, PWM:254
22:42:36.389 -> Water_Temp.:85.66, PWM:255
22:42:37.148 -> Water_Temp.:85.66, PWM:255
22:42:37.949 -> Water_Temp.:85.77, PWM:255
22:42:38.716 -> Water_Temp.:85.77, PWM:255
22:42:39.513 -> Water_Temp.:85.89, PWM:255
22:42:40.277 -> Water_Temp.:85.89, PWM:255
22:42:41.076 -> Water_Temp.:85.89, PWM:255
22:42:41.836 -> Water_Temp.:86.00, PWM:255
22:42:42.634 -> Water_Temp.:86.00, PWM:255
22:42:43.393 -> Water_Temp.:86.00, PWM:255
22:42:44.194 -> Water_Temp.:86.11, PWM:255
22:42:44.957 -> Water_Temp.:86.11, PWM:254
22:42:45.757 -> Water_Temp.:86.11, PWM:255
22:42:46.518 -> Water_Temp.:86.11, PWM:255
22:42:47.313 -> Water_Temp.:86.22, PWM:255
22:42:48.078 -> Water_Temp.:86.22, PWM:254
22:42:48.878 -> Water_Temp.:86.22, PWM:255
22:42:49.641 -> Water_Temp.:86.34, PWM:255
22:42:50.403 -> Water_Temp.:86.34, PWM:254
22:42:51.195 -> Water_Temp.:86.45, PWM:254
22:42:51.987 -> Water_Temp.:86.56, PWM:255
22:42:52.748 -> Water_Temp.:86.56, PWM:254
22:42:53.544 -> Water_Temp.:86.68, PWM:254
22:42:55.103 -> Water_Temp.:86.79, PWM:254
22:42:55.861 -> Water_Temp.:86.90, PWM:254
22:42:56.659 -> Water_Temp.:86.90, PWM:255
22:42:57.425 -> Water_Temp.:87.01, PWM:255
22:42:58.225 -> Water_Temp.:87.12, PWM:255
22:42:58.990 -> Water_Temp.:87.24, PWM:255
22:43:00.547 -> Water_Temp.:87.35, PWM:254
22:43:01.313 -> Water_Temp.:87.46, PWM:197
22:43:02.111 -> Water_Temp.:87.46, PWM:180
22:43:02.872 -> Water_Temp.:87.57, PWM:180
22:43:04.435 -> Water_Temp.:87.69, PWM:180
22:43:05.232 -> Water_Temp.:87.80, PWM:175
22:43:05.995 -> Water_Temp.:87.80, PWM:152
22:43:06.789 -> Water_Temp.:87.80, PWM:130
22:43:07.552 -> Water_Temp.:87.91, PWM:105
22:43:09.120 -> Water_Temp.:87.91, PWM:98
22:43:11.447 -> Water_Temp.:87.91, PWM:89
22:43:12.207 -> Water_Temp.:87.91, PWM:89
22:43:13.005 -> Water_Temp.:88.02, PWM:81
22:43:14.531 -> Water_Temp.:88.02, PWM:74
22:43:15.330 -> Water_Temp.:88.02, PWM:59
22:43:16.092 -> Water_Temp.:88.02, PWM:59
22:43:17.618 -> Water_Temp.:88.14, PWM:23
22:43:18.416 -> Water_Temp.:88.14, PWM:0
22:43:19.179 -> Water_Temp.:88.25, PWM:0
22:43:19.976 -> Water_Temp.:88.25, PWM:0
22:43:20.736 -> Water_Temp.:88.25, PWM:0
22:43:21.537 -> Water_Temp.:88.25, PWM:0
22:43:23.061 -> Water_Temp.:88.36, PWM:4
22:43:24.624 -> Water_Temp.:88.36, PWM:4
22:43:25.386 -> Water_Temp.:88.47, PWM:5
22:43:26.182 -> Water_Temp.:88.47, PWM:5
22:43:26.944 -> Water_Temp.:88.59, PWM:6
22:43:29.264 -> Water_Temp.:88.59, PWM:8
22:43:30.059 -> Water_Temp.:88.59, PWM:15
22:43:31.622 -> Water_Temp.:88.70, PWM:15
22:43:32.386 -> Water_Temp.:88.70, PWM:18
22:43:33.177 -> Water_Temp.:88.70, PWM:18
22:43:33.940 -> Water_Temp.:88.70, PWM:18
22:43:34.736 -> Water_Temp.:88.81, PWM:18
22:43:36.298 -> Water_Temp.:88.81, PWM:19
22:43:38.620 -> Water_Temp.:88.81, PWM:21
22:43:40.179 -> Water_Temp.:88.93, PWM:28
22:43:41.695 -> Water_Temp.:88.93, PWM:32
22:43:42.491 -> Water_Temp.:88.93, PWM:33
22:43:43.254 -> Water_Temp.:88.93, PWM:33
22:43:44.056 -> Water_Temp.:88.93, PWM:33
22:43:44.818 -> Water_Temp.:89.04, PWM:33
22:43:46.380 -> Water_Temp.:89.04, PWM:33
22:43:47.179 -> Water_Temp.:89.04, PWM:33
22:43:48.704 -> Water_Temp.:89.04, PWM:33
22:43:49.505 -> Water_Temp.:89.04, PWM:33
22:43:50.265 -> Water_Temp.:89.04, PWM:33
22:43:51.057 -> Water_Temp.:89.04, PWM:33
22:43:52.582 -> Water_Temp.:89.04, PWM:33
22:43:53.346 -> Water_Temp.:89.04, PWM:33
22:43:54.143 -> Water_Temp.:89.04, PWM:33
22:43:55.700 -> Water_Temp.:89.04, PWM:33
22:43:56.494 -> Water_Temp.:89.15, PWM:33
22:43:57.256 -> Water_Temp.:89.15, PWM:33
22:43:58.053 -> Water_Temp.:89.15, PWM:33
22:43:58.817 -> Water_Temp.:89.15, PWM:33
22:44:00.342 -> Water_Temp.:89.15, PWM:33
22:44:01.140 -> Water_Temp.:89.15, PWM:33
22:44:02.668 -> Water_Temp.:89.15, PWM:33
22:44:03.467 -> Water_Temp.:89.26, PWM:33
22:44:04.231 -> Water_Temp.:89.26, PWM:33
22:44:05.025 -> Water_Temp.:89.26, PWM:33
22:44:08.832 -> Water_Temp.:89.37, PWM:33
22:44:09.623 -> Water_Temp.:89.37, PWM:33
22:44:10.387 -> Water_Temp.:89.37, PWM:33
22:44:11.185 -> Water_Temp.:89.49, PWM:33
22:44:11.944 -> Water_Temp.:89.49, PWM:33
22:44:12.742 -> Water_Temp.:89.49, PWM:33
22:44:13.506 -> Water_Temp.:89.49, PWM:33
22:44:15.067 -> Water_Temp.:89.60, PWM:33
22:44:15.865 -> Water_Temp.:89.60, PWM:33
22:44:16.624 -> Water_Temp.:89.60, PWM:33
22:44:18.185 -> Water_Temp.:89.60, PWM:33
22:44:18.984 -> Water_Temp.:89.71, PWM:33
22:44:19.743 -> Water_Temp.:89.71, PWM:33
22:44:22.063 -> Water_Temp.:89.71, PWM:33
22:44:22.859 -> Water_Temp.:89.71, PWM:33
22:44:23.624 -> Water_Temp.:89.60, PWM:33
22:44:24.425 -> Water_Temp.:89.60, PWM:33
22:44:25.186 -> Water_Temp.:89.60, PWM:33
22:44:26.743 -> Water_Temp.:89.60, PWM:33
22:44:29.032 -> Water_Temp.:89.60, PWM:33
22:44:30.592 -> Water_Temp.:89.60, PWM:33
22:44:31.384 -> Water_Temp.:89.60, PWM:33
22:44:32.907 -> Water_Temp.:89.60, PWM:33
22:44:34.468 -> Water_Temp.:89.71, PWM:33
22:44:35.229 -> Water_Temp.:89.71, PWM:33
22:44:36.026 -> Water_Temp.:89.71, PWM:33
22:44:36.785 -> Water_Temp.:89.71, PWM:33
22:44:37.582 -> Water_Temp.:89.71, PWM:33
22:44:38.345 -> Water_Temp.:89.82, PWM:33
22:44:41.468 -> Water_Temp.:89.94, PWM:33
22:44:42.263 -> Water_Temp.:89.94, PWM:33
22:44:43.029 -> Water_Temp.:89.94, PWM:33
22:44:43.827 -> Water_Temp.:89.94, PWM:33
22:44:44.588 -> Water_Temp.:89.94, PWM:33
22:44:45.378 -> Water_Temp.:89.94, PWM:33
22:44:47.726 -> Water_Temp.:89.94, PWM:33
22:44:48.482 -> Water_Temp.:89.94, PWM:33
22:44:49.276 -> Water_Temp.:89.94, PWM:33
22:44:52.361 -> Water_Temp.:89.94, PWM:33
22:44:53.155 -> Water_Temp.:89.94, PWM:33
22:44:54.709 -> Water_Temp.:89.94, PWM:33
22:44:55.492 -> Water_Temp.:89.94, PWM:33
22:44:56.253 -> Water_Temp.:89.94, PWM:33
22:44:57.048 -> Water_Temp.:89.94, PWM:33
22:45:00.099 -> Water_Temp.:90.05, PWM:33
22:45:00.894 -> Water_Temp.:90.05, PWM:33
22:45:02.453 -> Water_Temp.:90.05, PWM:33
22:45:03.248 -> Water_Temp.:90.05, PWM:33
22:45:04.015 -> Water_Temp.:90.05, PWM:33
22:45:04.809 -> Water_Temp.:90.05, PWM:33
22:45:05.571 -> Water_Temp.:90.05, PWM:33
22:45:06.369 -> Water_Temp.:90.05, PWM:33
22:45:07.130 -> Water_Temp.:90.05, PWM:33
22:45:07.925 -> Water_Temp.:90.05, PWM:33
22:45:08.687 -> Water_Temp.:90.05, PWM:33
22:45:09.487 -> Water_Temp.:90.05, PWM:33
22:45:10.251 -> Water_Temp.:90.05, PWM:33
22:45:11.015 -> Water_Temp.:90.05, PWM:33
22:45:11.813 -> Water_Temp.:90.05, PWM:33
22:45:13.342 -> Water_Temp.:90.05, PWM:33
22:45:14.140 -> Water_Temp.:90.05, PWM:33
22:45:16.424 -> Water_Temp.:90.05, PWM:33
22:45:17.222 -> Water_Temp.:90.16, PWM:33
22:45:18.018 -> Water_Temp.:90.16, PWM:33
22:45:19.545 -> Water_Temp.:90.16, PWM:33
22:45:21.069 -> Water_Temp.:90.16, PWM:33
22:45:23.393 -> Water_Temp.:90.27, PWM:33
22:45:24.923 -> Water_Temp.:90.27, PWM:33
22:45:28.773 -> Water_Temp.:90.27, PWM:33
22:45:30.328 -> Water_Temp.:90.27, PWM:33
22:45:31.125 -> Water_Temp.:90.27, PWM:33
22:45:31.882 -> Water_Temp.:90.27, PWM:33
22:45:33.443 -> Water_Temp.:90.27, PWM:33
22:45:34.206 -> Water_Temp.:90.27, PWM:33
22:45:35.002 -> Water_Temp.:90.27, PWM:33
22:45:35.766 -> Water_Temp.:90.27, PWM:33
22:45:37.290 -> Water_Temp.:90.39, PWM:33
22:45:38.087 -> Water_Temp.:90.39, PWM:33
22:45:38.849 -> Water_Temp.:90.39, PWM:33
22:45:39.646 -> Water_Temp.:90.39, PWM:33
22:45:40.406 -> Water_Temp.:90.39, PWM:33
22:45:41.202 -> Water_Temp.:90.39, PWM:33
22:45:41.999 -> Water_Temp.:90.39, PWM:33
22:45:44.324 -> Water_Temp.:90.27, PWM:33
22:45:45.090 -> Water_Temp.:90.27, PWM:33
22:45:45.886 -> Water_Temp.:90.39, PWM:33
22:45:46.647 -> Water_Temp.:90.39, PWM:33
22:45:47.410 -> Water_Temp.:90.39, PWM:33
22:45:48.209 -> Water_Temp.:90.39, PWM:33
22:45:48.973 -> Water_Temp.:90.39, PWM:33
22:45:49.768 -> Water_Temp.:90.39, PWM:33
22:45:50.562 -> Water_Temp.:90.39, PWM:33
22:45:51.322 -> Water_Temp.:90.27, PWM:33
22:45:52.118 -> Water_Temp.:90.39, PWM:33
22:45:52.878 -> Water_Temp.:90.39, PWM:33
22:45:53.674 -> Water_Temp.:90.39, PWM:33
22:45:54.434 -> Water_Temp.:90.39, PWM:33
22:45:55.232 -> Water_Temp.:90.39, PWM:33
22:45:55.992 -> Water_Temp.:90.39, PWM:33
22:45:56.788 -> Water_Temp.:90.39, PWM:33
22:45:58.313 -> Water_Temp.:90.39, PWM:33
22:45:59.113 -> Water_Temp.:90.39, PWM:33
22:46:00.641 -> Water_Temp.:90.39, PWM:33
22:46:02.195 -> Water_Temp.:90.39, PWM:33
22:46:02.993 -> Water_Temp.:90.39, PWM:33
22:46:03.758 -> Water_Temp.:90.39, PWM:33
22:46:07.672 -> Water_Temp.:90.50, PWM:33
22:46:10.761 -> Water_Temp.:90.61, PWM:33
22:46:12.316 -> Water_Temp.:90.61, PWM:33
22:46:13.873 -> Water_Temp.:90.61, PWM:33
22:46:17.757 -> Water_Temp.:90.72, PWM:33
22:46:18.522 -> Water_Temp.:90.72, PWM:33
22:46:20.843 -> Water_Temp.:90.72, PWM:33
22:46:21.637 -> Water_Temp.:90.72, PWM:33
22:46:23.960 -> Water_Temp.:90.72, PWM:33
22:46:27.077 -> Water_Temp.:90.72, PWM:33
22:46:28.607 -> Water_Temp.:90.72, PWM:33
22:46:29.404 -> Water_Temp.:90.84, PWM:33
22:46:30.165 -> Water_Temp.:90.84, PWM:33
22:46:30.956 -> Water_Temp.:90.84, PWM:33
22:46:33.269 -> Water_Temp.:90.84, PWM:33
22:46:34.832 -> Water_Temp.:90.84, PWM:33
22:46:35.592 -> Water_Temp.:90.84, PWM:33
22:46:36.384 -> Water_Temp.:90.84, PWM:33
22:46:37.939 -> Water_Temp.:90.84, PWM:33
22:46:38.701 -> Water_Temp.:90.84, PWM:33
22:46:40.262 -> Water_Temp.:90.84, PWM:33
22:46:41.024 -> Water_Temp.:90.84, PWM:33
22:46:41.820 -> Water_Temp.:90.84, PWM:33
22:46:44.112 -> Water_Temp.:90.84, PWM:33
22:46:44.912 -> Water_Temp.:90.84, PWM:33
22:46:46.467 -> Water_Temp.:90.84, PWM:33
22:46:48.756 -> Water_Temp.:90.84, PWM:33
22:46:50.309 -> Water_Temp.:90.84, PWM:33
22:46:51.102 -> Water_Temp.:90.84, PWM:33
22:46:54.151 -> Water_Temp.:90.84, PWM:33
22:46:54.948 -> Water_Temp.:90.84, PWM:33
22:46:56.473 -> Water_Temp.:90.84, PWM:33
22:46:58.797 -> Water_Temp.:90.84, PWM:33
22:47:00.356 -> Water_Temp.:90.84, PWM:33
22:47:02.678 -> Water_Temp.:90.84, PWM:33
22:47:03.437 -> Water_Temp.:90.84, PWM:33
22:47:04.999 -> Water_Temp.:90.95, PWM:33
22:47:05.762 -> Water_Temp.:90.95, PWM:33
22:47:06.557 -> Water_Temp.:90.95, PWM:33
22:47:07.323 -> Water_Temp.:90.95, PWM:33
22:47:08.886 -> Water_Temp.:90.95, PWM:33
22:47:10.444 -> Water_Temp.:90.95, PWM:33
22:47:11.203 -> Water_Temp.:90.95, PWM:33
22:47:13.557 -> Water_Temp.:90.95, PWM:33
22:47:14.322 -> Water_Temp.:90.95, PWM:33
22:47:16.636 -> Water_Temp.:90.95, PWM:33
22:47:17.428 -> Water_Temp.:90.95, PWM:33
22:47:18.189 -> Water_Temp.:90.95, PWM:33
22:47:18.988 -> Water_Temp.:92.64, PWM:33
22:47:19.748 -> Water_Temp.:94.55, PWM:33
22:47:22.861 -> Water_Temp.:91.62, PWM:33
22:47:23.659 -> Water_Temp.:91.51, PWM:33
22:47:25.979 -> Water_Temp.:91.40, PWM:33
22:47:26.741 -> Water_Temp.:91.29, PWM:33
22:47:27.537 -> Water_Temp.:91.29, PWM:33
22:47:29.060 -> Water_Temp.:91.29, PWM:33
22:47:31.376 -> Water_Temp.:91.29, PWM:33
22:47:32.898 -> Water_Temp.:91.29, PWM:33
22:47:34.460 -> Water_Temp.:91.29, PWM:33
22:47:35.224 -> Water_Temp.:91.29, PWM:33
22:47:36.016 -> Water_Temp.:91.29, PWM:33
22:47:36.778 -> Water_Temp.:91.29, PWM:33
22:47:37.574 -> Water_Temp.:91.29, PWM:33
22:47:38.337 -> Water_Temp.:91.29, PWM:33
22:47:40.698 -> Water_Temp.:91.29, PWM:33
22:47:41.459 -> Water_Temp.:91.29, PWM:33
22:47:44.573 -> Water_Temp.:91.29, PWM:33
22:47:47.685 -> Water_Temp.:91.29, PWM:33
22:47:48.449 -> Water_Temp.:91.29, PWM:33
22:47:50.772 -> Water_Temp.:92.86, PWM:33
22:47:53.085 -> Water_Temp.:92.07, PWM:33
22:47:53.882 -> Water_Temp.:91.85, PWM:33
22:47:55.438 -> Water_Temp.:91.62, PWM:33
22:47:56.201 -> Water_Temp.:91.62, PWM:33
22:47:56.998 -> Water_Temp.:91.51, PWM:33
22:47:58.525 -> Water_Temp.:91.40, PWM:33
22:48:00.847 -> Water_Temp.:91.29, PWM:33
22:48:01.611 -> Water_Temp.:91.29, PWM:33
22:48:02.406 -> Water_Temp.:91.29, PWM:33
22:48:03.201 -> Water_Temp.:91.29, PWM:33
22:48:03.965 -> Water_Temp.:91.29, PWM:33
22:48:05.490 -> Water_Temp.:91.18, PWM:33
22:48:06.286 -> Water_Temp.:91.18, PWM:33
22:48:07.079 -> Water_Temp.:91.18, PWM:33
22:48:07.845 -> Water_Temp.:91.18, PWM:33
22:48:09.372 -> Water_Temp.:91.18, PWM:33
22:48:10.166 -> Water_Temp.:91.18, PWM:33
22:48:12.484 -> Water_Temp.:91.18, PWM:33
22:48:14.042 -> Water_Temp.:91.18, PWM:33
22:48:14.801 -> Water_Temp.:91.18, PWM:33
22:48:15.596 -> Water_Temp.:91.18, PWM:33
22:48:17.153 -> Water_Temp.:91.18, PWM:33
22:48:17.915 -> Water_Temp.:91.18, PWM:33
22:48:19.474 -> Water_Temp.:91.18, PWM:33
22:48:20.271 -> Water_Temp.:91.29, PWM:33
22:48:21.031 -> Water_Temp.:91.29, PWM:33
22:48:21.823 -> Water_Temp.:91.29, PWM:33
22:48:22.618 -> Water_Temp.:91.29, PWM:33
22:48:23.380 -> Water_Temp.:91.40, PWM:33
22:48:24.147 -> Water_Temp.:91.40, PWM:33
22:48:24.944 -> Water_Temp.:91.40, PWM:33
22:48:26.470 -> Water_Temp.:91.40, PWM:33
22:48:27.265 -> Water_Temp.:91.40, PWM:33
22:48:28.830 -> Water_Temp.:91.40, PWM:33
22:48:29.592 -> Water_Temp.:91.40, PWM:33
22:48:35.689 -> Water_Temp.:91.51, PWM:33
22:48:37.250 -> Water_Temp.:91.51, PWM:33
22:48:39.546 -> Water_Temp.:91.51, PWM:33
22:48:41.107 -> Water_Temp.:91.51, PWM:33
22:48:41.870 -> Water_Temp.:91.40, PWM:33
22:48:43.433 -> Water_Temp.:91.51, PWM:33
22:48:45.759 -> Water_Temp.:91.51, PWM:33
22:48:50.393 -> Water_Temp.:91.51, PWM:33
22:48:51.157 -> Water_Temp.:91.51, PWM:33
22:48:51.956 -> Water_Temp.:91.51, PWM:33
22:48:52.721 -> Water_Temp.:91.40, PWM:33
22:48:53.517 -> Water_Temp.:91.40, PWM:33
22:48:54.280 -> Water_Temp.:91.40, PWM:33
22:48:55.074 -> Water_Temp.:91.51, PWM:33
22:48:56.604 -> Water_Temp.:91.51, PWM:33
22:48:58.160 -> Water_Temp.:91.51, PWM:33
22:48:58.925 -> Water_Temp.:91.51, PWM:33
22:49:00.485 -> Water_Temp.:95.11, PWM:33
22:49:02.040 -> Water_Temp.:92.41, PWM:33
22:49:02.837 -> Water_Temp.:92.19, PWM:33
22:49:03.603 -> Water_Temp.:92.07, PWM:33
22:49:04.405 -> Water_Temp.:91.96, PWM:33
22:49:05.927 -> Water_Temp.:91.85, PWM:33
22:49:06.728 -> Water_Temp.:91.74, PWM:33
22:49:12.102 -> Water_Temp.:91.62, PWM:33
22:49:12.864 -> Water_Temp.:91.51, PWM:33
22:49:13.660 -> Water_Temp.:91.51, PWM:33
22:49:15.982 -> Water_Temp.:91.51, PWM:33
22:49:16.742 -> Water_Temp.:91.51, PWM:33
22:49:17.534 -> Water_Temp.:91.51, PWM:33
22:49:18.328 -> Water_Temp.:91.51, PWM:33
22:49:20.617 -> Water_Temp.:91.40, PWM:33
22:49:21.416 -> Water_Temp.:91.51, PWM:33
22:49:22.176 -> Water_Temp.:91.40, PWM:33
22:49:22.970 -> Water_Temp.:91.51, PWM:33
22:49:24.529 -> Water_Temp.:91.40, PWM:33
22:49:27.650 -> Water_Temp.:91.40, PWM:33
22:49:28.415 -> Water_Temp.:91.40, PWM:33
22:49:29.214 -> Water_Temp.:91.40, PWM:33
22:49:30.746 -> Water_Temp.:91.40, PWM:33
22:49:32.309 -> Water_Temp.:91.40, PWM:33
22:49:33.075 -> Water_Temp.:91.40, PWM:33
22:49:33.875 -> Water_Temp.:91.40, PWM:33
22:49:35.408 -> Water_Temp.:91.51, PWM:33
22:49:36.208 -> Water_Temp.:91.51, PWM:33
22:49:36.970 -> Water_Temp.:91.51, PWM:33
22:49:37.768 -> Water_Temp.:91.51, PWM:33
22:49:38.534 -> Water_Temp.:91.51, PWM:33
22:49:40.860 -> Water_Temp.:91.51, PWM:33
22:49:41.622 -> Water_Temp.:91.51, PWM:33
22:49:42.421 -> Water_Temp.:91.51, PWM:33
22:49:44.745 -> Water_Temp.:91.51, PWM:33
22:49:45.540 -> Water_Temp.:91.51, PWM:33
22:49:47.094 -> Water_Temp.:91.51, PWM:33
22:49:48.620 -> Water_Temp.:91.51, PWM:33
22:49:49.416 -> Water_Temp.:91.51, PWM:33
22:49:50.178 -> Water_Temp.:91.51, PWM:33
22:49:50.974 -> Water_Temp.:91.51, PWM:33
22:49:51.738 -> Water_Temp.:91.51, PWM:33
22:49:54.063 -> Water_Temp.:91.51, PWM:33
22:49:54.862 -> Water_Temp.:91.51, PWM:33
22:49:55.624 -> Water_Temp.:91.51, PWM:33
22:49:56.421 -> Water_Temp.:91.51, PWM:33
22:49:57.181 -> Water_Temp.:91.51, PWM:33
22:49:58.743 -> Water_Temp.:91.51, PWM:33
22:50:00.266 -> Water_Temp.:91.62, PWM:33
22:50:01.822 -> Water_Temp.:91.62, PWM:33
22:50:03.378 -> Water_Temp.:91.62, PWM:33
22:50:04.143 -> Water_Temp.:91.62, PWM:33
22:50:04.938 -> Water_Temp.:91.62, PWM:33
22:50:05.697 -> Water_Temp.:91.62, PWM:33
22:50:06.497 -> Water_Temp.:91.62, PWM:33
22:50:07.259 -> Water_Temp.:91.62, PWM:33
22:50:08.052 -> Water_Temp.:91.62, PWM:33
22:50:08.810 -> Water_Temp.:91.62, PWM:33
22:50:09.605 -> Water_Temp.:91.62, PWM:33
22:50:13.482 -> Water_Temp.:91.74, PWM:33
22:50:15.007 -> Water_Temp.:91.74, PWM:33
22:50:17.294 -> Water_Temp.:91.85, PWM:33
22:50:18.088 -> Water_Temp.:91.85, PWM:33
22:50:18.855 -> Water_Temp.:91.85, PWM:33
22:50:19.653 -> Water_Temp.:91.85, PWM:33
22:50:21.929 -> Water_Temp.:91.85, PWM:33
22:50:23.485 -> Water_Temp.:91.85, PWM:33
22:50:25.013 -> Water_Temp.:91.85, PWM:33
22:50:25.810 -> Water_Temp.:91.85, PWM:33
22:50:26.575 -> Water_Temp.:91.85, PWM:33
22:50:27.340 -> Water_Temp.:91.74, PWM:33
22:50:28.136 -> Water_Temp.:91.74, PWM:33
22:50:28.900 -> Water_Temp.:91.74, PWM:33
22:50:29.702 -> Water_Temp.:91.74, PWM:33
22:50:30.468 -> Water_Temp.:91.74, PWM:33
22:50:31.264 -> Water_Temp.:91.74, PWM:33
22:50:32.024 -> Water_Temp.:91.74, PWM:33
22:50:32.820 -> Water_Temp.:91.74, PWM:33
22:50:33.578 -> Water_Temp.:91.74, PWM:33
22:50:34.374 -> Water_Temp.:91.74, PWM:33
22:50:35.167 -> Water_Temp.:91.74, PWM:33
22:50:35.927 -> Water_Temp.:91.74, PWM:33
22:50:36.726 -> Water_Temp.:91.74, PWM:33
22:50:37.485 -> Water_Temp.:91.74, PWM:33
22:50:38.283 -> Water_Temp.:91.74, PWM:33
22:50:39.049 -> Water_Temp.:91.74, PWM:33
22:50:41.378 -> Water_Temp.:91.74, PWM:33
22:50:42.932 -> Water_Temp.:91.74, PWM:33
22:50:44.495 -> Water_Temp.:91.74, PWM:33
22:50:46.051 -> Water_Temp.:91.85, PWM:33
22:50:46.849 -> Water_Temp.:91.74, PWM:33
22:50:48.371 -> Water_Temp.:91.74, PWM:33
22:50:49.933 -> Water_Temp.:91.74, PWM:33
22:50:50.694 -> Water_Temp.:91.74, PWM:33
22:50:53.746 -> Water_Temp.:91.74, PWM:33
22:50:55.307 -> Water_Temp.:91.74, PWM:33
22:50:56.070 -> Water_Temp.:91.74, PWM:33
22:50:57.626 -> Water_Temp.:91.85, PWM:33
22:50:58.423 -> Water_Temp.:91.85, PWM:33
22:50:59.186 -> Water_Temp.:91.85, PWM:33
22:50:59.978 -> Water_Temp.:91.85, PWM:33
22:51:00.741 -> Water_Temp.:91.85, PWM:33
22:51:01.537 -> Water_Temp.:91.85, PWM:33
22:51:03.094 -> Water_Temp.:91.85, PWM:33
22:51:04.648 -> Water_Temp.:91.85, PWM:33
22:51:05.409 -> Water_Temp.:91.85, PWM:33
22:51:06.202 -> Water_Temp.:91.85, PWM:33
22:51:06.965 -> Water_Temp.:91.85, PWM:33
22:51:07.760 -> Water_Temp.:91.85, PWM:33
22:51:08.520 -> Water_Temp.:91.85, PWM:33
22:51:10.082 -> Water_Temp.:91.85, PWM:33
22:51:10.880 -> Water_Temp.:91.85, PWM:33
22:51:12.407 -> Water_Temp.:91.85, PWM:33
22:51:18.596 -> Water_Temp.:91.96, PWM:33
22:51:19.394 -> Water_Temp.:91.96, PWM:33
22:51:20.155 -> Water_Temp.:91.96, PWM:33
22:51:22.474 -> Water_Temp.:91.96, PWM:33
22:51:23.237 -> Water_Temp.:91.96, PWM:33
22:51:24.794 -> Water_Temp.:91.96, PWM:33
22:51:25.585 -> Water_Temp.:91.96, PWM:33
22:51:26.378 -> Water_Temp.:91.96, PWM:33
22:51:27.900 -> Water_Temp.:91.96, PWM:33
22:51:28.700 -> Water_Temp.:91.96, PWM:33
22:51:29.465 -> Water_Temp.:91.96, PWM:33
22:51:30.256 -> Water_Temp.:91.96, PWM:33
22:51:31.020 -> Water_Temp.:91.96, PWM:33
22:51:31.811 -> Water_Temp.:91.96, PWM:33
22:51:32.601 -> Water_Temp.:91.96, PWM:33
22:51:33.364 -> Water_Temp.:91.96, PWM:33
22:51:34.130 -> Water_Temp.:91.96, PWM:33
22:51:35.723 -> Water_Temp.:91.96, PWM:33
22:51:36.484 -> Water_Temp.:91.96, PWM:33
22:51:37.279 -> Water_Temp.:91.96, PWM:33
22:51:38.805 -> Water_Temp.:91.96, PWM:33
22:51:39.599 -> Water_Temp.:92.07, PWM:33
22:51:46.582 -> Water_Temp.:91.96, PWM:33
22:51:47.342 -> Water_Temp.:91.96, PWM:33
22:51:48.900 -> Water_Temp.:92.07, PWM:33
22:51:49.699 -> Water_Temp.:91.96, PWM:33
22:51:51.257 -> Water_Temp.:92.07, PWM:33
22:51:52.013 -> Water_Temp.:92.07, PWM:33
22:51:52.810 -> Water_Temp.:92.07, PWM:33
22:51:54.337 -> Water_Temp.:92.07, PWM:33
22:51:55.135 -> Water_Temp.:92.07, PWM:33
22:51:55.897 -> Water_Temp.:92.07, PWM:33
22:51:56.695 -> Water_Temp.:92.07, PWM:33
22:51:58.258 -> Water_Temp.:92.19, PWM:33
22:51:59.814 -> Water_Temp.:92.19, PWM:33
22:52:00.573 -> Water_Temp.:92.19, PWM:33
22:52:02.893 -> Water_Temp.:92.19, PWM:33
22:52:07.500 -> Water_Temp.:92.19, PWM:33
22:52:09.821 -> Water_Temp.:92.19, PWM:33
22:52:10.614 -> Water_Temp.:92.19, PWM:33
22:52:11.379 -> Water_Temp.:92.30, PWM:33
22:52:12.939 -> Water_Temp.:92.30, PWM:33
22:52:14.464 -> Water_Temp.:92.30, PWM:33
22:52:15.263 -> Water_Temp.:92.30, PWM:33
22:52:16.025 -> Water_Temp.:92.30, PWM:33
22:52:17.587 -> Water_Temp.:92.30, PWM:33
22:52:18.350 -> Water_Temp.:92.19, PWM:33
22:52:19.149 -> Water_Temp.:92.19, PWM:33
22:52:21.434 -> Water_Temp.:92.19, PWM:33
22:52:22.227 -> Water_Temp.:92.19, PWM:33
22:52:22.992 -> Water_Temp.:92.19, PWM:33
22:52:23.791 -> Water_Temp.:92.19, PWM:33
22:52:24.556 -> Water_Temp.:92.19, PWM:33
22:52:25.356 -> Water_Temp.:92.19, PWM:33
22:52:26.120 -> Water_Temp.:92.19, PWM:33
22:52:27.683 -> Water_Temp.:92.19, PWM:33
22:52:28.446 -> Water_Temp.:92.19, PWM:33
22:52:30.002 -> Water_Temp.:92.19, PWM:33
22:52:30.798 -> Water_Temp.:92.19, PWM:33
22:52:31.592 -> Water_Temp.:92.19, PWM:33
22:52:33.111 -> Water_Temp.:92.19, PWM:33
22:52:33.911 -> Water_Temp.:92.19, PWM:33
22:52:34.673 -> Water_Temp.:92.19, PWM:33
22:52:35.474 -> Water_Temp.:92.19, PWM:33
22:52:36.997 -> Water_Temp.:92.19, PWM:33
22:52:37.758 -> Water_Temp.:92.19, PWM:33
22:52:39.314 -> Water_Temp.:92.19, PWM:33
22:52:41.633 -> Water_Temp.:92.19, PWM:33
22:52:43.186 -> Water_Temp.:92.19, PWM:33
22:52:43.947 -> Water_Temp.:92.19, PWM:33
22:52:45.505 -> Water_Temp.:92.19, PWM:33
22:52:46.300 -> Water_Temp.:92.19, PWM:33
22:52:49.383 -> Water_Temp.:92.30, PWM:33
22:52:50.939 -> Water_Temp.:92.30, PWM:33
22:52:51.697 -> Water_Temp.:92.30, PWM:33
22:52:52.494 -> Water_Temp.:92.30, PWM:33
22:52:53.260 -> Water_Temp.:92.30, PWM:33
22:52:54.053 -> Water_Temp.:92.30, PWM:33
22:52:54.819 -> Water_Temp.:92.19, PWM:33
22:52:55.617 -> Water_Temp.:92.19, PWM:33
22:52:56.382 -> Water_Temp.:92.19, PWM:33
22:52:57.178 -> Water_Temp.:92.19, PWM:33
22:52:58.735 -> Water_Temp.:92.19, PWM:33
22:52:59.498 -> Water_Temp.:92.19, PWM:33
22:53:01.816 -> Water_Temp.:92.19, PWM:33
22:53:02.612 -> Water_Temp.:92.19, PWM:33
22:53:03.374 -> Water_Temp.:92.19, PWM:33
22:53:04.170 -> Water_Temp.:92.19, PWM:33
22:53:04.930 -> Water_Temp.:92.19, PWM:33
22:53:08.804 -> Water_Temp.:92.19, PWM:33
22:53:09.601 -> Water_Temp.:92.19, PWM:33
22:53:11.126 -> Water_Temp.:92.19, PWM:33
22:53:13.448 -> Water_Temp.:92.19, PWM:33
22:53:14.250 -> Water_Temp.:92.19, PWM:33
22:53:15.008 -> Water_Temp.:92.19, PWM:33
22:53:15.799 -> Water_Temp.:92.19, PWM:33
22:53:16.594 -> Water_Temp.:92.19, PWM:33
22:53:17.353 -> Water_Temp.:92.19, PWM:33
22:53:18.151 -> Water_Temp.:92.19, PWM:33
22:53:18.917 -> Water_Temp.:92.19, PWM:33
22:53:20.476 -> Water_Temp.:92.19, PWM:33
22:53:21.234 -> Water_Temp.:92.19, PWM:33
22:53:22.032 -> Water_Temp.:92.19, PWM:33
22:53:22.793 -> Water_Temp.:93.87, PWM:33
22:53:23.589 -> Water_Temp.:95.79, PWM:33
22:53:24.351 -> Water_Temp.:93.54, PWM:33
22:53:25.144 -> Water_Temp.:93.09, PWM:33
22:53:26.694 -> Water_Temp.:92.75, PWM:33
22:53:29.015 -> Water_Temp.:92.64, PWM:33
22:53:29.776 -> Water_Temp.:92.52, PWM:33
22:53:33.688 -> Water_Temp.:92.41, PWM:33
22:53:34.447 -> Water_Temp.:92.41, PWM:33
22:53:35.248 -> Water_Temp.:92.41, PWM:33
22:53:36.011 -> Water_Temp.:92.41, PWM:33
22:53:36.810 -> Water_Temp.:92.41, PWM:33
22:53:37.573 -> Water_Temp.:92.41, PWM:33
22:53:39.101 -> Water_Temp.:92.41, PWM:33
22:53:39.902 -> Water_Temp.:92.41, PWM:33
22:53:41.460 -> Water_Temp.:92.41, PWM:33
22:53:43.753 -> Water_Temp.:92.41, PWM:33
22:53:46.079 -> Water_Temp.:92.41, PWM:33
22:53:46.840 -> Water_Temp.:92.41, PWM:33
22:53:47.636 -> Water_Temp.:92.41, PWM:33
22:53:48.401 -> Water_Temp.:92.41, PWM:33
22:53:49.202 -> Water_Temp.:92.41, PWM:33
22:53:49.962 -> Water_Temp.:92.41, PWM:33
22:53:50.756 -> Water_Temp.:92.41, PWM:33
22:53:51.517 -> Water_Temp.:92.41, PWM:33
22:53:54.636 -> Water_Temp.:92.41, PWM:33
22:53:55.434 -> Water_Temp.:92.41, PWM:33
22:53:56.193 -> Water_Temp.:92.41, PWM:33
22:53:57.753 -> Water_Temp.:92.41, PWM:33
22:53:58.547 -> Water_Temp.:92.41, PWM:33
22:54:00.866 -> Water_Temp.:92.41, PWM:33
22:54:03.182 -> Water_Temp.:92.52, PWM:33
22:54:04.742 -> Water_Temp.:92.52, PWM:33
22:54:05.505 -> Water_Temp.:92.52, PWM:33
22:54:06.303 -> Water_Temp.:92.52, PWM:33
22:54:07.065 -> Water_Temp.:92.52, PWM:33
22:54:08.623 -> Water_Temp.:92.52, PWM:33
22:54:10.182 -> Water_Temp.:92.52, PWM:33
22:54:12.505 -> Water_Temp.:92.52, PWM:33
22:54:13.270 -> Water_Temp.:92.52, PWM:33
22:54:14.831 -> Water_Temp.:92.52, PWM:33
22:54:15.592 -> Water_Temp.:92.52, PWM:33
22:54:16.389 -> Water_Temp.:92.52, PWM:33
22:54:17.184 -> Water_Temp.:92.52, PWM:33
22:54:17.944 -> Water_Temp.:92.52, PWM:33
22:54:19.499 -> Water_Temp.:92.41, PWM:33
22:54:20.255 -> Water_Temp.:92.41, PWM:33
22:54:21.052 -> Water_Temp.:92.41, PWM:33
22:54:21.810 -> Water_Temp.:92.52, PWM:33
22:54:22.604 -> Water_Temp.:92.52, PWM:33
22:54:24.923 -> Water_Temp.:92.52, PWM:33
22:54:25.717 -> Water_Temp.:92.52, PWM:33
22:54:27.273 -> Water_Temp.:92.52, PWM:33
22:54:28.035 -> Water_Temp.:92.52, PWM:33
22:54:28.832 -> Water_Temp.:92.52, PWM:33
22:54:30.383 -> Water_Temp.:92.52, PWM:33
22:54:31.144 -> Water_Temp.:92.52, PWM:33
22:54:31.939 -> Water_Temp.:92.52, PWM:33
22:54:32.738 -> Water_Temp.:92.52, PWM:33
22:54:33.502 -> Water_Temp.:92.52, PWM:33
22:54:34.266 -> Water_Temp.:92.52, PWM:33
22:54:35.826 -> Water_Temp.:92.52, PWM:33
22:54:36.621 -> Water_Temp.:92.64, PWM:33
22:54:37.387 -> Water_Temp.:92.64, PWM:33
22:54:38.186 -> Water_Temp.:92.64, PWM:33
22:54:38.950 -> Water_Temp.:92.52, PWM:33
22:54:39.748 -> Water_Temp.:92.52, PWM:33
22:54:40.508 -> Water_Temp.:92.52, PWM:33
22:54:41.300 -> Water_Temp.:92.52, PWM:33
22:54:42.856 -> Water_Temp.:92.64, PWM:33
22:54:44.413 -> Water_Temp.:92.64, PWM:33
22:54:45.942 -> Water_Temp.:92.64, PWM:33
22:54:46.737 -> Water_Temp.:92.64, PWM:33
22:54:48.260 -> Water_Temp.:92.64, PWM:33
22:54:49.056 -> Water_Temp.:92.64, PWM:33
22:54:49.816 -> Water_Temp.:92.64, PWM:33
22:54:50.612 -> Water_Temp.:92.64, PWM:33
22:54:51.406 -> Water_Temp.:92.64, PWM:33
22:54:52.168 -> Water_Temp.:92.64, PWM:33
22:54:53.731 -> Water_Temp.:92.64, PWM:33
22:54:54.497 -> Water_Temp.:92.52, PWM:33
22:54:55.290 -> Water_Temp.:92.52, PWM:33
22:54:56.085 -> Water_Temp.:92.52, PWM:33
22:54:57.614 -> Water_Temp.:92.52, PWM:33
22:54:59.172 -> Water_Temp.:92.52, PWM:33
22:55:01.487 -> Water_Temp.:92.64, PWM:33
22:55:02.285 -> Water_Temp.:92.64, PWM:33
22:55:03.050 -> Water_Temp.:92.64, PWM:33
22:55:03.848 -> Water_Temp.:92.64, PWM:33
22:55:04.613 -> Water_Temp.:92.64, PWM:33
22:55:06.171 -> Water_Temp.:92.64, PWM:33
22:55:08.489 -> Water_Temp.:92.64, PWM:33
22:55:09.288 -> Water_Temp.:92.64, PWM:33
22:55:10.816 -> Water_Temp.:92.64, PWM:33
22:55:12.405 -> Water_Temp.:92.64, PWM:33
22:55:13.167 -> Water_Temp.:92.64, PWM:33
22:55:13.964 -> Water_Temp.:92.64, PWM:33
22:55:15.519 -> Water_Temp.:92.64, PWM:33
22:55:16.279 -> Water_Temp.:92.64, PWM:33
22:55:17.073 -> Water_Temp.:92.64, PWM:33
22:55:18.632 -> Water_Temp.:92.64, PWM:33
22:55:19.393 -> Water_Temp.:92.75, PWM:33
22:55:20.186 -> Water_Temp.:92.75, PWM:33
22:55:20.980 -> Water_Temp.:92.75, PWM:33
22:55:22.501 -> Water_Temp.:92.75, PWM:33
22:55:23.294 -> Water_Temp.:92.75, PWM:33
22:55:24.854 -> Water_Temp.:92.75, PWM:33
22:55:25.617 -> Water_Temp.:92.75, PWM:33
22:55:26.416 -> Water_Temp.:92.75, PWM:33
22:55:27.177 -> Water_Temp.:92.75, PWM:33
22:55:27.975 -> Water_Temp.:92.75, PWM:33
22:55:29.536 -> Water_Temp.:92.86, PWM:33
22:55:30.295 -> Water_Temp.:92.86, PWM:33
22:55:31.095 -> Water_Temp.:92.75, PWM:33
22:55:31.855 -> Water_Temp.:92.75, PWM:33
22:55:34.182 -> Water_Temp.:92.75, PWM:33
22:55:34.944 -> Water_Temp.:92.75, PWM:33
22:55:36.497 -> Water_Temp.:92.75, PWM:33
22:55:37.290 -> Water_Temp.:92.75, PWM:33
22:55:38.089 -> Water_Temp.:92.75, PWM:33
22:55:41.939 -> Water_Temp.:92.75, PWM:33
22:55:42.700 -> Water_Temp.:92.75, PWM:33
22:55:43.496 -> Water_Temp.:92.75, PWM:33
22:55:45.821 -> Water_Temp.:92.75, PWM:33
22:55:46.584 -> Water_Temp.:92.75, PWM:33
22:55:47.350 -> Water_Temp.:92.75, PWM:33
22:55:48.151 -> Water_Temp.:92.75, PWM:33
22:55:50.437 -> Water_Temp.:92.75, PWM:33
22:55:51.994 -> Water_Temp.:92.75, PWM:33
22:55:52.756 -> Water_Temp.:92.86, PWM:33
22:55:55.839 -> Water_Temp.:92.86, PWM:33
22:55:58.894 -> Water_Temp.:92.86, PWM:33
22:55:59.691 -> Water_Temp.:92.86, PWM:33
22:56:00.453 -> Water_Temp.:92.86, PWM:33
22:56:01.251 -> Water_Temp.:92.97, PWM:33
22:56:03.568 -> Water_Temp.:92.97, PWM:33
22:56:05.125 -> Water_Temp.:92.97, PWM:33
22:56:05.883 -> Water_Temp.:92.97, PWM:33
22:56:07.447 -> Water_Temp.:92.97, PWM:33
22:56:08.238 -> Water_Temp.:92.86, PWM:33
22:56:10.529 -> Water_Temp.:92.86, PWM:33
22:56:11.325 -> Water_Temp.:92.86, PWM:33
22:56:12.122 -> Water_Temp.:92.86, PWM:33
22:56:13.643 -> Water_Temp.:92.86, PWM:33
22:56:14.407 -> Water_Temp.:92.86, PWM:33
22:56:15.200 -> Water_Temp.:92.86, PWM:33
22:56:15.960 -> Water_Temp.:92.86, PWM:33
22:56:16.759 -> Water_Temp.:92.86, PWM:33
22:56:21.399 -> Water_Temp.:94.44, PWM:33
22:56:22.955 -> Water_Temp.:94.10, PWM:33
22:56:24.515 -> Water_Temp.:93.54, PWM:33
22:56:26.040 -> Water_Temp.:93.31, PWM:33
22:56:26.832 -> Water_Temp.:93.20, PWM:33
22:56:27.597 -> Water_Temp.:93.20, PWM:33
22:56:28.393 -> Water_Temp.:93.09, PWM:33
22:56:29.915 -> Water_Temp.:92.97, PWM:33
22:56:30.709 -> Water_Temp.:92.97, PWM:33
22:56:32.232 -> Water_Temp.:92.97, PWM:33
22:56:33.785 -> Water_Temp.:92.97, PWM:33
22:56:34.579 -> Water_Temp.:92.97, PWM:33
22:56:35.338 -> Water_Temp.:92.86, PWM:33
22:56:36.895 -> Water_Temp.:92.86, PWM:33
22:56:37.696 -> Water_Temp.:92.86, PWM:33
22:56:38.462 -> Water_Temp.:92.86, PWM:33
22:56:39.261 -> Water_Temp.:92.86, PWM:33
22:56:40.025 -> Water_Temp.:92.86, PWM:33
22:56:40.822 -> Water_Temp.:92.86, PWM:33
22:56:41.584 -> Water_Temp.:92.86, PWM:33
22:56:43.143 -> Water_Temp.:92.86, PWM:33
22:56:43.907 -> Water_Temp.:92.86, PWM:33
22:56:44.704 -> Water_Temp.:92.86, PWM:33
22:56:45.498 -> Water_Temp.:92.86, PWM:33
22:56:47.784 -> Water_Temp.:92.86, PWM:33
22:56:49.344 -> Water_Temp.:92.86, PWM:33
22:56:50.106 -> Water_Temp.:92.86, PWM:33
22:56:50.902 -> Water_Temp.:92.86, PWM:33
22:56:52.458 -> Water_Temp.:92.86, PWM:33
22:56:53.998 -> Water_Temp.:92.97, PWM:33
22:56:55.550 -> Water_Temp.:92.97, PWM:33
22:56:56.348 -> Water_Temp.:92.86, PWM:33
22:56:57.112 -> Water_Temp.:92.86, PWM:33
22:56:57.904 -> Water_Temp.:92.97, PWM:33
22:56:58.663 -> Water_Temp.:92.86, PWM:33
22:56:59.460 -> Water_Temp.:92.86, PWM:33
22:57:00.222 -> Water_Temp.:92.86, PWM:33
22:57:01.014 -> Water_Temp.:92.86, PWM:33
22:57:01.808 -> Water_Temp.:92.86, PWM:33
22:57:05.646 -> Water_Temp.:92.86, PWM:33
22:57:06.405 -> Water_Temp.:92.86, PWM:33
22:57:07.205 -> Water_Temp.:92.86, PWM:33
22:57:08.763 -> Water_Temp.:92.97, PWM:33
22:57:09.525 -> Water_Temp.:92.97, PWM:33
22:57:14.132 -> Water_Temp.:92.86, PWM:33
22:57:14.889 -> Water_Temp.:92.86, PWM:33
22:57:17.969 -> Water_Temp.:92.86, PWM:33
22:57:18.729 -> Water_Temp.:92.86, PWM:33
22:57:19.528 -> Water_Temp.:92.97, PWM:33
22:57:20.286 -> Water_Temp.:92.86, PWM:33
22:57:22.602 -> Water_Temp.:92.86, PWM:33
22:57:23.398 -> Water_Temp.:92.86, PWM:33
22:57:24.925 -> Water_Temp.:92.86, PWM:33
22:57:26.485 -> Water_Temp.:92.86, PWM:33
22:57:27.245 -> Water_Temp.:92.86, PWM:33
22:57:28.803 -> Water_Temp.:92.86, PWM:33
22:57:30.397 -> Water_Temp.:92.86, PWM:33
22:57:31.155 -> Water_Temp.:92.86, PWM:33
22:57:31.952 -> Water_Temp.:92.86, PWM:33
22:57:32.711 -> Water_Temp.:92.86, PWM:33
22:57:33.506 -> Water_Temp.:92.86, PWM:33
22:57:34.269 -> Water_Temp.:92.86, PWM:33
22:57:35.061 -> Water_Temp.:92.86, PWM:33
22:57:35.824 -> Water_Temp.:92.86, PWM:33
22:57:36.622 -> Water_Temp.:92.86, PWM:33
22:57:37.384 -> Water_Temp.:92.86, PWM:33
22:57:38.176 -> Water_Temp.:92.86, PWM:33
22:57:38.942 -> Water_Temp.:92.86, PWM:33
22:57:39.741 -> Water_Temp.:92.86, PWM:33
22:57:40.503 -> Water_Temp.:92.86, PWM:33
22:57:42.062 -> Water_Temp.:92.86, PWM:33
22:57:42.862 -> Water_Temp.:92.86, PWM:33
22:57:46.670 -> Water_Temp.:92.97, PWM:33
22:57:47.465 -> Water_Temp.:92.97, PWM:33
22:57:49.021 -> Water_Temp.:93.09, PWM:33
22:57:49.782 -> Water_Temp.:93.09, PWM:33
22:57:51.340 -> Water_Temp.:93.20, PWM:33
22:57:52.892 -> Water_Temp.:93.20, PWM:33
22:57:54.418 -> Water_Temp.:93.20, PWM:33
22:57:55.213 -> Water_Temp.:93.20, PWM:33
22:57:55.976 -> Water_Temp.:93.20, PWM:33
22:57:56.766 -> Water_Temp.:93.20, PWM:33
22:57:59.081 -> Water_Temp.:93.20, PWM:33
22:58:00.628 -> Water_Temp.:93.20, PWM:33
22:58:01.421 -> Water_Temp.:93.20, PWM:33
22:58:02.216 -> Water_Temp.:93.20, PWM:33
22:58:02.978 -> Water_Temp.:93.20, PWM:33
22:58:03.774 -> Water_Temp.:93.20, PWM:33
22:58:05.337 -> Water_Temp.:93.20, PWM:33
22:58:06.101 -> Water_Temp.:93.31, PWM:33
22:58:06.893 -> Water_Temp.:93.31, PWM:33
22:58:07.656 -> Water_Temp.:93.31, PWM:33
22:58:09.214 -> Water_Temp.:93.31, PWM:33
22:58:11.570 -> Water_Temp.:93.31, PWM:33
22:58:12.329 -> Water_Temp.:93.31, PWM:33
22:58:13.890 -> Water_Temp.:93.31, PWM:33
22:58:14.688 -> Water_Temp.:93.31, PWM:33
22:58:15.444 -> Water_Temp.:93.31, PWM:33
22:58:16.235 -> Water_Temp.:93.31, PWM:33
22:58:17.029 -> Water_Temp.:93.31, PWM:33
22:58:17.790 -> Water_Temp.:93.43, PWM:33
22:58:18.590 -> Water_Temp.:93.43, PWM:33
22:58:19.351 -> Water_Temp.:93.43, PWM:33
22:58:20.904 -> Water_Temp.:93.43, PWM:33
22:58:21.699 -> Water_Temp.:93.43, PWM:33
22:58:22.460 -> Water_Temp.:93.43, PWM:33
22:58:24.021 -> Water_Temp.:93.43, PWM:33
22:58:24.780 -> Water_Temp.:93.54, PWM:33
22:58:25.573 -> Water_Temp.:93.54, PWM:33
22:58:26.340 -> Water_Temp.:93.54, PWM:33
22:58:27.139 -> Water_Temp.:93.54, PWM:33
22:58:27.901 -> Water_Temp.:93.54, PWM:33
22:58:28.696 -> Water_Temp.:93.54, PWM:33
22:58:29.456 -> Water_Temp.:93.54, PWM:33
22:58:30.253 -> Water_Temp.:93.54, PWM:33
22:58:31.012 -> Water_Temp.:93.54, PWM:33
22:58:33.363 -> Water_Temp.:93.54, PWM:33
22:58:34.126 -> Water_Temp.:93.54, PWM:33
22:58:34.919 -> Water_Temp.:93.54, PWM:33
22:58:35.714 -> Water_Temp.:93.43, PWM:33
22:58:36.474 -> Water_Temp.:93.43, PWM:33
22:58:38.029 -> Water_Temp.:93.43, PWM:33
22:58:39.592 -> Water_Temp.:93.54, PWM:33
22:58:40.385 -> Water_Temp.:93.54, PWM:33
22:58:43.491 -> Water_Temp.:93.54, PWM:33
22:58:45.036 -> Water_Temp.:93.65, PWM:33
22:58:45.835 -> Water_Temp.:93.65, PWM:33
22:58:46.595 -> Water_Temp.:93.65, PWM:33
22:58:47.389 -> Water_Temp.:93.65, PWM:33
22:58:48.152 -> Water_Temp.:93.65, PWM:33
22:58:48.951 -> Water_Temp.:93.65, PWM:33
22:58:49.713 -> Water_Temp.:93.65, PWM:33
22:58:52.064 -> Water_Temp.:93.65, PWM:33
22:58:53.619 -> Water_Temp.:93.65, PWM:33
22:58:55.177 -> Water_Temp.:93.65, PWM:33
22:58:55.939 -> Water_Temp.:93.65, PWM:33
22:58:56.737 -> Water_Temp.:93.65, PWM:33
22:58:57.498 -> Water_Temp.:93.65, PWM:33
22:58:58.295 -> Water_Temp.:93.65, PWM:33
22:58:59.062 -> Water_Temp.:93.65, PWM:33
22:59:00.621 -> Water_Temp.:93.65, PWM:33
22:59:01.379 -> Water_Temp.:93.76, PWM:33
22:59:02.173 -> Water_Temp.:93.76, PWM:33
22:59:02.936 -> Water_Temp.:93.76, PWM:33
22:59:03.728 -> Water_Temp.:93.76, PWM:33
22:59:04.489 -> Water_Temp.:93.76, PWM:33
22:59:06.807 -> Water_Temp.:93.87, PWM:33
22:59:07.604 -> Water_Temp.:93.87, PWM:33
22:59:08.367 -> Water_Temp.:93.87, PWM:33
22:59:09.167 -> Water_Temp.:93.87, PWM:33
22:59:11.489 -> Water_Temp.:93.87, PWM:33
22:59:12.245 -> Water_Temp.:93.87, PWM:33
22:59:13.047 -> Water_Temp.:93.87, PWM:33
22:59:13.810 -> Water_Temp.:93.87, PWM:33
22:59:14.608 -> Water_Temp.:93.87, PWM:33
22:59:16.163 -> Water_Temp.:93.87, PWM:33
22:59:16.927 -> Water_Temp.:93.87, PWM:33
22:59:17.690 -> Water_Temp.:93.87, PWM:33
22:59:18.486 -> Water_Temp.:93.87, PWM:33
22:59:20.014 -> Water_Temp.:93.87, PWM:33
22:59:20.778 -> Water_Temp.:93.87, PWM:33
22:59:21.576 -> Water_Temp.:93.87, PWM:33
22:59:23.132 -> Water_Temp.:93.87, PWM:33
22:59:23.931 -> Water_Temp.:93.87, PWM:33
22:59:25.461 -> Water_Temp.:93.87, PWM:33
22:59:28.543 -> Water_Temp.:93.87, PWM:33
22:59:30.100 -> Water_Temp.:93.87, PWM:33
22:59:31.657 -> Water_Temp.:93.87, PWM:33
22:59:33.204 -> Water_Temp.:93.76, PWM:33
22:59:34.000 -> Water_Temp.:93.76, PWM:33
22:59:34.763 -> Water_Temp.:93.76, PWM:33
22:59:35.560 -> Water_Temp.:93.76, PWM:33
22:59:36.323 -> Water_Temp.:93.76, PWM:33
22:59:37.882 -> Water_Temp.:93.76, PWM:33
22:59:39.443 -> Water_Temp.:93.76, PWM:33
22:59:40.204 -> Water_Temp.:93.76, PWM:33
22:59:40.999 -> Water_Temp.:93.76, PWM:33
22:59:41.763 -> Water_Temp.:93.76, PWM:33
22:59:42.559 -> Water_Temp.:93.76, PWM:33
22:59:43.318 -> Water_Temp.:93.65, PWM:33
22:59:44.115 -> Water_Temp.:93.76, PWM:33
22:59:44.881 -> Water_Temp.:93.76, PWM:33
22:59:45.678 -> Water_Temp.:93.76, PWM:33
22:59:46.439 -> Water_Temp.:93.76, PWM:33
22:59:48.003 -> Water_Temp.:93.87, PWM:33
22:59:51.083 -> Water_Temp.:93.87, PWM:33
22:59:51.844 -> Water_Temp.:93.87, PWM:33
22:59:53.397 -> Water_Temp.:93.87, PWM:33
22:59:54.954 -> Water_Temp.:93.99, PWM:33
22:59:56.509 -> Water_Temp.:93.99, PWM:33
22:59:58.058 -> Water_Temp.:93.99, PWM:33
22:59:59.615 -> Water_Temp.:94.10, PWM:33
23:00:00.409 -> Water_Temp.:93.99, PWM:33
23:00:01.169 -> Water_Temp.:93.99, PWM:33
23:00:01.965 -> Water_Temp.:94.10, PWM:33
23:00:02.760 -> Water_Temp.:94.10, PWM:33
23:00:05.079 -> Water_Temp.:94.10, PWM:33
23:00:06.602 -> Water_Temp.:94.10, PWM:33
23:00:07.401 -> Water_Temp.:94.10, PWM:33
23:00:08.166 -> Water_Temp.:94.10, PWM:33
23:00:08.965 -> Water_Temp.:94.10, PWM:33
23:00:10.486 -> Water_Temp.:94.10, PWM:33
23:00:11.281 -> Water_Temp.:94.10, PWM:33
23:00:12.042 -> Water_Temp.:94.10, PWM:33
23:00:12.840 -> Water_Temp.:94.10, PWM:33
23:00:13.599 -> Water_Temp.:94.21, PWM:33
23:00:14.398 -> Water_Temp.:94.21, PWM:33
23:00:15.159 -> Water_Temp.:94.21, PWM:33
23:00:15.954 -> Water_Temp.:94.21, PWM:33
23:00:17.509 -> Water_Temp.:94.21, PWM:33
23:00:18.268 -> Water_Temp.:94.21, PWM:33
23:00:19.064 -> Water_Temp.:94.21, PWM:33
23:00:19.857 -> Water_Temp.:94.21, PWM:33
23:00:20.613 -> Water_Temp.:94.21, PWM:33
23:00:21.405 -> Water_Temp.:94.21, PWM:33
23:00:22.198 -> Water_Temp.:94.21, PWM:33
23:00:22.958 -> Water_Temp.:94.21, PWM:33
23:00:23.754 -> Water_Temp.:94.21, PWM:33
23:00:24.514 -> Water_Temp.:94.21, PWM:33
23:00:26.069 -> Water_Temp.:94.21, PWM:33
23:00:27.631 -> Water_Temp.:94.21, PWM:33
23:00:29.157 -> Water_Temp.:94.21, PWM:33
23:00:29.952 -> Water_Temp.:94.21, PWM:33
23:00:30.749 -> Water_Temp.:94.21, PWM:33
23:00:31.505 -> Water_Temp.:94.21, PWM:33
23:00:32.302 -> Water_Temp.:94.21, PWM:33
23:00:33.857 -> Water_Temp.:94.32, PWM:33
23:00:35.418 -> Water_Temp.:94.32, PWM:33
23:00:36.179 -> Water_Temp.:94.32, PWM:33
23:00:37.736 -> Water_Temp.:94.21, PWM:33
23:00:38.535 -> Water_Temp.:94.21, PWM:33
23:00:39.299 -> Water_Temp.:94.21, PWM:33
23:00:40.092 -> Water_Temp.:94.32, PWM:33
23:00:40.855 -> Water_Temp.:94.21, PWM:33
23:00:43.176 -> Water_Temp.:94.32, PWM:33
23:00:44.735 -> Water_Temp.:94.21, PWM:33
23:00:46.295 -> Water_Temp.:94.32, PWM:33
23:00:47.087 -> Water_Temp.:94.32, PWM:33
23:00:47.850 -> Water_Temp.:94.32, PWM:33
23:00:49.408 -> Water_Temp.:94.32, PWM:33
23:00:50.205 -> Water_Temp.:94.32, PWM:33
23:00:52.518 -> Water_Temp.:94.32, PWM:33
23:00:53.278 -> Water_Temp.:94.32, PWM:33
23:00:54.042 -> Water_Temp.:94.32, PWM:33
23:00:55.603 -> Water_Temp.:94.32, PWM:33
23:00:57.149 -> Water_Temp.:94.32, PWM:33
23:00:57.916 -> Water_Temp.:94.44, PWM:33
23:00:58.710 -> Water_Temp.:94.32, PWM:33
23:00:59.470 -> Water_Temp.:94.32, PWM:33
23:01:00.994 -> Water_Temp.:94.44, PWM:33
23:01:02.581 -> Water_Temp.:94.44, PWM:33
23:01:04.141 -> Water_Temp.:94.44, PWM:33
23:01:04.902 -> Water_Temp.:94.44, PWM:33
23:01:05.697 -> Water_Temp.:94.44, PWM:33
23:01:08.011 -> Water_Temp.:94.44, PWM:33
23:01:08.777 -> Water_Temp.:94.44, PWM:33
23:01:10.331 -> Water_Temp.:94.44, PWM:33
23:01:13.409 -> Water_Temp.:94.55, PWM:33
23:01:14.173 -> Water_Temp.:94.55, PWM:33
23:01:16.528 -> Water_Temp.:94.55, PWM:33
23:01:18.087 -> Water_Temp.:94.55, PWM:33
23:01:20.409 -> Water_Temp.:94.55, PWM:33
23:01:21.173 -> Water_Temp.:94.55, PWM:33
23:01:21.966 -> Water_Temp.:94.55, PWM:33
23:01:23.522 -> Water_Temp.:94.55, PWM:33
23:01:25.082 -> Water_Temp.:94.66, PWM:33
23:01:28.919 -> Water_Temp.:94.55, PWM:33
23:01:30.479 -> Water_Temp.:94.55, PWM:33
23:01:33.563 -> Water_Temp.:94.55, PWM:33
23:01:35.126 -> Water_Temp.:94.55, PWM:33
23:01:36.678 -> Water_Temp.:94.55, PWM:33
23:01:37.437 -> Water_Temp.:94.55, PWM:33
23:01:38.230 -> Water_Temp.:94.44, PWM:33
23:01:38.994 -> Water_Temp.:94.44, PWM:33
23:01:42.109 -> Water_Temp.:94.55, PWM:33
23:01:44.432 -> Water_Temp.:94.55, PWM:33
23:01:45.194 -> Water_Temp.:94.55, PWM:33
23:01:45.990 -> Water_Temp.:94.55, PWM:33
23:01:46.751 -> Water_Temp.:94.55, PWM:33
23:01:48.310 -> Water_Temp.:94.55, PWM:33
23:01:49.104 -> Water_Temp.:94.55, PWM:33
23:01:50.661 -> Water_Temp.:94.55, PWM:33
23:01:51.455 -> Water_Temp.:94.55, PWM:33
23:01:52.216 -> Water_Temp.:94.55, PWM:33
23:01:55.294 -> Water_Temp.:94.66, PWM:33
23:01:56.093 -> Water_Temp.:94.66, PWM:33
23:01:56.856 -> Water_Temp.:94.66, PWM:33
23:01:57.650 -> Water_Temp.:94.66, PWM:33
23:01:59.200 -> Water_Temp.:94.66, PWM:33
23:01:59.961 -> Water_Temp.:94.66, PWM:33
23:02:00.761 -> Water_Temp.:94.66, PWM:33
23:02:01.525 -> Water_Temp.:94.66, PWM:33
23:02:02.323 -> Water_Temp.:94.66, PWM:33
23:02:03.854 -> Water_Temp.:94.66, PWM:33
23:02:04.652 -> Water_Temp.:94.66, PWM:33
23:02:06.172 -> Water_Temp.:94.77, PWM:33
23:02:07.764 -> Water_Temp.:94.77, PWM:33
23:02:08.528 -> Water_Temp.:94.77, PWM:33
23:02:09.326 -> Water_Temp.:94.77, PWM:33
23:02:10.086 -> Water_Temp.:94.77, PWM:33
23:02:10.886 -> Water_Temp.:94.77, PWM:33
23:02:11.648 -> Water_Temp.:94.77, PWM:33
23:02:12.412 -> Water_Temp.:94.77, PWM:33
23:02:13.206 -> Water_Temp.:94.77, PWM:33
23:02:14.003 -> Water_Temp.:94.77, PWM:33
23:02:14.764 -> Water_Temp.:94.77, PWM:33
23:02:15.559 -> Water_Temp.:94.77, PWM:33
23:02:16.321 -> Water_Temp.:94.77, PWM:33
23:02:17.119 -> Water_Temp.:94.77, PWM:33
23:02:18.679 -> Water_Temp.:94.77, PWM:33
23:02:19.441 -> Water_Temp.:94.77, PWM:33
23:02:20.997 -> Water_Temp.:94.77, PWM:33
23:02:23.313 -> Water_Temp.:94.77, PWM:33
23:02:24.080 -> Water_Temp.:94.77, PWM:33
23:02:24.842 -> Water_Temp.:94.77, PWM:33
23:02:26.436 -> Water_Temp.:94.77, PWM:33
23:02:27.196 -> Water_Temp.:94.77, PWM:33
23:02:27.991 -> Water_Temp.:94.77, PWM:33
23:02:29.512 -> Water_Temp.:94.77, PWM:33
23:02:31.075 -> Water_Temp.:94.66, PWM:33
23:02:34.124 -> Water_Temp.:94.77, PWM:33
23:02:34.887 -> Water_Temp.:94.77, PWM:33
23:02:35.683 -> Water_Temp.:94.77, PWM:33
23:02:37.969 -> Water_Temp.:94.77, PWM:33
23:02:38.770 -> Water_Temp.:94.66, PWM:33
23:02:39.530 -> Water_Temp.:94.66, PWM:33
23:02:40.326 -> Water_Temp.:94.66, PWM:33
23:02:41.851 -> Water_Temp.:94.66, PWM:33
23:02:44.943 -> Water_Temp.:94.66, PWM:33
23:02:45.738 -> Water_Temp.:94.66, PWM:33
23:02:46.493 -> Water_Temp.:94.66, PWM:33
23:02:47.290 -> Water_Temp.:94.66, PWM:33
23:02:48.050 -> Water_Temp.:94.66, PWM:33
23:02:48.852 -> Water_Temp.:94.77, PWM:33
23:02:50.383 -> Water_Temp.:94.77, PWM:33
23:02:52.706 -> Water_Temp.:94.77, PWM:33
23:02:54.233 -> Water_Temp.:94.77, PWM:33
23:02:55.027 -> Water_Temp.:94.77, PWM:33
23:02:55.790 -> Water_Temp.:94.77, PWM:33
23:02:56.586 -> Water_Temp.:94.77, PWM:33
23:02:57.349 -> Water_Temp.:94.77, PWM:33
23:02:59.669 -> Water_Temp.:94.77, PWM:33
23:03:01.228 -> Water_Temp.:94.77, PWM:33
23:03:01.990 -> Water_Temp.:94.77, PWM:33
23:03:03.545 -> Water_Temp.:94.77, PWM:33
23:03:04.310 -> Water_Temp.:94.77, PWM:33
23:03:05.867 -> Water_Temp.:94.77, PWM:33
23:03:07.425 -> Water_Temp.:94.77, PWM:33
23:03:08.951 -> Water_Temp.:94.77, PWM:33
23:03:09.746 -> Water_Temp.:94.89, PWM:33
23:03:10.538 -> Water_Temp.:94.77, PWM:33
23:03:11.294 -> Water_Temp.:94.89, PWM:33
23:03:12.089 -> Water_Temp.:94.89, PWM:33
23:03:13.647 -> Water_Temp.:94.89, PWM:33
23:03:15.204 -> Water_Temp.:94.89, PWM:33
23:03:15.996 -> Water_Temp.:94.89, PWM:33
23:03:16.755 -> Water_Temp.:96.46, PWM:33
23:03:17.549 -> Water_Temp.:98.49, PWM:33
23:03:18.314 -> Water_Temp.:96.12, PWM:33
23:03:19.113 -> Water_Temp.:95.68, PWM:33
23:03:19.875 -> Water_Temp.:95.45, PWM:33
23:03:20.675 -> Water_Temp.:95.34, PWM:33
23:03:21.434 -> Water_Temp.:95.22, PWM:33
23:03:22.233 -> Water_Temp.:95.22, PWM:33
23:03:22.993 -> Water_Temp.:95.11, PWM:33
23:03:23.785 -> Water_Temp.:95.11, PWM:33
23:03:25.336 -> Water_Temp.:95.00, PWM:33
23:03:26.125 -> Water_Temp.:95.00, PWM:33
23:03:26.884 -> Water_Temp.:95.00, PWM:33
23:03:27.682 -> Water_Temp.:95.00, PWM:33
23:03:29.204 -> Water_Temp.:94.89, PWM:33
23:03:31.521 -> Water_Temp.:94.89, PWM:33
23:03:33.077 -> Water_Temp.:95.00, PWM:33
23:03:33.837 -> Water_Temp.:95.00, PWM:33
23:03:34.633 -> Water_Temp.:95.00, PWM:33
23:03:35.387 -> Water_Temp.:95.00, PWM:33
23:03:36.182 -> Water_Temp.:95.00, PWM:33
23:03:36.943 -> Water_Temp.:95.00, PWM:33
23:03:37.740 -> Water_Temp.:95.00, PWM:33
23:03:38.537 -> Water_Temp.:95.00, PWM:33
23:03:39.298 -> Water_Temp.:95.00, PWM:33
23:03:40.091 -> Water_Temp.:95.00, PWM:33
23:03:41.616 -> Water_Temp.:95.00, PWM:33
23:03:42.379 -> Water_Temp.:95.00, PWM:33
23:03:43.178 -> Water_Temp.:95.00, PWM:33
23:03:43.944 -> Water_Temp.:95.00, PWM:33
23:03:44.741 -> Water_Temp.:95.00, PWM:33
23:03:45.501 -> Water_Temp.:95.00, PWM:33
23:03:46.300 -> Water_Temp.:95.00, PWM:33
23:03:47.823 -> Water_Temp.:95.00, PWM:33
23:03:48.621 -> Water_Temp.:95.00, PWM:33
23:03:49.384 -> Water_Temp.:95.00, PWM:33
23:03:50.184 -> Water_Temp.:95.00, PWM:33
23:03:51.715 -> Water_Temp.:95.00, PWM:33
23:03:52.511 -> Water_Temp.:95.00, PWM:33
23:03:53.274 -> Water_Temp.:95.00, PWM:33
23:03:54.072 -> Water_Temp.:95.00, PWM:33
23:03:55.593 -> Water_Temp.:95.00, PWM:33
23:03:56.392 -> Water_Temp.:95.00, PWM:33
23:03:57.146 -> Water_Temp.:95.00, PWM:33
23:03:57.940 -> Water_Temp.:95.00, PWM:33
23:03:59.502 -> Water_Temp.:95.00, PWM:33
23:04:01.795 -> Water_Temp.:95.00, PWM:33
23:04:02.591 -> Water_Temp.:94.89, PWM:33
23:04:04.153 -> Water_Temp.:94.89, PWM:33
23:04:05.676 -> Water_Temp.:94.89, PWM:33
23:04:08.016 -> Water_Temp.:94.89, PWM:33
23:04:08.778 -> Water_Temp.:94.89, PWM:33
23:04:09.573 -> Water_Temp.:94.89, PWM:33
23:04:10.335 -> Water_Temp.:94.89, PWM:33
23:04:11.894 -> Water_Temp.:94.77, PWM:33
23:04:12.657 -> Water_Temp.:94.77, PWM:33
23:04:13.448 -> Water_Temp.:94.77, PWM:33
23:04:14.242 -> Water_Temp.:94.77, PWM:33
23:04:15.759 -> Water_Temp.:94.77, PWM:33
23:04:16.556 -> Water_Temp.:94.77, PWM:33
23:04:17.319 -> Water_Temp.:94.77, PWM:33
23:04:18.113 -> Water_Temp.:94.77, PWM:33
23:04:19.669 -> Water_Temp.:94.77, PWM:33
23:04:21.222 -> Water_Temp.:94.77, PWM:33
23:04:21.986 -> Water_Temp.:94.77, PWM:33
23:04:23.541 -> Water_Temp.:94.89, PWM:33
23:04:24.304 -> Water_Temp.:94.89, PWM:33
23:04:25.096 -> Water_Temp.:94.89, PWM:33
23:04:25.859 -> Water_Temp.:94.89, PWM:33
23:04:26.656 -> Water_Temp.:94.89, PWM:33
23:04:27.419 -> Water_Temp.:94.89, PWM:33
23:04:28.212 -> Water_Temp.:94.89, PWM:33
23:04:28.974 -> Water_Temp.:94.89, PWM:33
23:04:29.770 -> Water_Temp.:94.89, PWM:33
23:04:30.531 -> Water_Temp.:94.89, PWM:33
23:04:31.328 -> Water_Temp.:94.89, PWM:33
23:04:33.646 -> Water_Temp.:94.89, PWM:33
23:04:39.094 -> Water_Temp.:94.89, PWM:33
23:04:39.890 -> Water_Temp.:94.89, PWM:33
23:04:40.654 -> Water_Temp.:94.89, PWM:33
23:04:42.216 -> Water_Temp.:95.00, PWM:33
23:04:43.738 -> Water_Temp.:95.00, PWM:33
23:04:46.064 -> Water_Temp.:95.00, PWM:33
23:04:49.112 -> Water_Temp.:95.00, PWM:33
23:04:51.430 -> Water_Temp.:95.00, PWM:33
23:04:52.192 -> Water_Temp.:95.00, PWM:33
23:04:52.990 -> Water_Temp.:95.00, PWM:33
23:04:53.751 -> Water_Temp.:95.00, PWM:33
23:04:55.279 -> Water_Temp.:95.00, PWM:33
23:04:56.079 -> Water_Temp.:95.00, PWM:33
23:04:56.871 -> Water_Temp.:95.00, PWM:33
23:04:57.634 -> Water_Temp.:95.00, PWM:33
23:04:59.961 -> Water_Temp.:95.00, PWM:33
23:05:00.722 -> Water_Temp.:95.00, PWM:33
23:05:01.518 -> Water_Temp.:95.00, PWM:33
23:05:02.276 -> Water_Temp.:95.00, PWM:33
23:05:03.072 -> Water_Temp.:95.00, PWM:33
23:05:04.599 -> Water_Temp.:95.00, PWM:33
23:05:05.392 -> Water_Temp.:95.00, PWM:33
23:05:06.156 -> Water_Temp.:95.00, PWM:33
23:05:07.716 -> Water_Temp.:95.00, PWM:33
23:05:09.246 -> Water_Temp.:95.00, PWM:33
23:05:10.808 -> Water_Temp.:95.11, PWM:33
23:05:11.570 -> Water_Temp.:95.11, PWM:33
23:05:12.369 -> Water_Temp.:95.11, PWM:33
23:05:13.133 -> Water_Temp.:95.11, PWM:33
23:05:13.933 -> Water_Temp.:95.11, PWM:33
23:05:14.697 -> Water_Temp.:95.11, PWM:33
23:05:15.492 -> Water_Temp.:95.22, PWM:33
23:05:17.051 -> Water_Temp.:95.22, PWM:33
23:05:18.574 -> Water_Temp.:95.22, PWM:33
23:05:19.368 -> Water_Temp.:95.22, PWM:33
23:05:20.129 -> Water_Temp.:95.34, PWM:33
23:05:20.923 -> Water_Temp.:95.34, PWM:33
23:05:21.720 -> Water_Temp.:95.34, PWM:33
23:05:23.279 -> Water_Temp.:95.34, PWM:33
23:05:24.807 -> Water_Temp.:95.34, PWM:33
23:05:25.604 -> Water_Temp.:95.34, PWM:33
23:05:26.400 -> Water_Temp.:95.34, PWM:33
23:05:27.162 -> Water_Temp.:95.34, PWM:33
23:05:27.956 -> Water_Temp.:95.34, PWM:33
23:05:28.714 -> Water_Temp.:95.34, PWM:33
23:05:29.510 -> Water_Temp.:95.34, PWM:33
23:05:30.270 -> Water_Temp.:95.34, PWM:33
23:05:32.576 -> Water_Temp.:95.34, PWM:33
23:05:33.342 -> Water_Temp.:95.34, PWM:33
23:05:34.934 -> Water_Temp.:95.34, PWM:33
23:05:35.695 -> Water_Temp.:95.34, PWM:33
23:05:39.570 -> Water_Temp.:95.34, PWM:33
23:05:40.332 -> Water_Temp.:95.34, PWM:33
23:05:41.129 -> Water_Temp.:95.34, PWM:33
23:05:41.890 -> Water_Temp.:95.34, PWM:33
23:05:43.449 -> Water_Temp.:95.34, PWM:33
23:05:44.208 -> Water_Temp.:95.34, PWM:33
23:05:45.002 -> Water_Temp.:95.34, PWM:33
23:05:45.765 -> Water_Temp.:95.34, PWM:33
23:05:47.322 -> Water_Temp.:95.34, PWM:33
23:05:48.121 -> Water_Temp.:95.34, PWM:33
23:05:48.886 -> Water_Temp.:95.34, PWM:33
23:05:50.439 -> Water_Temp.:95.34, PWM:33
23:05:51.994 -> Water_Temp.:95.34, PWM:33
23:05:55.084 -> Water_Temp.:95.34, PWM:33
23:05:56.647 -> Water_Temp.:95.34, PWM:33
23:05:57.405 -> Water_Temp.:95.22, PWM:33
23:05:58.199 -> Water_Temp.:95.22, PWM:33
23:05:58.994 -> Water_Temp.:95.34, PWM:33
23:06:00.513 -> Water_Temp.:95.34, PWM:33
23:06:01.306 -> Water_Temp.:95.34, PWM:33
23:06:03.619 -> Water_Temp.:95.34, PWM:33
23:06:07.464 -> Water_Temp.:95.34, PWM:33
23:06:09.026 -> Water_Temp.:95.34, PWM:33
23:06:09.822 -> Water_Temp.:95.34, PWM:33
23:06:10.583 -> Water_Temp.:95.34, PWM:33
23:06:11.374 -> Water_Temp.:95.34, PWM:33
23:06:12.136 -> Water_Temp.:95.34, PWM:33
23:06:12.931 -> Water_Temp.:95.34, PWM:33
23:06:13.692 -> Water_Temp.:95.34, PWM:33
23:06:15.251 -> Water_Temp.:95.34, PWM:33
23:06:16.798 -> Water_Temp.:95.45, PWM:33
23:06:20.598 -> Water_Temp.:95.45, PWM:33
23:06:22.150 -> Water_Temp.:95.45, PWM:33
23:06:22.916 -> Water_Temp.:95.45, PWM:33
23:06:24.478 -> Water_Temp.:95.45, PWM:33
23:06:25.273 -> Water_Temp.:95.45, PWM:33
23:06:26.825 -> Water_Temp.:95.45, PWM:33
23:06:28.349 -> Water_Temp.:95.45, PWM:33
23:06:29.908 -> Water_Temp.:95.45, PWM:33
23:06:30.671 -> Water_Temp.:95.45, PWM:33
23:06:33.774 -> Water_Temp.:95.45, PWM:33
23:06:35.335 -> Water_Temp.:95.56, PWM:33
23:06:36.131 -> Water_Temp.:95.56, PWM:33
23:06:38.459 -> Water_Temp.:95.56, PWM:33
23:06:39.258 -> Water_Temp.:95.56, PWM:33
23:06:40.780 -> Water_Temp.:95.68, PWM:33
23:06:41.575 -> Water_Temp.:95.68, PWM:33
23:06:42.337 -> Water_Temp.:95.68, PWM:33
23:06:43.129 -> Water_Temp.:95.68, PWM:33
23:06:44.686 -> Water_Temp.:95.68, PWM:33
23:06:45.478 -> Water_Temp.:95.56, PWM:33
23:06:46.237 -> Water_Temp.:95.56, PWM:33
23:06:47.036 -> Water_Temp.:95.56, PWM:33
23:06:47.800 -> Water_Temp.:95.56, PWM:33
23:06:48.597 -> Water_Temp.:95.56, PWM:33
23:06:49.361 -> Water_Temp.:95.56, PWM:33
23:06:50.153 -> Water_Temp.:95.56, PWM:33
23:06:50.911 -> Water_Temp.:95.56, PWM:33
23:06:51.707 -> Water_Temp.:95.56, PWM:33
23:06:52.471 -> Water_Temp.:95.56, PWM:33
23:06:53.269 -> Water_Temp.:97.25, PWM:33
23:06:55.585 -> Water_Temp.:96.46, PWM:33
23:06:57.939 -> Water_Temp.:96.12, PWM:33
23:06:58.704 -> Water_Temp.:96.01, PWM:33
23:06:59.503 -> Water_Temp.:96.01, PWM:33
23:07:00.262 -> Water_Temp.:97.70, PWM:33
23:07:02.583 -> Water_Temp.:96.91, PWM:33
23:07:03.345 -> Water_Temp.:96.69, PWM:33
23:07:04.902 -> Water_Temp.:96.35, PWM:33
23:07:05.697 -> Water_Temp.:96.35, PWM:33
23:07:07.258 -> Water_Temp.:96.12, PWM:33
23:07:08.781 -> Water_Temp.:96.12, PWM:33
23:07:10.372 -> Water_Temp.:96.01, PWM:33
23:07:11.890 -> Water_Temp.:95.90, PWM:33
23:07:13.448 -> Water_Temp.:95.90, PWM:33
23:07:15.010 -> Water_Temp.:95.90, PWM:33
23:07:16.537 -> Water_Temp.:95.79, PWM:33
23:07:17.331 -> Water_Temp.:95.79, PWM:33
23:07:18.090 -> Water_Temp.:95.79, PWM:33
23:07:18.885 -> Water_Temp.:95.79, PWM:33
23:07:20.444 -> Water_Temp.:95.79, PWM:33
23:07:21.997 -> Water_Temp.:95.79, PWM:33
23:07:23.528 -> Water_Temp.:95.79, PWM:33
23:07:24.293 -> Water_Temp.:95.79, PWM:33
23:07:25.089 -> Water_Temp.:95.79, PWM:33
23:07:25.851 -> Water_Temp.:95.79, PWM:33
23:07:27.415 -> Water_Temp.:95.79, PWM:33
23:07:28.944 -> Water_Temp.:95.79, PWM:33
23:07:31.296 -> Water_Temp.:95.90, PWM:33
23:07:32.094 -> Water_Temp.:95.90, PWM:33
23:07:33.622 -> Water_Temp.:95.90, PWM:33
23:07:35.941 -> Water_Temp.:95.90, PWM:33
23:07:36.702 -> Water_Temp.:95.90, PWM:33
23:07:37.499 -> Water_Temp.:95.90, PWM:33
23:07:38.261 -> Water_Temp.:95.90, PWM:33
23:07:39.789 -> Water_Temp.:95.90, PWM:33
23:07:42.146 -> Water_Temp.:95.90, PWM:33
23:07:44.462 -> Water_Temp.:95.90, PWM:33
23:07:45.992 -> Water_Temp.:95.90, PWM:33
23:07:46.780 -> Water_Temp.:95.90, PWM:33
23:07:47.541 -> Water_Temp.:95.90, PWM:33
23:07:48.340 -> Water_Temp.:95.79, PWM:33
23:07:49.105 -> Water_Temp.:95.79, PWM:33
23:07:50.660 -> Water_Temp.:95.79, PWM:33
23:07:51.424 -> Water_Temp.:95.90, PWM:33
23:07:52.222 -> Water_Temp.:95.90, PWM:33
23:07:52.984 -> Water_Temp.:95.90, PWM:33
23:07:54.507 -> Water_Temp.:95.90, PWM:33
23:07:56.857 -> Water_Temp.:95.90, PWM:33
23:07:59.172 -> Water_Temp.:95.90, PWM:33
23:07:59.933 -> Water_Temp.:96.01, PWM:33
23:08:03.801 -> Water_Temp.:95.90, PWM:33
23:08:05.328 -> Water_Temp.:95.90, PWM:33
23:08:06.124 -> Water_Temp.:95.90, PWM:33
23:08:06.886 -> Water_Temp.:95.90, PWM:33
23:08:09.210 -> Water_Temp.:95.90, PWM:33
23:08:12.309 -> Water_Temp.:95.90, PWM:33
23:08:14.596 -> Water_Temp.:95.90, PWM:33
23:08:16.117 -> Water_Temp.:95.90, PWM:33
23:08:18.432 -> Water_Temp.:95.90, PWM:33
23:08:19.198 -> Water_Temp.:95.90, PWM:33
23:08:20.753 -> Water_Temp.:96.01, PWM:33
23:08:22.342 -> Water_Temp.:96.01, PWM:33
23:08:23.867 -> Water_Temp.:96.01, PWM:33
23:08:26.151 -> Water_Temp.:96.01, PWM:33
23:08:26.944 -> Water_Temp.:96.01, PWM:33
23:08:27.739 -> Water_Temp.:96.01, PWM:33
23:08:29.264 -> Water_Temp.:96.01, PWM:33
23:08:30.056 -> Water_Temp.:96.01, PWM:33
23:08:32.357 -> Water_Temp.:96.12, PWM:33
23:08:33.911 -> Water_Temp.:96.12, PWM:33
23:08:35.468 -> Water_Temp.:96.12, PWM:33
23:08:36.228 -> Water_Temp.:96.12, PWM:33
23:08:37.811 -> Water_Temp.:96.01, PWM:33
23:08:38.575 -> Water_Temp.:96.01, PWM:33
23:08:39.369 -> Water_Temp.:96.01, PWM:33
23:08:40.919 -> Water_Temp.:96.01, PWM:33
23:08:41.677 -> Water_Temp.:96.01, PWM:33
23:08:44.035 -> Water_Temp.:96.01, PWM:33
23:08:44.793 -> Water_Temp.:96.01, PWM:33
23:08:46.350 -> Water_Temp.:96.12, PWM:33
23:08:47.144 -> Water_Temp.:96.12, PWM:33
23:08:48.695 -> Water_Temp.:96.12, PWM:33
23:08:49.455 -> Water_Temp.:96.12, PWM:33
23:08:50.248 -> Water_Temp.:96.12, PWM:33
23:08:51.041 -> Water_Temp.:96.12, PWM:33
23:08:53.366 -> Water_Temp.:96.01, PWM:33
23:08:54.132 -> Water_Temp.:96.12, PWM:33
23:08:55.651 -> Water_Temp.:96.12, PWM:33
23:08:57.210 -> Water_Temp.:96.12, PWM:33
23:08:58.798 -> Water_Temp.:96.12, PWM:33
23:08:59.559 -> Water_Temp.:96.12, PWM:33
23:09:01.115 -> Water_Temp.:96.12, PWM:33
23:09:03.428 -> Water_Temp.:96.12, PWM:33
23:09:06.481 -> Water_Temp.:96.12, PWM:33
23:09:07.279 -> Water_Temp.:96.12, PWM:33
23:09:08.042 -> Water_Temp.:96.12, PWM:33
23:09:08.842 -> Water_Temp.:96.12, PWM:33
23:09:09.601 -> Water_Temp.:96.12, PWM:33
23:09:10.392 -> Water_Temp.:96.12, PWM:33
23:09:11.154 -> Water_Temp.:96.12, PWM:33
23:09:13.467 -> Water_Temp.:96.12, PWM:33
23:09:14.259 -> Water_Temp.:96.12, PWM:33
23:09:15.786 -> Water_Temp.:96.12, PWM:33
23:09:16.583 -> Water_Temp.:96.12, PWM:33
23:09:17.345 -> Water_Temp.:96.12, PWM:33
23:09:18.143 -> Water_Temp.:96.12, PWM:33
23:09:18.905 -> Water_Temp.:96.12, PWM:33
23:09:19.704 -> Water_Temp.:96.12, PWM:33
23:09:22.751 -> Water_Temp.:96.12, PWM:33
23:09:24.312 -> Water_Temp.:96.12, PWM:33
23:09:25.077 -> Water_Temp.:96.12, PWM:33
23:09:27.399 -> Water_Temp.:96.24, PWM:33
23:09:28.195 -> Water_Temp.:96.24, PWM:33
23:09:28.959 -> Water_Temp.:96.24, PWM:33
23:09:29.752 -> Water_Temp.:96.24, PWM:33
23:09:31.308 -> Water_Temp.:96.35, PWM:33
23:09:32.066 -> Water_Temp.:96.35, PWM:33
23:09:33.622 -> Water_Temp.:96.35, PWM:33
23:09:34.387 -> Water_Temp.:96.35, PWM:33
23:09:35.178 -> Water_Temp.:96.35, PWM:33
23:09:35.945 -> Water_Temp.:96.35, PWM:33
23:09:40.553 -> Water_Temp.:96.35, PWM:33
23:09:41.311 -> Water_Temp.:96.24, PWM:33
23:09:42.107 -> Water_Temp.:96.24, PWM:33
23:09:42.873 -> Water_Temp.:96.35, PWM:33
23:09:44.430 -> Water_Temp.:96.35, PWM:33
23:09:45.193 -> Water_Temp.:96.24, PWM:33
23:09:46.751 -> Water_Temp.:96.24, PWM:33
23:09:47.548 -> Water_Temp.:96.24, PWM:33
23:09:49.069 -> Water_Temp.:96.35, PWM:33
23:09:49.864 -> Water_Temp.:96.35, PWM:33
23:09:50.660 -> Water_Temp.:96.24, PWM:33
23:09:56.023 -> Water_Temp.:96.35, PWM:33
23:09:56.787 -> Water_Temp.:96.35, PWM:33
23:09:59.077 -> Water_Temp.:96.35, PWM:33
23:09:59.868 -> Water_Temp.:96.35, PWM:33
23:10:00.627 -> Water_Temp.:96.24, PWM:33
23:10:02.183 -> Water_Temp.:96.24, PWM:33
23:10:02.948 -> Water_Temp.:96.35, PWM:33
23:10:03.743 -> Water_Temp.:96.24, PWM:33
23:10:04.506 -> Water_Temp.:96.24, PWM:33
23:10:05.299 -> Water_Temp.:96.24, PWM:33
23:10:06.093 -> Water_Temp.:96.35, PWM:33
23:10:06.849 -> Water_Temp.:96.35, PWM:33
23:10:07.642 -> Water_Temp.:96.35, PWM:33
23:10:08.402 -> Water_Temp.:96.35, PWM:33
23:10:09.200 -> Water_Temp.:96.35, PWM:33
23:10:09.961 -> Water_Temp.:96.35, PWM:33
23:10:11.515 -> Water_Temp.:96.35, PWM:33
23:10:13.065 -> Water_Temp.:96.35, PWM:33
23:10:13.828 -> Water_Temp.:96.35, PWM:33
23:10:15.390 -> Water_Temp.:96.35, PWM:33
23:10:16.183 -> Water_Temp.:96.35, PWM:33
23:10:16.977 -> Water_Temp.:96.46, PWM:33
23:10:17.739 -> Water_Temp.:96.46, PWM:33
23:10:19.291 -> Water_Temp.:96.46, PWM:33
23:10:20.086 -> Water_Temp.:96.46, PWM:33
23:10:21.602 -> Water_Temp.:96.35, PWM:33
23:10:22.399 -> Water_Temp.:96.35, PWM:33
23:10:23.160 -> Water_Temp.:96.35, PWM:33
23:10:24.719 -> Water_Temp.:96.35, PWM:33
23:10:25.477 -> Water_Temp.:96.35, PWM:33
23:10:26.275 -> Water_Temp.:96.35, PWM:33
23:10:27.034 -> Water_Temp.:96.24, PWM:33
23:10:27.828 -> Water_Temp.:96.24, PWM:33
23:10:28.592 -> Water_Temp.:96.24, PWM:33
23:10:29.388 -> Water_Temp.:96.24, PWM:33
23:10:30.939 -> Water_Temp.:96.24, PWM:33
23:10:31.702 -> Water_Temp.:96.24, PWM:33
23:10:33.256 -> Water_Temp.:96.24, PWM:33
23:10:34.817 -> Water_Temp.:96.24, PWM:33
23:10:35.578 -> Water_Temp.:96.24, PWM:33
23:10:36.344 -> Water_Temp.:96.24, PWM:33
23:10:37.933 -> Water_Temp.:96.24, PWM:33
23:10:40.221 -> Water_Temp.:96.24, PWM:33
23:10:40.981 -> Water_Temp.:96.24, PWM:33
23:10:42.503 -> Water_Temp.:96.24, PWM:33
23:10:44.096 -> Water_Temp.:96.24, PWM:33
23:10:45.619 -> Water_Temp.:96.35, PWM:33
23:10:46.415 -> Water_Temp.:96.35, PWM:33
23:10:47.173 -> Water_Temp.:96.35, PWM:33
23:10:47.971 -> Water_Temp.:96.35, PWM:33
23:10:48.730 -> Water_Temp.:96.35, PWM:33
23:10:49.529 -> Water_Temp.:96.35, PWM:33
23:10:50.291 -> Water_Temp.:96.35, PWM:33
23:10:51.088 -> Water_Temp.:96.35, PWM:33
23:10:52.614 -> Water_Temp.:96.35, PWM:33
23:10:54.172 -> Water_Temp.:96.24, PWM:33
23:10:55.699 -> Water_Temp.:96.24, PWM:33
23:10:56.492 -> Water_Temp.:96.24, PWM:33
23:10:57.285 -> Water_Temp.:96.24, PWM:33
23:10:58.045 -> Water_Temp.:96.24, PWM:33
23:10:58.844 -> Water_Temp.:96.24, PWM:33
23:11:00.395 -> Water_Temp.:96.12, PWM:33
23:11:01.155 -> Water_Temp.:96.24, PWM:33
23:11:01.949 -> Water_Temp.:96.24, PWM:33
23:11:02.748 -> Water_Temp.:96.24, PWM:33
23:11:03.507 -> Water_Temp.:96.24, PWM:33
23:11:05.063 -> Water_Temp.:96.24, PWM:33
23:11:05.858 -> Water_Temp.:96.24, PWM:33
23:11:06.619 -> Water_Temp.:96.24, PWM:33
23:11:07.411 -> Water_Temp.:96.24, PWM:33
23:11:08.174 -> Water_Temp.:96.24, PWM:33
23:11:08.974 -> Water_Temp.:96.24, PWM:33
23:11:09.736 -> Water_Temp.:96.24, PWM:33
23:11:11.294 -> Water_Temp.:96.24, PWM:33
23:11:12.847 -> Water_Temp.:96.24, PWM:33
23:11:13.611 -> Water_Temp.:96.24, PWM:33
23:11:14.376 -> Water_Temp.:96.24, PWM:33
23:11:15.171 -> Water_Temp.:96.24, PWM:33
23:11:15.967 -> Water_Temp.:96.24, PWM:33
23:11:16.729 -> Water_Temp.:96.24, PWM:33
23:11:17.524 -> Water_Temp.:96.24, PWM:33
23:11:20.603 -> Water_Temp.:96.24, PWM:33
23:11:22.160 -> Water_Temp.:96.35, PWM:33
23:11:22.920 -> Water_Temp.:96.35, PWM:33
23:11:23.720 -> Water_Temp.:96.35, PWM:33
23:11:24.483 -> Water_Temp.:96.35, PWM:33
23:11:26.042 -> Water_Temp.:96.46, PWM:33
23:11:28.359 -> Water_Temp.:96.46, PWM:33
23:11:29.882 -> Water_Temp.:96.46, PWM:33
23:11:30.678 -> Water_Temp.:96.46, PWM:33
23:11:31.475 -> Water_Temp.:96.46, PWM:33
23:11:32.235 -> Water_Temp.:96.46, PWM:33
23:11:33.788 -> Water_Temp.:96.46, PWM:33
23:11:35.348 -> Water_Temp.:96.46, PWM:33
23:11:36.111 -> Water_Temp.:96.46, PWM:33
23:11:37.674 -> Water_Temp.:96.46, PWM:33
23:11:38.435 -> Water_Temp.:96.35, PWM:33
23:11:39.230 -> Water_Temp.:96.35, PWM:33
23:11:39.992 -> Water_Temp.:96.35, PWM:33
23:11:40.789 -> Water_Temp.:96.35, PWM:33
23:11:41.552 -> Water_Temp.:96.35, PWM:33
23:11:42.349 -> Water_Temp.:96.35, PWM:33
23:11:43.110 -> Water_Temp.:96.35, PWM:33
23:11:45.426 -> Water_Temp.:96.35, PWM:33
23:11:46.977 -> Water_Temp.:96.35, PWM:33
23:11:47.776 -> Water_Temp.:96.35, PWM:33
23:11:49.333 -> Water_Temp.:96.35, PWM:33
23:11:50.091 -> Water_Temp.:96.35, PWM:33
23:11:50.884 -> Water_Temp.:96.35, PWM:33
23:11:51.642 -> Water_Temp.:96.35, PWM:33
23:11:52.434 -> Water_Temp.:96.35, PWM:33
23:11:53.227 -> Water_Temp.:96.24, PWM:33
23:11:53.991 -> Water_Temp.:96.24, PWM:33
23:11:54.787 -> Water_Temp.:96.24, PWM:33
23:11:55.547 -> Water_Temp.:96.35, PWM:33
23:11:56.343 -> Water_Temp.:96.35, PWM:33
23:11:57.102 -> Water_Temp.:96.35, PWM:33
23:11:57.898 -> Water_Temp.:96.35, PWM:33
23:11:58.694 -> Water_Temp.:96.35, PWM:33
23:11:59.455 -> Water_Temp.:96.35, PWM:33
23:12:00.252 -> Water_Temp.:96.35, PWM:33
23:12:01.014 -> Water_Temp.:96.35, PWM:33
23:12:01.813 -> Water_Temp.:96.35, PWM:33
23:12:02.575 -> Water_Temp.:96.35, PWM:33
23:12:03.366 -> Water_Temp.:96.35, PWM:33
23:12:04.927 -> Water_Temp.:96.35, PWM:33
23:12:05.688 -> Water_Temp.:96.46, PWM:33
23:12:06.487 -> Water_Temp.:96.46, PWM:33
23:12:07.246 -> Water_Temp.:96.46, PWM:33
23:12:08.042 -> Water_Temp.:96.35, PWM:33
23:12:08.808 -> Water_Temp.:96.35, PWM:33
23:12:10.366 -> Water_Temp.:96.35, PWM:33
23:12:15.044 -> Water_Temp.:100.06, PWM:33
23:12:15.803 -> Water_Temp.:97.70, PWM:33
23:12:16.597 -> Water_Temp.:97.25, PWM:33
23:12:17.359 -> Water_Temp.:97.02, PWM:33
23:12:18.154 -> Water_Temp.:97.02, PWM:33
23:12:19.677 -> Water_Temp.:96.80, PWM:33
23:12:20.473 -> Water_Temp.:96.69, PWM:33
23:12:21.233 -> Water_Temp.:96.69, PWM:33
23:12:22.789 -> Water_Temp.:96.57, PWM:33
23:12:23.582 -> Water_Temp.:96.57, PWM:33
23:12:24.347 -> Water_Temp.:96.57, PWM:33
23:12:25.144 -> Water_Temp.:96.57, PWM:33
23:12:25.908 -> Water_Temp.:96.46, PWM:33
23:12:27.467 -> Water_Temp.:96.46, PWM:33
23:12:29.757 -> Water_Temp.:96.46, PWM:33
23:12:30.553 -> Water_Temp.:96.46, PWM:33
23:12:31.313 -> Water_Temp.:96.46, PWM:33
23:12:32.112 -> Water_Temp.:96.46, PWM:33
23:12:33.638 -> Water_Temp.:96.46, PWM:33
23:12:35.169 -> Water_Temp.:96.35, PWM:33
23:12:35.963 -> Water_Temp.:96.35, PWM:33
23:12:36.756 -> Water_Temp.:96.35, PWM:33
23:12:37.520 -> Water_Temp.:96.35, PWM:33
23:12:38.315 -> Water_Temp.:96.35, PWM:33
23:12:40.605 -> Water_Temp.:96.35, PWM:33
23:12:42.160 -> Water_Temp.:96.35, PWM:33
23:12:43.710 -> Water_Temp.:96.35, PWM:33
23:12:44.505 -> Water_Temp.:96.35, PWM:33
23:12:46.026 -> Water_Temp.:96.35, PWM:33
23:12:47.577 -> Water_Temp.:96.35, PWM:33
23:12:48.372 -> Water_Temp.:96.35, PWM:33
23:12:49.134 -> Water_Temp.:96.35, PWM:33
23:12:49.931 -> Water_Temp.:96.35, PWM:33
23:12:50.689 -> Water_Temp.:96.35, PWM:33
23:12:55.326 -> Water_Temp.:96.46, PWM:33
23:12:56.088 -> Water_Temp.:96.46, PWM:33
23:12:56.877 -> Water_Temp.:96.46, PWM:33
23:12:57.669 -> Water_Temp.:96.46, PWM:33
23:12:58.430 -> Water_Temp.:96.46, PWM:33
23:12:59.226 -> Water_Temp.:96.57, PWM:33
23:13:00.749 -> Water_Temp.:96.57, PWM:33
23:13:01.547 -> Water_Temp.:96.57, PWM:33
23:13:02.310 -> Water_Temp.:96.57, PWM:33
23:13:03.106 -> Water_Temp.:96.57, PWM:33
23:13:05.389 -> Water_Temp.:96.57, PWM:33
23:13:06.181 -> Water_Temp.:96.57, PWM:33
23:13:06.937 -> Water_Temp.:96.57, PWM:33
23:13:07.733 -> Water_Temp.:96.57, PWM:33
23:13:08.526 -> Water_Temp.:96.57, PWM:33
23:13:09.292 -> Water_Temp.:96.57, PWM:33
23:13:13.174 -> Water_Temp.:96.57, PWM:33
23:13:14.698 -> Water_Temp.:96.57, PWM:33
23:13:17.014 -> Water_Temp.:96.57, PWM:33
23:13:17.777 -> Water_Temp.:96.57, PWM:33
23:13:18.574 -> Water_Temp.:96.57, PWM:33
23:13:20.103 -> Water_Temp.:96.57, PWM:33
23:13:20.867 -> Water_Temp.:96.57, PWM:33
23:13:21.660 -> Water_Temp.:96.57, PWM:33
23:13:22.423 -> Water_Temp.:96.69, PWM:33
23:13:23.217 -> Water_Temp.:96.69, PWM:33
23:13:23.981 -> Water_Temp.:96.69, PWM:33
23:13:25.543 -> Water_Temp.:96.69, PWM:33
23:13:27.097 -> Water_Temp.:96.69, PWM:33
23:13:29.418 -> Water_Temp.:96.69, PWM:33
23:13:30.182 -> Water_Temp.:96.69, PWM:33
23:13:30.977 -> Water_Temp.:96.69, PWM:33
23:13:31.737 -> Water_Temp.:96.69, PWM:33
23:13:32.534 -> Water_Temp.:96.69, PWM:33
23:13:33.296 -> Water_Temp.:96.80, PWM:33
23:13:34.093 -> Water_Temp.:96.80, PWM:33
23:13:34.885 -> Water_Temp.:96.80, PWM:33
23:13:36.411 -> Water_Temp.:96.80, PWM:27
23:13:37.206 -> Water_Temp.:96.80, PWM:28
23:13:37.973 -> Water_Temp.:96.80, PWM:27
23:13:40.299 -> Water_Temp.:96.80, PWM:27
23:13:41.853 -> Water_Temp.:96.80, PWM:28
23:13:42.653 -> Water_Temp.:96.80, PWM:27
23:13:44.971 -> Water_Temp.:96.80, PWM:27
23:13:45.733 -> Water_Temp.:96.80, PWM:27
23:13:46.496 -> Water_Temp.:96.80, PWM:27
23:13:47.292 -> Water_Temp.:96.80, PWM:27
23:13:48.855 -> Water_Temp.:96.80, PWM:27
23:13:49.617 -> Water_Temp.:96.80, PWM:27
23:13:50.414 -> Water_Temp.:96.69, PWM:28
23:13:51.970 -> Water_Temp.:96.80, PWM:28
23:13:52.734 -> Water_Temp.:96.80, PWM:27
23:13:53.498 -> Water_Temp.:96.80, PWM:27
23:13:54.295 -> Water_Temp.:96.69, PWM:27
23:13:56.620 -> Water_Temp.:96.69, PWM:27
23:13:57.385 -> Water_Temp.:96.69, PWM:27
23:13:58.182 -> Water_Temp.:96.69, PWM:27
23:14:00.467 -> Water_Temp.:96.69, PWM:28
23:14:02.029 -> Water_Temp.:96.69, PWM:27
23:14:03.590 -> Water_Temp.:96.69, PWM:27
23:14:04.354 -> Water_Temp.:96.69, PWM:27
23:14:05.150 -> Water_Temp.:96.69, PWM:28
23:14:05.912 -> Water_Temp.:96.69, PWM:28
23:14:06.708 -> Water_Temp.:96.69, PWM:27
23:14:07.471 -> Water_Temp.:96.69, PWM:27
23:14:08.272 -> Water_Temp.:96.69, PWM:28
23:14:09.036 -> Water_Temp.:96.69, PWM:27
23:14:09.826 -> Water_Temp.:96.69, PWM:27
23:14:10.590 -> Water_Temp.:96.69, PWM:28
23:14:11.386 -> Water_Temp.:96.69, PWM:28
23:14:12.905 -> Water_Temp.:96.69, PWM:28
23:14:14.464 -> Water_Temp.:96.69, PWM:28
23:14:15.221 -> Water_Temp.:96.69, PWM:28
23:14:16.016 -> Water_Temp.:96.69, PWM:27
23:14:16.810 -> Water_Temp.:96.69, PWM:28
23:14:18.331 -> Water_Temp.:96.80, PWM:28
23:14:20.682 -> Water_Temp.:96.91, PWM:28
23:14:21.444 -> Water_Temp.:96.91, PWM:28
23:14:22.240 -> Water_Temp.:96.91, PWM:27
23:14:23.003 -> Water_Temp.:96.91, PWM:28
23:14:24.554 -> Water_Temp.:96.91, PWM:28
23:14:25.350 -> Water_Temp.:96.91, PWM:27
23:14:26.110 -> Water_Temp.:96.91, PWM:27
23:14:26.908 -> Water_Temp.:96.91, PWM:28
23:14:27.668 -> Water_Temp.:96.91, PWM:28
23:14:28.462 -> Water_Temp.:97.02, PWM:27
23:14:30.021 -> Water_Temp.:97.02, PWM:28
23:14:30.778 -> Water_Temp.:97.02, PWM:28
23:14:32.334 -> Water_Temp.:97.02, PWM:27
23:14:33.896 -> Water_Temp.:96.91, PWM:28
23:14:34.660 -> Water_Temp.:96.91, PWM:27
23:14:36.975 -> Water_Temp.:96.91, PWM:27
23:14:39.306 -> Water_Temp.:96.91, PWM:27
23:14:40.102 -> Water_Temp.:96.91, PWM:28
23:14:40.861 -> Water_Temp.:96.91, PWM:27
23:14:41.658 -> Water_Temp.:96.91, PWM:27
23:14:42.444 -> Water_Temp.:96.91, PWM:28
23:14:43.203 -> Water_Temp.:96.91, PWM:27
23:14:43.997 -> Water_Temp.:96.91, PWM:27
23:14:44.755 -> Water_Temp.:96.91, PWM:27
23:14:47.081 -> Water_Temp.:96.91, PWM:28
23:14:47.880 -> Water_Temp.:96.91, PWM:27
23:14:52.486 -> Water_Temp.:96.91, PWM:27
23:14:53.246 -> Water_Temp.:97.02, PWM:27
23:14:54.045 -> Water_Temp.:97.02, PWM:27
23:14:54.807 -> Water_Temp.:96.91, PWM:27
23:14:57.122 -> Water_Temp.:96.91, PWM:27
23:14:58.681 -> Water_Temp.:96.91, PWM:27
23:14:59.446 -> Water_Temp.:96.91, PWM:27
23:15:00.243 -> Water_Temp.:96.91, PWM:27
23:15:01.771 -> Water_Temp.:96.91, PWM:28
23:15:02.569 -> Water_Temp.:96.91, PWM:27
23:15:03.335 -> Water_Temp.:96.91, PWM:28
23:15:04.097 -> Water_Temp.:96.91, PWM:28
23:15:04.895 -> Water_Temp.:96.91, PWM:27
23:15:07.215 -> Water_Temp.:96.91, PWM:27
23:15:07.975 -> Water_Temp.:96.91, PWM:28
23:15:10.294 -> Water_Temp.:96.91, PWM:28
23:15:11.091 -> Water_Temp.:96.80, PWM:27
23:15:11.852 -> Water_Temp.:96.80, PWM:28
23:15:12.645 -> Water_Temp.:96.80, PWM:27
23:15:14.203 -> Water_Temp.:96.80, PWM:28
23:15:14.964 -> Water_Temp.:96.80, PWM:27
23:15:15.756 -> Water_Temp.:96.80, PWM:28
23:15:17.309 -> Water_Temp.:96.80, PWM:28
23:15:21.159 -> Water_Temp.:96.80, PWM:27
23:15:22.705 -> Water_Temp.:96.80, PWM:27
23:15:24.264 -> Water_Temp.:96.80, PWM:27
23:15:26.581 -> Water_Temp.:96.80, PWM:27
23:15:27.379 -> Water_Temp.:96.80, PWM:28
23:15:28.934 -> Water_Temp.:96.80, PWM:27
23:15:29.697 -> Water_Temp.:96.80, PWM:27
23:15:30.494 -> Water_Temp.:96.91, PWM:27
23:15:31.254 -> Water_Temp.:96.91, PWM:27
23:15:32.048 -> Water_Temp.:96.91, PWM:27
23:15:35.896 -> Water_Temp.:96.91, PWM:27
23:15:36.690 -> Water_Temp.:96.91, PWM:27
23:15:37.449 -> Water_Temp.:96.91, PWM:27
23:15:39.772 -> Water_Temp.:96.91, PWM:28
23:15:40.532 -> Water_Temp.:96.91, PWM:27
23:15:42.088 -> Water_Temp.:96.91, PWM:28
23:15:42.885 -> Water_Temp.:96.91, PWM:27
23:15:43.645 -> Water_Temp.:96.91, PWM:27
23:15:44.412 -> Water_Temp.:96.91, PWM:27
23:15:46.730 -> Water_Temp.:96.91, PWM:28
23:15:47.526 -> Water_Temp.:96.91, PWM:27
23:15:48.289 -> Water_Temp.:96.91, PWM:28
23:15:49.088 -> Water_Temp.:96.91, PWM:27
23:15:50.602 -> Water_Temp.:96.91, PWM:28
23:15:51.398 -> Water_Temp.:96.91, PWM:28
23:15:52.163 -> Water_Temp.:96.80, PWM:28
23:15:53.720 -> Water_Temp.:96.80, PWM:27
23:15:55.243 -> Water_Temp.:96.80, PWM:28
23:15:56.037 -> Water_Temp.:96.80, PWM:27
23:15:56.797 -> Water_Temp.:96.80, PWM:28
23:15:57.589 -> Water_Temp.:96.80, PWM:28
23:15:59.151 -> Water_Temp.:96.80, PWM:27
23:16:00.707 -> Water_Temp.:96.80, PWM:27
23:16:01.469 -> Water_Temp.:96.80, PWM:27
23:16:02.262 -> Water_Temp.:98.49, PWM:27
23:16:03.055 -> Water_Temp.:100.51, PWM:27
23:16:03.814 -> Water_Temp.:98.04, PWM:28
23:16:04.612 -> Water_Temp.:97.70, PWM:27
23:16:05.375 -> Water_Temp.:97.47, PWM:28
23:16:06.929 -> Water_Temp.:97.25, PWM:27
23:16:07.720 -> Water_Temp.:97.25, PWM:28
23:16:08.485 -> Water_Temp.:97.14, PWM:28
23:16:09.249 -> Water_Temp.:97.14, PWM:27
23:16:12.326 -> Water_Temp.:96.91, PWM:27
23:16:13.123 -> Water_Temp.:96.91, PWM:27
23:16:13.887 -> Water_Temp.:96.91, PWM:27
23:16:14.685 -> Water_Temp.:96.91, PWM:28
23:16:15.451 -> Water_Temp.:96.91, PWM:28
23:16:17.736 -> Water_Temp.:96.91, PWM:27
23:16:19.298 -> Water_Temp.:96.80, PWM:28
23:16:21.654 -> Water_Temp.:96.80, PWM:28
23:16:22.417 -> Water_Temp.:96.80, PWM:28
23:16:23.213 -> Water_Temp.:96.80, PWM:27
23:16:24.772 -> Water_Temp.:96.80, PWM:27
23:16:26.325 -> Water_Temp.:96.80, PWM:27
23:16:27.854 -> Water_Temp.:96.80, PWM:27
23:16:28.649 -> Water_Temp.:96.80, PWM:27
23:16:30.961 -> Water_Temp.:96.80, PWM:27
23:16:31.719 -> Water_Temp.:96.80, PWM:27
23:16:32.515 -> Water_Temp.:96.80, PWM:27
23:16:33.279 -> Water_Temp.:96.80, PWM:29
23:16:34.078 -> Water_Temp.:96.80, PWM:32
23:16:34.837 -> Water_Temp.:96.80, PWM:32
23:16:37.965 -> Water_Temp.:96.80, PWM:32
23:16:39.496 -> Water_Temp.:96.91, PWM:28
23:16:41.055 -> Water_Temp.:96.91, PWM:28
23:16:41.817 -> Water_Temp.:96.91, PWM:29
23:16:43.407 -> Water_Temp.:96.91, PWM:28
23:16:44.170 -> Water_Temp.:96.91, PWM:28
23:16:45.729 -> Water_Temp.:96.91, PWM:29
23:16:46.528 -> Water_Temp.:96.91, PWM:28
23:16:47.290 -> Water_Temp.:96.91, PWM:29
23:16:48.847 -> Water_Temp.:96.91, PWM:29
23:16:49.646 -> Water_Temp.:96.91, PWM:29
23:16:51.200 -> Water_Temp.:96.91, PWM:29
23:16:51.967 -> Water_Temp.:96.91, PWM:29
23:16:54.250 -> Water_Temp.:96.91, PWM:31
23:16:55.042 -> Water_Temp.:96.91, PWM:31
23:16:58.892 -> Water_Temp.:96.80, PWM:31
23:16:59.686 -> Water_Temp.:96.80, PWM:31
23:17:00.483 -> Water_Temp.:96.80, PWM:31
23:17:02.769 -> Water_Temp.:96.80, PWM:31
23:17:03.568 -> Water_Temp.:96.80, PWM:31
23:17:04.331 -> Water_Temp.:96.80, PWM:31
23:17:05.120 -> Water_Temp.:96.80, PWM:31
23:17:06.667 -> Water_Temp.:96.80, PWM:31
23:17:07.465 -> Water_Temp.:96.80, PWM:31
23:17:08.229 -> Water_Temp.:96.91, PWM:31
23:17:09.757 -> Water_Temp.:96.91, PWM:31
23:17:10.551 -> Water_Temp.:96.80, PWM:31
23:17:11.313 -> Water_Temp.:96.80, PWM:31
23:17:12.107 -> Water_Temp.:96.80, PWM:31
23:17:12.867 -> Water_Temp.:96.80, PWM:31
23:17:13.665 -> Water_Temp.:96.80, PWM:31
23:17:14.428 -> Water_Temp.:96.80, PWM:31
23:17:19.064 -> Water_Temp.:96.80, PWM:31
23:17:19.827 -> Water_Temp.:96.80, PWM:31
23:17:20.626 -> Water_Temp.:96.80, PWM:31
23:17:21.391 -> Water_Temp.:96.80, PWM:31
23:17:22.185 -> Water_Temp.:96.80, PWM:31
23:17:22.947 -> Water_Temp.:96.80, PWM:31
23:17:23.743 -> Water_Temp.:96.80, PWM:31
23:17:24.502 -> Water_Temp.:96.80, PWM:31
23:17:26.059 -> Water_Temp.:96.80, PWM:31
23:17:26.851 -> Water_Temp.:96.80, PWM:31
23:17:27.615 -> Water_Temp.:96.80, PWM:31
23:17:29.176 -> Water_Temp.:96.80, PWM:31
23:17:29.938 -> Water_Temp.:96.80, PWM:31
23:17:31.499 -> Water_Temp.:96.80, PWM:31
23:17:32.262 -> Water_Temp.:96.80, PWM:31
23:17:33.059 -> Water_Temp.:96.80, PWM:31
23:17:34.618 -> Water_Temp.:96.80, PWM:31
23:17:36.177 -> Water_Temp.:96.80, PWM:31
23:17:36.937 -> Water_Temp.:96.80, PWM:31
23:17:37.732 -> Water_Temp.:96.80, PWM:31
23:17:38.493 -> Water_Temp.:96.80, PWM:31
23:17:39.291 -> Water_Temp.:96.80, PWM:31
23:17:40.056 -> Water_Temp.:96.80, PWM:31
23:17:40.851 -> Water_Temp.:96.80, PWM:31
23:17:41.617 -> Water_Temp.:96.80, PWM:31
23:17:44.735 -> Water_Temp.:96.80, PWM:31
23:17:45.531 -> Water_Temp.:96.80, PWM:31
23:17:46.297 -> Water_Temp.:96.80, PWM:31
23:17:47.093 -> Water_Temp.:96.80, PWM:31
23:17:47.856 -> Water_Temp.:96.80, PWM:31
23:17:49.412 -> Water_Temp.:96.80, PWM:31
23:17:50.176 -> Water_Temp.:96.80, PWM:31
23:17:50.975 -> Water_Temp.:96.80, PWM:31
23:17:51.736 -> Water_Temp.:96.80, PWM:31
23:17:52.534 -> Water_Temp.:96.80, PWM:31
23:17:53.296 -> Water_Temp.:96.80, PWM:31
23:17:54.857 -> Water_Temp.:96.80, PWM:31
23:17:56.414 -> Water_Temp.:96.80, PWM:31
23:17:57.176 -> Water_Temp.:96.80, PWM:31
23:18:01.805 -> Water_Temp.:96.80, PWM:31
23:18:02.606 -> Water_Temp.:96.80, PWM:31
23:18:03.369 -> Water_Temp.:96.80, PWM:31
23:18:04.164 -> Water_Temp.:96.80, PWM:31
23:18:04.923 -> Water_Temp.:96.80, PWM:31
23:18:05.721 -> Water_Temp.:96.80, PWM:31
23:18:08.764 -> Water_Temp.:96.80, PWM:31
23:18:11.858 -> Water_Temp.:96.80, PWM:31
23:18:12.653 -> Water_Temp.:96.80, PWM:31
23:18:13.420 -> Water_Temp.:96.80, PWM:31
23:18:14.221 -> Water_Temp.:96.80, PWM:31
23:18:15.777 -> Water_Temp.:96.80, PWM:31
23:18:17.301 -> Water_Temp.:96.80, PWM:31
23:18:20.393 -> Water_Temp.:96.69, PWM:31
23:18:21.157 -> Water_Temp.:96.80, PWM:31
23:18:21.953 -> Water_Temp.:96.69, PWM:31
23:18:22.746 -> Water_Temp.:96.69, PWM:31
23:18:23.509 -> Water_Temp.:96.69, PWM:31
23:18:24.306 -> Water_Temp.:96.80, PWM:31
23:18:25.063 -> Water_Temp.:96.80, PWM:30
23:18:26.619 -> Water_Temp.:96.80, PWM:31
23:18:27.414 -> Water_Temp.:96.80, PWM:31
23:18:30.497 -> Water_Temp.:96.80, PWM:31
23:18:31.295 -> Water_Temp.:96.80, PWM:31
23:18:34.316 -> Water_Temp.:96.80, PWM:31
23:18:36.634 -> Water_Temp.:96.80, PWM:31
23:18:39.681 -> Water_Temp.:96.91, PWM:31
23:18:40.444 -> Water_Temp.:96.91, PWM:31
23:18:43.558 -> Water_Temp.:96.91, PWM:31
23:18:44.323 -> Water_Temp.:96.91, PWM:31
23:18:45.116 -> Water_Temp.:96.80, PWM:31
23:18:46.677 -> Water_Temp.:96.80, PWM:31
23:18:48.993 -> Water_Temp.:96.80, PWM:31
23:18:49.787 -> Water_Temp.:96.80, PWM:31
23:18:50.585 -> Water_Temp.:96.80, PWM:31
23:18:51.345 -> Water_Temp.:96.80, PWM:31
23:18:53.659 -> Water_Temp.:96.80, PWM:31
23:18:54.459 -> Water_Temp.:96.80, PWM:31
23:18:55.977 -> Water_Temp.:96.91, PWM:31
23:18:57.535 -> Water_Temp.:96.91, PWM:31
23:18:59.097 -> Water_Temp.:96.91, PWM:31
23:18:59.858 -> Water_Temp.:96.91, PWM:31
23:19:00.653 -> Water_Temp.:96.91, PWM:31
23:19:01.416 -> Water_Temp.:96.91, PWM:31
23:19:03.760 -> Water_Temp.:96.91, PWM:31
23:19:05.309 -> Water_Temp.:96.80, PWM:31
23:19:06.092 -> Water_Temp.:96.80, PWM:31
23:19:06.851 -> Water_Temp.:96.80, PWM:31
23:19:07.644 -> Water_Temp.:96.80, PWM:31
23:19:08.405 -> Water_Temp.:96.80, PWM:31
23:19:10.734 -> Water_Temp.:96.80, PWM:31
23:19:11.532 -> Water_Temp.:96.80, PWM:31
23:19:13.093 -> Water_Temp.:96.91, PWM:31
23:19:13.856 -> Water_Temp.:96.91, PWM:31
23:19:14.621 -> Water_Temp.:96.91, PWM:31
23:19:15.420 -> Water_Temp.:97.02, PWM:31
23:19:16.180 -> Water_Temp.:97.02, PWM:31
23:19:19.270 -> Water_Temp.:97.02, PWM:31
23:19:20.034 -> Water_Temp.:97.02, PWM:31
23:19:20.832 -> Water_Temp.:96.91, PWM:31
23:19:21.597 -> Water_Temp.:96.91, PWM:31
23:19:22.394 -> Water_Temp.:96.91, PWM:31
23:19:23.184 -> Water_Temp.:96.91, PWM:31
23:19:23.950 -> Water_Temp.:96.91, PWM:31
23:19:24.714 -> Water_Temp.:96.91, PWM:31
23:19:25.509 -> Water_Temp.:96.91, PWM:31
23:19:27.072 -> Water_Temp.:96.91, PWM:31
23:19:27.865 -> Water_Temp.:96.91, PWM:31
23:19:28.626 -> Water_Temp.:96.91, PWM:31
23:19:29.424 -> Water_Temp.:96.91, PWM:31
23:19:30.182 -> Water_Temp.:96.91, PWM:31
23:19:31.733 -> Water_Temp.:96.91, PWM:31
23:19:33.292 -> Water_Temp.:96.91, PWM:31
23:19:34.052 -> Water_Temp.:96.91, PWM:31
23:19:35.609 -> Water_Temp.:96.91, PWM:31
23:19:36.367 -> Water_Temp.:96.91, PWM:31
23:19:37.925 -> Water_Temp.:96.91, PWM:31
23:19:39.479 -> Water_Temp.:96.91, PWM:31
23:19:40.239 -> Water_Temp.:96.91, PWM:31
23:19:41.032 -> Water_Temp.:97.02, PWM:31
23:19:41.826 -> Water_Temp.:97.02, PWM:31
23:19:44.139 -> Water_Temp.:97.02, PWM:31
23:19:46.446 -> Water_Temp.:97.02, PWM:31
23:19:47.242 -> Water_Temp.:97.02, PWM:31
23:19:48.760 -> Water_Temp.:97.02, PWM:31
23:19:49.553 -> Water_Temp.:97.02, PWM:31
23:19:50.313 -> Water_Temp.:96.91, PWM:31
23:19:51.110 -> Water_Temp.:96.91, PWM:31
23:19:51.905 -> Water_Temp.:96.91, PWM:31
23:19:52.667 -> Water_Temp.:96.91, PWM:31
23:19:53.462 -> Water_Temp.:96.91, PWM:31
23:19:54.228 -> Water_Temp.:96.91, PWM:31
23:19:55.789 -> Water_Temp.:96.91, PWM:31
23:19:56.585 -> Water_Temp.:96.91, PWM:31
23:19:57.347 -> Water_Temp.:96.91, PWM:31
23:19:58.877 -> Water_Temp.:96.91, PWM:31
23:19:59.672 -> Water_Temp.:97.02, PWM:31
23:20:00.467 -> Water_Temp.:97.02, PWM:31
23:20:01.229 -> Water_Temp.:97.02, PWM:31
23:20:02.020 -> Water_Temp.:97.02, PWM:31
23:20:02.779 -> Water_Temp.:97.02, PWM:31
23:20:03.576 -> Water_Temp.:97.02, PWM:31
23:20:04.341 -> Water_Temp.:97.02, PWM:31
23:20:05.135 -> Water_Temp.:97.02, PWM:31
23:20:05.897 -> Water_Temp.:97.02, PWM:31
23:20:07.449 -> Water_Temp.:97.14, PWM:31
23:20:08.245 -> Water_Temp.:97.14, PWM:31
23:20:09.773 -> Water_Temp.:97.14, PWM:31
23:20:11.332 -> Water_Temp.:97.02, PWM:31
23:20:12.089 -> Water_Temp.:97.02, PWM:31
23:20:12.889 -> Water_Temp.:97.02, PWM:31
23:20:13.652 -> Water_Temp.:97.02, PWM:31
23:20:15.178 -> Water_Temp.:97.02, PWM:31
23:20:15.976 -> Water_Temp.:97.02, PWM:31
23:20:16.738 -> Water_Temp.:97.02, PWM:31
23:20:18.298 -> Water_Temp.:97.02, PWM:31
23:20:19.094 -> Water_Temp.:97.14, PWM:31
23:20:20.649 -> Water_Temp.:97.14, PWM:31
23:20:21.408 -> Water_Temp.:97.14, PWM:31
23:20:22.207 -> Water_Temp.:97.14, PWM:30
23:20:22.966 -> Water_Temp.:97.14, PWM:31
23:20:23.761 -> Water_Temp.:97.14, PWM:31
23:20:24.523 -> Water_Temp.:97.14, PWM:31
23:20:25.318 -> Water_Temp.:97.14, PWM:31
23:20:26.868 -> Water_Temp.:97.14, PWM:31
23:20:28.394 -> Water_Temp.:97.02, PWM:31
23:20:29.191 -> Water_Temp.:97.14, PWM:31
23:20:30.746 -> Water_Temp.:97.14, PWM:31
23:20:31.539 -> Water_Temp.:97.14, PWM:31
23:20:32.302 -> Water_Temp.:97.14, PWM:31
23:20:33.098 -> Water_Temp.:97.14, PWM:31
23:20:33.855 -> Water_Temp.:97.14, PWM:31
23:20:36.178 -> Water_Temp.:97.14, PWM:31
23:20:37.735 -> Water_Temp.:97.14, PWM:31
23:20:38.498 -> Water_Temp.:97.25, PWM:31
23:20:40.061 -> Water_Temp.:97.25, PWM:31
23:20:40.823 -> Water_Temp.:97.25, PWM:31
23:20:41.624 -> Water_Temp.:97.25, PWM:31
23:20:42.383 -> Water_Temp.:97.25, PWM:31
23:20:43.176 -> Water_Temp.:97.25, PWM:31
23:20:43.940 -> Water_Temp.:97.25, PWM:31
23:20:46.983 -> Water_Temp.:97.25, PWM:31
23:20:47.777 -> Water_Temp.:97.25, PWM:31
23:20:48.572 -> Water_Temp.:97.25, PWM:31
23:20:49.331 -> Water_Temp.:97.14, PWM:31
23:20:51.648 -> Water_Temp.:97.25, PWM:31
23:20:52.414 -> Water_Temp.:97.25, PWM:31
23:20:53.209 -> Water_Temp.:97.25, PWM:30
23:20:54.729 -> Water_Temp.:97.25, PWM:31
23:20:57.851 -> Water_Temp.:97.25, PWM:31
23:20:58.613 -> Water_Temp.:97.25, PWM:31
23:20:59.375 -> Water_Temp.:97.25, PWM:31
23:21:00.175 -> Water_Temp.:97.25, PWM:31
23:21:00.937 -> Water_Temp.:97.25, PWM:31
23:21:01.732 -> Water_Temp.:97.25, PWM:31
23:21:02.529 -> Water_Temp.:97.25, PWM:31
23:21:03.294 -> Water_Temp.:97.25, PWM:31
23:21:04.086 -> Water_Temp.:97.25, PWM:31
23:21:04.850 -> Water_Temp.:97.36, PWM:31
23:21:05.644 -> Water_Temp.:97.25, PWM:31
23:21:07.961 -> Water_Temp.:97.25, PWM:31
23:21:08.725 -> Water_Temp.:97.25, PWM:31
23:21:10.280 -> Water_Temp.:97.25, PWM:31
23:21:11.841 -> Water_Temp.:97.36, PWM:31
23:21:12.604 -> Water_Temp.:97.36, PWM:31
23:21:13.400 -> Water_Temp.:97.36, PWM:31
23:21:14.951 -> Water_Temp.:97.36, PWM:31
23:21:16.509 -> Water_Temp.:97.47, PWM:31
23:21:18.029 -> Water_Temp.:97.36, PWM:31
23:21:18.827 -> Water_Temp.:97.36, PWM:31
23:21:19.623 -> Water_Temp.:97.36, PWM:31
23:21:20.384 -> Water_Temp.:97.36, PWM:31
23:21:21.943 -> Water_Temp.:97.36, PWM:31
23:21:22.709 -> Water_Temp.:97.36, PWM:31
23:21:24.987 -> Water_Temp.:97.36, PWM:31
23:21:26.581 -> Water_Temp.:97.25, PWM:31
23:21:28.902 -> Water_Temp.:97.25, PWM:31
23:21:29.666 -> Water_Temp.:97.25, PWM:31
23:21:30.462 -> Water_Temp.:97.25, PWM:31
23:21:31.255 -> Water_Temp.:97.14, PWM:31
23:21:32.011 -> Water_Temp.:97.14, PWM:30
23:21:33.568 -> Water_Temp.:97.14, PWM:31
23:21:35.092 -> Water_Temp.:97.14, PWM:31
23:21:35.888 -> Water_Temp.:97.14, PWM:30
23:21:36.680 -> Water_Temp.:97.14, PWM:31
23:21:37.437 -> Water_Temp.:97.14, PWM:31
23:21:38.229 -> Water_Temp.:97.14, PWM:31
23:21:39.749 -> Water_Temp.:97.14, PWM:31
23:21:40.544 -> Water_Temp.:97.14, PWM:31
23:21:41.337 -> Water_Temp.:97.14, PWM:31
23:21:42.103 -> Water_Temp.:97.14, PWM:31
23:21:43.656 -> Water_Temp.:97.14, PWM:31
23:21:45.206 -> Water_Temp.:97.14, PWM:31
23:21:45.998 -> Water_Temp.:97.14, PWM:31
23:21:46.759 -> Water_Temp.:97.14, PWM:31
23:21:48.319 -> Water_Temp.:97.14, PWM:31
23:21:49.118 -> Water_Temp.:97.14, PWM:31
23:21:49.881 -> Water_Temp.:97.14, PWM:31
23:21:50.674 -> Water_Temp.:97.14, PWM:31
23:21:51.437 -> Water_Temp.:97.14, PWM:31
23:21:52.233 -> Water_Temp.:97.14, PWM:31
23:21:53.756 -> Water_Temp.:97.14, PWM:31
23:21:55.316 -> Water_Temp.:97.25, PWM:31
23:21:57.630 -> Water_Temp.:97.25, PWM:31
23:21:59.946 -> Water_Temp.:97.36, PWM:31
23:22:00.745 -> Water_Temp.:97.36, PWM:31
23:22:01.536 -> Water_Temp.:97.36, PWM:31
23:22:03.093 -> Water_Temp.:97.36, PWM:31
23:22:04.649 -> Water_Temp.:97.36, PWM:31
23:22:06.173 -> Water_Temp.:97.36, PWM:31
23:22:07.735 -> Water_Temp.:97.36, PWM:31
23:22:09.288 -> Water_Temp.:97.36, PWM:31
23:22:10.081 -> Water_Temp.:97.36, PWM:31
23:22:10.878 -> Water_Temp.:97.36, PWM:31
23:22:11.641 -> Water_Temp.:97.36, PWM:31
23:22:14.722 -> Water_Temp.:97.25, PWM:31
23:22:15.483 -> Water_Temp.:97.25, PWM:31
23:22:16.280 -> Water_Temp.:97.25, PWM:31
23:22:17.041 -> Water_Temp.:97.25, PWM:31
23:22:17.842 -> Water_Temp.:97.25, PWM:31
23:22:18.604 -> Water_Temp.:97.25, PWM:31
23:22:19.406 -> Water_Temp.:97.25, PWM:31
23:22:20.165 -> Water_Temp.:97.25, PWM:31
23:22:20.959 -> Water_Temp.:97.25, PWM:31
23:22:23.279 -> Water_Temp.:97.25, PWM:31
23:22:24.079 -> Water_Temp.:97.25, PWM:31
23:22:25.632 -> Water_Temp.:97.25, PWM:31
23:22:26.393 -> Water_Temp.:97.25, PWM:31
23:22:28.712 -> Water_Temp.:97.25, PWM:31
23:22:29.509 -> Water_Temp.:97.25, PWM:30
23:22:30.264 -> Water_Temp.:97.25, PWM:31
23:22:31.058 -> Water_Temp.:97.25, PWM:31
23:22:31.849 -> Water_Temp.:97.25, PWM:31
23:22:34.139 -> Water_Temp.:97.25, PWM:31
23:22:34.934 -> Water_Temp.:97.25, PWM:30
23:22:35.693 -> Water_Temp.:97.25, PWM:31
23:22:38.774 -> Water_Temp.:97.25, PWM:31
23:22:39.569 -> Water_Temp.:97.25, PWM:31
23:22:40.330 -> Water_Temp.:97.25, PWM:31
23:22:41.885 -> Water_Temp.:97.36, PWM:31
23:22:42.681 -> Water_Temp.:97.36, PWM:31
23:22:43.474 -> Water_Temp.:97.36, PWM:31
23:22:44.997 -> Water_Temp.:97.36, PWM:31
23:22:46.549 -> Water_Temp.:97.25, PWM:31
23:22:47.315 -> Water_Temp.:97.25, PWM:31
23:22:48.103 -> Water_Temp.:97.36, PWM:31
23:22:48.896 -> Water_Temp.:97.36, PWM:31
23:22:49.658 -> Water_Temp.:97.36, PWM:31
23:22:50.450 -> Water_Temp.:97.36, PWM:31
23:22:51.212 -> Water_Temp.:97.36, PWM:31
23:22:52.011 -> Water_Temp.:97.36, PWM:31
23:22:53.527 -> Water_Temp.:97.36, PWM:30
23:22:54.321 -> Water_Temp.:97.36, PWM:31
23:22:55.113 -> Water_Temp.:97.25, PWM:31
23:22:55.874 -> Water_Temp.:97.25, PWM:31
23:22:56.669 -> Water_Temp.:97.25, PWM:31
23:22:57.429 -> Water_Temp.:97.25, PWM:31
23:22:59.017 -> Water_Temp.:97.25, PWM:31
23:22:59.779 -> Water_Temp.:97.25, PWM:31
23:23:00.575 -> Water_Temp.:97.25, PWM:31
23:23:01.334 -> Water_Temp.:97.25, PWM:31
23:23:02.891 -> Water_Temp.:97.25, PWM:31
23:23:05.211 -> Water_Temp.:97.25, PWM:31
23:23:05.977 -> Water_Temp.:97.25, PWM:31
23:23:06.772 -> Water_Temp.:97.25, PWM:31
23:23:09.818 -> Water_Temp.:97.25, PWM:31
23:23:10.615 -> Water_Temp.:97.14, PWM:31
23:23:11.409 -> Water_Temp.:97.14, PWM:31
23:23:12.170 -> Water_Temp.:97.14, PWM:31
23:23:12.968 -> Water_Temp.:97.14, PWM:31
23:23:13.732 -> Water_Temp.:97.14, PWM:31
23:23:14.530 -> Water_Temp.:97.14, PWM:31
23:23:16.054 -> Water_Temp.:97.14, PWM:31
23:23:16.849 -> Water_Temp.:97.14, PWM:31
23:23:17.612 -> Water_Temp.:97.14, PWM:31
23:23:18.411 -> Water_Temp.:97.14, PWM:31
23:23:19.933 -> Water_Temp.:97.14, PWM:31
23:23:20.731 -> Water_Temp.:97.25, PWM:31
23:23:21.494 -> Water_Temp.:97.25, PWM:31
23:23:22.288 -> Water_Temp.:97.25, PWM:31
23:23:23.082 -> Water_Temp.:97.25, PWM:31
23:23:23.847 -> Water_Temp.:97.25, PWM:31
23:23:24.642 -> Water_Temp.:97.25, PWM:31
23:23:25.406 -> Water_Temp.:97.25, PWM:31
23:23:26.200 -> Water_Temp.:97.25, PWM:31
23:23:27.714 -> Water_Temp.:97.25, PWM:31
23:23:29.278 -> Water_Temp.:97.14, PWM:31
23:23:30.041 -> Water_Temp.:97.14, PWM:31
23:23:31.597 -> Water_Temp.:97.14, PWM:31
23:23:32.389 -> Water_Temp.:97.14, PWM:31
23:23:33.150 -> Water_Temp.:97.14, PWM:31
23:23:33.945 -> Water_Temp.:97.25, PWM:31
23:23:35.470 -> Water_Temp.:97.25, PWM:31
23:23:36.263 -> Water_Temp.:97.25, PWM:31
23:23:38.583 -> Water_Temp.:97.25, PWM:31
23:23:39.342 -> Water_Temp.:97.25, PWM:31
23:23:40.141 -> Water_Temp.:97.25, PWM:31
23:23:41.664 -> Water_Temp.:97.25, PWM:31
23:23:42.428 -> Water_Temp.:97.36, PWM:31
23:23:43.225 -> Water_Temp.:97.36, PWM:31
23:23:43.989 -> Water_Temp.:97.36, PWM:31
23:23:44.787 -> Water_Temp.:97.36, PWM:31
23:23:45.547 -> Water_Temp.:97.25, PWM:31
23:23:46.344 -> Water_Temp.:97.25, PWM:31
23:23:49.431 -> Water_Temp.:97.25, PWM:31
23:23:50.192 -> Water_Temp.:97.25, PWM:31
23:23:50.990 -> Water_Temp.:97.25, PWM:31
23:23:51.751 -> Water_Temp.:97.14, PWM:31
23:23:54.065 -> Water_Temp.:97.14, PWM:31
23:23:54.826 -> Water_Temp.:97.14, PWM:31
23:23:57.145 -> Water_Temp.:97.14, PWM:31
23:23:57.939 -> Water_Temp.:97.02, PWM:31
23:23:58.705 -> Water_Temp.:97.02, PWM:31
23:24:00.263 -> Water_Temp.:97.02, PWM:31
23:24:01.816 -> Water_Temp.:97.02, PWM:31
23:24:02.607 -> Water_Temp.:97.02, PWM:31
23:24:03.372 -> Water_Temp.:97.02, PWM:31
23:24:04.169 -> Water_Temp.:97.02, PWM:31
23:24:05.721 -> Water_Temp.:97.02, PWM:31
23:24:06.519 -> Water_Temp.:97.02, PWM:31
23:24:08.835 -> Water_Temp.:97.14, PWM:31
23:24:09.596 -> Water_Temp.:97.14, PWM:31
23:24:10.393 -> Water_Temp.:97.14, PWM:31
23:24:11.156 -> Water_Temp.:97.14, PWM:31
23:24:11.956 -> Water_Temp.:97.14, PWM:31
23:24:14.278 -> Water_Temp.:97.14, PWM:31
23:24:15.039 -> Water_Temp.:97.14, PWM:31
23:24:15.833 -> Water_Temp.:97.14, PWM:31
23:24:16.593 -> Water_Temp.:97.14, PWM:31
23:24:17.393 -> Water_Temp.:97.25, PWM:31
23:24:18.154 -> Water_Temp.:97.25, PWM:31
23:24:19.714 -> Water_Temp.:97.25, PWM:31
23:24:22.068 -> Water_Temp.:97.14, PWM:31
23:24:22.831 -> Water_Temp.:97.14, PWM:31
23:24:23.630 -> Water_Temp.:97.14, PWM:31
23:24:29.008 -> Water_Temp.:97.14, PWM:31
23:24:30.566 -> Water_Temp.:97.14, PWM:31
23:24:31.329 -> Water_Temp.:97.14, PWM:31
23:24:32.128 -> Water_Temp.:97.14, PWM:31
23:24:33.650 -> Water_Temp.:97.14, PWM:31
23:24:34.414 -> Water_Temp.:97.14, PWM:31
23:24:37.521 -> Water_Temp.:97.25, PWM:30
23:24:39.043 -> Water_Temp.:97.25, PWM:31
23:24:39.837 -> Water_Temp.:97.25, PWM:31
23:24:40.599 -> Water_Temp.:97.25, PWM:31
23:24:41.393 -> Water_Temp.:97.25, PWM:31
23:24:42.192 -> Water_Temp.:97.25, PWM:31
23:24:44.510 -> Water_Temp.:97.25, PWM:31
23:24:45.304 -> Water_Temp.:97.25, PWM:31
23:24:46.067 -> Water_Temp.:97.25, PWM:31
23:24:46.867 -> Water_Temp.:97.25, PWM:31
23:24:48.392 -> Water_Temp.:97.25, PWM:31
23:24:49.157 -> Water_Temp.:97.25, PWM:31
23:24:50.720 -> Water_Temp.:97.25, PWM:31
23:24:51.480 -> Water_Temp.:97.25, PWM:31
23:24:52.274 -> Water_Temp.:97.25, PWM:30
23:24:53.041 -> Water_Temp.:97.25, PWM:31
23:24:54.601 -> Water_Temp.:97.25, PWM:31
23:24:56.924 -> Water_Temp.:97.25, PWM:31
23:24:57.689 -> Water_Temp.:97.25, PWM:31
23:24:58.483 -> Water_Temp.:97.25, PWM:31
23:25:00.011 -> Water_Temp.:97.25, PWM:31
23:25:01.567 -> Water_Temp.:97.25, PWM:31
23:25:03.122 -> Water_Temp.:97.25, PWM:31
23:25:03.886 -> Water_Temp.:97.25, PWM:31
23:25:04.682 -> Water_Temp.:97.25, PWM:31
23:25:05.440 -> Water_Temp.:97.25, PWM:31
23:25:06.239 -> Water_Temp.:97.25, PWM:31
23:25:10.055 -> Water_Temp.:97.25, PWM:31
23:25:10.821 -> Water_Temp.:97.25, PWM:31
23:25:11.616 -> Water_Temp.:97.25, PWM:31
23:25:13.168 -> Water_Temp.:97.36, PWM:31
23:25:13.964 -> Water_Temp.:97.36, PWM:31
23:25:14.727 -> Water_Temp.:97.36, PWM:31
23:25:15.523 -> Water_Temp.:97.36, PWM:31
23:25:17.803 -> Water_Temp.:97.25, PWM:31
23:25:19.368 -> Water_Temp.:97.25, PWM:30
23:25:20.168 -> Water_Temp.:97.25, PWM:31
23:25:20.933 -> Water_Temp.:97.25, PWM:31
23:25:22.489 -> Water_Temp.:97.25, PWM:31
23:25:23.288 -> Water_Temp.:97.25, PWM:31
23:25:24.048 -> Water_Temp.:97.25, PWM:31
23:25:26.408 -> Water_Temp.:97.25, PWM:31
23:25:29.494 -> Water_Temp.:97.36, PWM:31
23:25:30.255 -> Water_Temp.:97.36, PWM:35
23:25:31.052 -> Water_Temp.:97.36, PWM:35
23:25:31.813 -> Water_Temp.:97.36, PWM:35
23:25:35.658 -> Water_Temp.:97.36, PWM:35
23:25:37.179 -> Water_Temp.:97.36, PWM:33
23:25:39.501 -> Water_Temp.:97.25, PWM:33
23:25:40.298 -> Water_Temp.:97.25, PWM:33
23:25:41.854 -> Water_Temp.:97.25, PWM:33
23:25:43.410 -> Water_Temp.:97.25, PWM:33
23:25:44.965 -> Water_Temp.:97.25, PWM:33
23:25:45.727 -> Water_Temp.:97.25, PWM:33
23:25:46.523 -> Water_Temp.:97.25, PWM:33
23:25:47.280 -> Water_Temp.:97.25, PWM:33
23:25:48.074 -> Water_Temp.:97.25, PWM:33
23:25:48.869 -> Water_Temp.:97.25, PWM:33
23:25:50.429 -> Water_Temp.:97.25, PWM:33
23:25:51.189 -> Water_Temp.:97.25, PWM:33
23:25:51.990 -> Water_Temp.:97.25, PWM:33
23:25:52.754 -> Water_Temp.:97.25, PWM:33
23:25:54.288 -> Water_Temp.:97.25, PWM:33
23:25:56.613 -> Water_Temp.:97.25, PWM:33
23:25:57.413 -> Water_Temp.:97.25, PWM:33
23:25:58.178 -> Water_Temp.:97.25, PWM:33
23:25:58.979 -> Water_Temp.:98.94, PWM:33
23:25:59.742 -> Water_Temp.:100.96, PWM:33
23:26:00.537 -> Water_Temp.:98.60, PWM:33
23:26:01.297 -> Water_Temp.:98.15, PWM:33
23:26:02.091 -> Water_Temp.:98.04, PWM:33
23:26:02.857 -> Water_Temp.:97.81, PWM:33
23:26:05.979 -> Water_Temp.:97.59, PWM:33
23:26:06.742 -> Water_Temp.:97.59, PWM:33
23:26:07.540 -> Water_Temp.:97.59, PWM:33
23:26:08.302 -> Water_Temp.:97.59, PWM:33
23:26:09.069 -> Water_Temp.:97.59, PWM:33
23:26:12.190 -> Water_Temp.:97.59, PWM:33
23:26:13.717 -> Water_Temp.:97.47, PWM:33
23:26:14.513 -> Water_Temp.:97.47, PWM:33
23:26:15.275 -> Water_Temp.:97.47, PWM:33
23:26:16.073 -> Water_Temp.:97.47, PWM:33
23:26:16.836 -> Water_Temp.:97.47, PWM:33
23:26:17.630 -> Water_Temp.:97.47, PWM:33
23:26:18.392 -> Water_Temp.:97.47, PWM:33
23:26:19.955 -> Water_Temp.:97.47, PWM:33
23:26:20.745 -> Water_Temp.:97.47, PWM:33
23:26:21.535 -> Water_Temp.:97.47, PWM:33
23:26:22.295 -> Water_Temp.:97.47, PWM:33
23:26:23.093 -> Water_Temp.:97.47, PWM:33
23:26:24.654 -> Water_Temp.:97.47, PWM:33
23:26:27.694 -> Water_Temp.:97.59, PWM:33
23:26:29.250 -> Water_Temp.:97.59, PWM:33
23:26:30.042 -> Water_Temp.:97.59, PWM:33
23:26:30.801 -> Water_Temp.:97.59, PWM:33
23:26:31.598 -> Water_Temp.:97.47, PWM:33
23:26:32.358 -> Water_Temp.:97.47, PWM:33
23:26:33.156 -> Water_Temp.:97.47, PWM:33
23:26:33.914 -> Water_Temp.:97.47, PWM:33
23:26:36.988 -> Water_Temp.:97.47, PWM:33
23:26:37.745 -> Water_Temp.:97.47, PWM:33
23:26:38.541 -> Water_Temp.:97.47, PWM:33
23:26:39.304 -> Water_Temp.:97.47, PWM:33
23:26:40.864 -> Water_Temp.:97.47, PWM:33
23:26:41.658 -> Water_Temp.:97.59, PWM:33
23:26:43.183 -> Water_Temp.:97.47, PWM:33
23:26:44.744 -> Water_Temp.:97.47, PWM:33
23:26:48.589 -> Water_Temp.:97.47, PWM:33
23:26:50.150 -> Water_Temp.:97.47, PWM:33
23:26:50.915 -> Water_Temp.:97.47, PWM:33
23:26:51.710 -> Water_Temp.:97.47, PWM:33
23:26:52.473 -> Water_Temp.:97.47, PWM:33
23:26:54.034 -> Water_Temp.:97.47, PWM:33
23:26:54.794 -> Water_Temp.:97.47, PWM:33
23:26:55.594 -> Water_Temp.:97.47, PWM:33
23:26:56.355 -> Water_Temp.:97.47, PWM:33
23:26:58.677 -> Water_Temp.:97.59, PWM:33
23:26:59.473 -> Water_Temp.:97.59, PWM:33
23:27:00.237 -> Water_Temp.:97.59, PWM:33
23:27:01.033 -> Water_Temp.:97.59, PWM:33
23:27:01.796 -> Water_Temp.:97.59, PWM:33
23:27:02.591 -> Water_Temp.:97.59, PWM:33
23:27:04.115 -> Water_Temp.:97.59, PWM:33
23:27:04.915 -> Water_Temp.:97.59, PWM:33
23:27:05.676 -> Water_Temp.:97.47, PWM:33
23:27:06.474 -> Water_Temp.:97.47, PWM:33
23:27:07.236 -> Water_Temp.:97.47, PWM:33
23:27:08.801 -> Water_Temp.:97.47, PWM:33
23:27:09.565 -> Water_Temp.:97.47, PWM:33
23:27:10.365 -> Water_Temp.:97.47, PWM:33
23:27:11.126 -> Water_Temp.:97.47, PWM:33
23:27:11.924 -> Water_Temp.:97.47, PWM:33
23:27:15.016 -> Water_Temp.:97.59, PWM:33
23:27:16.575 -> Water_Temp.:97.47, PWM:33
23:27:18.889 -> Water_Temp.:97.59, PWM:33
23:27:19.650 -> Water_Temp.:97.59, PWM:33
23:27:21.999 -> Water_Temp.:97.59, PWM:33
23:27:22.761 -> Water_Temp.:97.59, PWM:33
23:27:23.554 -> Water_Temp.:97.59, PWM:33
23:27:24.314 -> Water_Temp.:97.59, PWM:33
23:27:25.867 -> Water_Temp.:97.59, PWM:33
23:27:26.660 -> Water_Temp.:97.59, PWM:33
23:27:27.416 -> Water_Temp.:97.59, PWM:33
23:27:28.210 -> Water_Temp.:97.70, PWM:33
23:27:29.766 -> Water_Temp.:97.70, PWM:33
23:27:30.529 -> Water_Temp.:97.70, PWM:33
23:27:31.324 -> Water_Temp.:97.70, PWM:33
23:27:32.089 -> Water_Temp.:97.70, PWM:33
23:27:32.882 -> Water_Temp.:97.70, PWM:33
23:27:35.957 -> Water_Temp.:97.70, PWM:33
23:27:36.752 -> Water_Temp.:97.70, PWM:33
23:27:37.516 -> Water_Temp.:97.70, PWM:33
23:27:39.872 -> Water_Temp.:97.70, PWM:33
23:27:40.632 -> Water_Temp.:97.70, PWM:33
23:27:42.192 -> Water_Temp.:97.70, PWM:33
23:27:42.957 -> Water_Temp.:97.70, PWM:33
23:27:43.755 -> Water_Temp.:97.70, PWM:33
23:27:44.515 -> Water_Temp.:97.70, PWM:33
23:27:45.308 -> Water_Temp.:97.70, PWM:33
23:27:46.068 -> Water_Temp.:97.70, PWM:33
23:27:46.864 -> Water_Temp.:97.70, PWM:33
23:27:47.660 -> Water_Temp.:97.70, PWM:33
23:27:48.421 -> Water_Temp.:97.70, PWM:33
23:27:50.774 -> Water_Temp.:97.70, PWM:33
23:27:51.533 -> Water_Temp.:97.70, PWM:33
23:27:52.329 -> Water_Temp.:97.70, PWM:33
23:27:53.088 -> Water_Temp.:97.70, PWM:33
23:27:53.884 -> Water_Temp.:97.70, PWM:33
23:27:54.648 -> Water_Temp.:97.70, PWM:33
23:27:55.444 -> Water_Temp.:97.70, PWM:33
23:27:57.763 -> Water_Temp.:97.70, PWM:33
23:27:59.319 -> Water_Temp.:97.70, PWM:33
23:28:00.877 -> Water_Temp.:97.70, PWM:33
23:28:02.399 -> Water_Temp.:97.70, PWM:33
23:28:03.199 -> Water_Temp.:97.70, PWM:33
23:28:07.795 -> Water_Temp.:97.70, PWM:33
23:28:09.357 -> Water_Temp.:97.59, PWM:33
23:28:10.883 -> Water_Temp.:97.59, PWM:33
23:28:13.194 -> Water_Temp.:97.59, PWM:33
23:28:14.751 -> Water_Temp.:97.59, PWM:33
23:28:15.546 -> Water_Temp.:97.59, PWM:33
23:28:16.307 -> Water_Temp.:97.70, PWM:33
23:28:17.101 -> Water_Temp.:97.70, PWM:33
23:28:17.864 -> Water_Temp.:97.70, PWM:33
23:28:19.423 -> Water_Temp.:97.59, PWM:33
23:28:20.181 -> Water_Temp.:97.59, PWM:33
23:28:23.296 -> Water_Temp.:97.59, PWM:33
23:28:24.061 -> Water_Temp.:97.59, PWM:33
23:28:24.859 -> Water_Temp.:97.59, PWM:33
23:28:25.620 -> Water_Temp.:97.59, PWM:33
23:28:27.180 -> Water_Temp.:97.59, PWM:33
23:28:27.974 -> Water_Temp.:97.59, PWM:33
23:28:28.736 -> Water_Temp.:97.70, PWM:33
23:28:29.501 -> Water_Temp.:97.70, PWM:33
23:28:32.621 -> Water_Temp.:97.70, PWM:33
23:28:35.708 -> Water_Temp.:97.70, PWM:33
23:28:37.271 -> Water_Temp.:97.70, PWM:33
23:28:38.068 -> Water_Temp.:97.70, PWM:33
23:28:38.830 -> Water_Temp.:97.70, PWM:33
23:28:39.626 -> Water_Temp.:97.70, PWM:33
23:28:40.388 -> Water_Temp.:97.70, PWM:33
23:28:41.186 -> Water_Temp.:97.70, PWM:33
23:28:41.948 -> Water_Temp.:97.70, PWM:33
23:28:42.745 -> Water_Temp.:97.70, PWM:33
23:28:43.507 -> Water_Temp.:97.70, PWM:33
23:28:44.304 -> Water_Temp.:97.70, PWM:33
23:28:46.616 -> Water_Temp.:97.70, PWM:33
23:28:47.410 -> Water_Temp.:97.70, PWM:33
23:28:48.171 -> Water_Temp.:97.70, PWM:33
23:28:48.966 -> Water_Temp.:97.70, PWM:33
23:28:49.728 -> Water_Temp.:97.70, PWM:33
23:28:50.530 -> Water_Temp.:97.70, PWM:33
23:28:51.292 -> Water_Temp.:97.70, PWM:33
23:28:55.134 -> Water_Temp.:97.59, PWM:33
23:28:55.897 -> Water_Temp.:97.59, PWM:33
23:28:56.693 -> Water_Temp.:97.59, PWM:33
23:28:57.486 -> Water_Temp.:97.59, PWM:33
23:28:58.244 -> Water_Temp.:97.59, PWM:33
23:28:59.771 -> Water_Temp.:97.59, PWM:33
23:29:02.126 -> Water_Temp.:97.59, PWM:33
23:29:02.888 -> Water_Temp.:97.59, PWM:33
23:29:03.684 -> Water_Temp.:97.59, PWM:33
23:29:04.448 -> Water_Temp.:97.59, PWM:33
23:29:06.007 -> Water_Temp.:97.70, PWM:33
23:29:07.532 -> Water_Temp.:97.70, PWM:33
23:29:08.331 -> Water_Temp.:97.70, PWM:33
23:29:09.095 -> Water_Temp.:97.70, PWM:33
23:29:09.888 -> Water_Temp.:97.70, PWM:33
23:29:11.412 -> Water_Temp.:97.70, PWM:33
23:29:12.210 -> Water_Temp.:97.59, PWM:33
23:29:12.973 -> Water_Temp.:97.70, PWM:33
23:29:13.771 -> Water_Temp.:97.59, PWM:33
23:29:16.058 -> Water_Temp.:97.59, PWM:33
23:29:16.853 -> Water_Temp.:97.59, PWM:33
23:29:18.415 -> Water_Temp.:97.59, PWM:33
23:29:19.970 -> Water_Temp.:97.59, PWM:33
23:29:21.528 -> Water_Temp.:97.59, PWM:33
23:29:22.290 -> Water_Temp.:97.59, PWM:33
23:29:23.853 -> Water_Temp.:97.70, PWM:33
23:29:26.138 -> Water_Temp.:97.70, PWM:33
23:29:26.938 -> Water_Temp.:97.70, PWM:33
23:29:27.696 -> Water_Temp.:97.70, PWM:33
23:29:29.256 -> Water_Temp.:97.70, PWM:33
23:29:34.663 -> Water_Temp.:97.70, PWM:33
23:29:35.459 -> Water_Temp.:97.70, PWM:33
23:29:36.220 -> Water_Temp.:97.70, PWM:33
23:29:37.015 -> Water_Temp.:97.70, PWM:33
23:29:38.540 -> Water_Temp.:97.70, PWM:33
23:29:41.662 -> Water_Temp.:97.70, PWM:33
23:29:42.457 -> Water_Temp.:97.70, PWM:33
23:29:45.502 -> Water_Temp.:97.70, PWM:33
23:29:47.058 -> Water_Temp.:97.59, PWM:33
23:29:47.818 -> Water_Temp.:97.59, PWM:33
23:29:48.618 -> Water_Temp.:97.59, PWM:33
23:29:49.382 -> Water_Temp.:97.59, PWM:33
23:29:50.174 -> Water_Temp.:97.59, PWM:33
23:29:51.697 -> Water_Temp.:97.59, PWM:33
23:29:52.491 -> Water_Temp.:97.59, PWM:33
23:29:54.045 -> Water_Temp.:97.59, PWM:33
23:29:54.838 -> Water_Temp.:97.59, PWM:33
23:29:55.598 -> Water_Temp.:97.59, PWM:33
23:29:57.156 -> Water_Temp.:97.59, PWM:33
23:30:01.032 -> Water_Temp.:97.70, PWM:33
23:30:01.792 -> Water_Temp.:97.70, PWM:33
23:30:02.588 -> Water_Temp.:97.70, PWM:33
23:30:03.349 -> Water_Temp.:97.70, PWM:33
23:30:04.145 -> Water_Temp.:97.70, PWM:33
23:30:05.701 -> Water_Temp.:97.70, PWM:33
23:30:06.460 -> Water_Temp.:97.70, PWM:33
23:30:07.251 -> Water_Temp.:97.70, PWM:33
23:30:08.014 -> Water_Temp.:97.70, PWM:33
23:30:08.813 -> Water_Temp.:97.70, PWM:33
23:30:09.576 -> Water_Temp.:97.70, PWM:33
23:30:11.886 -> Water_Temp.:97.70, PWM:33
23:30:13.444 -> Water_Temp.:97.70, PWM:33
23:30:14.203 -> Water_Temp.:97.70, PWM:33
23:30:15.760 -> Water_Temp.:97.70, PWM:33
23:30:16.522 -> Water_Temp.:97.70, PWM:33
23:30:17.319 -> Water_Temp.:97.70, PWM:33
23:30:18.081 -> Water_Temp.:97.70, PWM:33
23:30:18.879 -> Water_Temp.:97.70, PWM:33
23:30:19.676 -> Water_Temp.:97.70, PWM:33
23:30:21.231 -> Water_Temp.:97.70, PWM:33
23:30:23.552 -> Water_Temp.:97.59, PWM:33
23:30:25.859 -> Water_Temp.:97.59, PWM:33
23:30:28.961 -> Water_Temp.:97.59, PWM:33
23:30:31.271 -> Water_Temp.:97.59, PWM:33
23:30:32.063 -> Water_Temp.:97.59, PWM:33
23:30:32.829 -> Water_Temp.:97.59, PWM:33
23:30:33.624 -> Water_Temp.:97.59, PWM:33
23:30:34.386 -> Water_Temp.:97.59, PWM:33
23:30:36.727 -> Water_Temp.:97.70, PWM:33
23:30:37.520 -> Water_Temp.:97.59, PWM:33
23:30:38.281 -> Water_Temp.:97.70, PWM:33
23:30:39.829 -> Water_Temp.:97.70, PWM:33
23:30:40.615 -> Water_Temp.:97.70, PWM:33
23:30:41.403 -> Water_Temp.:97.59, PWM:33
23:30:42.157 -> Water_Temp.:97.59, PWM:33
23:30:42.957 -> Water_Temp.:97.59, PWM:33
23:30:44.518 -> Water_Temp.:97.70, PWM:33
23:30:45.279 -> Water_Temp.:97.70, PWM:33
23:30:46.072 -> Water_Temp.:97.70, PWM:33
23:30:48.387 -> Water_Temp.:97.70, PWM:33
23:30:51.508 -> Water_Temp.:97.70, PWM:33
23:30:53.026 -> Water_Temp.:97.70, PWM:33
23:30:53.827 -> Water_Temp.:97.70, PWM:33
23:30:54.591 -> Water_Temp.:97.70, PWM:33
23:30:55.387 -> Water_Temp.:97.59, PWM:33
23:30:56.153 -> Water_Temp.:97.59, PWM:33
23:30:58.474 -> Water_Temp.:97.59, PWM:33
23:31:00.762 -> Water_Temp.:97.59, PWM:33
23:31:02.318 -> Water_Temp.:97.59, PWM:33
23:31:03.112 -> Water_Temp.:97.59, PWM:33
23:31:03.870 -> Water_Temp.:97.59, PWM:33
23:31:04.666 -> Water_Temp.:97.59, PWM:33
23:31:06.220 -> Water_Temp.:97.59, PWM:33
23:31:07.019 -> Water_Temp.:97.59, PWM:33
23:31:07.783 -> Water_Temp.:97.59, PWM:33
23:31:09.315 -> Water_Temp.:97.59, PWM:33
23:31:10.871 -> Water_Temp.:97.59, PWM:33
23:31:11.632 -> Water_Temp.:97.59, PWM:33
23:31:13.197 -> Water_Temp.:97.70, PWM:33
23:31:13.993 -> Water_Temp.:97.70, PWM:33
23:31:14.751 -> Water_Temp.:97.81, PWM:33
23:31:15.545 -> Water_Temp.:97.81, PWM:33
23:31:16.304 -> Water_Temp.:97.81, PWM:33
23:31:17.095 -> Water_Temp.:97.81, PWM:33
23:31:18.650 -> Water_Temp.:97.92, PWM:33
23:31:19.411 -> Water_Temp.:97.92, PWM:33
23:31:20.209 -> Water_Temp.:97.81, PWM:33
23:31:20.971 -> Water_Temp.:99.50, PWM:33
23:31:22.521 -> Water_Temp.:99.16, PWM:33
23:31:23.285 -> Water_Temp.:98.71, PWM:33
23:31:24.081 -> Water_Temp.:98.49, PWM:33
23:31:25.607 -> Water_Temp.:98.26, PWM:33
23:31:26.402 -> Water_Temp.:98.15, PWM:33
23:31:27.164 -> Water_Temp.:98.15, PWM:33
23:31:29.482 -> Water_Temp.:97.92, PWM:33
23:31:30.243 -> Water_Temp.:97.92, PWM:33
23:31:31.798 -> Water_Temp.:97.81, PWM:33
23:31:32.593 -> Water_Temp.:97.81, PWM:33
23:31:33.356 -> Water_Temp.:97.81, PWM:33
23:31:34.151 -> Water_Temp.:97.70, PWM:33
23:31:36.474 -> Water_Temp.:97.70, PWM:33
23:31:37.271 -> Water_Temp.:97.70, PWM:33
23:31:40.361 -> Water_Temp.:97.59, PWM:33
23:31:41.124 -> Water_Temp.:97.59, PWM:33
23:31:41.916 -> Water_Temp.:97.59, PWM:33
23:31:42.679 -> Water_Temp.:97.59, PWM:33
23:31:44.239 -> Water_Temp.:97.59, PWM:33
23:31:45.037 -> Water_Temp.:97.59, PWM:33
23:31:45.797 -> Water_Temp.:97.70, PWM:33
23:31:46.592 -> Water_Temp.:97.70, PWM:33
23:31:48.154 -> Water_Temp.:97.70, PWM:33
23:31:49.711 -> Water_Temp.:97.81, PWM:33
23:31:53.593 -> Water_Temp.:97.81, PWM:33
23:31:54.389 -> Water_Temp.:97.81, PWM:33
23:31:55.149 -> Water_Temp.:97.81, PWM:33
23:31:55.949 -> Water_Temp.:97.81, PWM:33
23:31:56.711 -> Water_Temp.:97.81, PWM:33
23:32:01.353 -> Water_Temp.:97.81, PWM:33
23:32:02.942 -> Water_Temp.:97.81, PWM:33
23:32:03.704 -> Water_Temp.:97.81, PWM:33
23:32:04.469 -> Water_Temp.:97.81, PWM:33
23:32:05.260 -> Water_Temp.:97.81, PWM:33
23:32:07.582 -> Water_Temp.:97.70, PWM:33
23:32:08.377 -> Water_Temp.:97.70, PWM:33
23:32:09.142 -> Water_Temp.:97.70, PWM:33
23:32:09.933 -> Water_Temp.:97.70, PWM:33
23:32:12.253 -> Water_Temp.:97.70, PWM:33
23:32:13.014 -> Water_Temp.:97.70, PWM:33
23:32:14.568 -> Water_Temp.:97.70, PWM:33
23:32:15.365 -> Water_Temp.:97.70, PWM:33
23:32:16.126 -> Water_Temp.:97.70, PWM:33
23:32:17.685 -> Water_Temp.:97.70, PWM:33
23:32:19.205 -> Water_Temp.:97.70, PWM:33
23:32:22.314 -> Water_Temp.:97.81, PWM:33
23:32:23.076 -> Water_Temp.:97.81, PWM:33
23:32:23.870 -> Water_Temp.:97.81, PWM:33
23:32:24.634 -> Water_Temp.:97.81, PWM:33
23:32:25.429 -> Water_Temp.:97.81, PWM:33
23:32:26.955 -> Water_Temp.:97.81, PWM:33
23:32:27.720 -> Water_Temp.:97.81, PWM:33
23:32:29.279 -> Water_Temp.:97.81, PWM:33
23:32:30.043 -> Water_Temp.:97.81, PWM:33
23:32:31.602 -> Water_Temp.:97.81, PWM:33
23:32:32.365 -> Water_Temp.:97.81, PWM:33
23:32:33.163 -> Water_Temp.:97.81, PWM:33
23:32:33.928 -> Water_Temp.:97.81, PWM:33
23:32:35.489 -> Water_Temp.:97.81, PWM:33
23:32:37.773 -> Water_Temp.:97.92, PWM:33
23:32:39.329 -> Water_Temp.:97.92, PWM:33
23:32:40.123 -> Water_Temp.:98.04, PWM:33
23:32:40.923 -> Water_Temp.:97.92, PWM:33
23:32:41.682 -> Water_Temp.:98.04, PWM:33
23:32:42.474 -> Water_Temp.:97.92, PWM:33
23:32:43.234 -> Water_Temp.:97.92, PWM:33
23:32:44.030 -> Water_Temp.:97.92, PWM:33
23:32:44.790 -> Water_Temp.:97.92, PWM:33
23:32:45.588 -> Water_Temp.:97.92, PWM:33
23:32:46.350 -> Water_Temp.:97.92, PWM:33
23:32:47.147 -> Water_Temp.:97.92, PWM:33
23:32:47.909 -> Water_Temp.:97.92, PWM:33
23:32:48.708 -> Water_Temp.:97.92, PWM:33
23:32:49.501 -> Water_Temp.:97.92, PWM:33
23:32:50.259 -> Water_Temp.:97.92, PWM:33
23:32:51.056 -> Water_Temp.:97.92, PWM:33
23:32:51.814 -> Water_Temp.:97.92, PWM:33
23:32:54.928 -> Water_Temp.:97.92, PWM:33
23:32:55.723 -> Water_Temp.:97.92, PWM:33
23:32:57.278 -> Water_Temp.:97.92, PWM:33
23:32:58.075 -> Water_Temp.:97.92, PWM:33
23:32:58.842 -> Water_Temp.:97.92, PWM:33
23:33:01.162 -> Water_Temp.:97.92, PWM:33
23:33:02.686 -> Water_Temp.:97.92, PWM:33
23:33:03.485 -> Water_Temp.:97.92, PWM:33
23:33:05.773 -> Water_Temp.:98.04, PWM:33
23:33:07.325 -> Water_Temp.:98.04, PWM:33
23:33:11.207 -> Water_Temp.:97.92, PWM:33
23:33:12.728 -> Water_Temp.:97.92, PWM:33
23:33:13.525 -> Water_Temp.:97.92, PWM:33
23:33:14.288 -> Water_Temp.:97.92, PWM:33
23:33:15.083 -> Water_Temp.:97.92, PWM:33
23:33:18.124 -> Water_Temp.:98.04, PWM:33
23:33:19.709 -> Water_Temp.:98.04, PWM:33
23:33:21.270 -> Water_Temp.:98.04, PWM:33
23:33:22.030 -> Water_Temp.:98.04, PWM:33
23:33:22.828 -> Water_Temp.:98.04, PWM:33
23:33:24.360 -> Water_Temp.:98.04, PWM:33
23:33:25.125 -> Water_Temp.:98.04, PWM:33
23:33:26.686 -> Water_Temp.:98.04, PWM:33
23:33:27.447 -> Water_Temp.:98.04, PWM:33
23:33:28.244 -> Water_Temp.:98.04, PWM:33
23:33:29.804 -> Water_Temp.:97.92, PWM:33
23:33:30.568 -> Water_Temp.:97.92, PWM:33
23:33:31.364 -> Water_Temp.:97.92, PWM:33
23:33:32.925 -> Water_Temp.:97.92, PWM:33
23:33:33.687 -> Water_Temp.:97.92, PWM:33
23:33:34.452 -> Water_Temp.:98.04, PWM:33
23:33:35.245 -> Water_Temp.:98.04, PWM:33
23:33:37.584 -> Water_Temp.:98.04, PWM:33
23:33:38.342 -> Water_Temp.:98.04, PWM:33
23:33:39.140 -> Water_Temp.:98.04, PWM:33
23:33:39.905 -> Water_Temp.:98.04, PWM:33
23:33:41.462 -> Water_Temp.:98.04, PWM:33
23:33:43.780 -> Water_Temp.:98.04, PWM:33
23:33:45.311 -> Water_Temp.:97.92, PWM:33
23:33:46.109 -> Water_Temp.:97.92, PWM:33
23:33:46.869 -> Water_Temp.:97.92, PWM:33
23:33:47.662 -> Water_Temp.:97.92, PWM:33
23:33:48.424 -> Water_Temp.:98.04, PWM:33
23:33:49.225 -> Water_Temp.:98.04, PWM:33
23:33:51.544 -> Water_Temp.:97.92, PWM:33
23:33:52.337 -> Water_Temp.:97.92, PWM:33
23:33:53.131 -> Water_Temp.:97.92, PWM:33
23:33:53.891 -> Water_Temp.:97.92, PWM:33
23:33:54.688 -> Water_Temp.:97.92, PWM:33
23:33:57.760 -> Water_Temp.:98.04, PWM:33
23:33:58.552 -> Water_Temp.:98.04, PWM:33
23:33:59.317 -> Water_Temp.:98.04, PWM:33
23:34:00.875 -> Water_Temp.:98.04, PWM:33
23:34:02.399 -> Water_Temp.:98.04, PWM:33
23:34:03.956 -> Water_Temp.:98.04, PWM:33
23:34:06.275 -> Water_Temp.:98.04, PWM:33
23:34:07.037 -> Water_Temp.:98.04, PWM:33
23:34:07.834 -> Water_Temp.:98.04, PWM:33
23:34:09.359 -> Water_Temp.:98.04, PWM:33
23:34:10.922 -> Water_Temp.:98.15, PWM:33
23:34:11.718 -> Water_Temp.:98.15, PWM:33
23:34:12.509 -> Water_Temp.:98.15, PWM:33
23:34:13.271 -> Water_Temp.:98.15, PWM:33
23:34:14.068 -> Water_Temp.:98.15, PWM:33
23:34:16.354 -> Water_Temp.:98.04, PWM:33
23:34:17.152 -> Water_Temp.:98.04, PWM:33
23:34:18.715 -> Water_Temp.:98.04, PWM:33
23:34:19.480 -> Water_Temp.:98.04, PWM:33
23:34:20.275 -> Water_Temp.:98.04, PWM:33
23:34:22.565 -> Water_Temp.:98.04, PWM:33
23:34:23.361 -> Water_Temp.:98.04, PWM:33
23:34:24.885 -> Water_Temp.:98.04, PWM:33
23:34:25.681 -> Water_Temp.:98.04, PWM:33
23:34:26.441 -> Water_Temp.:98.04, PWM:33
23:34:28.792 -> Water_Temp.:98.04, PWM:33
23:34:29.592 -> Water_Temp.:98.15, PWM:33
23:34:30.357 -> Water_Temp.:98.15, PWM:33
23:34:31.121 -> Water_Temp.:98.15, PWM:33
23:34:31.918 -> Water_Temp.:98.15, PWM:33
23:34:33.440 -> Water_Temp.:98.15, PWM:33
23:34:34.997 -> Water_Temp.:98.15, PWM:33
23:34:36.556 -> Water_Temp.:98.15, PWM:33
23:34:37.315 -> Water_Temp.:98.15, PWM:33
23:34:38.109 -> Water_Temp.:98.15, PWM:33
23:34:38.875 -> Water_Temp.:98.15, PWM:33
23:34:40.432 -> Water_Temp.:98.15, PWM:33
23:34:41.228 -> Water_Temp.:98.15, PWM:33
23:34:42.756 -> Water_Temp.:98.26, PWM:33
23:34:44.314 -> Water_Temp.:98.26, PWM:33
23:34:45.108 -> Water_Temp.:98.26, PWM:33
23:34:45.904 -> Water_Temp.:98.26, PWM:33
23:34:46.667 -> Water_Temp.:98.26, PWM:33
23:34:47.429 -> Water_Temp.:98.26, PWM:33
23:34:49.785 -> Water_Temp.:98.26, PWM:33
23:34:50.546 -> Water_Temp.:98.26, PWM:33
23:34:52.105 -> Water_Temp.:98.26, PWM:33
23:34:52.901 -> Water_Temp.:98.26, PWM:33
23:34:53.660 -> Water_Temp.:98.26, PWM:33
23:34:55.215 -> Water_Temp.:98.26, PWM:33
23:34:56.775 -> Water_Temp.:98.26, PWM:33
23:34:57.536 -> Water_Temp.:98.15, PWM:33
23:35:00.613 -> Water_Temp.:98.15, PWM:33
23:35:01.412 -> Water_Temp.:98.15, PWM:33
23:35:02.205 -> Water_Temp.:98.15, PWM:33
23:35:02.968 -> Water_Temp.:98.15, PWM:33
23:35:03.763 -> Water_Temp.:98.15, PWM:33
23:35:04.521 -> Water_Temp.:98.15, PWM:33
23:35:05.319 -> Water_Temp.:98.15, PWM:33
23:35:06.086 -> Water_Temp.:98.15, PWM:33
23:35:06.882 -> Water_Temp.:98.15, PWM:33
23:35:08.445 -> Water_Temp.:98.15, PWM:33
23:35:09.209 -> Water_Temp.:98.15, PWM:33
23:35:10.004 -> Water_Temp.:98.26, PWM:33
23:35:10.764 -> Water_Temp.:98.26, PWM:33
23:35:12.287 -> Water_Temp.:98.26, PWM:33
23:35:13.846 -> Water_Temp.:98.26, PWM:33
23:35:14.644 -> Water_Temp.:98.26, PWM:33
23:35:15.401 -> Water_Temp.:98.15, PWM:33
23:35:19.967 -> Water_Temp.:98.15, PWM:33
23:35:20.761 -> Water_Temp.:98.15, PWM:33
23:35:21.526 -> Water_Temp.:98.15, PWM:33
23:35:23.079 -> Water_Temp.:98.15, PWM:33
23:35:23.839 -> Water_Temp.:98.15, PWM:33
23:35:24.639 -> Water_Temp.:98.15, PWM:33
23:35:25.430 -> Water_Temp.:98.15, PWM:33
23:35:27.741 -> Water_Temp.:98.15, PWM:33
23:35:29.294 -> Water_Temp.:98.04, PWM:33
23:35:30.092 -> Water_Temp.:98.04, PWM:33
23:35:30.852 -> Water_Temp.:98.04, PWM:33
23:35:32.410 -> Water_Temp.:98.04, PWM:33
23:35:33.174 -> Water_Temp.:98.04, PWM:33
23:35:34.734 -> Water_Temp.:98.15, PWM:33
23:35:35.525 -> Water_Temp.:99.72, PWM:33
23:35:37.845 -> Water_Temp.:99.50, PWM:33
23:35:38.606 -> Water_Temp.:98.82, PWM:33
23:35:39.402 -> Water_Temp.:98.71, PWM:33
23:35:40.167 -> Water_Temp.:98.60, PWM:33
23:35:40.967 -> Water_Temp.:98.60, PWM:33
23:35:41.731 -> Water_Temp.:98.49, PWM:33
23:35:42.529 -> Water_Temp.:98.49, PWM:33
23:35:44.054 -> Water_Temp.:98.49, PWM:33
23:35:44.850 -> Water_Temp.:98.49, PWM:33
23:35:45.615 -> Water_Temp.:98.37, PWM:33
23:35:46.410 -> Water_Temp.:98.37, PWM:33
23:35:47.968 -> Water_Temp.:98.37, PWM:33
23:35:49.488 -> Water_Temp.:98.37, PWM:33
23:35:50.277 -> Water_Temp.:98.37, PWM:33
23:35:51.831 -> Water_Temp.:98.26, PWM:33
23:35:54.974 -> Water_Temp.:98.26, PWM:33
23:35:55.736 -> Water_Temp.:98.26, PWM:33
23:35:57.295 -> Water_Temp.:98.26, PWM:33
23:35:58.089 -> Water_Temp.:98.26, PWM:33
23:35:58.850 -> Water_Temp.:99.84, PWM:33
23:36:00.374 -> Water_Temp.:99.50, PWM:33
23:36:01.168 -> Water_Temp.:99.16, PWM:33
23:36:02.725 -> Water_Temp.:98.71, PWM:33
23:36:03.486 -> Water_Temp.:98.60, PWM:33
23:36:04.287 -> Water_Temp.:98.60, PWM:33
23:36:05.046 -> Water_Temp.:98.60, PWM:33
23:36:05.844 -> Water_Temp.:98.49, PWM:33
23:36:06.601 -> Water_Temp.:98.49, PWM:33
23:36:08.161 -> Water_Temp.:98.37, PWM:33
23:36:08.960 -> Water_Temp.:98.37, PWM:33
23:36:09.718 -> Water_Temp.:98.37, PWM:33
23:36:10.509 -> Water_Temp.:98.37, PWM:33
23:36:11.302 -> Water_Temp.:98.37, PWM:33
23:36:12.066 -> Water_Temp.:98.26, PWM:33
23:36:13.616 -> Water_Temp.:98.26, PWM:33
23:36:16.698 -> Water_Temp.:98.26, PWM:33
23:36:17.496 -> Water_Temp.:98.26, PWM:33
23:36:19.022 -> Water_Temp.:98.26, PWM:33
23:36:19.822 -> Water_Temp.:98.26, PWM:33
23:36:21.382 -> Water_Temp.:98.26, PWM:33
23:36:22.141 -> Water_Temp.:98.26, PWM:33
23:36:22.942 -> Water_Temp.:98.26, PWM:33
23:36:23.708 -> Water_Temp.:98.26, PWM:33
23:36:24.501 -> Water_Temp.:98.26, PWM:33
23:36:25.263 -> Water_Temp.:98.15, PWM:33
23:36:26.062 -> Water_Temp.:98.15, PWM:33
23:36:26.825 -> Water_Temp.:98.15, PWM:33
23:36:27.591 -> Water_Temp.:98.15, PWM:33
23:36:28.386 -> Water_Temp.:98.15, PWM:33
23:36:30.673 -> Water_Temp.:98.15, PWM:33
23:36:31.467 -> Water_Temp.:98.15, PWM:33
23:36:33.024 -> Water_Temp.:98.15, PWM:33
23:36:33.821 -> Water_Temp.:98.15, PWM:33
23:36:35.376 -> Water_Temp.:98.15, PWM:33
23:36:36.134 -> Water_Temp.:98.15, PWM:33
23:36:38.486 -> Water_Temp.:98.15, PWM:33
23:36:39.248 -> Water_Temp.:98.15, PWM:33
23:36:40.009 -> Water_Temp.:98.15, PWM:33
23:36:41.599 -> Water_Temp.:98.15, PWM:33
23:36:42.359 -> Water_Temp.:98.15, PWM:33
23:36:43.913 -> Water_Temp.:98.15, PWM:33
23:36:44.674 -> Water_Temp.:98.15, PWM:33
23:36:46.233 -> Water_Temp.:98.15, PWM:33
23:36:47.027 -> Water_Temp.:98.15, PWM:33
23:36:47.788 -> Water_Temp.:98.15, PWM:33
23:36:48.584 -> Water_Temp.:98.15, PWM:33
23:36:50.898 -> Water_Temp.:98.15, PWM:33
23:36:51.689 -> Water_Temp.:98.15, PWM:33
23:36:53.212 -> Water_Temp.:98.15, PWM:33
23:36:54.766 -> Water_Temp.:98.15, PWM:33
23:36:56.289 -> Water_Temp.:98.15, PWM:33
23:36:57.085 -> Water_Temp.:98.15, PWM:33
23:36:57.846 -> Water_Temp.:98.15, PWM:33
23:36:58.643 -> Water_Temp.:98.15, PWM:33
23:36:59.438 -> Water_Temp.:98.26, PWM:33
23:37:00.953 -> Water_Temp.:98.26, PWM:33
23:37:02.505 -> Water_Temp.:98.26, PWM:33
23:37:03.304 -> Water_Temp.:98.15, PWM:33
23:37:04.063 -> Water_Temp.:98.15, PWM:33
23:37:04.858 -> Water_Temp.:98.15, PWM:33
23:37:05.616 -> Water_Temp.:98.15, PWM:33
23:37:07.173 -> Water_Temp.:98.15, PWM:33
23:37:07.973 -> Water_Temp.:98.15, PWM:33
23:37:08.733 -> Water_Temp.:98.15, PWM:33
23:37:10.290 -> Water_Temp.:98.26, PWM:33
23:37:11.875 -> Water_Temp.:98.26, PWM:33
23:37:12.635 -> Water_Temp.:98.26, PWM:33
23:37:14.197 -> Water_Temp.:98.26, PWM:33
23:37:15.755 -> Water_Temp.:98.26, PWM:33
23:37:16.517 -> Water_Temp.:98.26, PWM:33
23:37:17.314 -> Water_Temp.:98.26, PWM:33
23:37:18.077 -> Water_Temp.:98.26, PWM:33
23:37:18.873 -> Water_Temp.:98.26, PWM:33
23:37:21.190 -> Water_Temp.:98.26, PWM:33
23:37:21.951 -> Water_Temp.:98.15, PWM:33
23:37:22.744 -> Water_Temp.:98.15, PWM:33
23:37:23.509 -> Water_Temp.:98.15, PWM:33
23:37:24.307 -> Water_Temp.:98.15, PWM:33
23:37:26.618 -> Water_Temp.:98.15, PWM:33
23:37:27.413 -> Water_Temp.:98.15, PWM:33
23:37:28.174 -> Water_Temp.:98.15, PWM:33
23:37:31.247 -> Water_Temp.:98.15, PWM:33
23:37:32.046 -> Water_Temp.:98.15, PWM:33
23:37:32.803 -> Water_Temp.:98.15, PWM:33
23:37:36.679 -> Water_Temp.:98.15, PWM:33
23:37:39.002 -> Water_Temp.:98.15, PWM:33
23:37:41.357 -> Water_Temp.:98.15, PWM:33
23:37:42.911 -> Water_Temp.:98.15, PWM:33
23:37:43.673 -> Water_Temp.:98.15, PWM:33
23:37:44.469 -> Water_Temp.:98.15, PWM:33
23:37:46.025 -> Water_Temp.:98.15, PWM:33
23:37:46.783 -> Water_Temp.:98.15, PWM:33
23:37:47.577 -> Water_Temp.:98.15, PWM:33
23:37:48.341 -> Water_Temp.:98.15, PWM:33
23:37:49.897 -> Water_Temp.:98.15, PWM:33
23:37:50.690 -> Water_Temp.:98.15, PWM:33
23:37:53.042 -> Water_Temp.:98.15, PWM:33
23:37:54.562 -> Water_Temp.:98.15, PWM:33
23:37:55.358 -> Water_Temp.:98.15, PWM:33
23:37:56.115 -> Water_Temp.:98.15, PWM:33
23:37:56.909 -> Water_Temp.:98.15, PWM:33
23:37:58.465 -> Water_Temp.:98.15, PWM:33
23:37:59.264 -> Water_Temp.:98.15, PWM:33
23:38:00.027 -> Water_Temp.:98.15, PWM:33
23:38:00.820 -> Water_Temp.:98.15, PWM:33
23:38:02.336 -> Water_Temp.:98.26, PWM:33
23:38:03.129 -> Water_Temp.:98.26, PWM:33
23:38:03.924 -> Water_Temp.:98.26, PWM:33
23:38:04.689 -> Water_Temp.:98.26, PWM:33
23:38:06.247 -> Water_Temp.:98.26, PWM:33
23:38:07.007 -> Water_Temp.:98.26, PWM:33
23:38:08.565 -> Water_Temp.:98.26, PWM:33
23:38:09.364 -> Water_Temp.:98.37, PWM:33
23:38:10.124 -> Water_Temp.:98.37, PWM:33
23:38:11.681 -> Water_Temp.:98.37, PWM:33
23:38:13.999 -> Water_Temp.:98.37, PWM:33
23:38:14.796 -> Water_Temp.:98.37, PWM:33
23:38:15.559 -> Water_Temp.:98.37, PWM:33
23:38:16.359 -> Water_Temp.:98.37, PWM:33
23:38:17.148 -> Water_Temp.:98.26, PWM:33
23:38:17.913 -> Water_Temp.:98.37, PWM:33
23:38:18.710 -> Water_Temp.:98.26, PWM:33
23:38:19.473 -> Water_Temp.:98.26, PWM:33
23:38:20.269 -> Water_Temp.:98.26, PWM:33
23:38:21.030 -> Water_Temp.:98.37, PWM:33
23:38:21.825 -> Water_Temp.:98.37, PWM:33
23:38:22.585 -> Water_Temp.:98.37, PWM:33
23:38:23.381 -> Water_Temp.:98.26, PWM:33
23:38:26.426 -> Water_Temp.:98.37, PWM:33
23:38:27.219 -> Water_Temp.:98.37, PWM:33
23:38:27.979 -> Water_Temp.:98.37, PWM:33
23:38:29.542 -> Water_Temp.:98.37, PWM:33
23:38:31.861 -> Water_Temp.:98.37, PWM:33
23:38:32.622 -> Water_Temp.:98.37, PWM:33
23:38:33.416 -> Water_Temp.:98.26, PWM:33
23:38:34.181 -> Water_Temp.:98.26, PWM:33
23:38:36.504 -> Water_Temp.:98.26, PWM:33
23:38:38.062 -> Water_Temp.:98.37, PWM:33
23:38:38.827 -> Water_Temp.:98.37, PWM:33
23:38:41.151 -> Water_Temp.:98.37, PWM:33
23:38:41.949 -> Water_Temp.:98.37, PWM:33
23:38:42.715 -> Water_Temp.:98.37, PWM:33
23:38:43.476 -> Water_Temp.:98.37, PWM:33
23:38:45.833 -> Water_Temp.:98.26, PWM:33
23:38:47.395 -> Water_Temp.:98.26, PWM:33
23:38:48.156 -> Water_Temp.:98.26, PWM:33
23:38:48.918 -> Water_Temp.:98.26, PWM:33
23:38:50.514 -> Water_Temp.:98.26, PWM:33
23:38:51.277 -> Water_Temp.:98.26, PWM:33
23:38:52.069 -> Water_Temp.:98.26, PWM:33
23:38:52.828 -> Water_Temp.:98.26, PWM:33
23:38:53.622 -> Water_Temp.:98.26, PWM:33
23:38:58.195 -> Water_Temp.:98.26, PWM:33
23:38:58.994 -> Water_Temp.:98.37, PWM:33
23:38:59.790 -> Water_Temp.:98.37, PWM:33
23:39:01.310 -> Water_Temp.:98.37, PWM:33
23:39:02.106 -> Water_Temp.:98.37, PWM:33
23:39:02.868 -> Water_Temp.:98.37, PWM:33
23:39:03.665 -> Water_Temp.:98.37, PWM:33
23:39:04.430 -> Water_Temp.:98.37, PWM:33
23:39:06.752 -> Water_Temp.:98.37, PWM:33
23:39:08.312 -> Water_Temp.:98.37, PWM:33
23:39:09.077 -> Water_Temp.:98.37, PWM:33
23:39:09.875 -> Water_Temp.:98.37, PWM:33
23:39:10.641 -> Water_Temp.:98.37, PWM:33
23:39:11.405 -> Water_Temp.:98.37, PWM:33
23:39:12.205 -> Water_Temp.:98.37, PWM:33
23:39:12.971 -> Water_Temp.:98.37, PWM:33
23:39:13.772 -> Water_Temp.:98.37, PWM:33
23:39:15.295 -> Water_Temp.:98.49, PWM:33
23:39:16.086 -> Water_Temp.:98.49, PWM:33
23:39:17.607 -> Water_Temp.:98.49, PWM:33
23:39:18.402 -> Water_Temp.:98.49, PWM:33
23:39:19.167 -> Water_Temp.:98.49, PWM:33
23:39:19.964 -> Water_Temp.:98.49, PWM:33
23:39:20.721 -> Water_Temp.:98.49, PWM:33
23:39:23.038 -> Water_Temp.:98.49, PWM:33
23:39:25.356 -> Water_Temp.:98.49, PWM:33
23:39:26.155 -> Water_Temp.:98.37, PWM:33
23:39:27.677 -> Water_Temp.:98.49, PWM:33
23:39:29.232 -> Water_Temp.:98.37, PWM:33
23:39:30.027 -> Water_Temp.:98.37, PWM:33
23:39:31.586 -> Water_Temp.:98.49, PWM:33
23:39:32.348 -> Water_Temp.:98.37, PWM:33
23:39:33.143 -> Water_Temp.:98.49, PWM:33
23:39:34.672 -> Water_Temp.:98.49, PWM:33
23:39:36.229 -> Water_Temp.:98.49, PWM:33
23:39:38.513 -> Water_Temp.:98.37, PWM:33
23:39:40.071 -> Water_Temp.:98.37, PWM:33
23:39:40.866 -> Water_Temp.:98.37, PWM:33
23:39:41.661 -> Water_Temp.:98.37, PWM:33
23:39:43.987 -> Water_Temp.:98.49, PWM:33
23:39:45.538 -> Water_Temp.:98.49, PWM:33
23:39:46.303 -> Water_Temp.:98.49, PWM:33
23:39:47.861 -> Water_Temp.:98.49, PWM:33
23:39:48.656 -> Water_Temp.:98.49, PWM:33
23:39:49.420 -> Water_Temp.:98.37, PWM:33
23:39:51.008 -> Water_Temp.:98.49, PWM:33
23:39:52.562 -> Water_Temp.:98.49, PWM:33
23:39:53.321 -> Water_Temp.:98.49, PWM:33
23:39:54.879 -> Water_Temp.:98.49, PWM:33
23:39:55.639 -> Water_Temp.:98.49, PWM:33
23:39:59.511 -> Water_Temp.:98.49, PWM:33
23:40:00.270 -> Water_Temp.:98.60, PWM:33
23:40:02.619 -> Water_Temp.:98.49, PWM:33
23:40:03.381 -> Water_Temp.:98.49, PWM:33
23:40:04.178 -> Water_Temp.:98.49, PWM:33
23:40:04.937 -> Water_Temp.:98.49, PWM:33
23:40:06.498 -> Water_Temp.:98.49, PWM:33
23:40:07.255 -> Water_Temp.:98.49, PWM:33
23:40:09.575 -> Water_Temp.:98.49, PWM:33
23:40:10.366 -> Water_Temp.:98.49, PWM:33
23:40:12.673 -> Water_Temp.:98.60, PWM:33
23:40:13.431 -> Water_Temp.:98.60, PWM:33
23:40:14.226 -> Water_Temp.:98.60, PWM:33
23:40:14.988 -> Water_Temp.:98.60, PWM:33
23:40:15.782 -> Water_Temp.:98.60, PWM:33
23:40:16.580 -> Water_Temp.:98.60, PWM:33
23:40:17.343 -> Water_Temp.:98.60, PWM:33
23:40:18.136 -> Water_Temp.:98.71, PWM:33
23:40:19.699 -> Water_Temp.:98.60, PWM:33
23:40:20.459 -> Water_Temp.:98.60, PWM:33
23:40:21.254 -> Water_Temp.:98.60, PWM:33
23:40:22.014 -> Water_Temp.:98.60, PWM:33
23:40:25.093 -> Water_Temp.:98.60, PWM:33
23:40:26.652 -> Water_Temp.:98.60, PWM:33
23:40:27.414 -> Water_Temp.:98.60, PWM:33
23:40:28.208 -> Water_Temp.:98.60, PWM:33
23:40:28.975 -> Water_Temp.:98.60, PWM:33
23:40:30.530 -> Water_Temp.:98.60, PWM:33
23:40:31.327 -> Water_Temp.:98.60, PWM:33
23:40:32.090 -> Water_Temp.:98.60, PWM:33
23:40:33.646 -> Water_Temp.:98.60, PWM:33
23:40:34.409 -> Water_Temp.:98.60, PWM:33
23:40:35.965 -> Water_Temp.:98.60, PWM:33
23:40:38.317 -> Water_Temp.:98.60, PWM:33
23:40:39.084 -> Water_Temp.:98.60, PWM:33
23:40:41.398 -> Water_Temp.:98.60, PWM:33
23:40:42.950 -> Water_Temp.:98.60, PWM:33
23:40:43.711 -> Water_Temp.:98.60, PWM:33
23:40:44.506 -> Water_Temp.:98.60, PWM:33
23:40:45.300 -> Water_Temp.:98.60, PWM:33
23:40:46.060 -> Water_Temp.:98.60, PWM:33
23:40:47.615 -> Water_Temp.:98.60, PWM:33
23:40:48.378 -> Water_Temp.:98.60, PWM:33
23:40:49.174 -> Water_Temp.:98.60, PWM:33
23:40:49.937 -> Water_Temp.:98.60, PWM:33
23:40:51.496 -> Water_Temp.:98.60, PWM:33
23:40:53.051 -> Water_Temp.:98.60, PWM:33
23:40:53.811 -> Water_Temp.:98.60, PWM:33
23:40:54.607 -> Water_Temp.:98.60, PWM:33
23:40:56.167 -> Water_Temp.:98.60, PWM:33
23:40:59.221 -> Water_Temp.:98.60, PWM:33
23:41:00.020 -> Water_Temp.:98.60, PWM:33
23:41:01.546 -> Water_Temp.:98.60, PWM:33
23:41:02.339 -> Water_Temp.:98.60, PWM:33
23:41:03.900 -> Water_Temp.:98.60, PWM:33
23:41:04.663 -> Water_Temp.:98.60, PWM:33
23:41:06.219 -> Water_Temp.:98.60, PWM:33
23:41:06.978 -> Water_Temp.:98.60, PWM:33
23:41:07.772 -> Water_Temp.:98.60, PWM:33
23:41:08.565 -> Water_Temp.:98.60, PWM:33
23:41:09.328 -> Water_Temp.:98.60, PWM:33
23:41:10.882 -> Water_Temp.:98.60, PWM:33
23:41:11.676 -> Water_Temp.:98.60, PWM:33
23:41:13.226 -> Water_Temp.:98.60, PWM:33
23:41:14.020 -> Water_Temp.:98.60, PWM:33
23:41:14.781 -> Water_Temp.:98.60, PWM:33
23:41:15.576 -> Water_Temp.:98.60, PWM:33
23:41:17.100 -> Water_Temp.:98.60, PWM:33
23:41:17.895 -> Water_Temp.:98.60, PWM:33
23:41:18.657 -> Water_Temp.:98.60, PWM:33
23:41:19.453 -> Water_Temp.:98.60, PWM:33
23:41:20.214 -> Water_Temp.:98.60, PWM:33
23:41:21.010 -> Water_Temp.:98.60, PWM:33
23:41:23.329 -> Water_Temp.:98.60, PWM:33
23:41:24.125 -> Water_Temp.:98.60, PWM:33
23:41:25.650 -> Water_Temp.:98.60, PWM:33
23:41:26.445 -> Water_Temp.:98.60, PWM:33
23:41:27.210 -> Water_Temp.:98.60, PWM:33
23:41:28.005 -> Water_Temp.:98.60, PWM:33
23:41:28.799 -> Water_Temp.:98.60, PWM:33
23:41:30.356 -> Water_Temp.:98.60, PWM:33
23:41:31.116 -> Water_Temp.:98.60, PWM:33
23:41:31.911 -> Water_Temp.:98.60, PWM:33
23:41:32.673 -> Water_Temp.:98.71, PWM:33
23:41:33.467 -> Water_Temp.:98.71, PWM:33
23:41:34.229 -> Water_Temp.:98.71, PWM:33
23:41:35.784 -> Water_Temp.:98.71, PWM:33
23:41:38.101 -> Water_Temp.:98.71, PWM:33
23:41:38.865 -> Water_Temp.:98.71, PWM:33
23:41:39.657 -> Water_Temp.:98.71, PWM:33
23:41:41.182 -> Water_Temp.:98.71, PWM:33
23:41:41.978 -> Water_Temp.:98.71, PWM:33
23:41:42.740 -> Water_Temp.:98.71, PWM:33
23:41:43.536 -> Water_Temp.:98.71, PWM:33
23:41:45.061 -> Water_Temp.:98.71, PWM:33
23:41:45.856 -> Water_Temp.:98.71, PWM:33
23:41:46.618 -> Water_Temp.:98.71, PWM:33
23:41:47.412 -> Water_Temp.:98.71, PWM:33
23:41:48.207 -> Water_Temp.:98.71, PWM:33
23:41:48.970 -> Water_Temp.:98.71, PWM:33
23:41:50.526 -> Water_Temp.:98.71, PWM:33
23:41:51.325 -> Water_Temp.:98.71, PWM:33
23:41:52.089 -> Water_Temp.:98.71, PWM:33
23:41:52.883 -> Water_Temp.:98.71, PWM:33
23:41:53.645 -> Water_Temp.:98.71, PWM:33
23:41:55.204 -> Water_Temp.:98.71, PWM:33
23:41:55.997 -> Water_Temp.:98.71, PWM:33
23:41:57.521 -> Water_Temp.:98.60, PWM:33
23:41:58.286 -> Water_Temp.:98.60, PWM:33
23:41:59.084 -> Water_Temp.:98.60, PWM:33
23:42:00.643 -> Water_Temp.:98.60, PWM:33
23:42:01.405 -> Water_Temp.:98.60, PWM:33
23:42:02.201 -> Water_Temp.:98.60, PWM:33
23:42:02.963 -> Water_Temp.:98.60, PWM:33
23:42:03.758 -> Water_Temp.:98.60, PWM:33
23:42:04.521 -> Water_Temp.:98.60, PWM:33
23:42:05.320 -> Water_Temp.:98.60, PWM:33
23:42:06.083 -> Water_Temp.:98.60, PWM:33
23:42:08.399 -> Water_Temp.:98.60, PWM:33
23:42:09.959 -> Water_Temp.:98.60, PWM:33
23:42:12.282 -> Water_Temp.:98.60, PWM:33
23:42:13.835 -> Water_Temp.:98.60, PWM:33
23:42:14.600 -> Water_Temp.:98.60, PWM:33
23:42:16.917 -> Water_Temp.:98.60, PWM:33
23:42:18.471 -> Water_Temp.:98.60, PWM:33
23:42:20.797 -> Water_Temp.:98.60, PWM:33
23:42:22.355 -> Water_Temp.:98.60, PWM:33
23:42:23.118 -> Water_Temp.:98.60, PWM:33
23:42:24.673 -> Water_Temp.:98.60, PWM:33
23:42:26.196 -> Water_Temp.:98.60, PWM:33
23:42:26.991 -> Water_Temp.:98.60, PWM:33
23:42:27.755 -> Water_Temp.:98.60, PWM:33
23:42:28.553 -> Water_Temp.:100.29, PWM:33
23:42:29.315 -> Water_Temp.:102.31, PWM:33
23:42:30.108 -> Water_Temp.:99.84, PWM:33
23:42:30.899 -> Water_Temp.:99.39, PWM:33
23:42:32.421 -> Water_Temp.:99.05, PWM:33
23:42:36.268 -> Water_Temp.:98.71, PWM:33
23:42:38.588 -> Water_Temp.:98.60, PWM:33
23:42:40.904 -> Water_Temp.:98.60, PWM:33
23:42:41.697 -> Water_Temp.:98.60, PWM:33
23:42:43.214 -> Water_Temp.:98.60, PWM:33
23:42:44.775 -> Water_Temp.:98.60, PWM:33
23:42:45.570 -> Water_Temp.:98.60, PWM:33
23:42:47.128 -> Water_Temp.:98.60, PWM:33
23:42:47.890 -> Water_Temp.:98.60, PWM:33
23:42:48.657 -> Water_Temp.:98.60, PWM:33
23:42:49.453 -> Water_Temp.:98.49, PWM:33
23:42:50.216 -> Water_Temp.:98.49, PWM:33
23:42:51.014 -> Water_Temp.:98.49, PWM:33
23:42:51.806 -> Water_Temp.:98.49, PWM:33
23:42:52.567 -> Water_Temp.:98.49, PWM:33
23:42:54.885 -> Water_Temp.:98.60, PWM:33
23:42:55.647 -> Water_Temp.:98.60, PWM:33
23:42:56.438 -> Water_Temp.:98.60, PWM:33
23:42:57.230 -> Water_Temp.:98.60, PWM:33
23:42:58.787 -> Water_Temp.:98.60, PWM:33
23:43:00.311 -> Water_Temp.:98.60, PWM:33
23:43:01.106 -> Water_Temp.:98.60, PWM:33
23:43:01.866 -> Water_Temp.:98.60, PWM:33
23:43:02.658 -> Water_Temp.:98.60, PWM:33
23:43:04.181 -> Water_Temp.:98.60, PWM:33
23:43:05.742 -> Water_Temp.:98.60, PWM:33
23:43:06.540 -> Water_Temp.:98.60, PWM:33
23:43:07.300 -> Water_Temp.:98.60, PWM:33
23:43:08.862 -> Water_Temp.:98.49, PWM:33
23:43:10.415 -> Water_Temp.:98.49, PWM:33
23:43:11.175 -> Water_Temp.:98.49, PWM:33
23:43:11.969 -> Water_Temp.:98.49, PWM:33
23:43:12.735 -> Water_Temp.:98.49, PWM:33
23:43:13.531 -> Water_Temp.:98.49, PWM:33
23:43:14.292 -> Water_Temp.:98.49, PWM:33
23:43:15.087 -> Water_Temp.:98.49, PWM:33
23:43:15.850 -> Water_Temp.:98.49, PWM:33
23:43:17.436 -> Water_Temp.:98.49, PWM:33
23:43:18.195 -> Water_Temp.:98.49, PWM:33
23:43:18.989 -> Water_Temp.:98.49, PWM:33
23:43:19.749 -> Water_Temp.:98.49, PWM:33
23:43:20.543 -> Water_Temp.:98.49, PWM:33
23:43:21.337 -> Water_Temp.:98.49, PWM:33
23:43:22.896 -> Water_Temp.:98.37, PWM:33
23:43:23.661 -> Water_Temp.:98.37, PWM:33
23:43:24.459 -> Water_Temp.:98.49, PWM:33
23:43:25.988 -> Water_Temp.:98.49, PWM:33
23:43:26.785 -> Water_Temp.:98.49, PWM:33
23:43:27.550 -> Water_Temp.:98.49, PWM:33
23:43:28.345 -> Water_Temp.:98.49, PWM:33
23:43:29.107 -> Water_Temp.:98.49, PWM:33
23:43:29.906 -> Water_Temp.:98.60, PWM:33
23:43:30.669 -> Water_Temp.:98.60, PWM:33
23:43:31.466 -> Water_Temp.:98.49, PWM:33
23:43:32.990 -> Water_Temp.:98.49, PWM:33
23:43:33.788 -> Water_Temp.:98.49, PWM:33
23:43:34.552 -> Water_Temp.:98.49, PWM:33
23:43:36.877 -> Water_Temp.:98.49, PWM:33
23:43:37.640 -> Water_Temp.:98.49, PWM:33
23:43:38.405 -> Water_Temp.:98.49, PWM:33
23:43:39.202 -> Water_Temp.:98.49, PWM:33
23:43:39.963 -> Water_Temp.:98.49, PWM:33
23:43:40.761 -> Water_Temp.:98.37, PWM:33
23:43:41.521 -> Water_Temp.:98.37, PWM:33
23:43:42.320 -> Water_Temp.:98.49, PWM:33
23:43:43.082 -> Water_Temp.:98.49, PWM:33
23:43:43.877 -> Water_Temp.:98.49, PWM:33
23:43:44.642 -> Water_Temp.:98.49, PWM:33
23:43:45.438 -> Water_Temp.:98.49, PWM:33
23:43:46.202 -> Water_Temp.:98.49, PWM:33
23:43:47.759 -> Water_Temp.:98.49, PWM:33
23:43:48.522 -> Water_Temp.:98.49, PWM:33
23:43:49.318 -> Water_Temp.:98.49, PWM:33
23:43:50.116 -> Water_Temp.:98.49, PWM:33
23:43:51.638 -> Water_Temp.:98.49, PWM:33
23:43:53.199 -> Water_Temp.:98.49, PWM:33
23:43:54.758 -> Water_Temp.:98.49, PWM:33
23:43:55.523 -> Water_Temp.:98.49, PWM:33
23:43:57.088 -> Water_Temp.:98.49, PWM:33
23:43:57.851 -> Water_Temp.:98.49, PWM:33
23:43:58.647 -> Water_Temp.:98.49, PWM:33
23:43:59.442 -> Water_Temp.:98.49, PWM:33
23:44:00.999 -> Water_Temp.:98.49, PWM:33
23:44:01.761 -> Water_Temp.:98.49, PWM:33
23:44:02.555 -> Water_Temp.:98.60, PWM:33
23:44:04.111 -> Water_Temp.:98.60, PWM:33
23:44:04.874 -> Water_Temp.:98.60, PWM:33
23:44:07.200 -> Water_Temp.:98.60, PWM:33
23:44:08.762 -> Water_Temp.:98.60, PWM:33
23:44:09.527 -> Water_Temp.:98.60, PWM:33
23:44:10.325 -> Water_Temp.:98.60, PWM:33
23:44:11.087 -> Water_Temp.:98.60, PWM:33
23:44:12.642 -> Water_Temp.:98.60, PWM:33
23:44:14.198 -> Water_Temp.:98.60, PWM:33
23:44:14.996 -> Water_Temp.:98.60, PWM:33
23:44:15.755 -> Water_Temp.:98.60, PWM:33
23:44:18.080 -> Water_Temp.:98.60, PWM:33
23:44:18.841 -> Water_Temp.:98.60, PWM:33
23:44:19.641 -> Water_Temp.:98.60, PWM:33
23:44:20.405 -> Water_Temp.:98.60, PWM:33
23:44:21.164 -> Water_Temp.:98.60, PWM:33
23:44:21.958 -> Water_Temp.:98.60, PWM:33
23:44:22.754 -> Water_Temp.:98.60, PWM:33
23:44:23.516 -> Water_Temp.:98.60, PWM:33
23:44:25.079 -> Water_Temp.:98.71, PWM:33
23:44:25.877 -> Water_Temp.:98.71, PWM:33
23:44:26.640 -> Water_Temp.:98.71, PWM:33
23:44:28.197 -> Water_Temp.:98.71, PWM:33
23:44:30.514 -> Water_Temp.:98.71, PWM:33
23:44:31.277 -> Water_Temp.:98.71, PWM:33
23:44:32.835 -> Water_Temp.:98.71, PWM:33
23:44:34.393 -> Water_Temp.:98.71, PWM:33
23:44:35.958 -> Water_Temp.:98.71, PWM:33
23:44:36.719 -> Water_Temp.:98.71, PWM:33
23:44:38.247 -> Water_Temp.:98.71, PWM:33
23:44:39.047 -> Water_Temp.:98.71, PWM:33
23:44:39.810 -> Water_Temp.:98.60, PWM:33
23:44:41.373 -> Water_Temp.:98.60, PWM:33
23:44:42.170 -> Water_Temp.:98.60, PWM:33
23:44:43.726 -> Water_Temp.:98.60, PWM:33
23:44:45.255 -> Water_Temp.:98.49, PWM:33
23:44:46.811 -> Water_Temp.:98.49, PWM:33
23:44:47.609 -> Water_Temp.:98.49, PWM:33
23:44:48.369 -> Water_Temp.:98.49, PWM:33
23:44:49.135 -> Water_Temp.:98.49, PWM:33
23:44:51.444 -> Water_Temp.:98.49, PWM:33
23:44:53.000 -> Water_Temp.:98.49, PWM:33
23:44:53.799 -> Water_Temp.:98.49, PWM:33
23:44:55.324 -> Water_Temp.:98.49, PWM:33
23:44:57.678 -> Water_Temp.:98.49, PWM:33
23:44:58.442 -> Water_Temp.:98.49, PWM:33
23:44:59.243 -> Water_Temp.:98.49, PWM:33
23:45:00.005 -> Water_Temp.:98.49, PWM:33
23:45:01.533 -> Water_Temp.:98.60, PWM:33
23:45:02.329 -> Water_Temp.:98.60, PWM:33
23:45:03.890 -> Water_Temp.:98.60, PWM:33
23:45:04.651 -> Water_Temp.:98.60, PWM:33
23:45:06.211 -> Water_Temp.:98.60, PWM:33
23:45:07.774 -> Water_Temp.:98.60, PWM:33
23:45:08.538 -> Water_Temp.:98.60, PWM:33
23:45:09.337 -> Water_Temp.:98.60, PWM:33
23:45:10.100 -> Water_Temp.:98.60, PWM:33
23:45:10.896 -> Water_Temp.:98.60, PWM:33
23:45:11.658 -> Water_Temp.:98.60, PWM:33
23:45:13.219 -> Water_Temp.:98.49, PWM:33
23:45:14.015 -> Water_Temp.:98.49, PWM:33
23:45:17.034 -> Water_Temp.:98.49, PWM:33
23:45:17.834 -> Water_Temp.:98.60, PWM:33
23:45:18.598 -> Water_Temp.:98.60, PWM:33
23:45:20.155 -> Water_Temp.:98.60, PWM:33
23:45:20.952 -> Water_Temp.:98.60, PWM:33
23:45:21.716 -> Water_Temp.:98.60, PWM:33
23:45:23.273 -> Water_Temp.:98.60, PWM:33
23:45:24.073 -> Water_Temp.:98.60, PWM:33
23:45:24.834 -> Water_Temp.:98.60, PWM:33
23:45:26.399 -> Water_Temp.:98.60, PWM:33
23:45:27.194 -> Water_Temp.:98.60, PWM:33
23:45:27.957 -> Water_Temp.:98.60, PWM:33
23:45:28.754 -> Water_Temp.:98.60, PWM:33
23:45:30.310 -> Water_Temp.:98.60, PWM:33
23:45:31.830 -> Water_Temp.:98.71, PWM:33
23:45:32.627 -> Water_Temp.:98.71, PWM:33
23:45:33.386 -> Water_Temp.:98.71, PWM:33
23:45:34.182 -> Water_Temp.:98.71, PWM:33
23:45:35.738 -> Water_Temp.:98.71, PWM:33
23:45:36.503 -> Water_Temp.:98.71, PWM:33
23:45:39.584 -> Water_Temp.:98.71, PWM:33
23:45:40.346 -> Water_Temp.:98.82, PWM:33
23:45:41.144 -> Water_Temp.:98.82, PWM:33
23:45:41.933 -> Water_Temp.:98.82, PWM:33
23:45:44.218 -> Water_Temp.:98.82, PWM:33
23:45:44.977 -> Water_Temp.:98.82, PWM:33
23:45:46.533 -> Water_Temp.:98.71, PWM:33
23:45:47.326 -> Water_Temp.:98.71, PWM:33
23:45:48.120 -> Water_Temp.:98.71, PWM:33
23:45:50.439 -> Water_Temp.:98.71, PWM:33
23:45:51.231 -> Water_Temp.:98.71, PWM:33
23:45:52.749 -> Water_Temp.:98.71, PWM:33
23:45:53.545 -> Water_Temp.:98.71, PWM:33
23:45:54.305 -> Water_Temp.:98.71, PWM:33
23:45:55.101 -> Water_Temp.:100.40, PWM:33
23:45:55.861 -> Water_Temp.:102.54, PWM:33
23:45:57.418 -> Water_Temp.:100.06, PWM:33
23:46:01.328 -> Water_Temp.:99.05, PWM:33
23:46:02.092 -> Water_Temp.:99.05, PWM:33
23:46:02.889 -> Water_Temp.:99.05, PWM:33
23:46:04.449 -> Water_Temp.:98.94, PWM:33
23:46:05.205 -> Water_Temp.:98.82, PWM:33
23:46:06.004 -> Water_Temp.:98.82, PWM:33
23:46:07.562 -> Water_Temp.:98.82, PWM:33
23:46:08.323 -> Water_Temp.:98.82, PWM:33
23:46:09.123 -> Water_Temp.:98.82, PWM:33
23:46:10.680 -> Water_Temp.:98.82, PWM:33
23:46:11.475 -> Water_Temp.:98.82, PWM:33
23:46:12.235 -> Water_Temp.:98.82, PWM:33
23:46:14.553 -> Water_Temp.:98.82, PWM:33
23:46:16.116 -> Water_Temp.:98.94, PWM:33
23:46:17.642 -> Water_Temp.:98.94, PWM:33
23:46:18.436 -> Water_Temp.:98.94, PWM:33
23:46:19.197 -> Water_Temp.:98.94, PWM:33
23:46:21.512 -> Water_Temp.:98.94, PWM:33
23:46:22.273 -> Water_Temp.:98.94, PWM:33
23:46:23.069 -> Water_Temp.:98.94, PWM:33
23:46:23.827 -> Water_Temp.:98.94, PWM:33
23:46:24.626 -> Water_Temp.:98.94, PWM:33
23:46:25.387 -> Water_Temp.:98.94, PWM:33
23:46:26.945 -> Water_Temp.:98.94, PWM:33
23:46:28.499 -> Water_Temp.:98.94, PWM:33
23:46:29.261 -> Water_Temp.:98.94, PWM:33
23:46:30.055 -> Water_Temp.:98.94, PWM:33
23:46:31.605 -> Water_Temp.:98.94, PWM:33
23:46:32.367 -> Water_Temp.:98.94, PWM:33
23:46:33.156 -> Water_Temp.:98.94, PWM:33
23:46:33.951 -> Water_Temp.:100.51, PWM:33
23:46:34.714 -> Water_Temp.:102.65, PWM:33
23:46:37.801 -> Water_Temp.:99.39, PWM:33
23:46:38.601 -> Water_Temp.:99.27, PWM:33
23:46:40.161 -> Water_Temp.:99.16, PWM:33
23:46:40.924 -> Water_Temp.:99.16, PWM:33
23:46:42.445 -> Water_Temp.:99.16, PWM:33
23:46:43.242 -> Water_Temp.:99.16, PWM:33
23:46:44.007 -> Water_Temp.:99.16, PWM:33
23:46:45.598 -> Water_Temp.:99.16, PWM:33
23:46:47.156 -> Water_Temp.:99.16, PWM:33
23:46:47.916 -> Water_Temp.:99.16, PWM:33
23:46:49.473 -> Water_Temp.:99.05, PWM:33
23:46:50.271 -> Water_Temp.:99.05, PWM:33
23:46:52.557 -> Water_Temp.:99.05, PWM:33
23:46:53.354 -> Water_Temp.:99.05, PWM:33
23:46:54.118 -> Water_Temp.:99.05, PWM:33
23:46:56.441 -> Water_Temp.:99.05, PWM:33
23:46:57.237 -> Water_Temp.:99.05, PWM:33
23:46:58.763 -> Water_Temp.:99.05, PWM:33
23:46:59.562 -> Water_Temp.:99.05, PWM:33
23:47:01.876 -> Water_Temp.:99.05, PWM:33
23:47:03.431 -> Water_Temp.:99.05, PWM:33
23:47:04.226 -> Water_Temp.:99.05, PWM:33
23:47:04.990 -> Water_Temp.:99.05, PWM:33
23:47:07.306 -> Water_Temp.:99.05, PWM:33
23:47:08.068 -> Water_Temp.:99.05, PWM:33
23:47:08.866 -> Water_Temp.:99.05, PWM:33
23:47:09.630 -> Water_Temp.:99.05, PWM:33
23:47:11.157 -> Water_Temp.:99.05, PWM:33
23:47:11.956 -> Water_Temp.:99.05, PWM:33
23:47:12.719 -> Water_Temp.:99.05, PWM:33
23:47:13.512 -> Water_Temp.:99.05, PWM:33
23:47:15.032 -> Water_Temp.:99.05, PWM:33
23:47:15.829 -> Water_Temp.:99.05, PWM:33
23:47:16.593 -> Water_Temp.:99.05, PWM:33
23:47:17.389 -> Water_Temp.:99.05, PWM:33
23:47:19.707 -> Water_Temp.:99.16, PWM:33
23:47:21.268 -> Water_Temp.:99.16, PWM:33
23:47:22.790 -> Water_Temp.:99.16, PWM:33
23:47:23.588 -> Water_Temp.:99.16, PWM:33
23:47:25.113 -> Water_Temp.:99.05, PWM:33
23:47:27.432 -> Water_Temp.:99.05, PWM:33
23:47:28.232 -> Water_Temp.:99.05, PWM:33
23:47:28.997 -> Water_Temp.:99.05, PWM:33
23:47:29.795 -> Water_Temp.:99.05, PWM:33
23:47:30.555 -> Water_Temp.:99.05, PWM:33
23:47:31.350 -> Water_Temp.:99.05, PWM:33
23:47:32.113 -> Water_Temp.:99.05, PWM:33
23:47:33.672 -> Water_Temp.:99.05, PWM:33
23:47:34.434 -> Water_Temp.:99.05, PWM:33
23:47:36.788 -> Water_Temp.:98.94, PWM:33
23:47:37.551 -> Water_Temp.:98.94, PWM:33
23:47:38.344 -> Water_Temp.:98.94, PWM:33
23:47:39.143 -> Water_Temp.:98.94, PWM:33
23:47:39.902 -> Water_Temp.:98.94, PWM:33
23:47:40.700 -> Water_Temp.:98.94, PWM:33
23:47:43.753 -> Water_Temp.:98.94, PWM:33
23:47:44.516 -> Water_Temp.:98.94, PWM:33
23:47:45.309 -> Water_Temp.:98.94, PWM:33
23:47:47.638 -> Water_Temp.:99.05, PWM:33
23:47:48.437 -> Water_Temp.:99.05, PWM:33
23:47:49.201 -> Water_Temp.:99.05, PWM:33
23:47:49.998 -> Water_Temp.:99.05, PWM:33
23:47:51.523 -> Water_Temp.:99.05, PWM:33
23:47:53.080 -> Water_Temp.:99.05, PWM:33
23:47:54.636 -> Water_Temp.:99.05, PWM:33
23:47:56.157 -> Water_Temp.:99.05, PWM:33
23:47:56.953 -> Water_Temp.:99.05, PWM:33
23:47:58.513 -> Water_Temp.:99.05, PWM:33
23:47:59.277 -> Water_Temp.:99.05, PWM:33
23:48:02.394 -> Water_Temp.:99.16, PWM:33
23:48:03.953 -> Water_Temp.:99.16, PWM:33
23:48:04.718 -> Water_Temp.:99.16, PWM:33
23:48:07.834 -> Water_Temp.:99.16, PWM:33
23:48:08.599 -> Water_Temp.:99.16, PWM:33
23:48:09.364 -> Water_Temp.:99.16, PWM:33
23:48:10.162 -> Water_Temp.:99.16, PWM:33
23:48:10.922 -> Water_Temp.:99.16, PWM:33
23:48:12.484 -> Water_Temp.:99.16, PWM:33
23:48:14.041 -> Water_Temp.:99.16, PWM:33
23:48:14.805 -> Water_Temp.:99.16, PWM:33
23:48:15.601 -> Water_Temp.:99.16, PWM:33
23:48:17.126 -> Water_Temp.:99.16, PWM:33
23:48:17.889 -> Water_Temp.:99.16, PWM:33
23:48:19.451 -> Water_Temp.:99.16, PWM:33
23:48:21.007 -> Water_Temp.:99.16, PWM:33
23:48:21.798 -> Water_Temp.:99.16, PWM:33
23:48:23.325 -> Water_Temp.:99.16, PWM:33
23:48:24.086 -> Water_Temp.:99.16, PWM:33
23:48:24.884 -> Water_Temp.:99.16, PWM:33
23:48:25.646 -> Water_Temp.:99.16, PWM:33
23:48:26.447 -> Water_Temp.:99.16, PWM:33
23:48:28.760 -> Water_Temp.:99.16, PWM:33
23:48:29.524 -> Water_Temp.:99.16, PWM:33
23:48:31.078 -> Water_Temp.:99.16, PWM:33
23:48:31.871 -> Water_Temp.:99.16, PWM:33
23:48:32.633 -> Water_Temp.:99.16, PWM:33
23:48:33.431 -> Water_Temp.:99.16, PWM:33
23:48:34.195 -> Water_Temp.:99.16, PWM:33
23:48:35.722 -> Water_Temp.:99.16, PWM:33
23:48:36.519 -> Water_Temp.:99.16, PWM:33
23:48:38.082 -> Water_Temp.:99.16, PWM:33
23:48:39.604 -> Water_Temp.:99.16, PWM:33
23:48:40.400 -> Water_Temp.:99.16, PWM:33
23:48:41.158 -> Water_Temp.:99.05, PWM:33
23:48:41.951 -> Water_Temp.:99.05, PWM:33
23:48:44.264 -> Water_Temp.:99.16, PWM:33
23:48:45.060 -> Water_Temp.:99.05, PWM:33
23:48:46.613 -> Water_Temp.:99.05, PWM:33
23:48:47.374 -> Water_Temp.:99.05, PWM:33
23:48:48.169 -> Water_Temp.:99.05, PWM:33
23:48:49.695 -> Water_Temp.:99.05, PWM:33
23:48:50.491 -> Water_Temp.:99.05, PWM:33
23:48:51.253 -> Water_Temp.:99.05, PWM:33
23:48:52.799 -> Water_Temp.:99.05, PWM:33
23:48:53.594 -> Water_Temp.:99.05, PWM:33
23:48:54.389 -> Water_Temp.:99.05, PWM:33
23:48:55.148 -> Water_Temp.:99.05, PWM:33
23:48:55.943 -> Water_Temp.:99.05, PWM:33
23:48:57.464 -> Water_Temp.:99.05, PWM:33
23:49:01.280 -> Water_Temp.:99.05, PWM:33
23:49:02.074 -> Water_Temp.:99.05, PWM:33
23:49:02.833 -> Water_Temp.:99.05, PWM:33
23:49:03.629 -> Water_Temp.:99.05, PWM:33
23:49:05.152 -> Water_Temp.:99.05, PWM:33
23:49:05.946 -> Water_Temp.:98.94, PWM:33
23:49:07.498 -> Water_Temp.:98.94, PWM:33
23:49:08.261 -> Water_Temp.:98.94, PWM:33
23:49:09.819 -> Water_Temp.:98.94, PWM:33
23:49:10.584 -> Water_Temp.:98.94, PWM:33
23:49:12.139 -> Water_Temp.:98.94, PWM:33
23:49:12.899 -> Water_Temp.:98.94, PWM:33
23:49:14.463 -> Water_Temp.:98.94, PWM:33
23:49:15.262 -> Water_Temp.:98.94, PWM:33
23:49:16.024 -> Water_Temp.:98.94, PWM:33
23:49:16.816 -> Water_Temp.:98.94, PWM:33
23:49:17.582 -> Water_Temp.:98.94, PWM:33
23:49:18.381 -> Water_Temp.:98.94, PWM:33
23:49:19.909 -> Water_Temp.:98.94, PWM:33
23:49:21.470 -> Water_Temp.:98.94, PWM:33
23:49:22.266 -> Water_Temp.:98.94, PWM:33
23:49:23.030 -> Water_Temp.:98.94, PWM:33
23:49:23.826 -> Water_Temp.:98.94, PWM:33
23:49:24.592 -> Water_Temp.:98.94, PWM:33
23:49:25.385 -> Water_Temp.:98.94, PWM:33
23:49:26.144 -> Water_Temp.:98.94, PWM:33
23:49:27.695 -> Water_Temp.:99.05, PWM:33
23:49:28.492 -> Water_Temp.:98.94, PWM:33
23:49:29.258 -> Water_Temp.:98.94, PWM:33
23:49:30.791 -> Water_Temp.:99.05, PWM:33
23:49:31.586 -> Water_Temp.:99.05, PWM:33
23:49:32.343 -> Water_Temp.:99.05, PWM:33
23:49:33.144 -> Water_Temp.:99.05, PWM:33
23:49:33.906 -> Water_Temp.:99.05, PWM:33
23:49:34.705 -> Water_Temp.:99.05, PWM:33
23:49:35.466 -> Water_Temp.:99.05, PWM:33
23:49:36.264 -> Water_Temp.:99.05, PWM:33
23:49:38.581 -> Water_Temp.:98.94, PWM:33
23:49:39.345 -> Water_Temp.:98.94, PWM:33
23:49:40.904 -> Water_Temp.:98.94, PWM:33
23:49:43.194 -> Water_Temp.:99.05, PWM:33
23:49:44.752 -> Water_Temp.:99.05, PWM:33
23:49:45.547 -> Water_Temp.:99.05, PWM:33
23:49:46.310 -> Water_Temp.:99.05, PWM:33
23:49:48.636 -> Water_Temp.:99.05, PWM:33
23:49:49.402 -> Water_Temp.:99.05, PWM:33
23:49:50.954 -> Water_Temp.:99.05, PWM:33
23:49:51.711 -> Water_Temp.:99.05, PWM:33
23:49:52.506 -> Water_Temp.:99.05, PWM:33
23:49:53.295 -> Water_Temp.:99.05, PWM:33
23:49:54.819 -> Water_Temp.:99.05, PWM:33
23:49:55.614 -> Water_Temp.:99.05, PWM:33
23:50:00.206 -> Water_Temp.:99.05, PWM:33
23:50:01.002 -> Water_Temp.:99.05, PWM:33
23:50:01.793 -> Water_Temp.:99.05, PWM:33
23:50:02.554 -> Water_Temp.:99.05, PWM:33
23:50:03.348 -> Water_Temp.:99.05, PWM:33
23:50:04.110 -> Water_Temp.:99.05, PWM:33
23:50:04.903 -> Water_Temp.:99.05, PWM:33
23:50:10.261 -> Water_Temp.:99.05, PWM:33
23:50:11.021 -> Water_Temp.:99.05, PWM:33
23:50:12.577 -> Water_Temp.:99.05, PWM:33
23:50:13.374 -> Water_Temp.:99.05, PWM:33
23:50:14.931 -> Water_Temp.:99.05, PWM:33
23:50:16.485 -> Water_Temp.:99.16, PWM:33
23:50:17.243 -> Water_Temp.:99.16, PWM:33
23:50:19.594 -> Water_Temp.:99.05, PWM:33
23:50:20.353 -> Water_Temp.:99.05, PWM:33
23:50:21.153 -> Water_Temp.:99.05, PWM:33
23:50:21.915 -> Water_Temp.:99.05, PWM:33
23:50:23.477 -> Water_Temp.:99.05, PWM:33
23:50:24.275 -> Water_Temp.:99.16, PWM:33
23:50:25.037 -> Water_Temp.:99.16, PWM:33
23:50:25.837 -> Water_Temp.:99.16, PWM:33
23:50:26.601 -> Water_Temp.:99.16, PWM:33
23:50:28.954 -> Water_Temp.:99.16, PWM:33
23:50:30.476 -> Water_Temp.:99.16, PWM:33
23:50:32.035 -> Water_Temp.:99.16, PWM:33
23:50:32.799 -> Water_Temp.:99.16, PWM:33
23:50:34.356 -> Water_Temp.:99.27, PWM:33
23:50:35.116 -> Water_Temp.:99.27, PWM:33
23:50:36.676 -> Water_Temp.:99.27, PWM:33
23:50:38.227 -> Water_Temp.:99.27, PWM:33
23:50:38.988 -> Water_Temp.:99.27, PWM:33
23:50:39.783 -> Water_Temp.:99.27, PWM:33
23:50:41.343 -> Water_Temp.:99.27, PWM:33
23:50:44.390 -> Water_Temp.:99.27, PWM:33
23:50:45.183 -> Water_Temp.:99.27, PWM:33
23:50:45.979 -> Water_Temp.:99.27, PWM:33
23:50:46.743 -> Water_Temp.:99.16, PWM:33
23:50:48.300 -> Water_Temp.:99.16, PWM:33
23:50:49.859 -> Water_Temp.:99.16, PWM:33
23:50:50.623 -> Water_Temp.:99.16, PWM:33
23:50:51.414 -> Water_Temp.:99.16, PWM:33
23:50:52.173 -> Water_Temp.:99.16, PWM:33
23:50:53.729 -> Water_Temp.:99.16, PWM:33
23:50:54.523 -> Water_Temp.:99.16, PWM:33
23:50:56.080 -> Water_Temp.:99.16, PWM:33
23:50:57.599 -> Water_Temp.:99.16, PWM:33
23:50:59.911 -> Water_Temp.:99.05, PWM:33
23:51:00.676 -> Water_Temp.:99.05, PWM:33
23:51:03.790 -> Water_Temp.:99.16, PWM:33
23:51:04.553 -> Water_Temp.:99.16, PWM:33
23:51:05.349 -> Water_Temp.:99.16, PWM:33
23:51:06.108 -> Water_Temp.:99.16, PWM:33
23:51:09.231 -> Water_Temp.:99.16, PWM:33
23:51:10.023 -> Water_Temp.:99.16, PWM:33
23:51:10.782 -> Water_Temp.:99.16, PWM:33
23:51:11.580 -> Water_Temp.:99.16, PWM:33
23:51:12.341 -> Water_Temp.:99.16, PWM:33
23:51:13.140 -> Water_Temp.:99.27, PWM:33
23:51:16.198 -> Water_Temp.:99.27, PWM:33
23:51:16.995 -> Water_Temp.:99.27, PWM:33
23:51:18.521 -> Water_Temp.:99.27, PWM:33
23:51:19.319 -> Water_Temp.:99.27, PWM:33
23:51:20.845 -> Water_Temp.:99.27, PWM:33
23:51:21.608 -> Water_Temp.:99.27, PWM:33
23:51:23.165 -> Water_Temp.:99.27, PWM:33
23:51:23.961 -> Water_Temp.:99.27, PWM:33
23:51:25.520 -> Water_Temp.:99.27, PWM:33
23:51:26.286 -> Water_Temp.:99.27, PWM:33
23:51:29.366 -> Water_Temp.:99.27, PWM:33
23:51:30.922 -> Water_Temp.:99.27, PWM:33
23:51:31.686 -> Water_Temp.:99.27, PWM:33
23:51:32.481 -> Water_Temp.:99.27, PWM:33
23:51:33.241 -> Water_Temp.:99.27, PWM:33
23:51:34.038 -> Water_Temp.:99.27, PWM:33
23:51:36.327 -> Water_Temp.:99.27, PWM:33
23:51:37.093 -> Water_Temp.:99.27, PWM:33
23:51:37.887 -> Water_Temp.:99.27, PWM:33
23:51:40.206 -> Water_Temp.:99.27, PWM:33
23:51:40.966 -> Water_Temp.:99.27, PWM:33
23:51:41.759 -> Water_Temp.:99.27, PWM:33
23:51:42.555 -> Water_Temp.:99.27, PWM:33
23:51:43.316 -> Water_Temp.:99.27, PWM:33
23:51:47.199 -> Water_Temp.:99.16, PWM:33
23:51:47.959 -> Water_Temp.:99.16, PWM:33
23:51:48.757 -> Water_Temp.:99.16, PWM:33
23:51:51.081 -> Water_Temp.:99.16, PWM:33
23:51:52.602 -> Water_Temp.:99.16, PWM:33
23:51:53.402 -> Water_Temp.:99.16, PWM:33
23:51:54.920 -> Water_Temp.:99.16, PWM:33
23:51:55.712 -> Water_Temp.:99.16, PWM:33
23:51:56.502 -> Water_Temp.:99.27, PWM:33
23:51:57.261 -> Water_Temp.:99.27, PWM:33
23:51:59.583 -> Water_Temp.:99.27, PWM:33
23:52:00.379 -> Water_Temp.:99.27, PWM:33
23:52:03.444 -> Water_Temp.:99.27, PWM:33
23:52:04.238 -> Water_Temp.:99.27, PWM:33
23:52:04.999 -> Water_Temp.:99.27, PWM:33
23:52:05.794 -> Water_Temp.:99.27, PWM:33
23:52:07.353 -> Water_Temp.:99.27, PWM:33
23:52:08.115 -> Water_Temp.:99.27, PWM:33
23:52:08.915 -> Water_Temp.:99.27, PWM:33
23:52:09.671 -> Water_Temp.:99.27, PWM:33
23:52:10.468 -> Water_Temp.:99.27, PWM:33
23:52:11.230 -> Water_Temp.:99.27, PWM:33
23:52:12.029 -> Water_Temp.:99.27, PWM:33
23:52:14.319 -> Water_Temp.:99.27, PWM:33
23:52:15.878 -> Water_Temp.:99.27, PWM:33
23:52:16.676 -> Water_Temp.:99.27, PWM:33
23:52:17.438 -> Water_Temp.:99.27, PWM:33
23:52:18.233 -> Water_Temp.:99.27, PWM:33
23:52:19.797 -> Water_Temp.:99.27, PWM:33
23:52:21.323 -> Water_Temp.:99.27, PWM:33
23:52:24.409 -> Water_Temp.:99.39, PWM:33
23:52:25.204 -> Water_Temp.:99.27, PWM:33
23:52:26.727 -> Water_Temp.:99.39, PWM:33
23:52:27.521 -> Water_Temp.:99.39, PWM:33
23:52:28.317 -> Water_Temp.:99.39, PWM:33
23:52:29.081 -> Water_Temp.:99.39, PWM:33
23:52:30.642 -> Water_Temp.:99.39, PWM:33
23:52:32.198 -> Water_Temp.:99.39, PWM:33
23:52:32.964 -> Water_Temp.:99.39, PWM:33
23:52:33.763 -> Water_Temp.:99.39, PWM:33
23:52:34.525 -> Water_Temp.:99.39, PWM:33
23:52:36.082 -> Water_Temp.:99.39, PWM:33
23:52:37.607 -> Water_Temp.:99.39, PWM:33
23:52:38.401 -> Water_Temp.:99.39, PWM:33
23:52:39.167 -> Water_Temp.:99.39, PWM:33
23:52:40.723 -> Water_Temp.:99.39, PWM:33
23:52:42.282 -> Water_Temp.:99.50, PWM:33
23:52:43.042 -> Water_Temp.:99.50, PWM:33
23:52:44.600 -> Water_Temp.:99.50, PWM:33
23:52:46.119 -> Water_Temp.:99.61, PWM:33
23:52:46.914 -> Water_Temp.:99.50, PWM:33
23:52:48.438 -> Water_Temp.:99.50, PWM:33
23:52:49.998 -> Water_Temp.:99.50, PWM:33
23:52:50.758 -> Water_Temp.:99.50, PWM:33
23:52:52.317 -> Water_Temp.:99.50, PWM:33
23:52:53.118 -> Water_Temp.:99.39, PWM:33
23:52:55.442 -> Water_Temp.:99.39, PWM:33
23:52:56.967 -> Water_Temp.:99.39, PWM:33
23:52:58.560 -> Water_Temp.:99.50, PWM:33
23:52:59.325 -> Water_Temp.:99.50, PWM:33
23:53:00.114 -> Water_Temp.:99.50, PWM:33
23:53:00.876 -> Water_Temp.:99.50, PWM:33
23:53:01.671 -> Water_Temp.:99.50, PWM:33
23:53:02.430 -> Water_Temp.:99.50, PWM:33
23:53:03.225 -> Water_Temp.:99.39, PWM:33
23:53:03.986 -> Water_Temp.:99.39, PWM:33
23:53:04.783 -> Water_Temp.:99.39, PWM:33
23:53:06.343 -> Water_Temp.:99.39, PWM:33
23:53:07.136 -> Water_Temp.:99.39, PWM:33
23:53:08.655 -> Water_Temp.:99.39, PWM:33
23:53:09.420 -> Water_Temp.:99.39, PWM:33
23:53:14.051 -> Water_Temp.:99.39, PWM:33
23:53:14.845 -> Water_Temp.:99.27, PWM:33
23:53:16.399 -> Water_Temp.:99.27, PWM:33
23:53:17.955 -> Water_Temp.:99.27, PWM:33
23:53:19.514 -> Water_Temp.:99.27, PWM:33
23:53:20.304 -> Water_Temp.:99.27, PWM:33
23:53:21.062 -> Water_Temp.:99.27, PWM:33
23:53:21.859 -> Water_Temp.:99.27, PWM:33
23:53:22.621 -> Water_Temp.:99.27, PWM:33
23:53:24.180 -> Water_Temp.:99.27, PWM:33
23:53:25.731 -> Water_Temp.:99.27, PWM:33
23:53:27.253 -> Water_Temp.:99.27, PWM:33
23:53:28.048 -> Water_Temp.:99.27, PWM:33
23:53:28.841 -> Water_Temp.:99.39, PWM:33
23:53:29.604 -> Water_Temp.:99.39, PWM:33
23:53:31.158 -> Water_Temp.:99.39, PWM:33
23:53:31.914 -> Water_Temp.:99.39, PWM:33
23:53:32.710 -> Water_Temp.:99.39, PWM:33
23:53:34.233 -> Water_Temp.:99.39, PWM:33
23:53:35.028 -> Water_Temp.:99.39, PWM:33
23:53:35.823 -> Water_Temp.:99.39, PWM:33
23:53:38.108 -> Water_Temp.:99.39, PWM:33
23:53:38.904 -> Water_Temp.:99.39, PWM:33
23:53:39.665 -> Water_Temp.:99.39, PWM:33
23:53:40.465 -> Water_Temp.:99.39, PWM:33
23:53:41.225 -> Water_Temp.:99.39, PWM:33
23:53:42.021 -> Water_Temp.:99.39, PWM:33
23:53:43.543 -> Water_Temp.:99.50, PWM:33
23:53:48.184 -> Water_Temp.:99.50, PWM:33
23:53:48.947 -> Water_Temp.:99.50, PWM:33
23:53:50.500 -> Water_Temp.:99.50, PWM:33
23:53:51.291 -> Water_Temp.:99.50, PWM:33
23:53:52.086 -> Water_Temp.:99.50, PWM:33
23:53:54.363 -> Water_Temp.:99.50, PWM:33
23:53:55.157 -> Water_Temp.:99.50, PWM:33
23:53:55.919 -> Water_Temp.:99.50, PWM:33
23:53:56.715 -> Water_Temp.:99.50, PWM:33
23:53:57.474 -> Water_Temp.:99.50, PWM:33
23:53:58.268 -> Water_Temp.:99.50, PWM:33
23:53:59.066 -> Water_Temp.:99.50, PWM:33
23:53:59.826 -> Water_Temp.:99.50, PWM:33
23:54:00.622 -> Water_Temp.:99.50, PWM:33
23:54:02.939 -> Water_Temp.:99.50, PWM:33
23:54:03.732 -> Water_Temp.:99.50, PWM:33
23:54:05.255 -> Water_Temp.:99.39, PWM:33
23:54:06.810 -> Water_Temp.:99.39, PWM:33
23:54:09.129 -> Water_Temp.:99.39, PWM:33
23:54:10.653 -> Water_Temp.:99.39, PWM:33
23:54:11.449 -> Water_Temp.:99.39, PWM:33
23:54:12.210 -> Water_Temp.:99.50, PWM:33
23:54:13.007 -> Water_Temp.:99.50, PWM:33
23:54:14.535 -> Water_Temp.:99.50, PWM:33
23:54:15.326 -> Water_Temp.:99.50, PWM:33
23:54:16.121 -> Water_Temp.:99.50, PWM:33
23:54:16.880 -> Water_Temp.:99.50, PWM:33
23:54:17.672 -> Water_Temp.:99.39, PWM:33
23:54:18.432 -> Water_Temp.:99.39, PWM:33
23:54:19.992 -> Water_Temp.:99.39, PWM:33
23:54:22.313 -> Water_Temp.:99.50, PWM:33
23:54:23.104 -> Water_Temp.:99.50, PWM:33
23:54:23.863 -> Water_Temp.:99.39, PWM:33
23:54:24.665 -> Water_Temp.:99.39, PWM:33
23:54:26.952 -> Water_Temp.:99.39, PWM:33
23:54:27.716 -> Water_Temp.:99.39, PWM:33
23:54:29.273 -> Water_Temp.:99.39, PWM:33
23:54:30.833 -> Water_Temp.:99.50, PWM:33
23:54:31.595 -> Water_Temp.:99.50, PWM:33
23:54:32.387 -> Water_Temp.:99.50, PWM:33
23:54:33.947 -> Water_Temp.:99.50, PWM:33
23:54:34.713 -> Water_Temp.:99.50, PWM:33
23:54:36.243 -> Water_Temp.:99.50, PWM:33
23:54:37.041 -> Water_Temp.:99.39, PWM:33
23:54:37.806 -> Water_Temp.:99.39, PWM:33
23:54:38.602 -> Water_Temp.:99.50, PWM:33
23:54:39.367 -> Water_Temp.:99.39, PWM:33
23:54:40.165 -> Water_Temp.:99.39, PWM:33
23:54:40.927 -> Water_Temp.:99.50, PWM:33
23:54:41.721 -> Water_Temp.:99.50, PWM:33
23:54:42.484 -> Water_Temp.:99.50, PWM:33
23:54:43.281 -> Water_Temp.:99.50, PWM:33
23:54:44.806 -> Water_Temp.:99.50, PWM:33
23:54:45.606 -> Water_Temp.:99.50, PWM:33
23:54:46.366 -> Water_Temp.:99.50, PWM:33
23:54:47.162 -> Water_Temp.:99.50, PWM:33
23:54:47.954 -> Water_Temp.:99.50, PWM:33
23:54:51.040 -> Water_Temp.:99.50, PWM:33
23:54:53.349 -> Water_Temp.:99.39, PWM:33
23:54:54.144 -> Water_Temp.:99.39, PWM:33
23:54:54.939 -> Water_Temp.:99.39, PWM:33
23:54:55.699 -> Water_Temp.:99.39, PWM:33
23:54:57.248 -> Water_Temp.:99.39, PWM:33
23:54:58.806 -> Water_Temp.:99.39, PWM:33
23:55:00.367 -> Water_Temp.:99.39, PWM:33
23:55:01.922 -> Water_Temp.:99.50, PWM:33
23:55:02.717 -> Water_Temp.:99.39, PWM:33
23:55:03.480 -> Water_Temp.:99.50, PWM:33
23:55:04.279 -> Water_Temp.:99.50, PWM:33
23:55:05.042 -> Water_Temp.:99.39, PWM:33
23:55:08.117 -> Water_Temp.:99.39, PWM:33
23:55:08.880 -> Water_Temp.:99.39, PWM:33
23:55:09.676 -> Water_Temp.:99.50, PWM:33
23:55:10.438 -> Water_Temp.:99.50, PWM:33
23:55:11.233 -> Water_Temp.:99.50, PWM:33
23:55:11.996 -> Water_Temp.:99.50, PWM:33
23:55:12.788 -> Water_Temp.:99.50, PWM:33
23:55:13.583 -> Water_Temp.:99.50, PWM:33
23:55:14.346 -> Water_Temp.:99.50, PWM:33
23:55:15.899 -> Water_Temp.:99.61, PWM:33
23:55:16.691 -> Water_Temp.:99.61, PWM:33
23:55:17.453 -> Water_Temp.:99.61, PWM:33
23:55:19.773 -> Water_Temp.:99.61, PWM:33
23:55:20.536 -> Water_Temp.:99.50, PWM:33
23:55:21.333 -> Water_Temp.:99.50, PWM:33
23:55:22.099 -> Water_Temp.:99.50, PWM:33
23:55:22.861 -> Water_Temp.:99.50, PWM:33
23:55:24.419 -> Water_Temp.:99.50, PWM:33
23:55:25.213 -> Water_Temp.:99.50, PWM:33
23:55:26.004 -> Water_Temp.:99.50, PWM:33
23:55:26.766 -> Water_Temp.:99.50, PWM:33
23:55:27.562 -> Water_Temp.:99.50, PWM:33
23:55:28.323 -> Water_Temp.:99.50, PWM:33
23:55:29.124 -> Water_Temp.:99.50, PWM:33
23:55:29.886 -> Water_Temp.:99.39, PWM:33
23:55:30.684 -> Water_Temp.:99.39, PWM:33
23:55:32.207 -> Water_Temp.:99.39, PWM:33
23:55:34.534 -> Water_Temp.:99.39, PWM:33
23:55:36.096 -> Water_Temp.:99.39, PWM:33
23:55:36.859 -> Water_Temp.:99.39, PWM:33
23:55:37.659 -> Water_Temp.:99.39, PWM:33
23:55:38.421 -> Water_Temp.:99.50, PWM:33
23:55:39.221 -> Water_Temp.:99.50, PWM:33
23:55:40.748 -> Water_Temp.:99.50, PWM:33
23:55:41.512 -> Water_Temp.:99.50, PWM:33
23:55:43.071 -> Water_Temp.:99.50, PWM:33
23:55:43.831 -> Water_Temp.:99.50, PWM:33
23:55:44.632 -> Water_Temp.:99.39, PWM:33
23:55:45.390 -> Water_Temp.:99.39, PWM:33
23:55:46.186 -> Water_Temp.:99.39, PWM:33
23:55:46.949 -> Water_Temp.:99.39, PWM:33
23:55:47.747 -> Water_Temp.:99.39, PWM:33
23:55:48.511 -> Water_Temp.:99.39, PWM:33
23:55:49.312 -> Water_Temp.:99.39, PWM:33
23:55:50.870 -> Water_Temp.:99.39, PWM:33
23:55:52.402 -> Water_Temp.:99.39, PWM:33
23:55:56.213 -> Water_Temp.:99.39, PWM:33
23:55:59.325 -> Water_Temp.:99.39, PWM:33
23:56:00.123 -> Water_Temp.:99.39, PWM:33
23:56:00.884 -> Water_Temp.:99.39, PWM:33
23:56:01.679 -> Water_Temp.:99.39, PWM:33
23:56:02.441 -> Water_Temp.:99.39, PWM:33
23:56:03.209 -> Water_Temp.:99.39, PWM:33
23:56:04.004 -> Water_Temp.:99.39, PWM:33
23:56:04.770 -> Water_Temp.:99.39, PWM:33
23:56:05.559 -> Water_Temp.:99.39, PWM:33
23:56:06.358 -> Water_Temp.:99.39, PWM:33
23:56:07.889 -> Water_Temp.:99.39, PWM:33
23:56:08.685 -> Water_Temp.:99.39, PWM:33
23:56:11.738 -> Water_Temp.:99.39, PWM:33
23:56:12.503 -> Water_Temp.:99.39, PWM:33
23:56:13.302 -> Water_Temp.:99.39, PWM:33
23:56:14.863 -> Water_Temp.:99.39, PWM:33
23:56:16.391 -> Water_Temp.:99.39, PWM:33
23:56:17.155 -> Water_Temp.:99.50, PWM:33
23:56:17.955 -> Water_Temp.:99.39, PWM:33
23:56:18.716 -> Water_Temp.:99.39, PWM:33
23:56:19.513 -> Water_Temp.:99.39, PWM:33
23:56:20.276 -> Water_Temp.:99.50, PWM:33
23:56:21.075 -> Water_Temp.:99.50, PWM:33
23:56:22.637 -> Water_Temp.:99.50, PWM:33
23:56:24.166 -> Water_Temp.:99.50, PWM:33
23:56:24.958 -> Water_Temp.:99.50, PWM:33
23:56:25.722 -> Water_Temp.:99.50, PWM:33
23:56:26.488 -> Water_Temp.:99.50, PWM:33
23:56:27.283 -> Water_Temp.:99.50, PWM:33
23:56:28.046 -> Water_Temp.:99.50, PWM:33
23:56:28.844 -> Water_Temp.:99.50, PWM:33
23:56:32.650 -> Water_Temp.:99.50, PWM:33
23:56:33.449 -> Water_Temp.:99.50, PWM:33
23:56:34.215 -> Water_Temp.:99.50, PWM:33
23:56:35.014 -> Water_Temp.:99.50, PWM:33
23:56:35.779 -> Water_Temp.:99.50, PWM:33
23:56:36.575 -> Water_Temp.:99.39, PWM:33
23:56:37.336 -> Water_Temp.:99.39, PWM:33
23:56:38.135 -> Water_Temp.:99.39, PWM:33
23:56:38.898 -> Water_Temp.:99.39, PWM:33
23:56:40.456 -> Water_Temp.:99.39, PWM:33
23:56:41.222 -> Water_Temp.:99.39, PWM:33
23:56:42.780 -> Water_Temp.:99.39, PWM:33
23:56:43.544 -> Water_Temp.:99.39, PWM:33
23:56:45.107 -> Water_Temp.:99.27, PWM:33
23:56:45.872 -> Water_Temp.:99.39, PWM:33
23:56:46.665 -> Water_Temp.:99.39, PWM:33
23:56:47.457 -> Water_Temp.:99.39, PWM:33
23:56:48.217 -> Water_Temp.:99.39, PWM:33
23:56:49.012 -> Water_Temp.:99.39, PWM:33
23:56:49.773 -> Water_Temp.:99.39, PWM:33
23:56:52.093 -> Water_Temp.:99.50, PWM:33
23:56:52.855 -> Water_Temp.:99.50, PWM:33
23:56:53.652 -> Water_Temp.:99.50, PWM:33
23:56:54.415 -> Water_Temp.:99.50, PWM:33
23:56:58.259 -> Water_Temp.:99.61, PWM:33
23:56:59.057 -> Water_Temp.:99.61, PWM:33
23:56:59.813 -> Water_Temp.:99.61, PWM:33
23:57:00.604 -> Water_Temp.:99.61, PWM:33
23:57:01.400 -> Water_Temp.:99.61, PWM:33
23:57:02.161 -> Water_Temp.:99.61, PWM:33
23:57:02.956 -> Water_Temp.:99.72, PWM:33
23:57:03.719 -> Water_Temp.:99.72, PWM:33
23:57:04.514 -> Water_Temp.:99.72, PWM:33
23:57:06.823 -> Water_Temp.:99.72, PWM:33
23:57:07.618 -> Water_Temp.:99.61, PWM:33
23:57:08.382 -> Water_Temp.:99.61, PWM:33
23:57:09.179 -> Water_Temp.:99.61, PWM:33
23:57:09.942 -> Water_Temp.:99.61, PWM:33
23:57:10.736 -> Water_Temp.:99.61, PWM:33
23:57:12.260 -> Water_Temp.:99.61, PWM:33
23:57:13.811 -> Water_Temp.:99.61, PWM:33
23:57:14.575 -> Water_Temp.:99.61, PWM:33
23:57:15.371 -> Water_Temp.:99.61, PWM:33
23:57:16.134 -> Water_Temp.:99.61, PWM:33
23:57:16.931 -> Water_Temp.:99.61, PWM:33
23:57:19.252 -> Water_Temp.:99.61, PWM:33
23:57:20.049 -> Water_Temp.:99.61, PWM:33
23:57:20.812 -> Water_Temp.:99.61, PWM:33
23:57:23.131 -> Water_Temp.:99.61, PWM:33
23:57:23.931 -> Water_Temp.:99.50, PWM:33
23:57:24.695 -> Water_Temp.:99.50, PWM:33
23:57:25.491 -> Water_Temp.:99.50, PWM:33
23:57:26.252 -> Water_Temp.:99.50, PWM:33
23:57:27.810 -> Water_Temp.:99.50, PWM:33
23:57:28.569 -> Water_Temp.:99.50, PWM:33
23:57:29.367 -> Water_Temp.:99.50, PWM:33
23:57:30.128 -> Water_Temp.:99.50, PWM:33
23:57:33.212 -> Water_Temp.:99.50, PWM:33
23:57:33.972 -> Water_Temp.:99.50, PWM:33
23:57:34.771 -> Water_Temp.:99.50, PWM:33
23:57:35.536 -> Water_Temp.:99.50, PWM:33
23:57:36.334 -> Water_Temp.:99.50, PWM:33
23:57:37.095 -> Water_Temp.:99.50, PWM:33
23:57:38.652 -> Water_Temp.:99.50, PWM:33
23:57:40.209 -> Water_Temp.:99.50, PWM:33
23:57:40.971 -> Water_Temp.:99.61, PWM:33
23:57:41.764 -> Water_Temp.:99.61, PWM:33
23:57:43.320 -> Water_Temp.:99.61, PWM:33
23:57:44.118 -> Water_Temp.:99.72, PWM:33
23:57:44.880 -> Water_Temp.:99.72, PWM:33
23:57:46.432 -> Water_Temp.:99.72, PWM:33
23:57:47.229 -> Water_Temp.:99.72, PWM:33
23:57:48.750 -> Water_Temp.:99.72, PWM:33
23:57:49.549 -> Water_Temp.:99.72, PWM:33
23:57:50.307 -> Water_Temp.:99.72, PWM:33
23:57:51.102 -> Water_Temp.:99.72, PWM:33
23:57:52.661 -> Water_Temp.:99.72, PWM:33
23:57:53.420 -> Water_Temp.:99.72, PWM:33
23:57:54.216 -> Water_Temp.:99.72, PWM:33
23:57:54.975 -> Water_Temp.:99.72, PWM:33
23:57:57.300 -> Water_Temp.:99.72, PWM:33
23:57:58.861 -> Water_Temp.:99.72, PWM:33
23:57:59.622 -> Water_Temp.:99.84, PWM:33
23:58:00.419 -> Water_Temp.:99.84, PWM:33
23:58:01.181 -> Water_Temp.:99.84, PWM:33
23:58:01.974 -> Water_Temp.:99.84, PWM:33
23:58:02.738 -> Water_Temp.:99.84, PWM:33
23:58:03.538 -> Water_Temp.:99.84, PWM:33
23:58:04.302 -> Water_Temp.:99.84, PWM:33
23:58:05.858 -> Water_Temp.:99.72, PWM:33
23:58:06.618 -> Water_Temp.:99.72, PWM:33
23:58:07.411 -> Water_Temp.:99.72, PWM:33
23:58:08.174 -> Water_Temp.:99.72, PWM:33
23:58:09.732 -> Water_Temp.:99.84, PWM:33
23:58:10.530 -> Water_Temp.:99.84, PWM:33
23:58:11.291 -> Water_Temp.:99.84, PWM:33
23:58:12.086 -> Water_Temp.:99.72, PWM:33
23:58:12.846 -> Water_Temp.:99.72, PWM:33
23:58:14.406 -> Water_Temp.:99.72, PWM:33
23:58:15.164 -> Water_Temp.:99.72, PWM:33
23:58:15.959 -> Water_Temp.:99.61, PWM:33
23:58:16.721 -> Water_Temp.:101.30, PWM:33
23:58:17.514 -> Water_Temp.:103.32, PWM:33
23:58:18.312 -> Water_Temp.:100.96, PWM:33
23:58:19.076 -> Water_Temp.:100.96, PWM:33
23:58:19.874 -> Water_Temp.:100.29, PWM:33
23:58:21.400 -> Water_Temp.:100.06, PWM:33
23:58:22.196 -> Water_Temp.:99.95, PWM:33
23:58:23.719 -> Water_Temp.:99.84, PWM:33
23:58:24.518 -> Water_Temp.:99.84, PWM:33
23:58:25.278 -> Water_Temp.:99.84, PWM:33
23:58:26.077 -> Water_Temp.:99.84, PWM:33
23:58:28.393 -> Water_Temp.:99.84, PWM:33
23:58:29.156 -> Water_Temp.:99.84, PWM:33
23:58:30.720 -> Water_Temp.:99.84, PWM:33
23:58:32.241 -> Water_Temp.:99.84, PWM:33
23:58:33.035 -> Water_Temp.:99.84, PWM:33
23:58:34.592 -> Water_Temp.:99.84, PWM:33
23:58:35.386 -> Water_Temp.:99.84, PWM:33
23:58:36.145 -> Water_Temp.:99.84, PWM:33
23:58:36.942 -> Water_Temp.:99.84, PWM:33
23:58:37.709 -> Water_Temp.:99.84, PWM:33
23:58:38.505 -> Water_Temp.:99.84, PWM:33
23:58:39.273 -> Water_Temp.:99.84, PWM:33
23:58:40.066 -> Water_Temp.:99.72, PWM:33
23:58:42.353 -> Water_Temp.:99.72, PWM:33
23:58:43.112 -> Water_Temp.:99.72, PWM:33
23:58:44.673 -> Water_Temp.:99.72, PWM:33
23:58:45.472 -> Water_Temp.:99.72, PWM:33
23:58:46.236 -> Water_Temp.:99.72, PWM:33
23:58:47.032 -> Water_Temp.:99.84, PWM:33
23:58:50.881 -> Water_Temp.:99.84, PWM:33
23:58:52.440 -> Water_Temp.:99.72, PWM:33
23:58:53.239 -> Water_Temp.:99.72, PWM:33
23:58:54.001 -> Water_Temp.:99.72, PWM:33
23:58:54.799 -> Water_Temp.:99.72, PWM:33
23:58:57.916 -> Water_Temp.:99.72, PWM:33
23:58:59.435 -> Water_Temp.:99.72, PWM:33
23:59:00.228 -> Water_Temp.:99.72, PWM:33
23:59:04.829 -> Water_Temp.:99.61, PWM:33
23:59:05.627 -> Water_Temp.:99.61, PWM:33
23:59:06.387 -> Water_Temp.:99.72, PWM:33
23:59:07.977 -> Water_Temp.:99.72, PWM:33
23:59:08.741 -> Water_Temp.:99.72, PWM:33
23:59:09.537 -> Water_Temp.:99.72, PWM:33
23:59:10.301 -> Water_Temp.:99.72, PWM:33
23:59:11.099 -> Water_Temp.:99.72, PWM:33
23:59:11.863 -> Water_Temp.:99.72, PWM:33
23:59:12.629 -> Water_Temp.:99.72, PWM:33
23:59:13.426 -> Water_Temp.:99.72, PWM:33
23:59:14.191 -> Water_Temp.:99.61, PWM:33
23:59:14.986 -> Water_Temp.:99.61, PWM:33
23:59:15.751 -> Water_Temp.:99.61, PWM:33
23:59:16.546 -> Water_Temp.:99.72, PWM:33
23:59:18.069 -> Water_Temp.:99.72, PWM:33
23:59:18.863 -> Water_Temp.:99.84, PWM:33
23:59:20.385 -> Water_Temp.:99.72, PWM:33
23:59:21.182 -> Water_Temp.:99.72, PWM:33
23:59:21.942 -> Water_Temp.:99.84, PWM:33
23:59:22.737 -> Water_Temp.:99.84, PWM:33
23:59:24.256 -> Water_Temp.:99.72, PWM:33
23:59:25.812 -> Water_Temp.:99.72, PWM:33
23:59:28.166 -> Water_Temp.:99.72, PWM:33
23:59:28.928 -> Water_Temp.:99.72, PWM:33
23:59:29.729 -> Water_Temp.:99.72, PWM:33
23:59:32.047 -> Water_Temp.:99.72, PWM:33
23:59:38.241 -> Water_Temp.:101.52, PWM:33
23:59:39.007 -> Water_Temp.:103.66, PWM:33
23:59:40.567 -> Water_Temp.:100.85, PWM:33
23:59:41.363 -> Water_Temp.:100.62, PWM:33
23:59:42.125 -> Water_Temp.:100.51, PWM:33
23:59:42.923 -> Water_Temp.:100.40, PWM:33
23:59:44.487 -> Water_Temp.:100.17, PWM:33
23:59:45.249 -> Water_Temp.:100.17, PWM:33
23:59:46.047 -> Water_Temp.:100.06, PWM:33
23:59:47.569 -> Water_Temp.:100.06, PWM:33
23:59:49.125 -> Water_Temp.:100.06, PWM:33
23:59:49.918 -> Water_Temp.:100.06, PWM:33
23:59:51.478 -> Water_Temp.:100.06, PWM:33
23:59:52.240 -> Water_Temp.:100.06, PWM:33
23:59:53.768 -> Water_Temp.:100.06, PWM:33
23:59:54.568 -> Water_Temp.:100.06, PWM:33
23:59:55.333 -> Water_Temp.:100.06, PWM:33
23:59:56.128 -> Water_Temp.:100.06, PWM:33
23:59:58.416 -> Water_Temp.:100.06, PWM:33
23:59:59.179 -> Water_Temp.:100.06, PWM:33
23:59:59.977 -> Water_Temp.:100.06, PWM:33
00:00:00.742 -> Water_Temp.:100.06, PWM:33
00:00:01.539 -> Water_Temp.:100.06, PWM:33
00:00:04.589 -> Water_Temp.:100.06, PWM:33
00:00:06.145 -> Water_Temp.:100.06, PWM:33
00:00:06.941 -> Water_Temp.:100.06, PWM:33
00:00:07.702 -> Water_Temp.:100.06, PWM:33
00:00:09.259 -> Water_Temp.:100.06, PWM:33
00:00:10.054 -> Water_Temp.:100.06, PWM:33
00:00:11.604 -> Water_Temp.:100.06, PWM:33
00:00:12.401 -> Water_Temp.:100.06, PWM:33
00:00:13.160 -> Water_Temp.:100.06, PWM:33
00:00:15.511 -> Water_Temp.:100.06, PWM:33
00:00:16.269 -> Water_Temp.:100.06, PWM:33
00:00:19.378 -> Water_Temp.:100.06, PWM:33
00:00:20.173 -> Water_Temp.:100.06, PWM:33
00:00:20.968 -> Water_Temp.:100.06, PWM:33
00:00:21.733 -> Water_Temp.:100.06, PWM:33
00:00:23.260 -> Water_Temp.:100.06, PWM:33
00:00:24.056 -> Water_Temp.:100.06, PWM:33
00:00:24.850 -> Water_Temp.:100.06, PWM:33
00:00:25.612 -> Water_Temp.:100.06, PWM:33
00:00:26.404 -> Water_Temp.:100.06, PWM:33
00:00:27.168 -> Water_Temp.:100.06, PWM:33
00:00:28.728 -> Water_Temp.:100.06, PWM:33
00:00:30.289 -> Water_Temp.:100.06, PWM:33
00:00:31.845 -> Water_Temp.:100.06, PWM:33
00:00:32.603 -> Water_Temp.:100.06, PWM:33
00:00:33.391 -> Water_Temp.:100.06, PWM:33
00:00:34.185 -> Water_Temp.:100.06, PWM:33
00:00:36.501 -> Water_Temp.:100.06, PWM:33
00:00:37.264 -> Water_Temp.:100.06, PWM:33
00:00:38.028 -> Water_Temp.:100.06, PWM:33
00:00:38.826 -> Water_Temp.:100.06, PWM:33
00:00:39.588 -> Water_Temp.:100.06, PWM:33
00:00:40.386 -> Water_Temp.:100.06, PWM:33
00:00:41.149 -> Water_Temp.:100.06, PWM:33
00:00:41.940 -> Water_Temp.:99.95, PWM:33
00:00:43.493 -> Water_Temp.:99.95, PWM:33
00:00:44.256 -> Water_Temp.:99.95, PWM:33
00:00:45.055 -> Water_Temp.:99.95, PWM:33
00:00:45.817 -> Water_Temp.:99.95, PWM:33
00:00:47.339 -> Water_Temp.:99.95, PWM:33
00:00:48.137 -> Water_Temp.:99.95, PWM:33
00:00:49.667 -> Water_Temp.:99.95, PWM:33
00:00:51.992 -> Water_Temp.:99.95, PWM:33
00:00:54.345 -> Water_Temp.:99.95, PWM:33
00:00:55.139 -> Water_Temp.:99.95, PWM:33
00:00:55.903 -> Water_Temp.:99.95, PWM:33
00:00:56.696 -> Water_Temp.:99.95, PWM:33
00:00:57.458 -> Water_Temp.:99.95, PWM:33
00:00:58.256 -> Water_Temp.:99.95, PWM:33
00:00:59.018 -> Water_Temp.:99.95, PWM:33
00:00:59.813 -> Water_Temp.:99.95, PWM:33
00:01:02.136 -> Water_Temp.:99.95, PWM:33
00:01:04.456 -> Water_Temp.:100.06, PWM:33
00:01:06.010 -> Water_Temp.:99.95, PWM:33
00:01:06.771 -> Water_Temp.:99.95, PWM:33
00:01:08.330 -> Water_Temp.:99.95, PWM:33
00:01:09.858 -> Water_Temp.:99.95, PWM:33
00:01:10.655 -> Water_Temp.:99.95, PWM:33
00:01:13.723 -> Water_Temp.:99.95, PWM:33
00:01:14.487 -> Water_Temp.:99.95, PWM:33
00:01:16.829 -> Water_Temp.:99.95, PWM:33
00:01:17.593 -> Water_Temp.:99.95, PWM:33
00:01:18.383 -> Water_Temp.:99.95, PWM:33
00:01:19.149 -> Water_Temp.:99.95, PWM:33
00:01:19.944 -> Water_Temp.:99.95, PWM:33
00:01:20.740 -> Water_Temp.:99.95, PWM:33
00:01:21.498 -> Water_Temp.:99.95, PWM:33
00:01:23.055 -> Water_Temp.:101.64, PWM:33
00:01:24.613 -> Water_Temp.:101.30, PWM:33
00:01:25.409 -> Water_Temp.:100.85, PWM:33
00:01:26.166 -> Water_Temp.:100.62, PWM:33
00:01:26.960 -> Water_Temp.:100.51, PWM:33
00:01:29.277 -> Water_Temp.:100.29, PWM:33
00:01:31.561 -> Water_Temp.:100.29, PWM:33
00:01:32.355 -> Water_Temp.:100.29, PWM:33
00:01:33.152 -> Water_Temp.:100.29, PWM:33
00:01:33.913 -> Water_Temp.:100.29, PWM:33
00:01:34.712 -> Water_Temp.:100.17, PWM:33
00:01:37.827 -> Water_Temp.:100.17, PWM:33
00:01:39.353 -> Water_Temp.:100.06, PWM:33
00:01:40.115 -> Water_Temp.:100.06, PWM:33
00:01:40.913 -> Water_Temp.:100.06, PWM:33
00:01:41.676 -> Water_Temp.:100.06, PWM:33
00:01:42.468 -> Water_Temp.:100.06, PWM:33
00:01:44.786 -> Water_Temp.:100.06, PWM:33
00:01:46.342 -> Water_Temp.:99.95, PWM:33
00:01:48.670 -> Water_Temp.:99.95, PWM:33
00:01:49.436 -> Water_Temp.:99.95, PWM:33
00:01:50.235 -> Water_Temp.:99.95, PWM:33
00:01:50.998 -> Water_Temp.:99.95, PWM:33
00:01:51.793 -> Water_Temp.:99.84, PWM:33
00:01:53.354 -> Water_Temp.:99.84, PWM:33
00:01:57.198 -> Water_Temp.:99.84, PWM:33
00:01:57.957 -> Water_Temp.:99.84, PWM:33
00:01:58.755 -> Water_Temp.:99.84, PWM:33
00:01:59.519 -> Water_Temp.:99.84, PWM:33
00:02:02.633 -> Water_Temp.:99.84, PWM:33
00:02:04.951 -> Water_Temp.:99.84, PWM:33
00:02:05.747 -> Water_Temp.:99.84, PWM:33
00:02:06.510 -> Water_Temp.:99.84, PWM:33
00:02:08.799 -> Water_Temp.:99.95, PWM:33
00:02:10.363 -> Water_Temp.:99.95, PWM:33
00:02:11.921 -> Water_Temp.:99.95, PWM:33
00:02:13.476 -> Water_Temp.:99.95, PWM:33
00:02:15.034 -> Water_Temp.:99.95, PWM:33
00:02:15.795 -> Water_Temp.:100.06, PWM:33
00:02:16.590 -> Water_Temp.:99.95, PWM:33
00:02:18.153 -> Water_Temp.:99.95, PWM:33
00:02:19.718 -> Water_Temp.:99.95, PWM:33
00:02:21.240 -> Water_Temp.:99.95, PWM:33
00:02:22.035 -> Water_Temp.:99.95, PWM:33
00:02:22.795 -> Water_Temp.:99.95, PWM:33
00:02:23.588 -> Water_Temp.:99.95, PWM:33
00:02:25.116 -> Water_Temp.:99.95, PWM:33
00:02:25.911 -> Water_Temp.:99.95, PWM:33
00:02:26.706 -> Water_Temp.:99.95, PWM:33
00:02:27.470 -> Water_Temp.:99.95, PWM:33
00:02:29.032 -> Water_Temp.:99.95, PWM:33
00:02:30.556 -> Water_Temp.:99.84, PWM:33
00:02:31.350 -> Water_Temp.:99.84, PWM:33
00:02:35.171 -> Water_Temp.:99.84, PWM:33
00:02:35.968 -> Water_Temp.:99.84, PWM:33
00:02:39.059 -> Water_Temp.:99.84, PWM:33
00:02:40.581 -> Water_Temp.:99.84, PWM:33
00:02:42.900 -> Water_Temp.:99.84, PWM:33
00:02:44.462 -> Water_Temp.:99.95, PWM:33
00:02:45.223 -> Water_Temp.:99.95, PWM:33
00:02:46.019 -> Water_Temp.:99.95, PWM:33
00:02:47.581 -> Water_Temp.:99.95, PWM:33
00:02:49.147 -> Water_Temp.:99.95, PWM:33
00:02:49.911 -> Water_Temp.:99.84, PWM:33
00:02:50.674 -> Water_Temp.:99.84, PWM:33
00:02:51.473 -> Water_Temp.:99.84, PWM:33
00:02:53.028 -> Water_Temp.:99.84, PWM:33
00:02:53.793 -> Water_Temp.:99.84, PWM:33
00:02:54.593 -> Water_Temp.:99.84, PWM:33
00:02:56.118 -> Water_Temp.:99.84, PWM:33
00:02:56.882 -> Water_Temp.:99.84, PWM:33
00:02:57.680 -> Water_Temp.:99.95, PWM:33
00:02:58.441 -> Water_Temp.:99.84, PWM:33
00:03:00.006 -> Water_Temp.:99.84, PWM:33
00:03:02.329 -> Water_Temp.:99.84, PWM:33
00:03:03.094 -> Water_Temp.:99.95, PWM:33
00:03:03.892 -> Water_Temp.:99.95, PWM:33
00:03:04.657 -> Water_Temp.:99.95, PWM:33
00:03:06.217 -> Water_Temp.:99.84, PWM:33
00:03:07.011 -> Water_Temp.:99.84, PWM:33
00:03:07.777 -> Water_Temp.:99.84, PWM:33
00:03:08.544 -> Water_Temp.:99.84, PWM:33
00:03:10.870 -> Water_Temp.:99.84, PWM:33
00:03:11.669 -> Water_Temp.:99.84, PWM:33
00:03:12.436 -> Water_Temp.:99.84, PWM:33
00:03:13.997 -> Water_Temp.:99.84, PWM:33
00:03:15.527 -> Water_Temp.:99.84, PWM:33
00:03:17.857 -> Water_Temp.:99.72, PWM:33
00:03:18.648 -> Water_Temp.:99.72, PWM:33
00:03:19.443 -> Water_Temp.:99.72, PWM:33
00:03:20.204 -> Water_Temp.:99.84, PWM:33
00:03:21.728 -> Water_Temp.:99.72, PWM:33
00:03:22.524 -> Water_Temp.:99.84, PWM:33
00:03:24.845 -> Water_Temp.:99.84, PWM:33
00:03:25.644 -> Water_Temp.:99.84, PWM:33
00:03:26.438 -> Water_Temp.:99.84, PWM:33
00:03:27.200 -> Water_Temp.:99.84, PWM:33
00:03:28.756 -> Water_Temp.:99.84, PWM:33
00:03:29.551 -> Water_Temp.:99.84, PWM:33
00:03:30.308 -> Water_Temp.:99.84, PWM:33
00:03:31.103 -> Water_Temp.:99.84, PWM:33
00:03:34.952 -> Water_Temp.:99.84, PWM:33
00:03:35.716 -> Water_Temp.:99.84, PWM:33
00:03:37.276 -> Water_Temp.:99.84, PWM:33
00:03:38.032 -> Water_Temp.:99.84, PWM:33
00:03:38.830 -> Water_Temp.:99.84, PWM:33
00:03:40.388 -> Water_Temp.:99.84, PWM:33
00:03:41.148 -> Water_Temp.:99.84, PWM:33
00:03:41.944 -> Water_Temp.:99.84, PWM:33
00:03:43.503 -> Water_Temp.:99.84, PWM:33
00:03:44.268 -> Water_Temp.:99.84, PWM:33
00:03:45.822 -> Water_Temp.:99.84, PWM:33
00:03:47.384 -> Water_Temp.:99.84, PWM:33
00:03:48.148 -> Water_Temp.:99.84, PWM:33
00:03:48.945 -> Water_Temp.:99.84, PWM:33
00:03:51.228 -> Water_Temp.:99.84, PWM:33
00:03:52.787 -> Water_Temp.:99.84, PWM:33
00:03:53.547 -> Water_Temp.:99.72, PWM:33
00:03:55.112 -> Water_Temp.:99.72, PWM:33
00:03:55.876 -> Water_Temp.:99.72, PWM:33
00:03:56.673 -> Water_Temp.:99.72, PWM:33
00:03:57.437 -> Water_Temp.:99.72, PWM:33
00:03:58.999 -> Water_Temp.:99.72, PWM:33
00:04:00.560 -> Water_Temp.:99.72, PWM:33
00:04:01.358 -> Water_Temp.:99.72, PWM:33
00:04:02.120 -> Water_Temp.:99.72, PWM:33
00:04:02.916 -> Water_Temp.:99.72, PWM:33
00:04:03.680 -> Water_Temp.:99.72, PWM:33
00:04:05.229 -> Water_Temp.:99.72, PWM:33
00:04:06.022 -> Water_Temp.:99.72, PWM:33
00:04:06.785 -> Water_Temp.:99.72, PWM:33
00:04:07.578 -> Water_Temp.:99.72, PWM:33
00:04:08.373 -> Water_Temp.:99.72, PWM:33
00:04:09.136 -> Water_Temp.:99.72, PWM:33
00:04:12.216 -> Water_Temp.:99.72, PWM:33
00:04:13.014 -> Water_Temp.:99.72, PWM:33
00:04:13.774 -> Water_Temp.:99.72, PWM:33
00:04:15.333 -> Water_Temp.:99.72, PWM:33
00:04:16.091 -> Water_Temp.:99.84, PWM:33
00:04:17.644 -> Water_Temp.:99.84, PWM:33
00:04:19.207 -> Water_Temp.:99.84, PWM:33
00:04:19.972 -> Water_Temp.:99.84, PWM:33
00:04:20.771 -> Water_Temp.:99.84, PWM:33
00:04:24.611 -> Water_Temp.:99.95, PWM:33
00:04:25.404 -> Water_Temp.:99.95, PWM:33
00:04:26.168 -> Water_Temp.:99.95, PWM:33
00:04:26.963 -> Water_Temp.:99.95, PWM:33
00:04:29.283 -> Water_Temp.:99.84, PWM:33
00:04:30.045 -> Water_Temp.:99.84, PWM:33
00:04:31.600 -> Water_Temp.:99.84, PWM:33
00:04:32.400 -> Water_Temp.:99.84, PWM:33
00:04:33.162 -> Water_Temp.:99.84, PWM:33
00:04:33.961 -> Water_Temp.:99.84, PWM:33
00:04:35.518 -> Water_Temp.:99.84, PWM:33
00:04:36.279 -> Water_Temp.:99.84, PWM:33
00:04:37.075 -> Water_Temp.:99.84, PWM:33
00:04:40.892 -> Water_Temp.:99.95, PWM:33
00:04:43.212 -> Water_Temp.:99.95, PWM:33
00:04:44.739 -> Water_Temp.:100.06, PWM:33
00:04:45.533 -> Water_Temp.:100.06, PWM:33
00:04:47.091 -> Water_Temp.:100.06, PWM:33
00:04:49.420 -> Water_Temp.:100.06, PWM:33
00:04:50.184 -> Water_Temp.:100.06, PWM:33
00:04:50.982 -> Water_Temp.:100.06, PWM:33
00:04:51.776 -> Water_Temp.:100.06, PWM:33
00:04:52.535 -> Water_Temp.:100.06, PWM:33
00:04:53.331 -> Water_Temp.:100.06, PWM:33
00:04:54.092 -> Water_Temp.:100.06, PWM:33
00:04:55.647 -> Water_Temp.:100.06, PWM:33
00:04:56.438 -> Water_Temp.:100.06, PWM:33
00:04:57.234 -> Water_Temp.:100.06, PWM:33
00:04:57.994 -> Water_Temp.:100.06, PWM:33
00:04:58.790 -> Water_Temp.:100.06, PWM:33
00:04:59.555 -> Water_Temp.:100.06, PWM:33
00:05:00.351 -> Water_Temp.:100.06, PWM:33
00:05:01.113 -> Water_Temp.:100.06, PWM:33
00:05:01.911 -> Water_Temp.:100.06, PWM:33
00:05:02.669 -> Water_Temp.:100.06, PWM:33
00:05:04.224 -> Water_Temp.:99.95, PWM:33
00:05:07.305 -> Water_Temp.:99.95, PWM:33
00:05:08.059 -> Water_Temp.:99.95, PWM:33
00:05:10.414 -> Water_Temp.:99.95, PWM:33
00:05:11.174 -> Water_Temp.:99.95, PWM:33
00:05:11.972 -> Water_Temp.:99.95, PWM:33
00:05:12.733 -> Water_Temp.:99.95, PWM:33
00:05:14.315 -> Water_Temp.:99.95, PWM:33
00:05:15.074 -> Water_Temp.:99.95, PWM:33
00:05:18.184 -> Water_Temp.:99.84, PWM:33
00:05:18.945 -> Water_Temp.:99.84, PWM:33
00:05:19.734 -> Water_Temp.:99.84, PWM:33
00:05:20.524 -> Water_Temp.:99.84, PWM:33
00:05:22.080 -> Water_Temp.:99.84, PWM:33
00:05:22.837 -> Water_Temp.:99.84, PWM:33
00:05:23.619 -> Water_Temp.:99.84, PWM:33
00:05:25.173 -> Water_Temp.:99.95, PWM:33
00:05:25.964 -> Water_Temp.:99.95, PWM:33
00:05:26.723 -> Water_Temp.:99.95, PWM:33
00:05:27.515 -> Water_Temp.:99.95, PWM:33
00:05:32.146 -> Water_Temp.:99.95, PWM:33
00:05:32.909 -> Water_Temp.:99.95, PWM:33
00:05:34.461 -> Water_Temp.:99.95, PWM:33
00:05:35.255 -> Water_Temp.:99.95, PWM:33
00:05:36.815 -> Water_Temp.:100.06, PWM:33
00:05:37.575 -> Water_Temp.:100.06, PWM:33
00:05:38.376 -> Water_Temp.:100.06, PWM:33
00:05:39.139 -> Water_Temp.:100.06, PWM:33
00:05:40.697 -> Water_Temp.:100.06, PWM:33
00:05:41.494 -> Water_Temp.:100.06, PWM:33
00:05:42.254 -> Water_Temp.:100.06, PWM:33
00:05:44.610 -> Water_Temp.:99.95, PWM:33
00:05:46.167 -> Water_Temp.:99.95, PWM:36
00:05:46.928 -> Water_Temp.:99.95, PWM:37
00:05:49.250 -> Water_Temp.:99.95, PWM:37
00:05:50.804 -> Water_Temp.:99.95, PWM:37
00:05:51.564 -> Water_Temp.:99.95, PWM:38
00:05:53.156 -> Water_Temp.:99.95, PWM:37
00:05:55.442 -> Water_Temp.:99.84, PWM:38
00:05:56.203 -> Water_Temp.:99.95, PWM:37
00:05:56.998 -> Water_Temp.:99.84, PWM:37
00:05:59.320 -> Water_Temp.:99.95, PWM:37
00:06:00.848 -> Water_Temp.:99.95, PWM:38
00:06:02.404 -> Water_Temp.:99.95, PWM:38
00:06:03.168 -> Water_Temp.:99.95, PWM:38
00:06:03.961 -> Water_Temp.:99.95, PWM:38
00:06:04.759 -> Water_Temp.:99.95, PWM:37
00:06:06.286 -> Water_Temp.:99.95, PWM:37
00:06:07.050 -> Water_Temp.:99.95, PWM:37
00:06:08.610 -> Water_Temp.:99.95, PWM:38
00:06:10.134 -> Water_Temp.:99.95, PWM:37
00:06:10.930 -> Water_Temp.:99.95, PWM:37
00:06:11.691 -> Water_Temp.:99.95, PWM:38
00:06:12.485 -> Water_Temp.:99.95, PWM:37
00:06:13.245 -> Water_Temp.:99.95, PWM:38
00:06:14.040 -> Water_Temp.:100.06, PWM:37
00:06:14.835 -> Water_Temp.:100.06, PWM:38
00:06:15.596 -> Water_Temp.:100.06, PWM:37
00:06:16.395 -> Water_Temp.:100.06, PWM:37
00:06:17.952 -> Water_Temp.:100.17, PWM:37
00:06:18.710 -> Water_Temp.:100.17, PWM:38
00:06:19.508 -> Water_Temp.:100.17, PWM:37
00:06:20.270 -> Water_Temp.:100.17, PWM:38
00:06:21.067 -> Water_Temp.:100.17, PWM:37
00:06:21.829 -> Water_Temp.:100.17, PWM:38
00:06:23.387 -> Water_Temp.:100.17, PWM:37
00:06:25.703 -> Water_Temp.:100.17, PWM:37
00:06:26.461 -> Water_Temp.:100.17, PWM:37
00:06:27.252 -> Water_Temp.:100.17, PWM:37
00:06:30.325 -> Water_Temp.:100.17, PWM:37
00:06:31.897 -> Water_Temp.:100.17, PWM:37
00:06:32.659 -> Water_Temp.:100.17, PWM:37
00:06:33.454 -> Water_Temp.:100.17, PWM:37
00:06:34.250 -> Water_Temp.:100.17, PWM:37
00:06:35.012 -> Water_Temp.:100.17, PWM:37
00:06:35.809 -> Water_Temp.:100.17, PWM:37
00:06:36.570 -> Water_Temp.:100.17, PWM:38
00:06:38.895 -> Water_Temp.:100.06, PWM:38
00:06:39.659 -> Water_Temp.:100.06, PWM:37
00:06:40.452 -> Water_Temp.:100.06, PWM:38
00:06:42.011 -> Water_Temp.:100.06, PWM:38
00:06:43.568 -> Water_Temp.:100.06, PWM:38
00:06:44.363 -> Water_Temp.:100.06, PWM:38
00:06:45.884 -> Water_Temp.:100.06, PWM:37
00:06:47.442 -> Water_Temp.:100.06, PWM:37
00:06:48.201 -> Water_Temp.:100.06, PWM:37
00:06:48.998 -> Water_Temp.:100.06, PWM:38
00:06:50.553 -> Water_Temp.:100.06, PWM:38
00:06:51.311 -> Water_Temp.:100.17, PWM:37
00:06:52.111 -> Water_Temp.:100.17, PWM:37
00:06:52.873 -> Water_Temp.:100.17, PWM:38
00:06:55.228 -> Water_Temp.:100.17, PWM:37
00:06:55.989 -> Water_Temp.:100.17, PWM:37
00:06:57.543 -> Water_Temp.:100.17, PWM:37
00:06:58.337 -> Water_Temp.:100.17, PWM:38
00:06:59.099 -> Water_Temp.:100.17, PWM:38
00:07:00.655 -> Water_Temp.:100.17, PWM:37
00:07:01.418 -> Water_Temp.:100.29, PWM:37
00:07:02.215 -> Water_Temp.:100.29, PWM:37
00:07:02.981 -> Water_Temp.:100.29, PWM:37
00:07:03.778 -> Water_Temp.:100.29, PWM:38
00:07:04.545 -> Water_Temp.:100.29, PWM:38
00:07:05.344 -> Water_Temp.:100.40, PWM:37
00:07:06.109 -> Water_Temp.:100.40, PWM:38
00:07:06.875 -> Water_Temp.:100.40, PWM:37
00:07:09.199 -> Water_Temp.:100.40, PWM:38
00:07:09.995 -> Water_Temp.:100.40, PWM:38
00:07:10.757 -> Water_Temp.:100.40, PWM:38
00:07:11.551 -> Water_Temp.:100.40, PWM:37
00:07:13.113 -> Water_Temp.:100.40, PWM:38
00:07:13.876 -> Water_Temp.:100.40, PWM:37
00:07:14.676 -> Water_Temp.:100.40, PWM:37
00:07:15.439 -> Water_Temp.:100.40, PWM:38
00:07:16.236 -> Water_Temp.:100.40, PWM:37
00:07:16.996 -> Water_Temp.:100.40, PWM:37
00:07:18.561 -> Water_Temp.:100.40, PWM:37
00:07:20.120 -> Water_Temp.:100.29, PWM:37
00:07:20.879 -> Water_Temp.:100.40, PWM:38
00:07:21.670 -> Water_Temp.:100.29, PWM:37
00:07:22.470 -> Water_Temp.:100.40, PWM:37
00:07:23.234 -> Water_Temp.:100.40, PWM:37
00:07:24.797 -> Water_Temp.:100.40, PWM:37
00:07:26.328 -> Water_Temp.:100.40, PWM:37
00:07:27.125 -> Water_Temp.:100.40, PWM:37
00:07:27.887 -> Water_Temp.:100.40, PWM:37
00:07:28.682 -> Water_Temp.:100.40, PWM:37
00:07:29.443 -> Water_Temp.:100.40, PWM:38
00:07:30.237 -> Water_Temp.:100.40, PWM:38
00:07:30.997 -> Water_Temp.:100.40, PWM:38
00:07:33.347 -> Water_Temp.:100.40, PWM:37
00:07:34.907 -> Water_Temp.:100.29, PWM:38
00:07:36.463 -> Water_Temp.:100.29, PWM:37
00:07:37.226 -> Water_Temp.:100.29, PWM:37
00:07:38.018 -> Water_Temp.:100.29, PWM:38
00:07:38.817 -> Water_Temp.:100.40, PWM:37
00:07:40.370 -> Water_Temp.:100.40, PWM:37
00:07:41.132 -> Water_Temp.:100.40, PWM:38
00:07:41.929 -> Water_Temp.:100.40, PWM:38
00:07:42.692 -> Water_Temp.:100.40, PWM:38
00:07:43.487 -> Water_Temp.:100.40, PWM:37
00:07:46.571 -> Water_Temp.:100.40, PWM:37
00:07:47.334 -> Water_Temp.:100.40, PWM:38
00:07:48.130 -> Water_Temp.:100.40, PWM:37
00:07:48.893 -> Water_Temp.:100.40, PWM:37
00:07:49.690 -> Water_Temp.:100.29, PWM:37
00:07:52.009 -> Water_Temp.:100.29, PWM:37
00:07:53.571 -> Water_Temp.:100.40, PWM:38
00:07:54.372 -> Water_Temp.:100.40, PWM:37
00:07:55.136 -> Water_Temp.:100.40, PWM:38
00:07:55.931 -> Water_Temp.:100.40, PWM:38
00:07:56.689 -> Water_Temp.:100.40, PWM:37
00:07:57.491 -> Water_Temp.:100.40, PWM:37
00:07:59.045 -> Water_Temp.:100.40, PWM:37
00:08:00.598 -> Water_Temp.:100.40, PWM:37
00:08:01.360 -> Water_Temp.:100.40, PWM:37
00:08:02.155 -> Water_Temp.:102.09, PWM:38
00:08:02.918 -> Water_Temp.:104.22, PWM:37
00:08:03.715 -> Water_Temp.:101.75, PWM:37
00:08:04.479 -> Water_Temp.:101.30, PWM:37
00:08:06.008 -> Water_Temp.:101.07, PWM:37
00:08:06.806 -> Water_Temp.:100.96, PWM:37
00:08:07.602 -> Water_Temp.:100.96, PWM:37
00:08:08.357 -> Water_Temp.:100.85, PWM:37
00:08:09.918 -> Water_Temp.:100.85, PWM:37
00:08:10.714 -> Water_Temp.:100.85, PWM:37
00:08:11.472 -> Water_Temp.:100.85, PWM:37
00:08:12.272 -> Water_Temp.:100.85, PWM:37
00:08:13.036 -> Water_Temp.:100.85, PWM:37
00:08:15.359 -> Water_Temp.:100.74, PWM:37
00:08:16.918 -> Water_Temp.:100.74, PWM:37
00:08:18.441 -> Water_Temp.:100.74, PWM:38
00:08:20.004 -> Water_Temp.:100.62, PWM:37
00:08:20.796 -> Water_Temp.:100.62, PWM:37
00:08:21.552 -> Water_Temp.:100.62, PWM:38
00:08:22.343 -> Water_Temp.:100.62, PWM:38
00:08:23.138 -> Water_Temp.:100.62, PWM:37
00:08:23.900 -> Water_Temp.:100.62, PWM:37
00:08:24.700 -> Water_Temp.:100.62, PWM:37
00:08:25.465 -> Water_Temp.:100.62, PWM:37
00:08:26.228 -> Water_Temp.:100.62, PWM:37
00:08:27.025 -> Water_Temp.:100.62, PWM:37
00:08:28.551 -> Water_Temp.:100.62, PWM:38
00:08:29.351 -> Water_Temp.:100.62, PWM:38
00:08:30.108 -> Water_Temp.:102.20, PWM:38
00:08:30.898 -> Water_Temp.:104.34, PWM:37
00:08:31.687 -> Water_Temp.:101.97, PWM:37
00:08:32.446 -> Water_Temp.:101.52, PWM:38
00:08:34.007 -> Water_Temp.:101.19, PWM:38
00:08:35.567 -> Water_Temp.:101.07, PWM:37
00:08:36.367 -> Water_Temp.:100.96, PWM:37
00:08:37.929 -> Water_Temp.:100.85, PWM:38
00:08:38.696 -> Water_Temp.:100.85, PWM:37
00:08:41.781 -> Water_Temp.:100.74, PWM:37
00:08:42.543 -> Water_Temp.:100.74, PWM:38
00:08:43.337 -> Water_Temp.:100.62, PWM:37
00:08:44.102 -> Water_Temp.:100.62, PWM:37
00:08:44.865 -> Water_Temp.:100.62, PWM:37
00:08:45.665 -> Water_Temp.:100.62, PWM:37
00:08:46.428 -> Water_Temp.:100.62, PWM:37
00:08:47.222 -> Water_Temp.:100.62, PWM:37
00:08:48.016 -> Water_Temp.:100.62, PWM:37
00:08:48.778 -> Water_Temp.:100.62, PWM:37
00:08:50.309 -> Water_Temp.:100.62, PWM:37
00:08:51.104 -> Water_Temp.:100.51, PWM:38
00:08:51.862 -> Water_Temp.:100.51, PWM:38
00:08:52.661 -> Water_Temp.:100.51, PWM:37
00:08:54.220 -> Water_Temp.:100.51, PWM:37
00:08:54.979 -> Water_Temp.:100.51, PWM:38
00:08:55.772 -> Water_Temp.:100.51, PWM:37
00:08:57.333 -> Water_Temp.:100.51, PWM:37
00:08:59.647 -> Water_Temp.:100.51, PWM:37
00:09:00.440 -> Water_Temp.:100.51, PWM:37
00:09:01.234 -> Water_Temp.:100.51, PWM:37
00:09:02.789 -> Water_Temp.:100.51, PWM:37
00:09:03.552 -> Water_Temp.:100.51, PWM:38
00:09:04.351 -> Water_Temp.:100.40, PWM:37
00:09:05.115 -> Water_Temp.:100.40, PWM:37
00:09:06.666 -> Water_Temp.:100.40, PWM:37
00:09:07.461 -> Water_Temp.:100.40, PWM:37
00:09:08.227 -> Water_Temp.:100.40, PWM:37
00:09:09.025 -> Water_Temp.:100.40, PWM:37
00:09:11.343 -> Water_Temp.:100.40, PWM:37
00:09:12.893 -> Water_Temp.:100.51, PWM:38
00:09:15.216 -> Water_Temp.:100.51, PWM:37
00:09:15.973 -> Water_Temp.:100.51, PWM:38
00:09:16.771 -> Water_Temp.:100.51, PWM:37
00:09:17.534 -> Water_Temp.:100.51, PWM:37
00:09:18.326 -> Water_Temp.:100.51, PWM:38
00:09:19.089 -> Water_Temp.:100.51, PWM:37
00:09:25.295 -> Water_Temp.:100.51, PWM:37
00:09:26.060 -> Water_Temp.:100.51, PWM:38
00:09:27.617 -> Water_Temp.:100.51, PWM:38
00:09:28.410 -> Water_Temp.:100.51, PWM:37
00:09:29.175 -> Water_Temp.:100.51, PWM:38
00:09:29.964 -> Water_Temp.:100.51, PWM:37
00:09:30.727 -> Water_Temp.:100.51, PWM:38
00:09:34.601 -> Water_Temp.:100.51, PWM:37
00:09:35.365 -> Water_Temp.:100.51, PWM:38
00:09:36.923 -> Water_Temp.:100.51, PWM:37
00:09:37.722 -> Water_Temp.:100.51, PWM:37
00:09:38.482 -> Water_Temp.:100.62, PWM:38
00:09:39.281 -> Water_Temp.:100.62, PWM:37
00:09:40.040 -> Water_Temp.:100.51, PWM:38
00:09:40.835 -> Water_Temp.:100.62, PWM:37
00:09:41.600 -> Water_Temp.:100.62, PWM:38
00:09:42.394 -> Water_Temp.:100.51, PWM:37
00:09:43.159 -> Water_Temp.:100.51, PWM:37
00:09:43.955 -> Water_Temp.:100.51, PWM:38
00:09:45.481 -> Water_Temp.:100.51, PWM:38
00:09:46.278 -> Water_Temp.:100.51, PWM:38
00:09:48.595 -> Water_Temp.:100.51, PWM:38
00:09:50.151 -> Water_Temp.:100.62, PWM:38
00:09:50.947 -> Water_Temp.:100.62, PWM:37
00:09:54.033 -> Water_Temp.:100.51, PWM:37
00:09:54.828 -> Water_Temp.:100.51, PWM:38
00:09:55.591 -> Water_Temp.:100.51, PWM:37
00:09:56.385 -> Water_Temp.:100.51, PWM:38
00:09:57.906 -> Water_Temp.:100.51, PWM:37
00:09:58.699 -> Water_Temp.:100.51, PWM:37
00:10:00.246 -> Water_Temp.:100.51, PWM:37
00:10:02.533 -> Water_Temp.:100.51, PWM:37
00:10:03.295 -> Water_Temp.:100.51, PWM:37
00:10:04.092 -> Water_Temp.:100.51, PWM:37
00:10:04.849 -> Water_Temp.:100.51, PWM:37
00:10:06.436 -> Water_Temp.:100.51, PWM:38
00:10:07.197 -> Water_Temp.:100.51, PWM:37
00:10:07.992 -> Water_Temp.:100.51, PWM:37
00:10:08.755 -> Water_Temp.:100.51, PWM:37
00:10:10.311 -> Water_Temp.:100.51, PWM:37
00:10:11.109 -> Water_Temp.:100.51, PWM:37
00:10:11.867 -> Water_Temp.:100.51, PWM:37
00:10:12.657 -> Water_Temp.:100.51, PWM:37
00:10:13.451 -> Water_Temp.:100.51, PWM:37
00:10:14.215 -> Water_Temp.:100.51, PWM:37
00:10:15.767 -> Water_Temp.:100.51, PWM:37
00:10:16.532 -> Water_Temp.:100.51, PWM:38
00:10:17.326 -> Water_Temp.:100.51, PWM:38
00:10:18.876 -> Water_Temp.:100.51, PWM:37
00:10:19.673 -> Water_Temp.:100.51, PWM:38
00:10:21.230 -> Water_Temp.:100.62, PWM:37
00:10:21.992 -> Water_Temp.:100.62, PWM:38
00:10:22.786 -> Water_Temp.:100.62, PWM:37
00:10:23.551 -> Water_Temp.:100.62, PWM:37
00:10:24.349 -> Water_Temp.:100.62, PWM:37
00:10:25.908 -> Water_Temp.:100.62, PWM:37
00:10:26.673 -> Water_Temp.:100.62, PWM:37
00:10:27.433 -> Water_Temp.:100.74, PWM:37
00:10:28.224 -> Water_Temp.:100.74, PWM:38
00:10:30.507 -> Water_Temp.:100.62, PWM:37
00:10:31.293 -> Water_Temp.:100.62, PWM:37
00:10:32.090 -> Water_Temp.:100.62, PWM:38
00:10:32.850 -> Water_Temp.:100.62, PWM:37
00:10:35.172 -> Water_Temp.:100.62, PWM:37
00:10:35.971 -> Water_Temp.:100.62, PWM:37
00:10:36.731 -> Water_Temp.:100.62, PWM:38
00:10:38.289 -> Water_Temp.:100.62, PWM:37
00:10:39.850 -> Water_Temp.:100.62, PWM:37
00:10:40.648 -> Water_Temp.:100.62, PWM:37
00:10:41.412 -> Water_Temp.:100.62, PWM:37
00:10:42.206 -> Water_Temp.:100.74, PWM:38
00:10:43.720 -> Water_Temp.:100.62, PWM:38
00:10:44.516 -> Water_Temp.:100.62, PWM:37
00:10:45.277 -> Water_Temp.:100.62, PWM:37
00:10:46.071 -> Water_Temp.:100.74, PWM:37
00:10:46.870 -> Water_Temp.:100.62, PWM:37
00:10:47.632 -> Water_Temp.:100.74, PWM:37
00:10:48.429 -> Water_Temp.:100.74, PWM:37
00:10:49.987 -> Water_Temp.:100.74, PWM:38
00:10:51.537 -> Water_Temp.:100.74, PWM:37
00:10:52.295 -> Water_Temp.:100.74, PWM:37
00:10:53.879 -> Water_Temp.:100.74, PWM:37
00:10:54.640 -> Water_Temp.:100.85, PWM:37
00:10:55.434 -> Water_Temp.:100.85, PWM:37
00:10:56.193 -> Water_Temp.:100.85, PWM:37
00:10:56.980 -> Water_Temp.:100.85, PWM:38
00:10:59.327 -> Water_Temp.:100.85, PWM:38
00:11:01.643 -> Water_Temp.:100.85, PWM:37
00:11:02.441 -> Water_Temp.:100.85, PWM:38
00:11:03.204 -> Water_Temp.:100.85, PWM:38
00:11:05.550 -> Water_Temp.:100.85, PWM:37
00:11:07.110 -> Water_Temp.:100.85, PWM:37
00:11:07.872 -> Water_Temp.:100.85, PWM:37
00:11:08.667 -> Water_Temp.:100.85, PWM:37
00:11:10.229 -> Water_Temp.:100.74, PWM:37
00:11:10.990 -> Water_Temp.:100.74, PWM:38
00:11:11.788 -> Water_Temp.:100.74, PWM:37
00:11:13.314 -> Water_Temp.:100.74, PWM:37
00:11:14.110 -> Water_Temp.:100.74, PWM:37
00:11:14.872 -> Water_Temp.:100.74, PWM:37
00:11:15.665 -> Water_Temp.:100.74, PWM:37
00:11:17.187 -> Water_Temp.:100.74, PWM:37
00:11:18.741 -> Water_Temp.:100.74, PWM:38
00:11:19.506 -> Water_Temp.:100.74, PWM:38
00:11:20.302 -> Water_Temp.:100.74, PWM:37
00:11:21.064 -> Water_Temp.:100.74, PWM:37
00:11:21.861 -> Water_Temp.:100.74, PWM:37
00:11:24.178 -> Water_Temp.:100.74, PWM:38
00:11:27.290 -> Water_Temp.:100.74, PWM:37
00:11:28.086 -> Water_Temp.:100.74, PWM:38
00:11:28.880 -> Water_Temp.:100.74, PWM:38
00:11:29.645 -> Water_Temp.:100.74, PWM:37
00:11:31.208 -> Water_Temp.:100.74, PWM:37
00:11:31.967 -> Water_Temp.:100.74, PWM:37
00:11:32.765 -> Water_Temp.:100.74, PWM:37
00:11:33.526 -> Water_Temp.:100.74, PWM:37
00:11:34.322 -> Water_Temp.:100.74, PWM:37
00:11:35.079 -> Water_Temp.:100.74, PWM:37
00:11:35.869 -> Water_Temp.:100.74, PWM:38
00:11:37.426 -> Water_Temp.:100.74, PWM:38
00:11:38.225 -> Water_Temp.:100.74, PWM:38
00:11:42.067 -> Water_Temp.:100.74, PWM:38
00:11:43.623 -> Water_Temp.:100.74, PWM:37
00:11:45.144 -> Water_Temp.:100.74, PWM:37
00:11:45.941 -> Water_Temp.:100.74, PWM:37
00:11:46.702 -> Water_Temp.:100.85, PWM:37
00:11:48.285 -> Water_Temp.:100.85, PWM:37
00:11:49.046 -> Water_Temp.:100.85, PWM:37
00:11:49.840 -> Water_Temp.:100.85, PWM:38
00:11:52.164 -> Water_Temp.:100.85, PWM:37
00:11:52.963 -> Water_Temp.:100.96, PWM:37
00:11:54.489 -> Water_Temp.:100.96, PWM:37
00:11:55.284 -> Water_Temp.:100.96, PWM:38
00:11:56.043 -> Water_Temp.:100.96, PWM:37
00:11:56.839 -> Water_Temp.:100.96, PWM:38
00:11:58.361 -> Water_Temp.:100.96, PWM:37
00:11:59.922 -> Water_Temp.:100.96, PWM:38
00:12:01.481 -> Water_Temp.:100.96, PWM:37
00:12:02.273 -> Water_Temp.:100.96, PWM:38
00:12:03.031 -> Water_Temp.:100.96, PWM:38
00:12:03.826 -> Water_Temp.:100.96, PWM:38
00:12:04.588 -> Water_Temp.:100.96, PWM:37
00:12:05.381 -> Water_Temp.:100.96, PWM:37
00:12:06.180 -> Water_Temp.:100.96, PWM:37
00:12:06.942 -> Water_Temp.:100.96, PWM:37
00:12:07.741 -> Water_Temp.:100.96, PWM:38
00:12:09.265 -> Water_Temp.:100.96, PWM:38
00:12:11.576 -> Water_Temp.:101.07, PWM:38
00:12:12.337 -> Water_Temp.:101.07, PWM:37
00:12:13.135 -> Water_Temp.:101.07, PWM:37
00:12:14.694 -> Water_Temp.:101.07, PWM:37
00:12:16.259 -> Water_Temp.:101.07, PWM:38
00:12:17.017 -> Water_Temp.:101.07, PWM:37
00:12:20.131 -> Water_Temp.:101.19, PWM:37
00:12:20.895 -> Water_Temp.:101.19, PWM:37
00:12:21.690 -> Water_Temp.:101.19, PWM:38
00:12:22.487 -> Water_Temp.:101.19, PWM:38
00:12:24.014 -> Water_Temp.:101.30, PWM:38
00:12:25.538 -> Water_Temp.:101.30, PWM:37
00:12:27.096 -> Water_Temp.:101.30, PWM:37
00:12:27.858 -> Water_Temp.:101.30, PWM:37
00:12:30.931 -> Water_Temp.:101.30, PWM:37
00:12:32.493 -> Water_Temp.:101.30, PWM:37
00:12:33.259 -> Water_Temp.:101.30, PWM:37
00:12:34.058 -> Water_Temp.:101.30, PWM:37
00:12:34.818 -> Water_Temp.:101.30, PWM:37
00:12:35.613 -> Water_Temp.:101.30, PWM:38
00:12:36.375 -> Water_Temp.:101.30, PWM:38
00:12:37.171 -> Water_Temp.:101.41, PWM:37
00:12:37.930 -> Water_Temp.:101.41, PWM:37
00:12:38.726 -> Water_Temp.:101.41, PWM:37
00:12:39.488 -> Water_Temp.:101.41, PWM:38
00:12:41.046 -> Water_Temp.:101.41, PWM:37
00:12:41.840 -> Water_Temp.:101.41, PWM:37
00:12:44.160 -> Water_Temp.:101.52, PWM:38
00:12:45.709 -> Water_Temp.:101.52, PWM:37
00:12:46.504 -> Water_Temp.:101.41, PWM:37
00:12:47.296 -> Water_Temp.:101.41, PWM:37
00:12:48.058 -> Water_Temp.:101.41, PWM:38
00:12:48.854 -> Water_Temp.:101.41, PWM:37
00:12:49.615 -> Water_Temp.:101.41, PWM:37
00:12:50.412 -> Water_Temp.:101.41, PWM:37
00:12:52.730 -> Water_Temp.:101.41, PWM:37
00:12:53.489 -> Water_Temp.:101.41, PWM:38
00:12:54.284 -> Water_Temp.:101.41, PWM:38
00:12:55.045 -> Water_Temp.:101.41, PWM:37
00:12:55.839 -> Water_Temp.:101.41, PWM:37
00:12:57.394 -> Water_Temp.:101.41, PWM:37
00:12:58.189 -> Water_Temp.:101.41, PWM:37
00:12:59.746 -> Water_Temp.:101.41, PWM:38
00:13:00.511 -> Water_Temp.:101.41, PWM:37
00:13:01.308 -> Water_Temp.:101.41, PWM:37
00:13:02.827 -> Water_Temp.:101.41, PWM:37
00:13:03.625 -> Water_Temp.:101.52, PWM:37
00:13:04.382 -> Water_Temp.:101.52, PWM:37
00:13:05.969 -> Water_Temp.:101.52, PWM:37
00:13:06.733 -> Water_Temp.:101.52, PWM:38
00:13:07.524 -> Water_Temp.:101.52, PWM:37
00:13:08.284 -> Water_Temp.:101.52, PWM:38
00:13:09.080 -> Water_Temp.:101.52, PWM:38
00:13:09.845 -> Water_Temp.:101.52, PWM:37
00:13:10.645 -> Water_Temp.:101.52, PWM:38
00:13:12.164 -> Water_Temp.:101.52, PWM:37
00:13:12.962 -> Water_Temp.:101.52, PWM:37
00:13:13.723 -> Water_Temp.:101.52, PWM:38
00:13:14.523 -> Water_Temp.:101.52, PWM:38
00:13:15.282 -> Water_Temp.:101.52, PWM:38
00:13:16.080 -> Water_Temp.:101.52, PWM:37
00:13:16.837 -> Water_Temp.:101.52, PWM:37
00:13:17.631 -> Water_Temp.:101.52, PWM:38
00:13:18.393 -> Water_Temp.:101.52, PWM:37
00:13:19.981 -> Water_Temp.:101.52, PWM:38
00:13:20.741 -> Water_Temp.:101.52, PWM:37
00:13:21.538 -> Water_Temp.:101.52, PWM:37
00:13:22.299 -> Water_Temp.:101.52, PWM:37
00:13:23.099 -> Water_Temp.:101.52, PWM:37
00:13:24.627 -> Water_Temp.:101.52, PWM:38
00:13:25.424 -> Water_Temp.:101.52, PWM:38
00:13:27.740 -> Water_Temp.:101.52, PWM:38
00:13:29.292 -> Water_Temp.:101.52, PWM:38
00:13:30.085 -> Water_Temp.:101.41, PWM:37
00:13:30.881 -> Water_Temp.:101.41, PWM:37
00:13:32.435 -> Water_Temp.:101.52, PWM:38
00:13:33.197 -> Water_Temp.:101.52, PWM:37
00:13:33.992 -> Water_Temp.:101.52, PWM:37
00:13:34.756 -> Water_Temp.:101.52, PWM:37
00:13:36.317 -> Water_Temp.:101.52, PWM:37
00:13:37.075 -> Water_Temp.:101.52, PWM:37
00:13:37.871 -> Water_Temp.:101.52, PWM:37
00:13:38.633 -> Water_Temp.:101.52, PWM:37
00:13:39.430 -> Water_Temp.:101.52, PWM:37
00:13:40.191 -> Water_Temp.:101.52, PWM:38
00:13:40.988 -> Water_Temp.:101.52, PWM:38
00:13:43.312 -> Water_Temp.:101.64, PWM:37
00:13:44.106 -> Water_Temp.:101.64, PWM:37
00:13:44.864 -> Water_Temp.:101.64, PWM:37
00:13:45.655 -> Water_Temp.:101.64, PWM:37
00:13:47.207 -> Water_Temp.:101.64, PWM:37
00:13:47.969 -> Water_Temp.:101.64, PWM:37
00:13:48.767 -> Water_Temp.:101.64, PWM:37
00:13:49.528 -> Water_Temp.:101.64, PWM:38
00:13:51.089 -> Water_Temp.:101.52, PWM:37
00:13:51.883 -> Water_Temp.:101.52, PWM:37
00:13:52.646 -> Water_Temp.:101.52, PWM:38
00:13:53.446 -> Water_Temp.:101.52, PWM:38
00:13:54.210 -> Water_Temp.:101.52, PWM:38
00:13:55.007 -> Water_Temp.:101.52, PWM:37
00:13:57.319 -> Water_Temp.:101.52, PWM:37
00:13:58.112 -> Water_Temp.:101.64, PWM:37
00:14:00.432 -> Water_Temp.:101.64, PWM:37
00:14:01.190 -> Water_Temp.:101.64, PWM:38
00:14:01.986 -> Water_Temp.:101.64, PWM:38
00:14:04.306 -> Water_Temp.:101.64, PWM:38
00:14:05.069 -> Water_Temp.:101.64, PWM:37
00:14:06.622 -> Water_Temp.:101.64, PWM:37
00:14:08.174 -> Water_Temp.:101.64, PWM:37
00:14:09.701 -> Water_Temp.:101.64, PWM:37
00:14:10.495 -> Water_Temp.:101.64, PWM:37
00:14:11.256 -> Water_Temp.:101.64, PWM:38
00:14:12.054 -> Water_Temp.:101.64, PWM:37
00:14:12.818 -> Water_Temp.:101.75, PWM:37
00:14:13.615 -> Water_Temp.:101.75, PWM:37
00:14:14.374 -> Water_Temp.:101.75, PWM:37
00:14:15.173 -> Water_Temp.:101.75, PWM:38
00:14:15.936 -> Water_Temp.:101.75, PWM:37
00:14:16.729 -> Water_Temp.:101.75, PWM:37
00:14:18.260 -> Water_Temp.:101.75, PWM:37
00:14:19.059 -> Water_Temp.:101.75, PWM:37
00:14:19.824 -> Water_Temp.:101.75, PWM:37
00:14:22.141 -> Water_Temp.:101.75, PWM:38
00:14:23.666 -> Water_Temp.:101.64, PWM:37
00:14:25.225 -> Water_Temp.:101.64, PWM:37
00:14:26.019 -> Water_Temp.:101.64, PWM:37
00:14:26.779 -> Water_Temp.:101.64, PWM:38
00:14:27.575 -> Water_Temp.:101.64, PWM:37
00:14:29.895 -> Water_Temp.:101.64, PWM:37
00:14:32.194 -> Water_Temp.:101.52, PWM:38
00:14:33.751 -> Water_Temp.:101.52, PWM:37
00:14:34.550 -> Water_Temp.:101.52, PWM:37
00:14:35.311 -> Water_Temp.:101.52, PWM:38
00:14:36.107 -> Water_Temp.:101.52, PWM:38
00:14:37.661 -> Water_Temp.:101.52, PWM:37
00:14:38.421 -> Water_Temp.:101.52, PWM:38
00:14:39.981 -> Water_Temp.:101.64, PWM:37
00:14:42.296 -> Water_Temp.:101.64, PWM:37
00:14:43.093 -> Water_Temp.:101.64, PWM:37
00:14:43.854 -> Water_Temp.:101.64, PWM:37
00:14:44.652 -> Water_Temp.:101.64, PWM:37
00:14:45.418 -> Water_Temp.:101.64, PWM:37
00:14:46.217 -> Water_Temp.:101.64, PWM:37
00:14:46.977 -> Water_Temp.:101.64, PWM:37
00:14:48.535 -> Water_Temp.:101.64, PWM:37
00:14:49.299 -> Water_Temp.:101.64, PWM:38
00:14:50.094 -> Water_Temp.:101.64, PWM:38
00:14:50.888 -> Water_Temp.:101.64, PWM:37
00:14:51.652 -> Water_Temp.:101.64, PWM:37
00:14:52.448 -> Water_Temp.:101.64, PWM:37
00:14:53.211 -> Water_Temp.:101.64, PWM:38
00:14:54.008 -> Water_Temp.:101.64, PWM:37
00:14:55.533 -> Water_Temp.:101.64, PWM:37
00:14:56.328 -> Water_Temp.:101.64, PWM:37
00:14:57.085 -> Water_Temp.:101.64, PWM:37
00:14:58.642 -> Water_Temp.:101.52, PWM:37
00:14:59.435 -> Water_Temp.:101.52, PWM:37
00:15:00.231 -> Water_Temp.:101.52, PWM:38
00:15:01.779 -> Water_Temp.:101.64, PWM:37
00:15:02.541 -> Water_Temp.:101.64, PWM:38
00:15:03.333 -> Water_Temp.:101.64, PWM:37
00:15:06.414 -> Water_Temp.:101.64, PWM:37
00:15:07.211 -> Water_Temp.:101.64, PWM:37
00:15:07.974 -> Water_Temp.:101.64, PWM:37
00:15:09.530 -> Water_Temp.:101.64, PWM:38
00:15:10.324 -> Water_Temp.:101.64, PWM:37
00:15:11.086 -> Water_Temp.:101.64, PWM:37
00:15:11.876 -> Water_Temp.:101.64, PWM:37
00:15:12.673 -> Water_Temp.:101.64, PWM:37
00:15:14.953 -> Water_Temp.:101.64, PWM:37
00:15:15.748 -> Water_Temp.:101.64, PWM:37
00:15:16.541 -> Water_Temp.:101.64, PWM:37
00:15:17.298 -> Water_Temp.:101.64, PWM:37
00:15:18.851 -> Water_Temp.:101.64, PWM:37
00:15:19.645 -> Water_Temp.:101.64, PWM:37
00:15:21.201 -> Water_Temp.:101.64, PWM:37
00:15:21.997 -> Water_Temp.:101.64, PWM:38
00:15:23.518 -> Water_Temp.:101.64, PWM:37
00:15:24.315 -> Water_Temp.:101.64, PWM:37
00:15:25.076 -> Water_Temp.:101.64, PWM:37
00:15:25.866 -> Water_Temp.:101.64, PWM:37
00:15:26.658 -> Water_Temp.:101.64, PWM:37
00:15:27.415 -> Water_Temp.:101.64, PWM:37
00:15:28.213 -> Water_Temp.:101.64, PWM:37
00:15:28.979 -> Water_Temp.:101.64, PWM:37
00:15:29.776 -> Water_Temp.:101.64, PWM:37
00:15:32.826 -> Water_Temp.:101.64, PWM:37
00:15:33.625 -> Water_Temp.:101.64, PWM:37
00:15:34.388 -> Water_Temp.:101.64, PWM:38
00:15:35.188 -> Water_Temp.:101.64, PWM:38
00:15:36.712 -> Water_Temp.:101.64, PWM:38
00:15:38.234 -> Water_Temp.:101.64, PWM:37
00:15:39.034 -> Water_Temp.:101.64, PWM:37
00:15:40.591 -> Water_Temp.:101.64, PWM:37
00:15:41.353 -> Water_Temp.:101.64, PWM:38
00:15:42.150 -> Water_Temp.:101.64, PWM:38
00:15:42.912 -> Water_Temp.:101.64, PWM:37
00:15:45.237 -> Water_Temp.:101.64, PWM:37
00:15:46.795 -> Water_Temp.:103.32, PWM:37
00:15:47.562 -> Water_Temp.:105.46, PWM:38
00:15:49.124 -> Water_Temp.:102.65, PWM:37
00:15:51.472 -> Water_Temp.:102.20, PWM:37
00:15:52.237 -> Water_Temp.:102.09, PWM:37
00:15:53.032 -> Water_Temp.:102.09, PWM:37
00:15:54.589 -> Water_Temp.:101.97, PWM:37
00:15:55.351 -> Water_Temp.:101.97, PWM:37
00:15:56.151 -> Water_Temp.:103.55, PWM:37
00:15:57.709 -> Water_Temp.:103.10, PWM:38
00:15:59.269 -> Water_Temp.:102.42, PWM:38
00:16:00.031 -> Water_Temp.:102.31, PWM:38
00:16:00.826 -> Water_Temp.:102.20, PWM:37
00:16:02.384 -> Water_Temp.:102.09, PWM:38
00:16:06.264 -> Water_Temp.:101.86, PWM:37
00:16:07.026 -> Water_Temp.:101.86, PWM:37
00:16:07.822 -> Water_Temp.:101.86, PWM:37
00:16:08.585 -> Water_Temp.:101.86, PWM:37
00:16:10.146 -> Water_Temp.:101.86, PWM:37
00:16:10.909 -> Water_Temp.:101.75, PWM:37
00:16:11.704 -> Water_Temp.:101.86, PWM:37
00:16:14.795 -> Water_Temp.:101.75, PWM:37
00:16:15.593 -> Water_Temp.:101.75, PWM:38
00:16:16.358 -> Water_Temp.:101.75, PWM:37
00:16:18.641 -> Water_Temp.:101.75, PWM:37
00:16:19.439 -> Water_Temp.:101.75, PWM:37
00:16:20.203 -> Water_Temp.:101.75, PWM:38
00:16:20.999 -> Water_Temp.:101.75, PWM:37
00:16:23.321 -> Water_Temp.:101.75, PWM:37
00:16:24.118 -> Water_Temp.:101.75, PWM:38
00:16:24.876 -> Water_Temp.:101.75, PWM:37
00:16:27.230 -> Water_Temp.:101.75, PWM:37
00:16:27.993 -> Water_Temp.:101.75, PWM:38
00:16:28.791 -> Water_Temp.:101.75, PWM:37
00:16:29.551 -> Water_Temp.:101.75, PWM:38
00:16:30.350 -> Water_Temp.:103.32, PWM:37
00:16:31.116 -> Water_Temp.:105.46, PWM:37
00:16:31.912 -> Water_Temp.:102.99, PWM:37
00:16:32.677 -> Water_Temp.:102.65, PWM:37
00:16:33.474 -> Water_Temp.:102.42, PWM:37
00:16:34.239 -> Water_Temp.:102.20, PWM:37
00:16:35.003 -> Water_Temp.:102.20, PWM:37
00:16:36.595 -> Water_Temp.:101.97, PWM:37
00:16:37.359 -> Water_Temp.:101.97, PWM:37
00:16:38.120 -> Water_Temp.:101.97, PWM:37
00:16:39.686 -> Water_Temp.:101.86, PWM:38
00:16:40.444 -> Water_Temp.:101.86, PWM:38
00:16:45.023 -> Water_Temp.:101.75, PWM:37
00:16:45.819 -> Water_Temp.:101.75, PWM:37
00:16:46.578 -> Water_Temp.:101.75, PWM:37
00:16:48.134 -> Water_Temp.:101.75, PWM:38
00:16:48.899 -> Water_Temp.:101.75, PWM:37
00:16:50.458 -> Water_Temp.:101.75, PWM:37
00:16:51.254 -> Water_Temp.:101.75, PWM:37
00:16:52.017 -> Water_Temp.:101.75, PWM:38
00:16:52.780 -> Water_Temp.:101.75, PWM:37
00:16:54.339 -> Water_Temp.:101.75, PWM:38
00:16:55.101 -> Water_Temp.:101.75, PWM:38
00:16:55.898 -> Water_Temp.:101.75, PWM:37
00:16:56.691 -> Water_Temp.:101.75, PWM:37
00:16:57.457 -> Water_Temp.:101.75, PWM:38
00:16:59.016 -> Water_Temp.:101.86, PWM:38
00:16:59.778 -> Water_Temp.:101.86, PWM:37
00:17:00.572 -> Water_Temp.:101.86, PWM:37
00:17:01.372 -> Water_Temp.:101.86, PWM:38
00:17:03.661 -> Water_Temp.:101.97, PWM:37
00:17:04.456 -> Water_Temp.:101.97, PWM:37
00:17:06.008 -> Water_Temp.:101.97, PWM:38
00:17:06.803 -> Water_Temp.:101.97, PWM:38
00:17:07.565 -> Water_Temp.:101.86, PWM:37
00:17:09.129 -> Water_Temp.:101.86, PWM:38
00:17:09.894 -> Water_Temp.:101.86, PWM:38
00:17:10.690 -> Water_Temp.:103.55, PWM:37
00:17:11.454 -> Water_Temp.:105.57, PWM:37
00:17:12.250 -> Water_Temp.:103.21, PWM:37
00:17:13.010 -> Water_Temp.:103.21, PWM:38
00:17:13.809 -> Water_Temp.:102.54, PWM:38
00:17:14.574 -> Water_Temp.:102.42, PWM:37
00:17:15.371 -> Water_Temp.:102.31, PWM:38
00:17:16.936 -> Water_Temp.:102.20, PWM:37
00:17:17.697 -> Water_Temp.:102.20, PWM:38
00:17:18.493 -> Water_Temp.:102.20, PWM:37
00:17:20.021 -> Water_Temp.:102.20, PWM:38
00:17:20.781 -> Water_Temp.:102.20, PWM:37
00:17:23.889 -> Water_Temp.:102.20, PWM:37
00:17:28.485 -> Water_Temp.:102.09, PWM:37
00:17:29.282 -> Water_Temp.:102.09, PWM:38
00:17:30.077 -> Water_Temp.:102.09, PWM:37
00:17:32.363 -> Water_Temp.:102.09, PWM:38
00:17:34.649 -> Water_Temp.:101.97, PWM:37
00:17:35.443 -> Water_Temp.:101.97, PWM:37
00:17:36.240 -> Water_Temp.:101.97, PWM:37
00:17:37.791 -> Water_Temp.:101.97, PWM:37
00:17:38.555 -> Water_Temp.:101.97, PWM:37
00:17:39.354 -> Water_Temp.:101.97, PWM:37
00:17:40.115 -> Water_Temp.:101.97, PWM:38
00:17:40.914 -> Water_Temp.:101.97, PWM:38
00:17:41.677 -> Water_Temp.:101.97, PWM:37
00:17:42.471 -> Water_Temp.:101.97, PWM:37
00:17:44.793 -> Water_Temp.:102.09, PWM:38
00:17:45.555 -> Water_Temp.:102.09, PWM:37
00:17:46.353 -> Water_Temp.:102.09, PWM:37
00:17:47.880 -> Water_Temp.:102.09, PWM:38
00:17:48.645 -> Water_Temp.:102.09, PWM:37
00:17:49.445 -> Water_Temp.:102.09, PWM:37
00:17:50.205 -> Water_Temp.:102.09, PWM:37
00:17:51.001 -> Water_Temp.:102.09, PWM:37
00:17:51.764 -> Water_Temp.:102.09, PWM:37
00:17:52.563 -> Water_Temp.:102.09, PWM:38
00:17:56.410 -> Water_Temp.:102.09, PWM:38
00:17:57.206 -> Water_Temp.:102.09, PWM:37
00:17:58.766 -> Water_Temp.:102.09, PWM:38
00:17:59.528 -> Water_Temp.:102.20, PWM:37
00:18:00.325 -> Water_Temp.:102.20, PWM:37
00:18:01.088 -> Water_Temp.:102.20, PWM:37
00:18:01.877 -> Water_Temp.:102.20, PWM:37
00:18:02.672 -> Water_Temp.:102.20, PWM:38
00:18:04.198 -> Water_Temp.:102.20, PWM:37
00:18:04.996 -> Water_Temp.:102.20, PWM:37
00:18:06.560 -> Water_Temp.:102.20, PWM:38
00:18:08.878 -> Water_Temp.:102.20, PWM:37
00:18:10.439 -> Water_Temp.:102.20, PWM:37
00:18:11.998 -> Water_Temp.:102.20, PWM:37
00:18:12.792 -> Water_Temp.:102.20, PWM:37
00:18:13.554 -> Water_Temp.:102.20, PWM:37
00:18:15.113 -> Water_Temp.:102.20, PWM:37
00:18:15.904 -> Water_Temp.:102.20, PWM:37
00:18:16.696 -> Water_Temp.:102.20, PWM:37
00:18:18.215 -> Water_Temp.:102.20, PWM:38
00:18:19.013 -> Water_Temp.:102.20, PWM:38
00:18:19.774 -> Water_Temp.:102.20, PWM:37
00:18:21.331 -> Water_Temp.:102.20, PWM:38
00:18:22.125 -> Water_Temp.:102.20, PWM:38
00:18:22.917 -> Water_Temp.:102.20, PWM:38
00:18:23.678 -> Water_Temp.:102.20, PWM:38
00:18:25.239 -> Water_Temp.:102.20, PWM:38
00:18:26.792 -> Water_Temp.:102.20, PWM:37
00:18:27.589 -> Water_Temp.:102.20, PWM:37
00:18:28.345 -> Water_Temp.:102.20, PWM:37
00:18:29.140 -> Water_Temp.:102.20, PWM:37
00:18:29.935 -> Water_Temp.:102.20, PWM:38
00:18:30.701 -> Water_Temp.:102.20, PWM:37
00:18:31.496 -> Water_Temp.:102.20, PWM:38
00:18:32.255 -> Water_Temp.:102.20, PWM:37
00:18:33.048 -> Water_Temp.:102.20, PWM:37
00:18:34.609 -> Water_Temp.:102.20, PWM:37
00:18:35.369 -> Water_Temp.:102.20, PWM:37
00:18:36.167 -> Water_Temp.:102.20, PWM:38
00:18:36.928 -> Water_Temp.:102.20, PWM:38
00:18:38.488 -> Water_Temp.:102.20, PWM:37
00:18:40.077 -> Water_Temp.:102.31, PWM:37
00:18:40.837 -> Water_Temp.:102.31, PWM:37
00:18:41.635 -> Water_Temp.:102.31, PWM:37
00:18:42.395 -> Water_Temp.:102.31, PWM:37
00:18:43.955 -> Water_Temp.:102.31, PWM:37
00:18:47.804 -> Water_Temp.:102.20, PWM:37
00:18:48.567 -> Water_Temp.:102.20, PWM:38
00:18:49.363 -> Water_Temp.:102.20, PWM:37
00:18:50.127 -> Water_Temp.:102.20, PWM:37
00:18:50.925 -> Water_Temp.:102.31, PWM:37
00:18:51.720 -> Water_Temp.:102.31, PWM:37
00:18:52.485 -> Water_Temp.:102.31, PWM:38
00:18:53.246 -> Water_Temp.:102.31, PWM:38
00:18:54.045 -> Water_Temp.:102.31, PWM:38
00:18:54.809 -> Water_Temp.:102.31, PWM:38
00:18:56.362 -> Water_Temp.:102.31, PWM:37
00:18:59.453 -> Water_Temp.:102.42, PWM:37
00:19:00.213 -> Water_Temp.:102.42, PWM:38
00:19:02.570 -> Water_Temp.:102.42, PWM:38
00:19:03.332 -> Water_Temp.:102.42, PWM:37
00:19:04.127 -> Water_Temp.:102.42, PWM:37
00:19:04.886 -> Water_Temp.:102.42, PWM:37
00:19:05.682 -> Water_Temp.:102.42, PWM:38
00:19:07.238 -> Water_Temp.:102.42, PWM:37
00:19:08.002 -> Water_Temp.:102.42, PWM:37
00:19:08.794 -> Water_Temp.:102.42, PWM:38
00:19:09.559 -> Water_Temp.:102.42, PWM:38
00:19:10.356 -> Water_Temp.:102.42, PWM:37
00:19:11.151 -> Water_Temp.:102.42, PWM:37
00:19:13.468 -> Water_Temp.:102.42, PWM:37
00:19:16.547 -> Water_Temp.:102.42, PWM:37
00:19:17.309 -> Water_Temp.:102.42, PWM:37
00:19:18.108 -> Water_Temp.:102.42, PWM:37
00:19:18.873 -> Water_Temp.:102.42, PWM:37
00:19:20.401 -> Water_Temp.:102.42, PWM:38
00:19:21.203 -> Water_Temp.:102.42, PWM:37
00:19:21.961 -> Water_Temp.:102.31, PWM:38
00:19:22.756 -> Water_Temp.:102.31, PWM:37
00:19:23.514 -> Water_Temp.:102.31, PWM:37
00:19:24.309 -> Water_Temp.:102.31, PWM:37
00:19:28.900 -> Water_Temp.:102.42, PWM:37
00:19:30.487 -> Water_Temp.:104.11, PWM:37
00:19:31.247 -> Water_Temp.:106.14, PWM:38
00:19:33.570 -> Water_Temp.:103.10, PWM:37
00:19:34.363 -> Water_Temp.:102.99, PWM:38
00:19:35.163 -> Water_Temp.:102.87, PWM:38
00:19:35.927 -> Water_Temp.:102.76, PWM:37
00:19:36.722 -> Water_Temp.:102.76, PWM:37
00:19:39.041 -> Water_Temp.:102.65, PWM:38
00:19:39.836 -> Water_Temp.:102.65, PWM:38
00:19:40.596 -> Water_Temp.:102.65, PWM:38
00:19:42.882 -> Water_Temp.:102.54, PWM:37
00:19:43.676 -> Water_Temp.:102.54, PWM:37
00:19:44.440 -> Water_Temp.:102.54, PWM:37
00:19:45.229 -> Water_Temp.:102.54, PWM:38
00:19:46.788 -> Water_Temp.:102.54, PWM:37
00:19:48.349 -> Water_Temp.:102.54, PWM:37
00:19:49.115 -> Water_Temp.:102.54, PWM:38
00:19:49.915 -> Water_Temp.:102.54, PWM:37
00:19:50.678 -> Water_Temp.:102.54, PWM:37
00:19:51.476 -> Water_Temp.:102.54, PWM:37
00:19:52.231 -> Water_Temp.:102.54, PWM:37
00:19:53.023 -> Water_Temp.:102.54, PWM:37
00:19:53.821 -> Water_Temp.:102.54, PWM:37
00:19:55.343 -> Water_Temp.:102.54, PWM:37
00:19:56.139 -> Water_Temp.:102.54, PWM:38
00:19:56.898 -> Water_Temp.:102.54, PWM:37
00:19:57.694 -> Water_Temp.:102.54, PWM:37
00:19:58.457 -> Water_Temp.:102.54, PWM:37
00:19:59.253 -> Water_Temp.:102.54, PWM:37
00:20:00.778 -> Water_Temp.:102.54, PWM:37
00:20:01.572 -> Water_Temp.:102.54, PWM:38
00:20:02.338 -> Water_Temp.:102.54, PWM:37
00:20:03.132 -> Water_Temp.:102.54, PWM:37
00:20:04.692 -> Water_Temp.:102.54, PWM:38
00:20:05.451 -> Water_Temp.:102.54, PWM:37
00:20:06.251 -> Water_Temp.:102.65, PWM:37
00:20:07.012 -> Water_Temp.:102.65, PWM:38
00:20:07.804 -> Water_Temp.:102.65, PWM:37
00:20:09.363 -> Water_Temp.:102.65, PWM:37
00:20:10.123 -> Water_Temp.:102.65, PWM:37
00:20:10.921 -> Water_Temp.:102.65, PWM:37
00:20:12.476 -> Water_Temp.:102.65, PWM:38
00:20:13.240 -> Water_Temp.:102.65, PWM:38
00:20:14.797 -> Water_Temp.:102.54, PWM:38
00:20:16.354 -> Water_Temp.:102.54, PWM:37
00:20:17.151 -> Water_Temp.:102.54, PWM:37
00:20:17.914 -> Water_Temp.:102.54, PWM:37
00:20:19.469 -> Water_Temp.:102.54, PWM:37
00:20:20.228 -> Water_Temp.:102.54, PWM:37
00:20:21.023 -> Water_Temp.:102.54, PWM:37
00:20:23.367 -> Water_Temp.:102.54, PWM:37
00:20:24.125 -> Water_Temp.:102.54, PWM:37
00:20:26.468 -> Water_Temp.:104.11, PWM:37
00:20:27.264 -> Water_Temp.:106.25, PWM:37
00:20:28.021 -> Water_Temp.:103.77, PWM:38
00:20:28.818 -> Water_Temp.:103.44, PWM:38
00:20:32.656 -> Water_Temp.:102.76, PWM:38
00:20:35.010 -> Water_Temp.:102.65, PWM:37
00:20:35.770 -> Water_Temp.:102.65, PWM:37
00:20:36.565 -> Water_Temp.:102.65, PWM:37
00:20:37.327 -> Water_Temp.:102.65, PWM:37
00:20:38.123 -> Water_Temp.:102.65, PWM:38
00:20:38.886 -> Water_Temp.:102.65, PWM:38
00:20:42.758 -> Water_Temp.:102.65, PWM:38
00:20:43.555 -> Water_Temp.:102.76, PWM:38
00:20:45.078 -> Water_Temp.:102.76, PWM:37
00:20:45.876 -> Water_Temp.:102.76, PWM:37
00:20:46.636 -> Water_Temp.:102.76, PWM:37
00:20:49.713 -> Water_Temp.:102.76, PWM:38
00:20:50.513 -> Water_Temp.:102.76, PWM:37
00:20:51.278 -> Water_Temp.:102.76, PWM:37
00:20:52.078 -> Water_Temp.:102.76, PWM:37
00:20:52.840 -> Water_Temp.:102.65, PWM:37
00:20:53.634 -> Water_Temp.:102.76, PWM:37
00:20:55.154 -> Water_Temp.:102.65, PWM:37
00:20:55.951 -> Water_Temp.:102.65, PWM:37
00:20:56.711 -> Water_Temp.:102.65, PWM:38
00:20:58.273 -> Water_Temp.:102.65, PWM:37
00:20:59.831 -> Water_Temp.:102.65, PWM:38
00:21:02.157 -> Water_Temp.:102.65, PWM:37
00:21:04.472 -> Water_Temp.:102.65, PWM:38
00:21:05.271 -> Water_Temp.:102.76, PWM:37
00:21:06.033 -> Water_Temp.:102.76, PWM:37
00:21:06.834 -> Water_Temp.:102.76, PWM:38
00:21:07.597 -> Water_Temp.:102.76, PWM:37
00:21:08.393 -> Water_Temp.:102.76, PWM:38
00:21:09.950 -> Water_Temp.:102.76, PWM:38
00:21:11.472 -> Water_Temp.:102.76, PWM:38
00:21:12.268 -> Water_Temp.:102.76, PWM:38
00:21:13.827 -> Water_Temp.:102.76, PWM:38
00:21:14.591 -> Water_Temp.:102.76, PWM:37
00:21:15.389 -> Water_Temp.:102.76, PWM:38
00:21:16.154 -> Water_Temp.:102.76, PWM:38
00:21:16.951 -> Water_Temp.:102.76, PWM:37
00:21:17.713 -> Water_Temp.:102.76, PWM:38
00:21:18.512 -> Water_Temp.:102.76, PWM:37
00:21:19.274 -> Water_Temp.:102.76, PWM:38
00:21:20.831 -> Water_Temp.:102.76, PWM:38
00:21:21.589 -> Water_Temp.:102.76, PWM:37
00:21:22.382 -> Water_Temp.:102.76, PWM:38
00:21:23.144 -> Water_Temp.:102.76, PWM:38
00:21:23.941 -> Water_Temp.:102.76, PWM:38
00:21:24.739 -> Water_Temp.:102.76, PWM:37
00:21:26.297 -> Water_Temp.:102.76, PWM:38
00:21:27.058 -> Water_Temp.:102.76, PWM:38
00:21:29.378 -> Water_Temp.:102.76, PWM:38
00:21:30.143 -> Water_Temp.:102.76, PWM:38
00:21:30.902 -> Water_Temp.:102.76, PWM:37
00:21:32.489 -> Water_Temp.:102.76, PWM:37
00:21:34.044 -> Water_Temp.:102.76, PWM:37
00:21:34.808 -> Water_Temp.:102.65, PWM:37
00:21:35.571 -> Water_Temp.:102.65, PWM:38
00:21:36.367 -> Water_Temp.:102.65, PWM:38
00:21:37.164 -> Water_Temp.:102.65, PWM:37
00:21:37.927 -> Water_Temp.:102.65, PWM:37
00:21:38.691 -> Water_Temp.:102.65, PWM:37
00:21:40.254 -> Water_Temp.:102.76, PWM:37
00:21:41.015 -> Water_Temp.:102.76, PWM:37
00:21:41.810 -> Water_Temp.:102.76, PWM:37
00:21:42.604 -> Water_Temp.:102.76, PWM:38
00:21:44.919 -> Water_Temp.:102.76, PWM:37
00:21:45.712 -> Water_Temp.:102.76, PWM:38
00:21:46.476 -> Water_Temp.:102.76, PWM:37
00:21:47.272 -> Water_Temp.:102.76, PWM:37
00:21:48.036 -> Water_Temp.:102.76, PWM:37
00:21:50.354 -> Water_Temp.:102.76, PWM:38
00:21:51.115 -> Water_Temp.:102.76, PWM:38
00:21:51.907 -> Water_Temp.:102.76, PWM:37
00:21:52.667 -> Water_Temp.:102.76, PWM:38
00:21:54.257 -> Water_Temp.:102.76, PWM:37
00:21:55.785 -> Water_Temp.:102.65, PWM:38
00:21:56.581 -> Water_Temp.:102.65, PWM:37
00:21:57.340 -> Water_Temp.:102.65, PWM:37
00:21:58.137 -> Water_Temp.:102.65, PWM:37
00:21:59.657 -> Water_Temp.:102.65, PWM:37
00:22:00.452 -> Water_Temp.:102.65, PWM:38
00:22:01.969 -> Water_Temp.:102.65, PWM:37
00:22:02.765 -> Water_Temp.:102.65, PWM:37
00:22:04.322 -> Water_Temp.:102.65, PWM:37
00:22:05.086 -> Water_Temp.:102.65, PWM:37
00:22:05.850 -> Water_Temp.:102.65, PWM:37
00:22:06.646 -> Water_Temp.:102.65, PWM:37
00:22:07.441 -> Water_Temp.:102.65, PWM:37
00:22:08.206 -> Water_Temp.:102.65, PWM:37
00:22:10.484 -> Water_Temp.:102.65, PWM:37
00:22:12.075 -> Water_Temp.:102.65, PWM:37
00:22:12.833 -> Water_Temp.:102.65, PWM:37
00:22:13.628 -> Water_Temp.:102.65, PWM:37
00:22:15.181 -> Water_Temp.:102.65, PWM:37
00:22:15.942 -> Water_Temp.:102.65, PWM:37
00:22:16.738 -> Water_Temp.:102.65, PWM:37
00:22:17.532 -> Water_Temp.:102.54, PWM:37
00:22:19.092 -> Water_Temp.:102.54, PWM:38
00:22:20.614 -> Water_Temp.:102.65, PWM:38
00:22:21.410 -> Water_Temp.:102.65, PWM:38
00:22:22.172 -> Water_Temp.:102.65, PWM:37
00:22:22.965 -> Water_Temp.:102.65, PWM:37
00:22:23.725 -> Water_Temp.:102.65, PWM:38
00:22:24.524 -> Water_Temp.:102.65, PWM:38
00:22:25.289 -> Water_Temp.:102.65, PWM:37
00:22:26.082 -> Water_Temp.:102.65, PWM:37
00:22:27.636 -> Water_Temp.:102.65, PWM:37
00:22:29.188 -> Water_Temp.:102.65, PWM:37
00:22:30.744 -> Water_Temp.:102.65, PWM:37
00:22:32.299 -> Water_Temp.:102.65, PWM:37
00:22:33.861 -> Water_Temp.:102.65, PWM:38
00:22:34.623 -> Water_Temp.:102.65, PWM:38
00:22:35.422 -> Water_Temp.:102.65, PWM:37
00:22:36.181 -> Water_Temp.:102.65, PWM:37
00:22:36.976 -> Water_Temp.:102.65, PWM:38
00:22:37.767 -> Water_Temp.:102.65, PWM:37
00:22:38.530 -> Water_Temp.:102.65, PWM:38
00:22:39.329 -> Water_Temp.:102.65, PWM:37
00:22:41.618 -> Water_Temp.:102.65, PWM:37
00:22:42.412 -> Water_Temp.:102.65, PWM:38
00:22:44.745 -> Water_Temp.:102.76, PWM:37
00:22:46.306 -> Water_Temp.:102.76, PWM:37
00:22:49.418 -> Water_Temp.:102.76, PWM:37
00:22:50.182 -> Water_Temp.:102.76, PWM:37
00:22:50.974 -> Water_Temp.:102.76, PWM:38
00:22:51.737 -> Water_Temp.:102.76, PWM:37
00:22:52.532 -> Water_Temp.:102.76, PWM:37
00:22:53.294 -> Water_Temp.:102.76, PWM:37
00:22:54.855 -> Water_Temp.:102.76, PWM:37
00:22:55.614 -> Water_Temp.:102.76, PWM:37
00:22:57.168 -> Water_Temp.:102.76, PWM:38
00:22:57.965 -> Water_Temp.:102.76, PWM:37
00:22:58.726 -> Water_Temp.:102.76, PWM:37
00:23:00.287 -> Water_Temp.:102.65, PWM:37
00:23:01.085 -> Water_Temp.:102.65, PWM:38
00:23:01.848 -> Water_Temp.:102.65, PWM:37
00:23:02.643 -> Water_Temp.:102.65, PWM:38
00:23:03.406 -> Water_Temp.:102.65, PWM:37
00:23:04.205 -> Water_Temp.:102.65, PWM:38
00:23:04.968 -> Water_Temp.:102.76, PWM:37
00:23:06.527 -> Water_Temp.:102.76, PWM:37
00:23:09.610 -> Water_Temp.:102.76, PWM:37
00:23:10.373 -> Water_Temp.:102.76, PWM:38
00:23:11.169 -> Water_Temp.:102.76, PWM:37
00:23:11.933 -> Water_Temp.:102.65, PWM:38
00:23:13.493 -> Water_Temp.:102.65, PWM:37
00:23:18.162 -> Water_Temp.:102.65, PWM:37
00:23:18.928 -> Water_Temp.:102.65, PWM:37
00:23:21.281 -> Water_Temp.:102.65, PWM:38
00:23:22.819 -> Water_Temp.:102.65, PWM:37
00:23:23.612 -> Water_Temp.:102.65, PWM:37
00:23:24.373 -> Water_Temp.:102.65, PWM:37
00:23:26.726 -> Water_Temp.:102.65, PWM:38
00:23:31.331 -> Water_Temp.:102.65, PWM:37
00:23:32.091 -> Water_Temp.:102.65, PWM:38
00:23:32.888 -> Water_Temp.:102.65, PWM:37
00:23:33.682 -> Water_Temp.:102.65, PWM:37
00:23:34.448 -> Water_Temp.:102.65, PWM:37
00:23:35.214 -> Water_Temp.:102.65, PWM:37
00:23:36.006 -> Water_Temp.:102.65, PWM:37
00:23:36.802 -> Water_Temp.:102.65, PWM:37
00:23:37.566 -> Water_Temp.:102.54, PWM:37
00:23:39.092 -> Water_Temp.:102.54, PWM:38
00:23:39.887 -> Water_Temp.:102.54, PWM:37
00:23:41.445 -> Water_Temp.:102.54, PWM:38
00:23:43.769 -> Water_Temp.:102.54, PWM:38
00:23:45.333 -> Water_Temp.:102.54, PWM:38
00:23:46.132 -> Water_Temp.:102.54, PWM:38
00:23:47.691 -> Water_Temp.:102.65, PWM:37
00:23:48.451 -> Water_Temp.:102.65, PWM:37
00:23:50.009 -> Water_Temp.:102.65, PWM:38
00:23:53.124 -> Water_Temp.:102.65, PWM:38
00:23:54.654 -> Water_Temp.:102.65, PWM:37
00:23:55.451 -> Water_Temp.:102.65, PWM:38
00:23:56.215 -> Water_Temp.:102.65, PWM:37
00:23:56.978 -> Water_Temp.:102.76, PWM:38
00:23:57.775 -> Water_Temp.:102.76, PWM:37
00:23:58.566 -> Water_Temp.:102.76, PWM:37
00:23:59.331 -> Water_Temp.:102.76, PWM:38
00:24:00.124 -> Water_Temp.:102.76, PWM:41
00:24:00.887 -> Water_Temp.:102.76, PWM:41
00:24:01.681 -> Water_Temp.:102.76, PWM:41
00:24:02.440 -> Water_Temp.:102.76, PWM:41
00:24:03.240 -> Water_Temp.:102.76, PWM:41
00:24:04.000 -> Water_Temp.:102.76, PWM:41
00:24:04.800 -> Water_Temp.:102.76, PWM:41
00:24:05.565 -> Water_Temp.:102.76, PWM:41
00:24:06.359 -> Water_Temp.:102.76, PWM:41
00:24:07.124 -> Water_Temp.:102.87, PWM:41
00:24:07.923 -> Water_Temp.:102.76, PWM:41
00:24:09.483 -> Water_Temp.:102.76, PWM:41
00:24:11.041 -> Water_Temp.:102.76, PWM:41
00:24:11.798 -> Water_Temp.:102.76, PWM:41
00:24:12.597 -> Water_Temp.:102.76, PWM:41
00:24:13.360 -> Water_Temp.:102.76, PWM:41
00:24:14.153 -> Water_Temp.:102.76, PWM:41
00:24:14.913 -> Water_Temp.:102.76, PWM:41
00:24:16.469 -> Water_Temp.:102.76, PWM:41
00:24:17.264 -> Water_Temp.:102.76, PWM:41
00:24:18.027 -> Water_Temp.:102.76, PWM:41
00:24:19.587 -> Water_Temp.:102.76, PWM:41
00:24:20.350 -> Water_Temp.:102.76, PWM:41
00:24:21.114 -> Water_Temp.:102.76, PWM:41
00:24:21.912 -> Water_Temp.:102.76, PWM:41
00:24:22.676 -> Water_Temp.:102.76, PWM:41
00:24:23.468 -> Water_Temp.:102.65, PWM:41
00:24:24.231 -> Water_Temp.:102.65, PWM:41
00:24:25.027 -> Water_Temp.:102.65, PWM:41
00:24:25.824 -> Water_Temp.:102.65, PWM:41
00:24:27.377 -> Water_Temp.:102.65, PWM:41
00:24:28.137 -> Water_Temp.:102.65, PWM:41
00:24:28.931 -> Water_Temp.:102.65, PWM:41
00:24:31.216 -> Water_Temp.:102.65, PWM:41
00:24:32.010 -> Water_Temp.:102.65, PWM:41
00:24:32.773 -> Water_Temp.:102.65, PWM:41
00:24:33.573 -> Water_Temp.:102.65, PWM:41
00:24:34.336 -> Water_Temp.:102.65, PWM:41
00:24:35.130 -> Water_Temp.:102.65, PWM:41
00:24:37.417 -> Water_Temp.:102.76, PWM:41
00:24:38.209 -> Water_Temp.:102.76, PWM:41
00:24:38.970 -> Water_Temp.:102.76, PWM:41
00:24:39.763 -> Water_Temp.:102.76, PWM:41
00:24:41.291 -> Water_Temp.:102.76, PWM:41
00:24:42.090 -> Water_Temp.:102.76, PWM:41
00:24:42.850 -> Water_Temp.:102.76, PWM:41
00:24:43.644 -> Water_Temp.:102.76, PWM:41
00:24:44.410 -> Water_Temp.:102.76, PWM:41
00:24:45.972 -> Water_Temp.:102.76, PWM:41
00:24:46.765 -> Water_Temp.:102.76, PWM:41
00:24:47.530 -> Water_Temp.:102.76, PWM:41
00:24:48.329 -> Water_Temp.:102.76, PWM:41
00:24:49.091 -> Water_Temp.:102.76, PWM:41
00:24:50.650 -> Water_Temp.:102.87, PWM:41
00:24:51.413 -> Water_Temp.:102.76, PWM:41
00:24:52.209 -> Water_Temp.:102.87, PWM:41
00:24:53.004 -> Water_Temp.:102.87, PWM:41
00:24:53.765 -> Water_Temp.:102.87, PWM:41
00:24:54.562 -> Water_Temp.:102.87, PWM:41
00:24:56.088 -> Water_Temp.:102.87, PWM:41
00:24:56.884 -> Water_Temp.:102.99, PWM:41
00:24:58.443 -> Water_Temp.:102.99, PWM:41
00:24:59.207 -> Water_Temp.:102.99, PWM:41
00:25:00.005 -> Water_Temp.:102.99, PWM:41
00:25:00.770 -> Water_Temp.:102.99, PWM:41
00:25:01.564 -> Water_Temp.:102.99, PWM:41
00:25:03.092 -> Water_Temp.:103.10, PWM:41
00:25:05.446 -> Water_Temp.:103.10, PWM:41
00:25:06.211 -> Water_Temp.:103.10, PWM:41
00:25:07.004 -> Water_Temp.:103.10, PWM:41
00:25:07.763 -> Water_Temp.:103.10, PWM:41
00:25:08.561 -> Water_Temp.:103.10, PWM:42
00:25:09.326 -> Water_Temp.:103.10, PWM:41
00:25:10.123 -> Water_Temp.:103.10, PWM:41
00:25:10.886 -> Water_Temp.:103.10, PWM:41
00:25:11.675 -> Water_Temp.:103.10, PWM:41
00:25:12.467 -> Water_Temp.:103.10, PWM:41
00:25:13.234 -> Water_Temp.:103.10, PWM:41
00:25:14.027 -> Water_Temp.:103.10, PWM:41
00:25:15.554 -> Water_Temp.:103.10, PWM:41
00:25:16.318 -> Water_Temp.:103.10, PWM:41
00:25:18.674 -> Water_Temp.:103.10, PWM:41
00:25:20.232 -> Water_Temp.:103.10, PWM:41
00:25:20.995 -> Water_Temp.:103.10, PWM:41
00:25:21.786 -> Water_Temp.:103.10, PWM:41
00:25:22.577 -> Water_Temp.:103.10, PWM:41
00:25:23.341 -> Water_Temp.:103.10, PWM:41
00:25:24.138 -> Water_Temp.:103.10, PWM:41
00:25:26.428 -> Water_Temp.:103.10, PWM:41
00:25:27.191 -> Water_Temp.:103.10, PWM:41
00:25:28.746 -> Water_Temp.:103.10, PWM:41
00:25:29.543 -> Water_Temp.:103.10, PWM:41
00:25:31.070 -> Water_Temp.:103.10, PWM:41
00:25:31.866 -> Water_Temp.:103.21, PWM:41
00:25:32.629 -> Water_Temp.:103.21, PWM:41
00:25:33.423 -> Water_Temp.:103.21, PWM:41
00:25:37.333 -> Water_Temp.:103.21, PWM:41
00:25:38.090 -> Water_Temp.:103.21, PWM:41
00:25:38.884 -> Water_Temp.:103.21, PWM:41
00:25:39.643 -> Water_Temp.:103.21, PWM:41
00:25:40.439 -> Water_Temp.:103.21, PWM:41
00:25:43.516 -> Water_Temp.:103.21, PWM:41
00:25:45.078 -> Water_Temp.:103.21, PWM:41
00:25:45.871 -> Water_Temp.:103.21, PWM:41
00:25:46.630 -> Water_Temp.:103.21, PWM:41
00:25:48.180 -> Water_Temp.:103.21, PWM:41
00:25:50.529 -> Water_Temp.:103.21, PWM:41
00:25:51.287 -> Water_Temp.:103.10, PWM:41
00:25:52.084 -> Water_Temp.:103.10, PWM:41
00:25:52.849 -> Water_Temp.:103.10, PWM:41
00:25:53.645 -> Water_Temp.:103.10, PWM:41
00:25:54.406 -> Water_Temp.:103.10, PWM:41
00:25:55.203 -> Water_Temp.:103.10, PWM:41
00:25:56.760 -> Water_Temp.:104.79, PWM:41
00:25:57.520 -> Water_Temp.:106.81, PWM:41
00:25:58.315 -> Water_Temp.:104.34, PWM:41
00:25:59.113 -> Water_Temp.:104.00, PWM:41
00:26:00.635 -> Water_Temp.:103.55, PWM:41
00:26:02.195 -> Water_Temp.:103.44, PWM:41
00:26:04.524 -> Water_Temp.:103.21, PWM:41
00:26:06.851 -> Water_Temp.:103.21, PWM:41
00:26:07.618 -> Water_Temp.:103.21, PWM:41
00:26:08.415 -> Water_Temp.:103.21, PWM:41
00:26:09.181 -> Water_Temp.:103.21, PWM:41
00:26:10.738 -> Water_Temp.:103.10, PWM:41
00:26:11.536 -> Water_Temp.:103.10, PWM:41
00:26:12.302 -> Water_Temp.:103.21, PWM:41
00:26:13.097 -> Water_Temp.:103.21, PWM:41
00:26:15.383 -> Water_Temp.:103.21, PWM:41
00:26:16.969 -> Water_Temp.:103.21, PWM:41
00:26:17.730 -> Water_Temp.:103.21, PWM:41
00:26:18.530 -> Water_Temp.:103.21, PWM:41
00:26:19.291 -> Water_Temp.:103.21, PWM:41
00:26:21.582 -> Water_Temp.:103.21, PWM:41
00:26:22.374 -> Water_Temp.:103.21, PWM:41
00:26:23.131 -> Water_Temp.:103.21, PWM:41
00:26:23.930 -> Water_Temp.:103.21, PWM:41
00:26:25.455 -> Water_Temp.:103.21, PWM:41
00:26:29.333 -> Water_Temp.:103.21, PWM:41
00:26:30.127 -> Water_Temp.:103.32, PWM:41
00:26:31.684 -> Water_Temp.:103.32, PWM:41
00:26:32.443 -> Water_Temp.:103.32, PWM:41
00:26:33.240 -> Water_Temp.:103.32, PWM:41
00:26:34.005 -> Water_Temp.:103.32, PWM:41
00:26:34.804 -> Water_Temp.:103.32, PWM:41
00:26:37.849 -> Water_Temp.:103.32, PWM:41
00:26:38.645 -> Water_Temp.:103.32, PWM:41
00:26:39.408 -> Water_Temp.:103.32, PWM:41
00:26:40.203 -> Water_Temp.:103.32, PWM:41
00:26:40.965 -> Water_Temp.:103.32, PWM:41
00:26:41.761 -> Water_Temp.:103.32, PWM:41
00:26:42.557 -> Water_Temp.:103.32, PWM:41
00:26:43.318 -> Water_Temp.:103.32, PWM:41
00:26:44.116 -> Water_Temp.:103.32, PWM:41
00:26:44.877 -> Water_Temp.:103.32, PWM:41
00:26:46.435 -> Water_Temp.:103.32, PWM:41
00:26:47.192 -> Water_Temp.:103.32, PWM:41
00:26:47.986 -> Water_Temp.:103.32, PWM:41
00:26:48.751 -> Water_Temp.:103.32, PWM:41
00:26:51.105 -> Water_Temp.:103.32, PWM:41
00:26:53.418 -> Water_Temp.:103.32, PWM:41
00:26:54.183 -> Water_Temp.:103.32, PWM:41
00:26:54.982 -> Water_Temp.:103.32, PWM:41
00:26:55.742 -> Water_Temp.:103.21, PWM:41
00:26:56.541 -> Water_Temp.:103.21, PWM:41
00:26:57.301 -> Water_Temp.:103.21, PWM:41
00:26:58.098 -> Water_Temp.:103.21, PWM:41
00:26:58.860 -> Water_Temp.:103.21, PWM:41
00:26:59.661 -> Water_Temp.:103.21, PWM:41
00:27:00.420 -> Water_Temp.:103.21, PWM:41
00:27:01.215 -> Water_Temp.:103.21, PWM:41
00:27:02.760 -> Water_Temp.:103.21, PWM:41
00:27:03.552 -> Water_Temp.:103.21, PWM:41
00:27:05.863 -> Water_Temp.:103.21, PWM:41
00:27:08.184 -> Water_Temp.:103.32, PWM:41
00:27:08.947 -> Water_Temp.:103.32, PWM:41
00:27:11.295 -> Water_Temp.:103.32, PWM:41
00:27:12.089 -> Water_Temp.:103.32, PWM:41
00:27:13.646 -> Water_Temp.:103.32, PWM:41
00:27:15.169 -> Water_Temp.:103.32, PWM:42
00:27:15.964 -> Water_Temp.:103.32, PWM:41
00:27:16.723 -> Water_Temp.:103.32, PWM:41
00:27:17.514 -> Water_Temp.:103.32, PWM:41
00:27:19.073 -> Water_Temp.:103.32, PWM:41
00:27:20.628 -> Water_Temp.:103.44, PWM:41
00:27:21.390 -> Water_Temp.:103.44, PWM:42
00:27:22.151 -> Water_Temp.:103.44, PWM:41
00:27:22.945 -> Water_Temp.:103.44, PWM:41
00:27:23.742 -> Water_Temp.:103.44, PWM:41
00:27:24.502 -> Water_Temp.:103.44, PWM:41
00:27:25.292 -> Water_Temp.:103.44, PWM:41
00:27:26.817 -> Water_Temp.:103.44, PWM:41
00:27:27.614 -> Water_Temp.:103.44, PWM:41
00:27:28.407 -> Water_Temp.:103.44, PWM:41
00:27:30.697 -> Water_Temp.:103.44, PWM:41
00:27:31.495 -> Water_Temp.:103.44, PWM:41
00:27:32.259 -> Water_Temp.:103.32, PWM:41
00:27:33.021 -> Water_Temp.:103.32, PWM:41
00:27:34.607 -> Water_Temp.:103.32, PWM:41
00:27:35.367 -> Water_Temp.:103.32, PWM:41
00:27:36.160 -> Water_Temp.:103.32, PWM:41
00:27:38.477 -> Water_Temp.:103.32, PWM:41
00:27:39.278 -> Water_Temp.:103.32, PWM:41
00:27:40.837 -> Water_Temp.:103.32, PWM:42
00:27:41.599 -> Water_Temp.:103.32, PWM:41
00:27:42.396 -> Water_Temp.:103.32, PWM:41
00:27:43.159 -> Water_Temp.:103.32, PWM:41
00:27:46.246 -> Water_Temp.:103.32, PWM:41
00:27:47.009 -> Water_Temp.:103.32, PWM:41
00:27:47.802 -> Water_Temp.:103.32, PWM:41
00:27:51.649 -> Water_Temp.:103.32, PWM:41
00:27:52.445 -> Water_Temp.:103.32, PWM:41
00:27:53.210 -> Water_Temp.:103.32, PWM:41
00:27:54.008 -> Water_Temp.:103.32, PWM:41
00:27:57.046 -> Water_Temp.:103.44, PWM:41
00:27:57.839 -> Water_Temp.:103.44, PWM:41
00:27:58.600 -> Water_Temp.:103.44, PWM:41
00:28:02.436 -> Water_Temp.:103.44, PWM:41
00:28:03.996 -> Water_Temp.:103.44, PWM:41
00:28:04.789 -> Water_Temp.:103.44, PWM:41
00:28:05.550 -> Water_Temp.:103.44, PWM:41
00:28:06.346 -> Water_Temp.:103.44, PWM:41
00:28:07.109 -> Water_Temp.:103.44, PWM:41
00:28:08.663 -> Water_Temp.:103.44, PWM:41
00:28:10.224 -> Water_Temp.:103.44, PWM:41
00:28:10.984 -> Water_Temp.:103.44, PWM:41
00:28:11.778 -> Water_Temp.:103.44, PWM:41
00:28:13.336 -> Water_Temp.:103.44, PWM:41
00:28:14.100 -> Water_Temp.:103.44, PWM:42
00:28:14.899 -> Water_Temp.:103.44, PWM:41
00:28:15.663 -> Water_Temp.:103.44, PWM:41
00:28:16.454 -> Water_Temp.:103.44, PWM:41
00:28:18.008 -> Water_Temp.:103.55, PWM:41
00:28:20.364 -> Water_Temp.:103.55, PWM:41
00:28:23.450 -> Water_Temp.:103.55, PWM:41
00:28:24.976 -> Water_Temp.:103.55, PWM:41
00:28:25.770 -> Water_Temp.:103.55, PWM:41
00:28:26.534 -> Water_Temp.:103.55, PWM:41
00:28:27.326 -> Water_Temp.:103.55, PWM:41
00:28:28.089 -> Water_Temp.:103.55, PWM:41
00:28:29.680 -> Water_Temp.:103.55, PWM:41
00:28:30.438 -> Water_Temp.:105.24, PWM:41
00:28:31.235 -> Water_Temp.:107.37, PWM:41
00:28:32.793 -> Water_Temp.:104.90, PWM:41
00:28:33.556 -> Water_Temp.:104.22, PWM:41
00:28:34.354 -> Water_Temp.:104.11, PWM:41
00:28:35.117 -> Water_Temp.:104.00, PWM:41
00:28:35.914 -> Water_Temp.:104.00, PWM:41
00:28:37.437 -> Water_Temp.:103.89, PWM:41
00:28:38.234 -> Water_Temp.:103.77, PWM:41
00:28:38.996 -> Water_Temp.:103.77, PWM:41
00:28:39.792 -> Water_Temp.:103.77, PWM:41
00:28:40.554 -> Water_Temp.:103.77, PWM:41
00:28:41.350 -> Water_Temp.:103.77, PWM:42
00:28:42.110 -> Water_Temp.:103.77, PWM:41
00:28:42.906 -> Water_Temp.:103.77, PWM:41
00:28:43.703 -> Water_Temp.:103.77, PWM:41
00:28:44.464 -> Water_Temp.:103.77, PWM:41
00:28:45.260 -> Water_Temp.:103.77, PWM:41
00:28:46.025 -> Water_Temp.:103.77, PWM:41
00:28:46.822 -> Water_Temp.:103.77, PWM:41
00:28:48.347 -> Water_Temp.:103.77, PWM:41
00:28:49.144 -> Water_Temp.:103.77, PWM:41
00:28:49.939 -> Water_Temp.:103.77, PWM:41
00:28:51.495 -> Water_Temp.:103.77, PWM:41
00:28:52.255 -> Water_Temp.:103.77, PWM:41
00:28:53.051 -> Water_Temp.:103.77, PWM:41
00:28:53.813 -> Water_Temp.:103.77, PWM:41
00:28:54.608 -> Water_Temp.:103.77, PWM:41
00:28:55.371 -> Water_Temp.:103.77, PWM:41
00:28:56.165 -> Water_Temp.:103.77, PWM:41
00:28:56.958 -> Water_Temp.:103.77, PWM:41
00:28:59.278 -> Water_Temp.:103.77, PWM:41
00:29:00.039 -> Water_Temp.:103.77, PWM:41
00:29:00.837 -> Water_Temp.:103.77, PWM:41
00:29:01.595 -> Water_Temp.:103.77, PWM:41
00:29:03.156 -> Water_Temp.:103.77, PWM:41
00:29:04.722 -> Water_Temp.:103.77, PWM:41
00:29:05.519 -> Water_Temp.:103.77, PWM:41
00:29:06.283 -> Water_Temp.:103.77, PWM:41
00:29:08.573 -> Water_Temp.:103.77, PWM:42
00:29:09.372 -> Water_Temp.:103.77, PWM:41
00:29:10.135 -> Water_Temp.:103.77, PWM:41
00:29:10.933 -> Water_Temp.:103.77, PWM:41
00:29:11.696 -> Water_Temp.:103.77, PWM:41
00:29:12.492 -> Water_Temp.:103.77, PWM:41
00:29:15.551 -> Water_Temp.:103.77, PWM:41
00:29:16.316 -> Water_Temp.:103.66, PWM:41
00:29:17.107 -> Water_Temp.:103.66, PWM:41
00:29:17.902 -> Water_Temp.:103.66, PWM:41
00:29:18.663 -> Water_Temp.:103.66, PWM:41
00:29:19.460 -> Water_Temp.:103.66, PWM:41
00:29:20.224 -> Water_Temp.:103.66, PWM:41
00:29:21.025 -> Water_Temp.:103.66, PWM:41
00:29:22.550 -> Water_Temp.:103.66, PWM:41
00:29:23.346 -> Water_Temp.:103.66, PWM:41
00:29:24.110 -> Water_Temp.:103.66, PWM:41
00:29:24.875 -> Water_Temp.:103.66, PWM:41
00:29:26.466 -> Water_Temp.:103.66, PWM:41
00:29:27.228 -> Water_Temp.:103.66, PWM:41
00:29:28.025 -> Water_Temp.:103.66, PWM:41
00:29:28.784 -> Water_Temp.:103.66, PWM:41
00:29:29.580 -> Water_Temp.:103.66, PWM:41
00:29:31.894 -> Water_Temp.:103.66, PWM:41
00:29:32.658 -> Water_Temp.:103.66, PWM:41
00:29:33.454 -> Water_Temp.:103.66, PWM:41
00:29:34.217 -> Water_Temp.:103.66, PWM:41
00:29:35.013 -> Water_Temp.:103.66, PWM:41
00:29:36.542 -> Water_Temp.:103.77, PWM:41
00:29:38.106 -> Water_Temp.:103.77, PWM:41
00:29:39.636 -> Water_Temp.:103.77, PWM:41
00:29:40.436 -> Water_Temp.:103.77, PWM:41
00:29:44.320 -> Water_Temp.:103.77, PWM:41
00:29:45.117 -> Water_Temp.:103.77, PWM:41
00:29:45.881 -> Water_Temp.:103.77, PWM:41
00:29:46.681 -> Water_Temp.:103.77, PWM:41
00:29:47.442 -> Water_Temp.:103.77, PWM:41
00:29:48.239 -> Water_Temp.:103.77, PWM:41
00:29:49.004 -> Water_Temp.:103.77, PWM:41
00:29:49.801 -> Water_Temp.:103.77, PWM:41
00:29:52.885 -> Water_Temp.:103.77, PWM:41
00:29:53.652 -> Water_Temp.:103.77, PWM:41
00:29:54.447 -> Water_Temp.:103.66, PWM:41
00:29:55.207 -> Water_Temp.:103.77, PWM:41
00:29:56.763 -> Water_Temp.:103.66, PWM:41
00:29:57.562 -> Water_Temp.:103.66, PWM:41
00:29:59.126 -> Water_Temp.:103.66, PWM:41
00:29:59.885 -> Water_Temp.:103.66, PWM:41
00:30:01.440 -> Water_Temp.:103.66, PWM:41
00:30:02.198 -> Water_Temp.:103.66, PWM:41
00:30:02.992 -> Water_Temp.:103.66, PWM:42
00:30:04.520 -> Water_Temp.:103.66, PWM:41
00:30:05.316 -> Water_Temp.:103.66, PWM:41
00:30:06.078 -> Water_Temp.:103.66, PWM:41
00:30:06.872 -> Water_Temp.:103.66, PWM:41
00:30:09.189 -> Water_Temp.:103.66, PWM:41
00:30:09.986 -> Water_Temp.:103.66, PWM:41
00:30:10.782 -> Water_Temp.:103.66, PWM:41
00:30:11.542 -> Water_Temp.:105.35, PWM:42
00:30:12.338 -> Water_Temp.:107.49, PWM:41
00:30:13.100 -> Water_Temp.:105.01, PWM:41
00:30:13.894 -> Water_Temp.:104.56, PWM:41
00:30:15.417 -> Water_Temp.:104.22, PWM:41
00:30:16.214 -> Water_Temp.:104.22, PWM:41
00:30:17.006 -> Water_Temp.:104.22, PWM:41
00:30:17.767 -> Water_Temp.:104.00, PWM:41
00:30:19.325 -> Water_Temp.:104.00, PWM:41
00:30:20.878 -> Water_Temp.:104.00, PWM:41
00:30:21.670 -> Water_Temp.:104.00, PWM:41
00:30:22.464 -> Water_Temp.:104.00, PWM:41
00:30:25.506 -> Water_Temp.:103.89, PWM:41
00:30:26.299 -> Water_Temp.:103.89, PWM:41
00:30:27.088 -> Water_Temp.:103.89, PWM:41
00:30:29.370 -> Water_Temp.:103.89, PWM:41
00:30:30.164 -> Water_Temp.:103.89, PWM:41
00:30:30.930 -> Water_Temp.:103.89, PWM:41
00:30:31.723 -> Water_Temp.:103.89, PWM:41
00:30:32.485 -> Water_Temp.:103.89, PWM:41
00:30:33.280 -> Water_Temp.:103.89, PWM:41
00:30:34.844 -> Water_Temp.:103.89, PWM:41
00:30:35.607 -> Water_Temp.:103.89, PWM:41
00:30:36.402 -> Water_Temp.:103.89, PWM:41
00:30:37.957 -> Water_Temp.:103.89, PWM:42
00:30:41.042 -> Water_Temp.:103.77, PWM:41
00:30:41.835 -> Water_Temp.:103.89, PWM:42
00:30:42.595 -> Water_Temp.:103.89, PWM:41
00:30:44.155 -> Water_Temp.:103.89, PWM:41
00:30:44.952 -> Water_Temp.:103.89, PWM:41
00:30:46.475 -> Water_Temp.:103.89, PWM:41
00:30:48.034 -> Water_Temp.:103.89, PWM:41
00:30:49.595 -> Water_Temp.:103.89, PWM:41
00:30:50.389 -> Water_Temp.:103.89, PWM:41
00:30:51.152 -> Water_Temp.:103.89, PWM:41
00:30:51.944 -> Water_Temp.:104.00, PWM:41
00:30:52.736 -> Water_Temp.:103.89, PWM:41
00:30:54.295 -> Water_Temp.:104.00, PWM:41
00:30:55.059 -> Water_Temp.:104.00, PWM:41
00:30:55.825 -> Water_Temp.:104.00, PWM:41
00:30:56.622 -> Water_Temp.:104.00, PWM:41
00:30:57.380 -> Water_Temp.:104.00, PWM:41
00:30:58.179 -> Water_Temp.:104.00, PWM:41
00:31:01.287 -> Water_Temp.:104.00, PWM:41
00:31:02.811 -> Water_Temp.:104.00, PWM:41
00:31:03.607 -> Water_Temp.:104.00, PWM:41
00:31:04.372 -> Water_Temp.:104.00, PWM:41
00:31:05.169 -> Water_Temp.:104.00, PWM:41
00:31:06.689 -> Water_Temp.:104.00, PWM:41
00:31:07.488 -> Water_Temp.:104.00, PWM:41
00:31:08.251 -> Water_Temp.:104.00, PWM:41
00:31:09.809 -> Water_Temp.:104.00, PWM:41
00:31:10.605 -> Water_Temp.:104.00, PWM:41
00:31:11.364 -> Water_Temp.:104.00, PWM:41
00:31:13.682 -> Water_Temp.:104.00, PWM:41
00:31:14.481 -> Water_Temp.:104.00, PWM:41
00:31:15.237 -> Water_Temp.:104.00, PWM:41
00:31:16.793 -> Water_Temp.:104.00, PWM:41
00:31:17.555 -> Water_Temp.:104.00, PWM:41
00:31:20.696 -> Water_Temp.:104.11, PWM:41
00:31:21.459 -> Water_Temp.:104.11, PWM:41
00:31:22.251 -> Water_Temp.:104.11, PWM:41
00:31:23.012 -> Water_Temp.:104.11, PWM:41
00:31:24.565 -> Water_Temp.:104.22, PWM:41
00:31:26.080 -> Water_Temp.:104.22, PWM:41
00:31:26.879 -> Water_Temp.:104.22, PWM:41
00:31:29.225 -> Water_Temp.:104.11, PWM:41
00:31:29.984 -> Water_Temp.:104.11, PWM:41
00:31:30.778 -> Water_Temp.:104.11, PWM:41
00:31:31.541 -> Water_Temp.:104.11, PWM:41
00:31:33.105 -> Water_Temp.:104.11, PWM:41
00:31:33.901 -> Water_Temp.:104.11, PWM:41
00:31:34.662 -> Water_Temp.:104.11, PWM:41
00:31:36.214 -> Water_Temp.:104.11, PWM:41
00:31:37.761 -> Water_Temp.:104.11, PWM:41
00:31:39.319 -> Water_Temp.:104.11, PWM:41
00:31:40.111 -> Water_Temp.:104.11, PWM:41
00:31:41.669 -> Water_Temp.:104.11, PWM:41
00:31:44.713 -> Water_Temp.:104.22, PWM:41
00:31:45.504 -> Water_Temp.:104.22, PWM:41
00:31:47.062 -> Water_Temp.:104.22, PWM:41
00:31:47.855 -> Water_Temp.:104.22, PWM:41
00:31:48.613 -> Water_Temp.:104.22, PWM:41
00:31:50.172 -> Water_Temp.:104.22, PWM:41
00:31:50.970 -> Water_Temp.:104.22, PWM:41
00:31:51.725 -> Water_Temp.:104.22, PWM:41
00:31:53.289 -> Water_Temp.:104.22, PWM:41
00:31:54.050 -> Water_Temp.:104.22, PWM:41
00:31:56.405 -> Water_Temp.:104.22, PWM:41
00:31:57.197 -> Water_Temp.:104.22, PWM:41
00:31:57.960 -> Water_Temp.:104.22, PWM:41
00:31:58.724 -> Water_Temp.:104.22, PWM:41
00:31:59.524 -> Water_Temp.:104.22, PWM:41
00:32:00.287 -> Water_Temp.:104.22, PWM:41
00:32:01.083 -> Water_Temp.:104.22, PWM:41
00:32:02.639 -> Water_Temp.:104.22, PWM:41
00:32:03.399 -> Water_Temp.:104.22, PWM:41
00:32:04.197 -> Water_Temp.:104.22, PWM:41
00:32:04.959 -> Water_Temp.:104.22, PWM:41
00:32:05.754 -> Water_Temp.:104.22, PWM:41
00:32:06.516 -> Water_Temp.:104.22, PWM:41
00:32:08.078 -> Water_Temp.:104.34, PWM:41
00:32:08.843 -> Water_Temp.:104.34, PWM:41
00:32:11.193 -> Water_Temp.:104.34, PWM:41
00:32:11.956 -> Water_Temp.:104.34, PWM:41
00:32:12.753 -> Water_Temp.:104.34, PWM:41
00:32:13.546 -> Water_Temp.:104.34, PWM:41
00:32:14.309 -> Water_Temp.:104.34, PWM:41
00:32:15.105 -> Water_Temp.:104.34, PWM:41
00:32:16.623 -> Water_Temp.:104.45, PWM:41
00:32:17.420 -> Water_Temp.:104.45, PWM:42
00:32:18.179 -> Water_Temp.:104.45, PWM:41
00:32:18.975 -> Water_Temp.:104.45, PWM:41
00:32:20.533 -> Water_Temp.:104.45, PWM:41
00:32:22.082 -> Water_Temp.:104.45, PWM:41
00:32:24.401 -> Water_Temp.:104.45, PWM:41
00:32:25.203 -> Water_Temp.:104.45, PWM:41
00:32:25.965 -> Water_Temp.:104.45, PWM:41
00:32:27.518 -> Water_Temp.:104.34, PWM:41
00:32:28.279 -> Water_Temp.:104.34, PWM:41
00:32:29.075 -> Water_Temp.:104.34, PWM:41
00:32:29.838 -> Water_Temp.:104.34, PWM:41
00:32:30.634 -> Water_Temp.:104.34, PWM:41
00:32:31.429 -> Water_Temp.:104.34, PWM:41
00:32:32.189 -> Water_Temp.:104.45, PWM:41
00:32:34.507 -> Water_Temp.:104.45, PWM:41
00:32:35.305 -> Water_Temp.:104.45, PWM:41
00:32:36.827 -> Water_Temp.:104.45, PWM:41
00:32:37.621 -> Water_Temp.:104.45, PWM:41
00:32:38.382 -> Water_Temp.:104.45, PWM:41
00:32:39.908 -> Water_Temp.:104.45, PWM:41
00:32:40.706 -> Water_Temp.:104.45, PWM:41
00:32:41.469 -> Water_Temp.:104.45, PWM:41
00:32:43.021 -> Water_Temp.:104.45, PWM:41
00:32:45.347 -> Water_Temp.:104.45, PWM:41
00:32:46.911 -> Water_Temp.:104.45, PWM:41
00:32:48.471 -> Water_Temp.:104.45, PWM:41
00:32:49.269 -> Water_Temp.:104.45, PWM:41
00:32:50.031 -> Water_Temp.:104.45, PWM:41
00:32:50.825 -> Water_Temp.:104.45, PWM:41
00:32:51.584 -> Water_Temp.:104.45, PWM:41
00:32:52.380 -> Water_Temp.:104.45, PWM:41
00:32:53.139 -> Water_Temp.:104.45, PWM:41
00:32:53.935 -> Water_Temp.:104.45, PWM:41
00:32:54.700 -> Water_Temp.:104.45, PWM:41
00:32:55.500 -> Water_Temp.:104.45, PWM:41
00:32:56.264 -> Water_Temp.:104.45, PWM:41
00:32:57.054 -> Water_Temp.:104.45, PWM:41
00:32:58.577 -> Water_Temp.:104.45, PWM:41
00:32:59.371 -> Water_Temp.:104.45, PWM:41
00:33:00.169 -> Water_Temp.:104.45, PWM:41
00:33:00.929 -> Water_Temp.:104.45, PWM:41
00:33:02.448 -> Water_Temp.:104.45, PWM:41
00:33:03.246 -> Water_Temp.:104.45, PWM:41
00:33:04.006 -> Water_Temp.:104.45, PWM:41
00:33:05.598 -> Water_Temp.:104.45, PWM:41
00:33:06.364 -> Water_Temp.:104.45, PWM:41
00:33:07.127 -> Water_Temp.:104.45, PWM:41
00:33:07.925 -> Water_Temp.:104.45, PWM:41
00:33:09.452 -> Water_Temp.:104.45, PWM:41
00:33:10.245 -> Water_Temp.:104.45, PWM:42
00:33:11.005 -> Water_Temp.:104.45, PWM:41
00:33:12.566 -> Water_Temp.:104.45, PWM:41
00:33:13.326 -> Water_Temp.:104.45, PWM:41
00:33:14.120 -> Water_Temp.:104.45, PWM:41
00:33:14.886 -> Water_Temp.:104.45, PWM:41
00:33:15.683 -> Water_Temp.:104.45, PWM:41
00:33:16.442 -> Water_Temp.:104.45, PWM:41
00:33:17.236 -> Water_Temp.:104.45, PWM:41
00:33:18.000 -> Water_Temp.:104.45, PWM:41
00:33:18.795 -> Water_Temp.:104.45, PWM:41
00:33:20.320 -> Water_Temp.:104.45, PWM:41
00:33:21.116 -> Water_Temp.:104.45, PWM:41
00:33:21.876 -> Water_Temp.:104.45, PWM:41
00:33:22.672 -> Water_Temp.:104.45, PWM:41
00:33:23.434 -> Water_Temp.:104.45, PWM:41
00:33:24.232 -> Water_Temp.:104.45, PWM:41
00:33:26.516 -> Water_Temp.:104.56, PWM:41
00:33:27.311 -> Water_Temp.:104.56, PWM:41
00:33:28.077 -> Water_Temp.:104.56, PWM:41
00:33:28.876 -> Water_Temp.:104.56, PWM:41
00:33:29.638 -> Water_Temp.:104.56, PWM:41
00:33:30.434 -> Water_Temp.:104.56, PWM:41
00:33:31.956 -> Water_Temp.:104.56, PWM:41
00:33:32.748 -> Water_Temp.:104.56, PWM:41
00:33:34.303 -> Water_Temp.:104.67, PWM:41
00:33:35.060 -> Water_Temp.:104.56, PWM:41
00:33:35.855 -> Water_Temp.:104.56, PWM:41
00:33:36.615 -> Water_Temp.:104.56, PWM:41
00:33:37.406 -> Water_Temp.:104.56, PWM:41
00:33:38.932 -> Water_Temp.:104.56, PWM:41
00:33:39.731 -> Water_Temp.:104.56, PWM:41
00:33:40.495 -> Water_Temp.:104.56, PWM:41
00:33:41.293 -> Water_Temp.:104.56, PWM:41
00:33:42.847 -> Water_Temp.:104.56, PWM:41
00:33:43.642 -> Water_Temp.:104.56, PWM:41
00:33:44.402 -> Water_Temp.:104.56, PWM:41
00:33:45.963 -> Water_Temp.:104.56, PWM:41
00:33:48.282 -> Water_Temp.:104.56, PWM:41
00:33:49.079 -> Water_Temp.:104.56, PWM:41
00:33:49.844 -> Water_Temp.:104.56, PWM:41
00:33:50.637 -> Water_Temp.:104.45, PWM:41
00:33:51.397 -> Water_Temp.:104.45, PWM:41
00:33:54.490 -> Water_Temp.:104.45, PWM:41
00:33:55.287 -> Water_Temp.:104.45, PWM:41
00:33:56.047 -> Water_Temp.:104.45, PWM:41
00:33:57.605 -> Water_Temp.:104.45, PWM:41
00:33:58.401 -> Water_Temp.:104.45, PWM:41
00:33:59.160 -> Water_Temp.:104.45, PWM:41
00:34:01.481 -> Water_Temp.:104.45, PWM:41
00:34:03.006 -> Water_Temp.:104.45, PWM:41
00:34:03.801 -> Water_Temp.:104.45, PWM:41
00:34:05.360 -> Water_Temp.:104.45, PWM:42
00:34:08.470 -> Water_Temp.:104.45, PWM:41
00:34:09.267 -> Water_Temp.:104.45, PWM:41
00:34:10.827 -> Water_Temp.:104.45, PWM:41
00:34:11.589 -> Water_Temp.:104.45, PWM:41
00:34:13.149 -> Water_Temp.:104.56, PWM:41
00:34:13.948 -> Water_Temp.:104.56, PWM:41
00:34:14.715 -> Water_Temp.:104.56, PWM:41
00:34:16.277 -> Water_Temp.:104.56, PWM:41
00:34:18.558 -> Water_Temp.:104.56, PWM:41
00:34:19.356 -> Water_Temp.:104.56, PWM:41
00:34:20.118 -> Water_Temp.:104.56, PWM:41
00:34:21.678 -> Water_Temp.:104.67, PWM:41
00:34:22.475 -> Water_Temp.:104.67, PWM:41
00:34:23.234 -> Water_Temp.:104.67, PWM:41
00:34:24.795 -> Water_Temp.:104.67, PWM:41
00:34:25.591 -> Water_Temp.:104.67, PWM:41
00:34:26.353 -> Water_Temp.:104.79, PWM:41
00:34:27.148 -> Water_Temp.:104.79, PWM:41
00:34:28.704 -> Water_Temp.:104.79, PWM:41
00:34:29.466 -> Water_Temp.:104.79, PWM:41
00:34:30.262 -> Water_Temp.:104.79, PWM:41
00:34:31.817 -> Water_Temp.:104.79, PWM:41
00:34:33.371 -> Water_Temp.:104.79, PWM:41
00:34:34.134 -> Water_Temp.:104.79, PWM:41
00:34:34.931 -> Water_Temp.:104.79, PWM:41
00:34:37.216 -> Water_Temp.:104.79, PWM:41
00:34:38.014 -> Water_Temp.:104.79, PWM:41
00:34:39.568 -> Water_Temp.:104.79, PWM:41
00:34:40.324 -> Water_Temp.:104.79, PWM:41
00:34:41.121 -> Water_Temp.:104.79, PWM:41
00:34:41.883 -> Water_Temp.:104.79, PWM:41
00:34:42.679 -> Water_Temp.:104.79, PWM:41
00:34:44.234 -> Water_Temp.:104.79, PWM:41
00:34:44.995 -> Water_Temp.:104.79, PWM:41
00:34:45.795 -> Water_Temp.:104.67, PWM:41
00:34:48.909 -> Water_Temp.:104.67, PWM:41
00:34:50.467 -> Water_Temp.:104.67, PWM:41
00:34:51.262 -> Water_Temp.:104.67, PWM:41
00:34:52.022 -> Water_Temp.:104.67, PWM:41
00:34:52.818 -> Water_Temp.:104.67, PWM:41
00:34:53.582 -> Water_Temp.:104.67, PWM:41
00:34:54.381 -> Water_Temp.:104.67, PWM:41
00:34:55.144 -> Water_Temp.:104.67, PWM:41
00:34:55.907 -> Water_Temp.:104.67, PWM:41
00:34:56.704 -> Water_Temp.:104.67, PWM:41
00:34:57.501 -> Water_Temp.:104.67, PWM:41
00:34:59.058 -> Water_Temp.:104.67, PWM:41
00:34:59.823 -> Water_Temp.:104.67, PWM:41
00:35:01.379 -> Water_Temp.:104.67, PWM:41
00:35:02.934 -> Water_Temp.:106.25, PWM:41
00:35:03.734 -> Water_Temp.:108.39, PWM:41
00:35:05.293 -> Water_Temp.:105.57, PWM:41
00:35:06.055 -> Water_Temp.:105.35, PWM:41
00:35:07.615 -> Water_Temp.:105.12, PWM:41
00:35:09.180 -> Water_Temp.:105.01, PWM:41
00:35:09.944 -> Water_Temp.:105.01, PWM:41
00:35:10.745 -> Water_Temp.:105.01, PWM:41
00:35:11.504 -> Water_Temp.:104.90, PWM:41
00:35:12.301 -> Water_Temp.:104.90, PWM:41
00:35:14.598 -> Water_Temp.:104.90, PWM:41
00:35:15.395 -> Water_Temp.:104.90, PWM:41
00:35:16.160 -> Water_Temp.:104.90, PWM:41
00:35:17.717 -> Water_Temp.:104.90, PWM:41
00:35:18.480 -> Water_Temp.:104.90, PWM:41
00:35:19.248 -> Water_Temp.:104.79, PWM:41
00:35:20.044 -> Water_Temp.:104.79, PWM:41
00:35:21.570 -> Water_Temp.:104.79, PWM:41
00:35:23.130 -> Water_Temp.:104.79, PWM:41
00:35:23.928 -> Water_Temp.:104.79, PWM:41
00:35:24.692 -> Water_Temp.:104.79, PWM:41
00:35:26.251 -> Water_Temp.:104.79, PWM:41
00:35:27.046 -> Water_Temp.:104.67, PWM:41
00:35:27.805 -> Water_Temp.:104.67, PWM:41
00:35:29.364 -> Water_Temp.:104.67, PWM:41
00:35:30.125 -> Water_Temp.:104.67, PWM:41
00:35:30.923 -> Water_Temp.:104.67, PWM:41
00:35:31.717 -> Water_Temp.:104.67, PWM:41
00:35:32.480 -> Water_Temp.:104.56, PWM:41
00:35:33.274 -> Water_Temp.:104.56, PWM:41
00:35:34.034 -> Water_Temp.:104.67, PWM:41
00:35:34.833 -> Water_Temp.:104.67, PWM:41
00:35:35.593 -> Water_Temp.:104.67, PWM:41
00:35:40.934 -> Water_Temp.:104.67, PWM:41
00:35:41.729 -> Water_Temp.:104.67, PWM:41
00:35:42.520 -> Water_Temp.:104.67, PWM:41
00:35:44.081 -> Water_Temp.:104.67, PWM:41
00:35:44.840 -> Water_Temp.:104.67, PWM:42
00:35:46.399 -> Water_Temp.:104.67, PWM:41
00:35:47.159 -> Water_Temp.:104.67, PWM:41
00:35:48.715 -> Water_Temp.:104.67, PWM:41
00:35:49.515 -> Water_Temp.:104.79, PWM:41
00:35:50.277 -> Water_Temp.:104.79, PWM:41
00:35:51.069 -> Water_Temp.:104.79, PWM:41
00:35:51.865 -> Water_Temp.:104.79, PWM:41
00:35:52.628 -> Water_Temp.:104.79, PWM:41
00:35:53.422 -> Water_Temp.:104.79, PWM:41
00:35:54.951 -> Water_Temp.:104.79, PWM:41
00:35:55.747 -> Water_Temp.:104.79, PWM:42
00:35:56.509 -> Water_Temp.:104.79, PWM:41
00:35:57.304 -> Water_Temp.:104.79, PWM:41
00:35:58.867 -> Water_Temp.:104.79, PWM:41
00:36:00.392 -> Water_Temp.:104.79, PWM:41
00:36:01.190 -> Water_Temp.:104.79, PWM:41
00:36:01.949 -> Water_Temp.:104.79, PWM:41
00:36:03.505 -> Water_Temp.:104.79, PWM:41
00:36:05.067 -> Water_Temp.:104.79, PWM:41
00:36:05.860 -> Water_Temp.:104.67, PWM:41
00:36:06.624 -> Water_Temp.:104.79, PWM:41
00:36:07.425 -> Water_Temp.:104.79, PWM:41
00:36:08.987 -> Water_Temp.:104.79, PWM:41
00:36:09.746 -> Water_Temp.:104.67, PWM:41
00:36:10.543 -> Water_Temp.:104.79, PWM:41
00:36:11.307 -> Water_Temp.:104.79, PWM:41
00:36:12.105 -> Water_Temp.:104.79, PWM:41
00:36:12.869 -> Water_Temp.:104.79, PWM:41
00:36:14.433 -> Water_Temp.:104.79, PWM:41
00:36:15.197 -> Water_Temp.:104.79, PWM:41
00:36:15.989 -> Water_Temp.:104.79, PWM:41
00:36:16.785 -> Water_Temp.:104.79, PWM:41
00:36:18.341 -> Water_Temp.:104.90, PWM:41
00:36:19.106 -> Water_Temp.:104.90, PWM:41
00:36:19.900 -> Water_Temp.:104.90, PWM:41
00:36:20.663 -> Water_Temp.:104.90, PWM:41
00:36:21.456 -> Water_Temp.:104.90, PWM:41
00:36:23.774 -> Water_Temp.:104.90, PWM:41
00:36:24.571 -> Water_Temp.:104.90, PWM:41
00:36:26.096 -> Water_Temp.:104.90, PWM:41
00:36:26.892 -> Water_Temp.:104.90, PWM:41
00:36:27.686 -> Water_Temp.:104.90, PWM:41
00:36:28.447 -> Water_Temp.:104.90, PWM:41
00:36:30.002 -> Water_Temp.:104.90, PWM:41
00:36:31.558 -> Water_Temp.:104.79, PWM:41
00:36:32.353 -> Water_Temp.:104.79, PWM:41
00:36:33.907 -> Water_Temp.:104.90, PWM:41
00:36:34.670 -> Water_Temp.:104.90, PWM:41
00:36:35.464 -> Water_Temp.:104.90, PWM:41
00:36:36.226 -> Water_Temp.:104.90, PWM:41
00:36:37.781 -> Water_Temp.:104.90, PWM:41
00:36:38.581 -> Water_Temp.:104.90, PWM:41
00:36:39.348 -> Water_Temp.:104.90, PWM:41
00:36:40.906 -> Water_Temp.:105.01, PWM:41
00:36:42.464 -> Water_Temp.:104.90, PWM:41
00:36:43.228 -> Water_Temp.:104.90, PWM:41
00:36:44.026 -> Water_Temp.:104.90, PWM:41
00:36:44.781 -> Water_Temp.:104.90, PWM:41
00:36:45.578 -> Water_Temp.:104.90, PWM:41
00:36:46.342 -> Water_Temp.:104.90, PWM:41
00:36:48.661 -> Water_Temp.:104.90, PWM:41
00:36:50.983 -> Water_Temp.:104.90, PWM:41
00:36:51.748 -> Water_Temp.:104.90, PWM:41
00:36:54.070 -> Water_Temp.:104.90, PWM:41
00:36:55.632 -> Water_Temp.:104.90, PWM:41
00:36:56.428 -> Water_Temp.:104.90, PWM:41
00:36:57.221 -> Water_Temp.:104.90, PWM:41
00:36:57.982 -> Water_Temp.:104.90, PWM:41
00:36:58.775 -> Water_Temp.:104.90, PWM:41
00:36:59.538 -> Water_Temp.:104.90, PWM:42
00:37:00.334 -> Water_Temp.:104.90, PWM:41
00:37:01.854 -> Water_Temp.:104.90, PWM:41
00:37:02.648 -> Water_Temp.:104.90, PWM:41
00:37:03.409 -> Water_Temp.:104.90, PWM:41
00:37:04.203 -> Water_Temp.:104.90, PWM:41
00:37:05.758 -> Water_Temp.:105.01, PWM:41
00:37:06.554 -> Water_Temp.:105.01, PWM:41
00:37:07.319 -> Water_Temp.:105.01, PWM:41
00:37:08.111 -> Water_Temp.:105.01, PWM:41
00:37:08.874 -> Water_Temp.:105.01, PWM:41
00:37:10.430 -> Water_Temp.:105.01, PWM:41
00:37:11.190 -> Water_Temp.:105.01, PWM:41
00:37:11.989 -> Water_Temp.:105.01, PWM:41
00:37:13.512 -> Water_Temp.:105.01, PWM:41
00:37:15.079 -> Water_Temp.:105.01, PWM:41
00:37:15.875 -> Water_Temp.:105.01, PWM:41
00:37:16.636 -> Water_Temp.:105.01, PWM:41
00:37:18.221 -> Water_Temp.:105.01, PWM:41
00:37:18.983 -> Water_Temp.:105.01, PWM:41
00:37:19.782 -> Water_Temp.:105.01, PWM:41
00:37:21.336 -> Water_Temp.:105.01, PWM:41
00:37:25.941 -> Water_Temp.:105.01, PWM:41
00:37:26.704 -> Water_Temp.:105.01, PWM:41
00:37:27.498 -> Water_Temp.:105.01, PWM:41
00:37:29.050 -> Water_Temp.:105.12, PWM:41
00:37:29.812 -> Water_Temp.:105.12, PWM:41
00:37:30.609 -> Water_Temp.:105.12, PWM:41
00:37:32.164 -> Water_Temp.:105.12, PWM:41
00:37:32.960 -> Water_Temp.:105.12, PWM:41
00:37:34.483 -> Water_Temp.:105.01, PWM:41
00:37:35.274 -> Water_Temp.:105.01, PWM:41
00:37:36.033 -> Water_Temp.:105.01, PWM:41
00:37:36.832 -> Water_Temp.:105.01, PWM:41
00:37:37.597 -> Water_Temp.:105.01, PWM:41
00:37:39.158 -> Water_Temp.:105.01, PWM:41
00:37:39.918 -> Water_Temp.:105.01, PWM:41
00:37:40.713 -> Water_Temp.:105.12, PWM:41
00:37:41.477 -> Water_Temp.:105.12, PWM:41
00:37:42.273 -> Water_Temp.:105.12, PWM:41
00:37:43.068 -> Water_Temp.:105.12, PWM:41
00:37:44.595 -> Water_Temp.:105.12, PWM:41
00:37:46.181 -> Water_Temp.:105.12, PWM:41
00:37:46.945 -> Water_Temp.:105.12, PWM:41
00:37:48.501 -> Water_Temp.:105.12, PWM:41
00:37:49.265 -> Water_Temp.:105.12, PWM:41
00:37:50.059 -> Water_Temp.:105.12, PWM:41
00:37:50.820 -> Water_Temp.:105.12, PWM:41
00:37:54.665 -> Water_Temp.:105.01, PWM:41
00:37:55.459 -> Water_Temp.:105.01, PWM:41
00:37:56.224 -> Water_Temp.:105.01, PWM:41
00:38:00.101 -> Water_Temp.:105.12, PWM:41
00:38:00.897 -> Water_Temp.:105.12, PWM:41
00:38:01.659 -> Water_Temp.:105.12, PWM:41
00:38:03.213 -> Water_Temp.:105.12, PWM:41
00:38:04.011 -> Water_Temp.:105.12, PWM:41
00:38:04.775 -> Water_Temp.:105.12, PWM:41
00:38:05.573 -> Water_Temp.:105.12, PWM:41
00:38:07.091 -> Water_Temp.:105.12, PWM:41
00:38:07.888 -> Water_Temp.:105.12, PWM:44
00:38:08.648 -> Water_Temp.:105.12, PWM:48
00:38:09.447 -> Water_Temp.:105.12, PWM:48
00:38:10.210 -> Water_Temp.:105.12, PWM:48
00:38:11.768 -> Water_Temp.:105.12, PWM:48
00:38:13.327 -> Water_Temp.:105.12, PWM:45
00:38:14.126 -> Water_Temp.:105.12, PWM:45
00:38:15.650 -> Water_Temp.:105.12, PWM:45
00:38:16.447 -> Water_Temp.:105.12, PWM:45
00:38:17.237 -> Water_Temp.:105.12, PWM:45
00:38:19.528 -> Water_Temp.:105.12, PWM:45
00:38:20.291 -> Water_Temp.:105.12, PWM:45
00:38:21.087 -> Water_Temp.:105.12, PWM:45
00:38:21.850 -> Water_Temp.:105.24, PWM:44
00:38:22.646 -> Water_Temp.:105.24, PWM:45
00:38:23.408 -> Water_Temp.:105.24, PWM:45
00:38:28.023 -> Water_Temp.:105.24, PWM:45
00:38:29.585 -> Water_Temp.:105.24, PWM:45
00:38:32.696 -> Water_Temp.:105.24, PWM:45
00:38:33.457 -> Water_Temp.:105.24, PWM:45
00:38:34.257 -> Water_Temp.:105.24, PWM:45
00:38:35.052 -> Water_Temp.:105.24, PWM:45
00:38:36.574 -> Water_Temp.:105.24, PWM:45
00:38:37.366 -> Water_Temp.:105.24, PWM:45
00:38:38.129 -> Water_Temp.:105.24, PWM:45
00:38:38.929 -> Water_Temp.:105.24, PWM:45
00:38:39.694 -> Water_Temp.:105.24, PWM:45
00:38:40.490 -> Water_Temp.:105.24, PWM:45
00:38:41.250 -> Water_Temp.:105.24, PWM:45
00:38:42.776 -> Water_Temp.:105.24, PWM:44
00:38:43.573 -> Water_Temp.:105.24, PWM:45
00:38:44.339 -> Water_Temp.:105.24, PWM:45
00:38:45.137 -> Water_Temp.:105.24, PWM:45
00:38:46.656 -> Water_Temp.:105.24, PWM:45
00:38:47.448 -> Water_Temp.:105.24, PWM:45
00:38:48.206 -> Water_Temp.:105.24, PWM:45
00:38:49.002 -> Water_Temp.:105.24, PWM:45
00:38:49.800 -> Water_Temp.:105.24, PWM:45
00:38:50.561 -> Water_Temp.:105.24, PWM:45
00:38:51.360 -> Water_Temp.:105.24, PWM:45
00:38:54.383 -> Water_Temp.:105.35, PWM:45
00:38:55.181 -> Water_Temp.:105.35, PWM:45
00:38:55.941 -> Water_Temp.:105.35, PWM:45
00:38:56.738 -> Water_Temp.:105.35, PWM:45
00:38:58.296 -> Water_Temp.:105.35, PWM:45
00:38:59.093 -> Water_Temp.:105.35, PWM:45
00:38:59.854 -> Water_Temp.:106.92, PWM:44
00:39:00.645 -> Water_Temp.:109.06, PWM:45
00:39:01.404 -> Water_Temp.:106.59, PWM:45
00:39:02.198 -> Water_Temp.:106.14, PWM:45
00:39:05.286 -> Water_Temp.:105.69, PWM:45
00:39:06.080 -> Water_Temp.:105.69, PWM:45
00:39:06.874 -> Water_Temp.:105.57, PWM:45
00:39:07.637 -> Water_Temp.:105.57, PWM:45
00:39:08.435 -> Water_Temp.:105.57, PWM:45
00:39:09.195 -> Water_Temp.:105.57, PWM:45
00:39:10.756 -> Water_Temp.:105.57, PWM:44
00:39:13.114 -> Water_Temp.:105.57, PWM:45
00:39:13.874 -> Water_Temp.:105.46, PWM:45
00:39:14.673 -> Water_Temp.:105.46, PWM:45
00:39:15.436 -> Water_Temp.:105.46, PWM:45
00:39:16.202 -> Water_Temp.:105.46, PWM:45
00:39:17.000 -> Water_Temp.:105.46, PWM:44
00:39:17.790 -> Water_Temp.:105.46, PWM:45
00:39:20.110 -> Water_Temp.:105.46, PWM:45
00:39:20.876 -> Water_Temp.:105.46, PWM:45
00:39:21.671 -> Water_Temp.:105.46, PWM:45
00:39:22.432 -> Water_Temp.:105.57, PWM:45
00:39:23.956 -> Water_Temp.:105.57, PWM:45
00:39:24.756 -> Water_Temp.:105.57, PWM:45
00:39:25.516 -> Water_Temp.:105.57, PWM:45
00:39:26.313 -> Water_Temp.:105.57, PWM:45
00:39:27.108 -> Water_Temp.:105.57, PWM:45
00:39:27.870 -> Water_Temp.:105.57, PWM:45
00:39:28.666 -> Water_Temp.:105.57, PWM:45
00:39:29.427 -> Water_Temp.:105.57, PWM:45
00:39:32.475 -> Water_Temp.:105.57, PWM:45
00:39:33.272 -> Water_Temp.:105.57, PWM:45
00:39:34.035 -> Water_Temp.:105.57, PWM:45
00:39:34.832 -> Water_Temp.:105.57, PWM:44
00:39:35.591 -> Water_Temp.:105.57, PWM:45
00:39:37.905 -> Water_Temp.:105.57, PWM:45
00:39:38.701 -> Water_Temp.:105.57, PWM:45
00:39:39.466 -> Water_Temp.:105.57, PWM:45
00:39:40.265 -> Water_Temp.:105.57, PWM:45
00:39:41.060 -> Water_Temp.:105.46, PWM:45
00:39:41.820 -> Water_Temp.:105.46, PWM:45
00:39:42.619 -> Water_Temp.:105.46, PWM:45
00:39:44.176 -> Water_Temp.:105.46, PWM:44
00:39:44.937 -> Water_Temp.:105.46, PWM:45
00:39:45.734 -> Water_Temp.:105.46, PWM:45
00:39:46.494 -> Water_Temp.:105.57, PWM:45
00:39:47.291 -> Water_Temp.:105.57, PWM:45
00:39:48.819 -> Water_Temp.:105.57, PWM:45
00:39:49.613 -> Water_Temp.:105.57, PWM:45
00:39:50.374 -> Water_Temp.:105.46, PWM:45
00:39:51.169 -> Water_Temp.:105.46, PWM:45
00:39:51.932 -> Water_Temp.:105.46, PWM:45
00:39:52.731 -> Water_Temp.:105.46, PWM:45
00:39:55.781 -> Water_Temp.:105.57, PWM:45
00:39:56.578 -> Water_Temp.:105.57, PWM:45
00:39:57.338 -> Water_Temp.:105.57, PWM:45
00:39:58.136 -> Water_Temp.:105.57, PWM:44
00:39:58.898 -> Water_Temp.:105.57, PWM:45
00:39:59.695 -> Water_Temp.:105.57, PWM:45
00:40:00.456 -> Water_Temp.:105.57, PWM:45
00:40:02.007 -> Water_Temp.:105.57, PWM:45
00:40:02.770 -> Water_Temp.:105.57, PWM:45
00:40:03.569 -> Water_Temp.:105.57, PWM:45
00:40:05.123 -> Water_Temp.:105.57, PWM:45
00:40:06.683 -> Water_Temp.:105.46, PWM:45
00:40:08.243 -> Water_Temp.:105.46, PWM:45
00:40:09.005 -> Water_Temp.:105.46, PWM:45
00:40:09.800 -> Water_Temp.:105.46, PWM:45
00:40:13.641 -> Water_Temp.:105.46, PWM:45
00:40:14.404 -> Water_Temp.:105.46, PWM:45
00:40:15.203 -> Water_Temp.:105.57, PWM:45
00:40:15.967 -> Water_Temp.:105.57, PWM:45
00:40:17.526 -> Water_Temp.:105.57, PWM:45
00:40:19.848 -> Water_Temp.:105.57, PWM:45
00:40:23.724 -> Water_Temp.:105.57, PWM:45
00:40:24.489 -> Water_Temp.:105.57, PWM:45
00:40:26.042 -> Water_Temp.:105.57, PWM:45
00:40:26.805 -> Water_Temp.:105.57, PWM:45
00:40:27.600 -> Water_Temp.:105.57, PWM:45
00:40:28.361 -> Water_Temp.:105.57, PWM:45
00:40:29.160 -> Water_Temp.:105.57, PWM:45
00:40:30.682 -> Water_Temp.:105.57, PWM:45
00:40:33.004 -> Water_Temp.:105.46, PWM:45
00:40:33.763 -> Water_Temp.:105.57, PWM:45
00:40:34.562 -> Water_Temp.:105.57, PWM:45
00:40:37.642 -> Water_Temp.:105.57, PWM:45
00:40:38.437 -> Water_Temp.:105.57, PWM:45
00:40:39.200 -> Water_Temp.:105.57, PWM:45
00:40:39.998 -> Water_Temp.:105.57, PWM:45
00:40:40.758 -> Water_Temp.:105.57, PWM:45
00:40:41.552 -> Water_Temp.:105.57, PWM:45
00:40:43.111 -> Water_Temp.:105.69, PWM:45
00:40:43.900 -> Water_Temp.:105.69, PWM:45
00:40:45.416 -> Water_Temp.:105.69, PWM:45
00:40:47.735 -> Water_Temp.:105.69, PWM:45
00:40:48.534 -> Water_Temp.:105.69, PWM:44
00:40:50.089 -> Water_Temp.:105.69, PWM:45
00:40:50.851 -> Water_Temp.:105.69, PWM:45
00:40:52.411 -> Water_Temp.:105.69, PWM:45
00:40:54.729 -> Water_Temp.:105.69, PWM:45
00:40:55.488 -> Water_Temp.:105.69, PWM:45
00:40:56.282 -> Water_Temp.:105.69, PWM:45
00:40:57.076 -> Water_Temp.:105.69, PWM:45
00:40:58.598 -> Water_Temp.:105.69, PWM:45
00:40:59.359 -> Water_Temp.:105.69, PWM:45
00:41:00.156 -> Water_Temp.:105.69, PWM:45
00:41:00.917 -> Water_Temp.:105.69, PWM:45
00:41:01.712 -> Water_Temp.:105.69, PWM:45
00:41:02.473 -> Water_Temp.:105.69, PWM:45
00:41:03.264 -> Water_Temp.:105.69, PWM:45
00:41:04.061 -> Water_Temp.:105.69, PWM:45
00:41:04.822 -> Water_Temp.:105.69, PWM:45
00:41:07.967 -> Water_Temp.:105.69, PWM:45
00:41:08.725 -> Water_Temp.:105.69, PWM:45
00:41:10.278 -> Water_Temp.:105.69, PWM:45
00:41:11.832 -> Water_Temp.:105.57, PWM:45
00:41:12.628 -> Water_Temp.:105.57, PWM:45
00:41:13.389 -> Water_Temp.:105.57, PWM:45
00:41:14.189 -> Water_Temp.:105.57, PWM:45
00:41:14.951 -> Water_Temp.:105.57, PWM:45
00:41:16.511 -> Water_Temp.:105.57, PWM:45
00:41:17.309 -> Water_Temp.:105.57, PWM:45
00:41:19.631 -> Water_Temp.:105.57, PWM:45
00:41:21.192 -> Water_Temp.:105.69, PWM:45
00:41:21.955 -> Water_Temp.:105.69, PWM:45
00:41:22.749 -> Water_Temp.:107.26, PWM:45
00:41:23.510 -> Water_Temp.:109.40, PWM:45
00:41:24.307 -> Water_Temp.:107.04, PWM:45
00:41:25.069 -> Water_Temp.:106.59, PWM:45
00:41:29.685 -> Water_Temp.:105.91, PWM:45
00:41:32.002 -> Water_Temp.:105.91, PWM:45
00:41:32.800 -> Water_Temp.:105.91, PWM:45
00:41:33.562 -> Water_Temp.:105.91, PWM:45
00:41:34.360 -> Water_Temp.:105.91, PWM:45
00:41:36.683 -> Water_Temp.:105.80, PWM:45
00:41:38.240 -> Water_Temp.:105.80, PWM:45
00:41:39.008 -> Water_Temp.:105.80, PWM:45
00:41:41.366 -> Water_Temp.:105.80, PWM:45
00:41:42.923 -> Water_Temp.:109.62, PWM:45
00:41:43.686 -> Water_Temp.:107.15, PWM:45
00:41:44.485 -> Water_Temp.:106.70, PWM:45
00:41:46.811 -> Water_Temp.:106.25, PWM:45
00:41:47.609 -> Water_Temp.:106.25, PWM:45
00:41:48.372 -> Water_Temp.:106.14, PWM:45
00:41:49.172 -> Water_Temp.:106.14, PWM:45
00:41:49.930 -> Water_Temp.:106.02, PWM:45
00:41:51.458 -> Water_Temp.:106.02, PWM:45
00:41:52.253 -> Water_Temp.:106.02, PWM:45
00:41:53.015 -> Water_Temp.:106.02, PWM:45
00:41:53.814 -> Water_Temp.:106.02, PWM:45
00:41:56.136 -> Water_Temp.:105.91, PWM:45
00:41:56.896 -> Water_Temp.:105.91, PWM:45
00:41:57.691 -> Water_Temp.:105.91, PWM:44
00:41:59.212 -> Water_Temp.:105.80, PWM:45
00:42:00.005 -> Water_Temp.:105.80, PWM:44
00:42:00.766 -> Water_Temp.:107.49, PWM:45
00:42:01.562 -> Water_Temp.:109.62, PWM:45
00:42:03.122 -> Water_Temp.:106.70, PWM:45
00:42:04.676 -> Water_Temp.:106.36, PWM:45
00:42:05.441 -> Water_Temp.:106.25, PWM:45
00:42:07.003 -> Water_Temp.:106.14, PWM:45
00:42:10.096 -> Water_Temp.:105.91, PWM:45
00:42:10.888 -> Water_Temp.:105.91, PWM:45
00:42:16.273 -> Water_Temp.:105.91, PWM:45
00:42:17.070 -> Water_Temp.:105.91, PWM:45
00:42:18.599 -> Water_Temp.:105.91, PWM:44
00:42:19.399 -> Water_Temp.:105.91, PWM:45
00:42:20.157 -> Water_Temp.:105.91, PWM:45
00:42:20.953 -> Water_Temp.:105.91, PWM:45
00:42:21.712 -> Water_Temp.:105.91, PWM:45
00:42:23.267 -> Water_Temp.:105.91, PWM:45
00:42:24.062 -> Water_Temp.:106.02, PWM:45
00:42:24.825 -> Water_Temp.:106.02, PWM:45
00:42:25.617 -> Water_Temp.:106.02, PWM:45
00:42:26.377 -> Water_Temp.:106.02, PWM:45
00:42:27.175 -> Water_Temp.:105.91, PWM:45
00:42:27.938 -> Water_Temp.:106.02, PWM:45
00:42:29.493 -> Water_Temp.:106.02, PWM:45
00:42:30.290 -> Water_Temp.:106.02, PWM:45
00:42:31.048 -> Water_Temp.:106.02, PWM:45
00:42:33.359 -> Water_Temp.:106.14, PWM:45
00:42:35.668 -> Water_Temp.:106.02, PWM:45
00:42:36.461 -> Water_Temp.:106.02, PWM:45
00:42:37.254 -> Water_Temp.:106.02, PWM:45
00:42:38.808 -> Water_Temp.:106.02, PWM:45
00:42:39.569 -> Water_Temp.:106.02, PWM:45
00:42:40.365 -> Water_Temp.:106.02, PWM:45
00:42:41.128 -> Water_Temp.:106.02, PWM:45
00:42:41.921 -> Water_Temp.:106.02, PWM:45
00:42:44.236 -> Water_Temp.:106.02, PWM:45
00:42:45.789 -> Water_Temp.:106.02, PWM:45
00:42:46.587 -> Water_Temp.:106.02, PWM:45
00:42:50.465 -> Water_Temp.:106.02, PWM:45
00:42:51.231 -> Water_Temp.:106.02, PWM:45
00:42:52.028 -> Water_Temp.:106.02, PWM:45
00:42:52.790 -> Water_Temp.:106.02, PWM:45
00:42:53.587 -> Water_Temp.:106.02, PWM:45
00:42:54.348 -> Water_Temp.:106.02, PWM:45
00:42:55.871 -> Water_Temp.:106.14, PWM:45
00:42:56.665 -> Water_Temp.:106.14, PWM:45
00:42:58.220 -> Water_Temp.:106.14, PWM:45
00:42:59.014 -> Water_Temp.:106.14, PWM:45
00:42:59.777 -> Water_Temp.:106.14, PWM:45
00:43:00.573 -> Water_Temp.:106.14, PWM:45
00:43:01.335 -> Water_Temp.:106.14, PWM:45
00:43:02.892 -> Water_Temp.:106.14, PWM:45
00:43:03.654 -> Water_Temp.:106.14, PWM:45
00:43:04.454 -> Water_Temp.:106.14, PWM:45
00:43:05.214 -> Water_Temp.:106.14, PWM:45
00:43:09.786 -> Water_Temp.:106.14, PWM:45
00:43:11.344 -> Water_Temp.:106.14, PWM:45
00:43:12.139 -> Water_Temp.:106.14, PWM:45
00:43:13.693 -> Water_Temp.:106.25, PWM:45
00:43:14.490 -> Water_Temp.:106.25, PWM:45
00:43:15.248 -> Water_Temp.:106.25, PWM:45
00:43:16.043 -> Water_Temp.:106.25, PWM:45
00:43:17.593 -> Water_Temp.:106.25, PWM:45
00:43:18.386 -> Water_Temp.:106.25, PWM:45
00:43:19.910 -> Water_Temp.:106.25, PWM:45
00:43:20.707 -> Water_Temp.:106.25, PWM:45
00:43:21.468 -> Water_Temp.:106.25, PWM:45
00:43:22.991 -> Water_Temp.:106.25, PWM:45
00:43:23.785 -> Water_Temp.:106.25, PWM:45
00:43:24.551 -> Water_Temp.:106.25, PWM:44
00:43:25.347 -> Water_Temp.:106.25, PWM:45
00:43:29.209 -> Water_Temp.:106.25, PWM:45
00:43:29.971 -> Water_Temp.:106.25, PWM:45
00:43:30.737 -> Water_Temp.:106.25, PWM:45
00:43:31.536 -> Water_Temp.:106.25, PWM:45
00:43:33.099 -> Water_Temp.:106.25, PWM:45
00:43:34.631 -> Water_Temp.:106.36, PWM:45
00:43:36.194 -> Water_Temp.:106.47, PWM:45
00:43:36.954 -> Water_Temp.:106.47, PWM:45
00:43:37.752 -> Water_Temp.:106.47, PWM:44
00:43:40.078 -> Water_Temp.:106.36, PWM:45
00:43:43.890 -> Water_Temp.:106.36, PWM:45
00:43:44.686 -> Water_Temp.:106.25, PWM:45
00:43:45.449 -> Water_Temp.:106.25, PWM:45
00:43:48.563 -> Water_Temp.:106.25, PWM:45
00:43:49.324 -> Water_Temp.:106.25, PWM:45
00:43:50.118 -> Water_Temp.:106.25, PWM:45
00:43:50.882 -> Water_Temp.:106.25, PWM:45
00:43:51.678 -> Water_Temp.:106.25, PWM:45
00:43:52.439 -> Water_Temp.:106.25, PWM:45
00:43:53.234 -> Water_Temp.:106.25, PWM:45
00:43:53.998 -> Water_Temp.:106.25, PWM:45
00:43:54.795 -> Water_Temp.:106.25, PWM:45
00:43:56.350 -> Water_Temp.:106.25, PWM:45
00:43:57.874 -> Water_Temp.:106.25, PWM:45
00:43:58.669 -> Water_Temp.:106.25, PWM:36
00:44:00.229 -> Water_Temp.:106.25, PWM:35
00:44:00.992 -> Water_Temp.:106.25, PWM:35
00:44:01.789 -> Water_Temp.:106.25, PWM:35
00:44:02.549 -> Water_Temp.:106.25, PWM:34
00:44:03.346 -> Water_Temp.:106.25, PWM:34
00:44:04.109 -> Water_Temp.:106.25, PWM:34
00:44:04.907 -> Water_Temp.:106.25, PWM:0
00:44:05.671 -> Water_Temp.:106.25, PWM:0
00:44:06.470 -> Water_Temp.:106.25, PWM:0
00:44:07.231 -> Water_Temp.:106.25, PWM:0
00:44:08.026 -> Water_Temp.:106.25, PWM:0
00:44:08.790 -> Water_Temp.:106.25, PWM:0
00:44:09.587 -> Water_Temp.:106.25, PWM:0
00:44:10.352 -> Water_Temp.:106.25, PWM:0
00:44:11.148 -> Water_Temp.:106.25, PWM:0
00:44:11.907 -> Water_Temp.:106.25, PWM:0
00:44:12.698 -> Water_Temp.:106.25, PWM:0
00:44:13.494 -> Water_Temp.:106.25, PWM:0
00:44:14.257 -> Water_Temp.:106.25, PWM:0
00:44:15.052 -> Water_Temp.:106.25, PWM:0
00:44:15.810 -> Water_Temp.:106.25, PWM:0
00:44:16.602 -> Water_Temp.:106.25, PWM:0
00:44:17.363 -> Water_Temp.:106.25, PWM:0
00:44:18.158 -> Water_Temp.:106.25, PWM:0
00:44:18.956 -> Water_Temp.:106.25, PWM:0
00:44:19.715 -> Water_Temp.:106.25, PWM:0
00:44:20.507 -> Water_Temp.:106.25, PWM:0
00:44:21.270 -> Water_Temp.:106.25, PWM:0
00:44:22.063 -> Water_Temp.:106.25, PWM:0
00:44:22.824 -> Water_Temp.:106.25, PWM:0
00:44:23.617 -> Water_Temp.:106.25, PWM:0
00:44:24.411 -> Water_Temp.:106.25, PWM:0
00:44:25.173 -> Water_Temp.:106.25, PWM:0
00:44:25.968 -> Water_Temp.:106.25, PWM:0
00:44:26.730 -> Water_Temp.:106.25, PWM:0
00:44:27.524 -> Water_Temp.:106.25, PWM:0
00:44:28.284 -> Water_Temp.:106.25, PWM:0
00:44:29.078 -> Water_Temp.:106.25, PWM:0
00:44:29.872 -> Water_Temp.:106.14, PWM:0
00:44:30.634 -> Water_Temp.:106.25, PWM:0
00:44:31.433 -> Water_Temp.:106.25, PWM:0
00:44:32.193 -> Water_Temp.:106.25, PWM:0
00:44:32.992 -> Water_Temp.:106.25, PWM:0
00:44:33.756 -> Water_Temp.:106.25, PWM:0
00:44:34.554 -> Water_Temp.:106.25, PWM:0
00:44:35.320 -> Water_Temp.:106.25, PWM:0
00:44:36.084 -> Water_Temp.:106.25, PWM:0
00:44:36.882 -> Water_Temp.:106.25, PWM:0
00:44:37.675 -> Water_Temp.:106.25, PWM:0
00:44:38.434 -> Water_Temp.:106.25, PWM:0
00:44:39.232 -> Water_Temp.:106.25, PWM:0
00:44:39.994 -> Water_Temp.:106.25, PWM:0
00:44:40.792 -> Water_Temp.:106.25, PWM:0
00:44:41.554 -> Water_Temp.:106.25, PWM:0
00:44:42.352 -> Water_Temp.:106.14, PWM:0
00:44:43.115 -> Water_Temp.:106.14, PWM:0
00:44:43.907 -> Water_Temp.:106.14, PWM:0
00:44:44.669 -> Water_Temp.:106.14, PWM:0
00:44:45.467 -> Water_Temp.:106.14, PWM:0
00:44:46.223 -> Water_Temp.:106.02, PWM:0
00:44:47.017 -> Water_Temp.:106.02, PWM:0
00:44:47.812 -> Water_Temp.:106.02, PWM:0
00:44:48.573 -> Water_Temp.:106.02, PWM:0
00:44:49.372 -> Water_Temp.:106.02, PWM:0
00:44:50.127 -> Water_Temp.:106.02, PWM:0
00:44:50.922 -> Water_Temp.:106.02, PWM:0
00:44:51.686 -> Water_Temp.:106.02, PWM:0
00:44:52.482 -> Water_Temp.:106.02, PWM:0
00:44:53.243 -> Water_Temp.:106.02, PWM:0
00:44:54.043 -> Water_Temp.:106.02, PWM:0
00:44:54.806 -> Water_Temp.:106.02, PWM:0
00:44:55.602 -> Water_Temp.:105.91, PWM:0
00:44:56.363 -> Water_Temp.:105.91, PWM:0
00:44:57.160 -> Water_Temp.:105.91, PWM:0
00:44:57.923 -> Water_Temp.:105.91, PWM:0
00:44:58.721 -> Water_Temp.:105.91, PWM:0
00:44:59.484 -> Water_Temp.:105.91, PWM:0
00:45:00.282 -> Water_Temp.:105.91, PWM:0
00:45:01.043 -> Water_Temp.:105.91, PWM:0
00:45:01.834 -> Water_Temp.:105.80, PWM:0
00:45:02.598 -> Water_Temp.:105.80, PWM:0
00:45:03.397 -> Water_Temp.:105.80, PWM:0
00:45:04.161 -> Water_Temp.:105.80, PWM:0
00:45:04.958 -> Water_Temp.:105.80, PWM:0
00:45:05.722 -> Water_Temp.:105.80, PWM:0
00:45:06.517 -> Water_Temp.:105.80, PWM:0
00:45:07.311 -> Water_Temp.:105.80, PWM:0
00:45:08.071 -> Water_Temp.:105.69, PWM:0
00:45:08.837 -> Water_Temp.:105.69, PWM:0
00:45:09.634 -> Water_Temp.:105.69, PWM:0
00:45:10.396 -> Water_Temp.:105.69, PWM:0
00:45:11.190 -> Water_Temp.:105.69, PWM:0
00:45:11.983 -> Water_Temp.:105.69, PWM:0
00:45:12.743 -> Water_Temp.:105.69, PWM:0
00:45:13.542 -> Water_Temp.:105.57, PWM:0
00:45:14.308 -> Water_Temp.:105.57, PWM:0
00:45:15.105 -> Water_Temp.:105.57, PWM:0
00:45:15.868 -> Water_Temp.:105.57, PWM:0
00:45:16.661 -> Water_Temp.:105.57, PWM:0
00:45:17.419 -> Water_Temp.:105.57, PWM:0
00:45:18.214 -> Water_Temp.:105.57, PWM:0
00:45:18.977 -> Water_Temp.:105.57, PWM:0
00:45:19.772 -> Water_Temp.:105.57, PWM:0
00:45:20.568 -> Water_Temp.:105.57, PWM:0
00:45:21.334 -> Water_Temp.:105.46, PWM:0
00:45:22.096 -> Water_Temp.:105.46, PWM:0
00:45:22.891 -> Water_Temp.:105.46, PWM:0
00:45:23.655 -> Water_Temp.:105.46, PWM:0
00:45:24.456 -> Water_Temp.:105.46, PWM:0
00:45:25.220 -> Water_Temp.:105.46, PWM:0
00:45:26.015 -> Water_Temp.:105.46, PWM:0
00:45:26.775 -> Water_Temp.:105.46, PWM:0
00:45:27.571 -> Water_Temp.:105.46, PWM:0
00:45:28.364 -> Water_Temp.:105.46, PWM:0
00:45:29.124 -> Water_Temp.:105.46, PWM:0
00:45:29.912 -> Water_Temp.:105.46, PWM:0
00:45:30.701 -> Water_Temp.:105.35, PWM:0

## RAW_CAL_MLX-3
22:40:00.354 -> Average Cal sensor Temp.: 81.67 Ambient Temp.: 80.44 F
22:40:01.909 -> Average Cal sensor Temp.: 81.82 Ambient Temp.: 80.37 F
22:40:03.469 -> Average Cal sensor Temp.: 81.78 Ambient Temp.: 80.33 F
22:40:05.023 -> Average Cal sensor Temp.: 81.92 Ambient Temp.: 80.40 F
22:40:06.583 -> Average Cal sensor Temp.: 81.96 Ambient Temp.: 80.37 F
22:40:08.135 -> Average Cal sensor Temp.: 81.75 Ambient Temp.: 80.37 F
22:40:09.695 -> Average Cal sensor Temp.: 81.83 Ambient Temp.: 80.40 F
22:40:11.251 -> Average Cal sensor Temp.: 81.68 Ambient Temp.: 80.33 F
22:40:12.807 -> Average Cal sensor Temp.: 81.69 Ambient Temp.: 80.29 F
22:40:14.362 -> Average Cal sensor Temp.: 81.86 Ambient Temp.: 80.33 F
22:40:15.922 -> Average Cal sensor Temp.: 81.82 Ambient Temp.: 80.37 F
22:40:17.446 -> Average Cal sensor Temp.: 81.81 Ambient Temp.: 80.29 F
22:40:19.005 -> Average Cal sensor Temp.: 81.84 Ambient Temp.: 80.33 F
22:40:20.560 -> Average Cal sensor Temp.: 81.54 Ambient Temp.: 80.33 F
22:40:22.120 -> Average Cal sensor Temp.: 81.46 Ambient Temp.: 80.29 F
22:40:23.674 -> Average Cal sensor Temp.: 81.49 Ambient Temp.: 80.29 F
22:40:25.228 -> Average Cal sensor Temp.: 81.50 Ambient Temp.: 80.29 F
22:40:26.810 -> Average Cal sensor Temp.: 81.71 Ambient Temp.: 80.26 F
22:40:28.359 -> Average Cal sensor Temp.: 81.68 Ambient Temp.: 80.26 F
22:40:29.917 -> Average Cal sensor Temp.: 81.47 Ambient Temp.: 80.26 F
22:40:31.474 -> Average Cal sensor Temp.: 81.73 Ambient Temp.: 80.22 F
22:40:33.023 -> Average Cal sensor Temp.: 81.75 Ambient Temp.: 80.22 F
22:40:34.583 -> Average Cal sensor Temp.: 81.74 Ambient Temp.: 80.22 F
22:40:36.135 -> Average Cal sensor Temp.: 81.76 Ambient Temp.: 80.19 F
22:40:37.698 -> Average Cal sensor Temp.: 81.88 Ambient Temp.: 80.22 F
22:40:39.258 -> Average Cal sensor Temp.: 81.83 Ambient Temp.: 80.22 F
22:40:40.784 -> Average Cal sensor Temp.: 81.46 Ambient Temp.: 80.22 F
22:40:42.346 -> Average Cal sensor Temp.: 81.64 Ambient Temp.: 80.22 F
22:40:43.906 -> Average Cal sensor Temp.: 81.54 Ambient Temp.: 80.19 F
22:40:45.467 -> Average Cal sensor Temp.: 81.46 Ambient Temp.: 80.11 F
22:40:47.022 -> Average Cal sensor Temp.: 81.51 Ambient Temp.: 80.11 F
22:40:48.582 -> Average Cal sensor Temp.: 81.69 Ambient Temp.: 80.22 F
22:40:50.143 -> Average Cal sensor Temp.: 81.55 Ambient Temp.: 80.19 F
22:40:51.669 -> Average Cal sensor Temp.: 81.39 Ambient Temp.: 80.22 F
22:40:53.225 -> Average Cal sensor Temp.: 81.44 Ambient Temp.: 80.19 F
22:40:54.788 -> Average Cal sensor Temp.: 81.31 Ambient Temp.: 80.11 F
22:40:56.342 -> Average Cal sensor Temp.: 81.45 Ambient Temp.: 80.19 F
22:40:57.893 -> Average Cal sensor Temp.: 81.37 Ambient Temp.: 80.19 F
22:40:59.454 -> Average Cal sensor Temp.: 81.27 Ambient Temp.: 80.19 F
22:41:01.019 -> Average Cal sensor Temp.: 81.41 Ambient Temp.: 80.22 F
22:41:02.575 -> Average Cal sensor Temp.: 81.46 Ambient Temp.: 80.19 F
22:41:04.134 -> Average Cal sensor Temp.: 81.45 Ambient Temp.: 80.19 F
22:41:05.689 -> Average Cal sensor Temp.: 81.42 Ambient Temp.: 80.11 F
22:41:07.240 -> Average Cal sensor Temp.: 81.63 Ambient Temp.: 80.19 F
22:41:08.801 -> Average Cal sensor Temp.: 81.37 Ambient Temp.: 80.11 F
22:41:10.359 -> Average Cal sensor Temp.: 81.33 Ambient Temp.: 80.19 F
22:41:11.917 -> Average Cal sensor Temp.: 81.28 Ambient Temp.: 80.11 F
22:41:13.443 -> Average Cal sensor Temp.: 81.27 Ambient Temp.: 80.08 F
22:41:15.002 -> Average Cal sensor Temp.: 81.34 Ambient Temp.: 80.04 F
22:41:16.564 -> Average Cal sensor Temp.: 81.42 Ambient Temp.: 80.04 F
22:41:18.119 -> Average Cal sensor Temp.: 81.58 Ambient Temp.: 80.11 F
22:41:19.678 -> Average Cal sensor Temp.: 81.48 Ambient Temp.: 80.08 F
22:41:21.244 -> Average Cal sensor Temp.: 81.41 Ambient Temp.: 80.11 F
22:41:22.803 -> Average Cal sensor Temp.: 81.37 Ambient Temp.: 80.04 F
22:41:24.361 -> Average Cal sensor Temp.: 81.30 Ambient Temp.: 79.97 F
22:41:25.919 -> Average Cal sensor Temp.: 81.46 Ambient Temp.: 80.01 F
22:41:27.443 -> Average Cal sensor Temp.: 81.20 Ambient Temp.: 80.08 F
22:41:29.005 -> Average Cal sensor Temp.: 81.18 Ambient Temp.: 80.01 F
22:41:30.554 -> Average Cal sensor Temp.: 81.27 Ambient Temp.: 80.04 F
22:41:32.112 -> Average Cal sensor Temp.: 81.29 Ambient Temp.: 80.01 F
22:41:33.669 -> Average Cal sensor Temp.: 81.74 Ambient Temp.: 80.04 F
22:41:35.229 -> Average Cal sensor Temp.: 82.26 Ambient Temp.: 79.97 F
22:41:36.790 -> Average Cal sensor Temp.: 82.52 Ambient Temp.: 79.97 F
22:41:38.348 -> Average Cal sensor Temp.: 82.75 Ambient Temp.: 80.01 F
22:41:39.913 -> Average Cal sensor Temp.: 82.84 Ambient Temp.: 80.04 F
22:41:41.478 -> Average Cal sensor Temp.: 82.96 Ambient Temp.: 79.97 F
22:41:43.005 -> Average Cal sensor Temp.: 83.07 Ambient Temp.: 79.90 F
22:41:44.570 -> Average Cal sensor Temp.: 83.49 Ambient Temp.: 79.93 F
22:41:46.129 -> Average Cal sensor Temp.: 83.42 Ambient Temp.: 79.93 F
22:41:47.683 -> Average Cal sensor Temp.: 83.44 Ambient Temp.: 79.93 F
22:41:49.242 -> Average Cal sensor Temp.: 83.49 Ambient Temp.: 79.93 F
22:41:50.805 -> Average Cal sensor Temp.: 83.44 Ambient Temp.: 79.90 F
22:41:52.359 -> Average Cal sensor Temp.: 83.18 Ambient Temp.: 79.93 F
22:41:53.921 -> Average Cal sensor Temp.: 83.31 Ambient Temp.: 79.93 F
22:41:55.479 -> Average Cal sensor Temp.: 83.25 Ambient Temp.: 79.90 F
22:41:57.029 -> Average Cal sensor Temp.: 83.32 Ambient Temp.: 79.93 F
22:41:58.590 -> Average Cal sensor Temp.: 83.50 Ambient Temp.: 79.90 F
22:42:00.150 -> Average Cal sensor Temp.: 83.41 Ambient Temp.: 79.93 F
22:42:01.673 -> Average Cal sensor Temp.: 83.64 Ambient Temp.: 79.90 F
22:42:03.229 -> Average Cal sensor Temp.: 83.72 Ambient Temp.: 79.86 F
22:42:04.786 -> Average Cal sensor Temp.: 83.93 Ambient Temp.: 79.86 F
22:42:06.348 -> Average Cal sensor Temp.: 84.08 Ambient Temp.: 79.83 F
22:42:07.904 -> Average Cal sensor Temp.: 84.20 Ambient Temp.: 79.79 F
22:42:09.464 -> Average Cal sensor Temp.: 84.43 Ambient Temp.: 79.83 F
22:42:11.022 -> Average Cal sensor Temp.: 84.40 Ambient Temp.: 79.72 F
22:42:12.579 -> Average Cal sensor Temp.: 84.66 Ambient Temp.: 79.68 F
22:42:14.144 -> Average Cal sensor Temp.: 84.68 Ambient Temp.: 79.72 F
22:42:15.695 -> Average Cal sensor Temp.: 84.60 Ambient Temp.: 79.68 F
22:42:17.248 -> Average Cal sensor Temp.: 84.64 Ambient Temp.: 79.65 F
22:42:18.806 -> Average Cal sensor Temp.: 84.92 Ambient Temp.: 79.61 F
22:42:20.361 -> Average Cal sensor Temp.: 85.00 Ambient Temp.: 79.65 F
22:42:21.925 -> Average Cal sensor Temp.: 85.03 Ambient Temp.: 79.65 F
22:42:23.483 -> Average Cal sensor Temp.: 84.89 Ambient Temp.: 79.57 F
22:42:25.040 -> Average Cal sensor Temp.: 84.91 Ambient Temp.: 79.57 F
22:42:26.592 -> Average Cal sensor Temp.: 84.89 Ambient Temp.: 79.54 F
22:42:28.156 -> Average Cal sensor Temp.: 84.90 Ambient Temp.: 79.50 F
22:42:29.715 -> Average Cal sensor Temp.: 85.12 Ambient Temp.: 79.47 F
22:42:31.268 -> Average Cal sensor Temp.: 85.33 Ambient Temp.: 79.47 F
22:42:32.825 -> Average Cal sensor Temp.: 85.52 Ambient Temp.: 79.39 F
22:42:34.382 -> Average Cal sensor Temp.: 85.67 Ambient Temp.: 79.43 F
22:42:35.939 -> Average Cal sensor Temp.: 85.47 Ambient Temp.: 79.39 F
22:42:37.495 -> Average Cal sensor Temp.: 85.55 Ambient Temp.: 79.39 F
22:42:39.030 -> Average Cal sensor Temp.: 85.61 Ambient Temp.: 79.39 F
22:42:40.589 -> Average Cal sensor Temp.: 85.86 Ambient Temp.: 79.29 F
22:42:42.150 -> Average Cal sensor Temp.: 86.09 Ambient Temp.: 79.29 F
22:42:43.706 -> Average Cal sensor Temp.: 86.05 Ambient Temp.: 79.25 F
22:42:45.268 -> Average Cal sensor Temp.: 86.23 Ambient Temp.: 79.29 F
22:42:46.826 -> Average Cal sensor Temp.: 86.19 Ambient Temp.: 79.21 F
22:42:48.391 -> Average Cal sensor Temp.: 86.28 Ambient Temp.: 79.21 F
22:42:49.919 -> Average Cal sensor Temp.: 86.64 Ambient Temp.: 79.21 F
22:42:51.505 -> Average Cal sensor Temp.: 86.82 Ambient Temp.: 79.18 F
22:42:53.061 -> Average Cal sensor Temp.: 86.80 Ambient Temp.: 79.25 F
22:42:54.587 -> Average Cal sensor Temp.: 86.73 Ambient Temp.: 79.25 F
22:42:56.175 -> Average Cal sensor Temp.: 87.22 Ambient Temp.: 79.18 F
22:42:57.703 -> Average Cal sensor Temp.: 87.60 Ambient Temp.: 79.21 F
22:42:59.266 -> Average Cal sensor Temp.: 87.50 Ambient Temp.: 79.18 F
22:43:00.826 -> Average Cal sensor Temp.: 87.77 Ambient Temp.: 79.18 F
22:43:02.388 -> Average Cal sensor Temp.: 87.97 Ambient Temp.: 79.18 F
22:43:03.949 -> Average Cal sensor Temp.: 88.01 Ambient Temp.: 79.18 F
22:43:05.475 -> Average Cal sensor Temp.: 88.11 Ambient Temp.: 79.25 F
22:43:07.033 -> Average Cal sensor Temp.: 87.90 Ambient Temp.: 79.18 F
22:43:08.598 -> Average Cal sensor Temp.: 88.07 Ambient Temp.: 79.25 F
22:43:10.162 -> Average Cal sensor Temp.: 87.92 Ambient Temp.: 79.18 F
22:43:11.723 -> Average Cal sensor Temp.: 87.98 Ambient Temp.: 79.25 F
22:43:13.283 -> Average Cal sensor Temp.: 88.14 Ambient Temp.: 79.21 F
22:43:14.808 -> Average Cal sensor Temp.: 88.33 Ambient Temp.: 79.21 F
22:43:16.370 -> Average Cal sensor Temp.: 88.41 Ambient Temp.: 79.25 F
22:43:17.931 -> Average Cal sensor Temp.: 88.65 Ambient Temp.: 79.25 F
22:43:19.492 -> Average Cal sensor Temp.: 88.89 Ambient Temp.: 79.29 F
22:43:21.048 -> Average Cal sensor Temp.: 88.83 Ambient Temp.: 79.32 F
22:43:22.611 -> Average Cal sensor Temp.: 88.76 Ambient Temp.: 79.29 F
22:43:24.174 -> Average Cal sensor Temp.: 88.97 Ambient Temp.: 79.29 F
22:43:25.733 -> Average Cal sensor Temp.: 88.93 Ambient Temp.: 79.32 F
22:43:27.292 -> Average Cal sensor Temp.: 88.68 Ambient Temp.: 79.39 F
22:43:28.849 -> Average Cal sensor Temp.: 88.79 Ambient Temp.: 79.43 F
22:43:30.405 -> Average Cal sensor Temp.: 88.88 Ambient Temp.: 79.39 F
22:43:31.935 -> Average Cal sensor Temp.: 88.87 Ambient Temp.: 79.43 F
22:43:33.490 -> Average Cal sensor Temp.: 88.98 Ambient Temp.: 79.50 F
22:43:35.050 -> Average Cal sensor Temp.: 88.88 Ambient Temp.: 79.50 F
22:43:36.610 -> Average Cal sensor Temp.: 88.97 Ambient Temp.: 79.57 F
22:43:38.167 -> Average Cal sensor Temp.: 88.96 Ambient Temp.: 79.61 F
22:43:39.729 -> Average Cal sensor Temp.: 89.05 Ambient Temp.: 79.65 F
22:43:41.284 -> Average Cal sensor Temp.: 89.10 Ambient Temp.: 79.65 F
22:43:42.842 -> Average Cal sensor Temp.: 89.00 Ambient Temp.: 79.68 F
22:43:44.402 -> Average Cal sensor Temp.: 88.75 Ambient Temp.: 79.72 F
22:43:45.963 -> Average Cal sensor Temp.: 88.93 Ambient Temp.: 79.83 F
22:43:47.526 -> Average Cal sensor Temp.: 88.99 Ambient Temp.: 79.79 F
22:43:49.093 -> Average Cal sensor Temp.: 89.00 Ambient Temp.: 79.86 F
22:43:50.649 -> Average Cal sensor Temp.: 89.05 Ambient Temp.: 79.86 F
22:43:52.205 -> Average Cal sensor Temp.: 88.84 Ambient Temp.: 79.86 F
22:43:53.767 -> Average Cal sensor Temp.: 89.03 Ambient Temp.: 79.93 F
22:43:55.324 -> Average Cal sensor Temp.: 89.03 Ambient Temp.: 79.90 F
22:43:56.848 -> Average Cal sensor Temp.: 89.33 Ambient Temp.: 80.04 F
22:43:58.438 -> Average Cal sensor Temp.: 89.66 Ambient Temp.: 80.01 F
22:43:59.995 -> Average Cal sensor Temp.: 89.41 Ambient Temp.: 80.01 F
22:44:01.521 -> Average Cal sensor Temp.: 89.46 Ambient Temp.: 80.11 F
22:44:03.085 -> Average Cal sensor Temp.: 89.32 Ambient Temp.: 80.08 F
22:44:04.645 -> Average Cal sensor Temp.: 89.32 Ambient Temp.: 80.11 F
22:44:06.199 -> Average Cal sensor Temp.: 89.21 Ambient Temp.: 80.22 F
22:44:06.929 -> Ambient Temp.: 80.29 F
22:44:07.484 -> Ambient Temp.: 80.29 F
22:44:09.041 -> Average Cal sensor Temp.: 89.15 Ambient Temp.: 80.26 F
22:44:10.596 -> Average Cal sensor Temp.: 89.24 Ambient Temp.: 80.37 F
22:44:12.155 -> Average Cal sensor Temp.: 89.19 Ambient Temp.: 80.33 F
22:44:13.716 -> Average Cal sensor Temp.: 89.20 Ambient Temp.: 80.37 F
22:44:15.276 -> Average Cal sensor Temp.: 89.19 Ambient Temp.: 80.44 F
22:44:16.837 -> Average Cal sensor Temp.: 89.23 Ambient Temp.: 80.47 F
22:44:18.395 -> Average Cal sensor Temp.: 89.23 Ambient Temp.: 80.51 F
22:44:19.951 -> Average Cal sensor Temp.: 89.24 Ambient Temp.: 80.51 F
22:44:21.511 -> Average Cal sensor Temp.: 89.20 Ambient Temp.: 80.65 F
22:44:23.068 -> Average Cal sensor Temp.: 89.29 Ambient Temp.: 80.69 F
22:44:24.597 -> Average Cal sensor Temp.: 89.33 Ambient Temp.: 80.65 F
22:44:26.155 -> Average Cal sensor Temp.: 89.28 Ambient Temp.: 80.80 F
22:44:27.714 -> Average Cal sensor Temp.: 89.23 Ambient Temp.: 80.80 F
22:44:29.274 -> Average Cal sensor Temp.: 89.22 Ambient Temp.: 80.80 F
22:44:30.836 -> Average Cal sensor Temp.: 89.24 Ambient Temp.: 80.87 F
22:44:32.395 -> Average Cal sensor Temp.: 89.32 Ambient Temp.: 80.94 F
22:44:33.953 -> Average Cal sensor Temp.: 89.27 Ambient Temp.: 80.98 F
22:44:35.507 -> Average Cal sensor Temp.: 89.41 Ambient Temp.: 80.98 F
22:44:37.066 -> Average Cal sensor Temp.: 89.40 Ambient Temp.: 81.05 F
22:44:38.624 -> Average Cal sensor Temp.: 89.37 Ambient Temp.: 81.01 F
22:44:40.182 -> Average Cal sensor Temp.: 89.37 Ambient Temp.: 81.05 F
22:44:41.742 -> Average Cal sensor Temp.: 89.30 Ambient Temp.: 81.09 F
22:44:43.303 -> Average Cal sensor Temp.: 89.31 Ambient Temp.: 81.12 F
22:44:44.864 -> Average Cal sensor Temp.: 89.33 Ambient Temp.: 81.12 F
22:44:46.414 -> Average Cal sensor Temp.: 89.41 Ambient Temp.: 81.16 F
22:44:47.975 -> Average Cal sensor Temp.: 89.49 Ambient Temp.: 81.19 F
22:44:49.527 -> Average Cal sensor Temp.: 89.42 Ambient Temp.: 81.19 F
22:44:51.081 -> Average Cal sensor Temp.: 89.75 Ambient Temp.: 81.19 F
22:44:52.640 -> Average Cal sensor Temp.: 89.72 Ambient Temp.: 81.30 F
22:44:54.193 -> Average Cal sensor Temp.: 89.82 Ambient Temp.: 81.37 F
22:44:55.731 -> Average Cal sensor Temp.: 89.85 Ambient Temp.: 81.37 F
22:44:57.291 -> Average Cal sensor Temp.: 89.84 Ambient Temp.: 81.37 F
22:44:58.856 -> Average Cal sensor Temp.: 89.84 Ambient Temp.: 81.37 F
22:45:00.412 -> Average Cal sensor Temp.: 89.75 Ambient Temp.: 81.45 F
22:45:01.968 -> Average Cal sensor Temp.: 89.64 Ambient Temp.: 81.48 F
22:45:03.527 -> Average Cal sensor Temp.: 89.59 Ambient Temp.: 81.48 F
22:45:05.085 -> Average Cal sensor Temp.: 89.54 Ambient Temp.: 81.55 F
22:45:06.644 -> Average Cal sensor Temp.: 89.41 Ambient Temp.: 81.59 F
22:45:08.201 -> Average Cal sensor Temp.: 89.28 Ambient Temp.: 81.66 F
22:45:09.765 -> Average Cal sensor Temp.: 89.28 Ambient Temp.: 81.59 F
22:45:11.293 -> Average Cal sensor Temp.: 89.47 Ambient Temp.: 81.66 F
22:45:12.857 -> Average Cal sensor Temp.: 89.41 Ambient Temp.: 81.70 F
22:45:14.417 -> Average Cal sensor Temp.: 89.65 Ambient Temp.: 81.77 F
22:45:15.976 -> Average Cal sensor Temp.: 89.78 Ambient Temp.: 81.77 F
22:45:17.533 -> Average Cal sensor Temp.: 89.56 Ambient Temp.: 81.73 F
22:45:19.092 -> Average Cal sensor Temp.: 89.61 Ambient Temp.: 81.84 F
22:45:20.653 -> Average Cal sensor Temp.: 89.41 Ambient Temp.: 81.81 F
22:45:22.211 -> Average Cal sensor Temp.: 89.71 Ambient Temp.: 81.88 F
22:45:23.742 -> Average Cal sensor Temp.: 89.77 Ambient Temp.: 81.91 F
22:45:25.304 -> Average Cal sensor Temp.: 89.70 Ambient Temp.: 81.91 F
22:45:26.867 -> Average Cal sensor Temp.: 89.64 Ambient Temp.: 81.95 F
22:45:28.425 -> Average Cal sensor Temp.: 89.74 Ambient Temp.: 81.99 F
22:45:29.985 -> Average Cal sensor Temp.: 89.76 Ambient Temp.: 81.99 F
22:45:31.537 -> Average Cal sensor Temp.: 89.77 Ambient Temp.: 82.06 F
22:45:33.097 -> Average Cal sensor Temp.: 89.70 Ambient Temp.: 82.06 F
22:45:34.657 -> Average Cal sensor Temp.: 89.72 Ambient Temp.: 82.20 F
22:45:36.217 -> Average Cal sensor Temp.: 89.72 Ambient Temp.: 82.17 F
22:45:37.773 -> Average Cal sensor Temp.: 89.60 Ambient Temp.: 82.17 F
22:45:39.302 -> Average Cal sensor Temp.: 89.65 Ambient Temp.: 82.20 F
22:45:40.862 -> Average Cal sensor Temp.: 89.96 Ambient Temp.: 82.27 F
22:45:42.417 -> Average Cal sensor Temp.: 89.70 Ambient Temp.: 82.31 F
22:45:43.977 -> Average Cal sensor Temp.: 89.64 Ambient Temp.: 82.35 F
22:45:45.539 -> Average Cal sensor Temp.: 89.78 Ambient Temp.: 82.38 F
22:45:47.100 -> Average Cal sensor Temp.: 89.68 Ambient Temp.: 82.38 F
22:45:48.659 -> Average Cal sensor Temp.: 89.66 Ambient Temp.: 82.38 F
22:45:50.218 -> Average Cal sensor Temp.: 89.70 Ambient Temp.: 82.49 F
22:45:51.774 -> Average Cal sensor Temp.: 89.69 Ambient Temp.: 82.49 F
22:45:53.334 -> Average Cal sensor Temp.: 89.78 Ambient Temp.: 82.53 F
22:45:54.892 -> Average Cal sensor Temp.: 89.87 Ambient Temp.: 82.56 F
22:45:56.450 -> Average Cal sensor Temp.: 89.77 Ambient Temp.: 82.53 F
22:45:58.003 -> Average Cal sensor Temp.: 89.75 Ambient Temp.: 82.56 F
22:45:59.565 -> Average Cal sensor Temp.: 89.75 Ambient Temp.: 82.60 F
22:46:01.124 -> Average Cal sensor Temp.: 89.78 Ambient Temp.: 82.63 F
22:46:02.683 -> Average Cal sensor Temp.: 89.74 Ambient Temp.: 82.60 F
22:46:04.246 -> Average Cal sensor Temp.: 89.87 Ambient Temp.: 82.63 F
22:46:05.801 -> Average Cal sensor Temp.: 89.87 Ambient Temp.: 82.74 F
22:46:07.360 -> Average Cal sensor Temp.: 89.79 Ambient Temp.: 82.71 F
22:46:08.887 -> Average Cal sensor Temp.: 89.91 Ambient Temp.: 82.71 F
22:46:10.448 -> Average Cal sensor Temp.: 89.73 Ambient Temp.: 82.78 F
22:46:12.007 -> Average Cal sensor Temp.: 89.79 Ambient Temp.: 82.78 F
22:46:13.567 -> Average Cal sensor Temp.: 89.52 Ambient Temp.: 82.78 F
22:46:15.126 -> Average Cal sensor Temp.: 89.48 Ambient Temp.: 82.74 F
22:46:16.687 -> Average Cal sensor Temp.: 89.74 Ambient Temp.: 82.78 F
22:46:18.247 -> Average Cal sensor Temp.: 89.82 Ambient Temp.: 82.78 F
22:46:19.804 -> Average Cal sensor Temp.: 89.75 Ambient Temp.: 82.78 F
22:46:21.363 -> Average Cal sensor Temp.: 89.65 Ambient Temp.: 82.89 F
22:46:22.918 -> Average Cal sensor Temp.: 89.71 Ambient Temp.: 82.78 F
22:46:24.481 -> Average Cal sensor Temp.: 89.96 Ambient Temp.: 82.85 F
22:46:26.008 -> Average Cal sensor Temp.: 89.87 Ambient Temp.: 82.92 F
22:46:27.569 -> Average Cal sensor Temp.: 90.07 Ambient Temp.: 82.96 F
22:46:29.133 -> Average Cal sensor Temp.: 90.00 Ambient Temp.: 82.92 F
22:46:30.685 -> Average Cal sensor Temp.: 89.85 Ambient Temp.: 82.96 F
22:46:32.268 -> Average Cal sensor Temp.: 90.03 Ambient Temp.: 82.96 F
22:46:33.825 -> Average Cal sensor Temp.: 90.07 Ambient Temp.: 82.92 F
22:46:35.385 -> Average Cal sensor Temp.: 90.00 Ambient Temp.: 82.96 F
22:46:36.909 -> Average Cal sensor Temp.: 89.78 Ambient Temp.: 82.92 F
22:46:38.469 -> Average Cal sensor Temp.: 89.77 Ambient Temp.: 82.89 F
22:46:40.056 -> Average Cal sensor Temp.: 89.80 Ambient Temp.: 82.96 F
22:46:41.611 -> Average Cal sensor Temp.: 89.91 Ambient Temp.: 82.96 F
22:46:43.141 -> Average Cal sensor Temp.: 90.17 Ambient Temp.: 82.99 F
22:46:44.705 -> Average Cal sensor Temp.: 90.07 Ambient Temp.: 82.96 F
22:46:46.260 -> Average Cal sensor Temp.: 90.01 Ambient Temp.: 82.89 F
22:46:47.822 -> Average Cal sensor Temp.: 89.88 Ambient Temp.: 82.92 F
22:46:49.378 -> Average Cal sensor Temp.: 90.03 Ambient Temp.: 82.89 F
22:46:50.933 -> Average Cal sensor Temp.: 89.65 Ambient Temp.: 82.85 F
22:46:52.490 -> Average Cal sensor Temp.: 89.75 Ambient Temp.: 82.78 F
22:46:54.050 -> Average Cal sensor Temp.: 89.81 Ambient Temp.: 82.89 F
22:46:55.608 -> Average Cal sensor Temp.: 89.97 Ambient Temp.: 82.89 F
22:46:57.168 -> Average Cal sensor Temp.: 89.88 Ambient Temp.: 82.85 F
22:46:58.728 -> Average Cal sensor Temp.: 89.80 Ambient Temp.: 82.89 F
22:47:00.287 -> Average Cal sensor Temp.: 89.75 Ambient Temp.: 82.78 F
22:47:01.845 -> Average Cal sensor Temp.: 89.60 Ambient Temp.: 82.78 F
22:47:03.402 -> Average Cal sensor Temp.: 89.76 Ambient Temp.: 82.71 F
22:47:04.964 -> Average Cal sensor Temp.: 89.55 Ambient Temp.: 82.74 F
22:47:06.522 -> Average Cal sensor Temp.: 89.83 Ambient Temp.: 82.71 F
22:47:08.081 -> Average Cal sensor Temp.: 89.91 Ambient Temp.: 82.71 F
22:47:09.641 -> Average Cal sensor Temp.: 89.92 Ambient Temp.: 82.67 F
22:47:11.166 -> Average Cal sensor Temp.: 89.84 Ambient Temp.: 82.67 F
22:47:12.725 -> Average Cal sensor Temp.: 89.81 Ambient Temp.: 82.63 F
22:47:14.288 -> Average Cal sensor Temp.: 89.93 Ambient Temp.: 82.63 F
22:47:15.844 -> Average Cal sensor Temp.: 89.90 Ambient Temp.: 82.60 F
22:47:17.395 -> Average Cal sensor Temp.: 89.94 Ambient Temp.: 82.63 F
22:47:18.984 -> Average Cal sensor Temp.: 90.09 Ambient Temp.: 82.63 F
22:47:20.511 -> Average Cal sensor Temp.: 89.86 Ambient Temp.: 82.56 F
22:47:22.100 -> Average Cal sensor Temp.: 89.85 Ambient Temp.: 82.56 F
22:47:23.659 -> Average Cal sensor Temp.: 90.29 Ambient Temp.: 82.60 F
22:47:25.183 -> Average Cal sensor Temp.: 90.04 Ambient Temp.: 82.56 F
22:47:26.743 -> Average Cal sensor Temp.: 89.67 Ambient Temp.: 82.49 F
22:47:28.303 -> Average Cal sensor Temp.: 89.92 Ambient Temp.: 82.53 F
22:47:29.868 -> Average Cal sensor Temp.: 90.28 Ambient Temp.: 82.56 F
22:47:31.428 -> Average Cal sensor Temp.: 90.14 Ambient Temp.: 82.53 F
22:47:32.984 -> Average Cal sensor Temp.: 90.27 Ambient Temp.: 82.49 F
22:47:34.542 -> Average Cal sensor Temp.: 90.19 Ambient Temp.: 82.53 F
22:47:36.103 -> Average Cal sensor Temp.: 90.18 Ambient Temp.: 82.49 F
22:47:37.653 -> Average Cal sensor Temp.: 90.14 Ambient Temp.: 82.49 F
22:47:39.214 -> Average Cal sensor Temp.: 90.30 Ambient Temp.: 82.49 F
22:47:40.777 -> Average Cal sensor Temp.: 90.29 Ambient Temp.: 82.45 F
22:47:42.339 -> Average Cal sensor Temp.: 90.11 Ambient Temp.: 82.49 F
22:47:43.862 -> Average Cal sensor Temp.: 90.14 Ambient Temp.: 82.45 F
22:47:45.422 -> Average Cal sensor Temp.: 90.08 Ambient Temp.: 82.53 F
22:47:46.982 -> Average Cal sensor Temp.: 89.97 Ambient Temp.: 82.38 F
22:47:48.544 -> Average Cal sensor Temp.: 90.27 Ambient Temp.: 82.35 F
22:47:50.104 -> Average Cal sensor Temp.: 90.21 Ambient Temp.: 82.38 F
22:47:51.669 -> Average Cal sensor Temp.: 90.05 Ambient Temp.: 82.35 F
22:47:53.226 -> Average Cal sensor Temp.: 90.37 Ambient Temp.: 82.38 F
22:47:54.789 -> Average Cal sensor Temp.: 90.30 Ambient Temp.: 82.38 F
22:47:56.319 -> Average Cal sensor Temp.: 90.37 Ambient Temp.: 82.38 F
22:47:57.877 -> Average Cal sensor Temp.: 90.31 Ambient Temp.: 82.38 F
22:47:59.438 -> Average Cal sensor Temp.: 90.22 Ambient Temp.: 82.38 F
22:48:01.002 -> Average Cal sensor Temp.: 90.19 Ambient Temp.: 82.35 F
22:48:02.557 -> Average Cal sensor Temp.: 90.03 Ambient Temp.: 82.31 F
22:48:04.119 -> Average Cal sensor Temp.: 90.00 Ambient Temp.: 82.27 F
22:48:05.681 -> Average Cal sensor Temp.: 90.28 Ambient Temp.: 82.24 F
22:48:07.236 -> Average Cal sensor Temp.: 90.47 Ambient Temp.: 82.24 F
22:48:08.764 -> Average Cal sensor Temp.: 90.38 Ambient Temp.: 82.24 F
22:48:10.325 -> Average Cal sensor Temp.: 90.53 Ambient Temp.: 82.24 F
22:48:11.884 -> Average Cal sensor Temp.: 90.53 Ambient Temp.: 82.20 F
22:48:13.434 -> Average Cal sensor Temp.: 90.47 Ambient Temp.: 82.24 F
22:48:14.990 -> Average Cal sensor Temp.: 90.19 Ambient Temp.: 82.24 F
22:48:16.553 -> Average Cal sensor Temp.: 90.17 Ambient Temp.: 82.20 F
22:48:18.114 -> Average Cal sensor Temp.: 90.07 Ambient Temp.: 82.17 F
22:48:19.674 -> Average Cal sensor Temp.: 90.09 Ambient Temp.: 82.20 F
22:48:21.232 -> Average Cal sensor Temp.: 89.94 Ambient Temp.: 82.13 F
22:48:22.787 -> Average Cal sensor Temp.: 89.93 Ambient Temp.: 82.13 F
22:48:24.348 -> Average Cal sensor Temp.: 90.19 Ambient Temp.: 82.06 F
22:48:25.909 -> Average Cal sensor Temp.: 90.11 Ambient Temp.: 82.09 F
22:48:27.464 -> Average Cal sensor Temp.: 90.30 Ambient Temp.: 82.06 F
22:48:28.993 -> Average Cal sensor Temp.: 90.55 Ambient Temp.: 82.09 F
22:48:30.575 -> Average Cal sensor Temp.: 90.63 Ambient Temp.: 82.06 F
22:48:32.129 -> Average Cal sensor Temp.: 90.17 Ambient Temp.: 81.95 F
22:48:33.692 -> Average Cal sensor Temp.: 90.27 Ambient Temp.: 81.95 F
22:48:35.219 -> Average Cal sensor Temp.: 90.27 Ambient Temp.: 81.91 F
22:48:36.778 -> Average Cal sensor Temp.: 90.41 Ambient Temp.: 81.91 F
22:48:38.337 -> Average Cal sensor Temp.: 90.35 Ambient Temp.: 81.95 F
22:48:39.895 -> Average Cal sensor Temp.: 90.37 Ambient Temp.: 81.88 F
22:48:41.454 -> Average Cal sensor Temp.: 90.56 Ambient Temp.: 81.84 F
22:48:43.017 -> Average Cal sensor Temp.: 90.63 Ambient Temp.: 81.88 F
22:48:44.580 -> Average Cal sensor Temp.: 90.75 Ambient Temp.: 81.91 F
22:48:46.138 -> Average Cal sensor Temp.: 90.41 Ambient Temp.: 81.88 F
22:48:47.696 -> Average Cal sensor Temp.: 90.44 Ambient Temp.: 81.88 F
22:48:49.254 -> Average Cal sensor Temp.: 90.25 Ambient Temp.: 81.84 F
22:48:50.811 -> Average Cal sensor Temp.: 89.93 Ambient Temp.: 81.77 F
22:48:52.339 -> Average Cal sensor Temp.: 90.24 Ambient Temp.: 81.77 F
22:48:53.899 -> Average Cal sensor Temp.: 90.50 Ambient Temp.: 81.73 F
22:48:55.456 -> Average Cal sensor Temp.: 90.47 Ambient Temp.: 81.81 F
22:48:57.018 -> Average Cal sensor Temp.: 90.45 Ambient Temp.: 81.77 F
22:48:58.578 -> Average Cal sensor Temp.: 90.51 Ambient Temp.: 81.73 F
22:49:00.140 -> Average Cal sensor Temp.: 90.39 Ambient Temp.: 81.77 F
22:49:01.695 -> Average Cal sensor Temp.: 90.31 Ambient Temp.: 81.66 F
22:49:03.254 -> Average Cal sensor Temp.: 90.48 Ambient Temp.: 81.66 F
22:49:04.786 -> Average Cal sensor Temp.: 90.58 Ambient Temp.: 81.59 F
22:49:06.345 -> Average Cal sensor Temp.: 90.52 Ambient Temp.: 81.66 F
22:49:07.903 -> Average Cal sensor Temp.: 90.48 Ambient Temp.: 81.55 F
22:49:09.472 -> Average Cal sensor Temp.: 90.77 Ambient Temp.: 81.59 F
22:49:11.033 -> Average Cal sensor Temp.: 90.50 Ambient Temp.: 81.59 F
22:49:12.590 -> Average Cal sensor Temp.: 90.41 Ambient Temp.: 81.59 F
22:49:14.114 -> Average Cal sensor Temp.: 90.66 Ambient Temp.: 81.55 F
22:49:15.673 -> Average Cal sensor Temp.: 90.31 Ambient Temp.: 81.55 F
22:49:17.257 -> Average Cal sensor Temp.: 90.29 Ambient Temp.: 81.48 F
22:49:18.814 -> Average Cal sensor Temp.: 90.36 Ambient Temp.: 81.45 F
22:49:20.339 -> Average Cal sensor Temp.: 90.68 Ambient Temp.: 81.45 F
22:49:21.900 -> Average Cal sensor Temp.: 90.38 Ambient Temp.: 81.45 F
22:49:23.464 -> Average Cal sensor Temp.: 90.80 Ambient Temp.: 81.48 F
22:49:25.023 -> Average Cal sensor Temp.: 90.50 Ambient Temp.: 81.41 F
22:49:26.583 -> Average Cal sensor Temp.: 90.61 Ambient Temp.: 81.45 F
22:49:28.145 -> Average Cal sensor Temp.: 90.25 Ambient Temp.: 81.37 F
22:49:29.677 -> Average Cal sensor Temp.: 90.52 Ambient Temp.: 81.41 F
22:49:31.237 -> Average Cal sensor Temp.: 90.33 Ambient Temp.: 81.41 F
22:49:32.798 -> Average Cal sensor Temp.: 90.28 Ambient Temp.: 81.34 F
22:49:34.365 -> Average Cal sensor Temp.: 90.36 Ambient Temp.: 81.30 F
22:49:35.930 -> Average Cal sensor Temp.: 90.24 Ambient Temp.: 81.30 F
22:49:37.457 -> Average Cal sensor Temp.: 90.54 Ambient Temp.: 81.30 F
22:49:39.023 -> Average Cal sensor Temp.: 90.44 Ambient Temp.: 81.30 F
22:49:40.580 -> Average Cal sensor Temp.: 90.05 Ambient Temp.: 81.16 F
22:49:42.144 -> Average Cal sensor Temp.: 90.41 Ambient Temp.: 81.19 F
22:49:43.705 -> Average Cal sensor Temp.: 90.36 Ambient Temp.: 81.27 F
22:49:45.262 -> Average Cal sensor Temp.: 90.47 Ambient Temp.: 81.16 F
22:49:46.820 -> Average Cal sensor Temp.: 90.42 Ambient Temp.: 81.19 F
22:49:48.376 -> Average Cal sensor Temp.: 90.45 Ambient Temp.: 81.16 F
22:49:49.902 -> Average Cal sensor Temp.: 90.46 Ambient Temp.: 81.12 F
22:49:51.461 -> Average Cal sensor Temp.: 90.58 Ambient Temp.: 81.16 F
22:49:53.022 -> Average Cal sensor Temp.: 90.60 Ambient Temp.: 81.12 F
22:49:54.587 -> Average Cal sensor Temp.: 90.13 Ambient Temp.: 81.12 F
22:49:56.148 -> Average Cal sensor Temp.: 90.44 Ambient Temp.: 81.09 F
22:49:57.707 -> Average Cal sensor Temp.: 90.76 Ambient Temp.: 81.05 F
22:49:59.261 -> Average Cal sensor Temp.: 90.87 Ambient Temp.: 81.05 F
22:50:00.821 -> Average Cal sensor Temp.: 90.90 Ambient Temp.: 81.09 F
22:50:02.374 -> Average Cal sensor Temp.: 90.74 Ambient Temp.: 81.09 F
22:50:03.934 -> Average Cal sensor Temp.: 90.96 Ambient Temp.: 81.09 F
22:50:05.489 -> Average Cal sensor Temp.: 90.97 Ambient Temp.: 81.09 F
22:50:07.050 -> Average Cal sensor Temp.: 90.87 Ambient Temp.: 81.12 F
22:50:08.601 -> Average Cal sensor Temp.: 90.84 Ambient Temp.: 81.16 F
22:50:10.158 -> Average Cal sensor Temp.: 90.85 Ambient Temp.: 81.09 F
22:50:11.715 -> Average Cal sensor Temp.: 90.52 Ambient Temp.: 81.09 F
22:50:13.273 -> Average Cal sensor Temp.: 90.81 Ambient Temp.: 81.12 F
22:50:14.835 -> Average Cal sensor Temp.: 90.80 Ambient Temp.: 81.12 F
22:50:16.360 -> Average Cal sensor Temp.: 90.77 Ambient Temp.: 81.12 F
22:50:17.915 -> Average Cal sensor Temp.: 91.04 Ambient Temp.: 81.12 F
22:50:19.479 -> Average Cal sensor Temp.: 91.09 Ambient Temp.: 81.12 F
22:50:21.039 -> Average Cal sensor Temp.: 91.16 Ambient Temp.: 81.19 F
22:50:22.596 -> Average Cal sensor Temp.: 91.30 Ambient Temp.: 81.19 F
22:50:24.157 -> Average Cal sensor Temp.: 90.97 Ambient Temp.: 81.16 F
22:50:25.715 -> Average Cal sensor Temp.: 91.08 Ambient Temp.: 81.27 F
22:50:27.274 -> Average Cal sensor Temp.: 90.85 Ambient Temp.: 81.27 F
22:50:28.803 -> Average Cal sensor Temp.: 90.98 Ambient Temp.: 81.27 F
22:50:30.366 -> Average Cal sensor Temp.: 91.15 Ambient Temp.: 81.34 F
22:50:31.929 -> Average Cal sensor Temp.: 91.23 Ambient Temp.: 81.37 F
22:50:33.486 -> Average Cal sensor Temp.: 90.86 Ambient Temp.: 81.45 F
22:50:35.048 -> Average Cal sensor Temp.: 90.82 Ambient Temp.: 81.41 F
22:50:36.605 -> Average Cal sensor Temp.: 90.86 Ambient Temp.: 81.41 F
22:50:38.132 -> Average Cal sensor Temp.: 90.83 Ambient Temp.: 81.45 F
22:50:39.693 -> Average Cal sensor Temp.: 91.06 Ambient Temp.: 81.45 F
22:50:41.259 -> Average Cal sensor Temp.: 91.02 Ambient Temp.: 81.48 F
22:50:42.817 -> Average Cal sensor Temp.: 91.24 Ambient Temp.: 81.48 F
22:50:44.381 -> Average Cal sensor Temp.: 90.82 Ambient Temp.: 81.55 F
22:50:45.938 -> Average Cal sensor Temp.: 91.01 Ambient Temp.: 81.55 F
22:50:47.495 -> Average Cal sensor Temp.: 91.12 Ambient Temp.: 81.70 F
22:50:49.018 -> Average Cal sensor Temp.: 91.04 Ambient Temp.: 81.70 F
22:50:50.576 -> Average Cal sensor Temp.: 91.01 Ambient Temp.: 81.70 F
22:50:52.133 -> Average Cal sensor Temp.: 91.02 Ambient Temp.: 81.77 F
22:50:53.695 -> Average Cal sensor Temp.: 90.93 Ambient Temp.: 81.59 F
22:50:55.250 -> Average Cal sensor Temp.: 90.94 Ambient Temp.: 81.77 F
22:50:56.807 -> Average Cal sensor Temp.: 91.05 Ambient Temp.: 81.77 F
22:50:58.363 -> Average Cal sensor Temp.: 90.95 Ambient Temp.: 81.77 F
22:50:59.923 -> Average Cal sensor Temp.: 90.98 Ambient Temp.: 81.73 F
22:51:01.477 -> Average Cal sensor Temp.: 91.15 Ambient Temp.: 81.81 F
22:51:03.026 -> Average Cal sensor Temp.: 91.20 Ambient Temp.: 81.88 F
22:51:04.579 -> Average Cal sensor Temp.: 91.13 Ambient Temp.: 81.91 F
22:51:06.132 -> Average Cal sensor Temp.: 91.03 Ambient Temp.: 81.88 F
22:51:07.692 -> Average Cal sensor Temp.: 91.12 Ambient Temp.: 81.95 F
22:51:09.250 -> Average Cal sensor Temp.: 91.27 Ambient Temp.: 81.95 F
22:51:10.811 -> Average Cal sensor Temp.: 91.22 Ambient Temp.: 81.95 F
22:51:12.339 -> Average Cal sensor Temp.: 91.20 Ambient Temp.: 82.09 F
22:51:13.898 -> Average Cal sensor Temp.: 91.23 Ambient Temp.: 81.99 F
22:51:15.458 -> Average Cal sensor Temp.: 91.06 Ambient Temp.: 81.99 F
22:51:17.010 -> Average Cal sensor Temp.: 91.22 Ambient Temp.: 82.06 F
22:51:18.562 -> Average Cal sensor Temp.: 91.24 Ambient Temp.: 82.09 F
22:51:20.151 -> Average Cal sensor Temp.: 91.21 Ambient Temp.: 82.13 F
22:51:21.674 -> Average Cal sensor Temp.: 91.30 Ambient Temp.: 82.09 F
22:51:23.234 -> Average Cal sensor Temp.: 91.36 Ambient Temp.: 82.13 F
22:51:24.794 -> Average Cal sensor Temp.: 91.42 Ambient Temp.: 82.17 F
22:51:26.343 -> Average Cal sensor Temp.: 91.23 Ambient Temp.: 82.24 F
22:51:27.900 -> Average Cal sensor Temp.: 91.12 Ambient Temp.: 82.27 F
22:51:29.465 -> Average Cal sensor Temp.: 91.13 Ambient Temp.: 82.27 F
22:51:31.020 -> Average Cal sensor Temp.: 91.25 Ambient Temp.: 82.27 F
22:51:32.571 -> Average Cal sensor Temp.: 91.39 Ambient Temp.: 82.31 F
22:51:34.130 -> Average Cal sensor Temp.: 91.18 Ambient Temp.: 82.35 F
22:51:35.688 -> Average Cal sensor Temp.: 91.01 Ambient Temp.: 82.38 F
22:51:37.244 -> Average Cal sensor Temp.: 91.13 Ambient Temp.: 82.38 F
22:51:38.770 -> Average Cal sensor Temp.: 91.17 Ambient Temp.: 82.35 F
22:51:40.323 -> Average Cal sensor Temp.: 91.26 Ambient Temp.: 82.38 F
22:51:41.881 -> Average Cal sensor Temp.: 90.95 Ambient Temp.: 82.45 F
22:51:43.442 -> Average Cal sensor Temp.: 91.04 Ambient Temp.: 82.49 F
22:51:45.002 -> Average Cal sensor Temp.: 90.91 Ambient Temp.: 82.53 F
22:51:46.557 -> Average Cal sensor Temp.: 91.35 Ambient Temp.: 82.53 F
22:51:48.118 -> Average Cal sensor Temp.: 91.06 Ambient Temp.: 82.53 F
22:51:49.679 -> Average Cal sensor Temp.: 91.19 Ambient Temp.: 82.56 F
22:51:51.239 -> Average Cal sensor Temp.: 90.98 Ambient Temp.: 82.56 F
22:51:52.796 -> Average Cal sensor Temp.: 90.98 Ambient Temp.: 82.60 F
22:51:54.356 -> Average Cal sensor Temp.: 91.03 Ambient Temp.: 82.60 F
22:51:55.910 -> Average Cal sensor Temp.: 91.29 Ambient Temp.: 82.60 F
22:51:57.472 -> Average Cal sensor Temp.: 91.22 Ambient Temp.: 82.67 F
22:51:58.996 -> Average Cal sensor Temp.: 91.12 Ambient Temp.: 82.71 F
22:52:00.556 -> Average Cal sensor Temp.: 91.15 Ambient Temp.: 82.67 F
22:52:02.112 -> Average Cal sensor Temp.: 91.39 Ambient Temp.: 82.71 F
22:52:03.664 -> Average Cal sensor Temp.: 91.19 Ambient Temp.: 82.71 F
22:52:05.216 -> Average Cal sensor Temp.: 91.24 Ambient Temp.: 82.74 F
22:52:06.774 -> Average Cal sensor Temp.: 91.32 Ambient Temp.: 82.78 F
22:52:08.331 -> Average Cal sensor Temp.: 91.20 Ambient Temp.: 82.78 F
22:52:09.890 -> Average Cal sensor Temp.: 91.48 Ambient Temp.: 82.78 F
22:52:11.449 -> Average Cal sensor Temp.: 91.48 Ambient Temp.: 82.89 F
22:52:13.008 -> Average Cal sensor Temp.: 91.35 Ambient Temp.: 82.89 F
22:52:14.568 -> Average Cal sensor Temp.: 91.31 Ambient Temp.: 82.92 F
22:52:16.097 -> Average Cal sensor Temp.: 91.53 Ambient Temp.: 82.99 F
22:52:17.657 -> Average Cal sensor Temp.: 91.59 Ambient Temp.: 82.99 F
22:52:19.222 -> Average Cal sensor Temp.: 91.49 Ambient Temp.: 82.96 F
22:52:20.786 -> Average Cal sensor Temp.: 91.35 Ambient Temp.: 83.03 F
22:52:22.347 -> Average Cal sensor Temp.: 91.53 Ambient Temp.: 83.03 F
22:52:23.907 -> Average Cal sensor Temp.: 91.62 Ambient Temp.: 83.03 F
22:52:25.435 -> Average Cal sensor Temp.: 91.44 Ambient Temp.: 83.10 F
22:52:26.997 -> Average Cal sensor Temp.: 91.44 Ambient Temp.: 83.14 F
22:52:28.561 -> Average Cal sensor Temp.: 91.43 Ambient Temp.: 83.10 F
22:52:30.122 -> Average Cal sensor Temp.: 91.48 Ambient Temp.: 83.14 F
22:52:31.679 -> Average Cal sensor Temp.: 91.51 Ambient Temp.: 83.14 F
22:52:33.233 -> Average Cal sensor Temp.: 91.40 Ambient Temp.: 83.25 F
22:52:34.762 -> Average Cal sensor Temp.: 91.44 Ambient Temp.: 83.17 F
22:52:36.321 -> Average Cal sensor Temp.: 91.39 Ambient Temp.: 83.28 F
22:52:37.876 -> Average Cal sensor Temp.: 91.31 Ambient Temp.: 83.25 F
22:52:39.438 -> Average Cal sensor Temp.: 91.36 Ambient Temp.: 83.25 F
22:52:40.993 -> Average Cal sensor Temp.: 91.28 Ambient Temp.: 83.32 F
22:52:42.545 -> Average Cal sensor Temp.: 91.45 Ambient Temp.: 83.32 F
22:52:44.100 -> Average Cal sensor Temp.: 91.30 Ambient Temp.: 83.35 F
22:52:45.659 -> Average Cal sensor Temp.: 91.35 Ambient Temp.: 83.39 F
22:52:47.215 -> Average Cal sensor Temp.: 91.29 Ambient Temp.: 83.39 F
22:52:48.766 -> Average Cal sensor Temp.: 91.29 Ambient Temp.: 83.39 F
22:52:50.327 -> Average Cal sensor Temp.: 91.56 Ambient Temp.: 83.39 F
22:52:51.877 -> Average Cal sensor Temp.: 91.65 Ambient Temp.: 83.50 F
22:52:53.432 -> Average Cal sensor Temp.: 91.38 Ambient Temp.: 83.46 F
22:52:54.991 -> Average Cal sensor Temp.: 91.54 Ambient Temp.: 83.50 F
22:52:56.552 -> Average Cal sensor Temp.: 91.40 Ambient Temp.: 83.46 F
22:52:58.112 -> Average Cal sensor Temp.: 91.64 Ambient Temp.: 83.50 F
22:52:59.672 -> Average Cal sensor Temp.: 91.36 Ambient Temp.: 83.57 F
22:53:01.227 -> Average Cal sensor Temp.: 91.48 Ambient Temp.: 83.53 F
22:53:02.752 -> Average Cal sensor Temp.: 91.36 Ambient Temp.: 83.53 F
22:53:04.308 -> Average Cal sensor Temp.: 91.50 Ambient Temp.: 83.57 F
22:53:05.863 -> Average Cal sensor Temp.: 91.80 Ambient Temp.: 83.57 F
22:53:07.422 -> Average Cal sensor Temp.: 91.77 Ambient Temp.: 83.57 F
22:53:08.979 -> Average Cal sensor Temp.: 91.50 Ambient Temp.: 83.57 F
22:53:10.536 -> Average Cal sensor Temp.: 91.48 Ambient Temp.: 83.57 F
22:53:12.100 -> Average Cal sensor Temp.: 91.40 Ambient Temp.: 83.68 F
22:53:13.661 -> Average Cal sensor Temp.: 91.44 Ambient Temp.: 83.57 F
22:53:15.220 -> Average Cal sensor Temp.: 91.42 Ambient Temp.: 83.57 F
22:53:16.775 -> Average Cal sensor Temp.: 91.52 Ambient Temp.: 83.64 F
22:53:18.302 -> Average Cal sensor Temp.: 91.61 Ambient Temp.: 83.57 F
22:53:19.865 -> Average Cal sensor Temp.: 91.48 Ambient Temp.: 83.68 F
22:53:21.428 -> Average Cal sensor Temp.: 91.59 Ambient Temp.: 83.71 F
22:53:22.977 -> Average Cal sensor Temp.: 91.54 Ambient Temp.: 83.57 F
22:53:24.536 -> Average Cal sensor Temp.: 91.61 Ambient Temp.: 83.68 F
22:53:26.091 -> Average Cal sensor Temp.: 91.76 Ambient Temp.: 83.71 F
22:53:27.646 -> Average Cal sensor Temp.: 91.91 Ambient Temp.: 83.68 F
22:53:29.208 -> Average Cal sensor Temp.: 91.80 Ambient Temp.: 83.71 F
22:53:30.765 -> Average Cal sensor Temp.: 91.75 Ambient Temp.: 83.71 F
22:53:32.323 -> Average Cal sensor Temp.: 91.53 Ambient Temp.: 83.75 F
22:53:33.847 -> Average Cal sensor Temp.: 91.51 Ambient Temp.: 83.71 F
22:53:35.408 -> Average Cal sensor Temp.: 91.60 Ambient Temp.: 83.71 F
22:53:36.957 -> Average Cal sensor Temp.: 91.80 Ambient Temp.: 83.71 F
22:53:38.518 -> Average Cal sensor Temp.: 91.76 Ambient Temp.: 83.71 F
22:53:40.074 -> Average Cal sensor Temp.: 91.64 Ambient Temp.: 83.71 F
22:53:41.633 -> Average Cal sensor Temp.: 91.63 Ambient Temp.: 83.75 F
22:53:43.197 -> Average Cal sensor Temp.: 91.72 Ambient Temp.: 83.68 F
22:53:44.759 -> Average Cal sensor Temp.: 92.10 Ambient Temp.: 83.68 F
22:53:46.286 -> Average Cal sensor Temp.: 91.96 Ambient Temp.: 83.71 F
22:53:47.844 -> Average Cal sensor Temp.: 91.66 Ambient Temp.: 83.68 F
22:53:49.411 -> Average Cal sensor Temp.: 91.52 Ambient Temp.: 83.68 F
22:53:50.963 -> Average Cal sensor Temp.: 91.62 Ambient Temp.: 83.68 F
22:53:52.520 -> Average Cal sensor Temp.: 91.56 Ambient Temp.: 83.64 F
22:53:54.079 -> Average Cal sensor Temp.: 91.31 Ambient Temp.: 83.53 F
22:53:55.642 -> Average Cal sensor Temp.: 91.17 Ambient Temp.: 83.53 F
22:53:57.198 -> Average Cal sensor Temp.: 91.08 Ambient Temp.: 83.57 F
22:53:58.754 -> Average Cal sensor Temp.: 91.42 Ambient Temp.: 83.53 F
22:54:00.313 -> Average Cal sensor Temp.: 91.53 Ambient Temp.: 83.46 F
22:54:01.868 -> Average Cal sensor Temp.: 91.34 Ambient Temp.: 83.50 F
22:54:03.391 -> Average Cal sensor Temp.: 91.48 Ambient Temp.: 83.50 F
22:54:04.950 -> Average Cal sensor Temp.: 91.44 Ambient Temp.: 83.53 F
22:54:06.512 -> Average Cal sensor Temp.: 91.38 Ambient Temp.: 83.50 F
22:54:08.066 -> Average Cal sensor Temp.: 91.28 Ambient Temp.: 83.50 F
22:54:09.630 -> Average Cal sensor Temp.: 91.35 Ambient Temp.: 83.46 F
22:54:11.188 -> Average Cal sensor Temp.: 91.35 Ambient Temp.: 83.39 F
22:54:12.748 -> Average Cal sensor Temp.: 91.58 Ambient Temp.: 83.43 F
22:54:14.312 -> Average Cal sensor Temp.: 91.68 Ambient Temp.: 83.46 F
22:54:15.835 -> Average Cal sensor Temp.: 91.61 Ambient Temp.: 83.43 F
22:54:17.390 -> Average Cal sensor Temp.: 91.56 Ambient Temp.: 83.35 F
22:54:18.946 -> Average Cal sensor Temp.: 91.52 Ambient Temp.: 83.35 F
22:54:20.531 -> Average Cal sensor Temp.: 91.65 Ambient Temp.: 83.39 F
22:54:22.053 -> Average Cal sensor Temp.: 91.76 Ambient Temp.: 83.35 F
22:54:23.639 -> Average Cal sensor Temp.: 91.55 Ambient Temp.: 83.35 F
22:54:25.166 -> Average Cal sensor Temp.: 91.56 Ambient Temp.: 83.32 F
22:54:26.719 -> Average Cal sensor Temp.: 91.76 Ambient Temp.: 83.32 F
22:54:28.278 -> Average Cal sensor Temp.: 91.71 Ambient Temp.: 83.28 F
22:54:29.835 -> Average Cal sensor Temp.: 91.53 Ambient Temp.: 83.32 F
22:54:31.394 -> Average Cal sensor Temp.: 91.57 Ambient Temp.: 83.32 F
22:54:32.954 -> Average Cal sensor Temp.: 91.48 Ambient Temp.: 83.17 F
22:54:34.514 -> Average Cal sensor Temp.: 91.10 Ambient Temp.: 83.17 F
22:54:36.042 -> Average Cal sensor Temp.: 90.94 Ambient Temp.: 83.17 F
22:54:37.630 -> Average Cal sensor Temp.: 91.42 Ambient Temp.: 83.17 F
22:54:39.159 -> Average Cal sensor Temp.: 91.41 Ambient Temp.: 83.14 F
22:54:40.715 -> Average Cal sensor Temp.: 91.71 Ambient Temp.: 83.07 F
22:54:42.273 -> Average Cal sensor Temp.: 91.32 Ambient Temp.: 83.10 F
22:54:43.831 -> Average Cal sensor Temp.: 91.51 Ambient Temp.: 83.03 F
22:54:45.390 -> Average Cal sensor Temp.: 91.29 Ambient Temp.: 83.07 F
22:54:46.944 -> Average Cal sensor Temp.: 91.00 Ambient Temp.: 83.03 F
22:54:48.502 -> Average Cal sensor Temp.: 91.31 Ambient Temp.: 82.99 F
22:54:50.058 -> Average Cal sensor Temp.: 91.66 Ambient Temp.: 82.96 F
22:54:51.614 -> Average Cal sensor Temp.: 91.31 Ambient Temp.: 82.96 F
22:54:53.176 -> Average Cal sensor Temp.: 91.71 Ambient Temp.: 82.96 F
22:54:54.704 -> Average Cal sensor Temp.: 91.63 Ambient Temp.: 82.92 F
22:54:56.260 -> Average Cal sensor Temp.: 91.78 Ambient Temp.: 82.89 F
22:54:57.821 -> Average Cal sensor Temp.: 91.58 Ambient Temp.: 82.96 F
22:54:59.384 -> Average Cal sensor Temp.: 91.80 Ambient Temp.: 82.89 F
22:55:00.937 -> Average Cal sensor Temp.: 91.43 Ambient Temp.: 82.89 F
22:55:02.497 -> Average Cal sensor Temp.: 91.34 Ambient Temp.: 82.85 F
22:55:04.062 -> Average Cal sensor Temp.: 91.43 Ambient Temp.: 82.74 F
22:55:05.622 -> Average Cal sensor Temp.: 91.55 Ambient Temp.: 82.74 F
22:55:07.145 -> Average Cal sensor Temp.: 91.55 Ambient Temp.: 82.71 F
22:55:08.700 -> Average Cal sensor Temp.: 91.44 Ambient Temp.: 82.67 F
22:55:10.260 -> Average Cal sensor Temp.: 91.28 Ambient Temp.: 82.67 F
22:55:11.814 -> Average Cal sensor Temp.: 91.67 Ambient Temp.: 82.60 F
22:55:13.374 -> Average Cal sensor Temp.: 91.95 Ambient Temp.: 82.63 F
22:55:14.934 -> Average Cal sensor Temp.: 92.15 Ambient Temp.: 82.60 F
22:55:16.488 -> Average Cal sensor Temp.: 91.70 Ambient Temp.: 82.67 F
22:55:18.042 -> Average Cal sensor Temp.: 91.89 Ambient Temp.: 82.60 F
22:55:19.600 -> Average Cal sensor Temp.: 91.59 Ambient Temp.: 82.63 F
22:55:21.154 -> Average Cal sensor Temp.: 91.51 Ambient Temp.: 82.60 F
22:55:22.707 -> Average Cal sensor Temp.: 91.54 Ambient Temp.: 82.60 F
22:55:24.265 -> Average Cal sensor Temp.: 91.67 Ambient Temp.: 82.56 F
22:55:25.827 -> Average Cal sensor Temp.: 92.11 Ambient Temp.: 82.56 F
22:55:27.385 -> Average Cal sensor Temp.: 91.58 Ambient Temp.: 82.56 F
22:55:28.949 -> Average Cal sensor Temp.: 91.76 Ambient Temp.: 82.53 F
22:55:30.504 -> Average Cal sensor Temp.: 91.42 Ambient Temp.: 82.53 F
22:55:32.032 -> Average Cal sensor Temp.: 91.71 Ambient Temp.: 82.53 F
22:55:33.592 -> Average Cal sensor Temp.: 91.97 Ambient Temp.: 82.49 F
22:55:35.153 -> Average Cal sensor Temp.: 91.58 Ambient Temp.: 82.45 F
22:55:36.703 -> Average Cal sensor Temp.: 91.39 Ambient Temp.: 82.38 F
22:55:38.264 -> Average Cal sensor Temp.: 91.42 Ambient Temp.: 82.35 F
22:55:39.821 -> Average Cal sensor Temp.: 91.57 Ambient Temp.: 82.35 F
22:55:41.382 -> Average Cal sensor Temp.: 91.70 Ambient Temp.: 82.31 F
22:55:42.943 -> Average Cal sensor Temp.: 91.71 Ambient Temp.: 82.27 F
22:55:44.467 -> Average Cal sensor Temp.: 92.04 Ambient Temp.: 82.31 F
22:55:46.028 -> Average Cal sensor Temp.: 91.71 Ambient Temp.: 82.31 F
22:55:47.594 -> Average Cal sensor Temp.: 91.83 Ambient Temp.: 82.27 F
22:55:49.159 -> Average Cal sensor Temp.: 92.13 Ambient Temp.: 82.27 F
22:55:50.681 -> Average Cal sensor Temp.: 92.97 Ambient Temp.: 82.27 F
22:55:52.243 -> Average Cal sensor Temp.: 93.06 Ambient Temp.: 82.27 F
22:55:53.805 -> Average Cal sensor Temp.: 92.91 Ambient Temp.: 82.27 F
22:55:55.363 -> Average Cal sensor Temp.: 92.56 Ambient Temp.: 82.27 F
22:55:56.914 -> Average Cal sensor Temp.: 92.73 Ambient Temp.: 82.24 F
22:55:58.475 -> Average Cal sensor Temp.: 92.64 Ambient Temp.: 82.27 F
22:56:00.036 -> Average Cal sensor Temp.: 92.43 Ambient Temp.: 82.24 F
22:56:01.597 -> Average Cal sensor Temp.: 92.40 Ambient Temp.: 82.17 F
22:56:03.150 -> Average Cal sensor Temp.: 92.20 Ambient Temp.: 82.24 F
22:56:04.710 -> Average Cal sensor Temp.: 92.12 Ambient Temp.: 82.20 F
22:56:06.263 -> Average Cal sensor Temp.: 92.14 Ambient Temp.: 82.24 F
22:56:07.792 -> Average Cal sensor Temp.: 92.07 Ambient Temp.: 82.24 F
22:56:09.350 -> Average Cal sensor Temp.: 91.91 Ambient Temp.: 82.17 F
22:56:10.909 -> Average Cal sensor Temp.: 92.13 Ambient Temp.: 82.17 F
22:56:12.469 -> Average Cal sensor Temp.: 92.39 Ambient Temp.: 82.20 F
22:56:14.027 -> Average Cal sensor Temp.: 92.44 Ambient Temp.: 82.24 F
22:56:15.580 -> Average Cal sensor Temp.: 92.44 Ambient Temp.: 82.27 F
22:56:17.141 -> Average Cal sensor Temp.: 92.31 Ambient Temp.: 82.24 F
22:56:18.701 -> Average Cal sensor Temp.: 92.29 Ambient Temp.: 82.31 F
22:56:20.263 -> Average Cal sensor Temp.: 92.22 Ambient Temp.: 82.27 F
22:56:21.818 -> Average Cal sensor Temp.: 92.21 Ambient Temp.: 82.27 F
22:56:23.375 -> Average Cal sensor Temp.: 92.29 Ambient Temp.: 82.35 F
22:56:24.900 -> Average Cal sensor Temp.: 92.17 Ambient Temp.: 82.27 F
22:56:26.460 -> Average Cal sensor Temp.: 92.05 Ambient Temp.: 82.31 F
22:56:28.012 -> Average Cal sensor Temp.: 92.18 Ambient Temp.: 82.27 F
22:56:29.569 -> Average Cal sensor Temp.: 92.15 Ambient Temp.: 82.35 F
22:56:31.125 -> Average Cal sensor Temp.: 92.17 Ambient Temp.: 82.31 F
22:56:32.681 -> Average Cal sensor Temp.: 92.14 Ambient Temp.: 82.31 F
22:56:34.236 -> Average Cal sensor Temp.: 92.02 Ambient Temp.: 82.31 F
22:56:35.789 -> Average Cal sensor Temp.: 92.21 Ambient Temp.: 82.35 F
22:56:37.350 -> Average Cal sensor Temp.: 91.95 Ambient Temp.: 82.35 F
22:56:38.915 -> Average Cal sensor Temp.: 92.07 Ambient Temp.: 82.38 F
22:56:40.474 -> Average Cal sensor Temp.: 92.21 Ambient Temp.: 82.35 F
22:56:41.998 -> Average Cal sensor Temp.: 92.14 Ambient Temp.: 82.31 F
22:56:43.558 -> Average Cal sensor Temp.: 91.96 Ambient Temp.: 82.38 F
22:56:45.117 -> Average Cal sensor Temp.: 91.79 Ambient Temp.: 82.38 F
22:56:46.673 -> Average Cal sensor Temp.: 91.89 Ambient Temp.: 82.35 F
22:56:48.234 -> Average Cal sensor Temp.: 92.04 Ambient Temp.: 82.31 F
22:56:49.793 -> Average Cal sensor Temp.: 91.87 Ambient Temp.: 82.31 F
22:56:51.349 -> Average Cal sensor Temp.: 91.79 Ambient Temp.: 82.24 F
22:56:52.911 -> Average Cal sensor Temp.: 92.12 Ambient Temp.: 82.27 F
22:56:54.470 -> Average Cal sensor Temp.: 91.91 Ambient Temp.: 82.27 F
22:56:56.026 -> Average Cal sensor Temp.: 92.18 Ambient Temp.: 82.35 F
22:56:57.552 -> Average Cal sensor Temp.: 91.89 Ambient Temp.: 82.27 F
22:56:59.109 -> Average Cal sensor Temp.: 91.91 Ambient Temp.: 82.24 F
22:57:00.669 -> Average Cal sensor Temp.: 92.13 Ambient Temp.: 82.27 F
22:57:02.227 -> Average Cal sensor Temp.: 92.17 Ambient Temp.: 82.24 F
22:57:03.783 -> Average Cal sensor Temp.: 92.18 Ambient Temp.: 82.20 F
22:57:05.340 -> Average Cal sensor Temp.: 92.51 Ambient Temp.: 82.27 F
22:57:06.897 -> Average Cal sensor Temp.: 92.54 Ambient Temp.: 82.31 F
22:57:08.450 -> Average Cal sensor Temp.: 92.32 Ambient Temp.: 82.27 F
22:57:10.010 -> Average Cal sensor Temp.: 92.26 Ambient Temp.: 82.31 F
22:57:11.563 -> Average Cal sensor Temp.: 92.52 Ambient Temp.: 82.31 F
22:57:13.124 -> Average Cal sensor Temp.: 92.58 Ambient Temp.: 82.31 F
22:57:14.651 -> Average Cal sensor Temp.: 92.69 Ambient Temp.: 82.38 F
22:57:16.241 -> Average Cal sensor Temp.: 92.70 Ambient Temp.: 82.45 F
22:57:17.795 -> Average Cal sensor Temp.: 92.99 Ambient Temp.: 82.45 F
22:57:19.320 -> Average Cal sensor Temp.: 92.73 Ambient Temp.: 82.38 F
22:57:20.879 -> Average Cal sensor Temp.: 92.57 Ambient Temp.: 82.49 F
22:57:22.439 -> Average Cal sensor Temp.: 92.88 Ambient Temp.: 82.49 F
22:57:23.992 -> Average Cal sensor Temp.: 92.76 Ambient Temp.: 82.45 F
22:57:25.554 -> Average Cal sensor Temp.: 92.80 Ambient Temp.: 82.56 F
22:57:27.113 -> Average Cal sensor Temp.: 92.87 Ambient Temp.: 82.53 F
22:57:28.665 -> Average Cal sensor Temp.: 92.97 Ambient Temp.: 82.56 F
22:57:30.223 -> Average Cal sensor Temp.: 92.94 Ambient Temp.: 82.60 F
22:57:31.779 -> Average Cal sensor Temp.: 92.71 Ambient Temp.: 82.63 F
22:57:33.333 -> Average Cal sensor Temp.: 92.67 Ambient Temp.: 82.56 F
22:57:34.888 -> Average Cal sensor Temp.: 92.64 Ambient Temp.: 82.71 F
22:57:36.449 -> Average Cal sensor Temp.: 92.87 Ambient Temp.: 82.67 F
22:57:38.004 -> Average Cal sensor Temp.: 92.79 Ambient Temp.: 82.67 F
22:57:39.568 -> Average Cal sensor Temp.: 92.80 Ambient Temp.: 82.74 F
22:57:41.126 -> Average Cal sensor Temp.: 92.52 Ambient Temp.: 82.78 F
22:57:42.689 -> Average Cal sensor Temp.: 92.41 Ambient Temp.: 82.78 F
22:57:44.248 -> Average Cal sensor Temp.: 92.66 Ambient Temp.: 82.85 F
22:57:45.806 -> Average Cal sensor Temp.: 93.09 Ambient Temp.: 82.85 F
22:57:47.361 -> Average Cal sensor Temp.: 92.84 Ambient Temp.: 82.89 F
22:57:48.916 -> Average Cal sensor Temp.: 93.01 Ambient Temp.: 82.89 F
22:57:50.472 -> Average Cal sensor Temp.: 92.91 Ambient Temp.: 82.89 F
22:57:52.028 -> Average Cal sensor Temp.: 92.66 Ambient Temp.: 82.96 F
22:57:53.585 -> Average Cal sensor Temp.: 92.67 Ambient Temp.: 82.92 F
22:57:55.144 -> Average Cal sensor Temp.: 92.73 Ambient Temp.: 83.03 F
22:57:56.697 -> Average Cal sensor Temp.: 92.64 Ambient Temp.: 83.03 F
22:57:58.254 -> Average Cal sensor Temp.: 92.85 Ambient Temp.: 83.03 F
22:57:59.814 -> Average Cal sensor Temp.: 92.81 Ambient Temp.: 83.03 F
22:58:01.370 -> Average Cal sensor Temp.: 92.59 Ambient Temp.: 83.10 F
22:58:02.893 -> Average Cal sensor Temp.: 92.72 Ambient Temp.: 83.14 F
22:58:04.454 -> Average Cal sensor Temp.: 92.80 Ambient Temp.: 83.25 F
22:58:06.008 -> Average Cal sensor Temp.: 92.68 Ambient Temp.: 83.25 F
22:58:07.566 -> Average Cal sensor Temp.: 92.66 Ambient Temp.: 83.28 F
22:58:09.124 -> Average Cal sensor Temp.: 92.76 Ambient Temp.: 83.35 F
22:58:10.687 -> Average Cal sensor Temp.: 92.75 Ambient Temp.: 83.28 F
22:58:12.249 -> Average Cal sensor Temp.: 92.75 Ambient Temp.: 83.39 F
22:58:13.806 -> Average Cal sensor Temp.: 92.61 Ambient Temp.: 83.46 F
22:58:15.359 -> Average Cal sensor Temp.: 92.70 Ambient Temp.: 83.43 F
22:58:16.909 -> Average Cal sensor Temp.: 92.59 Ambient Temp.: 83.50 F
22:58:18.472 -> Average Cal sensor Temp.: 92.69 Ambient Temp.: 83.50 F
22:58:20.031 -> Average Cal sensor Temp.: 92.78 Ambient Temp.: 83.50 F
22:58:21.591 -> Average Cal sensor Temp.: 92.75 Ambient Temp.: 83.53 F
22:58:23.117 -> Average Cal sensor Temp.: 92.92 Ambient Temp.: 83.64 F
22:58:24.677 -> Average Cal sensor Temp.: 92.94 Ambient Temp.: 83.71 F
22:58:26.235 -> Average Cal sensor Temp.: 92.80 Ambient Temp.: 83.64 F
22:58:27.797 -> Average Cal sensor Temp.: 92.90 Ambient Temp.: 83.75 F
22:58:29.355 -> Average Cal sensor Temp.: 92.85 Ambient Temp.: 83.75 F
22:58:30.908 -> Average Cal sensor Temp.: 92.68 Ambient Temp.: 83.79 F
22:58:32.465 -> Average Cal sensor Temp.: 92.58 Ambient Temp.: 83.82 F
22:58:34.023 -> Average Cal sensor Temp.: 92.63 Ambient Temp.: 83.86 F
22:58:35.579 -> Average Cal sensor Temp.: 92.78 Ambient Temp.: 83.89 F
22:58:37.139 -> Average Cal sensor Temp.: 92.66 Ambient Temp.: 83.89 F
22:58:38.696 -> Average Cal sensor Temp.: 92.76 Ambient Temp.: 83.93 F
22:58:40.253 -> Average Cal sensor Temp.: 92.70 Ambient Temp.: 84.00 F
22:58:41.804 -> Average Cal sensor Temp.: 92.87 Ambient Temp.: 84.07 F
22:58:43.359 -> Average Cal sensor Temp.: 92.74 Ambient Temp.: 84.07 F
22:58:44.916 -> Average Cal sensor Temp.: 92.80 Ambient Temp.: 84.11 F
22:58:46.475 -> Average Cal sensor Temp.: 92.81 Ambient Temp.: 84.15 F
22:58:48.030 -> Average Cal sensor Temp.: 92.92 Ambient Temp.: 84.18 F
22:58:49.589 -> Average Cal sensor Temp.: 92.92 Ambient Temp.: 84.18 F
22:58:51.141 -> Average Cal sensor Temp.: 92.92 Ambient Temp.: 84.25 F
22:58:52.700 -> Average Cal sensor Temp.: 92.95 Ambient Temp.: 84.33 F
22:58:54.252 -> Average Cal sensor Temp.: 92.92 Ambient Temp.: 84.25 F
22:58:55.807 -> Average Cal sensor Temp.: 92.91 Ambient Temp.: 84.33 F
22:58:57.364 -> Average Cal sensor Temp.: 93.02 Ambient Temp.: 84.36 F
22:58:58.928 -> Average Cal sensor Temp.: 93.02 Ambient Temp.: 84.36 F
22:59:00.489 -> Average Cal sensor Temp.: 92.97 Ambient Temp.: 84.40 F
22:59:02.043 -> Average Cal sensor Temp.: 93.07 Ambient Temp.: 84.40 F
22:59:03.598 -> Average Cal sensor Temp.: 93.06 Ambient Temp.: 84.47 F
22:59:05.154 -> Average Cal sensor Temp.: 92.94 Ambient Temp.: 84.51 F
22:59:06.711 -> Average Cal sensor Temp.: 92.78 Ambient Temp.: 84.54 F
22:59:08.268 -> Average Cal sensor Temp.: 92.90 Ambient Temp.: 84.58 F
22:59:09.823 -> Average Cal sensor Temp.: 93.00 Ambient Temp.: 84.61 F
22:59:11.385 -> Average Cal sensor Temp.: 93.20 Ambient Temp.: 84.61 F
22:59:12.942 -> Average Cal sensor Temp.: 93.22 Ambient Temp.: 84.65 F
22:59:14.504 -> Average Cal sensor Temp.: 93.17 Ambient Temp.: 84.69 F
22:59:16.062 -> Average Cal sensor Temp.: 92.93 Ambient Temp.: 84.72 F
22:59:17.591 -> Average Cal sensor Temp.: 93.04 Ambient Temp.: 84.76 F
22:59:19.151 -> Average Cal sensor Temp.: 93.16 Ambient Temp.: 84.76 F
22:59:20.713 -> Average Cal sensor Temp.: 92.99 Ambient Temp.: 84.87 F
22:59:22.275 -> Average Cal sensor Temp.: 93.13 Ambient Temp.: 84.87 F
22:59:23.836 -> Average Cal sensor Temp.: 93.27 Ambient Temp.: 84.87 F
22:59:25.392 -> Average Cal sensor Temp.: 93.02 Ambient Temp.: 84.94 F
22:59:26.955 -> Average Cal sensor Temp.: 92.95 Ambient Temp.: 84.94 F
22:59:28.516 -> Average Cal sensor Temp.: 93.07 Ambient Temp.: 84.94 F
22:59:30.072 -> Average Cal sensor Temp.: 93.15 Ambient Temp.: 85.01 F
22:59:31.595 -> Average Cal sensor Temp.: 93.10 Ambient Temp.: 84.97 F
22:59:33.150 -> Average Cal sensor Temp.: 93.13 Ambient Temp.: 85.05 F
22:59:34.707 -> Average Cal sensor Temp.: 93.17 Ambient Temp.: 85.08 F
22:59:36.276 -> Average Cal sensor Temp.: 93.21 Ambient Temp.: 85.08 F
22:59:37.831 -> Average Cal sensor Temp.: 93.16 Ambient Temp.: 85.15 F
22:59:39.397 -> Average Cal sensor Temp.: 92.96 Ambient Temp.: 85.12 F
22:59:40.957 -> Average Cal sensor Temp.: 93.16 Ambient Temp.: 85.19 F
22:59:42.484 -> Average Cal sensor Temp.: 93.01 Ambient Temp.: 85.26 F
22:59:44.043 -> Average Cal sensor Temp.: 93.29 Ambient Temp.: 85.19 F
22:59:45.591 -> Average Cal sensor Temp.: 93.37 Ambient Temp.: 85.26 F
22:59:47.154 -> Average Cal sensor Temp.: 93.17 Ambient Temp.: 85.26 F
22:59:48.713 -> Average Cal sensor Temp.: 93.33 Ambient Temp.: 85.30 F
22:59:50.272 -> Average Cal sensor Temp.: 93.23 Ambient Temp.: 85.30 F
22:59:51.836 -> Average Cal sensor Temp.: 93.22 Ambient Temp.: 85.30 F
22:59:53.388 -> Average Cal sensor Temp.: 93.15 Ambient Temp.: 85.37 F
22:59:54.939 -> Average Cal sensor Temp.: 93.24 Ambient Temp.: 85.37 F
22:59:56.494 -> Average Cal sensor Temp.: 93.42 Ambient Temp.: 85.37 F
22:59:58.050 -> Average Cal sensor Temp.: 93.35 Ambient Temp.: 85.41 F
22:59:59.604 -> Average Cal sensor Temp.: 93.17 Ambient Temp.: 85.48 F
23:00:01.158 -> Average Cal sensor Temp.: 93.22 Ambient Temp.: 85.48 F
23:00:02.712 -> Average Cal sensor Temp.: 93.41 Ambient Temp.: 85.51 F
23:00:04.272 -> Average Cal sensor Temp.: 93.36 Ambient Temp.: 85.48 F
23:00:05.834 -> Average Cal sensor Temp.: 93.39 Ambient Temp.: 85.48 F
23:00:07.391 -> Average Cal sensor Temp.: 93.33 Ambient Temp.: 85.59 F
23:00:08.920 -> Average Cal sensor Temp.: 93.29 Ambient Temp.: 85.51 F
23:00:10.481 -> Average Cal sensor Temp.: 93.48 Ambient Temp.: 85.51 F
23:00:12.032 -> Average Cal sensor Temp.: 93.46 Ambient Temp.: 85.59 F
23:00:13.595 -> Average Cal sensor Temp.: 93.45 Ambient Temp.: 85.59 F
23:00:15.154 -> Average Cal sensor Temp.: 93.46 Ambient Temp.: 85.66 F
23:00:16.704 -> Average Cal sensor Temp.: 93.47 Ambient Temp.: 85.69 F
23:00:18.261 -> Average Cal sensor Temp.: 93.46 Ambient Temp.: 85.66 F
23:00:19.817 -> Average Cal sensor Temp.: 93.51 Ambient Temp.: 85.69 F
23:00:21.369 -> Average Cal sensor Temp.: 93.46 Ambient Temp.: 85.66 F
23:00:22.953 -> Average Cal sensor Temp.: 93.49 Ambient Temp.: 85.73 F
23:00:24.510 -> Average Cal sensor Temp.: 93.41 Ambient Temp.: 85.69 F
23:00:26.065 -> Average Cal sensor Temp.: 93.33 Ambient Temp.: 85.77 F
23:00:27.594 -> Average Cal sensor Temp.: 93.34 Ambient Temp.: 85.73 F
23:00:29.155 -> Average Cal sensor Temp.: 93.40 Ambient Temp.: 85.77 F
23:00:30.715 -> Average Cal sensor Temp.: 93.49 Ambient Temp.: 85.80 F
23:00:32.266 -> Average Cal sensor Temp.: 93.69 Ambient Temp.: 85.80 F
23:00:33.823 -> Average Cal sensor Temp.: 93.32 Ambient Temp.: 85.77 F
23:00:35.383 -> Average Cal sensor Temp.: 93.42 Ambient Temp.: 85.77 F
23:00:36.941 -> Average Cal sensor Temp.: 93.47 Ambient Temp.: 85.77 F
23:00:38.500 -> Average Cal sensor Temp.: 93.03 Ambient Temp.: 85.73 F
23:00:40.057 -> Average Cal sensor Temp.: 93.14 Ambient Temp.: 85.77 F
23:00:41.621 -> Average Cal sensor Temp.: 93.31 Ambient Temp.: 85.73 F
23:00:43.141 -> Average Cal sensor Temp.: 93.13 Ambient Temp.: 85.77 F
23:00:44.706 -> Average Cal sensor Temp.: 93.25 Ambient Temp.: 85.80 F
23:00:46.266 -> Average Cal sensor Temp.: 93.25 Ambient Temp.: 85.77 F
23:00:47.822 -> Average Cal sensor Temp.: 93.31 Ambient Temp.: 85.77 F
23:00:49.384 -> Average Cal sensor Temp.: 93.36 Ambient Temp.: 85.77 F
23:00:50.939 -> Average Cal sensor Temp.: 93.17 Ambient Temp.: 85.77 F
23:00:52.490 -> Average Cal sensor Temp.: 93.13 Ambient Temp.: 85.73 F
23:00:54.047 -> Average Cal sensor Temp.: 92.88 Ambient Temp.: 85.69 F
23:00:55.609 -> Average Cal sensor Temp.: 93.34 Ambient Temp.: 85.66 F
23:00:57.149 -> Average Cal sensor Temp.: 93.52 Ambient Temp.: 85.69 F
23:00:58.716 -> Average Cal sensor Temp.: 93.38 Ambient Temp.: 85.69 F
23:01:00.273 -> Average Cal sensor Temp.: 93.34 Ambient Temp.: 85.69 F
23:01:01.831 -> Average Cal sensor Temp.: 92.95 Ambient Temp.: 85.66 F
23:01:03.387 -> Average Cal sensor Temp.: 93.39 Ambient Temp.: 85.69 F
23:01:04.947 -> Average Cal sensor Temp.: 93.37 Ambient Temp.: 85.69 F
23:01:06.505 -> Average Cal sensor Temp.: 93.52 Ambient Temp.: 85.66 F
23:01:08.029 -> Average Cal sensor Temp.: 93.31 Ambient Temp.: 85.66 F
23:01:09.584 -> Average Cal sensor Temp.: 93.12 Ambient Temp.: 85.55 F
23:01:11.139 -> Average Cal sensor Temp.: 93.22 Ambient Temp.: 85.59 F
23:01:12.697 -> Average Cal sensor Temp.: 93.30 Ambient Temp.: 85.55 F
23:01:14.253 -> Average Cal sensor Temp.: 93.41 Ambient Temp.: 85.55 F
23:01:15.817 -> Average Cal sensor Temp.: 93.21 Ambient Temp.: 85.51 F
23:01:17.377 -> Average Cal sensor Temp.: 93.51 Ambient Temp.: 85.55 F
23:01:18.940 -> Average Cal sensor Temp.: 93.65 Ambient Temp.: 85.55 F
23:01:20.497 -> Average Cal sensor Temp.: 93.54 Ambient Temp.: 85.51 F
23:01:22.049 -> Average Cal sensor Temp.: 93.53 Ambient Temp.: 85.51 F
23:01:23.573 -> Average Cal sensor Temp.: 93.62 Ambient Temp.: 85.48 F
23:01:25.132 -> Average Cal sensor Temp.: 93.28 Ambient Temp.: 85.44 F
23:01:26.719 -> Average Cal sensor Temp.: 93.35 Ambient Temp.: 85.44 F
23:01:28.268 -> Average Cal sensor Temp.: 93.25 Ambient Temp.: 85.41 F
23:01:29.830 -> Average Cal sensor Temp.: 93.47 Ambient Temp.: 85.44 F
23:01:31.385 -> Average Cal sensor Temp.: 93.58 Ambient Temp.: 85.37 F
23:01:32.912 -> Average Cal sensor Temp.: 93.65 Ambient Temp.: 85.41 F
23:01:34.469 -> Average Cal sensor Temp.: 93.38 Ambient Temp.: 85.37 F
23:01:36.021 -> Average Cal sensor Temp.: 93.45 Ambient Temp.: 85.33 F
23:01:37.575 -> Average Cal sensor Temp.: 93.40 Ambient Temp.: 85.33 F
23:01:39.133 -> Average Cal sensor Temp.: 93.41 Ambient Temp.: 85.33 F
23:01:40.689 -> Average Cal sensor Temp.: 93.55 Ambient Temp.: 85.33 F
23:01:42.247 -> Average Cal sensor Temp.: 93.48 Ambient Temp.: 85.30 F
23:01:43.808 -> Average Cal sensor Temp.: 93.38 Ambient Temp.: 85.26 F
23:01:45.368 -> Average Cal sensor Temp.: 93.47 Ambient Temp.: 85.19 F
23:01:46.923 -> Average Cal sensor Temp.: 93.65 Ambient Temp.: 85.19 F
23:01:48.482 -> Average Cal sensor Temp.: 93.53 Ambient Temp.: 85.19 F
23:01:50.039 -> Average Cal sensor Temp.: 93.50 Ambient Temp.: 85.19 F
23:01:51.593 -> Average Cal sensor Temp.: 93.71 Ambient Temp.: 85.15 F
23:01:53.150 -> Average Cal sensor Temp.: 93.73 Ambient Temp.: 85.15 F
23:01:54.709 -> Average Cal sensor Temp.: 93.65 Ambient Temp.: 85.12 F
23:01:56.235 -> Average Cal sensor Temp.: 93.73 Ambient Temp.: 85.15 F
23:01:57.793 -> Average Cal sensor Temp.: 93.81 Ambient Temp.: 85.08 F
23:01:59.357 -> Average Cal sensor Temp.: 93.67 Ambient Temp.: 85.05 F
23:02:00.910 -> Average Cal sensor Temp.: 93.70 Ambient Temp.: 85.08 F
23:02:02.462 -> Average Cal sensor Temp.: 93.71 Ambient Temp.: 85.08 F
23:02:04.019 -> Average Cal sensor Temp.: 93.46 Ambient Temp.: 85.05 F
23:02:05.578 -> Average Cal sensor Temp.: 93.65 Ambient Temp.: 85.01 F
23:02:07.136 -> Average Cal sensor Temp.: 93.64 Ambient Temp.: 84.97 F
23:02:08.702 -> Average Cal sensor Temp.: 93.56 Ambient Temp.: 85.01 F
23:02:10.260 -> Average Cal sensor Temp.: 93.53 Ambient Temp.: 85.01 F
23:02:11.822 -> Average Cal sensor Temp.: 93.60 Ambient Temp.: 84.97 F
23:02:13.378 -> Average Cal sensor Temp.: 93.63 Ambient Temp.: 84.90 F
23:02:14.937 -> Average Cal sensor Temp.: 93.51 Ambient Temp.: 84.90 F
23:02:16.461 -> Average Cal sensor Temp.: 93.56 Ambient Temp.: 84.87 F
23:02:18.020 -> Average Cal sensor Temp.: 93.62 Ambient Temp.: 84.87 F
23:02:19.580 -> Average Cal sensor Temp.: 93.68 Ambient Temp.: 84.79 F
23:02:21.141 -> Average Cal sensor Temp.: 93.78 Ambient Temp.: 84.76 F
23:02:22.703 -> Average Cal sensor Temp.: 93.67 Ambient Temp.: 84.76 F
23:02:24.268 -> Average Cal sensor Temp.: 93.35 Ambient Temp.: 84.72 F
23:02:25.828 -> Average Cal sensor Temp.: 93.38 Ambient Temp.: 84.69 F
23:02:27.382 -> Average Cal sensor Temp.: 93.42 Ambient Temp.: 84.69 F
23:02:28.941 -> Average Cal sensor Temp.: 93.87 Ambient Temp.: 84.72 F
23:02:30.495 -> Average Cal sensor Temp.: 93.62 Ambient Temp.: 84.65 F
23:02:32.054 -> Average Cal sensor Temp.: 93.46 Ambient Temp.: 84.65 F
23:02:33.614 -> Average Cal sensor Temp.: 93.61 Ambient Temp.: 84.58 F
23:02:35.142 -> Average Cal sensor Temp.: 93.62 Ambient Temp.: 84.61 F
23:02:36.698 -> Average Cal sensor Temp.: 93.65 Ambient Temp.: 84.61 F
23:02:38.259 -> Average Cal sensor Temp.: 93.34 Ambient Temp.: 84.58 F
23:02:39.821 -> Average Cal sensor Temp.: 93.68 Ambient Temp.: 84.54 F
23:02:41.381 -> Average Cal sensor Temp.: 93.43 Ambient Temp.: 84.54 F
23:02:42.941 -> Average Cal sensor Temp.: 93.35 Ambient Temp.: 84.54 F
23:02:44.506 -> Average Cal sensor Temp.: 93.27 Ambient Temp.: 84.51 F
23:02:46.033 -> Average Cal sensor Temp.: 93.26 Ambient Temp.: 84.40 F
23:02:47.593 -> Average Cal sensor Temp.: 93.15 Ambient Temp.: 84.36 F
23:02:49.153 -> Average Cal sensor Temp.: 93.45 Ambient Temp.: 84.36 F
23:02:50.712 -> Average Cal sensor Temp.: 93.67 Ambient Temp.: 84.33 F
23:02:52.276 -> Average Cal sensor Temp.: 93.57 Ambient Temp.: 84.36 F
23:02:53.836 -> Average Cal sensor Temp.: 93.44 Ambient Temp.: 84.25 F
23:02:55.393 -> Average Cal sensor Temp.: 93.38 Ambient Temp.: 84.22 F
23:02:56.948 -> Average Cal sensor Temp.: 93.83 Ambient Temp.: 84.22 F
23:02:58.507 -> Average Cal sensor Temp.: 93.55 Ambient Temp.: 84.18 F
23:03:00.064 -> Average Cal sensor Temp.: 93.34 Ambient Temp.: 84.18 F
23:03:01.588 -> Average Cal sensor Temp.: 93.57 Ambient Temp.: 84.18 F
23:03:03.142 -> Average Cal sensor Temp.: 93.59 Ambient Temp.: 84.18 F
23:03:04.698 -> Average Cal sensor Temp.: 93.63 Ambient Temp.: 84.15 F
23:03:06.260 -> Average Cal sensor Temp.: 93.66 Ambient Temp.: 84.07 F
23:03:07.824 -> Average Cal sensor Temp.: 93.78 Ambient Temp.: 84.11 F
23:03:09.381 -> Average Cal sensor Temp.: 93.92 Ambient Temp.: 84.11 F
23:03:10.934 -> Average Cal sensor Temp.: 93.74 Ambient Temp.: 84.00 F
23:03:12.483 -> Average Cal sensor Temp.: 93.81 Ambient Temp.: 84.00 F
23:03:14.047 -> Average Cal sensor Temp.: 93.70 Ambient Temp.: 83.97 F
23:03:15.606 -> Average Cal sensor Temp.: 93.76 Ambient Temp.: 83.93 F
23:03:17.168 -> Average Cal sensor Temp.: 94.06 Ambient Temp.: 83.97 F
23:03:18.731 -> Average Cal sensor Temp.: 94.12 Ambient Temp.: 83.97 F
23:03:20.291 -> Average Cal sensor Temp.: 94.15 Ambient Temp.: 84.00 F
23:03:21.817 -> Average Cal sensor Temp.: 94.04 Ambient Temp.: 83.97 F
23:03:23.373 -> Average Cal sensor Temp.: 93.85 Ambient Temp.: 83.93 F
23:03:24.931 -> Average Cal sensor Temp.: 93.64 Ambient Temp.: 83.93 F
23:03:26.518 -> Average Cal sensor Temp.: 93.73 Ambient Temp.: 83.86 F
23:03:28.072 -> Average Cal sensor Temp.: 93.92 Ambient Temp.: 83.86 F
23:03:29.601 -> Average Cal sensor Temp.: 93.83 Ambient Temp.: 83.86 F
23:03:31.159 -> Average Cal sensor Temp.: 93.96 Ambient Temp.: 83.86 F
23:03:32.716 -> Average Cal sensor Temp.: 93.77 Ambient Temp.: 83.82 F
23:03:34.276 -> Average Cal sensor Temp.: 93.79 Ambient Temp.: 83.79 F
23:03:35.828 -> Average Cal sensor Temp.: 93.96 Ambient Temp.: 83.75 F
23:03:37.385 -> Average Cal sensor Temp.: 93.91 Ambient Temp.: 83.71 F
23:03:38.939 -> Average Cal sensor Temp.: 93.63 Ambient Temp.: 83.71 F
23:03:40.501 -> Average Cal sensor Temp.: 93.67 Ambient Temp.: 83.71 F
23:03:42.061 -> Average Cal sensor Temp.: 94.02 Ambient Temp.: 83.71 F
23:03:43.615 -> Average Cal sensor Temp.: 94.01 Ambient Temp.: 83.71 F
23:03:45.175 -> Average Cal sensor Temp.: 93.87 Ambient Temp.: 83.64 F
23:03:46.731 -> Average Cal sensor Temp.: 94.08 Ambient Temp.: 83.64 F
23:03:48.292 -> Average Cal sensor Temp.: 93.65 Ambient Temp.: 83.64 F
23:03:49.855 -> Average Cal sensor Temp.: 94.08 Ambient Temp.: 83.64 F
23:03:51.415 -> Average Cal sensor Temp.: 93.96 Ambient Temp.: 83.57 F
23:03:52.968 -> Average Cal sensor Temp.: 93.40 Ambient Temp.: 83.46 F
23:03:54.526 -> Average Cal sensor Temp.: 93.59 Ambient Temp.: 83.46 F
23:03:56.083 -> Average Cal sensor Temp.: 93.46 Ambient Temp.: 83.43 F
23:03:57.629 -> Average Cal sensor Temp.: 94.01 Ambient Temp.: 83.43 F
23:03:59.190 -> Average Cal sensor Temp.: 93.88 Ambient Temp.: 83.43 F
23:04:00.721 -> Average Cal sensor Temp.: 94.15 Ambient Temp.: 83.39 F
23:04:02.279 -> Average Cal sensor Temp.: 93.99 Ambient Temp.: 83.46 F
23:04:03.839 -> Average Cal sensor Temp.: 93.94 Ambient Temp.: 83.39 F
23:04:05.401 -> Average Cal sensor Temp.: 94.28 Ambient Temp.: 83.39 F
23:04:06.953 -> Average Cal sensor Temp.: 94.18 Ambient Temp.: 83.43 F
23:04:08.507 -> Average Cal sensor Temp.: 94.16 Ambient Temp.: 83.39 F
23:04:10.069 -> Average Cal sensor Temp.: 93.99 Ambient Temp.: 83.32 F
23:04:11.628 -> Average Cal sensor Temp.: 93.72 Ambient Temp.: 83.39 F
23:04:13.185 -> Average Cal sensor Temp.: 93.92 Ambient Temp.: 83.35 F
23:04:14.739 -> Average Cal sensor Temp.: 94.10 Ambient Temp.: 83.32 F
23:04:16.290 -> Average Cal sensor Temp.: 94.15 Ambient Temp.: 83.32 F
23:04:17.846 -> Average Cal sensor Temp.: 94.10 Ambient Temp.: 83.35 F
23:04:19.403 -> Average Cal sensor Temp.: 94.09 Ambient Temp.: 83.35 F
23:04:20.955 -> Average Cal sensor Temp.: 94.11 Ambient Temp.: 83.35 F
23:04:22.514 -> Average Cal sensor Temp.: 94.13 Ambient Temp.: 83.43 F
23:04:24.074 -> Average Cal sensor Temp.: 94.11 Ambient Temp.: 83.39 F
23:04:25.634 -> Average Cal sensor Temp.: 94.01 Ambient Temp.: 83.39 F
23:04:27.194 -> Average Cal sensor Temp.: 93.88 Ambient Temp.: 83.35 F
23:04:28.752 -> Average Cal sensor Temp.: 94.02 Ambient Temp.: 83.39 F
23:04:30.310 -> Average Cal sensor Temp.: 94.06 Ambient Temp.: 83.35 F
23:04:31.866 -> Average Cal sensor Temp.: 94.12 Ambient Temp.: 83.35 F
23:04:33.424 -> Average Cal sensor Temp.: 94.38 Ambient Temp.: 83.39 F
23:04:34.979 -> Average Cal sensor Temp.: 94.59 Ambient Temp.: 83.39 F
23:04:36.538 -> Average Cal sensor Temp.: 94.39 Ambient Temp.: 83.43 F
23:04:38.062 -> Average Cal sensor Temp.: 93.94 Ambient Temp.: 83.43 F
23:04:39.624 -> Average Cal sensor Temp.: 94.01 Ambient Temp.: 83.39 F
23:04:41.184 -> Average Cal sensor Temp.: 94.29 Ambient Temp.: 83.46 F
23:04:42.741 -> Average Cal sensor Temp.: 94.23 Ambient Temp.: 83.46 F
23:04:44.302 -> Average Cal sensor Temp.: 94.31 Ambient Temp.: 83.53 F
23:04:45.858 -> Average Cal sensor Temp.: 94.22 Ambient Temp.: 83.50 F
23:04:47.413 -> Average Cal sensor Temp.: 94.25 Ambient Temp.: 83.57 F
23:04:48.972 -> Average Cal sensor Temp.: 94.33 Ambient Temp.: 83.53 F
23:04:50.535 -> Average Cal sensor Temp.: 94.28 Ambient Temp.: 83.57 F
23:04:52.089 -> Average Cal sensor Temp.: 94.28 Ambient Temp.: 83.57 F
23:04:53.650 -> Average Cal sensor Temp.: 94.34 Ambient Temp.: 83.64 F
23:04:55.178 -> Average Cal sensor Temp.: 94.62 Ambient Temp.: 83.68 F
23:04:56.734 -> Average Cal sensor Temp.: 94.41 Ambient Temp.: 83.71 F
23:04:58.294 -> Average Cal sensor Temp.: 94.42 Ambient Temp.: 83.71 F
23:04:59.857 -> Average Cal sensor Temp.: 94.35 Ambient Temp.: 83.71 F
23:05:01.415 -> Average Cal sensor Temp.: 94.31 Ambient Temp.: 83.79 F
23:05:02.973 -> Average Cal sensor Temp.: 94.16 Ambient Temp.: 83.79 F
23:05:04.534 -> Average Cal sensor Temp.: 94.19 Ambient Temp.: 83.79 F
23:05:06.087 -> Average Cal sensor Temp.: 94.41 Ambient Temp.: 83.82 F
23:05:07.644 -> Average Cal sensor Temp.: 94.45 Ambient Temp.: 83.82 F
23:05:09.207 -> Average Cal sensor Temp.: 94.52 Ambient Temp.: 83.82 F
23:05:10.735 -> Average Cal sensor Temp.: 94.71 Ambient Temp.: 83.89 F
23:05:12.297 -> Average Cal sensor Temp.: 94.64 Ambient Temp.: 83.82 F
23:05:13.863 -> Average Cal sensor Temp.: 94.66 Ambient Temp.: 83.86 F
23:05:15.425 -> Average Cal sensor Temp.: 94.78 Ambient Temp.: 83.93 F
23:05:16.982 -> Average Cal sensor Temp.: 94.93 Ambient Temp.: 83.93 F
23:05:18.536 -> Average Cal sensor Temp.: 94.59 Ambient Temp.: 84.00 F
23:05:20.094 -> Average Cal sensor Temp.: 94.38 Ambient Temp.: 83.97 F
23:05:21.652 -> Average Cal sensor Temp.: 94.55 Ambient Temp.: 84.07 F
23:05:23.215 -> Average Cal sensor Temp.: 94.28 Ambient Temp.: 84.11 F
23:05:24.772 -> Average Cal sensor Temp.: 94.28 Ambient Temp.: 84.11 F
23:05:26.331 -> Average Cal sensor Temp.: 94.26 Ambient Temp.: 84.11 F
23:05:27.887 -> Average Cal sensor Temp.: 94.17 Ambient Temp.: 84.11 F
23:05:29.441 -> Average Cal sensor Temp.: 94.36 Ambient Temp.: 84.15 F
23:05:30.994 -> Average Cal sensor Temp.: 94.38 Ambient Temp.: 84.18 F
23:05:32.546 -> Average Cal sensor Temp.: 94.39 Ambient Temp.: 84.18 F
23:05:34.101 -> Average Cal sensor Temp.: 94.38 Ambient Temp.: 84.22 F
23:05:35.662 -> Average Cal sensor Temp.: 94.46 Ambient Temp.: 84.22 F
23:05:37.219 -> Average Cal sensor Temp.: 94.44 Ambient Temp.: 84.18 F
23:05:38.778 -> Average Cal sensor Temp.: 94.49 Ambient Temp.: 84.22 F
23:05:40.336 -> Average Cal sensor Temp.: 94.91 Ambient Temp.: 84.33 F
23:05:41.890 -> Average Cal sensor Temp.: 94.31 Ambient Temp.: 84.33 F
23:05:43.449 -> Average Cal sensor Temp.: 94.31 Ambient Temp.: 84.36 F
23:05:45.002 -> Average Cal sensor Temp.: 94.42 Ambient Temp.: 84.47 F
23:05:46.559 -> Average Cal sensor Temp.: 94.43 Ambient Temp.: 84.36 F
23:05:48.123 -> Average Cal sensor Temp.: 94.24 Ambient Temp.: 84.40 F
23:05:49.655 -> Average Cal sensor Temp.: 94.50 Ambient Temp.: 84.40 F
23:05:51.208 -> Average Cal sensor Temp.: 94.48 Ambient Temp.: 84.40 F
23:05:52.762 -> Average Cal sensor Temp.: 94.63 Ambient Temp.: 84.47 F
23:05:54.325 -> Average Cal sensor Temp.: 94.81 Ambient Temp.: 84.54 F
23:05:55.886 -> Average Cal sensor Temp.: 94.73 Ambient Temp.: 84.54 F
23:05:57.446 -> Average Cal sensor Temp.: 94.59 Ambient Temp.: 84.54 F
23:05:59.007 -> Average Cal sensor Temp.: 94.70 Ambient Temp.: 84.54 F
23:06:00.564 -> Average Cal sensor Temp.: 94.69 Ambient Temp.: 84.58 F
23:06:02.120 -> Average Cal sensor Temp.: 94.46 Ambient Temp.: 84.58 F
23:06:03.673 -> Average Cal sensor Temp.: 94.59 Ambient Temp.: 84.65 F
23:06:05.237 -> Average Cal sensor Temp.: 94.67 Ambient Temp.: 84.61 F
23:06:06.796 -> Average Cal sensor Temp.: 94.71 Ambient Temp.: 84.61 F
23:06:08.354 -> Average Cal sensor Temp.: 94.55 Ambient Temp.: 84.72 F
23:06:09.908 -> Average Cal sensor Temp.: 94.68 Ambient Temp.: 84.72 F
23:06:11.464 -> Average Cal sensor Temp.: 94.67 Ambient Temp.: 84.69 F
23:06:13.019 -> Average Cal sensor Temp.: 94.86 Ambient Temp.: 84.72 F
23:06:14.574 -> Average Cal sensor Temp.: 94.79 Ambient Temp.: 84.72 F
23:06:16.125 -> Average Cal sensor Temp.: 94.80 Ambient Temp.: 84.87 F
23:06:17.677 -> Average Cal sensor Temp.: 94.74 Ambient Temp.: 84.87 F
23:06:19.260 -> Average Cal sensor Temp.: 94.58 Ambient Temp.: 84.79 F
23:06:20.814 -> Average Cal sensor Temp.: 94.59 Ambient Temp.: 84.87 F
23:06:22.365 -> Average Cal sensor Temp.: 94.73 Ambient Temp.: 84.90 F
23:06:23.922 -> Average Cal sensor Temp.: 94.81 Ambient Temp.: 84.90 F
23:06:25.479 -> Average Cal sensor Temp.: 94.70 Ambient Temp.: 84.87 F
23:06:27.037 -> Average Cal sensor Temp.: 94.57 Ambient Temp.: 84.97 F
23:06:28.592 -> Average Cal sensor Temp.: 94.59 Ambient Temp.: 84.94 F
23:06:30.151 -> Average Cal sensor Temp.: 94.90 Ambient Temp.: 84.94 F
23:06:31.696 -> Average Cal sensor Temp.: 94.87 Ambient Temp.: 84.97 F
23:06:33.253 -> Average Cal sensor Temp.: 95.18 Ambient Temp.: 85.08 F
23:06:34.816 -> Average Cal sensor Temp.: 95.01 Ambient Temp.: 85.08 F
23:06:36.373 -> Average Cal sensor Temp.: 95.13 Ambient Temp.: 85.08 F
23:06:37.937 -> Average Cal sensor Temp.: 95.11 Ambient Temp.: 85.12 F
23:06:39.467 -> Average Cal sensor Temp.: 94.96 Ambient Temp.: 85.08 F
23:06:41.021 -> Average Cal sensor Temp.: 94.77 Ambient Temp.: 85.12 F
23:06:42.578 -> Average Cal sensor Temp.: 94.59 Ambient Temp.: 85.12 F
23:06:44.134 -> Average Cal sensor Temp.: 95.06 Ambient Temp.: 85.12 F
23:06:45.718 -> Average Cal sensor Temp.: 94.99 Ambient Temp.: 85.15 F
23:06:47.277 -> Average Cal sensor Temp.: 94.90 Ambient Temp.: 85.19 F
23:06:48.800 -> Average Cal sensor Temp.: 95.06 Ambient Temp.: 85.30 F
23:06:50.363 -> Average Cal sensor Temp.: 94.81 Ambient Temp.: 85.19 F
23:06:51.915 -> Average Cal sensor Temp.: 94.71 Ambient Temp.: 85.26 F
23:06:53.476 -> Average Cal sensor Temp.: 94.99 Ambient Temp.: 85.30 F
23:06:55.041 -> Average Cal sensor Temp.: 95.00 Ambient Temp.: 85.33 F
23:06:56.608 -> Average Cal sensor Temp.: 94.96 Ambient Temp.: 85.33 F
23:06:58.169 -> Average Cal sensor Temp.: 94.68 Ambient Temp.: 85.37 F
23:06:59.728 -> Average Cal sensor Temp.: 95.05 Ambient Temp.: 85.41 F
23:07:01.250 -> Average Cal sensor Temp.: 94.99 Ambient Temp.: 85.44 F
23:07:02.841 -> Average Cal sensor Temp.: 94.99 Ambient Temp.: 85.41 F
23:07:04.369 -> Average Cal sensor Temp.: 94.91 Ambient Temp.: 85.41 F
23:07:05.928 -> Average Cal sensor Temp.: 94.77 Ambient Temp.: 85.44 F
23:07:07.484 -> Average Cal sensor Temp.: 94.95 Ambient Temp.: 85.44 F
23:07:09.044 -> Average Cal sensor Temp.: 94.87 Ambient Temp.: 85.48 F
23:07:10.599 -> Average Cal sensor Temp.: 94.70 Ambient Temp.: 85.48 F
23:07:12.161 -> Average Cal sensor Temp.: 94.88 Ambient Temp.: 85.48 F
23:07:13.718 -> Average Cal sensor Temp.: 94.77 Ambient Temp.: 85.48 F
23:07:15.276 -> Average Cal sensor Temp.: 94.82 Ambient Temp.: 85.48 F
23:07:16.837 -> Average Cal sensor Temp.: 95.10 Ambient Temp.: 85.51 F
23:07:18.402 -> Average Cal sensor Temp.: 95.17 Ambient Temp.: 85.55 F
23:07:19.927 -> Average Cal sensor Temp.: 94.83 Ambient Temp.: 85.59 F
23:07:21.514 -> Average Cal sensor Temp.: 95.06 Ambient Temp.: 85.51 F
23:07:23.039 -> Average Cal sensor Temp.: 95.01 Ambient Temp.: 85.59 F
23:07:24.604 -> Average Cal sensor Temp.: 94.64 Ambient Temp.: 85.66 F
23:07:26.164 -> Average Cal sensor Temp.: 94.71 Ambient Temp.: 85.59 F
23:07:27.728 -> Average Cal sensor Temp.: 94.97 Ambient Temp.: 85.66 F
23:07:29.292 -> Average Cal sensor Temp.: 95.05 Ambient Temp.: 85.69 F
23:07:30.848 -> Average Cal sensor Temp.: 94.72 Ambient Temp.: 85.59 F
23:07:32.376 -> Average Cal sensor Temp.: 95.02 Ambient Temp.: 85.59 F
23:07:33.937 -> Average Cal sensor Temp.: 95.16 Ambient Temp.: 85.73 F
23:07:35.497 -> Average Cal sensor Temp.: 95.01 Ambient Temp.: 85.66 F
23:07:37.062 -> Average Cal sensor Temp.: 94.66 Ambient Temp.: 85.66 F
23:07:38.626 -> Average Cal sensor Temp.: 94.72 Ambient Temp.: 85.69 F
23:07:40.190 -> Average Cal sensor Temp.: 95.11 Ambient Temp.: 85.69 F
23:07:41.718 -> Average Cal sensor Temp.: 95.19 Ambient Temp.: 85.59 F
23:07:43.279 -> Average Cal sensor Temp.: 95.26 Ambient Temp.: 85.66 F
23:07:44.841 -> Average Cal sensor Temp.: 94.80 Ambient Temp.: 85.66 F
23:07:46.399 -> Average Cal sensor Temp.: 95.04 Ambient Temp.: 85.73 F
23:07:47.957 -> Average Cal sensor Temp.: 94.67 Ambient Temp.: 85.66 F
23:07:49.520 -> Average Cal sensor Temp.: 94.86 Ambient Temp.: 85.69 F
23:07:51.078 -> Average Cal sensor Temp.: 94.83 Ambient Temp.: 85.66 F
23:07:52.638 -> Average Cal sensor Temp.: 94.92 Ambient Temp.: 85.59 F
23:07:54.196 -> Average Cal sensor Temp.: 95.15 Ambient Temp.: 85.59 F
23:07:55.751 -> Average Cal sensor Temp.: 95.23 Ambient Temp.: 85.66 F
23:07:57.308 -> Average Cal sensor Temp.: 95.23 Ambient Temp.: 85.69 F
23:07:58.862 -> Average Cal sensor Temp.: 95.04 Ambient Temp.: 85.69 F
23:08:00.420 -> Average Cal sensor Temp.: 95.16 Ambient Temp.: 85.69 F
23:08:01.972 -> Average Cal sensor Temp.: 95.04 Ambient Temp.: 85.69 F
23:08:03.524 -> Average Cal sensor Temp.: 95.00 Ambient Temp.: 85.73 F
23:08:05.085 -> Average Cal sensor Temp.: 95.22 Ambient Temp.: 85.69 F
23:08:06.614 -> Average Cal sensor Temp.: 95.08 Ambient Temp.: 85.66 F
23:08:08.203 -> Average Cal sensor Temp.: 95.15 Ambient Temp.: 85.66 F
23:08:09.757 -> Average Cal sensor Temp.: 94.83 Ambient Temp.: 85.66 F
23:08:11.313 -> Average Cal sensor Temp.: 94.45 Ambient Temp.: 85.59 F
23:08:12.864 -> Average Cal sensor Temp.: 94.64 Ambient Temp.: 85.59 F
23:08:14.425 -> Average Cal sensor Temp.: 94.58 Ambient Temp.: 85.59 F
23:08:15.982 -> Average Cal sensor Temp.: 94.84 Ambient Temp.: 85.55 F
23:08:17.530 -> Average Cal sensor Temp.: 95.19 Ambient Temp.: 85.59 F
23:08:19.094 -> Average Cal sensor Temp.: 95.40 Ambient Temp.: 85.66 F
23:08:20.653 -> Average Cal sensor Temp.: 95.10 Ambient Temp.: 85.59 F
23:08:22.181 -> Average Cal sensor Temp.: 94.77 Ambient Temp.: 85.59 F
23:08:23.735 -> Average Cal sensor Temp.: 94.47 Ambient Temp.: 85.55 F
23:08:25.293 -> Average Cal sensor Temp.: 94.84 Ambient Temp.: 85.55 F
23:08:26.855 -> Average Cal sensor Temp.: 94.62 Ambient Temp.: 85.51 F
23:08:28.417 -> Average Cal sensor Temp.: 95.02 Ambient Temp.: 85.44 F
23:08:29.974 -> Average Cal sensor Temp.: 94.99 Ambient Temp.: 85.44 F
23:08:31.523 -> Average Cal sensor Temp.: 95.07 Ambient Temp.: 85.48 F
23:08:33.075 -> Average Cal sensor Temp.: 94.82 Ambient Temp.: 85.44 F
23:08:34.636 -> Average Cal sensor Temp.: 94.95 Ambient Temp.: 85.44 F
23:08:36.194 -> Average Cal sensor Temp.: 95.15 Ambient Temp.: 85.44 F
23:08:37.743 -> Average Cal sensor Temp.: 95.27 Ambient Temp.: 85.44 F
23:08:39.299 -> Average Cal sensor Temp.: 95.11 Ambient Temp.: 85.41 F
23:08:40.850 -> Average Cal sensor Temp.: 95.05 Ambient Temp.: 85.41 F
23:08:42.405 -> Average Cal sensor Temp.: 95.09 Ambient Temp.: 85.41 F
23:08:43.966 -> Average Cal sensor Temp.: 94.93 Ambient Temp.: 85.37 F
23:08:45.519 -> Average Cal sensor Temp.: 94.96 Ambient Temp.: 85.33 F
23:08:47.074 -> Average Cal sensor Temp.: 95.05 Ambient Temp.: 85.33 F
23:08:48.626 -> Average Cal sensor Temp.: 94.58 Ambient Temp.: 85.30 F
23:08:50.184 -> Average Cal sensor Temp.: 94.79 Ambient Temp.: 85.33 F
23:08:51.771 -> Average Cal sensor Temp.: 95.21 Ambient Temp.: 85.30 F
23:08:53.296 -> Average Cal sensor Temp.: 95.05 Ambient Temp.: 85.30 F
23:08:54.858 -> Average Cal sensor Temp.: 95.05 Ambient Temp.: 85.26 F
23:08:56.412 -> Average Cal sensor Temp.: 95.23 Ambient Temp.: 85.30 F
23:08:57.973 -> Average Cal sensor Temp.: 95.00 Ambient Temp.: 85.26 F
23:08:59.525 -> Average Cal sensor Temp.: 95.01 Ambient Temp.: 85.19 F
23:09:01.083 -> Average Cal sensor Temp.: 95.40 Ambient Temp.: 85.15 F
23:09:02.640 -> Average Cal sensor Temp.: 95.42 Ambient Temp.: 85.19 F
23:09:04.196 -> Average Cal sensor Temp.: 94.80 Ambient Temp.: 85.12 F
23:09:05.753 -> Average Cal sensor Temp.: 94.78 Ambient Temp.: 85.08 F
23:09:07.314 -> Average Cal sensor Temp.: 94.80 Ambient Temp.: 85.12 F
23:09:08.877 -> Average Cal sensor Temp.: 94.42 Ambient Temp.: 85.08 F
23:09:10.427 -> Average Cal sensor Temp.: 94.49 Ambient Temp.: 85.01 F
23:09:11.984 -> Average Cal sensor Temp.: 94.53 Ambient Temp.: 84.94 F
23:09:13.537 -> Average Cal sensor Temp.: 95.00 Ambient Temp.: 84.97 F
23:09:15.095 -> Average Cal sensor Temp.: 94.85 Ambient Temp.: 84.97 F
23:09:16.651 -> Average Cal sensor Temp.: 94.96 Ambient Temp.: 84.94 F
23:09:18.212 -> Average Cal sensor Temp.: 94.93 Ambient Temp.: 84.94 F
23:09:19.772 -> Average Cal sensor Temp.: 94.89 Ambient Temp.: 84.94 F
23:09:21.330 -> Average Cal sensor Temp.: 94.96 Ambient Temp.: 84.87 F
23:09:22.892 -> Average Cal sensor Temp.: 95.23 Ambient Temp.: 84.94 F
23:09:24.452 -> Average Cal sensor Temp.: 94.89 Ambient Temp.: 84.87 F
23:09:26.014 -> Average Cal sensor Temp.: 95.29 Ambient Temp.: 84.97 F
23:09:27.571 -> Average Cal sensor Temp.: 94.96 Ambient Temp.: 84.87 F
23:09:29.132 -> Average Cal sensor Temp.: 94.64 Ambient Temp.: 84.79 F
23:09:30.686 -> Average Cal sensor Temp.: 94.72 Ambient Temp.: 84.76 F
23:09:32.242 -> Average Cal sensor Temp.: 94.79 Ambient Temp.: 84.72 F
23:09:33.803 -> Average Cal sensor Temp.: 94.71 Ambient Temp.: 84.65 F
23:09:35.327 -> Average Cal sensor Temp.: 95.17 Ambient Temp.: 84.69 F
23:09:36.883 -> Average Cal sensor Temp.: 94.89 Ambient Temp.: 84.65 F
23:09:38.448 -> Average Cal sensor Temp.: 95.12 Ambient Temp.: 84.61 F
23:09:40.007 -> Average Cal sensor Temp.: 95.18 Ambient Temp.: 84.61 F
23:09:41.570 -> Average Cal sensor Temp.: 94.90 Ambient Temp.: 84.61 F
23:09:43.131 -> Average Cal sensor Temp.: 94.92 Ambient Temp.: 84.61 F
23:09:44.689 -> Average Cal sensor Temp.: 94.55 Ambient Temp.: 84.51 F
23:09:46.244 -> Average Cal sensor Temp.: 94.67 Ambient Temp.: 84.47 F
23:09:47.796 -> Average Cal sensor Temp.: 94.67 Ambient Temp.: 84.47 F
23:09:49.358 -> Average Cal sensor Temp.: 94.67 Ambient Temp.: 84.36 F
23:09:50.913 -> Average Cal sensor Temp.: 94.78 Ambient Temp.: 84.36 F
23:09:52.473 -> Average Cal sensor Temp.: 95.14 Ambient Temp.: 84.33 F
23:09:54.000 -> Average Cal sensor Temp.: 95.17 Ambient Temp.: 84.33 F
23:09:55.555 -> Average Cal sensor Temp.: 95.05 Ambient Temp.: 84.33 F
23:09:57.112 -> Average Cal sensor Temp.: 94.92 Ambient Temp.: 84.25 F
23:09:58.671 -> Average Cal sensor Temp.: 94.63 Ambient Temp.: 84.22 F
23:10:00.227 -> Average Cal sensor Temp.: 94.88 Ambient Temp.: 84.15 F
23:10:01.782 -> Average Cal sensor Temp.: 95.09 Ambient Temp.: 84.07 F
23:10:03.368 -> Average Cal sensor Temp.: 94.86 Ambient Temp.: 84.07 F
23:10:04.925 -> Average Cal sensor Temp.: 95.07 Ambient Temp.: 84.11 F
23:10:06.452 -> Average Cal sensor Temp.: 94.77 Ambient Temp.: 84.07 F
23:10:08.037 -> Average Cal sensor Temp.: 94.73 Ambient Temp.: 84.07 F
23:10:09.591 -> Average Cal sensor Temp.: 94.77 Ambient Temp.: 83.97 F
23:10:11.145 -> Average Cal sensor Temp.: 94.91 Ambient Temp.: 83.93 F
23:10:12.706 -> Average Cal sensor Temp.: 95.26 Ambient Temp.: 83.97 F
23:10:14.230 -> Average Cal sensor Temp.: 95.14 Ambient Temp.: 83.93 F
23:10:15.789 -> Average Cal sensor Temp.: 95.18 Ambient Temp.: 83.93 F
23:10:17.347 -> Average Cal sensor Temp.: 95.09 Ambient Temp.: 83.93 F
23:10:18.907 -> Average Cal sensor Temp.: 95.33 Ambient Temp.: 83.86 F
23:10:20.460 -> Average Cal sensor Temp.: 94.81 Ambient Temp.: 83.86 F
23:10:22.014 -> Average Cal sensor Temp.: 94.79 Ambient Temp.: 83.86 F
23:10:23.569 -> Average Cal sensor Temp.: 95.06 Ambient Temp.: 83.79 F
23:10:25.126 -> Average Cal sensor Temp.: 95.29 Ambient Temp.: 83.75 F
23:10:26.682 -> Average Cal sensor Temp.: 94.99 Ambient Temp.: 83.79 F
23:10:28.243 -> Average Cal sensor Temp.: 95.03 Ambient Temp.: 83.71 F
23:10:29.800 -> Average Cal sensor Temp.: 95.28 Ambient Temp.: 83.75 F
23:10:31.354 -> Average Cal sensor Temp.: 94.83 Ambient Temp.: 83.75 F
23:10:32.916 -> Average Cal sensor Temp.: 95.08 Ambient Temp.: 83.71 F
23:10:34.473 -> Average Cal sensor Temp.: 95.05 Ambient Temp.: 83.71 F
23:10:36.029 -> Average Cal sensor Temp.: 94.97 Ambient Temp.: 83.64 F
23:10:37.588 -> Average Cal sensor Temp.: 95.00 Ambient Temp.: 83.64 F
23:10:39.147 -> Average Cal sensor Temp.: 95.01 Ambient Temp.: 83.64 F
23:10:40.706 -> Average Cal sensor Temp.: 95.25 Ambient Temp.: 83.57 F
23:10:42.259 -> Average Cal sensor Temp.: 95.14 Ambient Temp.: 83.53 F
23:10:43.816 -> Average Cal sensor Temp.: 95.00 Ambient Temp.: 83.50 F
23:10:45.377 -> Average Cal sensor Temp.: 95.19 Ambient Temp.: 83.50 F
23:10:46.936 -> Average Cal sensor Temp.: 95.06 Ambient Temp.: 83.46 F
23:10:48.459 -> Average Cal sensor Temp.: 94.78 Ambient Temp.: 83.53 F
23:10:50.021 -> Average Cal sensor Temp.: 95.20 Ambient Temp.: 83.46 F
23:10:51.580 -> Average Cal sensor Temp.: 94.85 Ambient Temp.: 83.46 F
23:10:53.139 -> Average Cal sensor Temp.: 95.08 Ambient Temp.: 83.46 F
23:10:54.699 -> Average Cal sensor Temp.: 95.17 Ambient Temp.: 83.46 F
23:10:56.257 -> Average Cal sensor Temp.: 95.02 Ambient Temp.: 83.39 F
23:10:57.815 -> Average Cal sensor Temp.: 95.08 Ambient Temp.: 83.35 F
23:10:59.377 -> Average Cal sensor Temp.: 95.29 Ambient Temp.: 83.43 F
23:11:00.933 -> Average Cal sensor Temp.: 95.26 Ambient Temp.: 83.43 F
23:11:02.493 -> Average Cal sensor Temp.: 95.35 Ambient Temp.: 83.43 F
23:11:04.019 -> Average Cal sensor Temp.: 95.20 Ambient Temp.: 83.39 F
23:11:05.575 -> Average Cal sensor Temp.: 95.09 Ambient Temp.: 83.39 F
23:11:07.136 -> Average Cal sensor Temp.: 95.22 Ambient Temp.: 83.39 F
23:11:08.697 -> Average Cal sensor Temp.: 95.22 Ambient Temp.: 83.43 F
23:11:10.257 -> Average Cal sensor Temp.: 95.23 Ambient Temp.: 83.39 F
23:11:11.810 -> Average Cal sensor Temp.: 95.05 Ambient Temp.: 83.39 F
23:11:13.370 -> Average Cal sensor Temp.: 95.19 Ambient Temp.: 83.39 F
23:11:14.929 -> Average Cal sensor Temp.: 95.21 Ambient Temp.: 83.39 F
23:11:16.485 -> Average Cal sensor Temp.: 95.45 Ambient Temp.: 83.39 F
23:11:18.044 -> Average Cal sensor Temp.: 95.53 Ambient Temp.: 83.43 F
23:11:19.601 -> Average Cal sensor Temp.: 95.40 Ambient Temp.: 83.43 F
23:11:21.160 -> Average Cal sensor Temp.: 95.30 Ambient Temp.: 83.46 F
23:11:22.712 -> Average Cal sensor Temp.: 95.27 Ambient Temp.: 83.43 F
23:11:24.273 -> Average Cal sensor Temp.: 95.36 Ambient Temp.: 83.43 F
23:11:25.799 -> Average Cal sensor Temp.: 95.39 Ambient Temp.: 83.43 F
23:11:27.356 -> Average Cal sensor Temp.: 95.68 Ambient Temp.: 83.50 F
23:11:28.915 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 83.46 F
23:11:30.477 -> Average Cal sensor Temp.: 95.14 Ambient Temp.: 83.50 F
23:11:32.034 -> Average Cal sensor Temp.: 95.46 Ambient Temp.: 83.50 F
23:11:33.596 -> Average Cal sensor Temp.: 95.40 Ambient Temp.: 83.46 F
23:11:35.155 -> Average Cal sensor Temp.: 95.42 Ambient Temp.: 83.53 F
23:11:36.718 -> Average Cal sensor Temp.: 95.67 Ambient Temp.: 83.50 F
23:11:38.245 -> Average Cal sensor Temp.: 95.39 Ambient Temp.: 83.50 F
23:11:39.810 -> Average Cal sensor Temp.: 95.36 Ambient Temp.: 83.57 F
23:11:41.369 -> Average Cal sensor Temp.: 95.64 Ambient Temp.: 83.64 F
23:11:42.927 -> Average Cal sensor Temp.: 95.34 Ambient Temp.: 83.57 F
23:11:44.484 -> Average Cal sensor Temp.: 95.45 Ambient Temp.: 83.68 F
23:11:46.038 -> Average Cal sensor Temp.: 95.54 Ambient Temp.: 83.68 F
23:11:47.595 -> Average Cal sensor Temp.: 95.65 Ambient Temp.: 83.68 F
23:11:49.152 -> Average Cal sensor Temp.: 95.51 Ambient Temp.: 83.71 F
23:11:50.712 -> Average Cal sensor Temp.: 95.72 Ambient Temp.: 83.71 F
23:11:52.262 -> Average Cal sensor Temp.: 95.79 Ambient Temp.: 83.75 F
23:11:53.816 -> Average Cal sensor Temp.: 95.60 Ambient Temp.: 83.79 F
23:11:55.373 -> Average Cal sensor Temp.: 95.85 Ambient Temp.: 83.75 F
23:11:56.928 -> Average Cal sensor Temp.: 96.00 Ambient Temp.: 83.82 F
23:11:58.485 -> Average Cal sensor Temp.: 95.78 Ambient Temp.: 83.82 F
23:12:00.045 -> Average Cal sensor Temp.: 95.64 Ambient Temp.: 83.86 F
23:12:01.605 -> Average Cal sensor Temp.: 95.79 Ambient Temp.: 83.86 F
23:12:03.160 -> Average Cal sensor Temp.: 95.66 Ambient Temp.: 83.86 F
23:12:04.723 -> Average Cal sensor Temp.: 95.62 Ambient Temp.: 83.93 F
23:12:06.281 -> Average Cal sensor Temp.: 95.54 Ambient Temp.: 83.93 F
23:12:07.838 -> Average Cal sensor Temp.: 95.69 Ambient Temp.: 83.97 F
23:12:09.363 -> Average Cal sensor Temp.: 95.39 Ambient Temp.: 84.00 F
23:12:10.924 -> Average Cal sensor Temp.: 95.62 Ambient Temp.: 84.07 F
23:12:12.479 -> Average Cal sensor Temp.: 95.78 Ambient Temp.: 84.00 F
23:12:14.040 -> Average Cal sensor Temp.: 95.76 Ambient Temp.: 84.07 F
23:12:15.594 -> Average Cal sensor Temp.: 95.71 Ambient Temp.: 84.11 F
23:12:17.151 -> Average Cal sensor Temp.: 95.48 Ambient Temp.: 84.15 F
23:12:18.710 -> Average Cal sensor Temp.: 95.52 Ambient Temp.: 84.11 F
23:12:20.265 -> Average Cal sensor Temp.: 95.60 Ambient Temp.: 84.15 F
23:12:21.820 -> Average Cal sensor Temp.: 95.83 Ambient Temp.: 84.22 F
23:12:23.377 -> Average Cal sensor Temp.: 95.83 Ambient Temp.: 84.18 F
23:12:24.936 -> Average Cal sensor Temp.: 95.33 Ambient Temp.: 84.22 F
23:12:26.496 -> Average Cal sensor Temp.: 95.72 Ambient Temp.: 84.22 F
23:12:28.056 -> Average Cal sensor Temp.: 95.55 Ambient Temp.: 84.22 F
23:12:29.618 -> Average Cal sensor Temp.: 95.72 Ambient Temp.: 84.33 F
23:12:31.174 -> Average Cal sensor Temp.: 95.50 Ambient Temp.: 84.25 F
23:12:32.700 -> Average Cal sensor Temp.: 95.53 Ambient Temp.: 84.29 F
23:12:34.264 -> Average Cal sensor Temp.: 95.67 Ambient Temp.: 84.33 F
23:12:35.824 -> Average Cal sensor Temp.: 95.91 Ambient Temp.: 84.36 F
23:12:37.381 -> Average Cal sensor Temp.: 95.85 Ambient Temp.: 84.40 F
23:12:38.938 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 84.36 F
23:12:40.501 -> Average Cal sensor Temp.: 95.75 Ambient Temp.: 84.47 F
23:12:42.056 -> Average Cal sensor Temp.: 95.83 Ambient Temp.: 84.51 F
23:12:43.607 -> Average Cal sensor Temp.: 95.82 Ambient Temp.: 84.54 F
23:12:45.166 -> Average Cal sensor Temp.: 95.74 Ambient Temp.: 84.58 F
23:12:46.724 -> Average Cal sensor Temp.: 95.25 Ambient Temp.: 84.58 F
23:12:48.285 -> Average Cal sensor Temp.: 95.53 Ambient Temp.: 84.58 F
23:12:49.832 -> Average Cal sensor Temp.: 95.41 Ambient Temp.: 84.58 F
23:12:51.387 -> Average Cal sensor Temp.: 95.78 Ambient Temp.: 84.65 F
23:12:52.943 -> Average Cal sensor Temp.: 95.80 Ambient Temp.: 84.61 F
23:12:54.500 -> Average Cal sensor Temp.: 95.39 Ambient Temp.: 84.61 F
23:12:56.059 -> Average Cal sensor Temp.: 95.39 Ambient Temp.: 84.69 F
23:12:57.615 -> Average Cal sensor Temp.: 95.39 Ambient Temp.: 84.69 F
23:12:59.179 -> Average Cal sensor Temp.: 95.65 Ambient Temp.: 84.65 F
23:13:00.740 -> Average Cal sensor Temp.: 95.75 Ambient Temp.: 84.72 F
23:13:02.265 -> Average Cal sensor Temp.: 95.88 Ambient Temp.: 84.72 F
23:13:03.825 -> Average Cal sensor Temp.: 95.77 Ambient Temp.: 84.79 F
23:13:05.386 -> Average Cal sensor Temp.: 95.92 Ambient Temp.: 84.76 F
23:13:06.937 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 84.79 F
23:13:08.492 -> Average Cal sensor Temp.: 95.86 Ambient Temp.: 84.90 F
23:13:10.051 -> Average Cal sensor Temp.: 95.94 Ambient Temp.: 84.79 F
23:13:11.610 -> Average Cal sensor Temp.: 95.96 Ambient Temp.: 84.90 F
23:13:13.174 -> Average Cal sensor Temp.: 96.10 Ambient Temp.: 84.97 F
23:13:14.731 -> Average Cal sensor Temp.: 96.02 Ambient Temp.: 85.01 F
23:13:16.289 -> Average Cal sensor Temp.: 95.84 Ambient Temp.: 84.97 F
23:13:17.847 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 84.97 F
23:13:19.408 -> Average Cal sensor Temp.: 95.69 Ambient Temp.: 85.01 F
23:13:20.970 -> Average Cal sensor Temp.: 95.62 Ambient Temp.: 85.12 F
23:13:22.521 -> Average Cal sensor Temp.: 95.62 Ambient Temp.: 85.08 F
23:13:24.076 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 85.05 F
23:13:25.637 -> Average Cal sensor Temp.: 95.75 Ambient Temp.: 85.12 F
23:13:27.192 -> Average Cal sensor Temp.: 95.82 Ambient Temp.: 85.15 F
23:13:28.718 -> Average Cal sensor Temp.: 95.98 Ambient Temp.: 85.12 F
23:13:30.281 -> Average Cal sensor Temp.: 95.80 Ambient Temp.: 85.26 F
23:13:31.841 -> Average Cal sensor Temp.: 95.76 Ambient Temp.: 85.19 F
23:13:33.401 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 85.19 F
23:13:34.961 -> Average Cal sensor Temp.: 95.86 Ambient Temp.: 85.26 F
23:13:36.518 -> Average Cal sensor Temp.: 95.93 Ambient Temp.: 85.30 F
23:13:38.078 -> Average Cal sensor Temp.: 95.75 Ambient Temp.: 85.30 F
23:13:39.642 -> Average Cal sensor Temp.: 95.67 Ambient Temp.: 85.26 F
23:13:41.164 -> Average Cal sensor Temp.: 95.92 Ambient Temp.: 85.33 F
23:13:42.726 -> Average Cal sensor Temp.: 95.87 Ambient Temp.: 85.33 F
23:13:44.281 -> Average Cal sensor Temp.: 95.86 Ambient Temp.: 85.41 F
23:13:45.837 -> Average Cal sensor Temp.: 95.78 Ambient Temp.: 85.37 F
23:13:47.395 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 85.41 F
23:13:48.960 -> Average Cal sensor Temp.: 95.51 Ambient Temp.: 85.37 F
23:13:50.517 -> Average Cal sensor Temp.: 95.55 Ambient Temp.: 85.48 F
23:13:52.074 -> Average Cal sensor Temp.: 95.62 Ambient Temp.: 85.48 F
23:13:53.633 -> Average Cal sensor Temp.: 95.61 Ambient Temp.: 85.44 F
23:13:55.162 -> Average Cal sensor Temp.: 95.51 Ambient Temp.: 85.44 F
23:13:56.719 -> Average Cal sensor Temp.: 95.59 Ambient Temp.: 85.44 F
23:13:58.281 -> Average Cal sensor Temp.: 95.71 Ambient Temp.: 85.48 F
23:13:59.840 -> Average Cal sensor Temp.: 95.76 Ambient Temp.: 85.48 F
23:14:01.399 -> Average Cal sensor Temp.: 96.02 Ambient Temp.: 85.51 F
23:14:02.959 -> Average Cal sensor Temp.: 95.69 Ambient Temp.: 85.51 F
23:14:04.517 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 85.51 F
23:14:06.075 -> Average Cal sensor Temp.: 95.97 Ambient Temp.: 85.55 F
23:14:07.630 -> Average Cal sensor Temp.: 95.77 Ambient Temp.: 85.59 F
23:14:09.188 -> Average Cal sensor Temp.: 95.57 Ambient Temp.: 85.51 F
23:14:10.744 -> Average Cal sensor Temp.: 95.56 Ambient Temp.: 85.55 F
23:14:12.268 -> Average Cal sensor Temp.: 95.73 Ambient Temp.: 85.55 F
23:14:13.831 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 85.66 F
23:14:15.395 -> Average Cal sensor Temp.: 96.03 Ambient Temp.: 85.59 F
23:14:16.955 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 85.69 F
23:14:18.505 -> Average Cal sensor Temp.: 96.09 Ambient Temp.: 85.66 F
23:14:20.063 -> Average Cal sensor Temp.: 95.86 Ambient Temp.: 85.59 F
23:14:21.618 -> Average Cal sensor Temp.: 96.00 Ambient Temp.: 85.69 F
23:14:23.175 -> Average Cal sensor Temp.: 95.66 Ambient Temp.: 85.73 F
23:14:24.730 -> Average Cal sensor Temp.: 95.64 Ambient Temp.: 85.73 F
23:14:26.284 -> Average Cal sensor Temp.: 95.82 Ambient Temp.: 85.73 F
23:14:27.838 -> Average Cal sensor Temp.: 95.79 Ambient Temp.: 85.73 F
23:14:29.396 -> Average Cal sensor Temp.: 95.73 Ambient Temp.: 85.73 F
23:14:30.951 -> Average Cal sensor Temp.: 96.05 Ambient Temp.: 85.80 F
23:14:32.509 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 85.80 F
23:14:34.067 -> Average Cal sensor Temp.: 95.99 Ambient Temp.: 85.80 F
23:14:35.624 -> Average Cal sensor Temp.: 96.13 Ambient Temp.: 85.77 F
23:14:37.183 -> Average Cal sensor Temp.: 96.14 Ambient Temp.: 85.77 F
23:14:38.751 -> Average Cal sensor Temp.: 95.95 Ambient Temp.: 85.80 F
23:14:40.310 -> Average Cal sensor Temp.: 95.86 Ambient Temp.: 85.80 F
23:14:41.867 -> Average Cal sensor Temp.: 95.84 Ambient Temp.: 85.84 F
23:14:43.409 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 85.84 F
23:14:44.965 -> Average Cal sensor Temp.: 95.84 Ambient Temp.: 85.84 F
23:14:46.527 -> Average Cal sensor Temp.: 95.64 Ambient Temp.: 85.84 F
23:14:48.088 -> Average Cal sensor Temp.: 95.68 Ambient Temp.: 85.80 F
23:14:49.647 -> Average Cal sensor Temp.: 95.77 Ambient Temp.: 85.80 F
23:14:51.208 -> Average Cal sensor Temp.: 95.53 Ambient Temp.: 85.73 F
23:14:52.764 -> Average Cal sensor Temp.: 95.66 Ambient Temp.: 85.77 F
23:14:54.322 -> Average Cal sensor Temp.: 95.50 Ambient Temp.: 85.66 F
23:14:55.847 -> Average Cal sensor Temp.: 95.58 Ambient Temp.: 85.77 F
23:14:57.434 -> Average Cal sensor Temp.: 95.57 Ambient Temp.: 85.73 F
23:14:58.958 -> Average Cal sensor Temp.: 95.78 Ambient Temp.: 85.69 F
23:15:00.521 -> Average Cal sensor Temp.: 95.58 Ambient Temp.: 85.69 F
23:15:02.083 -> Average Cal sensor Temp.: 95.79 Ambient Temp.: 85.73 F
23:15:03.647 -> Average Cal sensor Temp.: 95.77 Ambient Temp.: 85.73 F
23:15:05.206 -> Average Cal sensor Temp.: 95.67 Ambient Temp.: 85.69 F
23:15:06.764 -> Average Cal sensor Temp.: 95.53 Ambient Temp.: 85.66 F
23:15:08.323 -> Average Cal sensor Temp.: 95.61 Ambient Temp.: 85.66 F
23:15:09.846 -> Average Cal sensor Temp.: 95.69 Ambient Temp.: 85.66 F
23:15:11.404 -> Average Cal sensor Temp.: 95.71 Ambient Temp.: 85.69 F
23:15:12.990 -> Average Cal sensor Temp.: 95.56 Ambient Temp.: 85.59 F
23:15:14.516 -> Average Cal sensor Temp.: 95.45 Ambient Temp.: 85.55 F
23:15:16.076 -> Average Cal sensor Temp.: 95.66 Ambient Temp.: 85.59 F
23:15:17.630 -> Average Cal sensor Temp.: 95.64 Ambient Temp.: 85.51 F
23:15:19.188 -> Average Cal sensor Temp.: 95.65 Ambient Temp.: 85.48 F
23:15:20.750 -> Average Cal sensor Temp.: 95.59 Ambient Temp.: 85.48 F
23:15:22.309 -> Average Cal sensor Temp.: 95.41 Ambient Temp.: 85.44 F
23:15:23.863 -> Average Cal sensor Temp.: 95.57 Ambient Temp.: 85.44 F
23:15:25.419 -> Average Cal sensor Temp.: 95.46 Ambient Temp.: 85.41 F
23:15:26.975 -> Average Cal sensor Temp.: 95.64 Ambient Temp.: 85.41 F
23:15:28.537 -> Average Cal sensor Temp.: 95.38 Ambient Temp.: 85.41 F
23:15:30.099 -> Average Cal sensor Temp.: 95.67 Ambient Temp.: 85.30 F
23:15:31.656 -> Average Cal sensor Temp.: 95.53 Ambient Temp.: 85.33 F
23:15:33.185 -> Average Cal sensor Temp.: 95.56 Ambient Temp.: 85.26 F
23:15:34.737 -> Average Cal sensor Temp.: 95.56 Ambient Temp.: 85.30 F
23:15:36.297 -> Average Cal sensor Temp.: 95.25 Ambient Temp.: 85.19 F
23:15:37.858 -> Average Cal sensor Temp.: 95.36 Ambient Temp.: 85.26 F
23:15:39.419 -> Average Cal sensor Temp.: 95.57 Ambient Temp.: 85.19 F
23:15:40.972 -> Average Cal sensor Temp.: 95.69 Ambient Temp.: 85.19 F
23:15:42.530 -> Average Cal sensor Temp.: 95.60 Ambient Temp.: 85.19 F
23:15:44.091 -> Average Cal sensor Temp.: 95.50 Ambient Temp.: 85.15 F
23:15:45.645 -> Average Cal sensor Temp.: 95.56 Ambient Temp.: 85.12 F
23:15:47.203 -> Average Cal sensor Temp.: 95.21 Ambient Temp.: 85.05 F
23:15:48.761 -> Average Cal sensor Temp.: 95.61 Ambient Temp.: 85.05 F
23:15:50.317 -> Average Cal sensor Temp.: 95.46 Ambient Temp.: 85.05 F
23:15:51.873 -> Average Cal sensor Temp.: 95.67 Ambient Temp.: 85.01 F
23:15:53.432 -> Average Cal sensor Temp.: 95.61 Ambient Temp.: 85.01 F
23:15:54.961 -> Average Cal sensor Temp.: 95.55 Ambient Temp.: 84.94 F
23:15:56.523 -> Average Cal sensor Temp.: 95.48 Ambient Temp.: 84.94 F
23:15:58.080 -> Average Cal sensor Temp.: 95.59 Ambient Temp.: 84.90 F
23:15:59.642 -> Average Cal sensor Temp.: 95.37 Ambient Temp.: 84.87 F
23:16:01.196 -> Average Cal sensor Temp.: 95.62 Ambient Temp.: 84.87 F
23:16:02.751 -> Average Cal sensor Temp.: 95.24 Ambient Temp.: 84.87 F
23:16:04.307 -> Average Cal sensor Temp.: 95.31 Ambient Temp.: 84.72 F
23:16:05.871 -> Average Cal sensor Temp.: 95.11 Ambient Temp.: 84.69 F
23:16:07.426 -> Average Cal sensor Temp.: 95.51 Ambient Temp.: 84.72 F
23:16:08.985 -> Average Cal sensor Temp.: 95.58 Ambient Temp.: 84.69 F
23:16:10.542 -> Average Cal sensor Temp.: 95.74 Ambient Temp.: 84.72 F
23:16:12.098 -> Average Cal sensor Temp.: 95.44 Ambient Temp.: 84.65 F
23:16:13.659 -> Average Cal sensor Temp.: 95.58 Ambient Temp.: 84.65 F
23:16:15.184 -> Average Cal sensor Temp.: 95.47 Ambient Temp.: 84.65 F
23:16:16.745 -> Average Cal sensor Temp.: 95.54 Ambient Temp.: 84.61 F
23:16:18.305 -> Average Cal sensor Temp.: 95.44 Ambient Temp.: 84.61 F
23:16:19.868 -> Average Cal sensor Temp.: 95.49 Ambient Temp.: 84.58 F
23:16:21.423 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 84.58 F
23:16:22.976 -> Average Cal sensor Temp.: 95.68 Ambient Temp.: 84.51 F
23:16:24.535 -> Average Cal sensor Temp.: 95.70 Ambient Temp.: 84.54 F
23:16:26.089 -> Average Cal sensor Temp.: 95.75 Ambient Temp.: 84.40 F
23:16:27.648 -> Average Cal sensor Temp.: 95.96 Ambient Temp.: 84.54 F
23:16:29.201 -> Average Cal sensor Temp.: 96.11 Ambient Temp.: 84.54 F
23:16:30.754 -> Average Cal sensor Temp.: 96.02 Ambient Temp.: 84.54 F
23:16:32.305 -> Average Cal sensor Temp.: 95.72 Ambient Temp.: 84.54 F
23:16:33.868 -> Average Cal sensor Temp.: 95.73 Ambient Temp.: 84.51 F
23:16:35.428 -> Average Cal sensor Temp.: 95.62 Ambient Temp.: 84.51 F
23:16:36.991 -> Average Cal sensor Temp.: 95.87 Ambient Temp.: 84.47 F
23:16:38.523 -> Average Cal sensor Temp.: 95.68 Ambient Temp.: 84.47 F
23:16:40.084 -> Average Cal sensor Temp.: 95.24 Ambient Temp.: 84.47 F
23:16:41.644 -> Average Cal sensor Temp.: 95.60 Ambient Temp.: 84.40 F
23:16:43.199 -> Average Cal sensor Temp.: 95.85 Ambient Temp.: 84.36 F
23:16:44.761 -> Average Cal sensor Temp.: 95.87 Ambient Temp.: 84.40 F
23:16:46.319 -> Average Cal sensor Temp.: 95.84 Ambient Temp.: 84.40 F
23:16:47.880 -> Average Cal sensor Temp.: 95.80 Ambient Temp.: 84.40 F
23:16:49.437 -> Average Cal sensor Temp.: 95.88 Ambient Temp.: 84.40 F
23:16:50.993 -> Average Cal sensor Temp.: 95.65 Ambient Temp.: 84.47 F
23:16:52.554 -> Average Cal sensor Temp.: 95.62 Ambient Temp.: 84.47 F
23:16:54.111 -> Average Cal sensor Temp.: 95.27 Ambient Temp.: 84.40 F
23:16:55.661 -> Average Cal sensor Temp.: 95.56 Ambient Temp.: 84.36 F
23:16:57.222 -> Average Cal sensor Temp.: 95.52 Ambient Temp.: 84.36 F
23:16:58.787 -> Average Cal sensor Temp.: 95.79 Ambient Temp.: 84.40 F
23:17:00.340 -> Average Cal sensor Temp.: 95.52 Ambient Temp.: 84.36 F
23:17:01.902 -> Average Cal sensor Temp.: 95.69 Ambient Temp.: 84.40 F
23:17:03.459 -> Average Cal sensor Temp.: 95.28 Ambient Temp.: 84.29 F
23:17:05.010 -> Average Cal sensor Temp.: 95.39 Ambient Temp.: 84.33 F
23:17:06.555 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 84.33 F
23:17:08.117 -> Average Cal sensor Temp.: 95.36 Ambient Temp.: 84.33 F
23:17:09.674 -> Average Cal sensor Temp.: 96.04 Ambient Temp.: 84.33 F
23:17:11.233 -> Average Cal sensor Temp.: 95.66 Ambient Temp.: 84.33 F
23:17:12.790 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 84.25 F
23:17:14.351 -> Average Cal sensor Temp.: 95.96 Ambient Temp.: 84.29 F
23:17:15.906 -> Average Cal sensor Temp.: 95.91 Ambient Temp.: 84.29 F
23:17:17.464 -> Average Cal sensor Temp.: 96.03 Ambient Temp.: 84.29 F
23:17:19.024 -> Average Cal sensor Temp.: 95.73 Ambient Temp.: 84.29 F
23:17:20.549 -> Average Cal sensor Temp.: 95.70 Ambient Temp.: 84.29 F
23:17:22.137 -> Average Cal sensor Temp.: 95.82 Ambient Temp.: 84.22 F
23:17:23.689 -> Average Cal sensor Temp.: 95.87 Ambient Temp.: 84.29 F
23:17:25.247 -> Average Cal sensor Temp.: 95.97 Ambient Temp.: 84.25 F
23:17:26.774 -> Average Cal sensor Temp.: 95.94 Ambient Temp.: 84.22 F
23:17:28.363 -> Average Cal sensor Temp.: 95.68 Ambient Temp.: 84.22 F
23:17:29.890 -> Average Cal sensor Temp.: 95.83 Ambient Temp.: 84.25 F
23:17:31.453 -> Average Cal sensor Temp.: 95.87 Ambient Temp.: 84.25 F
23:17:33.014 -> Average Cal sensor Temp.: 95.84 Ambient Temp.: 84.22 F
23:17:34.571 -> Average Cal sensor Temp.: 96.06 Ambient Temp.: 84.25 F
23:17:36.131 -> Average Cal sensor Temp.: 95.94 Ambient Temp.: 84.25 F
23:17:37.688 -> Average Cal sensor Temp.: 95.80 Ambient Temp.: 84.22 F
23:17:39.247 -> Average Cal sensor Temp.: 95.74 Ambient Temp.: 84.25 F
23:17:40.804 -> Average Cal sensor Temp.: 96.15 Ambient Temp.: 84.25 F
23:17:42.367 -> Average Cal sensor Temp.: 96.05 Ambient Temp.: 84.33 F
23:17:43.898 -> Average Cal sensor Temp.: 95.68 Ambient Temp.: 84.29 F
23:17:45.453 -> Average Cal sensor Temp.: 95.62 Ambient Temp.: 84.29 F
23:17:47.014 -> Average Cal sensor Temp.: 95.90 Ambient Temp.: 84.29 F
23:17:48.571 -> Average Cal sensor Temp.: 95.61 Ambient Temp.: 84.29 F
23:17:50.129 -> Average Cal sensor Temp.: 96.01 Ambient Temp.: 84.29 F
23:17:51.686 -> Average Cal sensor Temp.: 95.85 Ambient Temp.: 84.36 F
23:17:53.241 -> Average Cal sensor Temp.: 95.79 Ambient Temp.: 84.29 F
23:17:54.802 -> Average Cal sensor Temp.: 95.90 Ambient Temp.: 84.33 F
23:17:56.356 -> Average Cal sensor Temp.: 96.06 Ambient Temp.: 84.36 F
23:17:57.913 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 84.40 F
23:17:59.475 -> Average Cal sensor Temp.: 95.80 Ambient Temp.: 84.40 F
23:18:01.038 -> Average Cal sensor Temp.: 96.10 Ambient Temp.: 84.40 F
23:18:02.603 -> Average Cal sensor Temp.: 96.05 Ambient Temp.: 84.47 F
23:18:04.161 -> Average Cal sensor Temp.: 96.16 Ambient Temp.: 84.47 F
23:18:05.718 -> Average Cal sensor Temp.: 96.23 Ambient Temp.: 84.51 F
23:18:07.271 -> Average Cal sensor Temp.: 95.87 Ambient Temp.: 84.47 F
23:18:08.829 -> Average Cal sensor Temp.: 95.87 Ambient Temp.: 84.47 F
23:18:10.358 -> Average Cal sensor Temp.: 95.94 Ambient Temp.: 84.51 F
23:18:11.918 -> Average Cal sensor Temp.: 95.97 Ambient Temp.: 84.58 F
23:18:13.477 -> Average Cal sensor Temp.: 95.89 Ambient Temp.: 84.54 F
23:18:15.039 -> Average Cal sensor Temp.: 95.65 Ambient Temp.: 84.58 F
23:18:16.596 -> Average Cal sensor Temp.: 95.86 Ambient Temp.: 84.61 F
23:18:18.156 -> Average Cal sensor Temp.: 95.75 Ambient Temp.: 84.58 F
23:18:19.719 -> Average Cal sensor Temp.: 96.02 Ambient Temp.: 84.61 F
23:18:21.282 -> Average Cal sensor Temp.: 96.23 Ambient Temp.: 84.65 F
23:18:22.837 -> Average Cal sensor Temp.: 96.13 Ambient Temp.: 84.65 F
23:18:24.363 -> Average Cal sensor Temp.: 95.93 Ambient Temp.: 84.65 F
23:18:25.920 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 84.65 F
23:18:27.484 -> Average Cal sensor Temp.: 96.24 Ambient Temp.: 84.69 F
23:18:29.042 -> Average Cal sensor Temp.: 95.99 Ambient Temp.: 84.65 F
23:18:30.601 -> Average Cal sensor Temp.: 96.16 Ambient Temp.: 84.69 F
23:18:32.162 -> Average Cal sensor Temp.: 95.96 Ambient Temp.: 84.69 F
23:18:33.723 -> Average Cal sensor Temp.: 96.06 Ambient Temp.: 84.76 F
23:18:35.287 -> Average Cal sensor Temp.: 96.04 Ambient Temp.: 84.72 F
23:18:36.845 -> Average Cal sensor Temp.: 96.03 Ambient Temp.: 84.76 F
23:18:38.398 -> Average Cal sensor Temp.: 95.98 Ambient Temp.: 84.79 F
23:18:39.957 -> Average Cal sensor Temp.: 95.88 Ambient Temp.: 84.72 F
23:18:41.516 -> Average Cal sensor Temp.: 95.87 Ambient Temp.: 84.79 F
23:18:43.074 -> Average Cal sensor Temp.: 95.94 Ambient Temp.: 84.87 F
23:18:44.632 -> Average Cal sensor Temp.: 96.00 Ambient Temp.: 84.87 F
23:18:46.194 -> Average Cal sensor Temp.: 95.94 Ambient Temp.: 84.87 F
23:18:47.751 -> Average Cal sensor Temp.: 95.96 Ambient Temp.: 84.87 F
23:18:49.309 -> Average Cal sensor Temp.: 96.02 Ambient Temp.: 84.94 F
23:18:50.867 -> Average Cal sensor Temp.: 95.97 Ambient Temp.: 84.94 F
23:18:52.429 -> Average Cal sensor Temp.: 95.89 Ambient Temp.: 85.01 F
23:18:53.958 -> Average Cal sensor Temp.: 95.88 Ambient Temp.: 84.97 F
23:18:55.519 -> Average Cal sensor Temp.: 96.07 Ambient Temp.: 85.01 F
23:18:57.080 -> Average Cal sensor Temp.: 96.04 Ambient Temp.: 85.01 F
23:18:58.645 -> Average Cal sensor Temp.: 96.19 Ambient Temp.: 85.01 F
23:19:00.204 -> Average Cal sensor Temp.: 96.05 Ambient Temp.: 85.01 F
23:19:01.763 -> Average Cal sensor Temp.: 96.11 Ambient Temp.: 85.05 F
23:19:03.312 -> Average Cal sensor Temp.: 96.11 Ambient Temp.: 85.12 F
23:19:04.863 -> Average Cal sensor Temp.: 95.94 Ambient Temp.: 85.12 F
23:19:06.402 -> Average Cal sensor Temp.: 96.11 Ambient Temp.: 85.12 F
23:19:07.954 -> Average Cal sensor Temp.: 95.99 Ambient Temp.: 85.12 F
23:19:09.520 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 85.26 F
23:19:11.082 -> Average Cal sensor Temp.: 95.90 Ambient Temp.: 85.15 F
23:19:12.643 -> Average Cal sensor Temp.: 96.11 Ambient Temp.: 85.19 F
23:19:14.205 -> Average Cal sensor Temp.: 95.89 Ambient Temp.: 85.26 F
23:19:15.734 -> Average Cal sensor Temp.: 95.86 Ambient Temp.: 85.19 F
23:19:17.321 -> Average Cal sensor Temp.: 95.83 Ambient Temp.: 85.26 F
23:19:18.852 -> Average Cal sensor Temp.: 95.74 Ambient Temp.: 85.30 F
23:19:20.415 -> Average Cal sensor Temp.: 95.84 Ambient Temp.: 85.33 F
23:19:21.980 -> Average Cal sensor Temp.: 96.00 Ambient Temp.: 85.33 F
23:19:23.533 -> Average Cal sensor Temp.: 95.91 Ambient Temp.: 85.37 F
23:19:25.094 -> Average Cal sensor Temp.: 95.83 Ambient Temp.: 85.41 F
23:19:26.657 -> Average Cal sensor Temp.: 95.90 Ambient Temp.: 85.37 F
23:19:28.211 -> Average Cal sensor Temp.: 95.88 Ambient Temp.: 85.41 F
23:19:29.735 -> Average Cal sensor Temp.: 95.94 Ambient Temp.: 85.48 F
23:19:31.317 -> Average Cal sensor Temp.: 96.23 Ambient Temp.: 85.48 F
23:19:32.875 -> Average Cal sensor Temp.: 96.15 Ambient Temp.: 85.48 F
23:19:34.434 -> Average Cal sensor Temp.: 96.22 Ambient Temp.: 85.48 F
23:19:35.989 -> Average Cal sensor Temp.: 96.07 Ambient Temp.: 85.51 F
23:19:37.541 -> Average Cal sensor Temp.: 95.76 Ambient Temp.: 85.55 F
23:19:39.100 -> Average Cal sensor Temp.: 95.92 Ambient Temp.: 85.48 F
23:19:40.653 -> Average Cal sensor Temp.: 96.03 Ambient Temp.: 85.51 F
23:19:42.207 -> Average Cal sensor Temp.: 96.12 Ambient Temp.: 85.66 F
23:19:43.760 -> Average Cal sensor Temp.: 96.09 Ambient Temp.: 85.55 F
23:19:45.317 -> Average Cal sensor Temp.: 96.01 Ambient Temp.: 85.59 F
23:19:46.871 -> Average Cal sensor Temp.: 96.04 Ambient Temp.: 85.59 F
23:19:48.422 -> Average Cal sensor Temp.: 96.00 Ambient Temp.: 85.69 F
23:19:49.974 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 85.66 F
23:19:51.525 -> Average Cal sensor Temp.: 96.19 Ambient Temp.: 85.69 F
23:19:53.084 -> Average Cal sensor Temp.: 95.97 Ambient Temp.: 85.69 F
23:19:54.645 -> Average Cal sensor Temp.: 95.98 Ambient Temp.: 85.77 F
23:19:56.204 -> Average Cal sensor Temp.: 96.19 Ambient Temp.: 85.80 F
23:19:57.765 -> Average Cal sensor Temp.: 96.18 Ambient Temp.: 85.80 F
23:19:59.329 -> Average Cal sensor Temp.: 96.21 Ambient Temp.: 85.80 F
23:20:00.885 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 85.77 F
23:20:02.435 -> Average Cal sensor Temp.: 96.11 Ambient Temp.: 85.84 F
23:20:03.995 -> Average Cal sensor Temp.: 96.29 Ambient Temp.: 85.84 F
23:20:05.550 -> Average Cal sensor Temp.: 96.28 Ambient Temp.: 85.91 F
23:20:07.104 -> Average Cal sensor Temp.: 96.06 Ambient Temp.: 85.84 F
23:20:08.667 -> Average Cal sensor Temp.: 96.27 Ambient Temp.: 85.91 F
23:20:10.225 -> Average Cal sensor Temp.: 96.33 Ambient Temp.: 85.91 F
23:20:11.779 -> Average Cal sensor Temp.: 96.31 Ambient Temp.: 85.91 F
23:20:13.340 -> Average Cal sensor Temp.: 96.35 Ambient Temp.: 85.98 F
23:20:14.901 -> Average Cal sensor Temp.: 96.15 Ambient Temp.: 85.98 F
23:20:16.425 -> Average Cal sensor Temp.: 96.28 Ambient Temp.: 85.98 F
23:20:17.987 -> Average Cal sensor Temp.: 96.12 Ambient Temp.: 86.05 F
23:20:19.544 -> Average Cal sensor Temp.: 96.09 Ambient Temp.: 86.05 F
23:20:21.097 -> Average Cal sensor Temp.: 95.76 Ambient Temp.: 86.05 F
23:20:22.654 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 85.95 F
23:20:24.211 -> Average Cal sensor Temp.: 96.03 Ambient Temp.: 85.98 F
23:20:25.799 -> Average Cal sensor Temp.: 96.09 Ambient Temp.: 86.09 F
23:20:27.353 -> Average Cal sensor Temp.: 96.13 Ambient Temp.: 86.09 F
23:20:28.914 -> Average Cal sensor Temp.: 96.15 Ambient Temp.: 86.13 F
23:20:30.436 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 86.13 F
23:20:32.023 -> Average Cal sensor Temp.: 96.25 Ambient Temp.: 86.13 F
23:20:33.580 -> Average Cal sensor Temp.: 96.21 Ambient Temp.: 86.09 F
23:20:35.137 -> Average Cal sensor Temp.: 96.21 Ambient Temp.: 86.13 F
23:20:36.663 -> Average Cal sensor Temp.: 96.00 Ambient Temp.: 86.09 F
23:20:38.221 -> Average Cal sensor Temp.: 96.07 Ambient Temp.: 86.13 F
23:20:39.783 -> Average Cal sensor Temp.: 96.25 Ambient Temp.: 86.16 F
23:20:41.344 -> Average Cal sensor Temp.: 95.98 Ambient Temp.: 86.16 F
23:20:42.902 -> Average Cal sensor Temp.: 96.23 Ambient Temp.: 86.13 F
23:20:44.460 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 86.16 F
23:20:46.017 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 86.16 F
23:20:47.572 -> Average Cal sensor Temp.: 96.10 Ambient Temp.: 86.13 F
23:20:49.131 -> Average Cal sensor Temp.: 96.09 Ambient Temp.: 86.13 F
23:20:50.688 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 86.16 F
23:20:52.250 -> Average Cal sensor Temp.: 96.35 Ambient Temp.: 86.13 F
23:20:53.806 -> Average Cal sensor Temp.: 96.07 Ambient Temp.: 86.13 F
23:20:55.362 -> Average Cal sensor Temp.: 96.48 Ambient Temp.: 86.13 F
23:20:56.918 -> Average Cal sensor Temp.: 96.10 Ambient Temp.: 86.13 F
23:20:58.480 -> Average Cal sensor Temp.: 96.11 Ambient Temp.: 86.16 F
23:21:00.038 -> Average Cal sensor Temp.: 96.27 Ambient Temp.: 86.09 F
23:21:01.568 -> Average Cal sensor Temp.: 96.62 Ambient Temp.: 86.16 F
23:21:03.155 -> Average Cal sensor Temp.: 96.36 Ambient Temp.: 86.20 F
23:21:04.684 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 86.20 F
23:21:06.244 -> Average Cal sensor Temp.: 96.22 Ambient Temp.: 86.16 F
23:21:07.799 -> Average Cal sensor Temp.: 96.05 Ambient Temp.: 86.20 F
23:21:09.354 -> Average Cal sensor Temp.: 95.86 Ambient Temp.: 86.16 F
23:21:10.914 -> Average Cal sensor Temp.: 95.88 Ambient Temp.: 86.16 F
23:21:12.470 -> Average Cal sensor Temp.: 95.62 Ambient Temp.: 86.13 F
23:21:14.022 -> Average Cal sensor Temp.: 95.75 Ambient Temp.: 86.13 F
23:21:15.577 -> Average Cal sensor Temp.: 95.83 Ambient Temp.: 86.13 F
23:21:17.136 -> Average Cal sensor Temp.: 95.87 Ambient Temp.: 86.05 F
23:21:18.688 -> Average Cal sensor Temp.: 95.75 Ambient Temp.: 85.98 F
23:21:20.246 -> Average Cal sensor Temp.: 95.72 Ambient Temp.: 85.98 F
23:21:21.807 -> Average Cal sensor Temp.: 95.87 Ambient Temp.: 85.98 F
23:21:23.366 -> Average Cal sensor Temp.: 95.69 Ambient Temp.: 85.91 F
23:21:24.917 -> Average Cal sensor Temp.: 95.74 Ambient Temp.: 85.87 F
23:21:26.475 -> Average Cal sensor Temp.: 95.91 Ambient Temp.: 85.84 F
23:21:28.030 -> Average Cal sensor Temp.: 95.74 Ambient Temp.: 85.87 F
23:21:29.586 -> Average Cal sensor Temp.: 95.54 Ambient Temp.: 85.84 F
23:21:31.152 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 85.84 F
23:21:32.706 -> Average Cal sensor Temp.: 95.75 Ambient Temp.: 85.77 F
23:21:34.263 -> Average Cal sensor Temp.: 96.22 Ambient Temp.: 85.73 F
23:21:35.820 -> Average Cal sensor Temp.: 96.18 Ambient Temp.: 85.84 F
23:21:37.371 -> Average Cal sensor Temp.: 96.12 Ambient Temp.: 85.77 F
23:21:38.929 -> Average Cal sensor Temp.: 96.44 Ambient Temp.: 85.69 F
23:21:40.486 -> Average Cal sensor Temp.: 96.39 Ambient Temp.: 85.73 F
23:21:42.042 -> Average Cal sensor Temp.: 96.37 Ambient Temp.: 85.73 F
23:21:43.590 -> Average Cal sensor Temp.: 96.16 Ambient Temp.: 85.69 F
23:21:45.145 -> Average Cal sensor Temp.: 95.95 Ambient Temp.: 85.66 F
23:21:46.703 -> Average Cal sensor Temp.: 96.32 Ambient Temp.: 85.69 F
23:21:48.265 -> Average Cal sensor Temp.: 95.83 Ambient Temp.: 85.73 F
23:21:49.826 -> Average Cal sensor Temp.: 95.76 Ambient Temp.: 85.66 F
23:21:51.384 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 85.59 F
23:21:52.948 -> Average Cal sensor Temp.: 95.95 Ambient Temp.: 85.55 F
23:21:54.474 -> Average Cal sensor Temp.: 95.82 Ambient Temp.: 85.51 F
23:21:56.030 -> Average Cal sensor Temp.: 96.26 Ambient Temp.: 85.55 F
23:21:57.589 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 85.51 F
23:21:59.145 -> Average Cal sensor Temp.: 96.32 Ambient Temp.: 85.48 F
23:22:00.702 -> Average Cal sensor Temp.: 96.14 Ambient Temp.: 85.51 F
23:22:02.253 -> Average Cal sensor Temp.: 96.11 Ambient Temp.: 85.51 F
23:22:03.810 -> Average Cal sensor Temp.: 95.90 Ambient Temp.: 85.48 F
23:22:05.367 -> Average Cal sensor Temp.: 95.98 Ambient Temp.: 85.48 F
23:22:06.924 -> Average Cal sensor Temp.: 95.86 Ambient Temp.: 85.41 F
23:22:08.482 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 85.44 F
23:22:10.040 -> Average Cal sensor Temp.: 95.59 Ambient Temp.: 85.33 F
23:22:11.598 -> Average Cal sensor Temp.: 95.59 Ambient Temp.: 85.33 F
23:22:13.154 -> Average Cal sensor Temp.: 96.05 Ambient Temp.: 85.30 F
23:22:14.719 -> Average Cal sensor Temp.: 95.96 Ambient Temp.: 85.33 F
23:22:16.280 -> Average Cal sensor Temp.: 95.61 Ambient Temp.: 85.26 F
23:22:17.842 -> Average Cal sensor Temp.: 95.64 Ambient Temp.: 85.15 F
23:22:19.371 -> Average Cal sensor Temp.: 96.02 Ambient Temp.: 85.15 F
23:22:20.925 -> Average Cal sensor Temp.: 95.71 Ambient Temp.: 85.15 F
23:22:22.483 -> Average Cal sensor Temp.: 95.82 Ambient Temp.: 85.12 F
23:22:24.046 -> Average Cal sensor Temp.: 96.14 Ambient Temp.: 85.08 F
23:22:25.603 -> Average Cal sensor Temp.: 95.89 Ambient Temp.: 85.05 F
23:22:27.158 -> Average Cal sensor Temp.: 95.85 Ambient Temp.: 85.01 F
23:22:28.716 -> Average Cal sensor Temp.: 95.59 Ambient Temp.: 84.97 F
23:22:30.274 -> Average Cal sensor Temp.: 95.82 Ambient Temp.: 85.01 F
23:22:31.827 -> Average Cal sensor Temp.: 95.64 Ambient Temp.: 84.97 F
23:22:33.386 -> Average Cal sensor Temp.: 95.76 Ambient Temp.: 84.90 F
23:22:34.941 -> Average Cal sensor Temp.: 95.85 Ambient Temp.: 84.90 F
23:22:36.493 -> Average Cal sensor Temp.: 95.91 Ambient Temp.: 84.90 F
23:22:38.051 -> Average Cal sensor Temp.: 95.62 Ambient Temp.: 84.79 F
23:22:39.612 -> Average Cal sensor Temp.: 96.01 Ambient Temp.: 84.79 F
23:22:41.170 -> Average Cal sensor Temp.: 96.12 Ambient Temp.: 84.79 F
23:22:42.722 -> Average Cal sensor Temp.: 96.14 Ambient Temp.: 84.79 F
23:22:44.286 -> Average Cal sensor Temp.: 95.84 Ambient Temp.: 84.79 F
23:22:45.838 -> Average Cal sensor Temp.: 95.89 Ambient Temp.: 84.76 F
23:22:47.390 -> Average Cal sensor Temp.: 95.75 Ambient Temp.: 84.69 F
23:22:48.943 -> Average Cal sensor Temp.: 95.72 Ambient Temp.: 84.69 F
23:22:50.500 -> Average Cal sensor Temp.: 96.11 Ambient Temp.: 84.79 F
23:22:52.054 -> Average Cal sensor Temp.: 96.24 Ambient Temp.: 84.72 F
23:22:53.611 -> Average Cal sensor Temp.: 96.37 Ambient Temp.: 84.72 F
23:22:55.168 -> Average Cal sensor Temp.: 95.94 Ambient Temp.: 84.76 F
23:22:56.727 -> Average Cal sensor Temp.: 96.18 Ambient Temp.: 84.72 F
23:22:58.282 -> Average Cal sensor Temp.: 96.04 Ambient Temp.: 84.69 F
23:22:59.841 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 84.69 F
23:23:01.397 -> Average Cal sensor Temp.: 96.21 Ambient Temp.: 84.69 F
23:23:02.955 -> Average Cal sensor Temp.: 96.29 Ambient Temp.: 84.65 F
23:23:04.511 -> Average Cal sensor Temp.: 96.09 Ambient Temp.: 84.65 F
23:23:06.070 -> Average Cal sensor Temp.: 96.39 Ambient Temp.: 84.65 F
23:23:07.630 -> Average Cal sensor Temp.: 96.47 Ambient Temp.: 84.58 F
23:23:09.191 -> Average Cal sensor Temp.: 95.79 Ambient Temp.: 84.61 F
23:23:10.747 -> Average Cal sensor Temp.: 95.86 Ambient Temp.: 84.58 F
23:23:12.308 -> Average Cal sensor Temp.: 95.82 Ambient Temp.: 84.58 F
23:23:13.871 -> Average Cal sensor Temp.: 95.98 Ambient Temp.: 84.51 F
23:23:15.398 -> Average Cal sensor Temp.: 96.06 Ambient Temp.: 84.51 F
23:23:16.958 -> Average Cal sensor Temp.: 96.32 Ambient Temp.: 84.54 F
23:23:18.514 -> Average Cal sensor Temp.: 96.44 Ambient Temp.: 84.47 F
23:23:20.070 -> Average Cal sensor Temp.: 96.06 Ambient Temp.: 84.51 F
23:23:21.633 -> Average Cal sensor Temp.: 95.66 Ambient Temp.: 84.40 F
23:23:23.196 -> Average Cal sensor Temp.: 95.75 Ambient Temp.: 84.40 F
23:23:24.754 -> Average Cal sensor Temp.: 95.62 Ambient Temp.: 84.40 F
23:23:26.318 -> Average Cal sensor Temp.: 96.01 Ambient Temp.: 84.40 F
23:23:27.874 -> Average Cal sensor Temp.: 95.75 Ambient Temp.: 84.40 F
23:23:29.437 -> Average Cal sensor Temp.: 95.74 Ambient Temp.: 84.33 F
23:23:30.996 -> Average Cal sensor Temp.: 95.60 Ambient Temp.: 84.33 F
23:23:32.551 -> Average Cal sensor Temp.: 96.07 Ambient Temp.: 84.29 F
23:23:34.107 -> Average Cal sensor Temp.: 95.99 Ambient Temp.: 84.25 F
23:23:35.666 -> Average Cal sensor Temp.: 96.14 Ambient Temp.: 84.22 F
23:23:37.221 -> Average Cal sensor Temp.: 96.39 Ambient Temp.: 84.25 F
23:23:38.776 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 84.22 F
23:23:40.335 -> Average Cal sensor Temp.: 96.29 Ambient Temp.: 84.25 F
23:23:41.893 -> Average Cal sensor Temp.: 95.88 Ambient Temp.: 84.29 F
23:23:43.420 -> Average Cal sensor Temp.: 95.89 Ambient Temp.: 84.18 F
23:23:44.980 -> Average Cal sensor Temp.: 96.31 Ambient Temp.: 84.18 F
23:23:46.536 -> Average Cal sensor Temp.: 96.24 Ambient Temp.: 84.18 F
23:23:48.093 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 84.15 F
23:23:49.656 -> Average Cal sensor Temp.: 95.98 Ambient Temp.: 84.11 F
23:23:51.208 -> Average Cal sensor Temp.: 96.09 Ambient Temp.: 84.11 F
23:23:52.771 -> Average Cal sensor Temp.: 96.16 Ambient Temp.: 84.07 F
23:23:54.335 -> Average Cal sensor Temp.: 96.11 Ambient Temp.: 84.07 F
23:23:55.888 -> Average Cal sensor Temp.: 96.28 Ambient Temp.: 84.11 F
23:23:57.450 -> Average Cal sensor Temp.: 96.10 Ambient Temp.: 84.07 F
23:23:59.012 -> Average Cal sensor Temp.: 96.09 Ambient Temp.: 84.11 F
23:24:00.567 -> Average Cal sensor Temp.: 95.96 Ambient Temp.: 84.07 F
23:24:02.120 -> Average Cal sensor Temp.: 96.25 Ambient Temp.: 84.07 F
23:24:03.676 -> Average Cal sensor Temp.: 96.22 Ambient Temp.: 84.07 F
23:24:05.236 -> Average Cal sensor Temp.: 96.28 Ambient Temp.: 84.07 F
23:24:06.794 -> Average Cal sensor Temp.: 96.25 Ambient Temp.: 84.07 F
23:24:08.354 -> Average Cal sensor Temp.: 96.20 Ambient Temp.: 84.11 F
23:24:09.878 -> Average Cal sensor Temp.: 96.32 Ambient Temp.: 84.11 F
23:24:11.436 -> Average Cal sensor Temp.: 96.22 Ambient Temp.: 84.11 F
23:24:13.002 -> Average Cal sensor Temp.: 96.26 Ambient Temp.: 84.15 F
23:24:14.559 -> Average Cal sensor Temp.: 96.53 Ambient Temp.: 84.11 F
23:24:16.120 -> Average Cal sensor Temp.: 96.62 Ambient Temp.: 84.15 F
23:24:17.682 -> Average Cal sensor Temp.: 96.40 Ambient Temp.: 84.11 F
23:24:19.240 -> Average Cal sensor Temp.: 96.24 Ambient Temp.: 84.15 F
23:24:20.799 -> Average Cal sensor Temp.: 96.42 Ambient Temp.: 84.15 F
23:24:22.325 -> Average Cal sensor Temp.: 96.35 Ambient Temp.: 84.18 F
23:24:23.888 -> Average Cal sensor Temp.: 96.21 Ambient Temp.: 84.18 F
23:24:25.454 -> Average Cal sensor Temp.: 96.27 Ambient Temp.: 84.22 F
23:24:27.013 -> Average Cal sensor Temp.: 96.43 Ambient Temp.: 84.18 F
23:24:28.573 -> Average Cal sensor Temp.: 96.39 Ambient Temp.: 84.22 F
23:24:30.136 -> Average Cal sensor Temp.: 96.65 Ambient Temp.: 84.22 F
23:24:31.662 -> Average Cal sensor Temp.: 96.41 Ambient Temp.: 84.29 F
23:24:33.219 -> Average Cal sensor Temp.: 96.58 Ambient Temp.: 84.25 F
23:24:34.782 -> Average Cal sensor Temp.: 96.39 Ambient Temp.: 84.29 F
23:24:36.336 -> Average Cal sensor Temp.: 96.54 Ambient Temp.: 84.25 F
23:24:37.893 -> Average Cal sensor Temp.: 96.66 Ambient Temp.: 84.33 F
23:24:39.453 -> Average Cal sensor Temp.: 96.35 Ambient Temp.: 84.29 F
23:24:41.009 -> Average Cal sensor Temp.: 96.09 Ambient Temp.: 84.29 F
23:24:42.567 -> Average Cal sensor Temp.: 96.16 Ambient Temp.: 84.33 F
23:24:44.129 -> Average Cal sensor Temp.: 96.18 Ambient Temp.: 87.03 F
23:24:45.686 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 84.36 F
23:24:47.247 -> Average Cal sensor Temp.: 96.29 Ambient Temp.: 84.33 F
23:24:48.809 -> Average Cal sensor Temp.: 96.14 Ambient Temp.: 84.33 F
23:24:50.376 -> Average Cal sensor Temp.: 96.36 Ambient Temp.: 84.40 F
23:24:51.901 -> Average Cal sensor Temp.: 96.19 Ambient Temp.: 84.40 F
23:24:53.461 -> Average Cal sensor Temp.: 96.21 Ambient Temp.: 84.40 F
23:24:55.025 -> Average Cal sensor Temp.: 96.42 Ambient Temp.: 84.40 F
23:24:56.587 -> Average Cal sensor Temp.: 96.54 Ambient Temp.: 84.40 F
23:24:58.143 -> Average Cal sensor Temp.: 96.57 Ambient Temp.: 84.47 F
23:24:59.705 -> Average Cal sensor Temp.: 96.55 Ambient Temp.: 84.47 F
23:25:01.265 -> Average Cal sensor Temp.: 96.41 Ambient Temp.: 84.51 F
23:25:02.826 -> Average Cal sensor Temp.: 96.39 Ambient Temp.: 84.47 F
23:25:04.386 -> Average Cal sensor Temp.: 96.69 Ambient Temp.: 84.54 F
23:25:05.940 -> Average Cal sensor Temp.: 96.40 Ambient Temp.: 84.51 F
23:25:07.496 -> Average Cal sensor Temp.: 96.48 Ambient Temp.: 84.54 F
23:25:09.060 -> Average Cal sensor Temp.: 96.53 Ambient Temp.: 84.58 F
23:25:10.587 -> Average Cal sensor Temp.: 96.26 Ambient Temp.: 84.54 F
23:25:12.146 -> Average Cal sensor Temp.: 96.65 Ambient Temp.: 84.61 F
23:25:13.703 -> Average Cal sensor Temp.: 96.65 Ambient Temp.: 84.61 F
23:25:15.261 -> Average Cal sensor Temp.: 96.76 Ambient Temp.: 84.61 F
23:25:16.813 -> Average Cal sensor Temp.: 96.54 Ambient Temp.: 84.58 F
23:25:18.378 -> Average Cal sensor Temp.: 96.35 Ambient Temp.: 84.72 F
23:25:19.941 -> Average Cal sensor Temp.: 96.43 Ambient Temp.: 84.69 F
23:25:21.503 -> Average Cal sensor Temp.: 96.53 Ambient Temp.: 84.69 F
23:25:23.055 -> Average Cal sensor Temp.: 96.49 Ambient Temp.: 84.72 F
23:25:24.613 -> Average Cal sensor Temp.: 96.49 Ambient Temp.: 84.76 F
23:25:26.174 -> Average Cal sensor Temp.: 96.55 Ambient Temp.: 84.79 F
23:25:27.729 -> Average Cal sensor Temp.: 96.55 Ambient Temp.: 84.79 F
23:25:29.291 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 84.76 F
23:25:30.852 -> Average Cal sensor Temp.: 96.67 Ambient Temp.: 84.79 F
23:25:32.380 -> Average Cal sensor Temp.: 96.96 Ambient Temp.: 84.87 F
23:25:33.939 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 84.90 F
23:25:35.498 -> Average Cal sensor Temp.: 96.55 Ambient Temp.: 84.79 F
23:25:37.055 -> Average Cal sensor Temp.: 96.50 Ambient Temp.: 84.97 F
23:25:38.617 -> Average Cal sensor Temp.: 96.54 Ambient Temp.: 84.97 F
23:25:40.178 -> Average Cal sensor Temp.: 96.53 Ambient Temp.: 84.94 F
23:25:41.736 -> Average Cal sensor Temp.: 96.01 Ambient Temp.: 84.97 F
23:25:43.290 -> Average Cal sensor Temp.: 95.95 Ambient Temp.: 85.01 F
23:25:44.847 -> Average Cal sensor Temp.: 96.04 Ambient Temp.: 84.97 F
23:25:46.403 -> Average Cal sensor Temp.: 96.25 Ambient Temp.: 85.01 F
23:25:47.958 -> Average Cal sensor Temp.: 96.40 Ambient Temp.: 85.01 F
23:25:49.517 -> Average Cal sensor Temp.: 96.22 Ambient Temp.: 85.01 F
23:25:51.075 -> Average Cal sensor Temp.: 96.39 Ambient Temp.: 85.05 F
23:25:52.635 -> Average Cal sensor Temp.: 96.23 Ambient Temp.: 85.08 F
23:25:54.195 -> Average Cal sensor Temp.: 96.29 Ambient Temp.: 85.05 F
23:25:55.755 -> Average Cal sensor Temp.: 96.50 Ambient Temp.: 85.15 F
23:25:57.317 -> Average Cal sensor Temp.: 96.35 Ambient Temp.: 85.05 F
23:25:58.848 -> Average Cal sensor Temp.: 96.47 Ambient Temp.: 85.15 F
23:26:00.404 -> Average Cal sensor Temp.: 96.48 Ambient Temp.: 85.12 F
23:26:01.960 -> Average Cal sensor Temp.: 96.46 Ambient Temp.: 85.15 F
23:26:03.523 -> Average Cal sensor Temp.: 96.69 Ambient Temp.: 85.19 F
23:26:05.085 -> Average Cal sensor Temp.: 96.64 Ambient Temp.: 85.19 F
23:26:06.647 -> Average Cal sensor Temp.: 96.36 Ambient Temp.: 85.30 F
23:26:08.208 -> Average Cal sensor Temp.: 96.36 Ambient Temp.: 85.26 F
23:26:09.739 -> Average Cal sensor Temp.: 96.37 Ambient Temp.: 85.30 F
23:26:11.301 -> Average Cal sensor Temp.: 96.28 Ambient Temp.: 85.33 F
23:26:12.858 -> Average Cal sensor Temp.: 96.14 Ambient Temp.: 85.33 F
23:26:14.417 -> Average Cal sensor Temp.: 96.28 Ambient Temp.: 85.37 F
23:26:15.976 -> Average Cal sensor Temp.: 96.47 Ambient Temp.: 85.33 F
23:26:17.526 -> Average Cal sensor Temp.: 96.69 Ambient Temp.: 85.37 F
23:26:19.089 -> Average Cal sensor Temp.: 96.42 Ambient Temp.: 85.44 F
23:26:20.641 -> Average Cal sensor Temp.: 96.15 Ambient Temp.: 85.41 F
23:26:22.195 -> Average Cal sensor Temp.: 96.02 Ambient Temp.: 85.41 F
23:26:23.756 -> Average Cal sensor Temp.: 96.27 Ambient Temp.: 85.41 F
23:26:25.315 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 85.44 F
23:26:26.873 -> Average Cal sensor Temp.: 96.39 Ambient Temp.: 85.51 F
23:26:28.429 -> Average Cal sensor Temp.: 96.19 Ambient Temp.: 85.44 F
23:26:29.987 -> Average Cal sensor Temp.: 96.28 Ambient Temp.: 85.51 F
23:26:31.545 -> Average Cal sensor Temp.: 96.46 Ambient Temp.: 85.48 F
23:26:33.099 -> Average Cal sensor Temp.: 96.18 Ambient Temp.: 85.51 F
23:26:34.655 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 85.55 F
23:26:36.214 -> Average Cal sensor Temp.: 96.36 Ambient Temp.: 85.55 F
23:26:37.768 -> Average Cal sensor Temp.: 96.35 Ambient Temp.: 85.66 F
23:26:39.330 -> Average Cal sensor Temp.: 96.42 Ambient Temp.: 85.59 F
23:26:40.893 -> Average Cal sensor Temp.: 96.22 Ambient Temp.: 85.55 F
23:26:42.440 -> Average Cal sensor Temp.: 96.40 Ambient Temp.: 85.59 F
23:26:43.998 -> Average Cal sensor Temp.: 96.48 Ambient Temp.: 85.69 F
23:26:45.559 -> Average Cal sensor Temp.: 96.64 Ambient Temp.: 85.69 F
23:26:47.120 -> Average Cal sensor Temp.: 96.49 Ambient Temp.: 85.69 F
23:26:48.680 -> Average Cal sensor Temp.: 96.58 Ambient Temp.: 85.73 F
23:26:50.205 -> Average Cal sensor Temp.: 96.68 Ambient Temp.: 85.73 F
23:26:51.766 -> Average Cal sensor Temp.: 96.61 Ambient Temp.: 85.73 F
23:26:53.325 -> Average Cal sensor Temp.: 96.53 Ambient Temp.: 85.73 F
23:26:54.883 -> Average Cal sensor Temp.: 96.43 Ambient Temp.: 85.80 F
23:26:56.447 -> Average Cal sensor Temp.: 96.53 Ambient Temp.: 85.73 F
23:26:58.005 -> Average Cal sensor Temp.: 96.32 Ambient Temp.: 85.73 F
23:26:59.566 -> Average Cal sensor Temp.: 96.65 Ambient Temp.: 85.80 F
23:27:01.124 -> Average Cal sensor Temp.: 96.51 Ambient Temp.: 85.77 F
23:27:02.685 -> Average Cal sensor Temp.: 96.68 Ambient Temp.: 85.77 F
23:27:04.209 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 85.80 F
23:27:05.773 -> Average Cal sensor Temp.: 96.73 Ambient Temp.: 85.84 F
23:27:07.330 -> Average Cal sensor Temp.: 96.66 Ambient Temp.: 85.80 F
23:27:08.894 -> Average Cal sensor Temp.: 96.63 Ambient Temp.: 85.84 F
23:27:10.457 -> Average Cal sensor Temp.: 96.37 Ambient Temp.: 85.84 F
23:27:12.018 -> Average Cal sensor Temp.: 96.68 Ambient Temp.: 85.87 F
23:27:13.582 -> Average Cal sensor Temp.: 96.55 Ambient Temp.: 85.80 F
23:27:15.109 -> Average Cal sensor Temp.: 96.24 Ambient Temp.: 85.84 F
23:27:16.670 -> Average Cal sensor Temp.: 96.40 Ambient Temp.: 85.84 F
23:27:18.223 -> Average Cal sensor Temp.: 96.61 Ambient Temp.: 85.84 F
23:27:19.781 -> Average Cal sensor Temp.: 96.91 Ambient Temp.: 85.91 F
23:27:21.367 -> Average Cal sensor Temp.: 96.80 Ambient Temp.: 85.95 F
23:27:22.892 -> Average Cal sensor Temp.: 96.91 Ambient Temp.: 85.91 F
23:27:24.448 -> Average Cal sensor Temp.: 96.45 Ambient Temp.: 85.91 F
23:27:26.037 -> Average Cal sensor Temp.: 96.59 Ambient Temp.: 85.95 F
23:27:27.589 -> Average Cal sensor Temp.: 96.59 Ambient Temp.: 85.91 F
23:27:29.145 -> Average Cal sensor Temp.: 96.79 Ambient Temp.: 85.98 F
23:27:30.702 -> Average Cal sensor Temp.: 96.59 Ambient Temp.: 85.95 F
23:27:32.256 -> Average Cal sensor Temp.: 96.52 Ambient Temp.: 85.98 F
23:27:33.809 -> Average Cal sensor Temp.: 96.63 Ambient Temp.: 85.98 F
23:27:35.364 -> Average Cal sensor Temp.: 96.48 Ambient Temp.: 86.05 F
23:27:36.918 -> Average Cal sensor Temp.: 96.47 Ambient Temp.: 86.05 F
23:27:38.472 -> Average Cal sensor Temp.: 96.41 Ambient Temp.: 86.13 F
23:27:40.029 -> Average Cal sensor Temp.: 96.49 Ambient Temp.: 85.98 F
23:27:41.587 -> Average Cal sensor Temp.: 96.54 Ambient Temp.: 85.98 F
23:27:43.144 -> Average Cal sensor Temp.: 96.50 Ambient Temp.: 85.98 F
23:27:44.701 -> Average Cal sensor Temp.: 96.39 Ambient Temp.: 86.05 F
23:27:46.255 -> Average Cal sensor Temp.: 96.13 Ambient Temp.: 86.05 F
23:27:47.814 -> Average Cal sensor Temp.: 95.99 Ambient Temp.: 85.98 F
23:27:49.375 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 85.95 F
23:27:50.928 -> Average Cal sensor Temp.: 96.16 Ambient Temp.: 85.91 F
23:27:52.480 -> Average Cal sensor Temp.: 96.61 Ambient Temp.: 85.98 F
23:27:54.041 -> Average Cal sensor Temp.: 96.32 Ambient Temp.: 85.91 F
23:27:55.598 -> Average Cal sensor Temp.: 96.10 Ambient Temp.: 85.91 F
23:27:57.159 -> Average Cal sensor Temp.: 96.23 Ambient Temp.: 85.91 F
23:27:58.724 -> Average Cal sensor Temp.: 96.49 Ambient Temp.: 85.87 F
23:28:00.284 -> Average Cal sensor Temp.: 96.70 Ambient Temp.: 85.91 F
23:28:01.838 -> Average Cal sensor Temp.: 96.66 Ambient Temp.: 85.87 F
23:28:03.399 -> Average Cal sensor Temp.: 96.28 Ambient Temp.: 85.84 F
23:28:04.924 -> Average Cal sensor Temp.: 96.34 Ambient Temp.: 85.87 F
23:28:06.513 -> Average Cal sensor Temp.: 96.14 Ambient Temp.: 85.84 F
23:28:08.038 -> Average Cal sensor Temp.: 95.95 Ambient Temp.: 85.84 F
23:28:09.601 -> Average Cal sensor Temp.: 95.64 Ambient Temp.: 85.80 F
23:28:11.160 -> Average Cal sensor Temp.: 96.01 Ambient Temp.: 85.73 F
23:28:12.713 -> Average Cal sensor Temp.: 96.25 Ambient Temp.: 85.77 F
23:28:14.274 -> Average Cal sensor Temp.: 96.26 Ambient Temp.: 85.73 F
23:28:15.837 -> Average Cal sensor Temp.: 96.31 Ambient Temp.: 85.73 F
23:28:17.397 -> Average Cal sensor Temp.: 96.61 Ambient Temp.: 85.73 F
23:28:18.960 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 85.69 F
23:28:20.485 -> Average Cal sensor Temp.: 96.40 Ambient Temp.: 85.69 F
23:28:22.046 -> Average Cal sensor Temp.: 96.35 Ambient Temp.: 85.66 F
23:28:23.605 -> Average Cal sensor Temp.: 96.57 Ambient Temp.: 85.66 F
23:28:25.165 -> Average Cal sensor Temp.: 96.45 Ambient Temp.: 85.59 F
23:28:26.724 -> Average Cal sensor Temp.: 96.31 Ambient Temp.: 85.59 F
23:28:28.285 -> Average Cal sensor Temp.: 96.27 Ambient Temp.: 85.51 F
23:28:29.847 -> Average Cal sensor Temp.: 96.61 Ambient Temp.: 85.55 F
23:28:31.376 -> Average Cal sensor Temp.: 96.00 Ambient Temp.: 85.51 F
23:28:32.934 -> Average Cal sensor Temp.: 95.94 Ambient Temp.: 85.48 F
23:28:34.495 -> Average Cal sensor Temp.: 96.19 Ambient Temp.: 85.41 F
23:28:36.055 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 85.44 F
23:28:37.617 -> Average Cal sensor Temp.: 96.03 Ambient Temp.: 85.41 F
23:28:39.177 -> Average Cal sensor Temp.: 95.99 Ambient Temp.: 85.33 F
23:28:40.734 -> Average Cal sensor Temp.: 95.86 Ambient Temp.: 85.30 F
23:28:42.295 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 85.33 F
23:28:43.853 -> Average Cal sensor Temp.: 96.06 Ambient Temp.: 85.33 F
23:28:45.414 -> Average Cal sensor Temp.: 95.96 Ambient Temp.: 85.26 F
23:28:46.969 -> Average Cal sensor Temp.: 96.15 Ambient Temp.: 85.26 F
23:28:48.530 -> Average Cal sensor Temp.: 96.41 Ambient Temp.: 85.15 F
23:28:50.055 -> Average Cal sensor Temp.: 96.18 Ambient Temp.: 85.15 F
23:28:51.618 -> Average Cal sensor Temp.: 96.41 Ambient Temp.: 85.08 F
23:28:53.176 -> Average Cal sensor Temp.: 96.56 Ambient Temp.: 85.12 F
23:28:54.732 -> Average Cal sensor Temp.: 96.47 Ambient Temp.: 85.12 F
23:28:56.288 -> Average Cal sensor Temp.: 96.36 Ambient Temp.: 85.05 F
23:28:57.843 -> Average Cal sensor Temp.: 96.54 Ambient Temp.: 85.01 F
23:28:59.403 -> Average Cal sensor Temp.: 96.54 Ambient Temp.: 85.05 F
23:29:00.960 -> Average Cal sensor Temp.: 96.55 Ambient Temp.: 85.05 F
23:29:02.519 -> Average Cal sensor Temp.: 96.48 Ambient Temp.: 85.01 F
23:29:04.080 -> Average Cal sensor Temp.: 96.34 Ambient Temp.: 84.97 F
23:29:05.634 -> Average Cal sensor Temp.: 96.12 Ambient Temp.: 85.01 F
23:29:07.191 -> Average Cal sensor Temp.: 96.24 Ambient Temp.: 84.90 F
23:29:08.748 -> Average Cal sensor Temp.: 96.57 Ambient Temp.: 84.94 F
23:29:10.301 -> Average Cal sensor Temp.: 96.65 Ambient Temp.: 84.87 F
23:29:11.865 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 84.79 F
23:29:13.424 -> Average Cal sensor Temp.: 96.51 Ambient Temp.: 84.87 F
23:29:14.985 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 84.79 F
23:29:16.543 -> Average Cal sensor Temp.: 96.42 Ambient Temp.: 84.76 F
23:29:18.103 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 84.76 F
23:29:19.627 -> Average Cal sensor Temp.: 96.59 Ambient Temp.: 84.69 F
23:29:21.216 -> Average Cal sensor Temp.: 96.52 Ambient Temp.: 84.69 F
23:29:22.775 -> Average Cal sensor Temp.: 96.29 Ambient Temp.: 84.69 F
23:29:24.303 -> Average Cal sensor Temp.: 96.71 Ambient Temp.: 84.65 F
23:29:25.863 -> Average Cal sensor Temp.: 96.48 Ambient Temp.: 84.72 F
23:29:27.422 -> Average Cal sensor Temp.: 96.47 Ambient Temp.: 84.65 F
23:29:28.980 -> Average Cal sensor Temp.: 96.32 Ambient Temp.: 84.61 F
23:29:30.539 -> Average Cal sensor Temp.: 96.19 Ambient Temp.: 84.58 F
23:29:32.099 -> Average Cal sensor Temp.: 96.01 Ambient Temp.: 84.54 F
23:29:33.655 -> Average Cal sensor Temp.: 96.16 Ambient Temp.: 84.54 F
23:29:35.216 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 84.51 F
23:29:36.773 -> Average Cal sensor Temp.: 95.87 Ambient Temp.: 84.40 F
23:29:38.332 -> Average Cal sensor Temp.: 96.23 Ambient Temp.: 84.40 F
23:29:39.894 -> Average Cal sensor Temp.: 96.16 Ambient Temp.: 84.36 F
23:29:41.420 -> Average Cal sensor Temp.: 95.91 Ambient Temp.: 84.29 F
23:29:42.980 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 84.29 F
23:29:44.535 -> Average Cal sensor Temp.: 96.48 Ambient Temp.: 84.33 F
23:29:46.099 -> Average Cal sensor Temp.: 96.38 Ambient Temp.: 84.29 F
23:29:47.655 -> Average Cal sensor Temp.: 96.26 Ambient Temp.: 84.29 F
23:29:49.219 -> Average Cal sensor Temp.: 96.10 Ambient Temp.: 84.22 F
23:29:50.783 -> Average Cal sensor Temp.: 96.46 Ambient Temp.: 84.22 F
23:29:52.333 -> Average Cal sensor Temp.: 96.80 Ambient Temp.: 84.22 F
23:29:53.889 -> Average Cal sensor Temp.: 96.60 Ambient Temp.: 84.18 F
23:29:55.443 -> Average Cal sensor Temp.: 96.49 Ambient Temp.: 84.22 F
23:29:57.026 -> Average Cal sensor Temp.: 96.67 Ambient Temp.: 84.18 F
23:29:58.552 -> Average Cal sensor Temp.: 96.36 Ambient Temp.: 84.22 F
23:30:00.107 -> Average Cal sensor Temp.: 95.90 Ambient Temp.: 84.15 F
23:30:01.663 -> Average Cal sensor Temp.: 96.34 Ambient Temp.: 84.18 F
23:30:03.251 -> Average Cal sensor Temp.: 96.18 Ambient Temp.: 84.07 F
23:30:04.806 -> Average Cal sensor Temp.: 96.41 Ambient Temp.: 84.11 F
23:30:06.365 -> Average Cal sensor Temp.: 96.47 Ambient Temp.: 84.11 F
23:30:07.891 -> Average Cal sensor Temp.: 96.69 Ambient Temp.: 84.11 F
23:30:09.445 -> Average Cal sensor Temp.: 96.78 Ambient Temp.: 84.15 F
23:30:11.001 -> Average Cal sensor Temp.: 96.52 Ambient Temp.: 84.11 F
23:30:12.589 -> Average Cal sensor Temp.: 96.05 Ambient Temp.: 84.07 F
23:30:14.144 -> Average Cal sensor Temp.: 96.53 Ambient Temp.: 84.11 F
23:30:15.666 -> Average Cal sensor Temp.: 96.65 Ambient Temp.: 84.07 F
23:30:17.225 -> Average Cal sensor Temp.: 96.56 Ambient Temp.: 84.11 F
23:30:18.786 -> Average Cal sensor Temp.: 96.69 Ambient Temp.: 84.07 F
23:30:20.352 -> Average Cal sensor Temp.: 96.66 Ambient Temp.: 84.07 F
23:30:21.905 -> Average Cal sensor Temp.: 96.60 Ambient Temp.: 84.11 F
23:30:23.460 -> Average Cal sensor Temp.: 96.77 Ambient Temp.: 84.11 F
23:30:25.019 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 84.00 F
23:30:26.569 -> Average Cal sensor Temp.: 96.81 Ambient Temp.: 84.11 F
23:30:28.131 -> Average Cal sensor Temp.: 96.72 Ambient Temp.: 84.11 F
23:30:29.690 -> Average Cal sensor Temp.: 96.73 Ambient Temp.: 84.11 F
23:30:31.242 -> Average Cal sensor Temp.: 96.72 Ambient Temp.: 84.15 F
23:30:32.805 -> Average Cal sensor Temp.: 96.64 Ambient Temp.: 84.11 F
23:30:34.365 -> Average Cal sensor Temp.: 96.74 Ambient Temp.: 84.15 F
23:30:35.920 -> Average Cal sensor Temp.: 96.69 Ambient Temp.: 84.11 F
23:30:37.473 -> Average Cal sensor Temp.: 96.67 Ambient Temp.: 84.15 F
23:30:39.029 -> Average Cal sensor Temp.: 96.73 Ambient Temp.: 84.18 F
23:30:40.580 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 84.15 F
23:30:42.127 -> Average Cal sensor Temp.: 96.85 Ambient Temp.: 84.11 F
23:30:43.687 -> Average Cal sensor Temp.: 96.71 Ambient Temp.: 84.18 F
23:30:45.244 -> Average Cal sensor Temp.: 96.74 Ambient Temp.: 84.11 F
23:30:46.801 -> Average Cal sensor Temp.: 96.93 Ambient Temp.: 84.22 F
23:30:48.354 -> Average Cal sensor Temp.: 96.84 Ambient Temp.: 84.18 F
23:30:49.915 -> Average Cal sensor Temp.: 96.50 Ambient Temp.: 84.25 F
23:30:51.475 -> Average Cal sensor Temp.: 96.58 Ambient Temp.: 84.25 F
23:30:53.026 -> Average Cal sensor Temp.: 96.70 Ambient Temp.: 84.18 F
23:30:54.583 -> Average Cal sensor Temp.: 97.00 Ambient Temp.: 84.29 F
23:30:56.139 -> Average Cal sensor Temp.: 96.92 Ambient Temp.: 84.33 F
23:30:57.698 -> Average Cal sensor Temp.: 96.50 Ambient Temp.: 84.29 F
23:30:59.248 -> Average Cal sensor Temp.: 96.74 Ambient Temp.: 84.25 F
23:31:00.810 -> Average Cal sensor Temp.: 96.50 Ambient Temp.: 84.33 F
23:31:02.359 -> Average Cal sensor Temp.: 96.53 Ambient Temp.: 84.36 F
23:31:03.923 -> Average Cal sensor Temp.: 96.55 Ambient Temp.: 84.36 F
23:31:05.475 -> Average Cal sensor Temp.: 96.90 Ambient Temp.: 84.40 F
23:31:07.029 -> Average Cal sensor Temp.: 96.81 Ambient Temp.: 84.40 F
23:31:08.592 -> Average Cal sensor Temp.: 96.74 Ambient Temp.: 84.47 F
23:31:10.152 -> Average Cal sensor Temp.: 96.73 Ambient Temp.: 84.40 F
23:31:11.706 -> Average Cal sensor Temp.: 96.82 Ambient Temp.: 84.51 F
23:31:13.268 -> Average Cal sensor Temp.: 96.57 Ambient Temp.: 84.47 F
23:31:14.824 -> Average Cal sensor Temp.: 96.71 Ambient Temp.: 84.54 F
23:31:16.385 -> Average Cal sensor Temp.: 96.66 Ambient Temp.: 84.51 F
23:31:17.940 -> Average Cal sensor Temp.: 96.94 Ambient Temp.: 84.51 F
23:31:19.498 -> Average Cal sensor Temp.: 97.12 Ambient Temp.: 84.54 F
23:31:21.058 -> Average Cal sensor Temp.: 96.87 Ambient Temp.: 84.54 F
23:31:22.619 -> Average Cal sensor Temp.: 96.76 Ambient Temp.: 84.54 F
23:31:24.151 -> Average Cal sensor Temp.: 96.78 Ambient Temp.: 84.58 F
23:31:25.710 -> Average Cal sensor Temp.: 96.73 Ambient Temp.: 84.58 F
23:31:27.266 -> Average Cal sensor Temp.: 96.88 Ambient Temp.: 84.61 F
23:31:28.824 -> Average Cal sensor Temp.: 96.70 Ambient Temp.: 84.58 F
23:31:30.381 -> Average Cal sensor Temp.: 96.57 Ambient Temp.: 84.65 F
23:31:31.935 -> Average Cal sensor Temp.: 96.65 Ambient Temp.: 84.72 F
23:31:33.494 -> Average Cal sensor Temp.: 96.64 Ambient Temp.: 84.72 F
23:31:35.049 -> Average Cal sensor Temp.: 96.55 Ambient Temp.: 84.72 F
23:31:36.613 -> Average Cal sensor Temp.: 96.38 Ambient Temp.: 84.69 F
23:31:38.172 -> Average Cal sensor Temp.: 96.69 Ambient Temp.: 84.72 F
23:31:39.735 -> Average Cal sensor Temp.: 96.46 Ambient Temp.: 84.72 F
23:31:41.296 -> Average Cal sensor Temp.: 96.55 Ambient Temp.: 84.79 F
23:31:42.852 -> Average Cal sensor Temp.: 96.89 Ambient Temp.: 84.76 F
23:31:44.413 -> Average Cal sensor Temp.: 96.82 Ambient Temp.: 84.79 F
23:31:45.970 -> Average Cal sensor Temp.: 96.84 Ambient Temp.: 84.87 F
23:31:47.533 -> Average Cal sensor Temp.: 96.85 Ambient Temp.: 84.87 F
23:31:49.093 -> Average Cal sensor Temp.: 96.92 Ambient Temp.: 84.87 F
23:31:50.648 -> Average Cal sensor Temp.: 96.86 Ambient Temp.: 84.90 F
23:31:52.205 -> Average Cal sensor Temp.: 96.97 Ambient Temp.: 84.90 F
23:31:53.766 -> Average Cal sensor Temp.: 96.90 Ambient Temp.: 84.94 F
23:31:55.324 -> Average Cal sensor Temp.: 96.85 Ambient Temp.: 84.94 F
23:31:56.883 -> Average Cal sensor Temp.: 96.91 Ambient Temp.: 84.94 F
23:31:58.443 -> Average Cal sensor Temp.: 96.77 Ambient Temp.: 84.97 F
23:31:59.969 -> Average Cal sensor Temp.: 96.89 Ambient Temp.: 85.01 F
23:32:01.532 -> Average Cal sensor Temp.: 96.56 Ambient Temp.: 85.01 F
23:32:03.085 -> Average Cal sensor Temp.: 96.64 Ambient Temp.: 85.05 F
23:32:04.646 -> Average Cal sensor Temp.: 96.72 Ambient Temp.: 85.05 F
23:32:06.203 -> Average Cal sensor Temp.: 96.98 Ambient Temp.: 85.08 F
23:32:07.756 -> Average Cal sensor Temp.: 96.95 Ambient Temp.: 85.05 F
23:32:09.314 -> Average Cal sensor Temp.: 96.89 Ambient Temp.: 85.08 F
23:32:10.875 -> Average Cal sensor Temp.: 96.78 Ambient Temp.: 85.08 F
23:32:12.429 -> Average Cal sensor Temp.: 96.80 Ambient Temp.: 85.12 F
23:32:13.988 -> Average Cal sensor Temp.: 96.69 Ambient Temp.: 85.12 F
23:32:15.547 -> Average Cal sensor Temp.: 96.53 Ambient Temp.: 85.15 F
23:32:17.098 -> Average Cal sensor Temp.: 96.74 Ambient Temp.: 85.19 F
23:32:18.652 -> Average Cal sensor Temp.: 96.85 Ambient Temp.: 85.19 F
23:32:20.207 -> Average Cal sensor Temp.: 96.74 Ambient Temp.: 85.19 F
23:32:21.763 -> Average Cal sensor Temp.: 96.77 Ambient Temp.: 85.26 F
23:32:23.318 -> Average Cal sensor Temp.: 96.96 Ambient Temp.: 85.30 F
23:32:24.877 -> Average Cal sensor Temp.: 97.00 Ambient Temp.: 85.26 F
23:32:26.436 -> Average Cal sensor Temp.: 97.11 Ambient Temp.: 85.37 F
23:32:27.998 -> Average Cal sensor Temp.: 97.16 Ambient Temp.: 85.33 F
23:32:29.557 -> Average Cal sensor Temp.: 97.09 Ambient Temp.: 85.33 F
23:32:31.115 -> Average Cal sensor Temp.: 97.41 Ambient Temp.: 85.37 F
23:32:32.678 -> Average Cal sensor Temp.: 97.15 Ambient Temp.: 85.48 F
23:32:34.236 -> Average Cal sensor Temp.: 97.07 Ambient Temp.: 85.41 F
23:32:35.765 -> Average Cal sensor Temp.: 97.21 Ambient Temp.: 85.48 F
23:32:37.322 -> Average Cal sensor Temp.: 97.11 Ambient Temp.: 85.48 F
23:32:38.880 -> Average Cal sensor Temp.: 97.16 Ambient Temp.: 85.55 F
23:32:40.435 -> Average Cal sensor Temp.: 97.03 Ambient Temp.: 85.51 F
23:32:41.990 -> Average Cal sensor Temp.: 96.77 Ambient Temp.: 85.55 F
23:32:43.546 -> Average Cal sensor Temp.: 96.96 Ambient Temp.: 85.55 F
23:32:45.102 -> Average Cal sensor Temp.: 96.81 Ambient Temp.: 85.55 F
23:32:46.661 -> Average Cal sensor Temp.: 96.93 Ambient Temp.: 85.66 F
23:32:48.221 -> Average Cal sensor Temp.: 97.02 Ambient Temp.: 85.69 F
23:32:49.783 -> Average Cal sensor Temp.: 97.15 Ambient Temp.: 85.69 F
23:32:51.343 -> Average Cal sensor Temp.: 97.10 Ambient Temp.: 85.73 F
23:32:52.900 -> Average Cal sensor Temp.: 97.00 Ambient Temp.: 85.73 F
23:32:54.428 -> Average Cal sensor Temp.: 97.03 Ambient Temp.: 85.69 F
23:32:55.988 -> Average Cal sensor Temp.: 96.88 Ambient Temp.: 85.69 F
23:32:57.544 -> Average Cal sensor Temp.: 96.70 Ambient Temp.: 85.73 F
23:32:59.108 -> Average Cal sensor Temp.: 96.73 Ambient Temp.: 85.73 F
23:33:00.669 -> Average Cal sensor Temp.: 96.87 Ambient Temp.: 85.77 F
23:33:02.226 -> Average Cal sensor Temp.: 96.60 Ambient Temp.: 85.77 F
23:33:03.784 -> Average Cal sensor Temp.: 96.69 Ambient Temp.: 85.73 F
23:33:05.340 -> Average Cal sensor Temp.: 96.83 Ambient Temp.: 85.80 F
23:33:06.888 -> Average Cal sensor Temp.: 96.77 Ambient Temp.: 85.77 F
23:33:08.445 -> Average Cal sensor Temp.: 96.61 Ambient Temp.: 85.77 F
23:33:10.008 -> Average Cal sensor Temp.: 96.87 Ambient Temp.: 85.84 F
23:33:11.564 -> Average Cal sensor Temp.: 96.66 Ambient Temp.: 85.80 F
23:33:13.125 -> Average Cal sensor Temp.: 96.74 Ambient Temp.: 85.80 F
23:33:14.689 -> Average Cal sensor Temp.: 96.71 Ambient Temp.: 85.80 F
23:33:16.217 -> Average Cal sensor Temp.: 96.51 Ambient Temp.: 85.77 F
23:33:17.778 -> Average Cal sensor Temp.: 96.54 Ambient Temp.: 85.80 F
23:33:19.329 -> Average Cal sensor Temp.: 96.76 Ambient Temp.: 85.80 F
23:33:20.889 -> Average Cal sensor Temp.: 96.72 Ambient Temp.: 85.80 F
23:33:22.446 -> Average Cal sensor Temp.: 96.85 Ambient Temp.: 85.84 F
23:33:24.013 -> Average Cal sensor Temp.: 97.02 Ambient Temp.: 85.84 F
23:33:25.571 -> Average Cal sensor Temp.: 96.97 Ambient Temp.: 85.80 F
23:33:27.125 -> Average Cal sensor Temp.: 97.22 Ambient Temp.: 85.80 F
23:33:28.684 -> Average Cal sensor Temp.: 97.20 Ambient Temp.: 85.84 F
23:33:30.240 -> Average Cal sensor Temp.: 96.78 Ambient Temp.: 85.84 F
23:33:31.797 -> Average Cal sensor Temp.: 96.76 Ambient Temp.: 85.80 F
23:33:33.361 -> Average Cal sensor Temp.: 96.76 Ambient Temp.: 85.87 F
23:33:34.917 -> Average Cal sensor Temp.: 96.93 Ambient Temp.: 85.87 F
23:33:36.476 -> Average Cal sensor Temp.: 96.93 Ambient Temp.: 85.84 F
23:33:38.033 -> Average Cal sensor Temp.: 96.93 Ambient Temp.: 85.91 F
23:33:39.593 -> Average Cal sensor Temp.: 96.97 Ambient Temp.: 85.80 F
23:33:41.151 -> Average Cal sensor Temp.: 96.54 Ambient Temp.: 85.84 F
23:33:42.707 -> Average Cal sensor Temp.: 96.79 Ambient Temp.: 85.87 F
23:33:44.267 -> Average Cal sensor Temp.: 96.86 Ambient Temp.: 85.87 F
23:33:45.797 -> Average Cal sensor Temp.: 96.97 Ambient Temp.: 85.84 F
23:33:47.353 -> Average Cal sensor Temp.: 96.89 Ambient Temp.: 85.80 F
23:33:48.920 -> Average Cal sensor Temp.: 96.97 Ambient Temp.: 85.80 F
23:33:50.476 -> Average Cal sensor Temp.: 97.10 Ambient Temp.: 85.87 F
23:33:52.035 -> Average Cal sensor Temp.: 96.83 Ambient Temp.: 85.84 F
23:33:53.592 -> Average Cal sensor Temp.: 97.22 Ambient Temp.: 85.87 F
23:33:55.154 -> Average Cal sensor Temp.: 96.80 Ambient Temp.: 85.84 F
23:33:56.713 -> Average Cal sensor Temp.: 97.02 Ambient Temp.: 85.84 F
23:33:58.278 -> Average Cal sensor Temp.: 97.00 Ambient Temp.: 85.91 F
23:33:59.834 -> Average Cal sensor Temp.: 96.98 Ambient Temp.: 85.87 F
23:34:01.395 -> Average Cal sensor Temp.: 97.18 Ambient Temp.: 85.87 F
23:34:02.951 -> Average Cal sensor Temp.: 97.24 Ambient Temp.: 85.91 F
23:34:04.509 -> Average Cal sensor Temp.: 96.91 Ambient Temp.: 85.91 F
23:34:06.036 -> Average Cal sensor Temp.: 96.77 Ambient Temp.: 85.87 F
23:34:07.590 -> Average Cal sensor Temp.: 96.97 Ambient Temp.: 85.87 F
23:34:09.152 -> Average Cal sensor Temp.: 97.06 Ambient Temp.: 85.87 F
23:34:10.713 -> Average Cal sensor Temp.: 96.57 Ambient Temp.: 85.87 F
23:34:12.267 -> Average Cal sensor Temp.: 96.69 Ambient Temp.: 85.84 F
23:34:13.827 -> Average Cal sensor Temp.: 96.90 Ambient Temp.: 85.84 F
23:34:15.388 -> Average Cal sensor Temp.: 97.09 Ambient Temp.: 85.84 F
23:34:16.943 -> Average Cal sensor Temp.: 96.91 Ambient Temp.: 85.87 F
23:34:18.506 -> Average Cal sensor Temp.: 96.90 Ambient Temp.: 85.87 F
23:34:20.069 -> Average Cal sensor Temp.: 97.27 Ambient Temp.: 85.87 F
23:34:21.627 -> Average Cal sensor Temp.: 97.25 Ambient Temp.: 85.91 F
23:34:23.188 -> Average Cal sensor Temp.: 97.03 Ambient Temp.: 85.87 F
23:34:24.713 -> Average Cal sensor Temp.: 97.01 Ambient Temp.: 85.87 F
23:34:26.270 -> Average Cal sensor Temp.: 97.16 Ambient Temp.: 85.87 F
23:34:27.825 -> Average Cal sensor Temp.: 97.07 Ambient Temp.: 85.87 F
23:34:29.385 -> Average Cal sensor Temp.: 97.14 Ambient Temp.: 85.84 F
23:34:30.946 -> Average Cal sensor Temp.: 97.17 Ambient Temp.: 85.84 F
23:34:32.507 -> Average Cal sensor Temp.: 97.07 Ambient Temp.: 85.87 F
23:34:34.067 -> Average Cal sensor Temp.: 96.89 Ambient Temp.: 85.84 F
23:34:35.630 -> Average Cal sensor Temp.: 97.12 Ambient Temp.: 85.87 F
23:34:37.184 -> Average Cal sensor Temp.: 96.64 Ambient Temp.: 85.84 F
23:34:38.747 -> Average Cal sensor Temp.: 96.58 Ambient Temp.: 85.77 F
23:34:40.271 -> Average Cal sensor Temp.: 96.92 Ambient Temp.: 85.84 F
23:34:41.832 -> Average Cal sensor Temp.: 96.93 Ambient Temp.: 85.80 F
23:34:43.395 -> Average Cal sensor Temp.: 96.89 Ambient Temp.: 85.77 F
23:34:44.959 -> Average Cal sensor Temp.: 97.03 Ambient Temp.: 85.77 F
23:34:46.515 -> Average Cal sensor Temp.: 96.53 Ambient Temp.: 85.77 F
23:34:48.073 -> Average Cal sensor Temp.: 96.90 Ambient Temp.: 85.77 F
23:34:49.634 -> Average Cal sensor Temp.: 96.68 Ambient Temp.: 85.77 F
23:34:51.190 -> Average Cal sensor Temp.: 96.50 Ambient Temp.: 85.73 F
23:34:52.748 -> Average Cal sensor Temp.: 96.57 Ambient Temp.: 85.69 F
23:34:54.302 -> Average Cal sensor Temp.: 96.62 Ambient Temp.: 85.69 F
23:34:55.862 -> Average Cal sensor Temp.: 96.72 Ambient Temp.: 85.69 F
23:34:57.420 -> Average Cal sensor Temp.: 96.65 Ambient Temp.: 85.69 F
23:34:58.976 -> Average Cal sensor Temp.: 96.83 Ambient Temp.: 85.59 F
23:35:00.538 -> Average Cal sensor Temp.: 96.58 Ambient Temp.: 85.59 F
23:35:02.092 -> Average Cal sensor Temp.: 96.57 Ambient Temp.: 85.59 F
23:35:03.622 -> Average Cal sensor Temp.: 96.63 Ambient Temp.: 85.59 F
23:35:05.179 -> Average Cal sensor Temp.: 96.58 Ambient Temp.: 85.51 F
23:35:06.744 -> Average Cal sensor Temp.: 96.24 Ambient Temp.: 85.51 F
23:35:08.305 -> Average Cal sensor Temp.: 96.11 Ambient Temp.: 85.48 F
23:35:09.865 -> Average Cal sensor Temp.: 96.79 Ambient Temp.: 85.48 F
23:35:11.423 -> Average Cal sensor Temp.: 96.64 Ambient Temp.: 85.44 F
23:35:12.983 -> Average Cal sensor Temp.: 96.78 Ambient Temp.: 85.41 F
23:35:14.542 -> Average Cal sensor Temp.: 96.71 Ambient Temp.: 85.41 F
23:35:16.091 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 85.37 F
23:35:17.647 -> Average Cal sensor Temp.: 96.90 Ambient Temp.: 85.41 F
23:35:19.205 -> Average Cal sensor Temp.: 96.03 Ambient Temp.: 85.33 F
23:35:20.761 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 85.33 F
23:35:22.321 -> Average Cal sensor Temp.: 96.43 Ambient Temp.: 85.26 F
23:35:23.880 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 85.15 F
23:35:25.435 -> Average Cal sensor Temp.: 96.91 Ambient Temp.: 85.26 F
23:35:26.990 -> Average Cal sensor Temp.: 97.00 Ambient Temp.: 85.19 F
23:35:28.545 -> Average Cal sensor Temp.: 97.04 Ambient Temp.: 85.15 F
23:35:30.100 -> Average Cal sensor Temp.: 96.88 Ambient Temp.: 85.19 F
23:35:31.659 -> Average Cal sensor Temp.: 96.48 Ambient Temp.: 85.19 F
23:35:33.218 -> Average Cal sensor Temp.: 97.09 Ambient Temp.: 85.15 F
23:35:34.742 -> Average Cal sensor Temp.: 96.84 Ambient Temp.: 85.19 F
23:35:36.325 -> Average Cal sensor Temp.: 96.72 Ambient Temp.: 85.08 F
23:35:37.885 -> Average Cal sensor Temp.: 96.64 Ambient Temp.: 85.08 F
23:35:39.444 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 85.12 F
23:35:40.969 -> Average Cal sensor Temp.: 96.76 Ambient Temp.: 85.08 F
23:35:42.529 -> Average Cal sensor Temp.: 96.94 Ambient Temp.: 85.05 F
23:35:44.088 -> Average Cal sensor Temp.: 97.24 Ambient Temp.: 85.08 F
23:35:45.649 -> Average Cal sensor Temp.: 96.83 Ambient Temp.: 85.05 F
23:35:47.208 -> Average Cal sensor Temp.: 96.82 Ambient Temp.: 85.05 F
23:35:48.769 -> Average Cal sensor Temp.: 96.71 Ambient Temp.: 85.01 F
23:35:50.319 -> Average Cal sensor Temp.: 96.80 Ambient Temp.: 84.97 F
23:35:51.873 -> Average Cal sensor Temp.: 96.62 Ambient Temp.: 85.01 F
23:35:53.431 -> Average Cal sensor Temp.: 96.52 Ambient Temp.: 84.94 F
23:35:54.990 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 84.90 F
23:35:56.552 -> Average Cal sensor Temp.: 96.88 Ambient Temp.: 84.87 F
23:35:58.113 -> Average Cal sensor Temp.: 96.99 Ambient Temp.: 84.87 F
23:35:59.671 -> Average Cal sensor Temp.: 97.04 Ambient Temp.: 84.90 F
23:36:01.226 -> Average Cal sensor Temp.: 96.71 Ambient Temp.: 84.79 F
23:36:02.786 -> Average Cal sensor Temp.: 96.60 Ambient Temp.: 84.79 F
23:36:04.344 -> Average Cal sensor Temp.: 96.62 Ambient Temp.: 84.72 F
23:36:05.902 -> Average Cal sensor Temp.: 96.79 Ambient Temp.: 84.72 F
23:36:07.458 -> Average Cal sensor Temp.: 97.01 Ambient Temp.: 84.72 F
23:36:09.014 -> Average Cal sensor Temp.: 97.43 Ambient Temp.: 84.69 F
23:36:10.569 -> Average Cal sensor Temp.: 97.13 Ambient Temp.: 84.72 F
23:36:12.121 -> Average Cal sensor Temp.: 97.19 Ambient Temp.: 84.72 F
23:36:13.677 -> Average Cal sensor Temp.: 97.51 Ambient Temp.: 84.76 F
23:36:15.233 -> Average Cal sensor Temp.: 97.36 Ambient Temp.: 84.69 F
23:36:16.782 -> Average Cal sensor Temp.: 97.18 Ambient Temp.: 84.69 F
23:36:18.337 -> Average Cal sensor Temp.: 96.72 Ambient Temp.: 84.61 F
23:36:19.899 -> Average Cal sensor Temp.: 97.21 Ambient Temp.: 84.65 F
23:36:21.456 -> Average Cal sensor Temp.: 97.33 Ambient Temp.: 84.65 F
23:36:23.018 -> Average Cal sensor Temp.: 97.33 Ambient Temp.: 84.72 F
23:36:24.551 -> Average Cal sensor Temp.: 97.03 Ambient Temp.: 84.65 F
23:36:26.114 -> Average Cal sensor Temp.: 97.12 Ambient Temp.: 84.65 F
23:36:27.675 -> Average Cal sensor Temp.: 97.12 Ambient Temp.: 84.58 F
23:36:29.232 -> Average Cal sensor Temp.: 97.20 Ambient Temp.: 84.61 F
23:36:30.792 -> Average Cal sensor Temp.: 97.22 Ambient Temp.: 84.61 F
23:36:32.345 -> Average Cal sensor Temp.: 97.12 Ambient Temp.: 84.61 F
23:36:33.903 -> Average Cal sensor Temp.: 97.31 Ambient Temp.: 84.61 F
23:36:35.462 -> Average Cal sensor Temp.: 97.21 Ambient Temp.: 84.65 F
23:36:37.011 -> Average Cal sensor Temp.: 96.97 Ambient Temp.: 84.58 F
23:36:38.566 -> Average Cal sensor Temp.: 96.87 Ambient Temp.: 84.58 F
23:36:40.123 -> Average Cal sensor Temp.: 97.10 Ambient Temp.: 84.58 F
23:36:41.682 -> Average Cal sensor Temp.: 97.10 Ambient Temp.: 84.65 F
23:36:43.236 -> Average Cal sensor Temp.: 97.18 Ambient Temp.: 84.72 F
23:36:44.789 -> Average Cal sensor Temp.: 97.31 Ambient Temp.: 84.69 F
23:36:46.342 -> Average Cal sensor Temp.: 97.27 Ambient Temp.: 84.65 F
23:36:47.894 -> Average Cal sensor Temp.: 97.56 Ambient Temp.: 84.65 F
23:36:49.451 -> Average Cal sensor Temp.: 97.16 Ambient Temp.: 84.65 F
23:36:51.010 -> Average Cal sensor Temp.: 97.13 Ambient Temp.: 84.69 F
23:36:52.563 -> Average Cal sensor Temp.: 97.19 Ambient Temp.: 84.69 F
23:36:54.117 -> Average Cal sensor Temp.: 97.23 Ambient Temp.: 84.72 F
23:36:55.674 -> Average Cal sensor Temp.: 97.35 Ambient Temp.: 84.69 F
23:36:57.237 -> Average Cal sensor Temp.: 97.20 Ambient Temp.: 84.72 F
23:36:58.799 -> Average Cal sensor Temp.: 97.21 Ambient Temp.: 84.76 F
23:37:00.349 -> Average Cal sensor Temp.: 97.06 Ambient Temp.: 84.76 F
23:37:01.901 -> Average Cal sensor Temp.: 97.21 Ambient Temp.: 84.76 F
23:37:03.455 -> Average Cal sensor Temp.: 97.09 Ambient Temp.: 84.72 F
23:37:05.012 -> Average Cal sensor Temp.: 97.21 Ambient Temp.: 84.76 F
23:37:06.574 -> Average Cal sensor Temp.: 97.56 Ambient Temp.: 84.79 F
23:37:08.134 -> Average Cal sensor Temp.: 97.50 Ambient Temp.: 84.87 F
23:37:09.662 -> Average Cal sensor Temp.: 97.16 Ambient Temp.: 84.87 F
23:37:11.216 -> Average Cal sensor Temp.: 97.21 Ambient Temp.: 84.87 F
23:37:12.799 -> Average Cal sensor Temp.: 97.01 Ambient Temp.: 84.90 F
23:37:14.362 -> Average Cal sensor Temp.: 97.42 Ambient Temp.: 84.97 F
23:37:15.890 -> Average Cal sensor Temp.: 97.36 Ambient Temp.: 84.94 F
23:37:17.448 -> Average Cal sensor Temp.: 96.92 Ambient Temp.: 84.90 F
23:37:19.002 -> Average Cal sensor Temp.: 96.91 Ambient Temp.: 84.97 F
23:37:20.560 -> Average Cal sensor Temp.: 97.13 Ambient Temp.: 84.94 F
23:37:22.112 -> Average Cal sensor Temp.: 97.40 Ambient Temp.: 84.97 F
23:37:23.701 -> Average Cal sensor Temp.: 97.48 Ambient Temp.: 84.97 F
23:37:25.225 -> Average Cal sensor Temp.: 97.36 Ambient Temp.: 85.01 F
23:37:26.780 -> Average Cal sensor Temp.: 97.27 Ambient Temp.: 85.05 F
23:37:28.336 -> Average Cal sensor Temp.: 97.11 Ambient Temp.: 84.97 F
23:37:29.896 -> Average Cal sensor Temp.: 97.45 Ambient Temp.: 85.05 F
23:37:31.457 -> Average Cal sensor Temp.: 97.52 Ambient Temp.: 85.05 F
23:37:33.010 -> Average Cal sensor Temp.: 97.31 Ambient Temp.: 85.05 F
23:37:34.572 -> Average Cal sensor Temp.: 97.29 Ambient Temp.: 85.05 F
23:37:36.131 -> Average Cal sensor Temp.: 97.30 Ambient Temp.: 85.08 F
23:37:37.684 -> Average Cal sensor Temp.: 96.84 Ambient Temp.: 85.12 F
23:37:39.245 -> Average Cal sensor Temp.: 97.15 Ambient Temp.: 85.08 F
23:37:40.800 -> Average Cal sensor Temp.: 97.29 Ambient Temp.: 85.08 F
23:37:42.356 -> Average Cal sensor Temp.: 97.36 Ambient Temp.: 85.08 F
23:37:43.916 -> Average Cal sensor Temp.: 97.53 Ambient Temp.: 85.15 F
23:37:45.473 -> Average Cal sensor Temp.: 97.29 Ambient Temp.: 85.12 F
23:37:47.024 -> Average Cal sensor Temp.: 97.20 Ambient Temp.: 85.15 F
23:37:48.584 -> Average Cal sensor Temp.: 97.12 Ambient Temp.: 85.12 F
23:37:50.115 -> Average Cal sensor Temp.: 96.99 Ambient Temp.: 85.12 F
23:37:51.674 -> Average Cal sensor Temp.: 97.21 Ambient Temp.: 85.12 F
23:37:53.236 -> Average Cal sensor Temp.: 97.26 Ambient Temp.: 85.19 F
23:37:54.792 -> Average Cal sensor Temp.: 97.30 Ambient Temp.: 85.19 F
23:37:56.348 -> Average Cal sensor Temp.: 97.41 Ambient Temp.: 85.19 F
23:37:57.904 -> Average Cal sensor Temp.: 97.02 Ambient Temp.: 85.26 F
23:37:59.459 -> Average Cal sensor Temp.: 97.20 Ambient Temp.: 85.19 F
23:38:01.021 -> Average Cal sensor Temp.: 97.24 Ambient Temp.: 85.15 F
23:38:02.573 -> Average Cal sensor Temp.: 97.21 Ambient Temp.: 85.26 F
23:38:04.130 -> Average Cal sensor Temp.: 97.29 Ambient Temp.: 85.33 F
23:38:05.683 -> Average Cal sensor Temp.: 97.34 Ambient Temp.: 85.30 F
23:38:07.235 -> Average Cal sensor Temp.: 97.07 Ambient Temp.: 85.30 F
23:38:08.794 -> Average Cal sensor Temp.: 97.24 Ambient Temp.: 85.33 F
23:38:10.353 -> Average Cal sensor Temp.: 97.38 Ambient Temp.: 85.33 F
23:38:11.911 -> Average Cal sensor Temp.: 97.27 Ambient Temp.: 85.37 F
23:38:13.435 -> Average Cal sensor Temp.: 97.43 Ambient Temp.: 85.41 F
23:38:14.989 -> Average Cal sensor Temp.: 97.21 Ambient Temp.: 85.41 F
23:38:16.546 -> Average Cal sensor Temp.: 97.34 Ambient Temp.: 85.44 F
23:38:18.105 -> Average Cal sensor Temp.: 97.33 Ambient Temp.: 85.44 F
23:38:19.665 -> Average Cal sensor Temp.: 97.32 Ambient Temp.: 85.44 F
23:38:21.219 -> Average Cal sensor Temp.: 97.47 Ambient Temp.: 85.48 F
23:38:22.777 -> Average Cal sensor Temp.: 97.33 Ambient Temp.: 85.44 F
23:38:24.334 -> Average Cal sensor Temp.: 97.26 Ambient Temp.: 85.51 F
23:38:25.894 -> Average Cal sensor Temp.: 97.38 Ambient Temp.: 85.51 F
23:38:27.453 -> Average Cal sensor Temp.: 97.69 Ambient Temp.: 85.55 F
23:38:29.013 -> Average Cal sensor Temp.: 97.75 Ambient Temp.: 85.59 F
23:38:30.572 -> Average Cal sensor Temp.: 97.69 Ambient Temp.: 85.59 F
23:38:32.124 -> Average Cal sensor Temp.: 97.49 Ambient Temp.: 85.66 F
23:38:33.686 -> Average Cal sensor Temp.: 97.47 Ambient Temp.: 85.59 F
23:38:35.212 -> Average Cal sensor Temp.: 97.43 Ambient Temp.: 85.69 F
23:38:36.769 -> Average Cal sensor Temp.: 97.20 Ambient Temp.: 85.59 F
23:38:38.325 -> Average Cal sensor Temp.: 97.30 Ambient Temp.: 85.69 F
23:38:39.912 -> Average Cal sensor Temp.: 97.42 Ambient Temp.: 85.69 F
23:38:41.436 -> Average Cal sensor Temp.: 97.34 Ambient Temp.: 85.69 F
23:38:42.995 -> Average Cal sensor Temp.: 97.29 Ambient Temp.: 85.73 F
23:38:44.559 -> Average Cal sensor Temp.: 97.40 Ambient Temp.: 85.73 F
23:38:46.119 -> Average Cal sensor Temp.: 97.34 Ambient Temp.: 85.80 F
23:38:47.681 -> Average Cal sensor Temp.: 97.26 Ambient Temp.: 85.73 F
23:38:49.242 -> Average Cal sensor Temp.: 97.54 Ambient Temp.: 85.84 F
23:38:50.770 -> Average Cal sensor Temp.: 97.53 Ambient Temp.: 85.80 F
23:38:52.328 -> Average Cal sensor Temp.: 97.47 Ambient Temp.: 85.87 F
23:38:53.884 -> Average Cal sensor Temp.: 97.39 Ambient Temp.: 85.87 F
23:38:55.445 -> Average Cal sensor Temp.: 97.42 Ambient Temp.: 85.87 F
23:38:56.996 -> Average Cal sensor Temp.: 97.40 Ambient Temp.: 85.87 F
23:38:58.550 -> Average Cal sensor Temp.: 97.52 Ambient Temp.: 85.87 F
23:39:00.114 -> Average Cal sensor Temp.: 97.70 Ambient Temp.: 85.95 F
23:39:01.669 -> Average Cal sensor Temp.: 97.54 Ambient Temp.: 85.95 F
23:39:03.225 -> Average Cal sensor Temp.: 97.57 Ambient Temp.: 85.95 F
23:39:04.787 -> Average Cal sensor Temp.: 97.57 Ambient Temp.: 85.95 F
23:39:06.343 -> Average Cal sensor Temp.: 97.51 Ambient Temp.: 85.98 F
23:39:07.905 -> Average Cal sensor Temp.: 97.55 Ambient Temp.: 85.98 F
23:39:09.436 -> Average Cal sensor Temp.: 97.51 Ambient Temp.: 85.98 F
23:39:10.994 -> Average Cal sensor Temp.: 97.64 Ambient Temp.: 86.05 F
23:39:12.558 -> Average Cal sensor Temp.: 97.44 Ambient Temp.: 86.13 F
23:39:14.119 -> Average Cal sensor Temp.: 97.31 Ambient Temp.: 86.09 F
23:39:15.674 -> Average Cal sensor Temp.: 97.34 Ambient Temp.: 86.09 F
23:39:17.225 -> Average Cal sensor Temp.: 97.66 Ambient Temp.: 86.13 F
23:39:18.786 -> Average Cal sensor Temp.: 97.65 Ambient Temp.: 86.20 F
23:39:20.344 -> Average Cal sensor Temp.: 97.31 Ambient Temp.: 86.16 F
23:39:21.898 -> Average Cal sensor Temp.: 97.19 Ambient Temp.: 86.20 F
23:39:23.459 -> Average Cal sensor Temp.: 97.16 Ambient Temp.: 86.20 F
23:39:25.011 -> Average Cal sensor Temp.: 97.36 Ambient Temp.: 86.20 F
23:39:26.564 -> Average Cal sensor Temp.: 97.72 Ambient Temp.: 86.27 F
23:39:28.116 -> Average Cal sensor Temp.: 97.85 Ambient Temp.: 86.27 F
23:39:29.680 -> Average Cal sensor Temp.: 97.80 Ambient Temp.: 86.27 F
23:39:31.238 -> Average Cal sensor Temp.: 97.56 Ambient Temp.: 86.27 F
23:39:32.796 -> Average Cal sensor Temp.: 97.58 Ambient Temp.: 86.31 F
23:39:34.360 -> Average Cal sensor Temp.: 97.87 Ambient Temp.: 86.31 F
23:39:35.918 -> Average Cal sensor Temp.: 97.51 Ambient Temp.: 86.34 F
23:39:37.439 -> Average Cal sensor Temp.: 97.65 Ambient Temp.: 86.34 F
23:39:39.001 -> Average Cal sensor Temp.: 97.37 Ambient Temp.: 86.34 F
23:39:40.560 -> Average Cal sensor Temp.: 97.41 Ambient Temp.: 86.31 F
23:39:42.119 -> Average Cal sensor Temp.: 97.33 Ambient Temp.: 86.23 F
23:39:43.673 -> Average Cal sensor Temp.: 97.54 Ambient Temp.: 86.27 F
23:39:45.229 -> Average Cal sensor Temp.: 97.23 Ambient Temp.: 86.34 F
23:39:46.787 -> Average Cal sensor Temp.: 97.40 Ambient Temp.: 86.34 F
23:39:48.344 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 86.31 F
23:39:49.903 -> Average Cal sensor Temp.: 97.74 Ambient Temp.: 86.38 F
23:39:51.457 -> Average Cal sensor Temp.: 97.62 Ambient Temp.: 86.31 F
23:39:53.011 -> Average Cal sensor Temp.: 97.48 Ambient Temp.: 86.31 F
23:39:54.569 -> Average Cal sensor Temp.: 97.49 Ambient Temp.: 86.31 F
23:39:56.124 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 86.34 F
23:39:57.679 -> Average Cal sensor Temp.: 97.38 Ambient Temp.: 86.38 F
23:39:59.235 -> Average Cal sensor Temp.: 96.97 Ambient Temp.: 86.34 F
23:40:00.788 -> Average Cal sensor Temp.: 97.43 Ambient Temp.: 86.34 F
23:40:02.345 -> Average Cal sensor Temp.: 97.35 Ambient Temp.: 86.31 F
23:40:03.905 -> Average Cal sensor Temp.: 97.31 Ambient Temp.: 86.31 F
23:40:05.466 -> Average Cal sensor Temp.: 97.62 Ambient Temp.: 86.31 F
23:40:07.020 -> Average Cal sensor Temp.: 97.40 Ambient Temp.: 86.31 F
23:40:08.578 -> Average Cal sensor Temp.: 97.23 Ambient Temp.: 86.31 F
23:40:10.135 -> Average Cal sensor Temp.: 97.45 Ambient Temp.: 86.38 F
23:40:11.697 -> Average Cal sensor Temp.: 97.39 Ambient Temp.: 86.27 F
23:40:13.220 -> Average Cal sensor Temp.: 96.86 Ambient Temp.: 86.20 F
23:40:14.780 -> Average Cal sensor Temp.: 96.80 Ambient Temp.: 86.20 F
23:40:16.336 -> Average Cal sensor Temp.: 96.95 Ambient Temp.: 86.16 F
23:40:17.897 -> Average Cal sensor Temp.: 97.42 Ambient Temp.: 86.13 F
23:40:19.461 -> Average Cal sensor Temp.: 97.41 Ambient Temp.: 86.16 F
23:40:21.021 -> Average Cal sensor Temp.: 97.27 Ambient Temp.: 86.16 F
23:40:22.572 -> Average Cal sensor Temp.: 97.36 Ambient Temp.: 86.13 F
23:40:24.131 -> Average Cal sensor Temp.: 97.15 Ambient Temp.: 86.13 F
23:40:25.690 -> Average Cal sensor Temp.: 97.29 Ambient Temp.: 86.09 F
23:40:27.250 -> Average Cal sensor Temp.: 96.91 Ambient Temp.: 86.09 F
23:40:28.809 -> Average Cal sensor Temp.: 96.90 Ambient Temp.: 86.09 F
23:40:30.372 -> Average Cal sensor Temp.: 96.66 Ambient Temp.: 85.95 F
23:40:31.922 -> Average Cal sensor Temp.: 97.15 Ambient Temp.: 85.98 F
23:40:33.473 -> Average Cal sensor Temp.: 96.92 Ambient Temp.: 85.95 F
23:40:35.032 -> Average Cal sensor Temp.: 97.00 Ambient Temp.: 85.91 F
23:40:36.593 -> Average Cal sensor Temp.: 97.38 Ambient Temp.: 85.95 F
23:40:38.149 -> Average Cal sensor Temp.: 97.36 Ambient Temp.: 85.91 F
23:40:39.709 -> Average Cal sensor Temp.: 97.48 Ambient Temp.: 85.91 F
23:40:41.259 -> Average Cal sensor Temp.: 97.64 Ambient Temp.: 85.87 F
23:40:42.814 -> Average Cal sensor Temp.: 97.54 Ambient Temp.: 85.91 F
23:40:44.368 -> Average Cal sensor Temp.: 97.67 Ambient Temp.: 85.91 F
23:40:45.922 -> Average Cal sensor Temp.: 97.47 Ambient Temp.: 85.87 F
23:40:47.477 -> Average Cal sensor Temp.: 97.49 Ambient Temp.: 85.84 F
23:40:49.034 -> Average Cal sensor Temp.: 97.24 Ambient Temp.: 85.84 F
23:40:50.597 -> Average Cal sensor Temp.: 97.48 Ambient Temp.: 85.80 F
23:40:52.154 -> Average Cal sensor Temp.: 97.38 Ambient Temp.: 85.84 F
23:40:53.709 -> Average Cal sensor Temp.: 97.75 Ambient Temp.: 85.80 F
23:40:55.265 -> Average Cal sensor Temp.: 97.23 Ambient Temp.: 85.77 F
23:40:56.827 -> Average Cal sensor Temp.: 97.24 Ambient Temp.: 85.73 F
23:40:58.353 -> Average Cal sensor Temp.: 96.89 Ambient Temp.: 85.73 F
23:40:59.915 -> Average Cal sensor Temp.: 97.34 Ambient Temp.: 85.66 F
23:41:01.477 -> Average Cal sensor Temp.: 97.33 Ambient Temp.: 85.66 F
23:41:03.033 -> Average Cal sensor Temp.: 97.40 Ambient Temp.: 85.66 F
23:41:04.594 -> Average Cal sensor Temp.: 97.60 Ambient Temp.: 85.69 F
23:41:06.151 -> Average Cal sensor Temp.: 97.16 Ambient Temp.: 85.69 F
23:41:07.710 -> Average Cal sensor Temp.: 97.20 Ambient Temp.: 85.66 F
23:41:09.272 -> Average Cal sensor Temp.: 97.36 Ambient Temp.: 85.59 F
23:41:10.826 -> Average Cal sensor Temp.: 97.25 Ambient Temp.: 85.51 F
23:41:12.385 -> Average Cal sensor Temp.: 97.18 Ambient Temp.: 85.55 F
23:41:13.941 -> Average Cal sensor Temp.: 97.35 Ambient Temp.: 85.55 F
23:41:15.467 -> Average Cal sensor Temp.: 96.79 Ambient Temp.: 85.48 F
23:41:17.028 -> Average Cal sensor Temp.: 97.17 Ambient Temp.: 85.48 F
23:41:18.582 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 85.48 F
23:41:20.141 -> Average Cal sensor Temp.: 97.59 Ambient Temp.: 85.48 F
23:41:21.697 -> Average Cal sensor Temp.: 97.48 Ambient Temp.: 85.44 F
23:41:23.250 -> Average Cal sensor Temp.: 97.10 Ambient Temp.: 85.37 F
23:41:24.805 -> Average Cal sensor Temp.: 97.04 Ambient Temp.: 85.41 F
23:41:26.359 -> Average Cal sensor Temp.: 97.20 Ambient Temp.: 85.37 F
23:41:27.911 -> Average Cal sensor Temp.: 97.38 Ambient Temp.: 85.37 F
23:41:29.468 -> Average Cal sensor Temp.: 97.49 Ambient Temp.: 85.30 F
23:41:31.027 -> Average Cal sensor Temp.: 97.13 Ambient Temp.: 85.30 F
23:41:32.584 -> Average Cal sensor Temp.: 97.60 Ambient Temp.: 85.33 F
23:41:34.137 -> Average Cal sensor Temp.: 97.19 Ambient Temp.: 85.26 F
23:41:35.694 -> Average Cal sensor Temp.: 97.52 Ambient Temp.: 85.26 F
23:41:37.248 -> Average Cal sensor Temp.: 97.33 Ambient Temp.: 85.19 F
23:41:38.808 -> Average Cal sensor Temp.: 97.49 Ambient Temp.: 85.19 F
23:41:40.358 -> Average Cal sensor Temp.: 97.85 Ambient Temp.: 85.26 F
23:41:41.912 -> Average Cal sensor Temp.: 97.31 Ambient Temp.: 85.26 F
23:41:43.468 -> Average Cal sensor Temp.: 97.16 Ambient Temp.: 85.19 F
23:41:45.030 -> Average Cal sensor Temp.: 97.27 Ambient Temp.: 85.15 F
23:41:46.592 -> Average Cal sensor Temp.: 96.94 Ambient Temp.: 85.12 F
23:41:48.152 -> Average Cal sensor Temp.: 97.14 Ambient Temp.: 85.12 F
23:41:49.713 -> Average Cal sensor Temp.: 97.41 Ambient Temp.: 85.08 F
23:41:51.241 -> Average Cal sensor Temp.: 97.38 Ambient Temp.: 85.12 F
23:41:52.799 -> Average Cal sensor Temp.: 97.23 Ambient Temp.: 85.08 F
23:41:54.354 -> Average Cal sensor Temp.: 97.11 Ambient Temp.: 85.05 F
23:41:55.907 -> Average Cal sensor Temp.: 97.21 Ambient Temp.: 84.97 F
23:41:57.464 -> Average Cal sensor Temp.: 97.54 Ambient Temp.: 85.01 F
23:41:59.026 -> Average Cal sensor Temp.: 97.11 Ambient Temp.: 84.97 F
23:42:00.583 -> Average Cal sensor Temp.: 97.35 Ambient Temp.: 84.97 F
23:42:02.139 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 84.97 F
23:42:03.695 -> Average Cal sensor Temp.: 97.16 Ambient Temp.: 84.97 F
23:42:05.252 -> Average Cal sensor Temp.: 96.80 Ambient Temp.: 84.90 F
23:42:06.809 -> Average Cal sensor Temp.: 96.85 Ambient Temp.: 84.90 F
23:42:08.364 -> Average Cal sensor Temp.: 96.88 Ambient Temp.: 84.79 F
23:42:09.926 -> Average Cal sensor Temp.: 97.08 Ambient Temp.: 84.76 F
23:42:11.485 -> Average Cal sensor Temp.: 97.41 Ambient Temp.: 84.76 F
23:42:13.042 -> Average Cal sensor Temp.: 97.11 Ambient Temp.: 84.76 F
23:42:14.600 -> Average Cal sensor Temp.: 97.41 Ambient Temp.: 84.72 F
23:42:16.157 -> Average Cal sensor Temp.: 97.42 Ambient Temp.: 84.72 F
23:42:17.713 -> Average Cal sensor Temp.: 97.73 Ambient Temp.: 84.69 F
23:42:19.268 -> Average Cal sensor Temp.: 97.71 Ambient Temp.: 84.72 F
23:42:20.797 -> Average Cal sensor Temp.: 97.42 Ambient Temp.: 84.65 F
23:42:22.355 -> Average Cal sensor Temp.: 96.90 Ambient Temp.: 84.61 F
23:42:23.911 -> Average Cal sensor Temp.: 97.03 Ambient Temp.: 84.65 F
23:42:25.470 -> Average Cal sensor Temp.: 97.24 Ambient Temp.: 84.58 F
23:42:27.029 -> Average Cal sensor Temp.: 97.23 Ambient Temp.: 84.58 F
23:42:28.587 -> Average Cal sensor Temp.: 97.50 Ambient Temp.: 84.58 F
23:42:30.142 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 84.58 F
23:42:31.703 -> Average Cal sensor Temp.: 97.85 Ambient Temp.: 84.58 F
23:42:33.263 -> Average Cal sensor Temp.: 97.17 Ambient Temp.: 84.54 F
23:42:34.825 -> Average Cal sensor Temp.: 97.52 Ambient Temp.: 84.51 F
23:42:36.378 -> Average Cal sensor Temp.: 97.11 Ambient Temp.: 84.54 F
23:42:37.932 -> Average Cal sensor Temp.: 97.24 Ambient Temp.: 84.51 F
23:42:39.489 -> Average Cal sensor Temp.: 97.38 Ambient Temp.: 84.47 F
23:42:41.043 -> Average Cal sensor Temp.: 97.21 Ambient Temp.: 84.47 F
23:42:42.595 -> Average Cal sensor Temp.: 97.18 Ambient Temp.: 84.36 F
23:42:44.154 -> Average Cal sensor Temp.: 97.41 Ambient Temp.: 84.40 F
23:42:45.714 -> Average Cal sensor Temp.: 97.16 Ambient Temp.: 84.33 F
23:42:47.271 -> Average Cal sensor Temp.: 97.33 Ambient Temp.: 84.33 F
23:42:48.799 -> Average Cal sensor Temp.: 97.34 Ambient Temp.: 84.29 F
23:42:50.361 -> Average Cal sensor Temp.: 96.96 Ambient Temp.: 84.25 F
23:42:51.914 -> Average Cal sensor Temp.: 97.29 Ambient Temp.: 84.22 F
23:42:53.470 -> Average Cal sensor Temp.: 97.56 Ambient Temp.: 84.29 F
23:42:55.025 -> Average Cal sensor Temp.: 97.40 Ambient Temp.: 84.29 F
23:42:56.611 -> Average Cal sensor Temp.: 97.33 Ambient Temp.: 84.25 F
23:42:58.134 -> Average Cal sensor Temp.: 97.69 Ambient Temp.: 84.25 F
23:42:59.723 -> Average Cal sensor Temp.: 97.37 Ambient Temp.: 84.25 F
23:43:01.247 -> Average Cal sensor Temp.: 97.52 Ambient Temp.: 84.25 F
23:43:02.804 -> Average Cal sensor Temp.: 97.38 Ambient Temp.: 84.22 F
23:43:04.362 -> Average Cal sensor Temp.: 97.31 Ambient Temp.: 84.22 F
23:43:05.916 -> Average Cal sensor Temp.: 97.16 Ambient Temp.: 84.22 F
23:43:07.474 -> Average Cal sensor Temp.: 96.88 Ambient Temp.: 84.22 F
23:43:09.036 -> Average Cal sensor Temp.: 97.19 Ambient Temp.: 84.18 F
23:43:10.589 -> Average Cal sensor Temp.: 97.39 Ambient Temp.: 84.22 F
23:43:12.143 -> Average Cal sensor Temp.: 97.48 Ambient Temp.: 84.15 F
23:43:13.703 -> Average Cal sensor Temp.: 97.70 Ambient Temp.: 84.22 F
23:43:15.261 -> Average Cal sensor Temp.: 97.21 Ambient Temp.: 84.22 F
23:43:16.821 -> Average Cal sensor Temp.: 97.36 Ambient Temp.: 84.22 F
23:43:18.380 -> Average Cal sensor Temp.: 97.44 Ambient Temp.: 84.18 F
23:43:19.933 -> Average Cal sensor Temp.: 97.63 Ambient Temp.: 84.15 F
23:43:21.483 -> Average Cal sensor Temp.: 97.47 Ambient Temp.: 84.25 F
23:43:23.042 -> Average Cal sensor Temp.: 97.04 Ambient Temp.: 84.18 F
23:43:24.600 -> Average Cal sensor Temp.: 97.34 Ambient Temp.: 84.25 F
23:43:26.162 -> Average Cal sensor Temp.: 97.60 Ambient Temp.: 84.22 F
23:43:27.724 -> Average Cal sensor Temp.: 96.93 Ambient Temp.: 84.25 F
23:43:29.281 -> Average Cal sensor Temp.: 97.16 Ambient Temp.: 84.18 F
23:43:30.807 -> Average Cal sensor Temp.: 96.97 Ambient Temp.: 84.29 F
23:43:32.367 -> Average Cal sensor Temp.: 97.33 Ambient Temp.: 84.25 F
23:43:33.928 -> Average Cal sensor Temp.: 97.53 Ambient Temp.: 84.25 F
23:43:35.489 -> Average Cal sensor Temp.: 97.59 Ambient Temp.: 84.25 F
23:43:37.051 -> Average Cal sensor Temp.: 97.41 Ambient Temp.: 84.29 F
23:43:38.613 -> Average Cal sensor Temp.: 97.49 Ambient Temp.: 84.33 F
23:43:40.140 -> Average Cal sensor Temp.: 97.56 Ambient Temp.: 84.33 F
23:43:41.701 -> Average Cal sensor Temp.: 97.50 Ambient Temp.: 84.36 F
23:43:43.265 -> Average Cal sensor Temp.: 97.39 Ambient Temp.: 84.36 F
23:43:44.826 -> Average Cal sensor Temp.: 97.45 Ambient Temp.: 84.47 F
23:43:46.382 -> Average Cal sensor Temp.: 97.82 Ambient Temp.: 84.47 F
23:43:47.933 -> Average Cal sensor Temp.: 97.67 Ambient Temp.: 84.40 F
23:43:49.492 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 84.51 F
23:43:51.051 -> Average Cal sensor Temp.: 97.85 Ambient Temp.: 84.51 F
23:43:52.610 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 84.65 F
23:43:54.170 -> Average Cal sensor Temp.: 97.57 Ambient Temp.: 84.58 F
23:43:55.698 -> Average Cal sensor Temp.: 97.68 Ambient Temp.: 84.65 F
23:43:57.260 -> Average Cal sensor Temp.: 97.48 Ambient Temp.: 84.65 F
23:43:58.818 -> Average Cal sensor Temp.: 97.58 Ambient Temp.: 84.69 F
23:44:00.377 -> Average Cal sensor Temp.: 97.42 Ambient Temp.: 84.69 F
23:44:01.935 -> Average Cal sensor Temp.: 97.93 Ambient Temp.: 84.72 F
23:44:03.494 -> Average Cal sensor Temp.: 97.77 Ambient Temp.: 84.72 F
23:44:05.058 -> Average Cal sensor Temp.: 97.50 Ambient Temp.: 84.76 F
23:44:06.618 -> Average Cal sensor Temp.: 97.59 Ambient Temp.: 84.76 F
23:44:08.178 -> Average Cal sensor Temp.: 97.67 Ambient Temp.: 84.87 F
23:44:09.734 -> Average Cal sensor Temp.: 97.85 Ambient Temp.: 84.87 F
23:44:11.261 -> Average Cal sensor Temp.: 97.57 Ambient Temp.: 84.87 F
23:44:12.849 -> Average Cal sensor Temp.: 97.68 Ambient Temp.: 84.90 F
23:44:14.403 -> Average Cal sensor Temp.: 98.00 Ambient Temp.: 84.90 F
23:44:15.964 -> Average Cal sensor Temp.: 98.08 Ambient Temp.: 84.97 F
23:44:17.493 -> Average Cal sensor Temp.: 97.93 Ambient Temp.: 85.01 F
23:44:19.050 -> Average Cal sensor Temp.: 97.84 Ambient Temp.: 84.97 F
23:44:20.613 -> Average Cal sensor Temp.: 97.57 Ambient Temp.: 84.97 F
23:44:22.166 -> Average Cal sensor Temp.: 97.82 Ambient Temp.: 85.01 F
23:44:23.724 -> Average Cal sensor Temp.: 97.69 Ambient Temp.: 85.01 F
23:44:25.286 -> Average Cal sensor Temp.: 97.43 Ambient Temp.: 85.05 F
23:44:26.849 -> Average Cal sensor Temp.: 97.28 Ambient Temp.: 85.01 F
23:44:28.405 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 85.08 F
23:44:29.960 -> Average Cal sensor Temp.: 97.58 Ambient Temp.: 85.08 F
23:44:31.519 -> Average Cal sensor Temp.: 97.55 Ambient Temp.: 85.12 F
23:44:33.078 -> Average Cal sensor Temp.: 97.72 Ambient Temp.: 85.15 F
23:44:34.601 -> Average Cal sensor Temp.: 97.38 Ambient Temp.: 85.12 F
23:44:36.166 -> Average Cal sensor Temp.: 97.43 Ambient Temp.: 85.15 F
23:44:37.726 -> Average Cal sensor Temp.: 97.10 Ambient Temp.: 85.19 F
23:44:39.291 -> Average Cal sensor Temp.: 97.37 Ambient Temp.: 85.19 F
23:44:40.853 -> Average Cal sensor Temp.: 97.42 Ambient Temp.: 85.19 F
23:44:42.377 -> Average Cal sensor Temp.: 97.44 Ambient Temp.: 85.19 F
23:44:43.935 -> Average Cal sensor Temp.: 97.49 Ambient Temp.: 85.26 F
23:44:45.498 -> Average Cal sensor Temp.: 97.59 Ambient Temp.: 85.26 F
23:44:47.052 -> Average Cal sensor Temp.: 97.82 Ambient Temp.: 85.33 F
23:44:48.613 -> Average Cal sensor Temp.: 97.47 Ambient Temp.: 85.37 F
23:44:50.170 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 85.37 F
23:44:51.721 -> Average Cal sensor Temp.: 97.87 Ambient Temp.: 85.33 F
23:44:53.278 -> Average Cal sensor Temp.: 97.84 Ambient Temp.: 85.41 F
23:44:54.839 -> Average Cal sensor Temp.: 97.51 Ambient Temp.: 85.41 F
23:44:56.395 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 85.37 F
23:44:57.955 -> Average Cal sensor Temp.: 97.70 Ambient Temp.: 85.41 F
23:44:59.484 -> Average Cal sensor Temp.: 97.78 Ambient Temp.: 85.41 F
23:45:01.046 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 85.44 F
23:45:02.607 -> Average Cal sensor Temp.: 97.97 Ambient Temp.: 85.44 F
23:45:04.167 -> Average Cal sensor Temp.: 97.94 Ambient Temp.: 85.51 F
23:45:05.725 -> Average Cal sensor Temp.: 97.60 Ambient Temp.: 85.51 F
23:45:07.288 -> Average Cal sensor Temp.: 97.34 Ambient Temp.: 85.55 F
23:45:08.816 -> Average Cal sensor Temp.: 97.49 Ambient Temp.: 85.59 F
23:45:10.377 -> Average Cal sensor Temp.: 97.31 Ambient Temp.: 85.55 F
23:45:11.936 -> Average Cal sensor Temp.: 97.54 Ambient Temp.: 85.55 F
23:45:13.497 -> Average Cal sensor Temp.: 97.59 Ambient Temp.: 85.59 F
23:45:15.056 -> Average Cal sensor Temp.: 97.58 Ambient Temp.: 85.66 F
23:45:16.618 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 85.69 F
23:45:18.180 -> Average Cal sensor Temp.: 97.55 Ambient Temp.: 85.69 F
23:45:19.707 -> Average Cal sensor Temp.: 97.51 Ambient Temp.: 85.73 F
23:45:21.264 -> Average Cal sensor Temp.: 97.38 Ambient Temp.: 85.77 F
23:45:22.820 -> Average Cal sensor Temp.: 97.69 Ambient Temp.: 85.77 F
23:45:24.378 -> Average Cal sensor Temp.: 97.73 Ambient Temp.: 85.77 F
23:45:25.942 -> Average Cal sensor Temp.: 98.08 Ambient Temp.: 85.84 F
23:45:27.497 -> Average Cal sensor Temp.: 98.13 Ambient Temp.: 85.84 F
23:45:29.060 -> Average Cal sensor Temp.: 97.76 Ambient Temp.: 85.80 F
23:45:30.622 -> Average Cal sensor Temp.: 97.67 Ambient Temp.: 85.87 F
23:45:32.180 -> Average Cal sensor Temp.: 97.56 Ambient Temp.: 85.91 F
23:45:33.731 -> Average Cal sensor Temp.: 97.43 Ambient Temp.: 85.95 F
23:45:35.290 -> Average Cal sensor Temp.: 97.56 Ambient Temp.: 85.95 F
23:45:36.845 -> Average Cal sensor Temp.: 97.46 Ambient Temp.: 85.95 F
23:45:38.404 -> Average Cal sensor Temp.: 97.41 Ambient Temp.: 85.98 F
23:45:39.930 -> Average Cal sensor Temp.: 97.59 Ambient Temp.: 85.98 F
23:45:41.488 -> Average Cal sensor Temp.: 97.65 Ambient Temp.: 86.05 F
23:45:43.050 -> Average Cal sensor Temp.: 97.81 Ambient Temp.: 86.09 F
23:45:44.608 -> Average Cal sensor Temp.: 97.94 Ambient Temp.: 86.05 F
23:45:46.160 -> Average Cal sensor Temp.: 97.71 Ambient Temp.: 86.16 F
23:45:47.721 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 86.13 F
23:45:49.281 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 86.16 F
23:45:50.834 -> Average Cal sensor Temp.: 97.84 Ambient Temp.: 86.20 F
23:45:52.393 -> Average Cal sensor Temp.: 97.81 Ambient Temp.: 86.20 F
23:45:53.955 -> Average Cal sensor Temp.: 97.65 Ambient Temp.: 86.20 F
23:45:55.506 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 86.20 F
23:45:57.062 -> Average Cal sensor Temp.: 97.73 Ambient Temp.: 86.23 F
23:45:58.620 -> Average Cal sensor Temp.: 97.46 Ambient Temp.: 86.20 F
23:46:00.181 -> Average Cal sensor Temp.: 97.66 Ambient Temp.: 86.23 F
23:46:01.739 -> Average Cal sensor Temp.: 97.82 Ambient Temp.: 86.20 F
23:46:03.298 -> Average Cal sensor Temp.: 97.39 Ambient Temp.: 86.23 F
23:46:04.852 -> Average Cal sensor Temp.: 97.56 Ambient Temp.: 86.27 F
23:46:06.405 -> Average Cal sensor Temp.: 97.45 Ambient Temp.: 86.27 F
23:46:07.955 -> Average Cal sensor Temp.: 97.20 Ambient Temp.: 86.27 F
23:46:09.520 -> Average Cal sensor Temp.: 97.69 Ambient Temp.: 86.23 F
23:46:11.086 -> Average Cal sensor Temp.: 97.67 Ambient Temp.: 86.31 F
23:46:12.609 -> Average Cal sensor Temp.: 97.75 Ambient Temp.: 86.31 F
23:46:14.168 -> Average Cal sensor Temp.: 97.77 Ambient Temp.: 86.27 F
23:46:15.726 -> Average Cal sensor Temp.: 97.73 Ambient Temp.: 86.31 F
23:46:17.287 -> Average Cal sensor Temp.: 97.38 Ambient Temp.: 86.31 F
23:46:18.849 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 86.31 F
23:46:20.403 -> Average Cal sensor Temp.: 97.40 Ambient Temp.: 86.31 F
23:46:21.960 -> Average Cal sensor Temp.: 97.75 Ambient Temp.: 86.34 F
23:46:23.518 -> Average Cal sensor Temp.: 97.83 Ambient Temp.: 86.34 F
23:46:25.078 -> Average Cal sensor Temp.: 98.08 Ambient Temp.: 86.38 F
23:46:26.635 -> Average Cal sensor Temp.: 97.73 Ambient Temp.: 86.38 F
23:46:28.188 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 86.31 F
23:46:29.747 -> Average Cal sensor Temp.: 98.08 Ambient Temp.: 86.38 F
23:46:31.295 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 86.34 F
23:46:32.849 -> Average Cal sensor Temp.: 97.87 Ambient Temp.: 86.38 F
23:46:34.413 -> Average Cal sensor Temp.: 97.68 Ambient Temp.: 86.38 F
23:46:35.975 -> Average Cal sensor Temp.: 97.47 Ambient Temp.: 86.34 F
23:46:37.539 -> Average Cal sensor Temp.: 97.77 Ambient Temp.: 86.38 F
23:46:39.069 -> Average Cal sensor Temp.: 97.72 Ambient Temp.: 86.38 F
23:46:40.629 -> Average Cal sensor Temp.: 97.47 Ambient Temp.: 86.38 F
23:46:42.188 -> Average Cal sensor Temp.: 98.10 Ambient Temp.: 86.45 F
23:46:43.748 -> Average Cal sensor Temp.: 97.94 Ambient Temp.: 86.38 F
23:46:45.307 -> Average Cal sensor Temp.: 97.62 Ambient Temp.: 86.45 F
23:46:46.869 -> Average Cal sensor Temp.: 97.39 Ambient Temp.: 86.45 F
23:46:48.428 -> Average Cal sensor Temp.: 97.46 Ambient Temp.: 86.45 F
23:46:49.989 -> Average Cal sensor Temp.: 97.37 Ambient Temp.: 86.34 F
23:46:51.543 -> Average Cal sensor Temp.: 97.51 Ambient Temp.: 86.31 F
23:46:53.067 -> Average Cal sensor Temp.: 97.73 Ambient Temp.: 86.34 F
23:46:54.630 -> Average Cal sensor Temp.: 97.58 Ambient Temp.: 86.34 F
23:46:56.189 -> Average Cal sensor Temp.: 97.60 Ambient Temp.: 86.31 F
23:46:57.745 -> Average Cal sensor Temp.: 97.76 Ambient Temp.: 86.45 F
23:46:59.299 -> Average Cal sensor Temp.: 97.75 Ambient Temp.: 86.34 F
23:47:00.852 -> Average Cal sensor Temp.: 97.66 Ambient Temp.: 86.31 F
23:47:02.409 -> Average Cal sensor Temp.: 97.84 Ambient Temp.: 86.31 F
23:47:03.968 -> Average Cal sensor Temp.: 97.69 Ambient Temp.: 86.31 F
23:47:05.527 -> Average Cal sensor Temp.: 97.43 Ambient Temp.: 86.27 F
23:47:07.089 -> Average Cal sensor Temp.: 97.57 Ambient Temp.: 86.23 F
23:47:08.651 -> Average Cal sensor Temp.: 97.93 Ambient Temp.: 86.27 F
23:47:10.180 -> Average Cal sensor Temp.: 97.62 Ambient Temp.: 86.27 F
23:47:11.739 -> Average Cal sensor Temp.: 97.55 Ambient Temp.: 86.27 F
23:47:13.296 -> Average Cal sensor Temp.: 97.36 Ambient Temp.: 86.20 F
23:47:14.853 -> Average Cal sensor Temp.: 97.38 Ambient Temp.: 86.20 F
23:47:16.409 -> Average Cal sensor Temp.: 97.49 Ambient Temp.: 86.13 F
23:47:17.961 -> Average Cal sensor Temp.: 98.09 Ambient Temp.: 86.13 F
23:47:19.521 -> Average Cal sensor Temp.: 97.72 Ambient Temp.: 86.16 F
23:47:21.075 -> Average Cal sensor Temp.: 97.78 Ambient Temp.: 86.13 F
23:47:22.631 -> Average Cal sensor Temp.: 97.96 Ambient Temp.: 86.16 F
23:47:24.185 -> Average Cal sensor Temp.: 97.68 Ambient Temp.: 86.13 F
23:47:25.744 -> Average Cal sensor Temp.: 97.89 Ambient Temp.: 86.13 F
23:47:27.300 -> Average Cal sensor Temp.: 97.98 Ambient Temp.: 86.13 F
23:47:28.857 -> Average Cal sensor Temp.: 97.83 Ambient Temp.: 86.13 F
23:47:30.421 -> Average Cal sensor Temp.: 98.07 Ambient Temp.: 86.13 F
23:47:31.977 -> Average Cal sensor Temp.: 97.20 Ambient Temp.: 86.09 F
23:47:33.534 -> Average Cal sensor Temp.: 97.76 Ambient Temp.: 86.05 F
23:47:35.092 -> Average Cal sensor Temp.: 98.02 Ambient Temp.: 86.09 F
23:47:36.650 -> Average Cal sensor Temp.: 97.70 Ambient Temp.: 86.05 F
23:47:38.205 -> Average Cal sensor Temp.: 97.67 Ambient Temp.: 85.95 F
23:47:39.762 -> Average Cal sensor Temp.: 97.81 Ambient Temp.: 85.91 F
23:47:41.327 -> Average Cal sensor Temp.: 97.83 Ambient Temp.: 85.95 F
23:47:42.888 -> Average Cal sensor Temp.: 97.69 Ambient Temp.: 85.91 F
23:47:44.446 -> Average Cal sensor Temp.: 97.94 Ambient Temp.: 85.91 F
23:47:45.975 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 85.91 F
23:47:47.537 -> Average Cal sensor Temp.: 98.13 Ambient Temp.: 85.95 F
23:47:49.099 -> Average Cal sensor Temp.: 98.12 Ambient Temp.: 85.84 F
23:47:50.666 -> Average Cal sensor Temp.: 98.42 Ambient Temp.: 85.91 F
23:47:52.221 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 85.95 F
23:47:53.782 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 85.91 F
23:47:55.335 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 85.91 F
23:47:56.892 -> Average Cal sensor Temp.: 97.70 Ambient Temp.: 85.87 F
23:47:58.453 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 85.91 F
23:47:59.978 -> Average Cal sensor Temp.: 97.64 Ambient Temp.: 85.91 F
23:48:01.543 -> Average Cal sensor Temp.: 97.40 Ambient Temp.: 85.80 F
23:48:03.103 -> Average Cal sensor Temp.: 97.48 Ambient Temp.: 85.77 F
23:48:04.663 -> Average Cal sensor Temp.: 97.94 Ambient Temp.: 85.80 F
23:48:06.229 -> Average Cal sensor Temp.: 97.66 Ambient Temp.: 85.77 F
23:48:07.789 -> Average Cal sensor Temp.: 97.43 Ambient Temp.: 85.69 F
23:48:09.318 -> Average Cal sensor Temp.: 97.75 Ambient Temp.: 85.69 F
23:48:10.876 -> Average Cal sensor Temp.: 97.75 Ambient Temp.: 85.69 F
23:48:12.434 -> Average Cal sensor Temp.: 97.46 Ambient Temp.: 85.59 F
23:48:13.992 -> Average Cal sensor Temp.: 97.57 Ambient Temp.: 85.55 F
23:48:15.549 -> Average Cal sensor Temp.: 97.58 Ambient Temp.: 85.51 F
23:48:17.106 -> Average Cal sensor Temp.: 97.94 Ambient Temp.: 85.51 F
23:48:18.671 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 85.51 F
23:48:20.231 -> Average Cal sensor Temp.: 98.16 Ambient Temp.: 85.48 F
23:48:21.791 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 85.55 F
23:48:23.316 -> Average Cal sensor Temp.: 98.03 Ambient Temp.: 85.48 F
23:48:24.874 -> Average Cal sensor Temp.: 98.28 Ambient Temp.: 85.48 F
23:48:26.435 -> Average Cal sensor Temp.: 98.22 Ambient Temp.: 85.48 F
23:48:28.020 -> Average Cal sensor Temp.: 97.70 Ambient Temp.: 85.48 F
23:48:29.545 -> Average Cal sensor Temp.: 98.05 Ambient Temp.: 85.44 F
23:48:31.102 -> Average Cal sensor Temp.: 97.85 Ambient Temp.: 85.48 F
23:48:32.687 -> Average Cal sensor Temp.: 97.98 Ambient Temp.: 85.44 F
23:48:34.215 -> Average Cal sensor Temp.: 98.02 Ambient Temp.: 85.51 F
23:48:35.773 -> Average Cal sensor Temp.: 97.69 Ambient Temp.: 85.44 F
23:48:37.332 -> Average Cal sensor Temp.: 97.86 Ambient Temp.: 85.44 F
23:48:38.889 -> Average Cal sensor Temp.: 98.04 Ambient Temp.: 85.41 F
23:48:40.441 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 85.48 F
23:48:41.999 -> Average Cal sensor Temp.: 97.87 Ambient Temp.: 85.41 F
23:48:43.585 -> Average Cal sensor Temp.: 97.91 Ambient Temp.: 85.41 F
23:48:45.141 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 85.41 F
23:48:46.665 -> Average Cal sensor Temp.: 97.98 Ambient Temp.: 85.37 F
23:48:48.223 -> Average Cal sensor Temp.: 97.64 Ambient Temp.: 85.41 F
23:48:49.782 -> Average Cal sensor Temp.: 97.44 Ambient Temp.: 85.37 F
23:48:51.346 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 85.37 F
23:48:52.903 -> Average Cal sensor Temp.: 97.63 Ambient Temp.: 85.33 F
23:48:54.459 -> Average Cal sensor Temp.: 97.62 Ambient Temp.: 85.30 F
23:48:56.018 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 85.30 F
23:48:57.575 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 85.26 F
23:48:59.134 -> Average Cal sensor Temp.: 97.75 Ambient Temp.: 85.15 F
23:49:00.692 -> Average Cal sensor Temp.: 97.75 Ambient Temp.: 85.26 F
23:49:02.220 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 85.19 F
23:49:03.775 -> Average Cal sensor Temp.: 97.69 Ambient Temp.: 85.26 F
23:49:05.335 -> Average Cal sensor Temp.: 97.84 Ambient Temp.: 85.19 F
23:49:06.891 -> Average Cal sensor Temp.: 97.74 Ambient Temp.: 85.19 F
23:49:08.453 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 85.19 F
23:49:10.014 -> Average Cal sensor Temp.: 97.96 Ambient Temp.: 85.19 F
23:49:11.578 -> Average Cal sensor Temp.: 97.85 Ambient Temp.: 85.15 F
23:49:13.138 -> Average Cal sensor Temp.: 97.74 Ambient Temp.: 85.19 F
23:49:14.672 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 85.15 F
23:49:16.231 -> Average Cal sensor Temp.: 97.76 Ambient Temp.: 85.26 F
23:49:17.789 -> Average Cal sensor Temp.: 97.86 Ambient Temp.: 85.26 F
23:49:19.355 -> Average Cal sensor Temp.: 97.72 Ambient Temp.: 85.19 F
23:49:20.914 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 85.26 F
23:49:22.441 -> Average Cal sensor Temp.: 97.93 Ambient Temp.: 85.26 F
23:49:24.004 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 85.26 F
23:49:25.557 -> Average Cal sensor Temp.: 97.66 Ambient Temp.: 85.19 F
23:49:27.112 -> Average Cal sensor Temp.: 97.54 Ambient Temp.: 85.33 F
23:49:28.675 -> Average Cal sensor Temp.: 97.84 Ambient Temp.: 85.33 F
23:49:30.243 -> Average Cal sensor Temp.: 97.52 Ambient Temp.: 85.37 F
23:49:31.804 -> Average Cal sensor Temp.: 97.85 Ambient Temp.: 85.37 F
23:49:33.362 -> Average Cal sensor Temp.: 97.99 Ambient Temp.: 85.41 F
23:49:34.889 -> Average Cal sensor Temp.: 97.97 Ambient Temp.: 85.44 F
23:49:36.454 -> Average Cal sensor Temp.: 98.05 Ambient Temp.: 85.41 F
23:49:38.012 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 85.48 F
23:49:39.573 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 85.44 F
23:49:41.133 -> Average Cal sensor Temp.: 97.75 Ambient Temp.: 85.48 F
23:49:42.693 -> Average Cal sensor Temp.: 98.17 Ambient Temp.: 85.48 F
23:49:44.251 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 85.48 F
23:49:45.808 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 85.48 F
23:49:47.360 -> Average Cal sensor Temp.: 97.57 Ambient Temp.: 85.51 F
23:49:48.889 -> Average Cal sensor Temp.: 97.91 Ambient Temp.: 85.48 F
23:49:50.447 -> Average Cal sensor Temp.: 97.55 Ambient Temp.: 85.51 F
23:49:51.998 -> Average Cal sensor Temp.: 97.66 Ambient Temp.: 85.51 F
23:49:53.554 -> Average Cal sensor Temp.: 97.89 Ambient Temp.: 85.55 F
23:49:55.112 -> Average Cal sensor Temp.: 97.96 Ambient Temp.: 85.55 F
23:49:56.696 -> Average Cal sensor Temp.: 97.93 Ambient Temp.: 85.55 F
23:49:58.249 -> Average Cal sensor Temp.: 97.91 Ambient Temp.: 85.59 F
23:49:59.802 -> Average Cal sensor Temp.: 97.98 Ambient Temp.: 85.59 F
23:50:01.356 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 85.51 F
23:50:02.906 -> Average Cal sensor Temp.: 98.02 Ambient Temp.: 85.66 F
23:50:04.466 -> Average Cal sensor Temp.: 97.81 Ambient Temp.: 85.59 F
23:50:06.026 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 85.69 F
23:50:07.579 -> Average Cal sensor Temp.: 97.77 Ambient Temp.: 85.66 F
23:50:09.137 -> Average Cal sensor Temp.: 97.76 Ambient Temp.: 85.69 F
23:50:10.691 -> Average Cal sensor Temp.: 97.95 Ambient Temp.: 85.73 F
23:50:12.249 -> Average Cal sensor Temp.: 97.82 Ambient Temp.: 85.59 F
23:50:13.795 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 85.77 F
23:50:15.350 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 85.69 F
23:50:16.902 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 85.73 F
23:50:18.460 -> Average Cal sensor Temp.: 97.82 Ambient Temp.: 85.66 F
23:50:20.011 -> Average Cal sensor Temp.: 97.87 Ambient Temp.: 85.77 F
23:50:21.571 -> Average Cal sensor Temp.: 98.18 Ambient Temp.: 85.73 F
23:50:23.131 -> Average Cal sensor Temp.: 98.05 Ambient Temp.: 85.73 F
23:50:24.690 -> Average Cal sensor Temp.: 98.00 Ambient Temp.: 85.77 F
23:50:26.255 -> Average Cal sensor Temp.: 98.09 Ambient Temp.: 85.80 F
23:50:27.815 -> Average Cal sensor Temp.: 98.14 Ambient Temp.: 85.84 F
23:50:29.369 -> Average Cal sensor Temp.: 98.08 Ambient Temp.: 85.80 F
23:50:30.928 -> Average Cal sensor Temp.: 98.32 Ambient Temp.: 85.87 F
23:50:32.486 -> Average Cal sensor Temp.: 98.05 Ambient Temp.: 85.87 F
23:50:34.017 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 85.87 F
23:50:35.575 -> Average Cal sensor Temp.: 98.25 Ambient Temp.: 85.87 F
23:50:37.136 -> Average Cal sensor Temp.: 98.17 Ambient Temp.: 85.84 F
23:50:38.693 -> Average Cal sensor Temp.: 97.98 Ambient Temp.: 85.87 F
23:50:40.248 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 85.91 F
23:50:41.803 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 85.87 F
23:50:43.359 -> Average Cal sensor Temp.: 98.00 Ambient Temp.: 85.84 F
23:50:44.927 -> Average Cal sensor Temp.: 97.86 Ambient Temp.: 85.91 F
23:50:46.494 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 85.84 F
23:50:48.049 -> Average Cal sensor Temp.: 97.93 Ambient Temp.: 85.95 F
23:50:49.608 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 85.91 F
23:50:51.168 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 85.98 F
23:50:52.723 -> Average Cal sensor Temp.: 97.94 Ambient Temp.: 85.95 F
23:50:54.283 -> Average Cal sensor Temp.: 98.05 Ambient Temp.: 85.98 F
23:50:55.837 -> Average Cal sensor Temp.: 97.95 Ambient Temp.: 85.98 F
23:50:57.392 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 85.98 F
23:50:58.945 -> Average Cal sensor Temp.: 98.04 Ambient Temp.: 86.05 F
23:51:00.502 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 86.05 F
23:51:02.060 -> Average Cal sensor Temp.: 97.85 Ambient Temp.: 86.09 F
23:51:03.618 -> Average Cal sensor Temp.: 97.97 Ambient Temp.: 86.09 F
23:51:05.176 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 86.05 F
23:51:06.733 -> Average Cal sensor Temp.: 97.97 Ambient Temp.: 86.13 F
23:51:08.296 -> Average Cal sensor Temp.: 97.83 Ambient Temp.: 86.13 F
23:51:09.823 -> Average Cal sensor Temp.: 98.26 Ambient Temp.: 86.13 F
23:51:11.380 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 86.20 F
23:51:12.935 -> Average Cal sensor Temp.: 98.09 Ambient Temp.: 86.20 F
23:51:14.496 -> Average Cal sensor Temp.: 98.30 Ambient Temp.: 86.13 F
23:51:16.059 -> Average Cal sensor Temp.: 98.30 Ambient Temp.: 86.20 F
23:51:17.622 -> Average Cal sensor Temp.: 98.12 Ambient Temp.: 86.16 F
23:51:19.181 -> Average Cal sensor Temp.: 97.93 Ambient Temp.: 86.23 F
23:51:20.740 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 86.23 F
23:51:22.298 -> Average Cal sensor Temp.: 98.05 Ambient Temp.: 86.27 F
23:51:23.857 -> Average Cal sensor Temp.: 97.95 Ambient Temp.: 86.27 F
23:51:25.417 -> Average Cal sensor Temp.: 98.05 Ambient Temp.: 86.27 F
23:51:26.945 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 86.27 F
23:51:28.535 -> Average Cal sensor Temp.: 97.93 Ambient Temp.: 86.31 F
23:51:30.090 -> Average Cal sensor Temp.: 98.08 Ambient Temp.: 86.38 F
23:51:31.649 -> Average Cal sensor Temp.: 97.87 Ambient Temp.: 86.34 F
23:51:33.199 -> Average Cal sensor Temp.: 98.04 Ambient Temp.: 86.34 F
23:51:34.760 -> Average Cal sensor Temp.: 98.13 Ambient Temp.: 86.38 F
23:51:36.317 -> Average Cal sensor Temp.: 98.25 Ambient Temp.: 86.45 F
23:51:37.843 -> Average Cal sensor Temp.: 98.23 Ambient Temp.: 86.45 F
23:51:39.404 -> Average Cal sensor Temp.: 97.93 Ambient Temp.: 86.45 F
23:51:40.963 -> Average Cal sensor Temp.: 98.10 Ambient Temp.: 86.45 F
23:51:42.520 -> Average Cal sensor Temp.: 97.83 Ambient Temp.: 86.52 F
23:51:44.080 -> Average Cal sensor Temp.: 97.87 Ambient Temp.: 86.52 F
23:51:45.638 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 86.52 F
23:51:47.199 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 86.49 F
23:51:48.757 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 86.56 F
23:51:50.315 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 86.56 F
23:51:51.879 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 86.59 F
23:51:53.436 -> Average Cal sensor Temp.: 98.08 Ambient Temp.: 86.52 F
23:51:54.990 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 86.59 F
23:51:56.543 -> Average Cal sensor Temp.: 98.18 Ambient Temp.: 86.56 F
23:51:58.089 -> Average Cal sensor Temp.: 98.23 Ambient Temp.: 86.63 F
23:51:59.652 -> Average Cal sensor Temp.: 98.18 Ambient Temp.: 86.63 F
23:52:01.205 -> Average Cal sensor Temp.: 98.21 Ambient Temp.: 86.67 F
23:52:02.753 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 86.67 F
23:52:04.340 -> Average Cal sensor Temp.: 98.23 Ambient Temp.: 86.70 F
23:52:05.898 -> Average Cal sensor Temp.: 98.23 Ambient Temp.: 86.70 F
23:52:07.424 -> Average Cal sensor Temp.: 98.27 Ambient Temp.: 86.74 F
23:52:08.985 -> Average Cal sensor Temp.: 98.34 Ambient Temp.: 86.74 F
23:52:10.538 -> Average Cal sensor Temp.: 98.22 Ambient Temp.: 86.74 F
23:52:12.099 -> Average Cal sensor Temp.: 98.29 Ambient Temp.: 86.74 F
23:52:13.661 -> Average Cal sensor Temp.: 98.27 Ambient Temp.: 86.74 F
23:52:15.221 -> Average Cal sensor Temp.: 98.26 Ambient Temp.: 86.70 F
23:52:16.780 -> Average Cal sensor Temp.: 98.24 Ambient Temp.: 86.74 F
23:52:18.338 -> Average Cal sensor Temp.: 98.33 Ambient Temp.: 86.74 F
23:52:19.901 -> Average Cal sensor Temp.: 98.24 Ambient Temp.: 86.81 F
23:52:21.462 -> Average Cal sensor Temp.: 98.19 Ambient Temp.: 86.77 F
23:52:23.023 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 86.88 F
23:52:24.581 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 86.77 F
23:52:26.138 -> Average Cal sensor Temp.: 97.96 Ambient Temp.: 86.81 F
23:52:27.694 -> Average Cal sensor Temp.: 97.95 Ambient Temp.: 86.74 F
23:52:29.220 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 86.88 F
23:52:30.785 -> Average Cal sensor Temp.: 98.18 Ambient Temp.: 86.77 F
23:52:32.343 -> Average Cal sensor Temp.: 98.22 Ambient Temp.: 86.77 F
23:52:33.904 -> Average Cal sensor Temp.: 98.31 Ambient Temp.: 86.88 F
23:52:35.460 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 86.81 F
23:52:37.016 -> Average Cal sensor Temp.: 98.16 Ambient Temp.: 86.77 F
23:52:38.574 -> Average Cal sensor Temp.: 98.17 Ambient Temp.: 86.95 F
23:52:40.133 -> Average Cal sensor Temp.: 97.87 Ambient Temp.: 86.92 F
23:52:41.692 -> Average Cal sensor Temp.: 97.82 Ambient Temp.: 86.92 F
23:52:43.250 -> Average Cal sensor Temp.: 97.91 Ambient Temp.: 86.88 F
23:52:44.808 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 86.88 F
23:52:46.361 -> Average Cal sensor Temp.: 98.07 Ambient Temp.: 86.81 F
23:52:47.920 -> Average Cal sensor Temp.: 98.31 Ambient Temp.: 86.92 F
23:52:49.478 -> Average Cal sensor Temp.: 98.19 Ambient Temp.: 86.88 F
23:52:51.034 -> Average Cal sensor Temp.: 98.08 Ambient Temp.: 86.88 F
23:52:52.597 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 86.88 F
23:52:54.161 -> Average Cal sensor Temp.: 97.76 Ambient Temp.: 86.81 F
23:52:55.722 -> Average Cal sensor Temp.: 97.84 Ambient Temp.: 86.81 F
23:52:57.281 -> Average Cal sensor Temp.: 97.96 Ambient Temp.: 86.81 F
23:52:58.805 -> Average Cal sensor Temp.: 97.89 Ambient Temp.: 86.81 F
23:53:00.363 -> Average Cal sensor Temp.: 97.91 Ambient Temp.: 86.81 F
23:53:01.917 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 86.77 F
23:53:03.475 -> Average Cal sensor Temp.: 97.74 Ambient Temp.: 86.70 F
23:53:05.030 -> Average Cal sensor Temp.: 97.66 Ambient Temp.: 86.70 F
23:53:06.586 -> Average Cal sensor Temp.: 97.54 Ambient Temp.: 86.67 F
23:53:08.148 -> Average Cal sensor Temp.: 97.27 Ambient Temp.: 86.63 F
23:53:09.708 -> Average Cal sensor Temp.: 97.96 Ambient Temp.: 86.63 F
23:53:11.264 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 86.67 F
23:53:12.819 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 86.67 F
23:53:14.372 -> Average Cal sensor Temp.: 98.33 Ambient Temp.: 86.63 F
23:53:15.934 -> Average Cal sensor Temp.: 97.95 Ambient Temp.: 86.63 F
23:53:17.488 -> Average Cal sensor Temp.: 97.91 Ambient Temp.: 86.59 F
23:53:19.044 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 86.63 F
23:53:20.595 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 86.63 F
23:53:22.150 -> Average Cal sensor Temp.: 97.95 Ambient Temp.: 86.63 F
23:53:23.710 -> Average Cal sensor Temp.: 97.99 Ambient Temp.: 86.59 F
23:53:25.267 -> Average Cal sensor Temp.: 98.07 Ambient Temp.: 86.59 F
23:53:26.822 -> Average Cal sensor Temp.: 97.74 Ambient Temp.: 86.56 F
23:53:28.410 -> Average Cal sensor Temp.: 97.91 Ambient Temp.: 86.56 F
23:53:29.933 -> Average Cal sensor Temp.: 98.14 Ambient Temp.: 86.56 F
23:53:31.491 -> Average Cal sensor Temp.: 98.04 Ambient Temp.: 86.49 F
23:53:33.080 -> Average Cal sensor Temp.: 98.10 Ambient Temp.: 86.52 F
23:53:34.630 -> Average Cal sensor Temp.: 98.03 Ambient Temp.: 86.49 F
23:53:36.190 -> Average Cal sensor Temp.: 97.99 Ambient Temp.: 86.38 F
23:53:37.746 -> Average Cal sensor Temp.: 98.02 Ambient Temp.: 86.45 F
23:53:39.310 -> Average Cal sensor Temp.: 98.27 Ambient Temp.: 86.45 F
23:53:40.834 -> Average Cal sensor Temp.: 98.24 Ambient Temp.: 86.45 F
23:53:42.422 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 86.45 F
23:53:43.981 -> Average Cal sensor Temp.: 98.16 Ambient Temp.: 86.49 F
23:53:45.541 -> Average Cal sensor Temp.: 97.96 Ambient Temp.: 86.45 F
23:53:47.064 -> Average Cal sensor Temp.: 97.75 Ambient Temp.: 86.34 F
23:53:48.620 -> Average Cal sensor Temp.: 97.82 Ambient Temp.: 86.34 F
23:53:50.181 -> Average Cal sensor Temp.: 97.91 Ambient Temp.: 86.31 F
23:53:51.734 -> Average Cal sensor Temp.: 97.99 Ambient Temp.: 86.31 F
23:53:53.289 -> Average Cal sensor Temp.: 98.40 Ambient Temp.: 86.31 F
23:53:54.845 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 86.31 F
23:53:56.405 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 86.23 F
23:53:57.964 -> Average Cal sensor Temp.: 97.85 Ambient Temp.: 86.27 F
23:53:59.528 -> Average Cal sensor Temp.: 98.19 Ambient Temp.: 86.23 F
23:54:01.086 -> Average Cal sensor Temp.: 98.14 Ambient Temp.: 86.16 F
23:54:02.646 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 86.23 F
23:54:04.203 -> Average Cal sensor Temp.: 97.72 Ambient Temp.: 86.16 F
23:54:05.762 -> Average Cal sensor Temp.: 97.63 Ambient Temp.: 86.09 F
23:54:07.319 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 86.13 F
23:54:08.877 -> Average Cal sensor Temp.: 98.00 Ambient Temp.: 86.09 F
23:54:10.434 -> Average Cal sensor Temp.: 98.33 Ambient Temp.: 86.13 F
23:54:11.984 -> Average Cal sensor Temp.: 98.00 Ambient Temp.: 86.09 F
23:54:13.534 -> Average Cal sensor Temp.: 98.21 Ambient Temp.: 85.98 F
23:54:15.096 -> Average Cal sensor Temp.: 98.27 Ambient Temp.: 86.05 F
23:54:16.651 -> Average Cal sensor Temp.: 97.48 Ambient Temp.: 85.98 F
23:54:18.208 -> Average Cal sensor Temp.: 97.64 Ambient Temp.: 85.98 F
23:54:19.767 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 86.05 F
23:54:21.321 -> Average Cal sensor Temp.: 97.71 Ambient Temp.: 85.87 F
23:54:22.877 -> Average Cal sensor Temp.: 97.99 Ambient Temp.: 85.95 F
23:54:24.444 -> Average Cal sensor Temp.: 98.07 Ambient Temp.: 85.95 F
23:54:26.003 -> Average Cal sensor Temp.: 98.27 Ambient Temp.: 85.87 F
23:54:27.561 -> Average Cal sensor Temp.: 98.26 Ambient Temp.: 85.87 F
23:54:29.119 -> Average Cal sensor Temp.: 97.95 Ambient Temp.: 85.87 F
23:54:30.678 -> Average Cal sensor Temp.: 97.98 Ambient Temp.: 85.91 F
23:54:32.206 -> Average Cal sensor Temp.: 98.12 Ambient Temp.: 85.84 F
23:54:33.764 -> Average Cal sensor Temp.: 98.29 Ambient Temp.: 85.84 F
23:54:35.327 -> Average Cal sensor Temp.: 98.60 Ambient Temp.: 85.84 F
23:54:36.885 -> Average Cal sensor Temp.: 97.96 Ambient Temp.: 85.77 F
23:54:38.446 -> Average Cal sensor Temp.: 98.03 Ambient Temp.: 85.80 F
23:54:40.008 -> Average Cal sensor Temp.: 98.18 Ambient Temp.: 85.77 F
23:54:41.568 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 85.77 F
23:54:43.127 -> Average Cal sensor Temp.: 97.99 Ambient Temp.: 85.73 F
23:54:44.686 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 85.69 F
23:54:46.240 -> Average Cal sensor Temp.: 97.93 Ambient Temp.: 85.69 F
23:54:47.801 -> Average Cal sensor Temp.: 98.04 Ambient Temp.: 85.69 F
23:54:49.359 -> Average Cal sensor Temp.: 98.17 Ambient Temp.: 85.59 F
23:54:50.911 -> Average Cal sensor Temp.: 98.24 Ambient Temp.: 85.69 F
23:54:52.465 -> Average Cal sensor Temp.: 97.98 Ambient Temp.: 85.59 F
23:54:54.020 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 85.55 F
23:54:55.567 -> Average Cal sensor Temp.: 98.24 Ambient Temp.: 85.59 F
23:54:57.122 -> Average Cal sensor Temp.: 98.41 Ambient Temp.: 85.55 F
23:54:58.679 -> Average Cal sensor Temp.: 98.22 Ambient Temp.: 85.55 F
23:55:00.240 -> Average Cal sensor Temp.: 98.33 Ambient Temp.: 85.55 F
23:55:01.794 -> Average Cal sensor Temp.: 98.51 Ambient Temp.: 85.55 F
23:55:03.351 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 85.55 F
23:55:04.911 -> Average Cal sensor Temp.: 98.10 Ambient Temp.: 85.55 F
23:55:06.467 -> Average Cal sensor Temp.: 98.45 Ambient Temp.: 85.51 F
23:55:08.021 -> Average Cal sensor Temp.: 98.03 Ambient Temp.: 85.51 F
23:55:09.579 -> Average Cal sensor Temp.: 98.19 Ambient Temp.: 85.44 F
23:55:11.130 -> Average Cal sensor Temp.: 97.93 Ambient Temp.: 85.44 F
23:55:12.685 -> Average Cal sensor Temp.: 98.04 Ambient Temp.: 85.41 F
23:55:14.243 -> Average Cal sensor Temp.: 97.97 Ambient Temp.: 85.44 F
23:55:15.795 -> Average Cal sensor Temp.: 97.83 Ambient Temp.: 85.33 F
23:55:17.349 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 85.37 F
23:55:18.910 -> Average Cal sensor Temp.: 98.33 Ambient Temp.: 85.30 F
23:55:20.468 -> Average Cal sensor Temp.: 98.19 Ambient Temp.: 85.37 F
23:55:22.030 -> Average Cal sensor Temp.: 97.81 Ambient Temp.: 85.26 F
23:55:23.586 -> Average Cal sensor Temp.: 97.91 Ambient Temp.: 85.26 F
23:55:25.143 -> Average Cal sensor Temp.: 97.72 Ambient Temp.: 85.19 F
23:55:26.664 -> Average Cal sensor Temp.: 97.95 Ambient Temp.: 85.26 F
23:55:28.253 -> Average Cal sensor Temp.: 97.80 Ambient Temp.: 85.15 F
23:55:29.781 -> Average Cal sensor Temp.: 98.22 Ambient Temp.: 85.15 F
23:55:31.344 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 85.15 F
23:55:32.902 -> Average Cal sensor Temp.: 98.36 Ambient Temp.: 85.12 F
23:55:34.466 -> Average Cal sensor Temp.: 98.46 Ambient Temp.: 85.19 F
23:55:36.026 -> Average Cal sensor Temp.: 98.29 Ambient Temp.: 85.12 F
23:55:37.589 -> Average Cal sensor Temp.: 98.18 Ambient Temp.: 85.12 F
23:55:39.116 -> Average Cal sensor Temp.: 98.10 Ambient Temp.: 85.08 F
23:55:40.680 -> Average Cal sensor Temp.: 98.25 Ambient Temp.: 85.15 F
23:55:42.244 -> Average Cal sensor Temp.: 97.76 Ambient Temp.: 85.08 F
23:55:43.800 -> Average Cal sensor Temp.: 97.82 Ambient Temp.: 85.05 F
23:55:45.361 -> Average Cal sensor Temp.: 98.21 Ambient Temp.: 85.08 F
23:55:46.916 -> Average Cal sensor Temp.: 98.23 Ambient Temp.: 85.08 F
23:55:48.476 -> Average Cal sensor Temp.: 98.36 Ambient Temp.: 85.05 F
23:55:50.004 -> Average Cal sensor Temp.: 98.37 Ambient Temp.: 85.08 F
23:55:51.561 -> Average Cal sensor Temp.: 98.24 Ambient Temp.: 85.05 F
23:55:53.123 -> Average Cal sensor Temp.: 98.12 Ambient Temp.: 85.01 F
23:55:54.688 -> Average Cal sensor Temp.: 98.33 Ambient Temp.: 85.08 F
23:55:56.247 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 85.01 F
23:55:57.804 -> Average Cal sensor Temp.: 98.09 Ambient Temp.: 85.01 F
23:55:59.360 -> Average Cal sensor Temp.: 98.27 Ambient Temp.: 84.94 F
23:56:00.919 -> Average Cal sensor Temp.: 98.39 Ambient Temp.: 85.01 F
23:56:02.475 -> Average Cal sensor Temp.: 98.28 Ambient Temp.: 84.94 F
23:56:04.038 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 85.01 F
23:56:05.594 -> Average Cal sensor Temp.: 98.24 Ambient Temp.: 85.01 F
23:56:07.159 -> Average Cal sensor Temp.: 98.24 Ambient Temp.: 85.05 F
23:56:08.685 -> Average Cal sensor Temp.: 98.02 Ambient Temp.: 85.01 F
23:56:10.248 -> Average Cal sensor Temp.: 98.00 Ambient Temp.: 84.97 F
23:56:11.808 -> Average Cal sensor Temp.: 98.08 Ambient Temp.: 84.97 F
23:56:13.371 -> Average Cal sensor Temp.: 97.95 Ambient Temp.: 84.97 F
23:56:14.931 -> Average Cal sensor Temp.: 98.07 Ambient Temp.: 84.90 F
23:56:16.496 -> Average Cal sensor Temp.: 98.04 Ambient Temp.: 84.94 F
23:56:18.058 -> Average Cal sensor Temp.: 98.18 Ambient Temp.: 84.94 F
23:56:19.617 -> Average Cal sensor Temp.: 98.27 Ambient Temp.: 84.97 F
23:56:21.149 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 84.94 F
23:56:22.710 -> Average Cal sensor Temp.: 97.81 Ambient Temp.: 84.87 F
23:56:24.277 -> Average Cal sensor Temp.: 97.94 Ambient Temp.: 84.94 F
23:56:25.830 -> Average Cal sensor Temp.: 98.54 Ambient Temp.: 84.94 F
23:56:27.392 -> Average Cal sensor Temp.: 98.51 Ambient Temp.: 84.94 F
23:56:28.953 -> Average Cal sensor Temp.: 98.42 Ambient Temp.: 85.01 F
23:56:30.514 -> Average Cal sensor Temp.: 98.51 Ambient Temp.: 84.97 F
23:56:32.042 -> Average Cal sensor Temp.: 98.12 Ambient Temp.: 84.97 F
23:56:33.604 -> Average Cal sensor Temp.: 97.95 Ambient Temp.: 84.97 F
23:56:35.172 -> Average Cal sensor Temp.: 98.18 Ambient Temp.: 85.01 F
23:56:36.734 -> Average Cal sensor Temp.: 97.98 Ambient Temp.: 84.97 F
23:56:38.298 -> Average Cal sensor Temp.: 98.12 Ambient Temp.: 85.01 F
23:56:39.825 -> Average Cal sensor Temp.: 98.24 Ambient Temp.: 85.08 F
23:56:41.390 -> Average Cal sensor Temp.: 98.10 Ambient Temp.: 85.01 F
23:56:42.954 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 85.05 F
23:56:44.518 -> Average Cal sensor Temp.: 98.32 Ambient Temp.: 85.05 F
23:56:46.079 -> Average Cal sensor Temp.: 98.23 Ambient Temp.: 85.05 F
23:56:47.629 -> Average Cal sensor Temp.: 98.16 Ambient Temp.: 85.08 F
23:56:49.185 -> Average Cal sensor Temp.: 98.27 Ambient Temp.: 85.05 F
23:56:50.747 -> Average Cal sensor Temp.: 98.39 Ambient Temp.: 85.08 F
23:56:52.300 -> Average Cal sensor Temp.: 98.27 Ambient Temp.: 85.12 F
23:56:53.859 -> Average Cal sensor Temp.: 98.18 Ambient Temp.: 85.08 F
23:56:55.418 -> Average Cal sensor Temp.: 98.27 Ambient Temp.: 85.15 F
23:56:56.977 -> Average Cal sensor Temp.: 98.25 Ambient Temp.: 85.08 F
23:56:58.503 -> Average Cal sensor Temp.: 98.58 Ambient Temp.: 85.19 F
23:57:00.064 -> Average Cal sensor Temp.: 98.49 Ambient Temp.: 85.19 F
23:57:01.620 -> Average Cal sensor Temp.: 98.43 Ambient Temp.: 85.15 F
23:57:03.181 -> Average Cal sensor Temp.: 98.61 Ambient Temp.: 85.15 F
23:57:04.732 -> Average Cal sensor Temp.: 98.32 Ambient Temp.: 85.26 F
23:57:06.292 -> Average Cal sensor Temp.: 98.26 Ambient Temp.: 85.19 F
23:57:07.852 -> Average Cal sensor Temp.: 98.34 Ambient Temp.: 85.15 F
23:57:09.411 -> Average Cal sensor Temp.: 98.36 Ambient Temp.: 85.26 F
23:57:10.974 -> Average Cal sensor Temp.: 98.25 Ambient Temp.: 85.19 F
23:57:12.534 -> Average Cal sensor Temp.: 98.48 Ambient Temp.: 85.33 F
23:57:14.093 -> Average Cal sensor Temp.: 98.45 Ambient Temp.: 85.30 F
23:57:15.652 -> Average Cal sensor Temp.: 98.18 Ambient Temp.: 85.26 F
23:57:17.211 -> Average Cal sensor Temp.: 98.49 Ambient Temp.: 85.37 F
23:57:18.766 -> Average Cal sensor Temp.: 98.44 Ambient Temp.: 85.33 F
23:57:20.292 -> Average Cal sensor Temp.: 98.37 Ambient Temp.: 85.37 F
23:57:21.853 -> Average Cal sensor Temp.: 98.17 Ambient Temp.: 85.41 F
23:57:23.409 -> Average Cal sensor Temp.: 98.47 Ambient Temp.: 85.44 F
23:57:24.969 -> Average Cal sensor Temp.: 98.44 Ambient Temp.: 85.41 F
23:57:26.527 -> Average Cal sensor Temp.: 98.07 Ambient Temp.: 85.41 F
23:57:28.090 -> Average Cal sensor Temp.: 98.19 Ambient Temp.: 85.44 F
23:57:29.646 -> Average Cal sensor Temp.: 98.47 Ambient Temp.: 85.48 F
23:57:31.209 -> Average Cal sensor Temp.: 98.36 Ambient Temp.: 85.55 F
23:57:32.773 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 85.55 F
23:57:34.298 -> Average Cal sensor Temp.: 98.35 Ambient Temp.: 85.51 F
23:57:35.853 -> Average Cal sensor Temp.: 98.10 Ambient Temp.: 85.59 F
23:57:37.409 -> Average Cal sensor Temp.: 98.33 Ambient Temp.: 85.51 F
23:57:38.968 -> Average Cal sensor Temp.: 98.26 Ambient Temp.: 85.55 F
23:57:40.527 -> Average Cal sensor Temp.: 98.31 Ambient Temp.: 85.59 F
23:57:42.083 -> Average Cal sensor Temp.: 98.19 Ambient Temp.: 85.59 F
23:57:43.638 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 85.59 F
23:57:45.200 -> Average Cal sensor Temp.: 98.26 Ambient Temp.: 85.51 F
23:57:46.755 -> Average Cal sensor Temp.: 98.24 Ambient Temp.: 85.55 F
23:57:48.316 -> Average Cal sensor Temp.: 98.50 Ambient Temp.: 85.55 F
23:57:49.878 -> Average Cal sensor Temp.: 98.57 Ambient Temp.: 85.69 F
23:57:51.439 -> Average Cal sensor Temp.: 98.55 Ambient Temp.: 85.69 F
23:57:52.995 -> Average Cal sensor Temp.: 98.60 Ambient Temp.: 85.69 F
23:57:54.557 -> Average Cal sensor Temp.: 98.58 Ambient Temp.: 85.69 F
23:57:56.114 -> Average Cal sensor Temp.: 98.41 Ambient Temp.: 85.73 F
23:57:57.642 -> Average Cal sensor Temp.: 98.56 Ambient Temp.: 85.80 F
23:57:59.204 -> Average Cal sensor Temp.: 98.54 Ambient Temp.: 85.77 F
23:58:00.757 -> Average Cal sensor Temp.: 98.39 Ambient Temp.: 85.80 F
23:58:02.313 -> Average Cal sensor Temp.: 98.34 Ambient Temp.: 85.80 F
23:58:03.874 -> Average Cal sensor Temp.: 98.41 Ambient Temp.: 85.77 F
23:58:05.442 -> Average Cal sensor Temp.: 98.51 Ambient Temp.: 85.84 F
23:58:06.997 -> Average Cal sensor Temp.: 98.44 Ambient Temp.: 85.87 F
23:58:08.556 -> Average Cal sensor Temp.: 98.45 Ambient Temp.: 85.87 F
23:58:10.116 -> Average Cal sensor Temp.: 98.44 Ambient Temp.: 85.91 F
23:58:11.676 -> Average Cal sensor Temp.: 98.48 Ambient Temp.: 85.87 F
23:58:13.238 -> Average Cal sensor Temp.: 98.29 Ambient Temp.: 85.91 F
23:58:14.799 -> Average Cal sensor Temp.: 98.41 Ambient Temp.: 85.91 F
23:58:16.322 -> Average Cal sensor Temp.: 98.82 Ambient Temp.: 85.95 F
23:58:17.883 -> Average Cal sensor Temp.: 98.61 Ambient Temp.: 85.95 F
23:58:19.444 -> Average Cal sensor Temp.: 98.47 Ambient Temp.: 85.95 F
23:58:21.007 -> Average Cal sensor Temp.: 98.57 Ambient Temp.: 85.98 F
23:58:22.565 -> Average Cal sensor Temp.: 98.18 Ambient Temp.: 85.98 F
23:58:24.124 -> Average Cal sensor Temp.: 98.19 Ambient Temp.: 86.05 F
23:58:25.686 -> Average Cal sensor Temp.: 98.39 Ambient Temp.: 86.05 F
23:58:27.243 -> Average Cal sensor Temp.: 98.36 Ambient Temp.: 86.09 F
23:58:28.801 -> Average Cal sensor Temp.: 98.39 Ambient Temp.: 86.05 F
23:58:30.363 -> Average Cal sensor Temp.: 98.34 Ambient Temp.: 86.05 F
23:58:31.923 -> Average Cal sensor Temp.: 98.54 Ambient Temp.: 86.13 F
23:58:33.481 -> Average Cal sensor Temp.: 98.60 Ambient Temp.: 86.09 F
23:58:35.037 -> Average Cal sensor Temp.: 98.59 Ambient Temp.: 86.13 F
23:58:36.588 -> Average Cal sensor Temp.: 98.47 Ambient Temp.: 86.13 F
23:58:38.149 -> Average Cal sensor Temp.: 98.47 Ambient Temp.: 86.16 F
23:58:39.709 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 86.16 F
23:58:41.238 -> Average Cal sensor Temp.: 98.10 Ambient Temp.: 86.16 F
23:58:42.794 -> Average Cal sensor Temp.: 98.38 Ambient Temp.: 86.20 F
23:58:44.353 -> Average Cal sensor Temp.: 98.55 Ambient Temp.: 86.20 F
23:58:45.909 -> Average Cal sensor Temp.: 98.69 Ambient Temp.: 86.23 F
23:58:47.470 -> Average Cal sensor Temp.: 98.46 Ambient Temp.: 86.20 F
23:58:49.034 -> Average Cal sensor Temp.: 98.34 Ambient Temp.: 86.27 F
23:58:50.591 -> Average Cal sensor Temp.: 98.53 Ambient Temp.: 86.23 F
23:58:52.149 -> Average Cal sensor Temp.: 98.70 Ambient Temp.: 86.31 F
23:58:53.716 -> Average Cal sensor Temp.: 98.77 Ambient Temp.: 86.27 F
23:58:55.277 -> Average Cal sensor Temp.: 98.85 Ambient Temp.: 86.31 F
23:58:56.805 -> Average Cal sensor Temp.: 98.31 Ambient Temp.: 86.38 F
23:58:58.360 -> Average Cal sensor Temp.: 98.57 Ambient Temp.: 86.34 F
23:58:59.917 -> Average Cal sensor Temp.: 98.70 Ambient Temp.: 86.34 F
23:59:01.477 -> Average Cal sensor Temp.: 98.88 Ambient Temp.: 86.38 F
23:59:03.037 -> Average Cal sensor Temp.: 98.81 Ambient Temp.: 86.45 F
23:59:04.595 -> Average Cal sensor Temp.: 98.62 Ambient Temp.: 86.38 F
23:59:06.159 -> Average Cal sensor Temp.: 98.85 Ambient Temp.: 86.52 F
23:59:07.707 -> Average Cal sensor Temp.: 98.90 Ambient Temp.: 86.49 F
23:59:09.264 -> Average Cal sensor Temp.: 98.79 Ambient Temp.: 86.49 F
23:59:10.826 -> Average Cal sensor Temp.: 98.40 Ambient Temp.: 86.49 F
23:59:12.383 -> Average Cal sensor Temp.: 98.50 Ambient Temp.: 86.52 F
23:59:13.944 -> Average Cal sensor Temp.: 98.58 Ambient Temp.: 86.52 F
23:59:15.502 -> Average Cal sensor Temp.: 98.69 Ambient Temp.: 86.52 F
23:59:17.059 -> Average Cal sensor Temp.: 98.88 Ambient Temp.: 86.59 F
23:59:18.615 -> Average Cal sensor Temp.: 98.98 Ambient Temp.: 86.52 F
23:59:20.174 -> Average Cal sensor Temp.: 98.68 Ambient Temp.: 86.63 F
23:59:21.729 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 86.59 F
23:59:23.284 -> Average Cal sensor Temp.: 98.23 Ambient Temp.: 86.59 F
23:59:24.841 -> Average Cal sensor Temp.: 98.47 Ambient Temp.: 86.59 F
23:59:26.397 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 86.56 F
23:59:27.959 -> Average Cal sensor Temp.: 98.14 Ambient Temp.: 86.59 F
23:59:29.520 -> Average Cal sensor Temp.: 98.08 Ambient Temp.: 86.56 F
23:59:31.079 -> Average Cal sensor Temp.: 98.45 Ambient Temp.: 86.59 F
23:59:32.635 -> Average Cal sensor Temp.: 97.84 Ambient Temp.: 86.56 F
23:59:34.193 -> Average Cal sensor Temp.: 98.35 Ambient Temp.: 86.56 F
23:59:35.748 -> Average Cal sensor Temp.: 98.53 Ambient Temp.: 86.56 F
23:59:37.306 -> Average Cal sensor Temp.: 98.41 Ambient Temp.: 86.56 F
23:59:38.870 -> Average Cal sensor Temp.: 98.12 Ambient Temp.: 86.52 F
23:59:40.428 -> Average Cal sensor Temp.: 98.21 Ambient Temp.: 86.56 F
23:59:41.985 -> Average Cal sensor Temp.: 98.46 Ambient Temp.: 86.56 F
23:59:43.549 -> Average Cal sensor Temp.: 98.45 Ambient Temp.: 86.52 F
23:59:45.110 -> Average Cal sensor Temp.: 98.13 Ambient Temp.: 86.52 F
23:59:46.669 -> Average Cal sensor Temp.: 98.37 Ambient Temp.: 86.56 F
23:59:48.228 -> Average Cal sensor Temp.: 98.56 Ambient Temp.: 86.52 F
23:59:49.751 -> Average Cal sensor Temp.: 98.92 Ambient Temp.: 86.52 F
23:59:51.312 -> Average Cal sensor Temp.: 98.39 Ambient Temp.: 86.52 F
23:59:52.869 -> Average Cal sensor Temp.: 98.71 Ambient Temp.: 86.56 F
23:59:54.430 -> Average Cal sensor Temp.: 98.83 Ambient Temp.: 86.56 F
23:59:55.990 -> Average Cal sensor Temp.: 98.39 Ambient Temp.: 86.52 F
23:59:57.551 -> Average Cal sensor Temp.: 98.65 Ambient Temp.: 86.52 F
23:59:59.109 -> Average Cal sensor Temp.: 98.74 Ambient Temp.: 86.56 F
00:00:00.672 -> Average Cal sensor Temp.: 98.71 Ambient Temp.: 86.59 F
00:00:02.231 -> Average Cal sensor Temp.: 98.46 Ambient Temp.: 86.59 F
00:00:03.791 -> Average Cal sensor Temp.: 98.46 Ambient Temp.: 86.59 F
00:00:05.351 -> Average Cal sensor Temp.: 98.53 Ambient Temp.: 86.59 F
00:00:06.906 -> Average Cal sensor Temp.: 98.16 Ambient Temp.: 86.56 F
00:00:08.461 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 86.49 F
00:00:10.023 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 86.45 F
00:00:11.577 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 86.49 F
00:00:13.129 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 86.49 F
00:00:14.690 -> Average Cal sensor Temp.: 98.33 Ambient Temp.: 86.45 F
00:00:16.248 -> Average Cal sensor Temp.: 98.23 Ambient Temp.: 86.38 F
00:00:17.806 -> Average Cal sensor Temp.: 98.05 Ambient Temp.: 86.38 F
00:00:19.365 -> Average Cal sensor Temp.: 98.12 Ambient Temp.: 86.38 F
00:00:20.924 -> Average Cal sensor Temp.: 98.44 Ambient Temp.: 86.45 F
00:00:22.484 -> Average Cal sensor Temp.: 98.36 Ambient Temp.: 86.34 F
00:00:24.042 -> Average Cal sensor Temp.: 98.51 Ambient Temp.: 86.38 F
00:00:25.598 -> Average Cal sensor Temp.: 98.54 Ambient Temp.: 86.34 F
00:00:27.156 -> Average Cal sensor Temp.: 98.61 Ambient Temp.: 86.34 F
00:00:28.684 -> Average Cal sensor Temp.: 98.66 Ambient Temp.: 86.38 F
00:00:30.242 -> Average Cal sensor Temp.: 98.44 Ambient Temp.: 86.38 F
00:00:31.803 -> Average Cal sensor Temp.: 98.55 Ambient Temp.: 86.34 F
00:00:33.356 -> Average Cal sensor Temp.: 98.78 Ambient Temp.: 86.38 F
00:00:34.942 -> Average Cal sensor Temp.: 98.89 Ambient Temp.: 86.34 F
00:00:36.501 -> Average Cal sensor Temp.: 98.77 Ambient Temp.: 86.31 F
00:00:38.028 -> Average Cal sensor Temp.: 98.86 Ambient Temp.: 86.45 F
00:00:39.588 -> Average Cal sensor Temp.: 98.74 Ambient Temp.: 86.38 F
00:00:41.149 -> Average Cal sensor Temp.: 98.72 Ambient Temp.: 86.38 F
00:00:42.708 -> Average Cal sensor Temp.: 98.72 Ambient Temp.: 86.45 F
00:00:44.262 -> Average Cal sensor Temp.: 98.65 Ambient Temp.: 86.45 F
00:00:45.823 -> Average Cal sensor Temp.: 98.69 Ambient Temp.: 86.45 F
00:00:47.388 -> Average Cal sensor Temp.: 98.41 Ambient Temp.: 86.34 F
00:00:48.951 -> Average Cal sensor Temp.: 98.47 Ambient Temp.: 86.31 F
00:00:50.482 -> Average Cal sensor Temp.: 98.82 Ambient Temp.: 86.34 F
00:00:52.044 -> Average Cal sensor Temp.: 98.51 Ambient Temp.: 86.34 F
00:00:53.603 -> Average Cal sensor Temp.: 98.46 Ambient Temp.: 86.34 F
00:00:55.159 -> Average Cal sensor Temp.: 98.58 Ambient Temp.: 86.31 F
00:00:56.714 -> Average Cal sensor Temp.: 98.55 Ambient Temp.: 86.31 F
00:00:58.271 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 86.23 F
00:00:59.858 -> Average Cal sensor Temp.: 98.08 Ambient Temp.: 86.23 F
00:01:01.413 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 86.20 F
00:01:02.941 -> Average Cal sensor Temp.: 98.19 Ambient Temp.: 86.13 F
00:01:04.503 -> Average Cal sensor Temp.: 98.24 Ambient Temp.: 86.20 F
00:01:06.060 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 86.16 F
00:01:07.623 -> Average Cal sensor Temp.: 98.22 Ambient Temp.: 86.13 F
00:01:09.183 -> Average Cal sensor Temp.: 98.45 Ambient Temp.: 86.16 F
00:01:10.738 -> Average Cal sensor Temp.: 98.63 Ambient Temp.: 86.16 F
00:01:12.288 -> Average Cal sensor Temp.: 98.66 Ambient Temp.: 86.13 F
00:01:13.841 -> Average Cal sensor Temp.: 98.47 Ambient Temp.: 86.13 F
00:01:15.397 -> Average Cal sensor Temp.: 98.57 Ambient Temp.: 86.13 F
00:01:16.954 -> Average Cal sensor Temp.: 98.60 Ambient Temp.: 86.09 F
00:01:18.512 -> Average Cal sensor Temp.: 98.72 Ambient Temp.: 86.05 F
00:01:20.069 -> Average Cal sensor Temp.: 98.56 Ambient Temp.: 85.98 F
00:01:21.625 -> Average Cal sensor Temp.: 98.58 Ambient Temp.: 86.05 F
00:01:23.184 -> Average Cal sensor Temp.: 98.62 Ambient Temp.: 86.05 F
00:01:24.741 -> Average Cal sensor Temp.: 98.33 Ambient Temp.: 85.98 F
00:01:26.298 -> Average Cal sensor Temp.: 98.79 Ambient Temp.: 85.98 F
00:01:27.883 -> Average Cal sensor Temp.: 98.52 Ambient Temp.: 85.98 F
00:01:29.411 -> Average Cal sensor Temp.: 98.56 Ambient Temp.: 85.98 F
00:01:30.999 -> Average Cal sensor Temp.: 98.42 Ambient Temp.: 85.95 F
00:01:32.524 -> Average Cal sensor Temp.: 98.72 Ambient Temp.: 85.95 F
00:01:34.112 -> Average Cal sensor Temp.: 98.29 Ambient Temp.: 85.95 F
00:01:35.667 -> Average Cal sensor Temp.: 98.40 Ambient Temp.: 85.91 F
00:01:37.221 -> Average Cal sensor Temp.: 98.90 Ambient Temp.: 85.95 F
00:01:38.777 -> Average Cal sensor Temp.: 98.69 Ambient Temp.: 85.95 F
00:01:40.331 -> Average Cal sensor Temp.: 98.49 Ambient Temp.: 85.95 F
00:01:41.884 -> Average Cal sensor Temp.: 98.63 Ambient Temp.: 85.95 F
00:01:43.439 -> Average Cal sensor Temp.: 98.41 Ambient Temp.: 85.91 F
00:01:44.994 -> Average Cal sensor Temp.: 98.60 Ambient Temp.: 85.87 F
00:01:46.549 -> Average Cal sensor Temp.: 98.45 Ambient Temp.: 85.87 F
00:01:48.114 -> Average Cal sensor Temp.: 98.72 Ambient Temp.: 85.84 F
00:01:49.679 -> Average Cal sensor Temp.: 98.68 Ambient Temp.: 85.80 F
00:01:51.239 -> Average Cal sensor Temp.: 98.39 Ambient Temp.: 85.84 F
00:01:52.798 -> Average Cal sensor Temp.: 98.45 Ambient Temp.: 85.80 F
00:01:54.327 -> Average Cal sensor Temp.: 98.33 Ambient Temp.: 85.80 F
00:01:55.884 -> Average Cal sensor Temp.: 98.26 Ambient Temp.: 85.80 F
00:01:57.439 -> Average Cal sensor Temp.: 98.39 Ambient Temp.: 85.80 F
00:01:58.998 -> Average Cal sensor Temp.: 98.33 Ambient Temp.: 85.80 F
00:02:00.559 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 85.73 F
00:02:02.117 -> Average Cal sensor Temp.: 98.36 Ambient Temp.: 85.73 F
00:02:03.681 -> Average Cal sensor Temp.: 98.52 Ambient Temp.: 85.69 F
00:02:05.243 -> Average Cal sensor Temp.: 98.68 Ambient Temp.: 85.66 F
00:02:06.802 -> Average Cal sensor Temp.: 98.67 Ambient Temp.: 85.69 F
00:02:08.361 -> Average Cal sensor Temp.: 98.42 Ambient Temp.: 85.66 F
00:02:09.924 -> Average Cal sensor Temp.: 98.46 Ambient Temp.: 85.66 F
00:02:11.455 -> Average Cal sensor Temp.: 98.62 Ambient Temp.: 85.69 F
00:02:13.013 -> Average Cal sensor Temp.: 98.64 Ambient Temp.: 85.66 F
00:02:14.580 -> Average Cal sensor Temp.: 98.46 Ambient Temp.: 85.59 F
00:02:16.142 -> Average Cal sensor Temp.: 98.14 Ambient Temp.: 85.55 F
00:02:17.703 -> Average Cal sensor Temp.: 98.30 Ambient Temp.: 85.55 F
00:02:19.266 -> Average Cal sensor Temp.: 98.02 Ambient Temp.: 85.51 F
00:02:20.825 -> Average Cal sensor Temp.: 97.95 Ambient Temp.: 85.48 F
00:02:22.350 -> Average Cal sensor Temp.: 98.53 Ambient Temp.: 85.44 F
00:02:23.913 -> Average Cal sensor Temp.: 98.55 Ambient Temp.: 85.44 F
00:02:25.470 -> Average Cal sensor Temp.: 98.62 Ambient Temp.: 85.44 F
00:02:27.030 -> Average Cal sensor Temp.: 98.66 Ambient Temp.: 85.41 F
00:02:28.597 -> Average Cal sensor Temp.: 98.91 Ambient Temp.: 85.41 F
00:02:30.152 -> Average Cal sensor Temp.: 98.87 Ambient Temp.: 85.41 F
00:02:31.712 -> Average Cal sensor Temp.: 98.63 Ambient Temp.: 85.37 F
00:02:33.270 -> Average Cal sensor Temp.: 98.29 Ambient Temp.: 85.41 F
00:02:34.803 -> Average Cal sensor Temp.: 98.25 Ambient Temp.: 85.37 F
00:02:36.365 -> Average Cal sensor Temp.: 98.43 Ambient Temp.: 85.33 F
00:02:37.926 -> Average Cal sensor Temp.: 98.65 Ambient Temp.: 85.33 F
00:02:39.492 -> Average Cal sensor Temp.: 98.96 Ambient Temp.: 85.37 F
00:02:41.049 -> Average Cal sensor Temp.: 98.90 Ambient Temp.: 85.33 F
00:02:42.612 -> Average Cal sensor Temp.: 98.83 Ambient Temp.: 85.30 F
00:02:44.172 -> Average Cal sensor Temp.: 98.80 Ambient Temp.: 85.37 F
00:02:45.697 -> Average Cal sensor Temp.: 98.78 Ambient Temp.: 85.33 F
00:02:47.255 -> Average Cal sensor Temp.: 98.56 Ambient Temp.: 85.33 F
00:02:48.818 -> Average Cal sensor Temp.: 98.53 Ambient Temp.: 85.33 F
00:02:50.384 -> Average Cal sensor Temp.: 98.45 Ambient Temp.: 85.30 F
00:02:51.942 -> Average Cal sensor Temp.: 98.54 Ambient Temp.: 85.26 F
00:02:53.501 -> Average Cal sensor Temp.: 98.78 Ambient Temp.: 85.26 F
00:02:55.066 -> Average Cal sensor Temp.: 98.53 Ambient Temp.: 85.26 F
00:02:56.624 -> Average Cal sensor Temp.: 98.73 Ambient Temp.: 85.19 F
00:02:58.152 -> Average Cal sensor Temp.: 98.73 Ambient Temp.: 85.26 F
00:02:59.716 -> Average Cal sensor Temp.: 98.74 Ambient Temp.: 85.26 F
00:03:01.280 -> Average Cal sensor Temp.: 98.52 Ambient Temp.: 85.19 F
00:03:02.840 -> Average Cal sensor Temp.: 98.48 Ambient Temp.: 85.15 F
00:03:04.402 -> Average Cal sensor Temp.: 98.47 Ambient Temp.: 85.19 F
00:03:05.961 -> Average Cal sensor Temp.: 98.67 Ambient Temp.: 85.19 F
00:03:07.490 -> Average Cal sensor Temp.: 98.31 Ambient Temp.: 85.19 F
00:03:09.050 -> Average Cal sensor Temp.: 98.10 Ambient Temp.: 85.15 F
00:03:10.611 -> Average Cal sensor Temp.: 98.29 Ambient Temp.: 85.19 F
00:03:12.172 -> Average Cal sensor Temp.: 98.56 Ambient Temp.: 85.15 F
00:03:13.732 -> Average Cal sensor Temp.: 98.60 Ambient Temp.: 85.12 F
00:03:15.295 -> Average Cal sensor Temp.: 98.60 Ambient Temp.: 85.12 F
00:03:16.855 -> Average Cal sensor Temp.: 98.75 Ambient Temp.: 85.15 F
00:03:18.409 -> Average Cal sensor Temp.: 98.80 Ambient Temp.: 85.15 F
00:03:19.963 -> Average Cal sensor Temp.: 98.79 Ambient Temp.: 85.15 F
00:03:21.519 -> Average Cal sensor Temp.: 98.79 Ambient Temp.: 85.26 F
00:03:23.077 -> Average Cal sensor Temp.: 98.84 Ambient Temp.: 85.19 F
00:03:24.605 -> Average Cal sensor Temp.: 98.73 Ambient Temp.: 85.15 F
00:03:26.163 -> Average Cal sensor Temp.: 98.75 Ambient Temp.: 85.19 F
00:03:27.720 -> Average Cal sensor Temp.: 98.88 Ambient Temp.: 85.26 F
00:03:29.277 -> Average Cal sensor Temp.: 98.91 Ambient Temp.: 85.26 F
00:03:30.833 -> Average Cal sensor Temp.: 98.69 Ambient Temp.: 85.26 F
00:03:32.392 -> Average Cal sensor Temp.: 98.75 Ambient Temp.: 85.30 F
00:03:33.952 -> Average Cal sensor Temp.: 98.51 Ambient Temp.: 85.30 F
00:03:35.512 -> Average Cal sensor Temp.: 98.38 Ambient Temp.: 85.30 F
00:03:37.072 -> Average Cal sensor Temp.: 98.60 Ambient Temp.: 85.26 F
00:03:38.629 -> Average Cal sensor Temp.: 98.48 Ambient Temp.: 85.26 F
00:03:40.188 -> Average Cal sensor Temp.: 98.37 Ambient Temp.: 85.30 F
00:03:41.751 -> Average Cal sensor Temp.: 98.34 Ambient Temp.: 85.30 F
00:03:43.307 -> Average Cal sensor Temp.: 98.47 Ambient Temp.: 85.26 F
00:03:44.869 -> Average Cal sensor Temp.: 98.65 Ambient Temp.: 85.30 F
00:03:46.428 -> Average Cal sensor Temp.: 98.90 Ambient Temp.: 85.33 F
00:03:47.983 -> Average Cal sensor Temp.: 98.71 Ambient Temp.: 85.30 F
00:03:49.512 -> Average Cal sensor Temp.: 98.79 Ambient Temp.: 85.33 F
00:03:51.070 -> Average Cal sensor Temp.: 98.87 Ambient Temp.: 85.33 F
00:03:52.627 -> Average Cal sensor Temp.: 98.86 Ambient Temp.: 85.41 F
00:03:54.182 -> Average Cal sensor Temp.: 99.20 Ambient Temp.: 85.41 F
00:03:55.744 -> Average Cal sensor Temp.: 98.78 Ambient Temp.: 85.41 F
00:03:57.299 -> Average Cal sensor Temp.: 98.99 Ambient Temp.: 85.37 F
00:03:58.860 -> Average Cal sensor Temp.: 99.05 Ambient Temp.: 85.48 F
00:04:00.420 -> Average Cal sensor Temp.: 98.83 Ambient Temp.: 85.48 F
00:04:01.982 -> Average Cal sensor Temp.: 98.68 Ambient Temp.: 85.48 F
00:04:03.540 -> Average Cal sensor Temp.: 98.55 Ambient Temp.: 85.51 F
00:04:05.093 -> Average Cal sensor Temp.: 98.43 Ambient Temp.: 85.51 F
00:04:06.647 -> Average Cal sensor Temp.: 98.61 Ambient Temp.: 85.51 F
00:04:08.206 -> Average Cal sensor Temp.: 98.53 Ambient Temp.: 85.44 F
00:04:09.768 -> Average Cal sensor Temp.: 98.48 Ambient Temp.: 85.51 F
00:04:11.327 -> Average Cal sensor Temp.: 98.59 Ambient Temp.: 85.55 F
00:04:12.883 -> Average Cal sensor Temp.: 98.57 Ambient Temp.: 85.51 F
00:04:14.444 -> Average Cal sensor Temp.: 98.66 Ambient Temp.: 85.59 F
00:04:15.971 -> Average Cal sensor Temp.: 98.60 Ambient Temp.: 85.59 F
00:04:17.529 -> Average Cal sensor Temp.: 98.68 Ambient Temp.: 85.66 F
00:04:19.090 -> Average Cal sensor Temp.: 98.82 Ambient Temp.: 85.69 F
00:04:20.648 -> Average Cal sensor Temp.: 98.60 Ambient Temp.: 85.66 F
00:04:22.208 -> Average Cal sensor Temp.: 98.78 Ambient Temp.: 85.66 F
00:04:23.767 -> Average Cal sensor Temp.: 98.88 Ambient Temp.: 85.73 F
00:04:25.323 -> Average Cal sensor Temp.: 98.82 Ambient Temp.: 85.69 F
00:04:26.880 -> Average Cal sensor Temp.: 98.85 Ambient Temp.: 85.69 F
00:04:28.439 -> Average Cal sensor Temp.: 98.79 Ambient Temp.: 85.77 F
00:04:30.004 -> Average Cal sensor Temp.: 98.67 Ambient Temp.: 85.77 F
00:04:31.560 -> Average Cal sensor Temp.: 99.12 Ambient Temp.: 85.77 F
00:04:33.115 -> Average Cal sensor Temp.: 99.08 Ambient Temp.: 85.80 F
00:04:34.671 -> Average Cal sensor Temp.: 98.96 Ambient Temp.: 85.80 F
00:04:36.227 -> Average Cal sensor Temp.: 98.61 Ambient Temp.: 85.84 F
00:04:37.783 -> Average Cal sensor Temp.: 98.72 Ambient Temp.: 85.77 F
00:04:39.343 -> Average Cal sensor Temp.: 98.45 Ambient Temp.: 85.84 F
00:04:40.903 -> Average Cal sensor Temp.: 98.77 Ambient Temp.: 85.84 F
00:04:42.461 -> Average Cal sensor Temp.: 98.96 Ambient Temp.: 85.87 F
00:04:44.016 -> Average Cal sensor Temp.: 99.09 Ambient Temp.: 85.87 F
00:04:45.575 -> Average Cal sensor Temp.: 99.13 Ambient Temp.: 85.91 F
00:04:47.135 -> Average Cal sensor Temp.: 99.27 Ambient Temp.: 85.87 F
00:04:48.701 -> Average Cal sensor Temp.: 98.92 Ambient Temp.: 85.95 F
00:04:50.257 -> Average Cal sensor Temp.: 98.80 Ambient Temp.: 85.91 F
00:04:51.813 -> Average Cal sensor Temp.: 99.05 Ambient Temp.: 85.98 F
00:04:53.366 -> Average Cal sensor Temp.: 98.68 Ambient Temp.: 85.95 F
00:04:54.925 -> Average Cal sensor Temp.: 98.59 Ambient Temp.: 85.98 F
00:04:56.473 -> Average Cal sensor Temp.: 98.73 Ambient Temp.: 86.09 F
00:04:58.028 -> Average Cal sensor Temp.: 98.70 Ambient Temp.: 86.05 F
00:04:59.590 -> Average Cal sensor Temp.: 98.57 Ambient Temp.: 86.05 F
00:05:01.148 -> Average Cal sensor Temp.: 99.02 Ambient Temp.: 86.05 F
00:05:02.703 -> Average Cal sensor Temp.: 98.91 Ambient Temp.: 86.09 F
00:05:04.259 -> Average Cal sensor Temp.: 98.96 Ambient Temp.: 86.13 F
00:05:05.816 -> Average Cal sensor Temp.: 99.07 Ambient Temp.: 86.13 F
00:05:07.372 -> Average Cal sensor Temp.: 99.19 Ambient Temp.: 86.16 F
00:05:08.929 -> Average Cal sensor Temp.: 99.16 Ambient Temp.: 86.16 F
00:05:10.487 -> Average Cal sensor Temp.: 98.78 Ambient Temp.: 86.16 F
00:05:12.042 -> Average Cal sensor Temp.: 98.74 Ambient Temp.: 86.23 F
00:05:13.593 -> Average Cal sensor Temp.: 98.89 Ambient Temp.: 86.23 F
00:05:15.143 -> Average Cal sensor Temp.: 99.12 Ambient Temp.: 86.23 F
00:05:16.699 -> Average Cal sensor Temp.: 99.06 Ambient Temp.: 86.23 F
00:05:18.258 -> Average Cal sensor Temp.: 99.13 Ambient Temp.: 86.34 F
00:05:19.843 -> Average Cal sensor Temp.: 98.97 Ambient Temp.: 86.34 F
00:05:21.390 -> Average Cal sensor Temp.: 98.94 Ambient Temp.: 86.34 F
00:05:22.943 -> Average Cal sensor Temp.: 98.98 Ambient Temp.: 86.34 F
00:05:24.496 -> Average Cal sensor Temp.: 99.05 Ambient Temp.: 86.38 F
00:05:26.076 -> Average Cal sensor Temp.: 99.01 Ambient Temp.: 86.38 F
00:05:27.628 -> Average Cal sensor Temp.: 99.03 Ambient Temp.: 86.45 F
00:05:29.183 -> Average Cal sensor Temp.: 98.94 Ambient Temp.: 86.45 F
00:05:30.740 -> Average Cal sensor Temp.: 99.07 Ambient Temp.: 86.52 F
00:05:32.292 -> Average Cal sensor Temp.: 98.96 Ambient Temp.: 86.52 F
00:05:33.847 -> Average Cal sensor Temp.: 98.65 Ambient Temp.: 86.56 F
00:05:35.399 -> Average Cal sensor Temp.: 98.73 Ambient Temp.: 86.52 F
00:05:36.953 -> Average Cal sensor Temp.: 98.91 Ambient Temp.: 86.56 F
00:05:38.514 -> Average Cal sensor Temp.: 98.78 Ambient Temp.: 86.59 F
00:05:40.074 -> Average Cal sensor Temp.: 98.95 Ambient Temp.: 86.63 F
00:05:41.633 -> Average Cal sensor Temp.: 98.86 Ambient Temp.: 86.59 F
00:05:43.189 -> Average Cal sensor Temp.: 98.86 Ambient Temp.: 86.59 F
00:05:44.746 -> Average Cal sensor Temp.: 98.82 Ambient Temp.: 86.67 F
00:05:46.306 -> Average Cal sensor Temp.: 98.58 Ambient Temp.: 86.70 F
00:05:47.864 -> Average Cal sensor Temp.: 98.83 Ambient Temp.: 86.70 F
00:05:49.421 -> Average Cal sensor Temp.: 99.01 Ambient Temp.: 86.74 F
00:05:50.975 -> Average Cal sensor Temp.: 98.73 Ambient Temp.: 86.74 F
00:05:52.532 -> Average Cal sensor Temp.: 98.67 Ambient Temp.: 86.74 F
00:05:54.091 -> Average Cal sensor Temp.: 98.76 Ambient Temp.: 86.70 F
00:05:55.650 -> Average Cal sensor Temp.: 98.86 Ambient Temp.: 86.77 F
00:05:57.209 -> Average Cal sensor Temp.: 98.99 Ambient Temp.: 86.81 F
00:05:58.770 -> Average Cal sensor Temp.: 98.84 Ambient Temp.: 86.77 F
00:06:00.332 -> Average Cal sensor Temp.: 98.98 Ambient Temp.: 86.77 F
00:06:01.885 -> Average Cal sensor Temp.: 98.66 Ambient Temp.: 86.77 F
00:06:03.436 -> Average Cal sensor Temp.: 98.82 Ambient Temp.: 86.77 F
00:06:04.996 -> Average Cal sensor Temp.: 98.75 Ambient Temp.: 86.77 F
00:06:06.553 -> Average Cal sensor Temp.: 98.77 Ambient Temp.: 86.81 F
00:06:08.110 -> Average Cal sensor Temp.: 98.69 Ambient Temp.: 86.81 F
00:06:09.672 -> Average Cal sensor Temp.: 99.08 Ambient Temp.: 86.88 F
00:06:11.235 -> Average Cal sensor Temp.: 99.08 Ambient Temp.: 86.92 F
00:06:12.789 -> Average Cal sensor Temp.: 98.95 Ambient Temp.: 86.88 F
00:06:14.345 -> Average Cal sensor Temp.: 99.09 Ambient Temp.: 86.88 F
00:06:15.905 -> Average Cal sensor Temp.: 99.02 Ambient Temp.: 86.88 F
00:06:17.464 -> Average Cal sensor Temp.: 98.82 Ambient Temp.: 86.88 F
00:06:19.021 -> Average Cal sensor Temp.: 99.14 Ambient Temp.: 86.88 F
00:06:20.582 -> Average Cal sensor Temp.: 98.75 Ambient Temp.: 86.88 F
00:06:22.140 -> Average Cal sensor Temp.: 98.84 Ambient Temp.: 86.88 F
00:06:23.700 -> Average Cal sensor Temp.: 99.33 Ambient Temp.: 86.92 F
00:06:25.256 -> Average Cal sensor Temp.: 99.06 Ambient Temp.: 86.92 F
00:06:26.814 -> Average Cal sensor Temp.: 99.02 Ambient Temp.: 86.95 F
00:06:28.370 -> Average Cal sensor Temp.: 98.77 Ambient Temp.: 86.88 F
00:06:29.928 -> Average Cal sensor Temp.: 98.79 Ambient Temp.: 86.99 F
00:06:31.467 -> Average Cal sensor Temp.: 98.81 Ambient Temp.: 86.88 F
00:06:33.045 -> Average Cal sensor Temp.: 98.84 Ambient Temp.: 86.81 F
00:06:34.600 -> Average Cal sensor Temp.: 99.02 Ambient Temp.: 86.92 F
00:06:36.153 -> Average Cal sensor Temp.: 98.85 Ambient Temp.: 86.88 F
00:06:37.717 -> Average Cal sensor Temp.: 98.91 Ambient Temp.: 86.92 F
00:06:39.277 -> Average Cal sensor Temp.: 98.66 Ambient Temp.: 86.92 F
00:06:40.832 -> Average Cal sensor Temp.: 98.69 Ambient Temp.: 86.92 F
00:06:42.391 -> Average Cal sensor Temp.: 98.52 Ambient Temp.: 86.81 F
00:06:43.952 -> Average Cal sensor Temp.: 98.68 Ambient Temp.: 86.81 F
00:06:45.509 -> Average Cal sensor Temp.: 98.44 Ambient Temp.: 86.77 F
00:06:47.065 -> Average Cal sensor Temp.: 98.32 Ambient Temp.: 86.81 F
00:06:48.625 -> Average Cal sensor Temp.: 98.70 Ambient Temp.: 86.77 F
00:06:50.182 -> Average Cal sensor Temp.: 98.28 Ambient Temp.: 86.74 F
00:06:51.740 -> Average Cal sensor Temp.: 98.46 Ambient Temp.: 86.81 F
00:06:53.301 -> Average Cal sensor Temp.: 98.66 Ambient Temp.: 86.77 F
00:06:54.858 -> Average Cal sensor Temp.: 98.87 Ambient Temp.: 86.77 F
00:06:56.415 -> Average Cal sensor Temp.: 98.82 Ambient Temp.: 86.74 F
00:06:57.970 -> Average Cal sensor Temp.: 99.04 Ambient Temp.: 86.74 F
00:06:59.526 -> Average Cal sensor Temp.: 98.73 Ambient Temp.: 86.77 F
00:07:01.087 -> Average Cal sensor Temp.: 98.80 Ambient Temp.: 86.70 F
00:07:02.650 -> Average Cal sensor Temp.: 98.69 Ambient Temp.: 86.74 F
00:07:04.209 -> Average Cal sensor Temp.: 98.69 Ambient Temp.: 86.70 F
00:07:05.736 -> Average Cal sensor Temp.: 98.82 Ambient Temp.: 86.70 F
00:07:07.299 -> Average Cal sensor Temp.: 98.89 Ambient Temp.: 86.70 F
00:07:08.863 -> Average Cal sensor Temp.: 98.92 Ambient Temp.: 86.67 F
00:07:10.424 -> Average Cal sensor Temp.: 98.60 Ambient Temp.: 86.67 F
00:07:11.981 -> Average Cal sensor Temp.: 98.40 Ambient Temp.: 86.63 F
00:07:13.538 -> Average Cal sensor Temp.: 98.69 Ambient Temp.: 86.67 F
00:07:15.100 -> Average Cal sensor Temp.: 98.56 Ambient Temp.: 86.59 F
00:07:16.653 -> Average Cal sensor Temp.: 98.78 Ambient Temp.: 86.59 F
00:07:18.213 -> Average Cal sensor Temp.: 98.92 Ambient Temp.: 86.59 F
00:07:19.777 -> Average Cal sensor Temp.: 98.90 Ambient Temp.: 86.56 F
00:07:21.330 -> Average Cal sensor Temp.: 98.74 Ambient Temp.: 86.52 F
00:07:22.893 -> Average Cal sensor Temp.: 98.37 Ambient Temp.: 86.52 F
00:07:24.458 -> Average Cal sensor Temp.: 99.05 Ambient Temp.: 86.52 F
00:07:25.986 -> Average Cal sensor Temp.: 98.77 Ambient Temp.: 86.49 F
00:07:27.551 -> Average Cal sensor Temp.: 98.65 Ambient Temp.: 86.52 F
00:07:29.102 -> Average Cal sensor Temp.: 98.77 Ambient Temp.: 86.52 F
00:07:30.657 -> Average Cal sensor Temp.: 98.78 Ambient Temp.: 86.49 F
00:07:32.218 -> Average Cal sensor Temp.: 99.11 Ambient Temp.: 86.45 F
00:07:33.767 -> Average Cal sensor Temp.: 98.84 Ambient Temp.: 86.45 F
00:07:35.357 -> Average Cal sensor Temp.: 98.95 Ambient Temp.: 86.49 F
00:07:36.913 -> Average Cal sensor Temp.: 98.89 Ambient Temp.: 86.45 F
00:07:38.471 -> Average Cal sensor Temp.: 99.05 Ambient Temp.: 86.45 F
00:07:39.998 -> Average Cal sensor Temp.: 99.01 Ambient Temp.: 86.52 F
00:07:41.559 -> Average Cal sensor Temp.: 98.91 Ambient Temp.: 86.38 F
00:07:43.120 -> Average Cal sensor Temp.: 99.03 Ambient Temp.: 86.38 F
00:07:44.679 -> Average Cal sensor Temp.: 98.92 Ambient Temp.: 86.38 F
00:07:46.231 -> Average Cal sensor Temp.: 99.37 Ambient Temp.: 86.34 F
00:07:47.785 -> Average Cal sensor Temp.: 99.03 Ambient Temp.: 86.38 F
00:07:49.344 -> Average Cal sensor Temp.: 99.18 Ambient Temp.: 86.34 F
00:07:50.901 -> Average Cal sensor Temp.: 99.32 Ambient Temp.: 86.34 F
00:07:52.461 -> Average Cal sensor Temp.: 99.11 Ambient Temp.: 86.34 F
00:07:54.024 -> Average Cal sensor Temp.: 99.23 Ambient Temp.: 86.31 F
00:07:55.585 -> Average Cal sensor Temp.: 98.82 Ambient Temp.: 86.31 F
00:07:57.143 -> Average Cal sensor Temp.: 98.88 Ambient Temp.: 86.27 F
00:07:58.699 -> Average Cal sensor Temp.: 98.59 Ambient Temp.: 86.27 F
00:08:00.251 -> Average Cal sensor Temp.: 98.99 Ambient Temp.: 86.23 F
00:08:01.811 -> Average Cal sensor Temp.: 98.92 Ambient Temp.: 86.20 F
00:08:03.368 -> Average Cal sensor Temp.: 98.82 Ambient Temp.: 86.20 F
00:08:04.927 -> Average Cal sensor Temp.: 98.53 Ambient Temp.: 86.16 F
00:08:06.489 -> Average Cal sensor Temp.: 98.68 Ambient Temp.: 86.13 F
00:08:08.049 -> Average Cal sensor Temp.: 98.86 Ambient Temp.: 86.16 F
00:08:09.608 -> Average Cal sensor Temp.: 98.90 Ambient Temp.: 86.16 F
00:08:11.161 -> Average Cal sensor Temp.: 99.02 Ambient Temp.: 86.13 F
00:08:12.719 -> Average Cal sensor Temp.: 99.16 Ambient Temp.: 86.13 F
00:08:14.279 -> Average Cal sensor Temp.: 99.05 Ambient Temp.: 86.13 F
00:08:15.811 -> Average Cal sensor Temp.: 98.46 Ambient Temp.: 86.13 F
00:08:17.370 -> Average Cal sensor Temp.: 98.83 Ambient Temp.: 86.05 F
00:08:18.929 -> Average Cal sensor Temp.: 99.09 Ambient Temp.: 85.98 F
00:08:20.488 -> Average Cal sensor Temp.: 98.97 Ambient Temp.: 86.05 F
00:08:22.067 -> Average Cal sensor Temp.: 98.65 Ambient Temp.: 85.98 F
00:08:23.623 -> Average Cal sensor Temp.: 98.50 Ambient Temp.: 85.91 F
00:08:25.187 -> Average Cal sensor Temp.: 98.66 Ambient Temp.: 85.91 F
00:08:26.719 -> Average Cal sensor Temp.: 98.44 Ambient Temp.: 85.84 F
00:08:28.276 -> Average Cal sensor Temp.: 98.65 Ambient Temp.: 85.80 F
00:08:29.838 -> Average Cal sensor Temp.: 99.08 Ambient Temp.: 85.87 F
00:08:31.413 -> Average Cal sensor Temp.: 98.77 Ambient Temp.: 85.77 F
00:08:32.965 -> Average Cal sensor Temp.: 99.05 Ambient Temp.: 85.84 F
00:08:34.528 -> Average Cal sensor Temp.: 99.32 Ambient Temp.: 85.77 F
00:08:36.087 -> Average Cal sensor Temp.: 99.47 Ambient Temp.: 85.84 F
00:08:37.617 -> Average Cal sensor Temp.: 99.23 Ambient Temp.: 85.80 F
00:08:39.184 -> Average Cal sensor Temp.: 98.99 Ambient Temp.: 85.80 F
00:08:40.741 -> Average Cal sensor Temp.: 99.19 Ambient Temp.: 85.77 F
00:08:42.301 -> Average Cal sensor Temp.: 99.21 Ambient Temp.: 85.73 F
00:08:43.859 -> Average Cal sensor Temp.: 98.96 Ambient Temp.: 85.77 F
00:08:45.421 -> Average Cal sensor Temp.: 99.07 Ambient Temp.: 85.73 F
00:08:46.981 -> Average Cal sensor Temp.: 99.14 Ambient Temp.: 85.77 F
00:08:48.535 -> Average Cal sensor Temp.: 98.85 Ambient Temp.: 85.73 F
00:08:50.065 -> Average Cal sensor Temp.: 98.87 Ambient Temp.: 85.69 F
00:08:51.621 -> Average Cal sensor Temp.: 99.00 Ambient Temp.: 85.66 F
00:08:53.180 -> Average Cal sensor Temp.: 98.96 Ambient Temp.: 85.66 F
00:08:54.738 -> Average Cal sensor Temp.: 99.16 Ambient Temp.: 85.59 F
00:08:56.291 -> Average Cal sensor Temp.: 99.00 Ambient Temp.: 85.55 F
00:08:57.853 -> Average Cal sensor Temp.: 98.78 Ambient Temp.: 85.55 F
00:08:59.405 -> Average Cal sensor Temp.: 99.13 Ambient Temp.: 85.51 F
00:09:00.993 -> Average Cal sensor Temp.: 99.37 Ambient Temp.: 85.55 F
00:09:02.547 -> Average Cal sensor Temp.: 99.18 Ambient Temp.: 85.51 F
00:09:04.075 -> Average Cal sensor Temp.: 99.25 Ambient Temp.: 85.55 F
00:09:05.630 -> Average Cal sensor Temp.: 99.33 Ambient Temp.: 85.59 F
00:09:07.185 -> Average Cal sensor Temp.: 99.13 Ambient Temp.: 85.55 F
00:09:08.746 -> Average Cal sensor Temp.: 99.00 Ambient Temp.: 85.51 F
00:09:10.303 -> Average Cal sensor Temp.: 99.10 Ambient Temp.: 85.44 F
00:09:11.862 -> Average Cal sensor Temp.: 99.10 Ambient Temp.: 85.48 F
00:09:13.420 -> Average Cal sensor Temp.: 99.62 Ambient Temp.: 85.44 F
00:09:14.981 -> Average Cal sensor Temp.: 99.33 Ambient Temp.: 85.48 F
00:09:16.538 -> Average Cal sensor Temp.: 99.59 Ambient Temp.: 85.51 F
00:09:18.092 -> Average Cal sensor Temp.: 99.50 Ambient Temp.: 85.51 F
00:09:19.649 -> Average Cal sensor Temp.: 99.38 Ambient Temp.: 85.55 F
00:09:21.211 -> Average Cal sensor Temp.: 99.27 Ambient Temp.: 85.59 F
00:09:22.765 -> Average Cal sensor Temp.: 99.13 Ambient Temp.: 85.59 F
00:09:24.325 -> Average Cal sensor Temp.: 99.29 Ambient Temp.: 85.55 F
00:09:25.886 -> Average Cal sensor Temp.: 99.45 Ambient Temp.: 85.51 F
00:09:27.446 -> Average Cal sensor Temp.: 99.13 Ambient Temp.: 85.55 F
00:09:29.007 -> Average Cal sensor Temp.: 99.18 Ambient Temp.: 85.55 F
00:09:30.533 -> Average Cal sensor Temp.: 99.04 Ambient Temp.: 85.59 F
00:09:32.112 -> Average Cal sensor Temp.: 99.40 Ambient Temp.: 85.59 F
00:09:33.664 -> Average Cal sensor Temp.: 99.23 Ambient Temp.: 85.69 F
00:09:35.226 -> Average Cal sensor Temp.: 99.26 Ambient Temp.: 85.66 F
00:09:36.786 -> Average Cal sensor Temp.: 99.08 Ambient Temp.: 85.69 F
00:09:38.343 -> Average Cal sensor Temp.: 99.28 Ambient Temp.: 85.69 F
00:09:39.905 -> Average Cal sensor Temp.: 99.41 Ambient Temp.: 85.69 F
00:09:41.461 -> Average Cal sensor Temp.: 99.06 Ambient Temp.: 85.69 F
00:09:43.022 -> Average Cal sensor Temp.: 99.62 Ambient Temp.: 85.73 F
00:09:44.581 -> Average Cal sensor Temp.: 99.45 Ambient Temp.: 85.73 F
00:09:46.140 -> Average Cal sensor Temp.: 99.37 Ambient Temp.: 85.77 F
00:09:47.692 -> Average Cal sensor Temp.: 99.23 Ambient Temp.: 85.77 F
00:09:49.251 -> Average Cal sensor Temp.: 99.17 Ambient Temp.: 85.77 F
00:09:50.774 -> Average Cal sensor Temp.: 99.09 Ambient Temp.: 85.77 F
00:09:52.334 -> Average Cal sensor Temp.: 99.12 Ambient Temp.: 85.80 F
00:09:53.894 -> Average Cal sensor Temp.: 99.45 Ambient Temp.: 85.80 F
00:09:55.451 -> Average Cal sensor Temp.: 99.36 Ambient Temp.: 85.80 F
00:09:57.005 -> Average Cal sensor Temp.: 99.36 Ambient Temp.: 85.80 F
00:09:58.561 -> Average Cal sensor Temp.: 99.27 Ambient Temp.: 85.87 F
00:10:00.143 -> Average Cal sensor Temp.: 99.41 Ambient Temp.: 85.87 F
00:10:01.700 -> Average Cal sensor Temp.: 99.29 Ambient Temp.: 85.84 F
00:10:03.260 -> Average Cal sensor Temp.: 99.42 Ambient Temp.: 85.84 F
00:10:04.816 -> Average Cal sensor Temp.: 99.44 Ambient Temp.: 85.87 F
00:10:06.374 -> Average Cal sensor Temp.: 99.50 Ambient Temp.: 85.87 F
00:10:07.930 -> Average Cal sensor Temp.: 99.50 Ambient Temp.: 85.87 F
00:10:09.488 -> Average Cal sensor Temp.: 99.45 Ambient Temp.: 85.87 F
00:10:11.043 -> Average Cal sensor Temp.: 99.70 Ambient Temp.: 85.95 F
00:10:12.598 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 85.91 F
00:10:14.157 -> Average Cal sensor Temp.: 99.52 Ambient Temp.: 85.95 F
00:10:15.714 -> Average Cal sensor Temp.: 99.50 Ambient Temp.: 86.05 F
00:10:17.272 -> Average Cal sensor Temp.: 99.59 Ambient Temp.: 86.05 F
00:10:18.856 -> Average Cal sensor Temp.: 99.45 Ambient Temp.: 86.05 F
00:10:20.417 -> Average Cal sensor Temp.: 99.35 Ambient Temp.: 86.05 F
00:10:21.973 -> Average Cal sensor Temp.: 99.17 Ambient Temp.: 86.09 F
00:10:23.498 -> Average Cal sensor Temp.: 99.42 Ambient Temp.: 86.09 F
00:10:25.064 -> Average Cal sensor Temp.: 99.37 Ambient Temp.: 86.09 F
00:10:26.626 -> Average Cal sensor Temp.: 99.58 Ambient Temp.: 86.09 F
00:10:28.185 -> Average Cal sensor Temp.: 99.59 Ambient Temp.: 86.13 F
00:10:29.750 -> Average Cal sensor Temp.: 99.65 Ambient Temp.: 86.13 F
00:10:31.327 -> Average Cal sensor Temp.: 99.64 Ambient Temp.: 86.13 F
00:10:32.885 -> Average Cal sensor Temp.: 99.68 Ambient Temp.: 86.20 F
00:10:34.446 -> Average Cal sensor Temp.: 99.82 Ambient Temp.: 86.20 F
00:10:36.005 -> Average Cal sensor Temp.: 99.54 Ambient Temp.: 86.20 F
00:10:37.566 -> Average Cal sensor Temp.: 99.55 Ambient Temp.: 86.27 F
00:10:39.123 -> Average Cal sensor Temp.: 99.59 Ambient Temp.: 86.20 F
00:10:40.683 -> Average Cal sensor Temp.: 99.59 Ambient Temp.: 86.31 F
00:10:42.240 -> Average Cal sensor Temp.: 99.67 Ambient Temp.: 86.23 F
00:10:43.828 -> Average Cal sensor Temp.: 99.34 Ambient Temp.: 86.27 F
00:10:45.357 -> Average Cal sensor Temp.: 99.30 Ambient Temp.: 86.31 F
00:10:46.918 -> Average Cal sensor Temp.: 99.54 Ambient Temp.: 86.27 F
00:10:48.475 -> Average Cal sensor Temp.: 99.60 Ambient Temp.: 86.27 F
00:10:50.069 -> Average Cal sensor Temp.: 99.75 Ambient Temp.: 86.31 F
00:10:51.619 -> Average Cal sensor Temp.: 99.41 Ambient Temp.: 86.34 F
00:10:53.171 -> Average Cal sensor Temp.: 99.57 Ambient Temp.: 86.38 F
00:10:54.724 -> Average Cal sensor Temp.: 99.61 Ambient Temp.: 86.38 F
00:10:56.312 -> Average Cal sensor Temp.: 99.48 Ambient Temp.: 86.38 F
00:10:57.856 -> Average Cal sensor Temp.: 99.24 Ambient Temp.: 86.45 F
00:10:59.410 -> Average Cal sensor Temp.: 99.46 Ambient Temp.: 86.45 F
00:11:00.962 -> Average Cal sensor Temp.: 99.54 Ambient Temp.: 86.49 F
00:11:02.543 -> Average Cal sensor Temp.: 99.45 Ambient Temp.: 86.49 F
00:11:04.103 -> Average Cal sensor Temp.: 99.40 Ambient Temp.: 86.49 F
00:11:05.654 -> Average Cal sensor Temp.: 99.48 Ambient Temp.: 86.45 F
00:11:07.214 -> Average Cal sensor Temp.: 99.53 Ambient Temp.: 86.52 F
00:11:08.772 -> Average Cal sensor Temp.: 99.60 Ambient Temp.: 86.56 F
00:11:10.333 -> Average Cal sensor Temp.: 99.57 Ambient Temp.: 86.56 F
00:11:11.894 -> Average Cal sensor Temp.: 99.61 Ambient Temp.: 86.52 F
00:11:13.455 -> Average Cal sensor Temp.: 99.45 Ambient Temp.: 86.59 F
00:11:15.010 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 86.59 F
00:11:16.568 -> Average Cal sensor Temp.: 99.85 Ambient Temp.: 86.59 F
00:11:18.121 -> Average Cal sensor Temp.: 99.65 Ambient Temp.: 86.63 F
00:11:19.716 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 86.67 F
00:11:21.273 -> Average Cal sensor Temp.: 99.57 Ambient Temp.: 86.70 F
00:11:22.837 -> Average Cal sensor Temp.: 99.80 Ambient Temp.: 86.70 F
00:11:24.396 -> Average Cal sensor Temp.: 100.00 Ambient Temp.: 86.70 F
00:11:25.949 -> Average Cal sensor Temp.: 99.99 Ambient Temp.: 86.74 F
00:11:27.507 -> Average Cal sensor Temp.: 99.79 Ambient Temp.: 86.74 F
00:11:29.070 -> Average Cal sensor Temp.: 99.79 Ambient Temp.: 86.88 F
00:11:30.633 -> Average Cal sensor Temp.: 99.73 Ambient Temp.: 86.81 F
00:11:32.191 -> Average Cal sensor Temp.: 99.72 Ambient Temp.: 86.88 F
00:11:33.755 -> Average Cal sensor Temp.: 99.85 Ambient Temp.: 86.95 F
00:11:35.342 -> Average Cal sensor Temp.: 99.82 Ambient Temp.: 86.88 F
00:11:36.899 -> Average Cal sensor Temp.: 100.04 Ambient Temp.: 86.99 F
00:11:38.457 -> Average Cal sensor Temp.: 99.64 Ambient Temp.: 86.92 F
00:11:40.017 -> Average Cal sensor Temp.: 99.79 Ambient Temp.: 86.99 F
00:11:41.607 -> Average Cal sensor Temp.: 99.67 Ambient Temp.: 86.95 F
00:11:43.160 -> Average Cal sensor Temp.: 99.67 Ambient Temp.: 87.03 F
00:11:44.712 -> Average Cal sensor Temp.: 99.72 Ambient Temp.: 87.03 F
00:11:46.266 -> Average Cal sensor Temp.: 99.87 Ambient Temp.: 87.03 F
00:11:47.822 -> Average Cal sensor Temp.: 99.62 Ambient Temp.: 86.99 F
00:11:49.414 -> Average Cal sensor Temp.: 99.61 Ambient Temp.: 87.06 F
00:11:50.973 -> Average Cal sensor Temp.: 99.68 Ambient Temp.: 87.10 F
00:11:52.528 -> Average Cal sensor Temp.: 99.81 Ambient Temp.: 87.10 F
00:11:54.086 -> Average Cal sensor Temp.: 99.60 Ambient Temp.: 87.10 F
00:11:55.648 -> Average Cal sensor Temp.: 99.60 Ambient Temp.: 87.10 F
00:11:57.201 -> Average Cal sensor Temp.: 99.53 Ambient Temp.: 87.13 F
00:11:58.761 -> Average Cal sensor Temp.: 99.62 Ambient Temp.: 87.10 F
00:12:00.322 -> Average Cal sensor Temp.: 99.60 Ambient Temp.: 87.13 F
00:12:01.913 -> Average Cal sensor Temp.: 99.81 Ambient Temp.: 87.13 F
00:12:03.473 -> Average Cal sensor Temp.: 99.68 Ambient Temp.: 87.17 F
00:12:05.036 -> Average Cal sensor Temp.: 99.54 Ambient Temp.: 87.17 F
00:12:06.596 -> Average Cal sensor Temp.: 99.85 Ambient Temp.: 87.21 F
00:12:08.157 -> Average Cal sensor Temp.: 99.95 Ambient Temp.: 87.21 F
00:12:09.713 -> Average Cal sensor Temp.: 100.14 Ambient Temp.: 87.21 F
00:12:11.268 -> Average Cal sensor Temp.: 99.99 Ambient Temp.: 87.21 F
00:12:12.823 -> Average Cal sensor Temp.: 97.13 Ambient Temp.: 87.28 F
00:12:14.380 -> Average Cal sensor Temp.: 99.49 Ambient Temp.: 87.35 F
00:12:15.945 -> Average Cal sensor Temp.: 100.36 Ambient Temp.: 87.31 F
00:12:17.502 -> Average Cal sensor Temp.: 100.25 Ambient Temp.: 87.35 F
00:12:19.057 -> Average Cal sensor Temp.: 100.46 Ambient Temp.: 87.35 F
00:12:20.618 -> Average Cal sensor Temp.: 100.13 Ambient Temp.: 87.35 F
00:12:22.209 -> Average Cal sensor Temp.: 100.15 Ambient Temp.: 87.39 F
00:12:23.771 -> Average Cal sensor Temp.: 100.17 Ambient Temp.: 87.35 F
00:12:25.332 -> Average Cal sensor Temp.: 100.15 Ambient Temp.: 87.39 F
00:12:26.889 -> Average Cal sensor Temp.: 100.58 Ambient Temp.: 87.42 F
00:12:28.444 -> Average Cal sensor Temp.: 100.39 Ambient Temp.: 87.35 F
00:12:30.034 -> Average Cal sensor Temp.: 100.28 Ambient Temp.: 87.39 F
00:12:31.596 -> Average Cal sensor Temp.: 100.22 Ambient Temp.: 87.42 F
00:12:33.156 -> Average Cal sensor Temp.: 100.14 Ambient Temp.: 87.42 F
00:12:34.713 -> Average Cal sensor Temp.: 100.35 Ambient Temp.: 87.42 F
00:12:36.271 -> Average Cal sensor Temp.: 100.06 Ambient Temp.: 87.39 F
00:12:37.828 -> Average Cal sensor Temp.: 100.07 Ambient Temp.: 87.42 F
00:12:39.422 -> Average Cal sensor Temp.: 100.08 Ambient Temp.: 87.39 F
00:12:40.980 -> Average Cal sensor Temp.: 100.05 Ambient Temp.: 87.35 F
00:12:42.536 -> Average Cal sensor Temp.: 100.07 Ambient Temp.: 87.35 F
00:12:44.098 -> Average Cal sensor Temp.: 99.88 Ambient Temp.: 87.35 F
00:12:45.682 -> Average Cal sensor Temp.: 99.68 Ambient Temp.: 87.31 F
00:12:47.235 -> Average Cal sensor Temp.: 100.24 Ambient Temp.: 87.31 F
00:12:48.788 -> Average Cal sensor Temp.: 100.12 Ambient Temp.: 87.28 F
00:12:50.342 -> Average Cal sensor Temp.: 99.85 Ambient Temp.: 87.28 F
00:12:51.935 -> Average Cal sensor Temp.: 100.19 Ambient Temp.: 87.21 F
00:12:53.489 -> Average Cal sensor Temp.: 99.66 Ambient Temp.: 87.28 F
00:12:55.045 -> Average Cal sensor Temp.: 100.02 Ambient Temp.: 87.17 F
00:12:56.601 -> Average Cal sensor Temp.: 100.24 Ambient Temp.: 87.21 F
00:12:58.189 -> Average Cal sensor Temp.: 100.14 Ambient Temp.: 87.28 F
00:12:59.748 -> Average Cal sensor Temp.: 100.09 Ambient Temp.: 87.21 F
00:13:01.308 -> Average Cal sensor Temp.: 100.07 Ambient Temp.: 87.21 F
00:13:02.862 -> Average Cal sensor Temp.: 100.13 Ambient Temp.: 87.17 F
00:13:04.452 -> Average Cal sensor Temp.: 100.45 Ambient Temp.: 87.13 F
00:13:06.004 -> Average Cal sensor Temp.: 100.29 Ambient Temp.: 87.21 F
00:13:07.565 -> Average Cal sensor Temp.: 100.38 Ambient Temp.: 87.17 F
00:13:09.129 -> Average Cal sensor Temp.: 100.09 Ambient Temp.: 87.17 F
00:13:10.693 -> Average Cal sensor Temp.: 100.25 Ambient Temp.: 87.21 F
00:13:12.255 -> Average Cal sensor Temp.: 100.18 Ambient Temp.: 87.13 F
00:13:13.813 -> Average Cal sensor Temp.: 100.09 Ambient Temp.: 87.13 F
00:13:15.381 -> Average Cal sensor Temp.: 100.27 Ambient Temp.: 87.13 F
00:13:16.943 -> Average Cal sensor Temp.: 100.69 Ambient Temp.: 87.10 F
00:13:18.531 -> Average Cal sensor Temp.: 100.67 Ambient Temp.: 87.17 F
00:13:20.084 -> Average Cal sensor Temp.: 100.50 Ambient Temp.: 87.13 F
00:13:21.641 -> Average Cal sensor Temp.: 100.65 Ambient Temp.: 87.10 F
00:13:23.203 -> Average Cal sensor Temp.: 100.37 Ambient Temp.: 87.10 F
00:13:24.767 -> Average Cal sensor Temp.: 100.53 Ambient Temp.: 87.06 F
00:13:26.326 -> Average Cal sensor Temp.: 100.48 Ambient Temp.: 87.13 F
00:13:27.919 -> Average Cal sensor Temp.: 100.22 Ambient Temp.: 87.13 F
00:13:29.478 -> Average Cal sensor Temp.: 100.33 Ambient Temp.: 87.10 F
00:13:31.035 -> Average Cal sensor Temp.: 100.45 Ambient Temp.: 87.06 F
00:13:32.593 -> Average Cal sensor Temp.: 100.43 Ambient Temp.: 87.06 F
00:13:34.151 -> Average Cal sensor Temp.: 99.92 Ambient Temp.: 87.03 F
00:13:35.713 -> Average Cal sensor Temp.: 100.04 Ambient Temp.: 86.95 F
00:13:37.302 -> Average Cal sensor Temp.: 100.34 Ambient Temp.: 86.95 F
00:13:38.857 -> Average Cal sensor Temp.: 99.96 Ambient Temp.: 86.95 F
00:13:40.414 -> Average Cal sensor Temp.: 100.04 Ambient Temp.: 86.92 F
00:13:42.004 -> Average Cal sensor Temp.: 100.35 Ambient Temp.: 86.92 F
00:13:43.564 -> Average Cal sensor Temp.: 100.49 Ambient Temp.: 86.92 F
00:13:45.121 -> Average Cal sensor Temp.: 100.37 Ambient Temp.: 86.81 F
00:13:46.679 -> Average Cal sensor Temp.: 100.33 Ambient Temp.: 86.92 F
00:13:48.235 -> Average Cal sensor Temp.: 100.55 Ambient Temp.: 86.77 F
00:13:49.830 -> Average Cal sensor Temp.: 100.24 Ambient Temp.: 86.81 F
00:13:51.387 -> Average Cal sensor Temp.: 99.99 Ambient Temp.: 86.74 F
00:13:52.950 -> Average Cal sensor Temp.: 99.96 Ambient Temp.: 86.70 F
00:13:54.509 -> Average Cal sensor Temp.: 100.24 Ambient Temp.: 86.70 F
00:13:56.066 -> Average Cal sensor Temp.: 100.53 Ambient Temp.: 86.70 F
00:13:57.624 -> Average Cal sensor Temp.: 100.38 Ambient Temp.: 86.67 F
00:13:59.216 -> Average Cal sensor Temp.: 100.54 Ambient Temp.: 86.67 F
00:14:00.771 -> Average Cal sensor Temp.: 100.53 Ambient Temp.: 86.63 F
00:14:02.327 -> Average Cal sensor Temp.: 100.20 Ambient Temp.: 86.63 F
00:14:03.914 -> Average Cal sensor Temp.: 100.30 Ambient Temp.: 86.59 F
00:14:05.471 -> Average Cal sensor Temp.: 100.37 Ambient Temp.: 86.56 F
00:14:07.025 -> Average Cal sensor Temp.: 100.30 Ambient Temp.: 86.56 F
00:14:08.617 -> Average Cal sensor Temp.: 100.41 Ambient Temp.: 86.59 F
00:14:10.175 -> Average Cal sensor Temp.: 99.88 Ambient Temp.: 86.49 F
00:14:11.730 -> Average Cal sensor Temp.: 100.06 Ambient Temp.: 86.49 F
00:14:13.285 -> Average Cal sensor Temp.: 100.09 Ambient Temp.: 86.45 F
00:14:14.879 -> Average Cal sensor Temp.: 100.31 Ambient Temp.: 86.45 F
00:14:16.439 -> Average Cal sensor Temp.: 100.15 Ambient Temp.: 86.34 F
00:14:17.998 -> Average Cal sensor Temp.: 100.26 Ambient Temp.: 86.38 F
00:14:19.559 -> Average Cal sensor Temp.: 100.30 Ambient Temp.: 86.34 F
00:14:21.118 -> Average Cal sensor Temp.: 99.94 Ambient Temp.: 86.34 F
00:14:22.679 -> Average Cal sensor Temp.: 100.02 Ambient Temp.: 86.27 F
00:14:24.269 -> Average Cal sensor Temp.: 100.03 Ambient Temp.: 86.27 F
00:14:25.831 -> Average Cal sensor Temp.: 99.79 Ambient Temp.: 86.20 F
00:14:27.386 -> Average Cal sensor Temp.: 100.07 Ambient Temp.: 86.16 F
00:14:28.944 -> Average Cal sensor Temp.: 100.34 Ambient Temp.: 86.20 F
00:14:30.499 -> Average Cal sensor Temp.: 100.72 Ambient Temp.: 86.16 F
00:14:32.082 -> Average Cal sensor Temp.: 100.44 Ambient Temp.: 86.16 F
00:14:33.639 -> Average Cal sensor Temp.: 100.27 Ambient Temp.: 86.05 F
00:14:35.196 -> Average Cal sensor Temp.: 100.42 Ambient Temp.: 86.09 F
00:14:36.788 -> Average Cal sensor Temp.: 100.45 Ambient Temp.: 86.05 F
00:14:38.345 -> Average Cal sensor Temp.: 100.62 Ambient Temp.: 86.05 F
00:14:39.903 -> Average Cal sensor Temp.: 100.44 Ambient Temp.: 85.98 F
00:14:41.459 -> Average Cal sensor Temp.: 100.72 Ambient Temp.: 85.98 F
00:14:43.053 -> Average Cal sensor Temp.: 100.48 Ambient Temp.: 85.95 F
00:14:44.617 -> Average Cal sensor Temp.: 100.54 Ambient Temp.: 85.87 F
00:14:46.182 -> Average Cal sensor Temp.: 100.50 Ambient Temp.: 85.87 F
00:14:47.743 -> Average Cal sensor Temp.: 100.09 Ambient Temp.: 85.91 F
00:14:49.299 -> Average Cal sensor Temp.: 100.39 Ambient Temp.: 85.91 F
00:14:50.854 -> Average Cal sensor Temp.: 100.10 Ambient Temp.: 85.80 F
00:14:52.448 -> Average Cal sensor Temp.: 99.86 Ambient Temp.: 85.77 F
00:14:54.008 -> Average Cal sensor Temp.: 100.11 Ambient Temp.: 85.77 F
00:14:55.567 -> Average Cal sensor Temp.: 100.37 Ambient Temp.: 85.73 F
00:14:57.120 -> Average Cal sensor Temp.: 100.24 Ambient Temp.: 85.77 F
00:14:58.679 -> Average Cal sensor Temp.: 100.49 Ambient Temp.: 85.66 F
00:15:00.266 -> Average Cal sensor Temp.: 100.67 Ambient Temp.: 85.69 F
00:15:01.814 -> Average Cal sensor Temp.: 100.76 Ambient Temp.: 85.66 F
00:15:03.403 -> Average Cal sensor Temp.: 100.54 Ambient Temp.: 85.59 F
00:15:04.960 -> Average Cal sensor Temp.: 100.27 Ambient Temp.: 85.59 F
00:15:06.517 -> Average Cal sensor Temp.: 100.38 Ambient Temp.: 85.55 F
00:15:08.075 -> Average Cal sensor Temp.: 100.56 Ambient Temp.: 85.55 F
00:15:09.632 -> Average Cal sensor Temp.: 100.82 Ambient Temp.: 85.59 F
00:15:11.224 -> Average Cal sensor Temp.: 100.52 Ambient Temp.: 85.55 F
00:15:12.776 -> Average Cal sensor Temp.: 100.84 Ambient Temp.: 85.51 F
00:15:14.331 -> Average Cal sensor Temp.: 100.83 Ambient Temp.: 85.55 F
00:15:15.921 -> Average Cal sensor Temp.: 101.08 Ambient Temp.: 85.51 F
00:15:17.471 -> Average Cal sensor Temp.: 101.01 Ambient Temp.: 85.51 F
00:15:19.026 -> Average Cal sensor Temp.: 100.83 Ambient Temp.: 85.51 F
00:15:20.581 -> Average Cal sensor Temp.: 100.73 Ambient Temp.: 85.51 F
00:15:22.172 -> Average Cal sensor Temp.: 100.73 Ambient Temp.: 85.51 F
00:15:23.727 -> Average Cal sensor Temp.: 100.63 Ambient Temp.: 85.48 F
00:15:25.282 -> Average Cal sensor Temp.: 100.66 Ambient Temp.: 85.48 F
00:15:26.865 -> Average Cal sensor Temp.: 100.81 Ambient Temp.: 85.48 F
00:15:28.423 -> Average Cal sensor Temp.: 100.78 Ambient Temp.: 85.44 F
00:15:29.984 -> Average Cal sensor Temp.: 100.92 Ambient Temp.: 85.51 F
00:15:31.546 -> Average Cal sensor Temp.: 100.64 Ambient Temp.: 85.48 F
00:15:33.102 -> Average Cal sensor Temp.: 100.76 Ambient Temp.: 85.48 F
00:15:34.701 -> Average Cal sensor Temp.: 100.66 Ambient Temp.: 85.48 F
00:15:36.259 -> Average Cal sensor Temp.: 100.44 Ambient Temp.: 85.41 F
00:15:37.819 -> Average Cal sensor Temp.: 100.74 Ambient Temp.: 85.48 F
00:15:39.380 -> Average Cal sensor Temp.: 100.76 Ambient Temp.: 85.48 F
00:15:40.941 -> Average Cal sensor Temp.: 100.80 Ambient Temp.: 85.48 F
00:15:42.531 -> Average Cal sensor Temp.: 100.83 Ambient Temp.: 85.51 F
00:15:44.092 -> Average Cal sensor Temp.: 100.76 Ambient Temp.: 85.48 F
00:15:45.650 -> Average Cal sensor Temp.: 100.56 Ambient Temp.: 85.44 F
00:15:47.214 -> Average Cal sensor Temp.: 100.65 Ambient Temp.: 85.48 F
00:15:48.778 -> Average Cal sensor Temp.: 100.65 Ambient Temp.: 85.48 F
00:15:50.334 -> Average Cal sensor Temp.: 100.53 Ambient Temp.: 85.51 F
00:15:51.894 -> Average Cal sensor Temp.: 100.76 Ambient Temp.: 85.55 F
00:15:53.456 -> Average Cal sensor Temp.: 100.44 Ambient Temp.: 85.55 F
00:15:55.050 -> Average Cal sensor Temp.: 100.36 Ambient Temp.: 85.51 F
00:15:56.602 -> Average Cal sensor Temp.: 101.00 Ambient Temp.: 85.51 F
00:15:58.164 -> Average Cal sensor Temp.: 100.33 Ambient Temp.: 85.48 F
00:15:59.724 -> Average Cal sensor Temp.: 100.50 Ambient Temp.: 85.51 F
00:16:01.289 -> Average Cal sensor Temp.: 100.41 Ambient Temp.: 85.48 F
00:16:02.848 -> Average Cal sensor Temp.: 100.58 Ambient Temp.: 85.59 F
00:16:04.443 -> Average Cal sensor Temp.: 100.61 Ambient Temp.: 85.51 F
00:16:05.999 -> Average Cal sensor Temp.: 100.67 Ambient Temp.: 85.51 F
00:16:07.556 -> Average Cal sensor Temp.: 100.78 Ambient Temp.: 85.51 F
00:16:09.121 -> Average Cal sensor Temp.: 100.65 Ambient Temp.: 85.59 F
00:16:10.680 -> Average Cal sensor Temp.: 100.80 Ambient Temp.: 85.55 F
00:16:12.240 -> Average Cal sensor Temp.: 101.08 Ambient Temp.: 85.59 F
00:16:13.804 -> Average Cal sensor Temp.: 101.27 Ambient Temp.: 85.59 F
00:16:15.364 -> Average Cal sensor Temp.: 101.00 Ambient Temp.: 85.66 F
00:16:16.957 -> Average Cal sensor Temp.: 100.95 Ambient Temp.: 85.66 F
00:16:18.519 -> Average Cal sensor Temp.: 100.90 Ambient Temp.: 85.69 F
00:16:20.078 -> Average Cal sensor Temp.: 100.94 Ambient Temp.: 85.77 F
00:16:21.637 -> Average Cal sensor Temp.: 100.90 Ambient Temp.: 85.73 F
00:16:23.195 -> Average Cal sensor Temp.: 100.96 Ambient Temp.: 85.77 F
00:16:24.756 -> Average Cal sensor Temp.: 100.67 Ambient Temp.: 85.73 F
00:16:26.348 -> Average Cal sensor Temp.: 101.02 Ambient Temp.: 85.80 F
00:16:27.903 -> Average Cal sensor Temp.: 100.91 Ambient Temp.: 85.87 F
00:16:29.464 -> Average Cal sensor Temp.: 100.58 Ambient Temp.: 85.84 F
00:16:31.023 -> Average Cal sensor Temp.: 101.13 Ambient Temp.: 85.87 F
00:16:32.583 -> Average Cal sensor Temp.: 100.98 Ambient Temp.: 85.87 F
00:16:34.176 -> Average Cal sensor Temp.: 101.02 Ambient Temp.: 85.87 F
00:16:35.740 -> Average Cal sensor Temp.: 100.76 Ambient Temp.: 85.87 F
00:16:37.304 -> Average Cal sensor Temp.: 100.64 Ambient Temp.: 85.91 F
00:16:38.861 -> Average Cal sensor Temp.: 100.83 Ambient Temp.: 85.95 F
00:16:40.419 -> Average Cal sensor Temp.: 100.73 Ambient Temp.: 85.98 F
00:16:41.980 -> Average Cal sensor Temp.: 100.86 Ambient Temp.: 85.98 F
00:16:43.540 -> Average Cal sensor Temp.: 100.89 Ambient Temp.: 86.05 F
00:16:45.133 -> Average Cal sensor Temp.: 101.10 Ambient Temp.: 86.05 F
00:16:46.688 -> Average Cal sensor Temp.: 101.04 Ambient Temp.: 86.09 F
00:16:48.250 -> Average Cal sensor Temp.: 100.74 Ambient Temp.: 86.05 F
00:16:49.812 -> Average Cal sensor Temp.: 100.73 Ambient Temp.: 86.13 F
00:16:51.367 -> Average Cal sensor Temp.: 100.88 Ambient Temp.: 86.13 F
00:16:52.962 -> Average Cal sensor Temp.: 100.77 Ambient Temp.: 86.16 F
00:16:54.523 -> Average Cal sensor Temp.: 100.89 Ambient Temp.: 86.16 F
00:16:56.081 -> Average Cal sensor Temp.: 100.84 Ambient Temp.: 86.16 F
00:16:57.639 -> Average Cal sensor Temp.: 100.89 Ambient Temp.: 86.20 F
00:16:59.200 -> Average Cal sensor Temp.: 100.95 Ambient Temp.: 86.23 F
00:17:00.761 -> Average Cal sensor Temp.: 100.94 Ambient Temp.: 86.27 F
00:17:02.358 -> Average Cal sensor Temp.: 101.14 Ambient Temp.: 86.27 F
00:17:03.917 -> Average Cal sensor Temp.: 101.11 Ambient Temp.: 86.27 F
00:17:05.475 -> Average Cal sensor Temp.: 101.09 Ambient Temp.: 86.27 F
00:17:07.029 -> Average Cal sensor Temp.: 101.31 Ambient Temp.: 86.34 F
00:17:08.589 -> Average Cal sensor Temp.: 101.21 Ambient Temp.: 86.34 F
00:17:10.144 -> Average Cal sensor Temp.: 100.98 Ambient Temp.: 86.38 F
00:17:11.735 -> Average Cal sensor Temp.: 100.92 Ambient Temp.: 86.38 F
00:17:13.287 -> Average Cal sensor Temp.: 101.12 Ambient Temp.: 86.38 F
00:17:14.853 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 86.45 F
00:17:16.416 -> Average Cal sensor Temp.: 101.14 Ambient Temp.: 86.52 F
00:17:17.975 -> Average Cal sensor Temp.: 100.96 Ambient Temp.: 86.49 F
00:17:19.536 -> Average Cal sensor Temp.: 101.02 Ambient Temp.: 86.52 F
00:17:21.128 -> Average Cal sensor Temp.: 101.16 Ambient Temp.: 86.56 F
00:17:22.681 -> Average Cal sensor Temp.: 100.85 Ambient Temp.: 86.59 F
00:17:24.242 -> Average Cal sensor Temp.: 100.94 Ambient Temp.: 86.59 F
00:17:25.798 -> Average Cal sensor Temp.: 100.89 Ambient Temp.: 86.63 F
00:17:27.388 -> Average Cal sensor Temp.: 100.97 Ambient Temp.: 86.63 F
00:17:28.947 -> Average Cal sensor Temp.: 101.06 Ambient Temp.: 86.59 F
00:17:30.503 -> Average Cal sensor Temp.: 101.10 Ambient Temp.: 86.63 F
00:17:32.056 -> Average Cal sensor Temp.: 100.92 Ambient Temp.: 86.67 F
00:17:33.650 -> Average Cal sensor Temp.: 100.82 Ambient Temp.: 86.70 F
00:17:35.211 -> Average Cal sensor Temp.: 100.98 Ambient Temp.: 86.67 F
00:17:36.762 -> Average Cal sensor Temp.: 100.96 Ambient Temp.: 86.67 F
00:17:38.324 -> Average Cal sensor Temp.: 100.99 Ambient Temp.: 86.74 F
00:17:39.919 -> Average Cal sensor Temp.: 100.91 Ambient Temp.: 86.70 F
00:17:41.481 -> Average Cal sensor Temp.: 101.12 Ambient Temp.: 86.74 F
00:17:43.043 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 86.77 F
00:17:44.603 -> Average Cal sensor Temp.: 101.18 Ambient Temp.: 86.81 F
00:17:46.164 -> Average Cal sensor Temp.: 101.12 Ambient Temp.: 86.77 F
00:17:47.722 -> Average Cal sensor Temp.: 101.15 Ambient Temp.: 86.88 F
00:17:49.286 -> Average Cal sensor Temp.: 101.21 Ambient Temp.: 86.92 F
00:17:50.846 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 86.95 F
00:17:52.405 -> Average Cal sensor Temp.: 101.39 Ambient Temp.: 86.95 F
00:17:53.966 -> Average Cal sensor Temp.: 101.08 Ambient Temp.: 86.92 F
00:17:55.527 -> Average Cal sensor Temp.: 101.06 Ambient Temp.: 86.92 F
00:17:57.122 -> Average Cal sensor Temp.: 101.01 Ambient Temp.: 86.95 F
00:17:58.681 -> Average Cal sensor Temp.: 100.96 Ambient Temp.: 86.92 F
00:18:00.238 -> Average Cal sensor Temp.: 101.09 Ambient Temp.: 86.99 F
00:18:01.795 -> Average Cal sensor Temp.: 101.27 Ambient Temp.: 87.03 F
00:18:03.383 -> Average Cal sensor Temp.: 101.20 Ambient Temp.: 87.06 F
00:18:04.944 -> Average Cal sensor Temp.: 101.25 Ambient Temp.: 87.10 F
00:18:06.503 -> Average Cal sensor Temp.: 101.12 Ambient Temp.: 87.06 F
00:18:08.066 -> Average Cal sensor Temp.: 101.01 Ambient Temp.: 87.06 F
00:18:09.628 -> Average Cal sensor Temp.: 100.81 Ambient Temp.: 87.10 F
00:18:11.188 -> Average Cal sensor Temp.: 100.98 Ambient Temp.: 87.13 F
00:18:12.746 -> Average Cal sensor Temp.: 100.91 Ambient Temp.: 87.13 F
00:18:14.303 -> Average Cal sensor Temp.: 100.94 Ambient Temp.: 87.17 F
00:18:15.893 -> Average Cal sensor Temp.: 101.14 Ambient Temp.: 87.21 F
00:18:17.447 -> Average Cal sensor Temp.: 101.04 Ambient Temp.: 87.17 F
00:18:19.005 -> Average Cal sensor Temp.: 101.17 Ambient Temp.: 87.17 F
00:18:20.562 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 87.21 F
00:18:22.115 -> Average Cal sensor Temp.: 101.08 Ambient Temp.: 87.21 F
00:18:23.704 -> Average Cal sensor Temp.: 101.25 Ambient Temp.: 87.17 F
00:18:25.256 -> Average Cal sensor Temp.: 101.29 Ambient Temp.: 87.28 F
00:18:26.815 -> Average Cal sensor Temp.: 101.25 Ambient Temp.: 87.21 F
00:18:28.372 -> Average Cal sensor Temp.: 101.02 Ambient Temp.: 87.28 F
00:18:29.930 -> Average Cal sensor Temp.: 101.03 Ambient Temp.: 87.31 F
00:18:31.518 -> Average Cal sensor Temp.: 101.10 Ambient Temp.: 87.28 F
00:18:33.074 -> Average Cal sensor Temp.: 101.17 Ambient Temp.: 87.31 F
00:18:34.638 -> Average Cal sensor Temp.: 101.21 Ambient Temp.: 87.35 F
00:18:36.200 -> Average Cal sensor Temp.: 101.28 Ambient Temp.: 87.31 F
00:18:37.761 -> Average Cal sensor Temp.: 101.18 Ambient Temp.: 87.35 F
00:18:39.320 -> Average Cal sensor Temp.: 101.41 Ambient Temp.: 87.35 F
00:18:40.872 -> Average Cal sensor Temp.: 101.48 Ambient Temp.: 87.35 F
00:18:42.430 -> Average Cal sensor Temp.: 101.47 Ambient Temp.: 87.35 F
00:18:43.990 -> Average Cal sensor Temp.: 101.56 Ambient Temp.: 87.35 F
00:18:45.555 -> Average Cal sensor Temp.: 101.33 Ambient Temp.: 87.35 F
00:18:47.116 -> Average Cal sensor Temp.: 100.85 Ambient Temp.: 87.35 F
00:18:48.710 -> Average Cal sensor Temp.: 100.66 Ambient Temp.: 87.35 F
00:18:50.265 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 87.39 F
00:18:51.825 -> Average Cal sensor Temp.: 100.71 Ambient Temp.: 87.31 F
00:18:53.384 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 87.31 F
00:18:54.947 -> Average Cal sensor Temp.: 101.27 Ambient Temp.: 87.35 F
00:18:56.502 -> Average Cal sensor Temp.: 101.08 Ambient Temp.: 87.39 F
00:18:58.063 -> Average Cal sensor Temp.: 101.20 Ambient Temp.: 87.35 F
00:18:59.626 -> Average Cal sensor Temp.: 101.10 Ambient Temp.: 87.39 F
00:19:01.184 -> Average Cal sensor Temp.: 101.43 Ambient Temp.: 87.39 F
00:19:02.777 -> Average Cal sensor Temp.: 101.05 Ambient Temp.: 87.39 F
00:19:04.334 -> Average Cal sensor Temp.: 101.47 Ambient Temp.: 87.42 F
00:19:05.895 -> Average Cal sensor Temp.: 100.97 Ambient Temp.: 87.42 F
00:19:07.451 -> Average Cal sensor Temp.: 100.92 Ambient Temp.: 87.35 F
00:19:09.038 -> Average Cal sensor Temp.: 101.06 Ambient Temp.: 87.35 F
00:19:10.599 -> Average Cal sensor Temp.: 100.79 Ambient Temp.: 87.31 F
00:19:12.155 -> Average Cal sensor Temp.: 101.71 Ambient Temp.: 87.39 F
00:19:13.712 -> Average Cal sensor Temp.: 101.27 Ambient Temp.: 87.35 F
00:19:15.268 -> Average Cal sensor Temp.: 101.36 Ambient Temp.: 87.39 F
00:19:16.823 -> Average Cal sensor Temp.: 100.90 Ambient Temp.: 87.31 F
00:19:18.386 -> Average Cal sensor Temp.: 101.13 Ambient Temp.: 87.35 F
00:19:19.950 -> Average Cal sensor Temp.: 101.27 Ambient Temp.: 87.31 F
00:19:21.512 -> Average Cal sensor Temp.: 101.76 Ambient Temp.: 87.39 F
00:19:23.102 -> Average Cal sensor Temp.: 101.34 Ambient Temp.: 87.35 F
00:19:24.656 -> Average Cal sensor Temp.: 100.93 Ambient Temp.: 87.35 F
00:19:26.208 -> Average Cal sensor Temp.: 101.17 Ambient Temp.: 87.31 F
00:19:27.766 -> Average Cal sensor Temp.: 101.15 Ambient Temp.: 87.35 F
00:19:29.324 -> Average Cal sensor Temp.: 101.49 Ambient Temp.: 87.31 F
00:19:30.913 -> Average Cal sensor Temp.: 101.92 Ambient Temp.: 87.35 F
00:19:32.473 -> Average Cal sensor Temp.: 101.20 Ambient Temp.: 87.31 F
00:19:34.033 -> Average Cal sensor Temp.: 101.27 Ambient Temp.: 87.31 F
00:19:35.595 -> Average Cal sensor Temp.: 101.28 Ambient Temp.: 87.31 F
00:19:37.148 -> Average Cal sensor Temp.: 101.34 Ambient Temp.: 87.28 F
00:19:38.707 -> Average Cal sensor Temp.: 101.62 Ambient Temp.: 87.31 F
00:19:40.266 -> Average Cal sensor Temp.: 101.50 Ambient Temp.: 87.31 F
00:19:41.854 -> Average Cal sensor Temp.: 101.42 Ambient Temp.: 87.31 F
00:19:43.411 -> Average Cal sensor Temp.: 101.32 Ambient Temp.: 87.31 F
00:19:44.970 -> Average Cal sensor Temp.: 101.23 Ambient Temp.: 87.28 F
00:19:46.533 -> Average Cal sensor Temp.: 101.47 Ambient Temp.: 87.21 F
00:19:48.090 -> Average Cal sensor Temp.: 100.97 Ambient Temp.: 87.21 F
00:19:49.655 -> Average Cal sensor Temp.: 100.97 Ambient Temp.: 87.31 F
00:19:51.217 -> Average Cal sensor Temp.: 101.12 Ambient Temp.: 87.13 F
00:19:52.769 -> Average Cal sensor Temp.: 100.94 Ambient Temp.: 87.13 F
00:19:54.330 -> Average Cal sensor Temp.: 101.27 Ambient Temp.: 87.17 F
00:19:55.924 -> Average Cal sensor Temp.: 100.80 Ambient Temp.: 87.10 F
00:19:57.484 -> Average Cal sensor Temp.: 101.16 Ambient Temp.: 87.13 F
00:19:59.039 -> Average Cal sensor Temp.: 101.05 Ambient Temp.: 87.06 F
00:20:00.604 -> Average Cal sensor Temp.: 101.10 Ambient Temp.: 87.10 F
00:20:02.163 -> Average Cal sensor Temp.: 101.38 Ambient Temp.: 87.06 F
00:20:03.724 -> Average Cal sensor Temp.: 101.39 Ambient Temp.: 87.03 F
00:20:05.279 -> Average Cal sensor Temp.: 101.04 Ambient Temp.: 87.06 F
00:20:06.839 -> Average Cal sensor Temp.: 101.23 Ambient Temp.: 87.03 F
00:20:08.428 -> Average Cal sensor Temp.: 101.03 Ambient Temp.: 87.03 F
00:20:09.986 -> Average Cal sensor Temp.: 101.43 Ambient Temp.: 86.99 F
00:20:11.545 -> Average Cal sensor Temp.: 100.85 Ambient Temp.: 86.92 F
00:20:13.101 -> Average Cal sensor Temp.: 101.16 Ambient Temp.: 86.99 F
00:20:14.659 -> Average Cal sensor Temp.: 101.16 Ambient Temp.: 86.88 F
00:20:16.251 -> Average Cal sensor Temp.: 101.16 Ambient Temp.: 86.88 F
00:20:17.809 -> Average Cal sensor Temp.: 101.44 Ambient Temp.: 86.81 F
00:20:19.366 -> Average Cal sensor Temp.: 101.33 Ambient Temp.: 86.81 F
00:20:20.920 -> Average Cal sensor Temp.: 101.35 Ambient Temp.: 86.81 F
00:20:22.506 -> Average Cal sensor Temp.: 101.36 Ambient Temp.: 86.81 F
00:20:24.064 -> Average Cal sensor Temp.: 101.68 Ambient Temp.: 86.81 F
00:20:25.612 -> Average Cal sensor Temp.: 101.25 Ambient Temp.: 86.77 F
00:20:27.202 -> Average Cal sensor Temp.: 101.15 Ambient Temp.: 86.77 F
00:20:28.760 -> Average Cal sensor Temp.: 101.47 Ambient Temp.: 86.74 F
00:20:30.315 -> Average Cal sensor Temp.: 101.49 Ambient Temp.: 86.74 F
00:20:31.872 -> Average Cal sensor Temp.: 101.68 Ambient Temp.: 86.74 F
00:20:33.430 -> Average Cal sensor Temp.: 101.56 Ambient Temp.: 86.70 F
00:20:34.991 -> Average Cal sensor Temp.: 101.59 Ambient Temp.: 86.63 F
00:20:36.585 -> Average Cal sensor Temp.: 101.56 Ambient Temp.: 86.67 F
00:20:38.141 -> Average Cal sensor Temp.: 101.19 Ambient Temp.: 86.63 F
00:20:39.698 -> Average Cal sensor Temp.: 101.21 Ambient Temp.: 86.59 F
00:20:41.249 -> Average Cal sensor Temp.: 101.50 Ambient Temp.: 86.59 F
00:20:42.839 -> Average Cal sensor Temp.: 101.14 Ambient Temp.: 86.59 F
00:20:44.399 -> Average Cal sensor Temp.: 100.48 Ambient Temp.: 86.52 F
00:20:45.951 -> Average Cal sensor Temp.: 100.98 Ambient Temp.: 86.49 F
00:20:47.539 -> Average Cal sensor Temp.: 101.33 Ambient Temp.: 86.45 F
00:20:49.100 -> Average Cal sensor Temp.: 101.05 Ambient Temp.: 86.49 F
00:20:50.664 -> Average Cal sensor Temp.: 101.56 Ambient Temp.: 86.38 F
00:20:52.223 -> Average Cal sensor Temp.: 101.39 Ambient Temp.: 86.45 F
00:20:53.781 -> Average Cal sensor Temp.: 101.00 Ambient Temp.: 86.31 F
00:20:55.341 -> Average Cal sensor Temp.: 101.14 Ambient Temp.: 86.31 F
00:20:56.899 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 86.31 F
00:20:58.459 -> Average Cal sensor Temp.: 101.42 Ambient Temp.: 86.27 F
00:21:00.022 -> Average Cal sensor Temp.: 101.46 Ambient Temp.: 86.27 F
00:21:01.586 -> Average Cal sensor Temp.: 101.35 Ambient Temp.: 86.27 F
00:21:03.144 -> Average Cal sensor Temp.: 101.17 Ambient Temp.: 86.20 F
00:21:04.705 -> Average Cal sensor Temp.: 101.31 Ambient Temp.: 86.16 F
00:21:06.296 -> Average Cal sensor Temp.: 101.18 Ambient Temp.: 86.13 F
00:21:07.858 -> Average Cal sensor Temp.: 101.40 Ambient Temp.: 86.16 F
00:21:09.421 -> Average Cal sensor Temp.: 101.21 Ambient Temp.: 86.05 F
00:21:10.982 -> Average Cal sensor Temp.: 101.18 Ambient Temp.: 86.09 F
00:21:12.540 -> Average Cal sensor Temp.: 101.18 Ambient Temp.: 86.05 F
00:21:14.104 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 86.05 F
00:21:15.666 -> Average Cal sensor Temp.: 100.82 Ambient Temp.: 85.98 F
00:21:17.233 -> Average Cal sensor Temp.: 100.95 Ambient Temp.: 85.95 F
00:21:18.791 -> Average Cal sensor Temp.: 101.29 Ambient Temp.: 85.91 F
00:21:20.349 -> Average Cal sensor Temp.: 101.36 Ambient Temp.: 85.87 F
00:21:21.937 -> Average Cal sensor Temp.: 101.23 Ambient Temp.: 85.91 F
00:21:23.496 -> Average Cal sensor Temp.: 101.36 Ambient Temp.: 85.87 F
00:21:25.053 -> Average Cal sensor Temp.: 101.19 Ambient Temp.: 85.91 F
00:21:26.613 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 85.84 F
00:21:28.174 -> Average Cal sensor Temp.: 101.17 Ambient Temp.: 85.84 F
00:21:29.734 -> Average Cal sensor Temp.: 101.21 Ambient Temp.: 85.73 F
00:21:31.294 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 85.73 F
00:21:32.855 -> Average Cal sensor Temp.: 101.62 Ambient Temp.: 85.73 F
00:21:34.414 -> Average Cal sensor Temp.: 101.20 Ambient Temp.: 85.80 F
00:21:36.008 -> Average Cal sensor Temp.: 100.97 Ambient Temp.: 85.69 F
00:21:37.564 -> Average Cal sensor Temp.: 101.51 Ambient Temp.: 85.69 F
00:21:39.126 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 85.73 F
00:21:40.684 -> Average Cal sensor Temp.: 101.18 Ambient Temp.: 85.66 F
00:21:42.240 -> Average Cal sensor Temp.: 101.41 Ambient Temp.: 85.66 F
00:21:43.796 -> Average Cal sensor Temp.: 100.91 Ambient Temp.: 85.66 F
00:21:45.391 -> Average Cal sensor Temp.: 101.34 Ambient Temp.: 85.55 F
00:21:46.947 -> Average Cal sensor Temp.: 101.12 Ambient Temp.: 85.51 F
00:21:48.502 -> Average Cal sensor Temp.: 101.23 Ambient Temp.: 85.51 F
00:21:50.060 -> Average Cal sensor Temp.: 101.82 Ambient Temp.: 85.51 F
00:21:51.622 -> Average Cal sensor Temp.: 101.27 Ambient Temp.: 85.44 F
00:21:53.177 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 85.51 F
00:21:54.768 -> Average Cal sensor Temp.: 101.22 Ambient Temp.: 85.48 F
00:21:56.329 -> Average Cal sensor Temp.: 101.35 Ambient Temp.: 85.44 F
00:21:57.888 -> Average Cal sensor Temp.: 101.08 Ambient Temp.: 85.41 F
00:21:59.446 -> Average Cal sensor Temp.: 101.71 Ambient Temp.: 85.44 F
00:22:01.009 -> Average Cal sensor Temp.: 101.60 Ambient Temp.: 85.44 F
00:22:02.561 -> Average Cal sensor Temp.: 101.88 Ambient Temp.: 85.44 F
00:22:04.122 -> Average Cal sensor Temp.: 101.67 Ambient Temp.: 85.44 F
00:22:05.716 -> Average Cal sensor Temp.: 101.62 Ambient Temp.: 85.44 F
00:22:07.278 -> Average Cal sensor Temp.: 101.88 Ambient Temp.: 85.48 F
00:22:08.834 -> Average Cal sensor Temp.: 101.47 Ambient Temp.: 85.44 F
00:22:10.390 -> Average Cal sensor Temp.: 101.41 Ambient Temp.: 85.44 F
00:22:11.948 -> Average Cal sensor Temp.: 101.25 Ambient Temp.: 85.48 F
00:22:13.505 -> Average Cal sensor Temp.: 101.41 Ambient Temp.: 85.44 F
00:22:15.086 -> Average Cal sensor Temp.: 101.48 Ambient Temp.: 85.51 F
00:22:16.642 -> Average Cal sensor Temp.: 101.36 Ambient Temp.: 85.44 F
00:22:18.195 -> Average Cal sensor Temp.: 101.51 Ambient Temp.: 85.48 F
00:22:19.785 -> Average Cal sensor Temp.: 101.26 Ambient Temp.: 85.44 F
00:22:21.341 -> Average Cal sensor Temp.: 101.62 Ambient Temp.: 85.44 F
00:22:22.897 -> Average Cal sensor Temp.: 101.45 Ambient Temp.: 85.48 F
00:22:24.455 -> Average Cal sensor Temp.: 101.46 Ambient Temp.: 85.44 F
00:22:26.015 -> Average Cal sensor Temp.: 101.15 Ambient Temp.: 85.51 F
00:22:27.601 -> Average Cal sensor Temp.: 101.27 Ambient Temp.: 85.51 F
00:22:29.153 -> Average Cal sensor Temp.: 101.39 Ambient Temp.: 85.51 F
00:22:30.710 -> Average Cal sensor Temp.: 101.52 Ambient Temp.: 85.55 F
00:22:32.266 -> Average Cal sensor Temp.: 101.64 Ambient Temp.: 85.55 F
00:22:33.826 -> Average Cal sensor Temp.: 101.65 Ambient Temp.: 85.59 F
00:22:35.388 -> Average Cal sensor Temp.: 101.77 Ambient Temp.: 85.59 F
00:22:36.976 -> Average Cal sensor Temp.: 101.65 Ambient Temp.: 85.66 F
00:22:38.523 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 85.59 F
00:22:40.080 -> Average Cal sensor Temp.: 101.34 Ambient Temp.: 85.59 F
00:22:41.641 -> Average Cal sensor Temp.: 101.31 Ambient Temp.: 85.66 F
00:22:43.235 -> Average Cal sensor Temp.: 101.39 Ambient Temp.: 85.66 F
00:22:44.793 -> Average Cal sensor Temp.: 101.48 Ambient Temp.: 85.59 F
00:22:46.354 -> Average Cal sensor Temp.: 101.48 Ambient Temp.: 85.66 F
00:22:47.913 -> Average Cal sensor Temp.: 101.56 Ambient Temp.: 85.66 F
00:22:49.471 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 85.69 F
00:22:51.028 -> Average Cal sensor Temp.: 101.43 Ambient Temp.: 85.73 F
00:22:52.589 -> Average Cal sensor Temp.: 101.60 Ambient Temp.: 85.73 F
00:22:54.153 -> Average Cal sensor Temp.: 101.80 Ambient Temp.: 85.77 F
00:22:55.715 -> Average Cal sensor Temp.: 101.49 Ambient Temp.: 85.77 F
00:22:57.306 -> Average Cal sensor Temp.: 101.43 Ambient Temp.: 85.77 F
00:22:58.866 -> Average Cal sensor Temp.: 101.64 Ambient Temp.: 85.77 F
00:23:00.432 -> Average Cal sensor Temp.: 101.97 Ambient Temp.: 85.80 F
00:23:01.995 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 85.77 F
00:23:03.554 -> Average Cal sensor Temp.: 101.27 Ambient Temp.: 85.80 F
00:23:05.114 -> Average Cal sensor Temp.: 101.62 Ambient Temp.: 85.87 F
00:23:06.671 -> Average Cal sensor Temp.: 101.80 Ambient Temp.: 85.91 F
00:23:08.225 -> Average Cal sensor Temp.: 101.74 Ambient Temp.: 85.87 F
00:23:09.789 -> Average Cal sensor Temp.: 101.63 Ambient Temp.: 85.95 F
00:23:11.349 -> Average Cal sensor Temp.: 101.63 Ambient Temp.: 85.95 F
00:23:12.943 -> Average Cal sensor Temp.: 101.88 Ambient Temp.: 85.98 F
00:23:14.505 -> Average Cal sensor Temp.: 101.75 Ambient Temp.: 85.95 F
00:23:16.053 -> Average Cal sensor Temp.: 101.62 Ambient Temp.: 85.95 F
00:23:17.613 -> Average Cal sensor Temp.: 101.73 Ambient Temp.: 86.05 F
00:23:19.178 -> Average Cal sensor Temp.: 101.39 Ambient Temp.: 85.95 F
00:23:20.736 -> Average Cal sensor Temp.: 101.46 Ambient Temp.: 86.05 F
00:23:22.314 -> Average Cal sensor Temp.: 101.59 Ambient Temp.: 86.05 F
00:23:23.863 -> Average Cal sensor Temp.: 101.89 Ambient Temp.: 86.09 F
00:23:25.457 -> Average Cal sensor Temp.: 101.87 Ambient Temp.: 86.09 F
00:23:27.014 -> Average Cal sensor Temp.: 101.62 Ambient Temp.: 86.09 F
00:23:28.576 -> Average Cal sensor Temp.: 101.70 Ambient Temp.: 86.09 F
00:23:30.137 -> Average Cal sensor Temp.: 101.84 Ambient Temp.: 86.13 F
00:23:31.701 -> Average Cal sensor Temp.: 101.69 Ambient Temp.: 86.13 F
00:23:33.267 -> Average Cal sensor Temp.: 101.56 Ambient Temp.: 86.16 F
00:23:34.831 -> Average Cal sensor Temp.: 101.68 Ambient Temp.: 86.16 F
00:23:36.388 -> Average Cal sensor Temp.: 101.61 Ambient Temp.: 86.20 F
00:23:37.946 -> Average Cal sensor Temp.: 101.96 Ambient Temp.: 86.20 F
00:23:39.510 -> Average Cal sensor Temp.: 101.90 Ambient Temp.: 86.23 F
00:23:41.101 -> Average Cal sensor Temp.: 101.80 Ambient Temp.: 86.23 F
00:23:42.661 -> Average Cal sensor Temp.: 101.72 Ambient Temp.: 86.27 F
00:23:44.225 -> Average Cal sensor Temp.: 101.64 Ambient Temp.: 86.23 F
00:23:45.781 -> Average Cal sensor Temp.: 101.88 Ambient Temp.: 86.31 F
00:23:47.343 -> Average Cal sensor Temp.: 102.13 Ambient Temp.: 86.38 F
00:23:48.901 -> Average Cal sensor Temp.: 102.04 Ambient Temp.: 86.34 F
00:23:50.461 -> Average Cal sensor Temp.: 101.91 Ambient Temp.: 86.34 F
00:23:52.047 -> Average Cal sensor Temp.: 101.71 Ambient Temp.: 86.34 F
00:23:53.611 -> Average Cal sensor Temp.: 101.97 Ambient Temp.: 86.34 F
00:23:55.174 -> Average Cal sensor Temp.: 101.82 Ambient Temp.: 86.38 F
00:23:56.736 -> Average Cal sensor Temp.: 101.68 Ambient Temp.: 86.38 F
00:23:58.296 -> Average Cal sensor Temp.: 101.92 Ambient Temp.: 86.45 F
00:23:59.858 -> Average Cal sensor Temp.: 102.05 Ambient Temp.: 86.45 F
00:24:01.415 -> Average Cal sensor Temp.: 102.06 Ambient Temp.: 86.52 F
00:24:03.002 -> Average Cal sensor Temp.: 101.79 Ambient Temp.: 86.52 F
00:24:04.565 -> Average Cal sensor Temp.: 101.69 Ambient Temp.: 86.52 F
00:24:06.116 -> Average Cal sensor Temp.: 101.76 Ambient Temp.: 86.52 F
00:24:07.681 -> Average Cal sensor Temp.: 101.54 Ambient Temp.: 86.56 F
00:24:09.241 -> Average Cal sensor Temp.: 101.88 Ambient Temp.: 86.56 F
00:24:10.799 -> Average Cal sensor Temp.: 101.65 Ambient Temp.: 86.52 F
00:24:12.355 -> Average Cal sensor Temp.: 101.84 Ambient Temp.: 86.56 F
00:24:13.948 -> Average Cal sensor Temp.: 101.69 Ambient Temp.: 86.59 F
00:24:15.503 -> Average Cal sensor Temp.: 101.84 Ambient Temp.: 86.63 F
00:24:17.056 -> Average Cal sensor Temp.: 101.73 Ambient Temp.: 86.63 F
00:24:18.616 -> Average Cal sensor Temp.: 101.83 Ambient Temp.: 86.63 F
00:24:20.211 -> Average Cal sensor Temp.: 101.85 Ambient Temp.: 86.70 F
00:24:21.773 -> Average Cal sensor Temp.: 101.95 Ambient Temp.: 86.67 F
00:24:23.331 -> Average Cal sensor Temp.: 101.79 Ambient Temp.: 86.70 F
00:24:24.891 -> Average Cal sensor Temp.: 101.85 Ambient Temp.: 86.67 F
00:24:26.445 -> Average Cal sensor Temp.: 101.93 Ambient Temp.: 86.74 F
00:24:28.000 -> Average Cal sensor Temp.: 101.97 Ambient Temp.: 86.77 F
00:24:29.595 -> Average Cal sensor Temp.: 101.85 Ambient Temp.: 86.70 F
00:24:31.155 -> Average Cal sensor Temp.: 101.98 Ambient Temp.: 86.81 F
00:24:32.711 -> Average Cal sensor Temp.: 101.79 Ambient Temp.: 86.77 F
00:24:34.275 -> Average Cal sensor Temp.: 101.94 Ambient Temp.: 86.81 F
00:24:35.830 -> Average Cal sensor Temp.: 101.89 Ambient Temp.: 86.81 F
00:24:37.387 -> Average Cal sensor Temp.: 101.96 Ambient Temp.: 86.88 F
00:24:38.976 -> Average Cal sensor Temp.: 101.95 Ambient Temp.: 86.88 F
00:24:40.533 -> Average Cal sensor Temp.: 101.72 Ambient Temp.: 86.88 F
00:24:42.093 -> Average Cal sensor Temp.: 101.77 Ambient Temp.: 86.92 F
00:24:43.650 -> Average Cal sensor Temp.: 101.74 Ambient Temp.: 86.95 F
00:24:45.210 -> Average Cal sensor Temp.: 101.65 Ambient Temp.: 86.99 F
00:24:46.799 -> Average Cal sensor Temp.: 102.11 Ambient Temp.: 86.99 F
00:24:48.364 -> Average Cal sensor Temp.: 102.14 Ambient Temp.: 87.06 F
00:24:49.922 -> Average Cal sensor Temp.: 101.87 Ambient Temp.: 87.03 F
00:24:51.482 -> Average Cal sensor Temp.: 102.15 Ambient Temp.: 87.06 F
00:24:53.039 -> Average Cal sensor Temp.: 101.85 Ambient Temp.: 87.10 F
00:24:54.597 -> Average Cal sensor Temp.: 102.02 Ambient Temp.: 87.10 F
00:24:56.156 -> Average Cal sensor Temp.: 101.73 Ambient Temp.: 87.06 F
00:24:57.750 -> Average Cal sensor Temp.: 101.82 Ambient Temp.: 87.13 F
00:24:59.311 -> Average Cal sensor Temp.: 101.91 Ambient Temp.: 87.06 F
00:25:00.874 -> Average Cal sensor Temp.: 101.84 Ambient Temp.: 87.06 F
00:25:02.433 -> Average Cal sensor Temp.: 102.13 Ambient Temp.: 87.13 F
00:25:03.991 -> Average Cal sensor Temp.: 102.16 Ambient Temp.: 87.10 F
00:25:05.555 -> Average Cal sensor Temp.: 101.99 Ambient Temp.: 87.17 F
00:25:07.115 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 87.17 F
00:25:08.672 -> Average Cal sensor Temp.: 102.15 Ambient Temp.: 87.17 F
00:25:10.229 -> Average Cal sensor Temp.: 101.92 Ambient Temp.: 87.17 F
00:25:11.783 -> Average Cal sensor Temp.: 101.77 Ambient Temp.: 87.17 F
00:25:13.371 -> Average Cal sensor Temp.: 101.96 Ambient Temp.: 87.17 F
00:25:14.931 -> Average Cal sensor Temp.: 102.15 Ambient Temp.: 87.21 F
00:25:16.491 -> Average Cal sensor Temp.: 102.25 Ambient Temp.: 87.28 F
00:25:18.049 -> Average Cal sensor Temp.: 102.16 Ambient Temp.: 87.31 F
00:25:19.609 -> Average Cal sensor Temp.: 102.13 Ambient Temp.: 87.31 F
00:25:21.168 -> Average Cal sensor Temp.: 102.18 Ambient Temp.: 87.35 F
00:25:22.751 -> Average Cal sensor Temp.: 102.08 Ambient Temp.: 87.39 F
00:25:24.310 -> Average Cal sensor Temp.: 102.07 Ambient Temp.: 87.35 F
00:25:25.874 -> Average Cal sensor Temp.: 101.95 Ambient Temp.: 87.31 F
00:25:27.436 -> Average Cal sensor Temp.: 102.08 Ambient Temp.: 87.39 F
00:25:28.989 -> Average Cal sensor Temp.: 102.08 Ambient Temp.: 87.39 F
00:25:30.551 -> Average Cal sensor Temp.: 101.89 Ambient Temp.: 87.35 F
00:25:32.112 -> Average Cal sensor Temp.: 101.88 Ambient Temp.: 87.31 F
00:25:33.701 -> Average Cal sensor Temp.: 101.93 Ambient Temp.: 87.42 F
00:25:35.259 -> Average Cal sensor Temp.: 102.21 Ambient Temp.: 87.35 F
00:25:36.813 -> Average Cal sensor Temp.: 102.10 Ambient Temp.: 87.42 F
00:25:38.366 -> Average Cal sensor Temp.: 102.19 Ambient Temp.: 87.42 F
00:25:39.955 -> Average Cal sensor Temp.: 102.28 Ambient Temp.: 87.39 F
00:25:41.511 -> Average Cal sensor Temp.: 102.21 Ambient Temp.: 87.39 F
00:25:43.064 -> Average Cal sensor Temp.: 101.97 Ambient Temp.: 87.42 F
00:25:44.625 -> Average Cal sensor Temp.: 101.97 Ambient Temp.: 87.35 F
00:25:46.182 -> Average Cal sensor Temp.: 102.10 Ambient Temp.: 87.35 F
00:25:47.767 -> Average Cal sensor Temp.: 102.06 Ambient Temp.: 87.35 F
00:25:49.319 -> Average Cal sensor Temp.: 102.32 Ambient Temp.: 87.42 F
00:25:50.874 -> Average Cal sensor Temp.: 102.23 Ambient Temp.: 87.46 F
00:25:52.431 -> Average Cal sensor Temp.: 102.16 Ambient Temp.: 87.39 F
00:25:54.024 -> Average Cal sensor Temp.: 102.02 Ambient Temp.: 87.39 F
00:25:55.584 -> Average Cal sensor Temp.: 102.13 Ambient Temp.: 87.35 F
00:25:57.141 -> Average Cal sensor Temp.: 102.02 Ambient Temp.: 87.39 F
00:25:58.697 -> Average Cal sensor Temp.: 102.46 Ambient Temp.: 87.39 F
00:26:00.254 -> Average Cal sensor Temp.: 102.15 Ambient Temp.: 87.39 F
00:26:01.814 -> Average Cal sensor Temp.: 102.28 Ambient Temp.: 87.39 F
00:26:03.410 -> Average Cal sensor Temp.: 102.20 Ambient Temp.: 87.39 F
00:26:04.976 -> Average Cal sensor Temp.: 101.87 Ambient Temp.: 87.39 F
00:26:06.538 -> Average Cal sensor Temp.: 101.74 Ambient Temp.: 87.31 F
00:26:08.102 -> Average Cal sensor Temp.: 101.89 Ambient Temp.: 87.31 F
00:26:09.633 -> Average Cal sensor Temp.: 101.98 Ambient Temp.: 87.31 F
00:26:11.198 -> Average Cal sensor Temp.: 101.79 Ambient Temp.: 87.31 F
00:26:12.763 -> Average Cal sensor Temp.: 101.99 Ambient Temp.: 87.28 F
00:26:14.356 -> Average Cal sensor Temp.: 101.86 Ambient Temp.: 87.21 F
00:26:15.886 -> Average Cal sensor Temp.: 102.26 Ambient Temp.: 87.28 F
00:26:17.481 -> Average Cal sensor Temp.: 101.66 Ambient Temp.: 87.28 F
00:26:19.046 -> Average Cal sensor Temp.: 101.63 Ambient Temp.: 87.21 F
00:26:20.607 -> Average Cal sensor Temp.: 102.03 Ambient Temp.: 87.21 F
00:26:22.165 -> Average Cal sensor Temp.: 101.82 Ambient Temp.: 87.13 F
00:26:23.724 -> Average Cal sensor Temp.: 101.83 Ambient Temp.: 87.13 F
00:26:25.286 -> Average Cal sensor Temp.: 101.85 Ambient Temp.: 87.13 F
00:26:26.843 -> Average Cal sensor Temp.: 101.89 Ambient Temp.: 87.13 F
00:26:28.405 -> Average Cal sensor Temp.: 102.00 Ambient Temp.: 87.10 F
00:26:29.960 -> Average Cal sensor Temp.: 101.94 Ambient Temp.: 87.06 F
00:26:31.549 -> Average Cal sensor Temp.: 102.08 Ambient Temp.: 87.03 F
00:26:33.113 -> Average Cal sensor Temp.: 102.16 Ambient Temp.: 87.03 F
00:26:34.679 -> Average Cal sensor Temp.: 102.28 Ambient Temp.: 87.06 F
00:26:36.239 -> Average Cal sensor Temp.: 102.02 Ambient Temp.: 87.03 F
00:26:37.789 -> Average Cal sensor Temp.: 101.93 Ambient Temp.: 86.88 F
00:26:39.347 -> Average Cal sensor Temp.: 102.32 Ambient Temp.: 86.99 F
00:26:40.937 -> Average Cal sensor Temp.: 102.17 Ambient Temp.: 86.95 F
00:26:42.498 -> Average Cal sensor Temp.: 102.44 Ambient Temp.: 87.03 F
00:26:44.050 -> Average Cal sensor Temp.: 101.97 Ambient Temp.: 86.99 F
00:26:45.604 -> Average Cal sensor Temp.: 102.20 Ambient Temp.: 86.92 F
00:26:47.195 -> Average Cal sensor Temp.: 102.00 Ambient Temp.: 86.92 F
00:26:48.755 -> Average Cal sensor Temp.: 102.02 Ambient Temp.: 86.88 F
00:26:50.313 -> Average Cal sensor Temp.: 102.40 Ambient Temp.: 86.92 F
00:26:51.867 -> Average Cal sensor Temp.: 102.19 Ambient Temp.: 86.95 F
00:26:53.458 -> Average Cal sensor Temp.: 102.12 Ambient Temp.: 86.81 F
00:26:55.016 -> Average Cal sensor Temp.: 101.97 Ambient Temp.: 86.88 F
00:26:56.575 -> Average Cal sensor Temp.: 101.81 Ambient Temp.: 86.77 F
00:26:58.132 -> Average Cal sensor Temp.: 102.27 Ambient Temp.: 86.81 F
00:26:59.695 -> Average Cal sensor Temp.: 102.22 Ambient Temp.: 86.77 F
00:27:01.248 -> Average Cal sensor Temp.: 102.21 Ambient Temp.: 86.77 F
00:27:02.829 -> Average Cal sensor Temp.: 102.44 Ambient Temp.: 86.81 F
00:27:04.382 -> Average Cal sensor Temp.: 102.10 Ambient Temp.: 86.74 F
00:27:05.941 -> Average Cal sensor Temp.: 101.90 Ambient Temp.: 86.70 F
00:27:07.536 -> Average Cal sensor Temp.: 101.93 Ambient Temp.: 86.74 F
00:27:09.097 -> Average Cal sensor Temp.: 102.04 Ambient Temp.: 86.74 F
00:27:10.656 -> Average Cal sensor Temp.: 101.78 Ambient Temp.: 86.70 F
00:27:12.212 -> Average Cal sensor Temp.: 101.79 Ambient Temp.: 86.67 F
00:27:13.765 -> Average Cal sensor Temp.: 101.76 Ambient Temp.: 86.59 F
00:27:15.359 -> Average Cal sensor Temp.: 102.17 Ambient Temp.: 86.59 F
00:27:16.914 -> Average Cal sensor Temp.: 101.93 Ambient Temp.: 86.59 F
00:27:18.469 -> Average Cal sensor Temp.: 102.11 Ambient Temp.: 86.56 F
00:27:20.020 -> Average Cal sensor Temp.: 102.25 Ambient Temp.: 86.59 F
00:27:21.610 -> Average Cal sensor Temp.: 102.26 Ambient Temp.: 86.56 F
00:27:23.167 -> Average Cal sensor Temp.: 102.38 Ambient Temp.: 86.56 F
00:27:24.724 -> Average Cal sensor Temp.: 101.83 Ambient Temp.: 86.52 F
00:27:26.283 -> Average Cal sensor Temp.: 102.11 Ambient Temp.: 86.49 F
00:27:27.844 -> Average Cal sensor Temp.: 102.14 Ambient Temp.: 86.59 F
00:27:29.404 -> Average Cal sensor Temp.: 102.17 Ambient Temp.: 86.49 F
00:27:31.001 -> Average Cal sensor Temp.: 102.38 Ambient Temp.: 86.38 F
00:27:32.549 -> Average Cal sensor Temp.: 102.57 Ambient Temp.: 86.45 F
00:27:34.108 -> Average Cal sensor Temp.: 102.51 Ambient Temp.: 86.49 F
00:27:35.665 -> Average Cal sensor Temp.: 102.47 Ambient Temp.: 86.49 F
00:27:37.226 -> Average Cal sensor Temp.: 102.47 Ambient Temp.: 86.45 F
00:27:38.790 -> Average Cal sensor Temp.: 102.33 Ambient Temp.: 86.52 F
00:27:40.385 -> Average Cal sensor Temp.: 101.82 Ambient Temp.: 86.38 F
00:27:41.945 -> Average Cal sensor Temp.: 102.21 Ambient Temp.: 86.45 F
00:27:43.507 -> Average Cal sensor Temp.: 102.61 Ambient Temp.: 86.38 F
00:27:45.068 -> Average Cal sensor Temp.: 102.20 Ambient Temp.: 86.38 F
00:27:46.628 -> Average Cal sensor Temp.: 102.27 Ambient Temp.: 86.49 F
00:27:48.181 -> Average Cal sensor Temp.: 102.51 Ambient Temp.: 86.38 F
00:27:49.742 -> Average Cal sensor Temp.: 102.66 Ambient Temp.: 86.38 F
00:27:51.336 -> Average Cal sensor Temp.: 102.52 Ambient Temp.: 86.31 F
00:27:52.897 -> Average Cal sensor Temp.: 102.27 Ambient Temp.: 86.38 F
00:27:54.456 -> Average Cal sensor Temp.: 102.28 Ambient Temp.: 86.38 F
00:27:56.010 -> Average Cal sensor Temp.: 101.93 Ambient Temp.: 86.27 F
00:27:57.563 -> Average Cal sensor Temp.: 102.12 Ambient Temp.: 86.34 F
00:27:59.157 -> Average Cal sensor Temp.: 102.38 Ambient Temp.: 86.31 F
00:28:00.707 -> Average Cal sensor Temp.: 102.47 Ambient Temp.: 86.34 F
00:28:02.263 -> Average Cal sensor Temp.: 102.02 Ambient Temp.: 86.34 F
00:28:03.823 -> Average Cal sensor Temp.: 102.41 Ambient Temp.: 86.34 F
00:28:05.410 -> Average Cal sensor Temp.: 102.80 Ambient Temp.: 86.45 F
00:28:06.969 -> Average Cal sensor Temp.: 102.40 Ambient Temp.: 86.34 F
00:28:08.524 -> Average Cal sensor Temp.: 102.37 Ambient Temp.: 86.34 F
00:28:10.086 -> Average Cal sensor Temp.: 102.58 Ambient Temp.: 86.34 F
00:28:11.674 -> Average Cal sensor Temp.: 102.61 Ambient Temp.: 86.45 F
00:28:13.236 -> Average Cal sensor Temp.: 102.76 Ambient Temp.: 86.34 F
00:28:14.800 -> Average Cal sensor Temp.: 102.44 Ambient Temp.: 86.38 F
00:28:16.361 -> Average Cal sensor Temp.: 102.12 Ambient Temp.: 86.34 F
00:28:17.921 -> Average Cal sensor Temp.: 102.13 Ambient Temp.: 86.38 F
00:28:19.478 -> Average Cal sensor Temp.: 102.51 Ambient Temp.: 86.31 F
00:28:21.037 -> Average Cal sensor Temp.: 102.68 Ambient Temp.: 86.34 F
00:28:22.601 -> Average Cal sensor Temp.: 102.86 Ambient Temp.: 86.45 F
00:28:24.163 -> Average Cal sensor Temp.: 102.55 Ambient Temp.: 86.45 F
00:28:25.721 -> Average Cal sensor Temp.: 102.63 Ambient Temp.: 86.38 F
00:28:27.306 -> Average Cal sensor Temp.: 102.44 Ambient Temp.: 86.34 F
00:28:28.867 -> Average Cal sensor Temp.: 102.40 Ambient Temp.: 86.38 F
00:28:30.427 -> Average Cal sensor Temp.: 102.63 Ambient Temp.: 86.34 F
00:28:31.983 -> Average Cal sensor Temp.: 102.64 Ambient Temp.: 86.45 F
00:28:33.547 -> Average Cal sensor Temp.: 102.62 Ambient Temp.: 86.45 F
00:28:35.146 -> Average Cal sensor Temp.: 102.56 Ambient Temp.: 86.38 F
00:28:36.708 -> Average Cal sensor Temp.: 102.66 Ambient Temp.: 86.45 F
00:28:38.269 -> Average Cal sensor Temp.: 102.61 Ambient Temp.: 86.49 F
00:28:39.829 -> Average Cal sensor Temp.: 102.62 Ambient Temp.: 86.49 F
00:28:41.385 -> Average Cal sensor Temp.: 102.87 Ambient Temp.: 86.52 F
00:28:42.940 -> Average Cal sensor Temp.: 102.65 Ambient Temp.: 86.49 F
00:28:44.533 -> Average Cal sensor Temp.: 102.54 Ambient Temp.: 86.52 F
00:28:46.095 -> Average Cal sensor Temp.: 102.41 Ambient Temp.: 86.63 F
00:28:47.657 -> Average Cal sensor Temp.: 102.50 Ambient Temp.: 86.52 F
00:28:49.212 -> Average Cal sensor Temp.: 102.67 Ambient Temp.: 86.56 F
00:28:50.772 -> Average Cal sensor Temp.: 102.66 Ambient Temp.: 86.49 F
00:28:52.328 -> Average Cal sensor Temp.: 102.48 Ambient Temp.: 86.56 F
00:28:53.921 -> Average Cal sensor Temp.: 102.61 Ambient Temp.: 86.52 F
00:28:55.480 -> Average Cal sensor Temp.: 102.53 Ambient Temp.: 86.59 F
00:28:57.041 -> Average Cal sensor Temp.: 102.29 Ambient Temp.: 86.56 F
00:28:58.598 -> Average Cal sensor Temp.: 102.70 Ambient Temp.: 86.63 F
00:29:00.159 -> Average Cal sensor Temp.: 102.66 Ambient Temp.: 86.59 F
00:29:01.716 -> Average Cal sensor Temp.: 102.77 Ambient Temp.: 86.59 F
00:29:03.308 -> Average Cal sensor Temp.: 102.81 Ambient Temp.: 86.63 F
00:29:04.871 -> Average Cal sensor Temp.: 102.86 Ambient Temp.: 86.63 F
00:29:06.426 -> Average Cal sensor Temp.: 102.41 Ambient Temp.: 86.67 F
00:29:07.989 -> Average Cal sensor Temp.: 102.82 Ambient Temp.: 86.63 F
00:29:09.551 -> Average Cal sensor Temp.: 102.64 Ambient Temp.: 86.70 F
00:29:11.110 -> Average Cal sensor Temp.: 102.74 Ambient Temp.: 86.63 F
00:29:12.700 -> Average Cal sensor Temp.: 102.78 Ambient Temp.: 86.67 F
00:29:14.265 -> Average Cal sensor Temp.: 102.63 Ambient Temp.: 86.74 F
00:29:15.828 -> Average Cal sensor Temp.: 102.88 Ambient Temp.: 86.70 F
00:29:17.382 -> Average Cal sensor Temp.: 102.75 Ambient Temp.: 86.70 F
00:29:18.938 -> Average Cal sensor Temp.: 102.46 Ambient Temp.: 86.70 F
00:29:20.502 -> Average Cal sensor Temp.: 102.99 Ambient Temp.: 86.74 F
00:29:22.066 -> Average Cal sensor Temp.: 102.84 Ambient Temp.: 86.77 F
00:29:23.626 -> Average Cal sensor Temp.: 102.66 Ambient Temp.: 86.74 F
00:29:25.223 -> Average Cal sensor Temp.: 102.72 Ambient Temp.: 86.77 F
00:29:26.778 -> Average Cal sensor Temp.: 102.47 Ambient Temp.: 86.77 F
00:29:28.336 -> Average Cal sensor Temp.: 102.35 Ambient Temp.: 86.77 F
00:29:29.891 -> Average Cal sensor Temp.: 102.51 Ambient Temp.: 86.81 F
00:29:31.450 -> Average Cal sensor Temp.: 102.73 Ambient Temp.: 86.81 F
00:29:33.007 -> Average Cal sensor Temp.: 102.93 Ambient Temp.: 86.81 F
00:29:34.598 -> Average Cal sensor Temp.: 103.06 Ambient Temp.: 86.88 F
00:29:36.159 -> Average Cal sensor Temp.: 102.59 Ambient Temp.: 86.81 F
00:29:37.724 -> Average Cal sensor Temp.: 102.69 Ambient Temp.: 86.88 F
00:29:39.292 -> Average Cal sensor Temp.: 102.83 Ambient Temp.: 86.88 F
00:29:40.851 -> Average Cal sensor Temp.: 102.59 Ambient Temp.: 86.88 F
00:29:42.412 -> Average Cal sensor Temp.: 102.72 Ambient Temp.: 86.92 F
00:29:43.975 -> Average Cal sensor Temp.: 102.74 Ambient Temp.: 86.92 F
00:29:45.535 -> Average Cal sensor Temp.: 102.44 Ambient Temp.: 86.92 F
00:29:47.096 -> Average Cal sensor Temp.: 102.80 Ambient Temp.: 86.92 F
00:29:48.690 -> Average Cal sensor Temp.: 102.64 Ambient Temp.: 86.95 F
00:29:50.252 -> Average Cal sensor Temp.: 102.88 Ambient Temp.: 86.95 F
00:29:51.814 -> Average Cal sensor Temp.: 102.57 Ambient Temp.: 86.99 F
00:29:53.373 -> Average Cal sensor Temp.: 102.66 Ambient Temp.: 86.95 F
00:29:54.933 -> Average Cal sensor Temp.: 102.66 Ambient Temp.: 86.99 F
00:29:56.520 -> Average Cal sensor Temp.: 102.81 Ambient Temp.: 87.03 F
00:29:58.084 -> Average Cal sensor Temp.: 102.79 Ambient Temp.: 87.06 F
00:29:59.643 -> Average Cal sensor Temp.: 103.07 Ambient Temp.: 87.06 F
00:30:01.203 -> Average Cal sensor Temp.: 103.18 Ambient Temp.: 87.10 F
00:30:02.754 -> Average Cal sensor Temp.: 103.23 Ambient Temp.: 87.10 F
00:30:04.312 -> Average Cal sensor Temp.: 103.18 Ambient Temp.: 87.13 F
00:30:05.905 -> Average Cal sensor Temp.: 102.75 Ambient Temp.: 87.13 F
00:30:07.462 -> Average Cal sensor Temp.: 102.90 Ambient Temp.: 87.17 F
00:30:09.021 -> Average Cal sensor Temp.: 103.00 Ambient Temp.: 87.17 F
00:30:10.577 -> Average Cal sensor Temp.: 102.92 Ambient Temp.: 87.13 F
00:30:12.136 -> Average Cal sensor Temp.: 102.80 Ambient Temp.: 87.28 F
00:30:13.694 -> Average Cal sensor Temp.: 102.68 Ambient Temp.: 87.21 F
00:30:15.288 -> Average Cal sensor Temp.: 102.68 Ambient Temp.: 87.28 F
00:30:16.851 -> Average Cal sensor Temp.: 102.95 Ambient Temp.: 87.28 F
00:30:18.410 -> Average Cal sensor Temp.: 102.78 Ambient Temp.: 87.31 F
00:30:19.967 -> Average Cal sensor Temp.: 102.87 Ambient Temp.: 87.31 F
00:30:21.525 -> Average Cal sensor Temp.: 102.77 Ambient Temp.: 87.31 F
00:30:23.082 -> Average Cal sensor Temp.: 102.83 Ambient Temp.: 87.31 F
00:30:24.671 -> Average Cal sensor Temp.: 102.96 Ambient Temp.: 87.42 F
00:30:26.229 -> Average Cal sensor Temp.: 102.98 Ambient Temp.: 87.39 F
00:30:27.779 -> Average Cal sensor Temp.: 102.98 Ambient Temp.: 87.42 F
00:30:29.336 -> Average Cal sensor Temp.: 102.82 Ambient Temp.: 87.42 F
00:30:30.930 -> Average Cal sensor Temp.: 102.81 Ambient Temp.: 87.42 F
00:30:32.485 -> Average Cal sensor Temp.: 102.67 Ambient Temp.: 87.42 F
00:30:34.041 -> Average Cal sensor Temp.: 102.82 Ambient Temp.: 87.42 F
00:30:35.602 -> Average Cal sensor Temp.: 103.05 Ambient Temp.: 87.49 F
00:30:37.164 -> Average Cal sensor Temp.: 103.09 Ambient Temp.: 87.49 F
00:30:38.721 -> Average Cal sensor Temp.: 102.96 Ambient Temp.: 87.53 F
00:30:40.314 -> Average Cal sensor Temp.: 102.93 Ambient Temp.: 87.57 F
00:30:41.870 -> Average Cal sensor Temp.: 103.08 Ambient Temp.: 87.60 F
00:30:43.428 -> Average Cal sensor Temp.: 103.21 Ambient Temp.: 87.60 F
00:30:44.987 -> Average Cal sensor Temp.: 103.07 Ambient Temp.: 87.67 F
00:30:46.545 -> Average Cal sensor Temp.: 102.86 Ambient Temp.: 87.67 F
00:30:48.137 -> Average Cal sensor Temp.: 102.63 Ambient Temp.: 87.71 F
00:30:49.699 -> Average Cal sensor Temp.: 102.78 Ambient Temp.: 87.71 F
00:30:51.247 -> Average Cal sensor Temp.: 102.74 Ambient Temp.: 87.67 F
00:30:52.799 -> Average Cal sensor Temp.: 102.93 Ambient Temp.: 87.71 F
00:30:54.391 -> Average Cal sensor Temp.: 102.75 Ambient Temp.: 87.75 F
00:30:55.954 -> Average Cal sensor Temp.: 102.56 Ambient Temp.: 87.67 F
00:30:57.501 -> Average Cal sensor Temp.: 102.75 Ambient Temp.: 87.71 F
00:30:59.059 -> Average Cal sensor Temp.: 102.73 Ambient Temp.: 87.78 F
00:31:00.643 -> Average Cal sensor Temp.: 102.95 Ambient Temp.: 87.75 F
00:31:02.197 -> Average Cal sensor Temp.: 102.96 Ambient Temp.: 87.78 F
00:31:03.752 -> Average Cal sensor Temp.: 103.19 Ambient Temp.: 87.75 F
00:31:05.311 -> Average Cal sensor Temp.: 102.72 Ambient Temp.: 87.82 F
00:31:06.899 -> Average Cal sensor Temp.: 102.39 Ambient Temp.: 87.75 F
00:31:08.458 -> Average Cal sensor Temp.: 102.99 Ambient Temp.: 87.78 F
00:31:10.016 -> Average Cal sensor Temp.: 103.46 Ambient Temp.: 87.75 F
00:31:11.571 -> Average Cal sensor Temp.: 103.44 Ambient Temp.: 87.82 F
00:31:13.127 -> Average Cal sensor Temp.: 103.30 Ambient Temp.: 87.78 F
00:31:14.691 -> Average Cal sensor Temp.: 103.34 Ambient Temp.: 87.82 F
00:31:16.278 -> Average Cal sensor Temp.: 103.09 Ambient Temp.: 87.85 F
00:31:17.838 -> Average Cal sensor Temp.: 102.92 Ambient Temp.: 87.85 F
00:31:19.397 -> Average Cal sensor Temp.: 102.89 Ambient Temp.: 87.85 F
00:31:20.953 -> Average Cal sensor Temp.: 102.60 Ambient Temp.: 87.82 F
00:31:22.544 -> Average Cal sensor Temp.: 102.93 Ambient Temp.: 87.82 F
00:31:24.101 -> Average Cal sensor Temp.: 102.68 Ambient Temp.: 87.89 F
00:31:25.664 -> Average Cal sensor Temp.: 103.05 Ambient Temp.: 87.93 F
00:31:27.225 -> Average Cal sensor Temp.: 102.71 Ambient Temp.: 87.96 F
00:31:28.782 -> Average Cal sensor Temp.: 102.54 Ambient Temp.: 87.89 F
00:31:30.366 -> Average Cal sensor Temp.: 102.69 Ambient Temp.: 87.93 F
00:31:31.923 -> Average Cal sensor Temp.: 102.59 Ambient Temp.: 87.93 F
00:31:33.486 -> Average Cal sensor Temp.: 103.04 Ambient Temp.: 87.93 F
00:31:35.042 -> Average Cal sensor Temp.: 103.06 Ambient Temp.: 87.89 F
00:31:36.599 -> Average Cal sensor Temp.: 103.17 Ambient Temp.: 87.96 F
00:31:38.193 -> Average Cal sensor Temp.: 102.99 Ambient Temp.: 87.93 F
00:31:39.743 -> Average Cal sensor Temp.: 102.78 Ambient Temp.: 87.93 F
00:31:41.299 -> Average Cal sensor Temp.: 103.18 Ambient Temp.: 87.96 F
00:31:42.855 -> Average Cal sensor Temp.: 103.01 Ambient Temp.: 87.89 F
00:31:44.451 -> Average Cal sensor Temp.: 102.86 Ambient Temp.: 87.93 F
00:31:46.009 -> Average Cal sensor Temp.: 102.61 Ambient Temp.: 87.96 F
00:31:47.570 -> Average Cal sensor Temp.: 103.00 Ambient Temp.: 87.89 F
00:31:49.134 -> Average Cal sensor Temp.: 102.99 Ambient Temp.: 87.85 F
00:31:50.692 -> Average Cal sensor Temp.: 102.93 Ambient Temp.: 87.85 F
00:31:52.248 -> Average Cal sensor Temp.: 102.84 Ambient Temp.: 87.85 F
00:31:53.808 -> Average Cal sensor Temp.: 102.85 Ambient Temp.: 87.89 F
00:31:55.398 -> Average Cal sensor Temp.: 102.93 Ambient Temp.: 87.85 F
00:31:56.957 -> Average Cal sensor Temp.: 103.01 Ambient Temp.: 87.85 F
00:31:58.515 -> Average Cal sensor Temp.: 103.08 Ambient Temp.: 87.82 F
00:32:00.079 -> Average Cal sensor Temp.: 102.94 Ambient Temp.: 87.89 F
00:32:01.637 -> Average Cal sensor Temp.: 103.04 Ambient Temp.: 87.85 F
00:32:03.201 -> Average Cal sensor Temp.: 103.24 Ambient Temp.: 87.82 F
00:32:04.794 -> Average Cal sensor Temp.: 102.65 Ambient Temp.: 87.82 F
00:32:06.351 -> Average Cal sensor Temp.: 102.74 Ambient Temp.: 87.78 F
00:32:07.910 -> Average Cal sensor Temp.: 102.96 Ambient Temp.: 87.82 F
00:32:09.468 -> Average Cal sensor Temp.: 102.78 Ambient Temp.: 87.82 F
00:32:11.019 -> Average Cal sensor Temp.: 102.91 Ambient Temp.: 87.75 F
00:32:12.614 -> Average Cal sensor Temp.: 102.56 Ambient Temp.: 87.71 F
00:32:14.170 -> Average Cal sensor Temp.: 102.92 Ambient Temp.: 87.75 F
00:32:15.727 -> Average Cal sensor Temp.: 102.94 Ambient Temp.: 87.75 F
00:32:17.282 -> Average Cal sensor Temp.: 102.82 Ambient Temp.: 87.75 F
00:32:18.836 -> Average Cal sensor Temp.: 102.54 Ambient Temp.: 87.71 F
00:32:20.431 -> Average Cal sensor Temp.: 102.79 Ambient Temp.: 87.71 F
00:32:21.988 -> Average Cal sensor Temp.: 102.87 Ambient Temp.: 87.67 F
00:32:23.542 -> Average Cal sensor Temp.: 102.74 Ambient Temp.: 87.67 F
00:32:25.100 -> Average Cal sensor Temp.: 103.09 Ambient Temp.: 87.75 F
00:32:26.655 -> Average Cal sensor Temp.: 103.01 Ambient Temp.: 87.60 F
00:32:28.244 -> Average Cal sensor Temp.: 102.84 Ambient Temp.: 87.71 F
00:32:29.803 -> Average Cal sensor Temp.: 102.79 Ambient Temp.: 87.67 F
00:32:31.359 -> Average Cal sensor Temp.: 102.76 Ambient Temp.: 87.60 F
00:32:32.912 -> Average Cal sensor Temp.: 102.88 Ambient Temp.: 87.57 F
00:32:34.507 -> Average Cal sensor Temp.: 103.27 Ambient Temp.: 87.57 F
00:32:36.068 -> Average Cal sensor Temp.: 103.39 Ambient Temp.: 87.57 F
00:32:37.627 -> Average Cal sensor Temp.: 103.13 Ambient Temp.: 87.53 F
00:32:39.185 -> Average Cal sensor Temp.: 103.08 Ambient Temp.: 87.57 F
00:32:40.746 -> Average Cal sensor Temp.: 103.01 Ambient Temp.: 87.57 F
00:32:42.300 -> Average Cal sensor Temp.: 102.85 Ambient Temp.: 87.53 F
00:32:43.860 -> Average Cal sensor Temp.: 103.26 Ambient Temp.: 87.49 F
00:32:45.423 -> Average Cal sensor Temp.: 103.29 Ambient Temp.: 87.46 F
00:32:46.985 -> Average Cal sensor Temp.: 103.08 Ambient Temp.: 87.49 F
00:32:48.582 -> Average Cal sensor Temp.: 103.15 Ambient Temp.: 87.46 F
00:32:50.146 -> Average Cal sensor Temp.: 102.80 Ambient Temp.: 87.46 F
00:32:51.703 -> Average Cal sensor Temp.: 103.19 Ambient Temp.: 87.49 F
00:32:53.261 -> Average Cal sensor Temp.: 103.34 Ambient Temp.: 87.46 F
00:32:54.830 -> Average Cal sensor Temp.: 103.29 Ambient Temp.: 87.46 F
00:32:56.387 -> Average Cal sensor Temp.: 103.19 Ambient Temp.: 87.49 F
00:32:57.945 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 87.46 F
00:32:59.505 -> Average Cal sensor Temp.: 103.78 Ambient Temp.: 87.49 F
00:33:01.098 -> Average Cal sensor Temp.: 103.52 Ambient Temp.: 87.46 F
00:33:02.658 -> Average Cal sensor Temp.: 103.01 Ambient Temp.: 87.46 F
00:33:04.214 -> Average Cal sensor Temp.: 103.18 Ambient Temp.: 87.46 F
00:33:05.778 -> Average Cal sensor Temp.: 103.22 Ambient Temp.: 87.42 F
00:33:07.342 -> Average Cal sensor Temp.: 103.41 Ambient Temp.: 87.42 F
00:33:08.899 -> Average Cal sensor Temp.: 103.17 Ambient Temp.: 87.42 F
00:33:10.460 -> Average Cal sensor Temp.: 103.32 Ambient Temp.: 87.49 F
00:33:12.019 -> Average Cal sensor Temp.: 103.21 Ambient Temp.: 87.42 F
00:33:13.581 -> Average Cal sensor Temp.: 103.23 Ambient Temp.: 87.39 F
00:33:15.141 -> Average Cal sensor Temp.: 103.22 Ambient Temp.: 87.42 F
00:33:16.704 -> Average Cal sensor Temp.: 102.01 Ambient Temp.: 87.42 F
00:33:18.292 -> Average Cal sensor Temp.: 100.42 Ambient Temp.: 87.35 F
00:33:19.851 -> Average Cal sensor Temp.: 100.80 Ambient Temp.: 87.28 F
00:33:21.410 -> Average Cal sensor Temp.: 102.18 Ambient Temp.: 87.17 F
00:33:22.966 -> Average Cal sensor Temp.: 103.72 Ambient Temp.: 87.13 F
00:33:24.528 -> Average Cal sensor Temp.: 103.43 Ambient Temp.: 87.10 F
00:33:26.084 -> Average Cal sensor Temp.: 103.54 Ambient Temp.: 87.13 F
00:33:27.679 -> Average Cal sensor Temp.: 103.31 Ambient Temp.: 87.06 F
00:33:29.242 -> Average Cal sensor Temp.: 102.72 Ambient Temp.: 86.99 F
00:33:30.798 -> Average Cal sensor Temp.: 103.57 Ambient Temp.: 86.99 F
00:33:32.351 -> Average Cal sensor Temp.: 103.53 Ambient Temp.: 86.95 F
00:33:33.902 -> Average Cal sensor Temp.: 103.17 Ambient Temp.: 86.95 F
00:33:35.492 -> Average Cal sensor Temp.: 103.32 Ambient Temp.: 86.92 F
00:33:37.048 -> Average Cal sensor Temp.: 102.99 Ambient Temp.: 86.92 F
00:33:38.603 -> Average Cal sensor Temp.: 103.07 Ambient Temp.: 86.88 F
00:33:40.170 -> Average Cal sensor Temp.: 103.20 Ambient Temp.: 86.81 F
00:33:41.760 -> Average Cal sensor Temp.: 103.08 Ambient Temp.: 86.81 F
00:33:43.315 -> Average Cal sensor Temp.: 103.56 Ambient Temp.: 86.88 F
00:33:44.872 -> Average Cal sensor Temp.: 103.40 Ambient Temp.: 86.81 F
00:33:46.426 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 86.77 F
00:33:48.019 -> Average Cal sensor Temp.: 103.42 Ambient Temp.: 86.77 F
00:33:49.582 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 86.74 F
00:33:51.115 -> Average Cal sensor Temp.: 103.39 Ambient Temp.: 86.81 F
00:33:52.680 -> Average Cal sensor Temp.: 103.59 Ambient Temp.: 86.77 F
00:33:54.246 -> Average Cal sensor Temp.: 103.13 Ambient Temp.: 86.74 F
00:33:55.807 -> Average Cal sensor Temp.: 103.19 Ambient Temp.: 86.74 F
00:33:57.399 -> Average Cal sensor Temp.: 103.26 Ambient Temp.: 86.70 F
00:33:58.952 -> Average Cal sensor Temp.: 103.33 Ambient Temp.: 86.70 F
00:34:00.511 -> Average Cal sensor Temp.: 103.76 Ambient Temp.: 86.74 F
00:34:02.070 -> Average Cal sensor Temp.: 103.36 Ambient Temp.: 86.74 F
00:34:03.662 -> Average Cal sensor Temp.: 103.23 Ambient Temp.: 86.74 F
00:34:05.221 -> Average Cal sensor Temp.: 103.38 Ambient Temp.: 86.70 F
00:34:06.777 -> Average Cal sensor Temp.: 103.13 Ambient Temp.: 86.59 F
00:34:08.332 -> Average Cal sensor Temp.: 103.59 Ambient Temp.: 86.63 F
00:34:09.891 -> Average Cal sensor Temp.: 103.38 Ambient Temp.: 86.59 F
00:34:11.485 -> Average Cal sensor Temp.: 103.33 Ambient Temp.: 86.59 F
00:34:13.045 -> Average Cal sensor Temp.: 103.28 Ambient Temp.: 86.59 F
00:34:14.612 -> Average Cal sensor Temp.: 103.66 Ambient Temp.: 86.59 F
00:34:16.174 -> Average Cal sensor Temp.: 103.41 Ambient Temp.: 86.56 F
00:34:17.728 -> Average Cal sensor Temp.: 103.21 Ambient Temp.: 86.56 F
00:34:19.286 -> Average Cal sensor Temp.: 103.05 Ambient Temp.: 86.49 F
00:34:20.851 -> Average Cal sensor Temp.: 103.10 Ambient Temp.: 86.52 F
00:34:22.411 -> Average Cal sensor Temp.: 103.33 Ambient Temp.: 86.56 F
00:34:23.967 -> Average Cal sensor Temp.: 103.35 Ambient Temp.: 86.52 F
00:34:25.563 -> Average Cal sensor Temp.: 103.49 Ambient Temp.: 86.56 F
00:34:27.125 -> Average Cal sensor Temp.: 103.16 Ambient Temp.: 86.49 F
00:34:28.689 -> Average Cal sensor Temp.: 103.33 Ambient Temp.: 86.49 F
00:34:30.243 -> Average Cal sensor Temp.: 103.29 Ambient Temp.: 86.52 F
00:34:31.804 -> Average Cal sensor Temp.: 103.24 Ambient Temp.: 86.52 F
00:34:33.365 -> Average Cal sensor Temp.: 103.34 Ambient Temp.: 86.45 F
00:34:34.924 -> Average Cal sensor Temp.: 103.31 Ambient Temp.: 86.52 F
00:34:36.482 -> Average Cal sensor Temp.: 103.46 Ambient Temp.: 86.45 F
00:34:38.074 -> Average Cal sensor Temp.: 103.30 Ambient Temp.: 86.52 F
00:34:39.636 -> Average Cal sensor Temp.: 103.18 Ambient Temp.: 86.49 F
00:34:41.190 -> Average Cal sensor Temp.: 103.33 Ambient Temp.: 86.49 F
00:34:42.748 -> Average Cal sensor Temp.: 103.39 Ambient Temp.: 86.52 F
00:34:44.338 -> Average Cal sensor Temp.: 103.36 Ambient Temp.: 86.49 F
00:34:45.897 -> Average Cal sensor Temp.: 103.47 Ambient Temp.: 86.52 F
00:34:47.458 -> Average Cal sensor Temp.: 103.23 Ambient Temp.: 86.52 F
00:34:49.020 -> Average Cal sensor Temp.: 103.08 Ambient Temp.: 86.49 F
00:34:50.581 -> Average Cal sensor Temp.: 103.39 Ambient Temp.: 86.52 F
00:34:52.140 -> Average Cal sensor Temp.: 103.28 Ambient Temp.: 86.52 F
00:34:53.698 -> Average Cal sensor Temp.: 103.40 Ambient Temp.: 86.49 F
00:34:55.258 -> Average Cal sensor Temp.: 103.49 Ambient Temp.: 86.56 F
00:34:56.857 -> Average Cal sensor Temp.: 103.17 Ambient Temp.: 86.56 F
00:34:58.412 -> Average Cal sensor Temp.: 103.46 Ambient Temp.: 86.56 F
00:34:59.972 -> Average Cal sensor Temp.: 103.54 Ambient Temp.: 86.63 F
00:35:01.525 -> Average Cal sensor Temp.: 103.59 Ambient Temp.: 86.59 F
00:35:03.082 -> Average Cal sensor Temp.: 103.63 Ambient Temp.: 86.63 F
00:35:04.678 -> Average Cal sensor Temp.: 103.58 Ambient Temp.: 86.63 F
00:35:06.233 -> Average Cal sensor Temp.: 103.83 Ambient Temp.: 86.67 F
00:35:07.795 -> Average Cal sensor Temp.: 103.56 Ambient Temp.: 86.67 F
00:35:09.358 -> Average Cal sensor Temp.: 103.37 Ambient Temp.: 86.67 F
00:35:10.926 -> Average Cal sensor Temp.: 103.42 Ambient Temp.: 86.63 F
00:35:12.481 -> Average Cal sensor Temp.: 103.60 Ambient Temp.: 86.74 F
00:35:14.043 -> Average Cal sensor Temp.: 103.41 Ambient Temp.: 86.70 F
00:35:15.604 -> Average Cal sensor Temp.: 103.28 Ambient Temp.: 86.67 F
00:35:17.164 -> Average Cal sensor Temp.: 103.31 Ambient Temp.: 86.70 F
00:35:18.759 -> Average Cal sensor Temp.: 103.33 Ambient Temp.: 86.74 F
00:35:20.319 -> Average Cal sensor Temp.: 103.62 Ambient Temp.: 86.74 F
00:35:21.883 -> Average Cal sensor Temp.: 103.49 Ambient Temp.: 86.74 F
00:35:23.442 -> Average Cal sensor Temp.: 103.92 Ambient Temp.: 86.77 F
00:35:25.003 -> Average Cal sensor Temp.: 103.75 Ambient Temp.: 86.81 F
00:35:26.563 -> Average Cal sensor Temp.: 103.47 Ambient Temp.: 86.88 F
00:35:28.121 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 86.88 F
00:35:29.683 -> Average Cal sensor Temp.: 103.43 Ambient Temp.: 86.92 F
00:35:31.275 -> Average Cal sensor Temp.: 103.27 Ambient Temp.: 86.95 F
00:35:32.829 -> Average Cal sensor Temp.: 103.18 Ambient Temp.: 86.95 F
00:35:34.389 -> Average Cal sensor Temp.: 103.33 Ambient Temp.: 86.99 F
00:35:35.949 -> Average Cal sensor Temp.: 102.95 Ambient Temp.: 87.03 F
00:35:37.509 -> Average Cal sensor Temp.: 103.16 Ambient Temp.: 86.99 F
00:35:39.076 -> Average Cal sensor Temp.: 103.15 Ambient Temp.: 87.03 F
00:35:40.640 -> Average Cal sensor Temp.: 103.47 Ambient Temp.: 87.03 F
00:35:42.194 -> Average Cal sensor Temp.: 103.58 Ambient Temp.: 87.06 F
00:35:43.789 -> Average Cal sensor Temp.: 103.70 Ambient Temp.: 87.10 F
00:35:45.347 -> Average Cal sensor Temp.: 103.69 Ambient Temp.: 87.06 F
00:35:46.902 -> Average Cal sensor Temp.: 103.55 Ambient Temp.: 87.10 F
00:35:48.461 -> Average Cal sensor Temp.: 103.64 Ambient Temp.: 87.17 F
00:35:50.019 -> Average Cal sensor Temp.: 103.51 Ambient Temp.: 87.10 F
00:35:51.578 -> Average Cal sensor Temp.: 103.62 Ambient Temp.: 87.13 F
00:35:53.173 -> Average Cal sensor Temp.: 103.67 Ambient Temp.: 87.17 F
00:35:54.701 -> Average Cal sensor Temp.: 103.59 Ambient Temp.: 87.28 F
00:35:56.300 -> Average Cal sensor Temp.: 103.52 Ambient Temp.: 87.21 F
00:35:57.829 -> Average Cal sensor Temp.: 103.33 Ambient Temp.: 87.28 F
00:35:59.424 -> Average Cal sensor Temp.: 103.31 Ambient Temp.: 87.21 F
00:36:00.983 -> Average Cal sensor Temp.: 103.65 Ambient Temp.: 87.28 F
00:36:02.538 -> Average Cal sensor Temp.: 103.51 Ambient Temp.: 87.28 F
00:36:04.096 -> Average Cal sensor Temp.: 103.55 Ambient Temp.: 87.28 F
00:36:05.653 -> Average Cal sensor Temp.: 103.43 Ambient Temp.: 87.31 F
00:36:07.216 -> Average Cal sensor Temp.: 103.42 Ambient Temp.: 87.35 F
00:36:08.777 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 87.35 F
00:36:10.335 -> Average Cal sensor Temp.: 103.72 Ambient Temp.: 87.42 F
00:36:11.899 -> Average Cal sensor Temp.: 103.63 Ambient Temp.: 87.39 F
00:36:13.493 -> Average Cal sensor Temp.: 103.36 Ambient Temp.: 87.39 F
00:36:15.059 -> Average Cal sensor Temp.: 103.47 Ambient Temp.: 87.42 F
00:36:16.617 -> Average Cal sensor Temp.: 103.58 Ambient Temp.: 87.42 F
00:36:18.173 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 87.42 F
00:36:19.736 -> Average Cal sensor Temp.: 103.57 Ambient Temp.: 87.42 F
00:36:21.291 -> Average Cal sensor Temp.: 103.83 Ambient Temp.: 87.46 F
00:36:22.845 -> Average Cal sensor Temp.: 103.74 Ambient Temp.: 87.49 F
00:36:24.436 -> Average Cal sensor Temp.: 103.74 Ambient Temp.: 87.49 F
00:36:25.992 -> Average Cal sensor Temp.: 104.01 Ambient Temp.: 87.49 F
00:36:27.548 -> Average Cal sensor Temp.: 103.84 Ambient Temp.: 87.53 F
00:36:29.108 -> Average Cal sensor Temp.: 103.69 Ambient Temp.: 87.57 F
00:36:30.694 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 87.57 F
00:36:32.250 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 87.67 F
00:36:33.804 -> Average Cal sensor Temp.: 103.57 Ambient Temp.: 87.57 F
00:36:35.360 -> Average Cal sensor Temp.: 103.77 Ambient Temp.: 87.67 F
00:36:36.918 -> Average Cal sensor Temp.: 103.80 Ambient Temp.: 87.60 F
00:36:38.512 -> Average Cal sensor Temp.: 103.88 Ambient Temp.: 87.75 F
00:36:40.077 -> Average Cal sensor Temp.: 103.87 Ambient Temp.: 87.75 F
00:36:41.635 -> Average Cal sensor Temp.: 103.84 Ambient Temp.: 87.78 F
00:36:43.193 -> Average Cal sensor Temp.: 103.82 Ambient Temp.: 87.75 F
00:36:44.749 -> Average Cal sensor Temp.: 103.90 Ambient Temp.: 87.78 F
00:36:46.308 -> Average Cal sensor Temp.: 103.95 Ambient Temp.: 87.82 F
00:36:47.899 -> Average Cal sensor Temp.: 103.84 Ambient Temp.: 87.82 F
00:36:49.463 -> Average Cal sensor Temp.: 103.76 Ambient Temp.: 87.85 F
00:36:51.028 -> Average Cal sensor Temp.: 104.03 Ambient Temp.: 87.93 F
00:36:52.592 -> Average Cal sensor Temp.: 103.60 Ambient Temp.: 87.89 F
00:36:54.155 -> Average Cal sensor Temp.: 103.87 Ambient Temp.: 87.93 F
00:36:55.720 -> Average Cal sensor Temp.: 103.54 Ambient Temp.: 87.93 F
00:36:57.271 -> Average Cal sensor Temp.: 103.60 Ambient Temp.: 87.89 F
00:36:58.829 -> Average Cal sensor Temp.: 103.56 Ambient Temp.: 87.89 F
00:37:00.420 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 87.93 F
00:37:01.973 -> Average Cal sensor Temp.: 103.85 Ambient Temp.: 87.96 F
00:37:03.529 -> Average Cal sensor Temp.: 103.53 Ambient Temp.: 87.96 F
00:37:05.091 -> Average Cal sensor Temp.: 103.46 Ambient Temp.: 88.00 F
00:37:06.652 -> Average Cal sensor Temp.: 103.88 Ambient Temp.: 88.00 F
00:37:08.246 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 88.07 F
00:37:09.800 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 88.07 F
00:37:11.360 -> Average Cal sensor Temp.: 104.05 Ambient Temp.: 88.07 F
00:37:12.925 -> Average Cal sensor Temp.: 103.79 Ambient Temp.: 88.07 F
00:37:14.487 -> Average Cal sensor Temp.: 103.76 Ambient Temp.: 88.07 F
00:37:16.048 -> Average Cal sensor Temp.: 103.89 Ambient Temp.: 88.11 F
00:37:17.635 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 88.11 F
00:37:19.191 -> Average Cal sensor Temp.: 103.51 Ambient Temp.: 88.11 F
00:37:20.747 -> Average Cal sensor Temp.: 104.01 Ambient Temp.: 88.18 F
00:37:22.305 -> Average Cal sensor Temp.: 103.96 Ambient Temp.: 88.11 F
00:37:23.864 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 88.14 F
00:37:25.454 -> Average Cal sensor Temp.: 103.63 Ambient Temp.: 88.18 F
00:37:27.016 -> Average Cal sensor Temp.: 103.33 Ambient Temp.: 88.18 F
00:37:28.574 -> Average Cal sensor Temp.: 103.33 Ambient Temp.: 88.18 F
00:37:30.130 -> Average Cal sensor Temp.: 103.76 Ambient Temp.: 88.14 F
00:37:31.686 -> Average Cal sensor Temp.: 104.05 Ambient Temp.: 88.14 F
00:37:33.278 -> Average Cal sensor Temp.: 103.82 Ambient Temp.: 88.18 F
00:37:34.827 -> Average Cal sensor Temp.: 104.10 Ambient Temp.: 88.14 F
00:37:36.382 -> Average Cal sensor Temp.: 104.25 Ambient Temp.: 88.18 F
00:37:37.939 -> Average Cal sensor Temp.: 104.13 Ambient Temp.: 88.21 F
00:37:39.536 -> Average Cal sensor Temp.: 103.46 Ambient Temp.: 88.21 F
00:37:41.094 -> Average Cal sensor Temp.: 103.95 Ambient Temp.: 88.25 F
00:37:42.657 -> Average Cal sensor Temp.: 103.70 Ambient Temp.: 88.18 F
00:37:44.218 -> Average Cal sensor Temp.: 103.97 Ambient Temp.: 88.21 F
00:37:45.775 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 88.29 F
00:37:47.329 -> Average Cal sensor Temp.: 103.58 Ambient Temp.: 88.29 F
00:37:48.922 -> Average Cal sensor Temp.: 103.65 Ambient Temp.: 88.25 F
00:37:50.485 -> Average Cal sensor Temp.: 103.84 Ambient Temp.: 88.25 F
00:37:52.044 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 88.29 F
00:37:53.608 -> Average Cal sensor Temp.: 103.80 Ambient Temp.: 88.21 F
00:37:55.169 -> Average Cal sensor Temp.: 103.79 Ambient Temp.: 88.25 F
00:37:56.733 -> Average Cal sensor Temp.: 104.31 Ambient Temp.: 88.25 F
00:37:58.283 -> Average Cal sensor Temp.: 103.80 Ambient Temp.: 88.21 F
00:37:59.878 -> Average Cal sensor Temp.: 103.67 Ambient Temp.: 88.36 F
00:38:01.430 -> Average Cal sensor Temp.: 103.51 Ambient Temp.: 88.25 F
00:38:02.987 -> Average Cal sensor Temp.: 103.37 Ambient Temp.: 88.25 F
00:38:04.548 -> Average Cal sensor Temp.: 103.20 Ambient Temp.: 88.21 F
00:38:06.103 -> Average Cal sensor Temp.: 103.40 Ambient Temp.: 88.18 F
00:38:07.697 -> Average Cal sensor Temp.: 103.62 Ambient Temp.: 88.18 F
00:38:09.260 -> Average Cal sensor Temp.: 103.78 Ambient Temp.: 88.14 F
00:38:10.816 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 88.14 F
00:38:12.375 -> Average Cal sensor Temp.: 104.06 Ambient Temp.: 88.18 F
00:38:13.936 -> Average Cal sensor Temp.: 104.03 Ambient Temp.: 88.21 F
00:38:15.497 -> Average Cal sensor Temp.: 103.97 Ambient Temp.: 88.21 F
00:38:17.058 -> Average Cal sensor Temp.: 103.95 Ambient Temp.: 88.25 F
00:38:18.620 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 88.21 F
00:38:20.183 -> Average Cal sensor Temp.: 103.78 Ambient Temp.: 88.21 F
00:38:21.780 -> Average Cal sensor Temp.: 103.91 Ambient Temp.: 88.21 F
00:38:23.338 -> Average Cal sensor Temp.: 103.54 Ambient Temp.: 88.25 F
00:38:24.903 -> Average Cal sensor Temp.: 103.83 Ambient Temp.: 88.18 F
00:38:26.461 -> Average Cal sensor Temp.: 103.68 Ambient Temp.: 88.21 F
00:38:28.024 -> Average Cal sensor Temp.: 103.71 Ambient Temp.: 88.18 F
00:38:29.585 -> Average Cal sensor Temp.: 103.77 Ambient Temp.: 88.14 F
00:38:31.147 -> Average Cal sensor Temp.: 103.68 Ambient Temp.: 88.14 F
00:38:32.708 -> Average Cal sensor Temp.: 103.58 Ambient Temp.: 88.07 F
00:38:34.266 -> Average Cal sensor Temp.: 103.48 Ambient Temp.: 88.11 F
00:38:35.825 -> Average Cal sensor Temp.: 103.88 Ambient Temp.: 88.00 F
00:38:37.417 -> Average Cal sensor Temp.: 103.94 Ambient Temp.: 88.07 F
00:38:38.975 -> Average Cal sensor Temp.: 103.82 Ambient Temp.: 87.96 F
00:38:40.540 -> Average Cal sensor Temp.: 103.78 Ambient Temp.: 88.00 F
00:38:42.095 -> Average Cal sensor Temp.: 103.79 Ambient Temp.: 88.07 F
00:38:43.685 -> Average Cal sensor Temp.: 104.01 Ambient Temp.: 88.07 F
00:38:45.246 -> Average Cal sensor Temp.: 104.02 Ambient Temp.: 87.96 F
00:38:46.801 -> Average Cal sensor Temp.: 104.04 Ambient Temp.: 87.96 F
00:38:48.357 -> Average Cal sensor Temp.: 103.99 Ambient Temp.: 87.93 F
00:38:49.917 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 87.93 F
00:38:51.510 -> Average Cal sensor Temp.: 103.58 Ambient Temp.: 87.89 F
00:38:53.069 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 87.85 F
00:38:54.628 -> Average Cal sensor Temp.: 103.77 Ambient Temp.: 87.82 F
00:38:56.184 -> Average Cal sensor Temp.: 103.98 Ambient Temp.: 87.78 F
00:38:57.743 -> Average Cal sensor Temp.: 103.89 Ambient Temp.: 87.78 F
00:38:59.303 -> Average Cal sensor Temp.: 104.01 Ambient Temp.: 87.78 F
00:39:00.886 -> Average Cal sensor Temp.: 104.15 Ambient Temp.: 87.78 F
00:39:02.442 -> Average Cal sensor Temp.: 103.92 Ambient Temp.: 87.75 F
00:39:04.002 -> Average Cal sensor Temp.: 104.17 Ambient Temp.: 87.78 F
00:39:05.563 -> Average Cal sensor Temp.: 104.25 Ambient Temp.: 87.78 F
00:39:07.120 -> Average Cal sensor Temp.: 104.42 Ambient Temp.: 87.78 F
00:39:08.711 -> Average Cal sensor Temp.: 104.21 Ambient Temp.: 87.75 F
00:39:10.267 -> Average Cal sensor Temp.: 104.13 Ambient Temp.: 87.75 F
00:39:11.827 -> Average Cal sensor Temp.: 104.25 Ambient Temp.: 87.71 F
00:39:13.382 -> Average Cal sensor Temp.: 103.91 Ambient Temp.: 87.67 F
00:39:14.937 -> Average Cal sensor Temp.: 104.17 Ambient Temp.: 87.67 F
00:39:16.532 -> Average Cal sensor Temp.: 103.59 Ambient Temp.: 87.57 F
00:39:18.090 -> Average Cal sensor Temp.: 103.88 Ambient Temp.: 87.67 F
00:39:19.647 -> Average Cal sensor Temp.: 103.91 Ambient Temp.: 87.60 F
00:39:21.204 -> Average Cal sensor Temp.: 103.37 Ambient Temp.: 87.60 F
00:39:22.766 -> Average Cal sensor Temp.: 103.56 Ambient Temp.: 87.60 F
00:39:24.329 -> Average Cal sensor Temp.: 103.57 Ambient Temp.: 87.53 F
00:39:25.889 -> Average Cal sensor Temp.: 103.64 Ambient Temp.: 87.53 F
00:39:27.447 -> Average Cal sensor Temp.: 103.89 Ambient Temp.: 87.53 F
00:39:29.036 -> Average Cal sensor Temp.: 104.06 Ambient Temp.: 87.49 F
00:39:30.593 -> Average Cal sensor Temp.: 103.62 Ambient Temp.: 87.46 F
00:39:32.153 -> Average Cal sensor Temp.: 103.68 Ambient Temp.: 87.42 F
00:39:33.719 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 87.42 F
00:39:35.280 -> Average Cal sensor Temp.: 103.98 Ambient Temp.: 87.42 F
00:39:36.870 -> Average Cal sensor Temp.: 103.79 Ambient Temp.: 87.35 F
00:39:38.424 -> Average Cal sensor Temp.: 103.59 Ambient Temp.: 87.31 F
00:39:39.987 -> Average Cal sensor Temp.: 104.07 Ambient Temp.: 87.31 F
00:39:41.544 -> Average Cal sensor Temp.: 103.76 Ambient Temp.: 87.31 F
00:39:43.102 -> Average Cal sensor Temp.: 103.54 Ambient Temp.: 87.28 F
00:39:44.661 -> Average Cal sensor Temp.: 104.02 Ambient Temp.: 87.21 F
00:39:46.219 -> Average Cal sensor Temp.: 104.23 Ambient Temp.: 87.21 F
00:39:47.814 -> Average Cal sensor Temp.: 104.23 Ambient Temp.: 87.17 F
00:39:49.373 -> Average Cal sensor Temp.: 104.19 Ambient Temp.: 87.17 F
00:39:50.932 -> Average Cal sensor Temp.: 103.96 Ambient Temp.: 87.10 F
00:39:52.489 -> Average Cal sensor Temp.: 103.95 Ambient Temp.: 87.13 F
00:39:54.048 -> Average Cal sensor Temp.: 104.25 Ambient Temp.: 87.10 F
00:39:55.609 -> Average Cal sensor Temp.: 104.03 Ambient Temp.: 87.10 F
00:39:57.197 -> Average Cal sensor Temp.: 104.17 Ambient Temp.: 87.06 F
00:39:58.757 -> Average Cal sensor Temp.: 104.23 Ambient Temp.: 86.99 F
00:40:00.319 -> Average Cal sensor Temp.: 104.23 Ambient Temp.: 87.03 F
00:40:01.869 -> Average Cal sensor Temp.: 103.74 Ambient Temp.: 86.99 F
00:40:03.429 -> Average Cal sensor Temp.: 104.04 Ambient Temp.: 86.92 F
00:40:05.019 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 86.95 F
00:40:06.575 -> Average Cal sensor Temp.: 104.12 Ambient Temp.: 86.99 F
00:40:08.129 -> Average Cal sensor Temp.: 104.77 Ambient Temp.: 86.95 F
00:40:09.696 -> Average Cal sensor Temp.: 103.65 Ambient Temp.: 86.88 F
00:40:11.257 -> Average Cal sensor Temp.: 103.63 Ambient Temp.: 86.88 F
00:40:12.815 -> Average Cal sensor Temp.: 103.87 Ambient Temp.: 86.77 F
00:40:14.410 -> Average Cal sensor Temp.: 104.01 Ambient Temp.: 86.77 F
00:40:15.968 -> Average Cal sensor Temp.: 104.09 Ambient Temp.: 86.81 F
00:40:17.526 -> Average Cal sensor Temp.: 104.18 Ambient Temp.: 86.74 F
00:40:19.087 -> Average Cal sensor Temp.: 104.20 Ambient Temp.: 86.74 F
00:40:20.644 -> Average Cal sensor Temp.: 104.18 Ambient Temp.: 86.74 F
00:40:22.234 -> Average Cal sensor Temp.: 103.99 Ambient Temp.: 86.70 F
00:40:23.793 -> Average Cal sensor Temp.: 104.07 Ambient Temp.: 86.77 F
00:40:25.353 -> Average Cal sensor Temp.: 103.98 Ambient Temp.: 86.74 F
00:40:26.909 -> Average Cal sensor Temp.: 104.11 Ambient Temp.: 86.74 F
00:40:28.465 -> Average Cal sensor Temp.: 104.32 Ambient Temp.: 86.70 F
00:40:30.059 -> Average Cal sensor Temp.: 104.33 Ambient Temp.: 86.67 F
00:40:31.618 -> Average Cal sensor Temp.: 104.54 Ambient Temp.: 86.70 F
00:40:33.178 -> Average Cal sensor Temp.: 104.51 Ambient Temp.: 86.74 F
00:40:34.736 -> Average Cal sensor Temp.: 104.38 Ambient Temp.: 86.70 F
00:40:36.291 -> Average Cal sensor Temp.: 104.30 Ambient Temp.: 86.70 F
00:40:37.882 -> Average Cal sensor Temp.: 104.04 Ambient Temp.: 86.70 F
00:40:39.442 -> Average Cal sensor Temp.: 103.97 Ambient Temp.: 86.63 F
00:40:41.002 -> Average Cal sensor Temp.: 104.04 Ambient Temp.: 86.70 F
00:40:42.557 -> Average Cal sensor Temp.: 104.13 Ambient Temp.: 86.67 F
00:40:44.142 -> Average Cal sensor Temp.: 104.09 Ambient Temp.: 86.63 F
00:40:45.696 -> Average Cal sensor Temp.: 104.35 Ambient Temp.: 86.70 F
00:40:47.252 -> Average Cal sensor Temp.: 104.43 Ambient Temp.: 86.70 F
00:40:48.815 -> Average Cal sensor Temp.: 104.34 Ambient Temp.: 86.70 F
00:40:50.406 -> Average Cal sensor Temp.: 104.23 Ambient Temp.: 86.67 F
00:40:51.960 -> Average Cal sensor Temp.: 104.79 Ambient Temp.: 86.74 F
00:40:53.517 -> Average Cal sensor Temp.: 104.65 Ambient Temp.: 86.74 F
00:40:55.071 -> Average Cal sensor Temp.: 104.37 Ambient Temp.: 86.70 F
00:40:56.657 -> Average Cal sensor Temp.: 104.47 Ambient Temp.: 86.74 F
00:40:58.207 -> Average Cal sensor Temp.: 104.36 Ambient Temp.: 86.74 F
00:40:59.770 -> Average Cal sensor Temp.: 104.41 Ambient Temp.: 86.77 F
00:41:01.328 -> Average Cal sensor Temp.: 104.12 Ambient Temp.: 86.74 F
00:41:02.918 -> Average Cal sensor Temp.: 104.05 Ambient Temp.: 86.81 F
00:41:04.478 -> Average Cal sensor Temp.: 104.41 Ambient Temp.: 86.77 F
00:41:06.030 -> Average Cal sensor Temp.: 104.64 Ambient Temp.: 86.88 F
00:41:07.620 -> Average Cal sensor Temp.: 104.55 Ambient Temp.: 86.88 F
00:41:09.182 -> Average Cal sensor Temp.: 104.29 Ambient Temp.: 86.88 F
00:41:10.739 -> Average Cal sensor Temp.: 104.29 Ambient Temp.: 86.92 F
00:41:12.295 -> Average Cal sensor Temp.: 104.36 Ambient Temp.: 86.92 F
00:41:13.851 -> Average Cal sensor Temp.: 104.29 Ambient Temp.: 86.92 F
00:41:15.449 -> Average Cal sensor Temp.: 104.35 Ambient Temp.: 86.92 F
00:41:17.009 -> Average Cal sensor Temp.: 104.15 Ambient Temp.: 86.92 F
00:41:18.569 -> Average Cal sensor Temp.: 104.23 Ambient Temp.: 86.95 F
00:41:20.133 -> Average Cal sensor Temp.: 104.26 Ambient Temp.: 86.95 F
00:41:21.697 -> Average Cal sensor Temp.: 104.19 Ambient Temp.: 86.95 F
00:41:23.252 -> Average Cal sensor Temp.: 104.19 Ambient Temp.: 86.95 F
00:41:24.815 -> Average Cal sensor Temp.: 104.50 Ambient Temp.: 87.03 F
00:41:26.380 -> Average Cal sensor Temp.: 104.35 Ambient Temp.: 87.03 F
00:41:27.937 -> Average Cal sensor Temp.: 104.40 Ambient Temp.: 86.99 F
00:41:29.533 -> Average Cal sensor Temp.: 104.42 Ambient Temp.: 87.06 F
00:41:31.065 -> Average Cal sensor Temp.: 104.62 Ambient Temp.: 87.10 F
00:41:32.660 -> Average Cal sensor Temp.: 104.67 Ambient Temp.: 87.03 F
00:41:34.223 -> Average Cal sensor Temp.: 104.60 Ambient Temp.: 87.06 F
00:41:35.782 -> Average Cal sensor Temp.: 104.66 Ambient Temp.: 87.06 F
00:41:37.343 -> Average Cal sensor Temp.: 104.69 Ambient Temp.: 87.13 F
00:41:38.908 -> Average Cal sensor Temp.: 104.69 Ambient Temp.: 87.17 F
00:41:40.471 -> Average Cal sensor Temp.: 104.61 Ambient Temp.: 87.13 F
00:41:42.034 -> Average Cal sensor Temp.: 104.67 Ambient Temp.: 87.13 F
00:41:43.591 -> Average Cal sensor Temp.: 104.63 Ambient Temp.: 87.17 F
00:41:45.148 -> Average Cal sensor Temp.: 104.39 Ambient Temp.: 87.17 F
00:41:46.710 -> Average Cal sensor Temp.: 105.44 Ambient Temp.: 87.17 F
00:41:48.306 -> Average Cal sensor Temp.: 104.82 Ambient Temp.: 87.21 F
00:41:49.862 -> Average Cal sensor Temp.: 104.69 Ambient Temp.: 87.31 F
00:41:51.423 -> Average Cal sensor Temp.: 104.85 Ambient Temp.: 87.28 F
00:41:52.980 -> Average Cal sensor Temp.: 104.90 Ambient Temp.: 87.31 F
00:41:54.542 -> Average Cal sensor Temp.: 105.11 Ambient Temp.: 87.39 F
00:41:56.101 -> Average Cal sensor Temp.: 104.72 Ambient Temp.: 87.35 F
00:41:57.691 -> Average Cal sensor Temp.: 104.63 Ambient Temp.: 87.39 F
00:41:59.247 -> Average Cal sensor Temp.: 104.72 Ambient Temp.: 87.35 F
00:42:00.801 -> Average Cal sensor Temp.: 104.91 Ambient Temp.: 87.46 F
00:42:02.358 -> Average Cal sensor Temp.: 104.97 Ambient Temp.: 87.42 F
00:42:03.917 -> Average Cal sensor Temp.: 104.91 Ambient Temp.: 87.46 F
00:42:05.511 -> Average Cal sensor Temp.: 104.58 Ambient Temp.: 87.53 F
00:42:07.072 -> Average Cal sensor Temp.: 104.75 Ambient Temp.: 87.49 F
00:42:08.634 -> Average Cal sensor Temp.: 104.79 Ambient Temp.: 87.49 F
00:42:10.199 -> Average Cal sensor Temp.: 104.89 Ambient Temp.: 87.57 F
00:42:11.754 -> Average Cal sensor Temp.: 104.90 Ambient Temp.: 87.57 F
00:42:13.306 -> Average Cal sensor Temp.: 104.80 Ambient Temp.: 87.57 F
00:42:14.892 -> Average Cal sensor Temp.: 104.82 Ambient Temp.: 87.67 F
00:42:16.445 -> Average Cal sensor Temp.: 104.67 Ambient Temp.: 87.67 F
00:42:18.008 -> Average Cal sensor Temp.: 104.54 Ambient Temp.: 87.71 F
00:42:19.570 -> Average Cal sensor Temp.: 104.54 Ambient Temp.: 87.67 F
00:42:21.126 -> Average Cal sensor Temp.: 104.62 Ambient Temp.: 87.71 F
00:42:22.713 -> Average Cal sensor Temp.: 104.71 Ambient Temp.: 87.71 F
00:42:24.270 -> Average Cal sensor Temp.: 104.50 Ambient Temp.: 87.71 F
00:42:25.825 -> Average Cal sensor Temp.: 104.90 Ambient Temp.: 87.71 F
00:42:27.383 -> Average Cal sensor Temp.: 104.85 Ambient Temp.: 87.75 F
00:42:28.977 -> Average Cal sensor Temp.: 104.94 Ambient Temp.: 87.78 F
00:42:30.532 -> Average Cal sensor Temp.: 104.73 Ambient Temp.: 87.78 F
00:42:32.082 -> Average Cal sensor Temp.: 104.95 Ambient Temp.: 87.78 F
00:42:33.666 -> Average Cal sensor Temp.: 104.51 Ambient Temp.: 87.82 F
00:42:35.220 -> Average Cal sensor Temp.: 104.61 Ambient Temp.: 87.82 F
00:42:36.780 -> Average Cal sensor Temp.: 105.02 Ambient Temp.: 87.82 F
00:42:38.338 -> Average Cal sensor Temp.: 104.89 Ambient Temp.: 87.89 F
00:42:39.927 -> Average Cal sensor Temp.: 104.49 Ambient Temp.: 87.89 F
00:42:41.486 -> Average Cal sensor Temp.: 104.76 Ambient Temp.: 87.93 F
00:42:43.041 -> Average Cal sensor Temp.: 104.67 Ambient Temp.: 88.00 F
00:42:44.603 -> Average Cal sensor Temp.: 105.47 Ambient Temp.: 87.96 F
00:42:46.158 -> Average Cal sensor Temp.: 105.33 Ambient Temp.: 87.96 F
00:42:47.723 -> Average Cal sensor Temp.: 104.68 Ambient Temp.: 88.07 F
00:42:49.288 -> Average Cal sensor Temp.: 104.90 Ambient Temp.: 88.07 F
00:42:50.883 -> Average Cal sensor Temp.: 105.01 Ambient Temp.: 88.00 F
00:42:52.444 -> Average Cal sensor Temp.: 104.96 Ambient Temp.: 88.07 F
00:42:54.001 -> Average Cal sensor Temp.: 104.88 Ambient Temp.: 88.07 F
00:42:55.561 -> Average Cal sensor Temp.: 104.71 Ambient Temp.: 88.11 F
00:42:57.116 -> Average Cal sensor Temp.: 104.58 Ambient Temp.: 88.11 F
00:42:58.712 -> Average Cal sensor Temp.: 104.69 Ambient Temp.: 88.14 F
00:43:00.270 -> Average Cal sensor Temp.: 105.21 Ambient Temp.: 88.18 F
00:43:01.825 -> Average Cal sensor Temp.: 105.10 Ambient Temp.: 88.18 F
00:43:03.386 -> Average Cal sensor Temp.: 105.38 Ambient Temp.: 88.18 F
00:43:04.939 -> Average Cal sensor Temp.: 105.33 Ambient Temp.: 88.25 F
00:43:06.526 -> Average Cal sensor Temp.: 105.09 Ambient Temp.: 88.25 F
00:43:08.085 -> Average Cal sensor Temp.: 105.12 Ambient Temp.: 88.32 F
00:43:09.647 -> Average Cal sensor Temp.: 105.02 Ambient Temp.: 88.32 F
00:43:11.204 -> Average Cal sensor Temp.: 104.95 Ambient Temp.: 88.29 F
00:43:12.760 -> Average Cal sensor Temp.: 104.94 Ambient Temp.: 88.29 F
00:43:14.350 -> Average Cal sensor Temp.: 104.94 Ambient Temp.: 88.32 F
00:43:15.905 -> Average Cal sensor Temp.: 104.86 Ambient Temp.: 88.29 F
00:43:17.492 -> Average Cal sensor Temp.: 105.07 Ambient Temp.: 88.39 F
00:43:19.045 -> Average Cal sensor Temp.: 104.69 Ambient Temp.: 88.32 F
00:43:20.603 -> Average Cal sensor Temp.: 104.68 Ambient Temp.: 88.36 F
00:43:22.162 -> Average Cal sensor Temp.: 105.08 Ambient Temp.: 88.36 F
00:43:23.752 -> Average Cal sensor Temp.: 104.85 Ambient Temp.: 88.36 F
00:43:25.313 -> Average Cal sensor Temp.: 105.10 Ambient Temp.: 88.39 F
00:43:26.860 -> Average Cal sensor Temp.: 105.11 Ambient Temp.: 88.39 F
00:43:28.445 -> Average Cal sensor Temp.: 105.16 Ambient Temp.: 88.39 F
00:43:30.006 -> Average Cal sensor Temp.: 104.94 Ambient Temp.: 88.39 F
00:43:31.570 -> Average Cal sensor Temp.: 105.16 Ambient Temp.: 88.54 F
00:43:33.133 -> Average Cal sensor Temp.: 105.40 Ambient Temp.: 88.50 F
00:43:34.700 -> Average Cal sensor Temp.: 104.85 Ambient Temp.: 88.54 F
00:43:36.263 -> Average Cal sensor Temp.: 104.99 Ambient Temp.: 88.54 F
00:43:37.822 -> Average Cal sensor Temp.: 104.58 Ambient Temp.: 88.50 F
00:43:39.387 -> Average Cal sensor Temp.: 105.28 Ambient Temp.: 88.57 F
00:43:40.944 -> Average Cal sensor Temp.: 105.18 Ambient Temp.: 88.57 F
00:43:42.503 -> Average Cal sensor Temp.: 105.18 Ambient Temp.: 88.61 F
00:43:44.097 -> Average Cal sensor Temp.: 105.46 Ambient Temp.: 88.65 F
00:43:45.657 -> Average Cal sensor Temp.: 105.17 Ambient Temp.: 88.65 F
00:43:47.214 -> Average Cal sensor Temp.: 105.37 Ambient Temp.: 88.65 F
00:43:48.769 -> Average Cal sensor Temp.: 105.06 Ambient Temp.: 88.65 F
00:43:50.327 -> Average Cal sensor Temp.: 104.83 Ambient Temp.: 88.68 F
00:43:51.923 -> Average Cal sensor Temp.: 105.03 Ambient Temp.: 88.65 F
00:43:53.483 -> Average Cal sensor Temp.: 105.32 Ambient Temp.: 88.68 F
00:43:55.043 -> Average Cal sensor Temp.: 105.00 Ambient Temp.: 88.68 F
00:43:56.603 -> Average Cal sensor Temp.: 105.00 Ambient Temp.: 88.75 F
00:43:58.161 -> Average Cal sensor Temp.: 104.76 Ambient Temp.: 88.75 F
00:43:59.727 -> Average Cal sensor Temp.: 104.93 Ambient Temp.: 88.68 F
00:44:01.285 -> Average Cal sensor Temp.: 105.05 Ambient Temp.: 88.72 F
00:44:02.847 -> Average Cal sensor Temp.: 104.74 Ambient Temp.: 88.72 F
00:44:04.412 -> Average Cal sensor Temp.: 104.71 Ambient Temp.: 88.65 F
00:44:06.007 -> Average Cal sensor Temp.: 105.17 Ambient Temp.: 88.72 F
00:44:07.568 -> Average Cal sensor Temp.: 105.07 Ambient Temp.: 88.72 F
00:44:09.131 -> Average Cal sensor Temp.: 105.18 Ambient Temp.: 88.75 F
00:44:10.686 -> Average Cal sensor Temp.: 105.40 Ambient Temp.: 88.72 F
00:44:12.242 -> Average Cal sensor Temp.: 105.21 Ambient Temp.: 88.72 F
00:44:13.802 -> Average Cal sensor Temp.: 104.97 Ambient Temp.: 88.72 F
00:44:15.391 -> Average Cal sensor Temp.: 105.66 Ambient Temp.: 88.75 F
00:44:16.942 -> Average Cal sensor Temp.: 105.01 Ambient Temp.: 88.86 F
00:44:18.500 -> Average Cal sensor Temp.: 104.36 Ambient Temp.: 88.75 F
00:44:20.058 -> Average Cal sensor Temp.: 104.25 Ambient Temp.: 88.72 F
00:44:21.647 -> Average Cal sensor Temp.: 104.94 Ambient Temp.: 88.75 F
00:44:23.199 -> Average Cal sensor Temp.: 105.15 Ambient Temp.: 88.72 F
00:44:24.755 -> Average Cal sensor Temp.: 104.85 Ambient Temp.: 88.75 F
00:44:26.314 -> Average Cal sensor Temp.: 104.86 Ambient Temp.: 88.72 F
00:44:27.903 -> Average Cal sensor Temp.: 104.53 Ambient Temp.: 88.72 F
00:44:29.459 -> Average Cal sensor Temp.: 104.44 Ambient Temp.: 88.72 F
00:44:31.017 -> Average Cal sensor Temp.: 104.17 Ambient Temp.: 88.68 F
00:44:32.574 -> Average Cal sensor Temp.: 104.01 Ambient Temp.: 88.65 F
00:44:34.138 -> Average Cal sensor Temp.: 104.08 Ambient Temp.: 88.65 F
00:44:35.703 -> Average Cal sensor Temp.: 104.36 Ambient Temp.: 88.65 F
00:44:37.296 -> Average Cal sensor Temp.: 104.03 Ambient Temp.: 88.65 F
00:44:38.850 -> Average Cal sensor Temp.: 104.40 Ambient Temp.: 88.68 F
00:44:40.410 -> Average Cal sensor Temp.: 104.76 Ambient Temp.: 88.65 F
00:44:41.967 -> Average Cal sensor Temp.: 105.11 Ambient Temp.: 88.68 F
00:44:43.555 -> Average Cal sensor Temp.: 104.47 Ambient Temp.: 88.72 F
00:44:45.111 -> Average Cal sensor Temp.: 104.50 Ambient Temp.: 88.65 F
00:44:46.671 -> Average Cal sensor Temp.: 104.58 Ambient Temp.: 88.61 F
00:44:48.229 -> Average Cal sensor Temp.: 104.42 Ambient Temp.: 88.68 F
00:44:49.820 -> Average Cal sensor Temp.: 104.30 Ambient Temp.: 88.61 F
00:44:51.375 -> Average Cal sensor Temp.: 104.21 Ambient Temp.: 88.61 F
00:44:52.931 -> Average Cal sensor Temp.: 104.65 Ambient Temp.: 88.65 F
00:44:54.496 -> Average Cal sensor Temp.: 104.68 Ambient Temp.: 88.65 F
00:44:56.053 -> Average Cal sensor Temp.: 104.55 Ambient Temp.: 88.65 F
00:44:57.644 -> Average Cal sensor Temp.: 104.17 Ambient Temp.: 88.61 F
00:44:59.208 -> Average Cal sensor Temp.: 104.31 Ambient Temp.: 88.57 F
00:45:00.767 -> Average Cal sensor Temp.: 104.22 Ambient Temp.: 88.57 F
00:45:02.327 -> Average Cal sensor Temp.: 104.26 Ambient Temp.: 88.57 F
00:45:03.884 -> Average Cal sensor Temp.: 104.11 Ambient Temp.: 88.57 F
00:45:05.444 -> Average Cal sensor Temp.: 104.33 Ambient Temp.: 88.54 F
00:45:07.002 -> Average Cal sensor Temp.: 103.74 Ambient Temp.: 88.54 F
00:45:08.594 -> Average Cal sensor Temp.: 103.78 Ambient Temp.: 88.50 F
00:45:10.153 -> Average Cal sensor Temp.: 103.95 Ambient Temp.: 88.50 F
00:45:11.718 -> Average Cal sensor Temp.: 104.59 Ambient Temp.: 88.57 F
00:45:13.272 -> Average Cal sensor Temp.: 104.11 Ambient Temp.: 88.47 F
00:45:14.836 -> Average Cal sensor Temp.: 103.55 Ambient Temp.: 88.47 F
00:45:16.395 -> Average Cal sensor Temp.: 103.88 Ambient Temp.: 88.39 F
00:45:17.955 -> Average Cal sensor Temp.: 103.59 Ambient Temp.: 88.39 F
00:45:19.543 -> Average Cal sensor Temp.: 103.75 Ambient Temp.: 88.39 F
00:45:21.106 -> Average Cal sensor Temp.: 103.47 Ambient Temp.: 88.36 F
00:45:22.666 -> Average Cal sensor Temp.: 103.55 Ambient Temp.: 88.36 F
00:45:24.221 -> Average Cal sensor Temp.: 103.85 Ambient Temp.: 88.36 F
00:45:25.779 -> Average Cal sensor Temp.: 103.96 Ambient Temp.: 88.32 F
00:45:27.369 -> Average Cal sensor Temp.: 104.06 Ambient Temp.: 88.29 F
00:45:28.925 -> Average Cal sensor Temp.: 103.83 Ambient Temp.: 88.32 F
00:45:30.482 -> Average Cal sensor Temp.: 103.74 Ambient Temp.: 88.25 F