## Sheet1
Reference Clinical Thermometer Readout (Y var.) | SET-4 |  |  | SET-5 |  |  | SET-6 |  |  | SET-7 |  |  | SET-AVERAGE |  |  |  | Average Regression
| Reference Water Temperature (Sensor) | Calibrating contactless sensor Readout (2-3 cm) | Calibrating contactless sensor AMBIENT Readout | Reference Water Temperature (Sensor) | Calibrating contactless sensor Readout (2-3 cm) | Calibrating contactless sensor AMBIENT Readout | Reference Water Temperature (Sensor) | Calibrating contactless sensor Readout (2-3 cm) | Calibrating contactless sensor AMBIENT Readout | Reference Water Temperature (Sensor) | Calibrating contactless sensor Readout (2-3 cm) | Calibrating contactless sensor AMBIENT Readout | Reference Water Temperature (Sensor) | Calibrating contactless sensor Readout (2-3 cm) (X var.) | DIFF between Water temp and clinical | DIFF bw clinical and sensor | Regressed values Calibrating sensor (ŷ = 1.03983X - 0.66635) | Diff between Clinical and Regressed Calibrating sensor
Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F)
94 | 93.09 | 91.19 | 94.23 | 92.41 | 90.06 | 93.9 | 93.2 | 91.43 | 95.27 | 93.31 | 91.2 | 94.59 | 93.0025 | 90.97 | 0.9975 | 3.03 | 93.927 | 0.073015
94.4 | 93.65 | 91.75 | 94.33 | 92.75 | 90.55 | 94.12 | 93.54 | 91.89 | 95.34 | 93.65 | 91.44 | 94.66 | 93.3975 | 91.4075 | 1.0025 | 2.9925 | 94.3819 | 0.018089
94.8 | 93.99 | 92.15 | 94.37 | 92.97 | 91.09 | 94.15 | 93.76 | 92.06 | 95.34 | 93.99 | 91.94 | 94.8 | 93.6775 | 91.81 | 1.1225 | 2.99 | 94.8004 | -0.0004423
95 | 94.21 | 92.28 | 94.37 | 93.99 | 91.4 | 94.3 | 93.99 | 92.26 | 95.41 | 94.21 | 92.21 | 94.77 | 94.1 | 92.0375 | 0.9 | 2.9625 | 95.037 | -0.037004
95.4 | 94.55 | 92.75 | 94.51 | 94.44 | 92.11 | 94.48 | 94.44 | 92.87 | 95.49 | 94.44 | 92.33 | 94.8 | 94.4675 | 92.515 | 0.9325 | 2.885 | 95.5335 | -0.13352
95.8 | 94.89 | 93.15 | 94.55 | 94.77 | 92.51 | 94.48 | 94.89 | 93.14 | 95.56 | 94.77 | 92.69 | 94.95 | 94.83 | 92.8725 | 0.97 | 2.9275 | 95.9053 | -0.10526
96 | 95 | 93.36 | 94.66 | 94.89 | 92.71 | 94.55 | 95.11 | 93.24 | 95.56 | 95 | 93.06 | 94.98 | 95 | 93.0925 | 1 | 2.9075 | 96.134 | -0.13402
96.4 | 95.35 | 93.65 | 94.66 | 95.34 | 93.23 | 94.77 | 95.56 | 93.38 | 95.67 | 95.45 | 93.45 | 95.05 | 95.425 | 93.4275 | 0.975 | 2.9725 | 96.4824 | -0.082367
96.8 | 95.68 | 96.94 | 94.66 | 95.56 | 93.47 | 94.8 | 95.79 | 93.56 | 95.77 | 95.68 | 93.81 | 95.16 | 95.6775 | 94.445 | 1.1225 | 2.355 | 97.5404 | -0.74039
97 | 95.9 | 94.17 | 94.73 | 95.79 | 93.55 | 94.8 | 96.01 | 93.82 | 95.81 | 96.01 | 94.17 | 95.27 | 95.9275 | 93.9275 | 1.0725 | 3.0725 | 97.0023 | -0.0022823
97.4 | 96.35 | 94.35 | 94.8 | 96.12 | 93.9 | 94.91 | 96.46 | 94.44 | 95.88 | 96.35 | 94.35 | 95.34 | 96.32 | 94.26 | 1.08 | 3.14 | 97.348 | 0.051974
97.8 | 96.8 | 94.69 | 94.91 | 96.57 | 94.43 | 94.98 | 96.69 | 94.76 | 95.92 | 96.69 | 94.74 | 95.45 | 96.6875 | 94.655 | 1.1125 | 3.145 | 97.7588 | 0.041241
98 | 96.91 | 95.07 | 94.98 | 96.8 | 94.86 | 94.98 | 96.91 | 95.02 | 95.99 | 97.02 | 94.95 | 95.52 | 96.91 | 94.975 | 1.09 | 3.025 | 98.0915 | -0.091504
98.2 | 97.14 | 95.23 | 95.05 | 97.02 | 94.99 | 95.02 | 97.14 | 95.17 | 96.03 | 97.25 | 95.16 | 95.56 | 97.1375 | 95.1375 | 1.0625 | 3.0625 | 98.2605 | -0.060477
98.4 | 97.25 | 95.36 | 95.05 | 97.25 | 95.15 | 94.98 | 97.36 | 95.38 | 96.21 | 97.36 | 95.37 | 95.59 | 97.305 | 95.315 | 1.095 | 3.085 | 98.445 | -0.045046
98.6 | 97.47 | 95.55 | 95.13 | 97.47 | 95.28 | 95.05 | 97.47 | 95.62 | 96.17 | 97.47 | 95.5 | 95.67 | 97.47 | 95.4875 | 1.13 | 3.1125 | 98.6244 | -0.024417
98.8 | 97.7 | 95.7 | 95.16 | 97.7 | 95.45 | 95.13 | 97.7 | 95.8 | 96.21 | 97.7 | 95.67 | 95.7 | 97.7 | 95.655 | 1.1 | 3.145 | 98.7986 | 0.0014113
99 | 97.81 | 95.68 | 95.16 | 97.81 | 95.63 | 95.16 | 97.92 | 95.93 | 96.24 | 97.92 | 95.74 | 95.7 | 97.865 | 95.745 | 1.135 | 3.255 | 98.8922 | 0.10783
99.2 | 98.04 | 95.86 | 95.16 | 98.15 | 95.81 | 95.16 | 98.04 | 96.15 | 96.24 | 98.15 | 95.93 | 95.77 | 98.095 | 95.9375 | 1.105 | 3.2625 | 99.0923 | 0.10766
99.4 | 98.26 | 96.56 | 95.2 | 98.26 | 95.92 | 95.27 | 98.37 | 96.33 | 96.35 | 98.37 | 96.16 | 95.81 | 98.315 | 96.2425 | 1.085 | 3.1575 | 99.4095 | -0.0094888
99.6 | 98.49 | 96.84 | 95.31 | 98.37 | 96.06 | 95.31 | 98.49 | 96.48 | 96.35 | 98.49 | 96.26 | 95.77 | 98.46 | 96.41 | 1.14 | 3.19 | 99.5837 | 0.01634
99.8 | 98.6 | 96.9 | 95.31 | 98.6 | 96.21 | 95.34 | 98.82 | 96.86 | 96.42 | 98.6 | 96.52 | 95.81 | 98.655 | 96.6225 | 1.145 | 3.1775 | 99.8046 | -0.0046242
100 | 98.82 | 96.86 | 95.34 | 98.82 | 96.49 | 95.41 | 98.71 | 97 | 96.42 | 98.82 | 96.74 | 95.85 | 98.7925 | 96.7725 | 1.2075 | 3.2275 | 99.9606 | 0.039401
100.2 | 99.05 | 97.15 | 95.34 | 99.05 | 96.63 | 95.52 | 99.05 | 96.93 | 96.53 | 99.05 | 96.89 | 95.92 | 99.05 | 96.9 | 1.15 | 3.3 | 100.0932 | 0.10682
100.4 | 99.16 | 97.35 | 95.34 | 99.16 | 96.88 | 95.52 | 99.16 | 97.14 | 96.53 | 99.27 | 97.09 | 95.95 | 99.1875 | 97.115 | 1.2125 | 3.285 | 100.3167 | 0.08326
100.6 | 99.39 | 97.48 | 95.41 | 99.39 | 97.07 | 95.59 | 99.27 | 97.23 | 96.53 | 99.5 | 97.26 | 96.03 | 99.3875 | 97.26 | 1.2125 | 3.34 | 100.4675 | 0.13248
100.8 | 99.61 | 97.65 | 95.52 | 99.5 | 97.28 | 95.59 | 99.61 | 97.45 | 96.64 | 99.61 | 97.33 | 95.99 | 99.5825 | 97.4275 | 1.2175 | 3.3725 | 100.6417 | 0.15831
101 | 99.84 | 97.88 | 95.41 | 99.61 | 97.52 | 95.56 | 99.72 | 97.65 | 96.67 | 99.84 | 97.65 | 96.1 | 99.7525 | 97.675 | 1.2475 | 3.325 | 100.899 | 0.10095
101.2 | 100.06 | 98.09 | 95.49 | 99.95 | 97.65 | 95.67 | 99.72 | 97.88 | 96.64 | 100.06 | 97.85 | 96.13 | 99.9475 | 97.8675 | 1.2525 | 3.3325 | 101.0992 | 0.10079
101.4 | 100.17 | 98.4 | 95.52 | 100.06 | 97.85 | 95.77 | 99.84 | 98.02 | 96.71 | 100.17 | 98.07 | 96.17 | 100.06 | 98.085 | 1.34 | 3.315 | 101.3254 | 0.074624
101.6 | 100.51 | 98.53 | 95.42 | 100.29 | 98 | 95.7 | 100.29 | 98.18 | 96.75 | 100.51 | 98.16 | 96.17 | 100.4 | 98.2175 | 1.2 | 3.3825 | 101.4632 | 0.13685
101.8 | 100.62 | 98.74 | 95.56 | 100.51 | 98.15 | 95.81 | 100.62 | 98.24 | 96.78 | 100.62 | 98.45 | 96.21 | 100.5925 | 98.395 | 1.2075 | 3.405 | 101.6477 | 0.15228
102 | 100.74 | 99.03 | 95.7 | 100.74 | 98.4 | 95.85 | 100.74 | 98.65 | 96.78 | 100.74 | 98.62 | 96.21 | 100.74 | 98.675 | 1.26 | 3.325 | 101.9389 | 0.061125
102.2 | 101.07 | 99.15 | 95.7 | 100.85 | 98.65 | 95.88 | 101.07 | 98.81 | 96.82 | 100.96 | 98.86 | 96.21 | 100.9875 | 98.8675 | 1.2125 | 3.3325 | 102.139 | 0.060957
102.4 | 101.19 | 99.29 | 95.85 | 100.96 | 98.75 | 95.92 | 101.19 | 99.02 | 96.93 | 101.3 | 99.1 | 96.35 | 101.16 | 99.04 | 1.24 | 3.36 | 102.3184 | 0.081587
102.6 | 101.3 | 99.45 | 95.85 | 101.19 | 98.92 | 96.13 | 101.52 | 99.19 | 96.96 | 101.41 | 99.24 | 96.42 | 101.355 | 99.2 | 1.245 | 3.4 | 102.4848 | 0.11521
102.8 | 101.52 | 99.76 | 95.85 | 101.41 | 99.12 | 96.1 | 101.64 | 99.39 | 96.96 | 101.75 | 99.42 | 96.53 | 101.58 | 99.4225 | 1.22 | 3.3775 | 102.7161 | 0.083852
103 | 101.64 | 99.6 | 95.81 | 101.75 | 99.37 | 95.99 | 101.75 | 99.45 | 97.07 | 101.86 | 99.59 | 96.57 | 101.75 | 99.5025 | 1.25 | 3.4975 | 102.7993 | 0.20067
103.4 | 101.97 | 100.01 | 95.85 | 102.2 | 99.77 | 96.21 | 102.09 | 100 | 97.14 | 102.2 | 100.08 | 96.49 | 102.115 | 99.965 | 1.285 | 3.435 | 103.2803 | 0.11974
103.8 | 102.54 | 100.59 | 96.03 | 102.42 | 100.17 | 96.28 | 102.2 | 100.22 | 97.11 | 102.65 | 100.45 | 96.71 | 102.4525 | 100.3575 | 1.3475 | 3.4425 | 103.6884 | 0.11161
104 | 102.76 | 100.89 | 96.1 | 102.54 | 100.56 | 96.31 | 102.31 | 100.3 | 97.18 | 102.76 | 100.58 | 96.78 | 102.5925 | 100.5825 | 1.4075 | 3.4175 | 103.9224 | 0.077649
104.4 | 103.1 | 101.46 | 96.13 | 102.99 | 100.69 | 96.35 | 102.65 | 100.62 | 97.11 | 102.65 | 100.85 | 96.67 | 102.8475 | 100.905 | 1.5525 | 3.495 | 104.2577 | 0.1423
104.8 | 103.55 | 101.95 | 96.31 | 103.44 | 101.16 | 96.53 | 103.55 | 101.09 | 97.36 | 102.87 | 101.05 | 96.75 | 103.3525 | 101.3125 | 1.4475 | 3.4875 | 104.6814 | 0.11857
105 | 103.89 | 102.22 | 96.03 | 103.66 | 101.48 | 96.53 | 103.66 | 101.43 | 97.39 | 103.21 | 101.25 | 96.82 | 103.605 | 101.595 | 1.395 | 3.405 | 104.9752 | 0.024821
105.4 | 104.67 | 102.47 | 96.13 | 104.11 | 101.88 | 96.64 | 104.34 | 102.27 | 97.47 | 103.66 | 101.56 | 96.93 | 104.195 | 102.045 | 1.205 | 3.355 | 105.4431 | -0.043102
105.8 | 105.12 | 102.9 | 96.21 | 104.45 | 102.25 | 96.64 | 104.45 | 102.65 | 97.47 | 103.77 | 101.81 | 96.89 | 104.4475 | 102.4025 | 1.3525 | 3.3975 | 105.8148 | -0.014842
106 | 105.35 | 103.36 | 96.31 | 104.67 | 102.58 | 96.64 | 104.56 | 102.78 | 97.61 | 103.89 | 102.22 | 97 | 104.6175 | 102.735 | 1.3825 | 3.265 | 106.1606 | -0.16059
106.4 | 105.69 | 103.75 | 96.35 | 105.01 | 102.87 | 96.78 | 105.01 | 103.25 | 97.79 | 104.9 | 102.5 | 97.03 | 105.1525 | 103.0925 | 1.2475 | 3.3075 | 106.5323 | -0.13232
106.8 | 105.8 | 104.08 | 96.35 | 105.46 | 103.11 | 96.82 | 105.35 | 103.57 | 97.86 | 105.35 | 103.38 | 97.14 | 105.49 | 103.535 | 1.31 | 3.265 | 106.9924 | -0.19245
107 | 106.02 | 104.23 | 96.6 | 105.8 | 103.47 | 96.89 | 105.69 | 103.84 | 97.9 | 105.46 | 103.5 | 97.18 | 105.7425 | 103.76 | 1.2575 | 3.24 | 107.2264 | -0.22641
107.4 | 106.47 | 104.51 | 96.53 | 106.14 | 103.82 | 97.03 | 105.8 | 104.16 | 97.93 | 105.8 | 103.88 | 97.21 | 106.0525 | 104.0925 | 1.3475 | 3.3075 | 107.5722 | -0.17215
107.8 | 106.7 | 104.75 | 96.53 | 106.47 | 104.12 | 97.14 | 106.14 | 104.55 | 97.9 | 106.25 | 104.14 | 97.21 | 106.39 | 104.39 | 1.41 | 3.41 | 107.8815 | -0.081504
108 | 106.92 | 104.97 | 96.57 | 106.59 | 104.3 | 97.11 | 106.47 | 104.65 | 98.08 | 106.47 | 104.54 | 97.39 | 106.6125 | 104.615 | 1.3875 | 3.385 | 108.1155 | -0.11547
[14 empty] | AVERAGE OFFSET of | 1.1959 | 3.2273
ŷ = 1.03983X - 0.66635