## SET-6
Reference Clinical Thermometer Readout | Reference DIGI Thermometer Readout | Reference Water Temperature (Sensor) | Calibrating contactless sensor Readout (2-3 cm) | Calibrating contactless sensor AMBIENT Readout | DIFF bw clinical and sensor | DIFF bw DIGI and sensor | DIFF BETWEEN CLINICAL AND DIGI
Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F)
94 | 93.7 | 93.2 | 91.43 | 95.27 | 2.57 | 2.27 | 0.3
94.4 | 94.2 | 93.54 | 91.89 | 95.34 | 2.51 | 2.31 | 0.2
94.8 | 94.4 | 93.76 | 92.06 | 95.34 | 2.74 | 2.34 | 0.4
95 | 95.1 | 93.99 | 92.26 | 95.41 | 2.74 | 2.84 | -0.1
95.4 | 95.3 | 94.44 | 92.87 | 95.49 | 2.53 | 2.43 | 0.1
95.8 | 95.5 | 94.89 | 93.14 | 95.56 | 2.66 | 2.36 | 0.3
96 | 95.7 | 95.11 | 93.24 | 95.56 | 2.76 | 2.46 | 0.3
96.4 | 96.2 | 95.56 | 93.38 | 95.67 | 3.02 | 2.82 | 0.2
96.8 | 96.8 | 95.79 | 93.56 | 95.77 | 3.24 | 3.24 | 0
97 | 96.9 | 96.01 | 93.82 | 95.81 | 3.18 | 3.08 | 0.1
97.4 | 97.4 | 96.46 | 94.44 | 95.88 | 2.96 | 2.96 | 0
97.8 | 97.5 | 96.69 | 94.76 | 95.92 | 3.04 | 2.74 | 0.3
98 | 98 | 96.91 | 95.02 | 95.99 | 2.98 | 2.98 | 0
98.2 | 98.2 | 97.14 | 95.17 | 96.03 | 3.03 | 3.03 | 0
98.4 | 98.3 | 97.36 | 95.38 | 96.21 | 3.02 | 2.92 | 0.1
98.6 | 98.5 | 97.47 | 95.62 | 96.17 | 2.98 | 2.88 | 0.1
98.8 | 99 | 97.7 | 95.8 | 96.21 | 3 | 3.2 | -0.2
99 | 99.1 | 97.92 | 95.93 | 96.24 | 3.07 | 3.17 | -0.1
99.2 | 99.1 | 98.04 | 96.15 | 96.24 | 3.05 | 2.95 | 0.1
99.4 | 99.6 | 98.37 | 96.33 | 96.35 | 3.07 | 3.27 | -0.2
99.6 | 99.7 | 98.49 | 96.48 | 96.35 | 3.12 | 3.22 | -0.1
99.8 | 100 | 98.82 | 96.86 | 96.42 | 2.94 | 3.14 | -0.2
100 | 100.3 | 98.71 | 97 | 96.42 | 3 | 3.3 | -0.3
100.2 | 100.3 | 99.05 | 96.93 | 96.53 | 3.27 | 3.37 | -0.1
100.4 | 100.5 | 99.16 | 97.14 | 96.53 | 3.26 | 3.36 | -0.1
100.6 | 100.7 | 99.27 | 97.23 | 96.53 | 3.37 | 3.47 | -0.1
100.8 | 100.8 | 99.61 | 97.45 | 96.64 | 3.35 | 3.35 | 0
101 | 101.1 | 99.72 | 97.65 | 96.67 | 3.35 | 3.45 | -0.1
101.2 | 101.2 | 99.72 | 97.88 | 96.64 | 3.32 | 3.32 | 0
101.4 | 101.4 | 99.84 | 98.02 | 96.71 | 3.38 | 3.38 | 0
101.6 | 101.5 | 100.29 | 98.18 | 96.75 | 3.42 | 3.32 | 0.1
101.8 | 101.8 | 100.62 | 98.24 | 96.78 | 3.56 | 3.56 | 0
102 | 102.2 | 100.74 | 98.65 | 96.78 | 3.35 | 3.55 | -0.2
102.2 | 102.4 | 101.07 | 98.81 | 96.82 | 3.39 | 3.59 | -0.2
102.4 | 102.5 | 101.19 | 99.02 | 96.93 | 3.38 | 3.48 | -0.1
102.6 | 102.8 | 101.52 | 99.19 | 96.96 | 3.41 | 3.61 | -0.2
102.8 | 103.1 | 101.64 | 99.39 | 96.96 | 3.41 | 3.71 | -0.3
103 | 103.2 | 101.75 | 99.45 | 97.07 | 3.55 | 3.75 | -0.2
103.4 | 103.7 | 102.09 | 100 | 97.14 | 3.4 | 3.7 | -0.3
103.8 | 104 | 102.2 | 100.22 | 97.11 | 3.58 | 3.78 | -0.2
104 | 104.2 | 102.31 | 100.3 | 97.18 | 3.7 | 3.9 | -0.2
104.4 | 104.4 | 102.65 | 100.62 | 97.11 | 3.78 | 3.78 | 0
104.8 | 104.9 | 103.55 | 101.09 | 97.36 | 3.71 | 3.81 | -0.1
105 | 105.2 | 103.66 | 101.43 | 97.39 | 3.57 | 3.77 | -0.2
105.4 | 105.9 | 104.34 | 102.27 | 97.47 | 3.13 | 3.63 | -0.5
105.8 | 106.3 | 104.45 | 102.65 | 97.47 | 3.15 | 3.65 | -0.5
106 | 106.3 | 104.56 | 102.78 | 97.61 | 3.22 | 3.52 | -0.3
106.4 | 107 | 105.01 | 103.25 | 97.79 | 3.15 | 3.75 | -0.6
106.8 | 107 | 105.35 | 103.57 | 97.86 | 3.23 | 3.43 | -0.2
107 | 107.4 | 105.69 | 103.84 | 97.9 | 3.16 | 3.56 | -0.4
107.4 | 107.9 | 105.8 | 104.16 | 97.93 | 3.24 | 3.74 | -0.5
107.8 | 108.3 | 106.14 | 104.55 | 97.9 | 3.25 | 3.75 | -0.5
108 | 108.7 | 106.47 | 104.65 | 98.08 | 3.35 | 4.05 | -0.7

## RAW_HEATER-6
22:06:54.394 -> Water_Temp.:88.02, PWM:19
22:06:55.178 -> Water_Temp.:88.02, PWM:19
22:06:55.965 -> Water_Temp.:88.02, PWM:19
22:06:56.725 -> Water_Temp.:88.02, PWM:19
22:06:57.515 -> Water_Temp.:88.02, PWM:19
22:06:58.306 -> Water_Temp.:88.02, PWM:19
22:06:59.064 -> Water_Temp.:88.02, PWM:19
22:06:59.863 -> Water_Temp.:88.02, PWM:19
22:07:00.622 -> Water_Temp.:88.02, PWM:19
22:07:01.413 -> Water_Temp.:88.02, PWM:19
22:07:02.205 -> Water_Temp.:88.02, PWM:19
22:07:02.964 -> Water_Temp.:88.02, PWM:19
22:07:03.756 -> Water_Temp.:88.02, PWM:19
22:07:04.518 -> Water_Temp.:88.02, PWM:19
22:07:05.310 -> Water_Temp.:88.02, PWM:19
22:07:06.107 -> Water_Temp.:88.02, PWM:19
22:07:06.863 -> Water_Temp.:88.02, PWM:19
22:07:07.656 -> Water_Temp.:88.02, PWM:19
22:07:08.449 -> Water_Temp.:88.02, PWM:19
22:07:09.210 -> Water_Temp.:88.02, PWM:19
22:07:10.002 -> Water_Temp.:88.02, PWM:19
22:07:10.766 -> Water_Temp.:88.02, PWM:19
22:07:11.559 -> Water_Temp.:88.02, PWM:19
22:07:12.317 -> Water_Temp.:88.02, PWM:19
22:07:13.111 -> Water_Temp.:88.02, PWM:19
22:07:13.899 -> Water_Temp.:88.02, PWM:19
22:07:14.658 -> Water_Temp.:88.02, PWM:19
22:07:16.243 -> Water_Temp.:88.02, PWM:19
22:07:17.005 -> Water_Temp.:88.02, PWM:19
22:07:17.799 -> Water_Temp.:88.02, PWM:19
22:07:18.555 -> Water_Temp.:88.02, PWM:19
22:07:19.350 -> Water_Temp.:88.02, PWM:19
22:07:20.144 -> Water_Temp.:88.02, PWM:18
22:07:20.898 -> Water_Temp.:88.02, PWM:19
22:07:21.692 -> Water_Temp.:88.02, PWM:19
22:07:22.487 -> Water_Temp.:88.14, PWM:19
22:07:23.243 -> Water_Temp.:88.14, PWM:19
22:07:24.034 -> Water_Temp.:88.14, PWM:19
22:07:24.822 -> Water_Temp.:88.14, PWM:19
22:07:25.577 -> Water_Temp.:88.14, PWM:19
22:07:26.369 -> Water_Temp.:88.14, PWM:19
22:07:27.161 -> Water_Temp.:88.14, PWM:19
22:07:27.918 -> Water_Temp.:88.14, PWM:19
22:07:28.709 -> Water_Temp.:88.14, PWM:19
22:07:29.504 -> Water_Temp.:88.14, PWM:19
22:07:30.266 -> Water_Temp.:88.14, PWM:19
22:07:31.061 -> Water_Temp.:88.14, PWM:19
22:07:31.820 -> Water_Temp.:88.14, PWM:18
22:07:32.615 -> Water_Temp.:88.14, PWM:19
22:07:33.373 -> Water_Temp.:88.14, PWM:19
22:07:34.170 -> Water_Temp.:88.14, PWM:19
22:07:35.725 -> Water_Temp.:88.14, PWM:19
22:07:36.486 -> Water_Temp.:88.14, PWM:19
22:07:37.277 -> Water_Temp.:88.14, PWM:18
22:07:38.033 -> Water_Temp.:88.14, PWM:19
22:07:39.586 -> Water_Temp.:88.25, PWM:19
22:07:40.384 -> Water_Temp.:88.25, PWM:19
22:07:41.174 -> Water_Temp.:88.25, PWM:18
22:07:41.936 -> Water_Temp.:88.25, PWM:19
22:07:42.730 -> Water_Temp.:88.25, PWM:19
22:07:44.276 -> Water_Temp.:88.25, PWM:19
22:07:45.038 -> Water_Temp.:88.25, PWM:19
22:07:45.834 -> Water_Temp.:88.25, PWM:19
22:07:46.627 -> Water_Temp.:88.25, PWM:19
22:07:47.382 -> Water_Temp.:88.25, PWM:19
22:07:48.175 -> Water_Temp.:88.25, PWM:19
22:07:48.968 -> Water_Temp.:88.25, PWM:19
22:07:49.731 -> Water_Temp.:88.25, PWM:19
22:07:50.525 -> Water_Temp.:88.25, PWM:19
22:07:51.284 -> Water_Temp.:88.25, PWM:19
22:07:52.080 -> Water_Temp.:88.25, PWM:19
22:07:52.869 -> Water_Temp.:88.25, PWM:19
22:07:53.630 -> Water_Temp.:88.25, PWM:19
22:07:54.417 -> Water_Temp.:88.25, PWM:19
22:07:55.180 -> Water_Temp.:88.25, PWM:19
22:07:55.974 -> Water_Temp.:88.25, PWM:19
22:07:56.769 -> Water_Temp.:88.25, PWM:19
22:07:57.527 -> Water_Temp.:88.25, PWM:19
22:07:58.317 -> Water_Temp.:88.36, PWM:19
22:07:59.079 -> Water_Temp.:88.36, PWM:19
22:07:59.869 -> Water_Temp.:88.36, PWM:19
22:08:00.666 -> Water_Temp.:88.36, PWM:19
22:08:01.428 -> Water_Temp.:88.36, PWM:19
22:08:02.224 -> Water_Temp.:88.36, PWM:19
22:08:02.984 -> Water_Temp.:88.36, PWM:19
22:08:03.772 -> Water_Temp.:88.36, PWM:18
22:08:04.568 -> Water_Temp.:88.36, PWM:19
22:08:05.327 -> Water_Temp.:88.36, PWM:19
22:08:06.125 -> Water_Temp.:88.36, PWM:19
22:08:06.881 -> Water_Temp.:88.36, PWM:19
22:08:07.677 -> Water_Temp.:88.36, PWM:19
22:08:08.436 -> Water_Temp.:88.47, PWM:19
22:08:09.232 -> Water_Temp.:88.47, PWM:19
22:08:10.028 -> Water_Temp.:88.36, PWM:19
22:08:10.790 -> Water_Temp.:88.47, PWM:19
22:08:11.582 -> Water_Temp.:88.47, PWM:19
22:08:12.340 -> Water_Temp.:88.47, PWM:19
22:08:13.899 -> Water_Temp.:88.47, PWM:19
22:08:14.659 -> Water_Temp.:88.47, PWM:18
22:08:15.455 -> Water_Temp.:88.47, PWM:19
22:08:16.219 -> Water_Temp.:88.47, PWM:19
22:08:17.016 -> Water_Temp.:88.47, PWM:19
22:08:17.808 -> Water_Temp.:88.47, PWM:18
22:08:18.568 -> Water_Temp.:88.47, PWM:19
22:08:19.362 -> Water_Temp.:88.47, PWM:19
22:08:20.122 -> Water_Temp.:88.47, PWM:19
22:08:20.920 -> Water_Temp.:88.47, PWM:18
22:08:21.679 -> Water_Temp.:88.47, PWM:18
22:08:22.472 -> Water_Temp.:88.47, PWM:19
22:08:23.264 -> Water_Temp.:88.47, PWM:19
22:08:26.341 -> Water_Temp.:88.59, PWM:19
22:08:27.890 -> Water_Temp.:88.59, PWM:18
22:08:28.678 -> Water_Temp.:88.59, PWM:19
22:08:29.474 -> Water_Temp.:88.59, PWM:19
22:08:30.236 -> Water_Temp.:88.59, PWM:19
22:08:31.033 -> Water_Temp.:88.59, PWM:19
22:08:31.787 -> Water_Temp.:88.59, PWM:19
22:08:32.580 -> Water_Temp.:88.59, PWM:19
22:08:33.377 -> Water_Temp.:88.59, PWM:19
22:08:34.137 -> Water_Temp.:88.59, PWM:19
22:08:34.935 -> Water_Temp.:88.59, PWM:19
22:08:35.697 -> Water_Temp.:88.59, PWM:19
22:08:36.497 -> Water_Temp.:88.59, PWM:19
22:08:37.254 -> Water_Temp.:88.59, PWM:19
22:08:38.051 -> Water_Temp.:88.70, PWM:19
22:08:38.811 -> Water_Temp.:88.70, PWM:19
22:08:39.604 -> Water_Temp.:88.70, PWM:19
22:08:40.366 -> Water_Temp.:88.70, PWM:19
22:08:41.161 -> Water_Temp.:88.70, PWM:19
22:08:42.682 -> Water_Temp.:88.70, PWM:19
22:08:43.472 -> Water_Temp.:88.70, PWM:18
22:08:44.269 -> Water_Temp.:88.70, PWM:19
22:08:45.026 -> Water_Temp.:88.70, PWM:19
22:08:45.821 -> Water_Temp.:88.70, PWM:19
22:08:46.580 -> Water_Temp.:88.70, PWM:19
22:08:47.377 -> Water_Temp.:88.70, PWM:19
22:08:48.173 -> Water_Temp.:88.70, PWM:19
22:08:49.727 -> Water_Temp.:88.70, PWM:19
22:08:50.488 -> Water_Temp.:88.70, PWM:19
22:08:51.277 -> Water_Temp.:88.70, PWM:19
22:08:52.072 -> Water_Temp.:88.70, PWM:19
22:08:52.831 -> Water_Temp.:88.70, PWM:19
22:08:53.627 -> Water_Temp.:88.70, PWM:19
22:08:54.390 -> Water_Temp.:88.81, PWM:19
22:08:55.179 -> Water_Temp.:88.70, PWM:19
22:08:55.943 -> Water_Temp.:88.70, PWM:19
22:08:56.741 -> Water_Temp.:88.81, PWM:19
22:08:57.506 -> Water_Temp.:88.81, PWM:19
22:08:58.300 -> Water_Temp.:88.81, PWM:19
22:08:59.064 -> Water_Temp.:88.81, PWM:19
22:08:59.858 -> Water_Temp.:88.81, PWM:19
22:09:00.619 -> Water_Temp.:88.81, PWM:19
22:09:01.415 -> Water_Temp.:88.81, PWM:19
22:09:02.210 -> Water_Temp.:88.81, PWM:19
22:09:02.971 -> Water_Temp.:88.81, PWM:19
22:09:03.764 -> Water_Temp.:88.81, PWM:19
22:09:04.525 -> Water_Temp.:88.81, PWM:19
22:09:05.324 -> Water_Temp.:88.81, PWM:19
22:09:06.087 -> Water_Temp.:88.81, PWM:19
22:09:06.883 -> Water_Temp.:88.81, PWM:19
22:09:07.644 -> Water_Temp.:88.81, PWM:19
22:09:08.436 -> Water_Temp.:88.81, PWM:19
22:09:09.231 -> Water_Temp.:88.81, PWM:19
22:09:09.986 -> Water_Temp.:88.81, PWM:19
22:09:10.780 -> Water_Temp.:88.81, PWM:19
22:09:11.541 -> Water_Temp.:88.81, PWM:19
22:09:12.330 -> Water_Temp.:88.81, PWM:19
22:09:13.124 -> Water_Temp.:88.81, PWM:19
22:09:13.882 -> Water_Temp.:88.81, PWM:19
22:09:14.676 -> Water_Temp.:88.81, PWM:19
22:09:15.436 -> Water_Temp.:88.81, PWM:19
22:09:16.230 -> Water_Temp.:88.81, PWM:23
22:09:17.022 -> Water_Temp.:88.81, PWM:23
22:09:17.784 -> Water_Temp.:88.81, PWM:23
22:09:18.577 -> Water_Temp.:88.81, PWM:23
22:09:19.338 -> Water_Temp.:88.81, PWM:23
22:09:20.129 -> Water_Temp.:88.93, PWM:23
22:09:20.926 -> Water_Temp.:88.93, PWM:23
22:09:21.685 -> Water_Temp.:88.93, PWM:23
22:09:22.478 -> Water_Temp.:88.93, PWM:23
22:09:23.242 -> Water_Temp.:88.93, PWM:23
22:09:24.033 -> Water_Temp.:88.93, PWM:23
22:09:24.826 -> Water_Temp.:88.93, PWM:23
22:09:25.586 -> Water_Temp.:88.93, PWM:23
22:09:26.380 -> Water_Temp.:88.93, PWM:23
22:09:27.141 -> Water_Temp.:88.93, PWM:23
22:09:27.935 -> Water_Temp.:88.93, PWM:23
22:09:28.726 -> Water_Temp.:88.93, PWM:23
22:09:29.487 -> Water_Temp.:88.93, PWM:23
22:09:30.282 -> Water_Temp.:88.93, PWM:23
22:09:31.040 -> Water_Temp.:88.93, PWM:23
22:09:31.834 -> Water_Temp.:88.93, PWM:23
22:09:32.625 -> Water_Temp.:88.93, PWM:23
22:09:34.150 -> Water_Temp.:88.93, PWM:23
22:09:34.943 -> Water_Temp.:89.04, PWM:23
22:09:35.705 -> Water_Temp.:89.04, PWM:23
22:09:36.503 -> Water_Temp.:89.04, PWM:23
22:09:37.261 -> Water_Temp.:89.04, PWM:23
22:09:38.056 -> Water_Temp.:89.04, PWM:23
22:09:38.816 -> Water_Temp.:89.04, PWM:23
22:09:39.614 -> Water_Temp.:89.04, PWM:23
22:09:40.373 -> Water_Temp.:89.04, PWM:23
22:09:41.165 -> Water_Temp.:89.04, PWM:23
22:09:41.957 -> Water_Temp.:89.04, PWM:23
22:09:42.718 -> Water_Temp.:89.04, PWM:23
22:09:43.512 -> Water_Temp.:89.04, PWM:23
22:09:44.273 -> Water_Temp.:89.04, PWM:23
22:09:45.069 -> Water_Temp.:89.04, PWM:23
22:09:46.623 -> Water_Temp.:89.04, PWM:23
22:09:47.415 -> Water_Temp.:89.04, PWM:23
22:09:48.173 -> Water_Temp.:89.04, PWM:23
22:09:48.968 -> Water_Temp.:89.15, PWM:23
22:09:49.759 -> Water_Temp.:89.15, PWM:23
22:09:50.521 -> Water_Temp.:89.15, PWM:23
22:09:51.315 -> Water_Temp.:89.15, PWM:23
22:09:52.074 -> Water_Temp.:89.15, PWM:23
22:09:52.871 -> Water_Temp.:89.15, PWM:23
22:09:53.628 -> Water_Temp.:89.15, PWM:23
22:09:54.422 -> Water_Temp.:89.15, PWM:23
22:09:55.220 -> Water_Temp.:89.15, PWM:23
22:09:55.977 -> Water_Temp.:89.15, PWM:23
22:09:56.771 -> Water_Temp.:89.15, PWM:23
22:09:57.532 -> Water_Temp.:89.15, PWM:23
22:09:58.327 -> Water_Temp.:89.15, PWM:23
22:09:59.120 -> Water_Temp.:89.15, PWM:23
22:09:59.878 -> Water_Temp.:89.15, PWM:23
22:10:00.674 -> Water_Temp.:89.15, PWM:23
22:10:02.195 -> Water_Temp.:89.15, PWM:23
22:10:03.745 -> Water_Temp.:89.15, PWM:23
22:10:05.334 -> Water_Temp.:89.15, PWM:23
22:10:06.092 -> Water_Temp.:89.15, PWM:23
22:10:06.884 -> Water_Temp.:89.26, PWM:23
22:10:07.645 -> Water_Temp.:89.26, PWM:23
22:10:08.441 -> Water_Temp.:89.26, PWM:23
22:10:09.233 -> Water_Temp.:89.26, PWM:23
22:10:09.987 -> Water_Temp.:89.26, PWM:23
22:10:10.787 -> Water_Temp.:89.26, PWM:23
22:10:11.546 -> Water_Temp.:89.26, PWM:23
22:10:12.337 -> Water_Temp.:89.26, PWM:23
22:10:13.889 -> Water_Temp.:89.26, PWM:23
22:10:14.656 -> Water_Temp.:89.26, PWM:23
22:10:15.451 -> Water_Temp.:89.26, PWM:23
22:10:16.213 -> Water_Temp.:89.26, PWM:23
22:10:17.009 -> Water_Temp.:89.26, PWM:23
22:10:17.770 -> Water_Temp.:89.26, PWM:23
22:10:18.567 -> Water_Temp.:89.26, PWM:23
22:10:19.326 -> Water_Temp.:89.26, PWM:23
22:10:20.123 -> Water_Temp.:89.37, PWM:23
22:10:20.884 -> Water_Temp.:89.37, PWM:23
22:10:21.679 -> Water_Temp.:89.37, PWM:23
22:10:22.436 -> Water_Temp.:89.37, PWM:23
22:10:23.232 -> Water_Temp.:89.37, PWM:23
22:10:24.026 -> Water_Temp.:89.37, PWM:23
22:10:24.784 -> Water_Temp.:89.37, PWM:23
22:10:25.580 -> Water_Temp.:89.37, PWM:23
22:10:26.344 -> Water_Temp.:89.37, PWM:23
22:10:27.140 -> Water_Temp.:89.37, PWM:23
22:10:27.898 -> Water_Temp.:89.37, PWM:23
22:10:28.694 -> Water_Temp.:89.37, PWM:23
22:10:29.459 -> Water_Temp.:89.37, PWM:23
22:10:30.254 -> Water_Temp.:89.37, PWM:23
22:10:31.017 -> Water_Temp.:89.49, PWM:23
22:10:31.811 -> Water_Temp.:89.37, PWM:23
22:10:32.608 -> Water_Temp.:89.37, PWM:23
22:10:33.367 -> Water_Temp.:89.37, PWM:23
22:10:34.166 -> Water_Temp.:89.49, PWM:23
22:10:34.926 -> Water_Temp.:89.49, PWM:23
22:10:35.724 -> Water_Temp.:89.49, PWM:23
22:10:36.483 -> Water_Temp.:89.49, PWM:23
22:10:37.277 -> Water_Temp.:89.49, PWM:23
22:10:38.034 -> Water_Temp.:89.49, PWM:23
22:10:38.833 -> Water_Temp.:89.49, PWM:23
22:10:39.628 -> Water_Temp.:89.49, PWM:23
22:10:41.147 -> Water_Temp.:89.49, PWM:23
22:10:41.935 -> Water_Temp.:89.49, PWM:23
22:10:42.700 -> Water_Temp.:89.49, PWM:23
22:10:44.256 -> Water_Temp.:89.49, PWM:23
22:10:45.055 -> Water_Temp.:89.49, PWM:23
22:10:45.813 -> Water_Temp.:89.49, PWM:23
22:10:46.605 -> Water_Temp.:89.49, PWM:23
22:10:47.399 -> Water_Temp.:89.49, PWM:23
22:10:48.163 -> Water_Temp.:89.49, PWM:23
22:10:49.721 -> Water_Temp.:89.60, PWM:23
22:10:50.517 -> Water_Temp.:89.49, PWM:23
22:10:51.281 -> Water_Temp.:89.60, PWM:23
22:10:52.077 -> Water_Temp.:89.60, PWM:23
22:10:52.837 -> Water_Temp.:89.60, PWM:23
22:10:53.630 -> Water_Temp.:89.60, PWM:23
22:10:54.426 -> Water_Temp.:89.60, PWM:23
22:10:55.187 -> Water_Temp.:89.60, PWM:23
22:10:55.980 -> Water_Temp.:89.60, PWM:23
22:10:56.742 -> Water_Temp.:89.60, PWM:23
22:10:57.534 -> Water_Temp.:89.60, PWM:23
22:10:58.295 -> Water_Temp.:89.60, PWM:23
22:10:59.089 -> Water_Temp.:89.60, PWM:23
22:10:59.886 -> Water_Temp.:89.60, PWM:23
22:11:00.651 -> Water_Temp.:89.60, PWM:23
22:11:01.413 -> Water_Temp.:89.60, PWM:23
22:11:02.203 -> Water_Temp.:89.60, PWM:23
22:11:03.000 -> Water_Temp.:89.60, PWM:23
22:11:03.761 -> Water_Temp.:89.60, PWM:23
22:11:05.316 -> Water_Temp.:89.71, PWM:23
22:11:06.109 -> Water_Temp.:89.71, PWM:23
22:11:06.867 -> Water_Temp.:89.71, PWM:23
22:11:07.662 -> Water_Temp.:89.71, PWM:23
22:11:08.417 -> Water_Temp.:89.71, PWM:23
22:11:09.211 -> Water_Temp.:89.71, PWM:23
22:11:10.006 -> Water_Temp.:89.71, PWM:23
22:11:10.773 -> Water_Temp.:89.71, PWM:23
22:11:11.566 -> Water_Temp.:89.71, PWM:23
22:11:12.330 -> Water_Temp.:89.82, PWM:23
22:11:13.127 -> Water_Temp.:89.82, PWM:23
22:11:13.886 -> Water_Temp.:89.82, PWM:23
22:11:14.683 -> Water_Temp.:89.82, PWM:23
22:11:15.447 -> Water_Temp.:89.82, PWM:23
22:11:16.242 -> Water_Temp.:89.82, PWM:23
22:11:17.000 -> Water_Temp.:89.82, PWM:23
22:11:17.797 -> Water_Temp.:89.82, PWM:23
22:11:18.557 -> Water_Temp.:89.82, PWM:23
22:11:19.356 -> Water_Temp.:89.82, PWM:23
22:11:20.118 -> Water_Temp.:89.82, PWM:23
22:11:20.916 -> Water_Temp.:89.82, PWM:23
22:11:21.681 -> Water_Temp.:89.82, PWM:23
22:11:22.470 -> Water_Temp.:89.82, PWM:23
22:11:23.263 -> Water_Temp.:89.82, PWM:23
22:11:24.025 -> Water_Temp.:89.82, PWM:23
22:11:24.823 -> Water_Temp.:89.82, PWM:23
22:11:25.584 -> Water_Temp.:89.82, PWM:23
22:11:26.378 -> Water_Temp.:89.82, PWM:23
22:11:27.139 -> Water_Temp.:89.94, PWM:23
22:11:27.933 -> Water_Temp.:89.94, PWM:23
22:11:28.729 -> Water_Temp.:89.94, PWM:23
22:11:29.491 -> Water_Temp.:89.94, PWM:23
22:11:30.285 -> Water_Temp.:89.94, PWM:23
22:11:31.047 -> Water_Temp.:89.94, PWM:23
22:11:31.842 -> Water_Temp.:89.94, PWM:23
22:11:32.600 -> Water_Temp.:89.94, PWM:23
22:11:33.393 -> Water_Temp.:89.94, PWM:23
22:11:34.187 -> Water_Temp.:89.94, PWM:23
22:11:34.945 -> Water_Temp.:89.94, PWM:23
22:11:35.737 -> Water_Temp.:89.94, PWM:23
22:11:36.499 -> Water_Temp.:89.94, PWM:23
22:11:37.293 -> Water_Temp.:89.94, PWM:23
22:11:38.846 -> Water_Temp.:89.94, PWM:23
22:11:39.639 -> Water_Temp.:89.94, PWM:23
22:11:40.405 -> Water_Temp.:89.94, PWM:23
22:11:41.200 -> Water_Temp.:89.94, PWM:23
22:11:41.956 -> Water_Temp.:89.94, PWM:23
22:11:42.749 -> Water_Temp.:89.94, PWM:23
22:11:43.546 -> Water_Temp.:89.94, PWM:23
22:11:44.308 -> Water_Temp.:89.94, PWM:23
22:11:45.104 -> Water_Temp.:90.05, PWM:23
22:11:45.865 -> Water_Temp.:89.94, PWM:23
22:11:46.661 -> Water_Temp.:90.05, PWM:23
22:11:47.423 -> Water_Temp.:90.05, PWM:23
22:11:48.217 -> Water_Temp.:90.05, PWM:23
22:11:48.979 -> Water_Temp.:90.05, PWM:23
22:11:51.334 -> Water_Temp.:90.05, PWM:23
22:11:52.094 -> Water_Temp.:90.05, PWM:23
22:11:52.888 -> Water_Temp.:90.05, PWM:23
22:11:53.646 -> Water_Temp.:90.05, PWM:23
22:11:54.438 -> Water_Temp.:90.05, PWM:23
22:11:55.199 -> Water_Temp.:90.05, PWM:23
22:11:55.993 -> Water_Temp.:90.05, PWM:23
22:11:56.788 -> Water_Temp.:90.05, PWM:23
22:11:57.547 -> Water_Temp.:90.05, PWM:23
22:11:58.342 -> Water_Temp.:90.05, PWM:23
22:11:59.898 -> Water_Temp.:90.16, PWM:23
22:12:00.659 -> Water_Temp.:90.16, PWM:23
22:12:02.215 -> Water_Temp.:90.16, PWM:23
22:12:03.011 -> Water_Temp.:90.16, PWM:23
22:12:03.764 -> Water_Temp.:90.16, PWM:23
22:12:04.561 -> Water_Temp.:90.16, PWM:23
22:12:05.357 -> Water_Temp.:90.16, PWM:23
22:12:06.117 -> Water_Temp.:90.16, PWM:23
22:12:06.912 -> Water_Temp.:90.16, PWM:23
22:12:07.669 -> Water_Temp.:90.16, PWM:23
22:12:08.465 -> Water_Temp.:90.16, PWM:23
22:12:09.224 -> Water_Temp.:90.16, PWM:23
22:12:10.018 -> Water_Temp.:90.27, PWM:23
22:12:10.815 -> Water_Temp.:90.16, PWM:23
22:12:11.578 -> Water_Temp.:90.27, PWM:23
22:12:12.369 -> Water_Temp.:90.27, PWM:23
22:12:13.129 -> Water_Temp.:90.27, PWM:23
22:12:13.923 -> Water_Temp.:90.27, PWM:23
22:12:14.716 -> Water_Temp.:90.27, PWM:23
22:12:15.476 -> Water_Temp.:90.27, PWM:23
22:12:16.272 -> Water_Temp.:90.27, PWM:23
22:12:17.030 -> Water_Temp.:90.27, PWM:23
22:12:17.824 -> Water_Temp.:90.27, PWM:23
22:12:18.616 -> Water_Temp.:90.27, PWM:23
22:12:19.373 -> Water_Temp.:90.27, PWM:23
22:12:20.168 -> Water_Temp.:90.27, PWM:23
22:12:20.933 -> Water_Temp.:90.27, PWM:23
22:12:21.727 -> Water_Temp.:90.27, PWM:23
22:12:22.486 -> Water_Temp.:90.39, PWM:23
22:12:23.282 -> Water_Temp.:90.39, PWM:23
22:12:24.044 -> Water_Temp.:90.39, PWM:23
22:12:24.838 -> Water_Temp.:90.39, PWM:23
22:12:25.630 -> Water_Temp.:90.39, PWM:23
22:12:26.386 -> Water_Temp.:90.39, PWM:23
22:12:27.181 -> Water_Temp.:90.39, PWM:23
22:12:27.941 -> Water_Temp.:90.39, PWM:23
22:12:28.730 -> Water_Temp.:90.39, PWM:23
22:12:29.526 -> Water_Temp.:90.39, PWM:23
22:12:30.287 -> Water_Temp.:90.39, PWM:23
22:12:31.083 -> Water_Temp.:90.39, PWM:23
22:12:31.873 -> Water_Temp.:90.39, PWM:23
22:12:32.633 -> Water_Temp.:90.39, PWM:23
22:12:34.191 -> Water_Temp.:90.39, PWM:23
22:12:34.949 -> Water_Temp.:90.39, PWM:23
22:12:35.749 -> Water_Temp.:90.39, PWM:23
22:12:36.509 -> Water_Temp.:90.50, PWM:23
22:12:37.303 -> Water_Temp.:90.50, PWM:23
22:12:38.095 -> Water_Temp.:90.39, PWM:23
22:12:38.858 -> Water_Temp.:90.39, PWM:23
22:12:39.654 -> Water_Temp.:90.50, PWM:23
22:12:40.414 -> Water_Temp.:90.50, PWM:23
22:12:41.211 -> Water_Temp.:90.50, PWM:23
22:12:41.970 -> Water_Temp.:90.50, PWM:23
22:12:42.766 -> Water_Temp.:90.50, PWM:23
22:12:43.529 -> Water_Temp.:90.50, PWM:23
22:12:44.324 -> Water_Temp.:90.50, PWM:23
22:12:45.089 -> Water_Temp.:90.50, PWM:23
22:12:45.886 -> Water_Temp.:90.50, PWM:23
22:12:46.679 -> Water_Temp.:90.50, PWM:23
22:12:47.437 -> Water_Temp.:90.50, PWM:23
22:12:48.234 -> Water_Temp.:90.50, PWM:23
22:12:48.991 -> Water_Temp.:90.50, PWM:23
22:12:49.783 -> Water_Temp.:90.50, PWM:23
22:12:50.576 -> Water_Temp.:90.50, PWM:23
22:12:52.884 -> Water_Temp.:90.50, PWM:23
22:12:53.641 -> Water_Temp.:90.50, PWM:23
22:12:54.439 -> Water_Temp.:90.61, PWM:23
22:12:55.198 -> Water_Temp.:90.61, PWM:23
22:12:55.994 -> Water_Temp.:90.61, PWM:23
22:12:56.789 -> Water_Temp.:90.61, PWM:23
22:12:57.549 -> Water_Temp.:90.61, PWM:23
22:12:58.342 -> Water_Temp.:90.61, PWM:23
22:12:59.100 -> Water_Temp.:90.61, PWM:23
22:12:59.895 -> Water_Temp.:90.61, PWM:23
22:13:01.451 -> Water_Temp.:90.72, PWM:23
22:13:02.210 -> Water_Temp.:90.72, PWM:23
22:13:03.002 -> Water_Temp.:90.72, PWM:23
22:13:03.795 -> Water_Temp.:90.72, PWM:23
22:13:04.557 -> Water_Temp.:90.72, PWM:23
22:13:05.353 -> Water_Temp.:90.72, PWM:23
22:13:06.109 -> Water_Temp.:90.72, PWM:23
22:13:06.898 -> Water_Temp.:90.72, PWM:23
22:13:07.661 -> Water_Temp.:90.72, PWM:23
22:13:08.454 -> Water_Temp.:90.72, PWM:23
22:13:09.253 -> Water_Temp.:90.72, PWM:23
22:13:10.014 -> Water_Temp.:90.72, PWM:23
22:13:10.809 -> Water_Temp.:90.72, PWM:23
22:13:11.569 -> Water_Temp.:90.72, PWM:23
22:13:13.116 -> Water_Temp.:90.72, PWM:23
22:13:13.910 -> Water_Temp.:90.72, PWM:23
22:13:14.674 -> Water_Temp.:90.72, PWM:23
22:13:15.471 -> Water_Temp.:90.84, PWM:23
22:13:16.229 -> Water_Temp.:90.84, PWM:23
22:13:17.016 -> Water_Temp.:90.84, PWM:23
22:13:17.809 -> Water_Temp.:90.84, PWM:23
22:13:18.568 -> Water_Temp.:90.84, PWM:23
22:13:19.359 -> Water_Temp.:90.84, PWM:23
22:13:20.152 -> Water_Temp.:90.84, PWM:23
22:13:20.910 -> Water_Temp.:90.84, PWM:23
22:13:21.702 -> Water_Temp.:90.84, PWM:23
22:13:22.496 -> Water_Temp.:90.84, PWM:23
22:13:23.255 -> Water_Temp.:90.84, PWM:23
22:13:24.047 -> Water_Temp.:90.84, PWM:23
22:13:24.813 -> Water_Temp.:90.84, PWM:23
22:13:25.604 -> Water_Temp.:90.84, PWM:23
22:13:26.363 -> Water_Temp.:90.95, PWM:23
22:13:27.153 -> Water_Temp.:90.95, PWM:23
22:13:27.947 -> Water_Temp.:90.95, PWM:23
22:13:28.707 -> Water_Temp.:90.95, PWM:23
22:13:29.507 -> Water_Temp.:90.95, PWM:23
22:13:30.268 -> Water_Temp.:90.95, PWM:23
22:13:31.819 -> Water_Temp.:90.95, PWM:23
22:13:32.614 -> Water_Temp.:90.95, PWM:23
22:13:33.374 -> Water_Temp.:90.95, PWM:23
22:13:34.168 -> Water_Temp.:90.95, PWM:23
22:13:35.691 -> Water_Temp.:90.95, PWM:23
22:13:36.486 -> Water_Temp.:90.95, PWM:23
22:13:37.246 -> Water_Temp.:90.95, PWM:23
22:13:38.034 -> Water_Temp.:90.95, PWM:23
22:13:38.827 -> Water_Temp.:90.95, PWM:23
22:13:39.588 -> Water_Temp.:90.95, PWM:23
22:13:40.380 -> Water_Temp.:90.95, PWM:23
22:13:41.174 -> Water_Temp.:90.95, PWM:23
22:13:41.933 -> Water_Temp.:90.95, PWM:23
22:13:42.726 -> Water_Temp.:90.95, PWM:23
22:13:43.514 -> Water_Temp.:90.95, PWM:23
22:13:44.275 -> Water_Temp.:90.95, PWM:23
22:13:45.070 -> Water_Temp.:90.95, PWM:23
22:13:45.827 -> Water_Temp.:90.95, PWM:23
22:13:46.621 -> Water_Temp.:90.95, PWM:23
22:13:47.383 -> Water_Temp.:90.95, PWM:23
22:13:48.174 -> Water_Temp.:90.95, PWM:23
22:13:48.965 -> Water_Temp.:91.06, PWM:23
22:13:49.728 -> Water_Temp.:91.06, PWM:23
22:13:50.522 -> Water_Temp.:91.06, PWM:23
22:13:51.315 -> Water_Temp.:91.06, PWM:23
22:13:52.072 -> Water_Temp.:91.06, PWM:23
22:13:52.868 -> Water_Temp.:91.06, PWM:23
22:13:53.627 -> Water_Temp.:91.06, PWM:23
22:13:54.420 -> Water_Temp.:91.06, PWM:23
22:13:55.182 -> Water_Temp.:91.06, PWM:23
22:13:55.974 -> Water_Temp.:91.06, PWM:23
22:13:56.769 -> Water_Temp.:91.06, PWM:23
22:13:57.525 -> Water_Temp.:91.06, PWM:23
22:13:58.314 -> Water_Temp.:91.06, PWM:23
22:13:59.103 -> Water_Temp.:91.06, PWM:23
22:13:59.866 -> Water_Temp.:91.06, PWM:23
22:14:00.665 -> Water_Temp.:91.18, PWM:23
22:14:01.426 -> Water_Temp.:91.18, PWM:23
22:14:02.218 -> Water_Temp.:91.18, PWM:23
22:14:03.012 -> Water_Temp.:91.18, PWM:23
22:14:03.774 -> Water_Temp.:91.18, PWM:23
22:14:04.572 -> Water_Temp.:91.18, PWM:23
22:14:05.332 -> Water_Temp.:91.18, PWM:23
22:14:06.128 -> Water_Temp.:91.18, PWM:23
22:14:08.409 -> Water_Temp.:91.18, PWM:23
22:14:09.962 -> Water_Temp.:91.18, PWM:23
22:14:10.754 -> Water_Temp.:91.29, PWM:23
22:14:11.547 -> Water_Temp.:91.29, PWM:23
22:14:12.305 -> Water_Temp.:91.29, PWM:23
22:14:13.100 -> Water_Temp.:91.29, PWM:23
22:14:13.861 -> Water_Temp.:91.29, PWM:23
22:14:14.660 -> Water_Temp.:91.29, PWM:23
22:14:15.422 -> Water_Temp.:91.29, PWM:23
22:14:16.215 -> Water_Temp.:91.29, PWM:23
22:14:17.004 -> Water_Temp.:91.29, PWM:23
22:14:17.761 -> Water_Temp.:91.29, PWM:23
22:14:18.556 -> Water_Temp.:91.29, PWM:23
22:14:19.318 -> Water_Temp.:91.29, PWM:23
22:14:20.114 -> Water_Temp.:91.40, PWM:23
22:14:20.875 -> Water_Temp.:91.40, PWM:23
22:14:21.667 -> Water_Temp.:91.40, PWM:23
22:14:22.462 -> Water_Temp.:91.40, PWM:23
22:14:23.223 -> Water_Temp.:91.40, PWM:23
22:14:24.019 -> Water_Temp.:91.40, PWM:23
22:14:24.780 -> Water_Temp.:91.40, PWM:23
22:14:25.577 -> Water_Temp.:91.40, PWM:23
22:14:26.335 -> Water_Temp.:91.40, PWM:23
22:14:27.128 -> Water_Temp.:91.40, PWM:23
22:14:27.922 -> Water_Temp.:91.40, PWM:23
22:14:28.682 -> Water_Temp.:91.40, PWM:23
22:14:29.476 -> Water_Temp.:91.40, PWM:23
22:14:30.234 -> Water_Temp.:91.40, PWM:23
22:14:31.028 -> Water_Temp.:91.40, PWM:23
22:14:31.825 -> Water_Temp.:91.40, PWM:23
22:14:32.585 -> Water_Temp.:91.40, PWM:23
22:14:33.379 -> Water_Temp.:91.40, PWM:23
22:14:34.142 -> Water_Temp.:91.40, PWM:23
22:14:34.939 -> Water_Temp.:91.40, PWM:23
22:14:35.700 -> Water_Temp.:91.40, PWM:23
22:14:36.493 -> Water_Temp.:91.40, PWM:23
22:14:37.255 -> Water_Temp.:91.40, PWM:23
22:14:38.051 -> Water_Temp.:91.40, PWM:23
22:14:38.845 -> Water_Temp.:91.40, PWM:23
22:14:39.601 -> Water_Temp.:91.40, PWM:23
22:14:40.397 -> Water_Temp.:91.40, PWM:23
22:14:41.163 -> Water_Temp.:91.40, PWM:23
22:14:41.957 -> Water_Temp.:91.40, PWM:23
22:14:42.717 -> Water_Temp.:91.40, PWM:23
22:14:43.512 -> Water_Temp.:91.40, PWM:23
22:14:45.031 -> Water_Temp.:91.51, PWM:23
22:14:45.823 -> Water_Temp.:91.51, PWM:23
22:14:46.616 -> Water_Temp.:91.40, PWM:23
22:14:47.379 -> Water_Temp.:91.51, PWM:23
22:14:48.930 -> Water_Temp.:91.51, PWM:23
22:14:49.727 -> Water_Temp.:91.51, PWM:23
22:14:50.489 -> Water_Temp.:91.51, PWM:23
22:14:51.285 -> Water_Temp.:91.51, PWM:23
22:14:52.048 -> Water_Temp.:91.51, PWM:23
22:14:52.842 -> Water_Temp.:91.51, PWM:23
22:14:53.601 -> Water_Temp.:91.51, PWM:23
22:14:54.399 -> Water_Temp.:91.51, PWM:23
22:14:55.162 -> Water_Temp.:91.51, PWM:23
22:14:55.959 -> Water_Temp.:91.51, PWM:23
22:14:56.747 -> Water_Temp.:91.62, PWM:23
22:14:58.301 -> Water_Temp.:91.62, PWM:23
22:14:59.064 -> Water_Temp.:91.62, PWM:23
22:14:59.861 -> Water_Temp.:91.62, PWM:23
22:15:00.624 -> Water_Temp.:91.62, PWM:23
22:15:01.424 -> Water_Temp.:91.62, PWM:23
22:15:02.183 -> Water_Temp.:91.62, PWM:23
22:15:02.977 -> Water_Temp.:91.62, PWM:23
22:15:03.736 -> Water_Temp.:91.62, PWM:23
22:15:04.530 -> Water_Temp.:91.62, PWM:23
22:15:05.324 -> Water_Temp.:91.62, PWM:23
22:15:06.085 -> Water_Temp.:91.62, PWM:23
22:15:06.881 -> Water_Temp.:91.62, PWM:23
22:15:07.640 -> Water_Temp.:91.62, PWM:23
22:15:08.439 -> Water_Temp.:91.62, PWM:23
22:15:09.198 -> Water_Temp.:91.62, PWM:23
22:15:09.992 -> Water_Temp.:91.62, PWM:23
22:15:10.787 -> Water_Temp.:91.62, PWM:23
22:15:11.548 -> Water_Temp.:91.62, PWM:23
22:15:12.345 -> Water_Temp.:91.62, PWM:23
22:15:13.103 -> Water_Temp.:91.62, PWM:23
22:15:13.895 -> Water_Temp.:91.62, PWM:23
22:15:14.660 -> Water_Temp.:91.62, PWM:23
22:15:15.456 -> Water_Temp.:91.62, PWM:23
22:15:16.218 -> Water_Temp.:91.62, PWM:23
22:15:17.014 -> Water_Temp.:91.62, PWM:23
22:15:17.807 -> Water_Temp.:91.74, PWM:23
22:15:18.570 -> Water_Temp.:91.74, PWM:23
22:15:19.362 -> Water_Temp.:91.74, PWM:23
22:15:20.122 -> Water_Temp.:91.74, PWM:23
22:15:20.919 -> Water_Temp.:91.74, PWM:23
22:15:22.437 -> Water_Temp.:91.74, PWM:23
22:15:23.232 -> Water_Temp.:91.74, PWM:23
22:15:23.990 -> Water_Temp.:91.74, PWM:23
22:15:24.787 -> Water_Temp.:91.74, PWM:23
22:15:25.551 -> Water_Temp.:91.74, PWM:23
22:15:26.346 -> Water_Temp.:91.74, PWM:23
22:15:27.136 -> Water_Temp.:93.43, PWM:23
22:15:27.896 -> Water_Temp.:95.45, PWM:23
22:15:28.694 -> Water_Temp.:93.09, PWM:23
22:15:29.454 -> Water_Temp.:92.64, PWM:23
22:15:30.248 -> Water_Temp.:92.41, PWM:23
22:15:31.043 -> Water_Temp.:92.30, PWM:23
22:15:31.800 -> Water_Temp.:92.19, PWM:23
22:15:32.594 -> Water_Temp.:92.19, PWM:23
22:15:33.352 -> Water_Temp.:92.07, PWM:23
22:15:34.148 -> Water_Temp.:92.07, PWM:23
22:15:34.943 -> Water_Temp.:92.07, PWM:23
22:15:35.702 -> Water_Temp.:91.96, PWM:23
22:15:36.495 -> Water_Temp.:91.96, PWM:23
22:15:37.257 -> Water_Temp.:91.96, PWM:23
22:15:38.809 -> Water_Temp.:91.96, PWM:23
22:15:39.567 -> Water_Temp.:91.96, PWM:23
22:15:40.362 -> Water_Temp.:91.96, PWM:23
22:15:41.156 -> Water_Temp.:91.96, PWM:23
22:15:41.914 -> Water_Temp.:91.96, PWM:23
22:15:43.464 -> Water_Temp.:91.96, PWM:23
22:15:44.224 -> Water_Temp.:91.96, PWM:23
22:15:45.020 -> Water_Temp.:91.96, PWM:23
22:15:45.817 -> Water_Temp.:91.96, PWM:23
22:15:46.574 -> Water_Temp.:91.96, PWM:23
22:15:47.368 -> Water_Temp.:91.96, PWM:23
22:15:48.128 -> Water_Temp.:91.96, PWM:23
22:15:48.922 -> Water_Temp.:91.96, PWM:23
22:15:49.688 -> Water_Temp.:91.96, PWM:23
22:15:50.481 -> Water_Temp.:91.96, PWM:23
22:15:51.244 -> Water_Temp.:91.96, PWM:23
22:15:52.036 -> Water_Temp.:91.96, PWM:23
22:15:52.830 -> Water_Temp.:91.96, PWM:23
22:15:53.589 -> Water_Temp.:91.96, PWM:23
22:15:54.385 -> Water_Temp.:91.96, PWM:23
22:15:55.148 -> Water_Temp.:92.07, PWM:23
22:15:55.939 -> Water_Temp.:92.07, PWM:23
22:15:56.734 -> Water_Temp.:92.07, PWM:23
22:15:57.492 -> Water_Temp.:92.07, PWM:23
22:15:58.289 -> Water_Temp.:92.07, PWM:23
22:15:59.047 -> Water_Temp.:92.07, PWM:23
22:15:59.839 -> Water_Temp.:92.07, PWM:23
22:16:00.602 -> Water_Temp.:92.07, PWM:23
22:16:01.398 -> Water_Temp.:92.07, PWM:23
22:16:02.194 -> Water_Temp.:92.19, PWM:23
22:16:02.956 -> Water_Temp.:92.19, PWM:27
22:16:03.753 -> Water_Temp.:92.19, PWM:28
22:16:04.510 -> Water_Temp.:92.19, PWM:28
22:16:05.304 -> Water_Temp.:92.19, PWM:29
22:16:06.064 -> Water_Temp.:92.19, PWM:28
22:16:06.857 -> Water_Temp.:92.19, PWM:29
22:16:08.407 -> Water_Temp.:92.19, PWM:29
22:16:09.199 -> Water_Temp.:92.19, PWM:29
22:16:10.756 -> Water_Temp.:92.19, PWM:28
22:16:11.518 -> Water_Temp.:92.19, PWM:29
22:16:12.312 -> Water_Temp.:92.19, PWM:29
22:16:13.107 -> Water_Temp.:92.19, PWM:29
22:16:13.867 -> Water_Temp.:92.19, PWM:29
22:16:14.664 -> Water_Temp.:92.19, PWM:28
22:16:15.424 -> Water_Temp.:92.19, PWM:29
22:16:16.218 -> Water_Temp.:92.19, PWM:29
22:16:17.769 -> Water_Temp.:92.19, PWM:28
22:16:18.528 -> Water_Temp.:92.19, PWM:29
22:16:19.324 -> Water_Temp.:92.19, PWM:29
22:16:20.082 -> Water_Temp.:92.19, PWM:28
22:16:20.880 -> Water_Temp.:92.19, PWM:29
22:16:21.673 -> Water_Temp.:92.19, PWM:29
22:16:22.433 -> Water_Temp.:92.19, PWM:28
22:16:23.225 -> Water_Temp.:92.19, PWM:29
22:16:23.983 -> Water_Temp.:92.19, PWM:28
22:16:24.779 -> Water_Temp.:92.19, PWM:29
22:16:25.541 -> Water_Temp.:92.19, PWM:28
22:16:26.334 -> Water_Temp.:92.19, PWM:28
22:16:27.126 -> Water_Temp.:92.30, PWM:29
22:16:27.888 -> Water_Temp.:92.30, PWM:28
22:16:28.682 -> Water_Temp.:92.30, PWM:28
22:16:29.442 -> Water_Temp.:92.30, PWM:28
22:16:30.239 -> Water_Temp.:92.30, PWM:29
22:16:31.005 -> Water_Temp.:92.30, PWM:29
22:16:31.797 -> Water_Temp.:92.30, PWM:29
22:16:32.594 -> Water_Temp.:92.30, PWM:29
22:16:33.358 -> Water_Temp.:92.30, PWM:29
22:16:34.153 -> Water_Temp.:92.30, PWM:29
22:16:34.917 -> Water_Temp.:92.30, PWM:29
22:16:36.478 -> Water_Temp.:92.30, PWM:28
22:16:37.272 -> Water_Temp.:92.30, PWM:28
22:16:38.825 -> Water_Temp.:92.30, PWM:29
22:16:39.587 -> Water_Temp.:92.30, PWM:28
22:16:40.380 -> Water_Temp.:92.30, PWM:29
22:16:41.142 -> Water_Temp.:92.30, PWM:29
22:16:41.934 -> Water_Temp.:92.30, PWM:29
22:16:42.730 -> Water_Temp.:92.30, PWM:29
22:16:43.492 -> Water_Temp.:92.30, PWM:28
22:16:44.288 -> Water_Temp.:92.30, PWM:29
22:16:45.053 -> Water_Temp.:92.30, PWM:29
22:16:45.849 -> Water_Temp.:92.30, PWM:29
22:16:46.609 -> Water_Temp.:92.41, PWM:29
22:16:47.406 -> Water_Temp.:92.41, PWM:28
22:16:48.166 -> Water_Temp.:92.41, PWM:28
22:16:48.962 -> Water_Temp.:92.41, PWM:28
22:16:49.723 -> Water_Temp.:92.41, PWM:29
22:16:50.518 -> Water_Temp.:92.41, PWM:28
22:16:51.283 -> Water_Temp.:92.41, PWM:28
22:16:52.075 -> Water_Temp.:92.41, PWM:29
22:16:52.866 -> Water_Temp.:92.41, PWM:29
22:16:53.627 -> Water_Temp.:92.41, PWM:29
22:16:54.423 -> Water_Temp.:92.41, PWM:29
22:16:55.977 -> Water_Temp.:92.41, PWM:28
22:16:56.770 -> Water_Temp.:92.41, PWM:29
22:16:57.531 -> Water_Temp.:92.41, PWM:28
22:16:58.296 -> Water_Temp.:92.41, PWM:29
22:16:59.093 -> Water_Temp.:92.41, PWM:29
22:16:59.887 -> Water_Temp.:92.41, PWM:28
22:17:00.651 -> Water_Temp.:92.41, PWM:28
22:17:01.450 -> Water_Temp.:92.41, PWM:28
22:17:02.212 -> Water_Temp.:92.52, PWM:28
22:17:03.006 -> Water_Temp.:92.41, PWM:29
22:17:03.767 -> Water_Temp.:92.52, PWM:29
22:17:04.565 -> Water_Temp.:92.52, PWM:28
22:17:05.321 -> Water_Temp.:92.52, PWM:29
22:17:06.121 -> Water_Temp.:92.52, PWM:29
22:17:07.647 -> Water_Temp.:92.52, PWM:28
22:17:08.446 -> Water_Temp.:92.52, PWM:29
22:17:09.209 -> Water_Temp.:92.52, PWM:29
22:17:10.006 -> Water_Temp.:92.64, PWM:29
22:17:10.768 -> Water_Temp.:92.64, PWM:28
22:17:11.561 -> Water_Temp.:92.64, PWM:29
22:17:12.323 -> Water_Temp.:92.64, PWM:29
22:17:13.118 -> Water_Temp.:92.64, PWM:28
22:17:13.880 -> Water_Temp.:92.64, PWM:29
22:17:14.677 -> Water_Temp.:92.64, PWM:29
22:17:15.437 -> Water_Temp.:92.64, PWM:28
22:17:16.228 -> Water_Temp.:92.64, PWM:29
22:17:17.019 -> Water_Temp.:92.64, PWM:28
22:17:17.775 -> Water_Temp.:92.64, PWM:29
22:17:18.572 -> Water_Temp.:92.64, PWM:29
22:17:19.335 -> Water_Temp.:92.64, PWM:29
22:17:20.129 -> Water_Temp.:92.64, PWM:29
22:17:20.894 -> Water_Temp.:92.64, PWM:29
22:17:21.688 -> Water_Temp.:92.64, PWM:28
22:17:22.482 -> Water_Temp.:92.64, PWM:29
22:17:23.244 -> Water_Temp.:92.64, PWM:28
22:17:24.038 -> Water_Temp.:92.64, PWM:28
22:17:25.557 -> Water_Temp.:92.64, PWM:28
22:17:26.351 -> Water_Temp.:92.64, PWM:29
22:17:27.142 -> Water_Temp.:92.64, PWM:28
22:17:27.901 -> Water_Temp.:92.64, PWM:29
22:17:28.694 -> Water_Temp.:92.64, PWM:28
22:17:29.489 -> Water_Temp.:92.64, PWM:28
22:17:30.251 -> Water_Temp.:92.64, PWM:29
22:17:31.044 -> Water_Temp.:92.75, PWM:28
22:17:31.804 -> Water_Temp.:92.75, PWM:28
22:17:32.601 -> Water_Temp.:92.75, PWM:29
22:17:33.361 -> Water_Temp.:92.75, PWM:28
22:17:34.156 -> Water_Temp.:92.75, PWM:29
22:17:34.918 -> Water_Temp.:92.75, PWM:28
22:17:35.712 -> Water_Temp.:92.75, PWM:29
22:17:36.508 -> Water_Temp.:92.75, PWM:28
22:17:37.265 -> Water_Temp.:92.75, PWM:29
22:17:38.058 -> Water_Temp.:92.75, PWM:29
22:17:38.816 -> Water_Temp.:92.86, PWM:29
22:17:39.609 -> Water_Temp.:92.86, PWM:28
22:17:40.403 -> Water_Temp.:92.86, PWM:29
22:17:41.166 -> Water_Temp.:92.86, PWM:29
22:17:41.959 -> Water_Temp.:92.86, PWM:28
22:17:42.720 -> Water_Temp.:92.86, PWM:28
22:17:44.276 -> Water_Temp.:92.86, PWM:29
22:17:45.074 -> Water_Temp.:92.86, PWM:29
22:17:45.866 -> Water_Temp.:92.86, PWM:28
22:17:46.626 -> Water_Temp.:92.86, PWM:28
22:17:47.420 -> Water_Temp.:92.86, PWM:29
22:17:48.183 -> Water_Temp.:92.86, PWM:28
22:17:49.740 -> Water_Temp.:92.86, PWM:29
22:17:51.294 -> Water_Temp.:92.86, PWM:29
22:17:52.053 -> Water_Temp.:92.86, PWM:29
22:17:52.849 -> Water_Temp.:92.86, PWM:29
22:17:53.608 -> Water_Temp.:92.97, PWM:29
22:17:54.402 -> Water_Temp.:92.97, PWM:28
22:17:55.198 -> Water_Temp.:92.97, PWM:29
22:17:55.960 -> Water_Temp.:92.97, PWM:28
22:17:56.754 -> Water_Temp.:92.97, PWM:29
22:17:57.513 -> Water_Temp.:92.97, PWM:28
22:17:58.311 -> Water_Temp.:92.97, PWM:29
22:17:59.072 -> Water_Temp.:92.97, PWM:28
22:17:59.865 -> Water_Temp.:92.97, PWM:29
22:18:00.627 -> Water_Temp.:92.97, PWM:28
22:18:01.424 -> Water_Temp.:92.97, PWM:28
22:18:02.217 -> Water_Temp.:92.97, PWM:28
22:18:02.976 -> Water_Temp.:92.97, PWM:29
22:18:03.770 -> Water_Temp.:92.97, PWM:29
22:18:04.531 -> Water_Temp.:93.09, PWM:29
22:18:05.328 -> Water_Temp.:93.09, PWM:29
22:18:06.087 -> Water_Temp.:93.09, PWM:29
22:18:07.639 -> Water_Temp.:93.09, PWM:29
22:18:08.436 -> Water_Temp.:93.09, PWM:28
22:18:09.199 -> Water_Temp.:93.09, PWM:29
22:18:09.958 -> Water_Temp.:93.09, PWM:29
22:18:10.754 -> Water_Temp.:93.09, PWM:28
22:18:11.519 -> Water_Temp.:93.09, PWM:29
22:18:12.315 -> Water_Temp.:93.09, PWM:29
22:18:13.113 -> Water_Temp.:93.09, PWM:29
22:18:13.874 -> Water_Temp.:93.09, PWM:28
22:18:14.669 -> Water_Temp.:93.09, PWM:29
22:18:15.428 -> Water_Temp.:93.09, PWM:29
22:18:16.224 -> Water_Temp.:93.20, PWM:28
22:18:16.980 -> Water_Temp.:93.20, PWM:28
22:18:17.773 -> Water_Temp.:93.20, PWM:28
22:18:18.569 -> Water_Temp.:93.20, PWM:29
22:18:19.331 -> Water_Temp.:93.20, PWM:29
22:18:20.125 -> Water_Temp.:93.20, PWM:29
22:18:20.885 -> Water_Temp.:93.20, PWM:29
22:18:21.679 -> Water_Temp.:93.20, PWM:28
22:18:22.440 -> Water_Temp.:93.20, PWM:29
22:18:23.234 -> Water_Temp.:93.20, PWM:29
22:18:24.029 -> Water_Temp.:93.20, PWM:28
22:18:24.795 -> Water_Temp.:93.20, PWM:28
22:18:25.590 -> Water_Temp.:93.20, PWM:29
22:18:26.351 -> Water_Temp.:93.20, PWM:28
22:18:27.145 -> Water_Temp.:93.20, PWM:29
22:18:27.906 -> Water_Temp.:93.20, PWM:29
22:18:28.703 -> Water_Temp.:93.20, PWM:29
22:18:29.464 -> Water_Temp.:93.31, PWM:28
22:18:30.256 -> Water_Temp.:93.31, PWM:29
22:18:31.051 -> Water_Temp.:93.31, PWM:28
22:18:31.815 -> Water_Temp.:93.31, PWM:28
22:18:32.609 -> Water_Temp.:93.31, PWM:29
22:18:33.367 -> Water_Temp.:95.00, PWM:28
22:18:34.160 -> Water_Temp.:97.02, PWM:29
22:18:34.924 -> Water_Temp.:94.66, PWM:28
22:18:35.719 -> Water_Temp.:94.21, PWM:28
22:18:38.025 -> Water_Temp.:93.76, PWM:28
22:18:38.818 -> Water_Temp.:93.76, PWM:28
22:18:39.609 -> Water_Temp.:93.65, PWM:29
22:18:40.370 -> Water_Temp.:93.65, PWM:28
22:18:41.166 -> Water_Temp.:93.65, PWM:29
22:18:41.925 -> Water_Temp.:93.54, PWM:28
22:18:43.480 -> Water_Temp.:93.54, PWM:28
22:18:45.035 -> Water_Temp.:93.54, PWM:28
22:18:45.793 -> Water_Temp.:93.43, PWM:28
22:18:46.586 -> Water_Temp.:93.43, PWM:29
22:18:48.140 -> Water_Temp.:93.43, PWM:29
22:18:48.931 -> Water_Temp.:93.43, PWM:28
22:18:49.689 -> Water_Temp.:93.43, PWM:28
22:18:51.248 -> Water_Temp.:93.43, PWM:29
22:18:52.006 -> Water_Temp.:93.43, PWM:29
22:18:52.800 -> Water_Temp.:93.43, PWM:29
22:18:53.562 -> Water_Temp.:93.54, PWM:29
22:18:54.355 -> Water_Temp.:93.43, PWM:29
22:18:55.117 -> Water_Temp.:93.54, PWM:29
22:18:55.910 -> Water_Temp.:93.54, PWM:29
22:18:56.707 -> Water_Temp.:93.54, PWM:29
22:18:57.464 -> Water_Temp.:93.54, PWM:29
22:18:58.258 -> Water_Temp.:93.54, PWM:28
22:18:59.018 -> Water_Temp.:93.54, PWM:28
22:18:59.808 -> Water_Temp.:93.43, PWM:29
22:19:00.606 -> Water_Temp.:93.54, PWM:29
22:19:01.367 -> Water_Temp.:93.54, PWM:29
22:19:02.161 -> Water_Temp.:93.54, PWM:28
22:19:02.922 -> Water_Temp.:93.43, PWM:28
22:19:03.719 -> Water_Temp.:93.43, PWM:28
22:19:04.481 -> Water_Temp.:93.43, PWM:28
22:19:05.280 -> Water_Temp.:93.43, PWM:29
22:19:06.043 -> Water_Temp.:93.43, PWM:28
22:19:06.837 -> Water_Temp.:93.43, PWM:28
22:19:07.596 -> Water_Temp.:93.43, PWM:29
22:19:08.389 -> Water_Temp.:93.43, PWM:29
22:19:09.185 -> Water_Temp.:93.43, PWM:29
22:19:09.942 -> Water_Temp.:93.43, PWM:28
22:19:10.737 -> Water_Temp.:93.43, PWM:29
22:19:11.496 -> Water_Temp.:93.43, PWM:28
22:19:12.292 -> Water_Temp.:93.43, PWM:29
22:19:13.050 -> Water_Temp.:93.43, PWM:29
22:19:13.845 -> Water_Temp.:93.43, PWM:29
22:19:14.639 -> Water_Temp.:93.43, PWM:29
22:19:15.403 -> Water_Temp.:93.43, PWM:29
22:19:16.197 -> Water_Temp.:93.43, PWM:29
22:19:16.987 -> Water_Temp.:93.43, PWM:28
22:19:17.749 -> Water_Temp.:93.43, PWM:29
22:19:18.511 -> Water_Temp.:93.43, PWM:29
22:19:19.307 -> Water_Temp.:93.43, PWM:28
22:19:20.073 -> Water_Temp.:93.43, PWM:29
22:19:21.627 -> Water_Temp.:93.43, PWM:28
22:19:23.182 -> Water_Temp.:93.43, PWM:28
22:19:23.942 -> Water_Temp.:93.54, PWM:28
22:19:24.736 -> Water_Temp.:93.54, PWM:29
22:19:25.531 -> Water_Temp.:93.54, PWM:29
22:19:26.294 -> Water_Temp.:93.54, PWM:28
22:19:27.088 -> Water_Temp.:93.54, PWM:29
22:19:27.847 -> Water_Temp.:93.54, PWM:28
22:19:28.640 -> Water_Temp.:93.54, PWM:29
22:19:29.403 -> Water_Temp.:93.54, PWM:29
22:19:30.195 -> Water_Temp.:93.54, PWM:28
22:19:30.989 -> Water_Temp.:93.54, PWM:28
22:19:31.749 -> Water_Temp.:93.54, PWM:28
22:19:32.545 -> Water_Temp.:93.54, PWM:28
22:19:33.305 -> Water_Temp.:93.54, PWM:29
22:19:34.097 -> Water_Temp.:93.65, PWM:29
22:19:34.859 -> Water_Temp.:93.65, PWM:29
22:19:35.654 -> Water_Temp.:93.65, PWM:28
22:19:36.447 -> Water_Temp.:93.65, PWM:29
22:19:37.209 -> Water_Temp.:93.65, PWM:28
22:19:38.002 -> Water_Temp.:93.65, PWM:28
22:19:38.762 -> Water_Temp.:93.65, PWM:28
22:19:39.561 -> Water_Temp.:93.65, PWM:29
22:19:40.321 -> Water_Temp.:93.76, PWM:28
22:19:41.116 -> Water_Temp.:93.76, PWM:29
22:19:41.879 -> Water_Temp.:93.76, PWM:29
22:19:42.673 -> Water_Temp.:93.76, PWM:29
22:19:43.467 -> Water_Temp.:93.76, PWM:29
22:19:44.225 -> Water_Temp.:93.76, PWM:28
22:19:45.019 -> Water_Temp.:93.76, PWM:29
22:19:45.776 -> Water_Temp.:93.76, PWM:29
22:19:46.568 -> Water_Temp.:93.76, PWM:28
22:19:47.366 -> Water_Temp.:93.87, PWM:28
22:19:48.124 -> Water_Temp.:93.76, PWM:29
22:19:48.917 -> Water_Temp.:93.76, PWM:28
22:19:49.679 -> Water_Temp.:93.76, PWM:28
22:19:52.001 -> Water_Temp.:93.76, PWM:28
22:19:52.798 -> Water_Temp.:93.76, PWM:29
22:19:53.562 -> Water_Temp.:93.76, PWM:29
22:19:54.360 -> Water_Temp.:93.76, PWM:28
22:19:55.119 -> Water_Temp.:93.76, PWM:29
22:19:55.911 -> Water_Temp.:93.76, PWM:28
22:19:56.707 -> Water_Temp.:93.76, PWM:29
22:19:57.464 -> Water_Temp.:93.76, PWM:29
22:19:58.257 -> Water_Temp.:93.76, PWM:28
22:19:59.046 -> Water_Temp.:93.76, PWM:28
22:19:59.806 -> Water_Temp.:93.76, PWM:28
22:20:00.606 -> Water_Temp.:93.76, PWM:28
22:20:01.366 -> Water_Temp.:93.76, PWM:28
22:20:02.160 -> Water_Temp.:93.76, PWM:29
22:20:02.918 -> Water_Temp.:93.76, PWM:29
22:20:03.715 -> Water_Temp.:93.76, PWM:29
22:20:04.506 -> Water_Temp.:93.76, PWM:29
22:20:05.268 -> Water_Temp.:93.76, PWM:29
22:20:06.060 -> Water_Temp.:93.76, PWM:29
22:20:06.846 -> Water_Temp.:93.76, PWM:29
22:20:08.363 -> Water_Temp.:93.76, PWM:28
22:20:09.157 -> Water_Temp.:93.76, PWM:29
22:20:09.921 -> Water_Temp.:93.87, PWM:29
22:20:10.718 -> Water_Temp.:93.87, PWM:28
22:20:11.510 -> Water_Temp.:93.87, PWM:29
22:20:12.266 -> Water_Temp.:93.87, PWM:29
22:20:13.061 -> Water_Temp.:93.87, PWM:29
22:20:13.818 -> Water_Temp.:93.87, PWM:28
22:20:14.611 -> Water_Temp.:93.87, PWM:28
22:20:15.410 -> Water_Temp.:93.87, PWM:28
22:20:16.171 -> Water_Temp.:93.87, PWM:28
22:20:16.967 -> Water_Temp.:93.87, PWM:29
22:20:17.727 -> Water_Temp.:93.87, PWM:29
22:20:18.521 -> Water_Temp.:93.87, PWM:29
22:20:19.280 -> Water_Temp.:93.87, PWM:29
22:20:20.077 -> Water_Temp.:93.87, PWM:29
22:20:20.837 -> Water_Temp.:93.87, PWM:29
22:20:21.633 -> Water_Temp.:93.87, PWM:29
22:20:22.428 -> Water_Temp.:93.87, PWM:29
22:20:23.185 -> Water_Temp.:93.87, PWM:29
22:20:23.981 -> Water_Temp.:93.87, PWM:28
22:20:24.743 -> Water_Temp.:93.87, PWM:28
22:20:25.537 -> Water_Temp.:93.87, PWM:29
22:20:26.327 -> Water_Temp.:93.87, PWM:28
22:20:27.089 -> Water_Temp.:93.87, PWM:29
22:20:27.885 -> Water_Temp.:93.87, PWM:29
22:20:28.646 -> Water_Temp.:93.87, PWM:28
22:20:29.439 -> Water_Temp.:93.87, PWM:29
22:20:30.198 -> Water_Temp.:93.87, PWM:28
22:20:30.996 -> Water_Temp.:93.87, PWM:29
22:20:31.787 -> Water_Temp.:93.87, PWM:28
22:20:33.341 -> Water_Temp.:93.99, PWM:29
22:20:34.104 -> Water_Temp.:93.99, PWM:29
22:20:34.896 -> Water_Temp.:93.99, PWM:28
22:20:35.655 -> Water_Temp.:93.99, PWM:29
22:20:36.449 -> Water_Temp.:93.99, PWM:28
22:20:37.248 -> Water_Temp.:93.99, PWM:28
22:20:38.008 -> Water_Temp.:93.99, PWM:29
22:20:38.804 -> Water_Temp.:93.99, PWM:28
22:20:39.562 -> Water_Temp.:93.99, PWM:29
22:20:40.357 -> Water_Temp.:93.99, PWM:29
22:20:41.913 -> Water_Temp.:93.99, PWM:29
22:20:42.708 -> Water_Temp.:93.99, PWM:28
22:20:43.468 -> Water_Temp.:93.99, PWM:28
22:20:44.260 -> Water_Temp.:93.99, PWM:29
22:20:45.024 -> Water_Temp.:93.99, PWM:29
22:20:45.821 -> Water_Temp.:93.99, PWM:28
22:20:46.577 -> Water_Temp.:93.99, PWM:29
22:20:47.370 -> Water_Temp.:93.99, PWM:28
22:20:48.165 -> Water_Temp.:93.99, PWM:29
22:20:48.929 -> Water_Temp.:93.99, PWM:28
22:20:49.724 -> Water_Temp.:93.99, PWM:28
22:20:50.484 -> Water_Temp.:93.99, PWM:29
22:20:52.008 -> Water_Temp.:93.99, PWM:28
22:20:52.804 -> Water_Temp.:94.10, PWM:29
22:20:53.596 -> Water_Temp.:94.10, PWM:29
22:20:54.356 -> Water_Temp.:94.10, PWM:29
22:20:55.155 -> Water_Temp.:94.10, PWM:29
22:20:55.912 -> Water_Temp.:94.10, PWM:28
22:20:56.700 -> Water_Temp.:94.10, PWM:29
22:20:57.497 -> Water_Temp.:94.10, PWM:28
22:20:58.256 -> Water_Temp.:94.10, PWM:28
22:20:59.048 -> Water_Temp.:94.10, PWM:28
22:20:59.841 -> Water_Temp.:94.10, PWM:29
22:21:00.599 -> Water_Temp.:94.10, PWM:28
22:21:02.154 -> Water_Temp.:94.10, PWM:29
22:21:02.946 -> Water_Temp.:94.10, PWM:29
22:21:04.497 -> Water_Temp.:94.10, PWM:29
22:21:05.258 -> Water_Temp.:94.10, PWM:28
22:21:06.055 -> Water_Temp.:94.10, PWM:28
22:21:06.817 -> Water_Temp.:94.10, PWM:29
22:21:07.613 -> Water_Temp.:94.10, PWM:28
22:21:08.377 -> Water_Temp.:94.10, PWM:29
22:21:09.171 -> Water_Temp.:94.10, PWM:28
22:21:09.961 -> Water_Temp.:94.10, PWM:29
22:21:10.723 -> Water_Temp.:94.10, PWM:29
22:21:11.519 -> Water_Temp.:94.10, PWM:29
22:21:12.278 -> Water_Temp.:94.10, PWM:28
22:21:13.074 -> Water_Temp.:94.10, PWM:28
22:21:13.833 -> Water_Temp.:94.10, PWM:29
22:21:14.628 -> Water_Temp.:94.10, PWM:28
22:21:15.423 -> Water_Temp.:94.10, PWM:29
22:21:16.184 -> Water_Temp.:94.10, PWM:28
22:21:16.978 -> Water_Temp.:94.10, PWM:29
22:21:17.738 -> Water_Temp.:94.21, PWM:28
22:21:18.533 -> Water_Temp.:94.21, PWM:28
22:21:19.294 -> Water_Temp.:94.21, PWM:29
22:21:20.084 -> Water_Temp.:94.21, PWM:28
22:21:21.633 -> Water_Temp.:94.21, PWM:29
22:21:22.394 -> Water_Temp.:94.21, PWM:28
22:21:23.982 -> Water_Temp.:94.21, PWM:28
22:21:24.742 -> Water_Temp.:94.21, PWM:29
22:21:25.540 -> Water_Temp.:94.21, PWM:29
22:21:26.296 -> Water_Temp.:94.21, PWM:29
22:21:27.089 -> Water_Temp.:94.32, PWM:29
22:21:27.881 -> Water_Temp.:94.32, PWM:28
22:21:28.641 -> Water_Temp.:94.32, PWM:29
22:21:29.439 -> Water_Temp.:94.32, PWM:29
22:21:30.198 -> Water_Temp.:94.32, PWM:29
22:21:30.996 -> Water_Temp.:94.32, PWM:28
22:21:31.753 -> Water_Temp.:94.32, PWM:29
22:21:32.545 -> Water_Temp.:94.32, PWM:28
22:21:33.339 -> Water_Temp.:94.32, PWM:28
22:21:34.099 -> Water_Temp.:94.32, PWM:29
22:21:34.893 -> Water_Temp.:94.32, PWM:28
22:21:36.451 -> Water_Temp.:94.32, PWM:29
22:21:37.211 -> Water_Temp.:94.44, PWM:29
22:21:38.007 -> Water_Temp.:94.44, PWM:28
22:21:38.802 -> Water_Temp.:94.44, PWM:29
22:21:39.560 -> Water_Temp.:94.44, PWM:28
22:21:40.352 -> Water_Temp.:94.44, PWM:28
22:21:41.116 -> Water_Temp.:94.44, PWM:29
22:21:41.907 -> Water_Temp.:94.44, PWM:29
22:21:42.700 -> Water_Temp.:94.44, PWM:28
22:21:44.225 -> Water_Temp.:94.44, PWM:29
22:21:44.989 -> Water_Temp.:94.44, PWM:28
22:21:45.783 -> Water_Temp.:94.44, PWM:29
22:21:46.545 -> Water_Temp.:94.44, PWM:29
22:21:47.340 -> Water_Temp.:94.44, PWM:29
22:21:48.132 -> Water_Temp.:94.44, PWM:28
22:21:48.891 -> Water_Temp.:94.44, PWM:29
22:21:49.688 -> Water_Temp.:94.44, PWM:28
22:21:50.448 -> Water_Temp.:94.44, PWM:29
22:21:51.240 -> Water_Temp.:94.44, PWM:28
22:21:52.035 -> Water_Temp.:94.44, PWM:29
22:21:52.793 -> Water_Temp.:94.44, PWM:28
22:21:53.588 -> Water_Temp.:94.44, PWM:28
22:21:54.348 -> Water_Temp.:94.44, PWM:29
22:21:55.144 -> Water_Temp.:94.44, PWM:29
22:21:55.909 -> Water_Temp.:94.44, PWM:29
22:21:56.701 -> Water_Temp.:94.44, PWM:29
22:21:57.464 -> Water_Temp.:94.44, PWM:28
22:21:58.254 -> Water_Temp.:94.44, PWM:29
22:21:59.047 -> Water_Temp.:94.44, PWM:28
22:21:59.806 -> Water_Temp.:94.44, PWM:29
22:22:00.601 -> Water_Temp.:94.44, PWM:28
22:22:01.394 -> Water_Temp.:94.44, PWM:29
22:22:02.149 -> Water_Temp.:94.44, PWM:28
22:22:02.947 -> Water_Temp.:94.44, PWM:28
22:22:03.704 -> Water_Temp.:94.44, PWM:28
22:22:04.499 -> Water_Temp.:94.44, PWM:29
22:22:05.295 -> Water_Temp.:94.44, PWM:29
22:22:06.055 -> Water_Temp.:94.44, PWM:29
22:22:06.848 -> Water_Temp.:94.44, PWM:29
22:22:07.607 -> Water_Temp.:94.44, PWM:29
22:22:08.401 -> Water_Temp.:94.44, PWM:28
22:22:09.162 -> Water_Temp.:94.44, PWM:28
22:22:09.954 -> Water_Temp.:94.44, PWM:29
22:22:10.749 -> Water_Temp.:94.44, PWM:29
22:22:11.510 -> Water_Temp.:94.44, PWM:28
22:22:12.304 -> Water_Temp.:94.44, PWM:29
22:22:13.063 -> Water_Temp.:94.44, PWM:29
22:22:13.855 -> Water_Temp.:94.44, PWM:28
22:22:14.652 -> Water_Temp.:94.44, PWM:28
22:22:15.409 -> Water_Temp.:94.44, PWM:29
22:22:16.203 -> Water_Temp.:94.44, PWM:29
22:22:16.964 -> Water_Temp.:94.44, PWM:29
22:22:17.761 -> Water_Temp.:94.44, PWM:28
22:22:18.522 -> Water_Temp.:94.44, PWM:28
22:22:19.317 -> Water_Temp.:94.55, PWM:29
22:22:20.111 -> Water_Temp.:94.55, PWM:29
22:22:20.872 -> Water_Temp.:94.55, PWM:29
22:22:21.667 -> Water_Temp.:94.55, PWM:29
22:22:22.428 -> Water_Temp.:94.55, PWM:28
22:22:23.217 -> Water_Temp.:94.55, PWM:28
22:22:24.010 -> Water_Temp.:94.55, PWM:29
22:22:24.770 -> Water_Temp.:94.55, PWM:28
22:22:25.571 -> Water_Temp.:94.55, PWM:28
22:22:26.329 -> Water_Temp.:94.55, PWM:28
22:22:27.120 -> Water_Temp.:94.55, PWM:29
22:22:27.909 -> Water_Temp.:94.55, PWM:28
22:22:28.665 -> Water_Temp.:94.55, PWM:28
22:22:29.458 -> Water_Temp.:94.55, PWM:28
22:22:30.249 -> Water_Temp.:94.55, PWM:28
22:22:31.805 -> Water_Temp.:94.55, PWM:28
22:22:33.319 -> Water_Temp.:94.66, PWM:28
22:22:34.113 -> Water_Temp.:94.66, PWM:28
22:22:34.877 -> Water_Temp.:94.66, PWM:28
22:22:35.671 -> Water_Temp.:94.66, PWM:28
22:22:36.463 -> Water_Temp.:94.66, PWM:29
22:22:37.213 -> Water_Temp.:94.66, PWM:29
22:22:38.009 -> Water_Temp.:94.66, PWM:29
22:22:38.801 -> Water_Temp.:94.77, PWM:29
22:22:39.560 -> Water_Temp.:94.77, PWM:28
22:22:40.352 -> Water_Temp.:94.77, PWM:28
22:22:41.112 -> Water_Temp.:94.77, PWM:28
22:22:41.902 -> Water_Temp.:94.77, PWM:28
22:22:42.698 -> Water_Temp.:94.89, PWM:28
22:22:43.457 -> Water_Temp.:94.89, PWM:28
22:22:44.250 -> Water_Temp.:94.89, PWM:28
22:22:45.012 -> Water_Temp.:94.89, PWM:29
22:22:45.810 -> Water_Temp.:94.89, PWM:29
22:22:46.601 -> Water_Temp.:94.89, PWM:29
22:22:47.358 -> Water_Temp.:94.89, PWM:29
22:22:48.155 -> Water_Temp.:95.00, PWM:28
22:22:48.912 -> Water_Temp.:95.00, PWM:28
22:22:49.707 -> Water_Temp.:95.00, PWM:29
22:22:50.504 -> Water_Temp.:95.00, PWM:29
22:22:51.262 -> Water_Temp.:95.00, PWM:28
22:22:52.057 -> Water_Temp.:95.00, PWM:29
22:22:52.817 -> Water_Temp.:96.57, PWM:29
22:22:53.612 -> Water_Temp.:98.71, PWM:29
22:22:54.371 -> Water_Temp.:96.24, PWM:28
22:22:55.169 -> Water_Temp.:95.90, PWM:28
22:22:55.962 -> Water_Temp.:95.68, PWM:29
22:22:56.720 -> Water_Temp.:95.56, PWM:29
22:22:58.274 -> Water_Temp.:95.34, PWM:29
22:22:59.071 -> Water_Temp.:95.34, PWM:28
22:22:59.831 -> Water_Temp.:95.22, PWM:29
22:23:00.625 -> Water_Temp.:95.22, PWM:28
22:23:01.421 -> Water_Temp.:95.22, PWM:28
22:23:02.182 -> Water_Temp.:95.22, PWM:29
22:23:02.975 -> Water_Temp.:95.22, PWM:28
22:23:03.735 -> Water_Temp.:95.11, PWM:28
22:23:04.531 -> Water_Temp.:95.11, PWM:28
22:23:05.291 -> Water_Temp.:95.11, PWM:29
22:23:06.089 -> Water_Temp.:95.11, PWM:28
22:23:06.881 -> Water_Temp.:95.11, PWM:28
22:23:07.645 -> Water_Temp.:95.11, PWM:28
22:23:08.437 -> Water_Temp.:95.11, PWM:29
22:23:09.198 -> Water_Temp.:95.11, PWM:29
22:23:09.996 -> Water_Temp.:95.11, PWM:29
22:23:10.756 -> Water_Temp.:95.11, PWM:29
22:23:11.549 -> Water_Temp.:95.11, PWM:29
22:23:12.340 -> Water_Temp.:95.11, PWM:29
22:23:13.100 -> Water_Temp.:95.11, PWM:29
22:23:13.894 -> Water_Temp.:95.11, PWM:28
22:23:14.655 -> Water_Temp.:95.11, PWM:29
22:23:16.242 -> Water_Temp.:95.11, PWM:28
22:23:17.003 -> Water_Temp.:95.11, PWM:29
22:23:17.797 -> Water_Temp.:95.11, PWM:29
22:23:19.345 -> Water_Temp.:95.11, PWM:29
22:23:20.139 -> Water_Temp.:95.11, PWM:28
22:23:20.904 -> Water_Temp.:95.11, PWM:28
22:23:21.665 -> Water_Temp.:95.11, PWM:28
22:23:22.456 -> Water_Temp.:95.11, PWM:29
22:23:23.252 -> Water_Temp.:95.22, PWM:28
22:23:24.011 -> Water_Temp.:95.22, PWM:28
22:23:24.801 -> Water_Temp.:95.22, PWM:29
22:23:25.598 -> Water_Temp.:95.22, PWM:29
22:23:26.355 -> Water_Temp.:95.22, PWM:29
22:23:27.144 -> Water_Temp.:95.22, PWM:28
22:23:27.937 -> Water_Temp.:95.22, PWM:28
22:23:28.697 -> Water_Temp.:95.22, PWM:28
22:23:29.486 -> Water_Temp.:95.22, PWM:29
22:23:30.277 -> Water_Temp.:95.22, PWM:29
22:23:31.043 -> Water_Temp.:95.22, PWM:29
22:23:31.806 -> Water_Temp.:95.22, PWM:29
22:23:32.603 -> Water_Temp.:95.34, PWM:28
22:23:33.395 -> Water_Temp.:95.34, PWM:28
22:23:34.154 -> Water_Temp.:95.34, PWM:29
22:23:34.950 -> Water_Temp.:95.34, PWM:29
22:23:35.708 -> Water_Temp.:95.34, PWM:28
22:23:36.501 -> Water_Temp.:95.34, PWM:29
22:23:37.291 -> Water_Temp.:95.34, PWM:29
22:23:38.051 -> Water_Temp.:95.34, PWM:28
22:23:38.842 -> Water_Temp.:95.34, PWM:28
22:23:39.604 -> Water_Temp.:95.34, PWM:29
22:23:40.402 -> Water_Temp.:95.34, PWM:29
22:23:41.166 -> Water_Temp.:95.34, PWM:28
22:23:41.959 -> Water_Temp.:95.34, PWM:29
22:23:42.724 -> Water_Temp.:95.34, PWM:28
22:23:43.518 -> Water_Temp.:95.34, PWM:29
22:23:44.313 -> Water_Temp.:95.34, PWM:29
22:23:45.074 -> Water_Temp.:95.34, PWM:29
22:23:45.871 -> Water_Temp.:95.34, PWM:28
22:23:47.394 -> Water_Temp.:95.34, PWM:29
22:23:48.187 -> Water_Temp.:95.34, PWM:28
22:23:48.948 -> Water_Temp.:95.34, PWM:29
22:23:49.746 -> Water_Temp.:95.34, PWM:29
22:23:50.506 -> Water_Temp.:95.34, PWM:29
22:23:51.302 -> Water_Temp.:95.34, PWM:28
22:23:52.859 -> Water_Temp.:95.34, PWM:28
22:23:53.620 -> Water_Temp.:95.34, PWM:29
22:23:54.417 -> Water_Temp.:95.34, PWM:29
22:23:55.184 -> Water_Temp.:95.34, PWM:28
22:23:55.981 -> Water_Temp.:95.34, PWM:29
22:23:56.739 -> Water_Temp.:95.34, PWM:28
22:23:57.533 -> Water_Temp.:95.34, PWM:29
22:23:58.330 -> Water_Temp.:95.34, PWM:29
22:23:59.093 -> Water_Temp.:95.45, PWM:28
22:23:59.889 -> Water_Temp.:95.45, PWM:29
22:24:00.651 -> Water_Temp.:95.45, PWM:29
22:24:01.446 -> Water_Temp.:95.45, PWM:29
22:24:02.206 -> Water_Temp.:95.45, PWM:28
22:24:02.999 -> Water_Temp.:95.45, PWM:28
22:24:03.789 -> Water_Temp.:95.45, PWM:28
22:24:04.551 -> Water_Temp.:95.45, PWM:29
22:24:05.348 -> Water_Temp.:95.45, PWM:29
22:24:06.108 -> Water_Temp.:95.45, PWM:29
22:24:07.659 -> Water_Temp.:95.45, PWM:29
22:24:08.422 -> Water_Temp.:95.45, PWM:29
22:24:09.219 -> Water_Temp.:95.45, PWM:28
22:24:09.985 -> Water_Temp.:95.45, PWM:29
22:24:10.780 -> Water_Temp.:95.45, PWM:28
22:24:11.538 -> Water_Temp.:95.45, PWM:29
22:24:12.335 -> Water_Temp.:95.45, PWM:28
22:24:13.128 -> Water_Temp.:95.45, PWM:29
22:24:13.891 -> Water_Temp.:95.45, PWM:28
22:24:14.686 -> Water_Temp.:95.56, PWM:29
22:24:15.449 -> Water_Temp.:95.56, PWM:28
22:24:16.245 -> Water_Temp.:95.56, PWM:29
22:24:16.999 -> Water_Temp.:95.56, PWM:29
22:24:17.790 -> Water_Temp.:95.56, PWM:28
22:24:18.587 -> Water_Temp.:95.56, PWM:28
22:24:19.344 -> Water_Temp.:95.56, PWM:29
22:24:20.140 -> Water_Temp.:95.56, PWM:28
22:24:20.905 -> Water_Temp.:95.56, PWM:28
22:24:21.705 -> Water_Temp.:95.56, PWM:29
22:24:22.468 -> Water_Temp.:95.56, PWM:28
22:24:23.262 -> Water_Temp.:95.56, PWM:28
22:24:24.022 -> Water_Temp.:95.56, PWM:29
22:24:24.820 -> Water_Temp.:95.56, PWM:29
22:24:25.585 -> Water_Temp.:95.56, PWM:28
22:24:26.373 -> Water_Temp.:95.56, PWM:28
22:24:27.169 -> Water_Temp.:95.56, PWM:28
22:24:27.930 -> Water_Temp.:95.56, PWM:29
22:24:28.726 -> Water_Temp.:95.56, PWM:28
22:24:29.486 -> Water_Temp.:95.56, PWM:28
22:24:31.043 -> Water_Temp.:95.56, PWM:29
22:24:31.834 -> Water_Temp.:95.56, PWM:29
22:24:32.596 -> Water_Temp.:95.56, PWM:28
22:24:33.389 -> Water_Temp.:97.14, PWM:29
22:24:34.146 -> Water_Temp.:99.27, PWM:29
22:24:34.939 -> Water_Temp.:96.80, PWM:28
22:24:35.700 -> Water_Temp.:96.46, PWM:29
22:24:36.491 -> Water_Temp.:96.24, PWM:29
22:24:37.284 -> Water_Temp.:96.12, PWM:29
22:24:38.046 -> Water_Temp.:96.01, PWM:28
22:24:38.841 -> Water_Temp.:95.90, PWM:29
22:24:39.633 -> Water_Temp.:95.90, PWM:29
22:24:40.393 -> Water_Temp.:95.79, PWM:29
22:24:41.186 -> Water_Temp.:95.79, PWM:28
22:24:41.946 -> Water_Temp.:95.79, PWM:28
22:24:42.741 -> Water_Temp.:95.79, PWM:28
22:24:43.501 -> Water_Temp.:95.68, PWM:29
22:24:44.296 -> Water_Temp.:95.68, PWM:29
22:24:45.060 -> Water_Temp.:97.36, PWM:28
22:24:45.860 -> Water_Temp.:99.39, PWM:29
22:24:46.619 -> Water_Temp.:96.91, PWM:28
22:24:47.409 -> Water_Temp.:96.57, PWM:28
22:24:48.203 -> Water_Temp.:96.35, PWM:28
22:24:48.960 -> Water_Temp.:96.24, PWM:29
22:24:49.751 -> Water_Temp.:96.12, PWM:28
22:24:50.547 -> Water_Temp.:96.01, PWM:28
22:24:51.307 -> Water_Temp.:96.01, PWM:28
22:24:52.102 -> Water_Temp.:95.90, PWM:28
22:24:52.862 -> Water_Temp.:95.90, PWM:29
22:24:53.656 -> Water_Temp.:95.79, PWM:28
22:24:54.420 -> Water_Temp.:95.79, PWM:29
22:24:55.217 -> Water_Temp.:95.79, PWM:28
22:24:56.011 -> Water_Temp.:95.79, PWM:29
22:24:56.769 -> Water_Temp.:95.79, PWM:28
22:24:57.568 -> Water_Temp.:95.79, PWM:29
22:24:58.328 -> Water_Temp.:95.79, PWM:29
22:24:59.122 -> Water_Temp.:95.79, PWM:29
22:24:59.882 -> Water_Temp.:95.79, PWM:28
22:25:00.675 -> Water_Temp.:95.79, PWM:29
22:25:01.471 -> Water_Temp.:95.79, PWM:29
22:25:02.996 -> Water_Temp.:95.79, PWM:29
22:25:03.789 -> Water_Temp.:95.79, PWM:28
22:25:04.549 -> Water_Temp.:95.79, PWM:28
22:25:05.341 -> Water_Temp.:95.79, PWM:29
22:25:06.135 -> Water_Temp.:95.79, PWM:29
22:25:06.891 -> Water_Temp.:95.79, PWM:29
22:25:07.684 -> Water_Temp.:95.79, PWM:28
22:25:08.442 -> Water_Temp.:95.79, PWM:29
22:25:09.237 -> Water_Temp.:95.79, PWM:29
22:25:10.029 -> Water_Temp.:95.79, PWM:29
22:25:10.787 -> Water_Temp.:95.79, PWM:29
22:25:11.582 -> Water_Temp.:95.79, PWM:29
22:25:12.342 -> Water_Temp.:95.79, PWM:28
22:25:13.138 -> Water_Temp.:95.79, PWM:29
22:25:13.929 -> Water_Temp.:95.79, PWM:29
22:25:14.690 -> Water_Temp.:95.79, PWM:29
22:25:15.487 -> Water_Temp.:95.79, PWM:29
22:25:17.045 -> Water_Temp.:95.79, PWM:29
22:25:17.803 -> Water_Temp.:95.79, PWM:28
22:25:19.360 -> Water_Temp.:95.79, PWM:29
22:25:20.151 -> Water_Temp.:95.79, PWM:28
22:25:20.945 -> Water_Temp.:95.79, PWM:28
22:25:21.711 -> Water_Temp.:95.79, PWM:28
22:25:22.503 -> Water_Temp.:95.79, PWM:28
22:25:23.263 -> Water_Temp.:95.79, PWM:29
22:25:24.818 -> Water_Temp.:95.79, PWM:28
22:25:25.610 -> Water_Temp.:95.79, PWM:29
22:25:26.403 -> Water_Temp.:95.79, PWM:29
22:25:27.959 -> Water_Temp.:95.79, PWM:28
22:25:28.720 -> Water_Temp.:95.79, PWM:29
22:25:29.516 -> Water_Temp.:95.79, PWM:29
22:25:30.309 -> Water_Temp.:95.79, PWM:29
22:25:31.071 -> Water_Temp.:95.79, PWM:29
22:25:31.867 -> Water_Temp.:95.79, PWM:29
22:25:32.628 -> Water_Temp.:95.79, PWM:29
22:25:34.185 -> Water_Temp.:95.79, PWM:29
22:25:34.981 -> Water_Temp.:95.79, PWM:29
22:25:35.742 -> Water_Temp.:95.79, PWM:29
22:25:36.535 -> Water_Temp.:95.79, PWM:29
22:25:37.295 -> Water_Temp.:95.79, PWM:29
22:25:38.092 -> Water_Temp.:95.79, PWM:29
22:25:38.887 -> Water_Temp.:95.79, PWM:28
22:25:39.648 -> Water_Temp.:95.79, PWM:28
22:25:40.413 -> Water_Temp.:95.79, PWM:29
22:25:41.209 -> Water_Temp.:95.79, PWM:29
22:25:42.759 -> Water_Temp.:95.79, PWM:29
22:25:43.554 -> Water_Temp.:95.79, PWM:29
22:25:44.314 -> Water_Temp.:95.90, PWM:28
22:25:45.112 -> Water_Temp.:95.90, PWM:28
22:25:45.905 -> Water_Temp.:95.90, PWM:29
22:25:46.663 -> Water_Temp.:95.90, PWM:28
22:25:47.460 -> Water_Temp.:95.90, PWM:28
22:25:48.984 -> Water_Temp.:95.90, PWM:29
22:25:49.777 -> Water_Temp.:95.90, PWM:28
22:25:50.538 -> Water_Temp.:95.90, PWM:28
22:25:51.332 -> Water_Temp.:95.90, PWM:28
22:25:52.089 -> Water_Temp.:95.90, PWM:29
22:25:52.883 -> Water_Temp.:95.90, PWM:28
22:25:53.680 -> Water_Temp.:95.90, PWM:29
22:25:54.442 -> Water_Temp.:95.90, PWM:29
22:25:55.240 -> Water_Temp.:95.90, PWM:28
22:25:55.999 -> Water_Temp.:95.90, PWM:28
22:25:56.790 -> Water_Temp.:95.90, PWM:28
22:25:57.552 -> Water_Temp.:95.90, PWM:29
22:25:58.340 -> Water_Temp.:95.90, PWM:29
22:25:59.134 -> Water_Temp.:95.90, PWM:29
22:25:59.892 -> Water_Temp.:95.90, PWM:29
22:26:01.446 -> Water_Temp.:95.90, PWM:29
22:26:02.206 -> Water_Temp.:95.90, PWM:29
22:26:02.998 -> Water_Temp.:95.90, PWM:29
22:26:03.790 -> Water_Temp.:95.90, PWM:29
22:26:04.550 -> Water_Temp.:95.90, PWM:29
22:26:05.346 -> Water_Temp.:95.90, PWM:29
22:26:06.106 -> Water_Temp.:95.90, PWM:28
22:26:06.902 -> Water_Temp.:95.90, PWM:28
22:26:07.699 -> Water_Temp.:95.90, PWM:28
22:26:08.462 -> Water_Temp.:95.90, PWM:29
22:26:09.252 -> Water_Temp.:96.01, PWM:29
22:26:10.016 -> Water_Temp.:95.90, PWM:29
22:26:10.809 -> Water_Temp.:95.90, PWM:29
22:26:11.566 -> Water_Temp.:95.90, PWM:29
22:26:12.359 -> Water_Temp.:95.90, PWM:29
22:26:13.152 -> Water_Temp.:95.90, PWM:29
22:26:14.676 -> Water_Temp.:95.90, PWM:29
22:26:15.475 -> Water_Temp.:96.01, PWM:29
22:26:16.232 -> Water_Temp.:95.90, PWM:28
22:26:17.030 -> Water_Temp.:95.90, PWM:28
22:26:17.792 -> Water_Temp.:96.01, PWM:29
22:26:19.347 -> Water_Temp.:96.01, PWM:28
22:26:20.104 -> Water_Temp.:96.01, PWM:29
22:26:20.903 -> Water_Temp.:96.01, PWM:28
22:26:22.454 -> Water_Temp.:96.01, PWM:29
22:26:23.246 -> Water_Temp.:96.01, PWM:29
22:26:24.006 -> Water_Temp.:96.01, PWM:29
22:26:24.799 -> Water_Temp.:96.01, PWM:28
22:26:25.561 -> Water_Temp.:96.01, PWM:29
22:26:26.353 -> Water_Temp.:96.01, PWM:29
22:26:27.143 -> Water_Temp.:96.01, PWM:29
22:26:28.699 -> Water_Temp.:96.01, PWM:29
22:26:29.459 -> Water_Temp.:96.01, PWM:28
22:26:30.257 -> Water_Temp.:96.01, PWM:28
22:26:31.024 -> Water_Temp.:96.01, PWM:29
22:26:31.819 -> Water_Temp.:96.01, PWM:28
22:26:32.610 -> Water_Temp.:96.01, PWM:28
22:26:33.368 -> Water_Temp.:96.01, PWM:29
22:26:34.164 -> Water_Temp.:96.01, PWM:29
22:26:34.925 -> Water_Temp.:96.01, PWM:29
22:26:35.721 -> Water_Temp.:96.01, PWM:29
22:26:36.482 -> Water_Temp.:97.70, PWM:28
22:26:37.277 -> Water_Temp.:99.72, PWM:29
22:26:38.073 -> Water_Temp.:97.36, PWM:29
22:26:38.832 -> Water_Temp.:96.91, PWM:28
22:26:39.627 -> Water_Temp.:96.69, PWM:29
22:26:40.389 -> Water_Temp.:96.57, PWM:28
22:26:41.186 -> Water_Temp.:96.46, PWM:28
22:26:41.946 -> Water_Temp.:96.46, PWM:28
22:26:42.736 -> Water_Temp.:96.35, PWM:29
22:26:43.526 -> Water_Temp.:96.35, PWM:29
22:26:44.284 -> Water_Temp.:96.35, PWM:29
22:26:45.078 -> Water_Temp.:96.24, PWM:28
22:26:45.872 -> Water_Temp.:96.24, PWM:28
22:26:46.630 -> Water_Temp.:96.24, PWM:28
22:26:47.424 -> Water_Temp.:96.24, PWM:28
22:26:48.185 -> Water_Temp.:97.81, PWM:29
22:26:48.976 -> Water_Temp.:99.84, PWM:28
22:26:49.770 -> Water_Temp.:97.47, PWM:29
22:26:50.529 -> Water_Temp.:97.02, PWM:28
22:26:51.326 -> Water_Temp.:97.02, PWM:29
22:26:52.085 -> Water_Temp.:96.69, PWM:29
22:26:52.880 -> Water_Temp.:96.57, PWM:29
22:26:54.433 -> Water_Temp.:96.46, PWM:28
22:26:55.227 -> Water_Temp.:96.46, PWM:28
22:26:55.985 -> Water_Temp.:96.35, PWM:28
22:26:56.778 -> Water_Temp.:96.35, PWM:29
22:26:57.569 -> Water_Temp.:96.35, PWM:29
22:26:58.329 -> Water_Temp.:96.35, PWM:29
22:26:59.124 -> Water_Temp.:96.35, PWM:29
22:26:59.884 -> Water_Temp.:96.35, PWM:28
22:27:00.680 -> Water_Temp.:96.35, PWM:29
22:27:01.440 -> Water_Temp.:96.35, PWM:28
22:27:02.231 -> Water_Temp.:96.24, PWM:29
22:27:03.024 -> Water_Temp.:96.35, PWM:28
22:27:03.780 -> Water_Temp.:96.24, PWM:29
22:27:04.571 -> Water_Temp.:96.24, PWM:29
22:27:05.366 -> Water_Temp.:96.24, PWM:29
22:27:06.129 -> Water_Temp.:96.24, PWM:29
22:27:06.922 -> Water_Temp.:96.24, PWM:29
22:27:07.679 -> Water_Temp.:96.24, PWM:28
22:27:08.475 -> Water_Temp.:96.24, PWM:28
22:27:09.237 -> Water_Temp.:96.24, PWM:29
22:27:10.031 -> Water_Temp.:96.24, PWM:28
22:27:10.796 -> Water_Temp.:96.24, PWM:29
22:27:11.587 -> Water_Temp.:96.24, PWM:29
22:27:12.383 -> Water_Temp.:96.24, PWM:28
22:27:13.140 -> Water_Temp.:96.24, PWM:29
22:27:13.936 -> Water_Temp.:96.24, PWM:29
22:27:14.695 -> Water_Temp.:96.24, PWM:28
22:27:15.485 -> Water_Temp.:96.24, PWM:28
22:27:16.281 -> Water_Temp.:96.24, PWM:29
22:27:17.040 -> Water_Temp.:96.24, PWM:29
22:27:17.832 -> Water_Temp.:96.24, PWM:29
22:27:18.626 -> Water_Temp.:96.24, PWM:29
22:27:19.382 -> Water_Temp.:96.24, PWM:28
22:27:20.178 -> Water_Temp.:96.24, PWM:28
22:27:20.943 -> Water_Temp.:96.24, PWM:29
22:27:21.734 -> Water_Temp.:97.92, PWM:29
22:27:22.526 -> Water_Temp.:99.95, PWM:28
22:27:24.079 -> Water_Temp.:97.14, PWM:29
22:27:24.839 -> Water_Temp.:96.91, PWM:29
22:27:25.633 -> Water_Temp.:96.80, PWM:29
22:27:27.185 -> Water_Temp.:96.69, PWM:28
22:27:27.941 -> Water_Temp.:96.57, PWM:28
22:27:28.735 -> Water_Temp.:96.57, PWM:29
22:27:30.289 -> Water_Temp.:96.57, PWM:29
22:27:31.053 -> Water_Temp.:96.46, PWM:29
22:27:31.845 -> Water_Temp.:96.46, PWM:28
22:27:32.638 -> Water_Temp.:96.46, PWM:29
22:27:33.400 -> Water_Temp.:96.46, PWM:28
22:27:34.195 -> Water_Temp.:96.46, PWM:28
22:27:34.953 -> Water_Temp.:96.46, PWM:28
22:27:36.537 -> Water_Temp.:96.46, PWM:29
22:27:37.300 -> Water_Temp.:96.46, PWM:29
22:27:38.094 -> Water_Temp.:96.46, PWM:28
22:27:38.855 -> Water_Temp.:96.46, PWM:29
22:27:39.648 -> Water_Temp.:96.46, PWM:29
22:27:40.412 -> Water_Temp.:96.35, PWM:29
22:27:41.209 -> Water_Temp.:96.35, PWM:29
22:27:41.969 -> Water_Temp.:96.35, PWM:29
22:27:42.765 -> Water_Temp.:96.35, PWM:29
22:27:43.562 -> Water_Temp.:96.35, PWM:28
22:27:44.316 -> Water_Temp.:96.46, PWM:29
22:27:45.113 -> Water_Temp.:96.46, PWM:29
22:27:45.869 -> Water_Temp.:96.46, PWM:28
22:27:47.424 -> Water_Temp.:96.46, PWM:29
22:27:48.185 -> Water_Temp.:96.46, PWM:29
22:27:48.983 -> Water_Temp.:96.46, PWM:28
22:27:49.775 -> Water_Temp.:96.46, PWM:29
22:27:50.536 -> Water_Temp.:96.46, PWM:28
22:27:51.332 -> Water_Temp.:96.46, PWM:28
22:27:52.088 -> Water_Temp.:96.46, PWM:29
22:27:52.881 -> Water_Temp.:96.46, PWM:29
22:27:53.679 -> Water_Temp.:96.46, PWM:29
22:27:54.439 -> Water_Temp.:96.46, PWM:28
22:27:55.232 -> Water_Temp.:96.46, PWM:28
22:27:55.995 -> Water_Temp.:96.46, PWM:29
22:27:56.789 -> Water_Temp.:96.46, PWM:28
22:27:57.553 -> Water_Temp.:96.46, PWM:29
22:27:58.347 -> Water_Temp.:96.46, PWM:28
22:27:59.897 -> Water_Temp.:96.46, PWM:28
22:28:00.692 -> Water_Temp.:96.46, PWM:29
22:28:01.456 -> Water_Temp.:96.46, PWM:29
22:28:03.014 -> Water_Temp.:96.46, PWM:29
22:28:04.567 -> Water_Temp.:96.46, PWM:29
22:28:05.362 -> Water_Temp.:96.46, PWM:29
22:28:08.438 -> Water_Temp.:96.46, PWM:29
22:28:09.235 -> Water_Temp.:96.46, PWM:29
22:28:09.996 -> Water_Temp.:96.46, PWM:29
22:28:10.785 -> Water_Temp.:96.46, PWM:28
22:28:11.549 -> Water_Temp.:96.46, PWM:29
22:28:12.345 -> Water_Temp.:96.46, PWM:29
22:28:13.896 -> Water_Temp.:96.46, PWM:28
22:28:14.690 -> Water_Temp.:96.46, PWM:29
22:28:15.452 -> Water_Temp.:96.46, PWM:29
22:28:16.247 -> Water_Temp.:96.46, PWM:29
22:28:17.040 -> Water_Temp.:96.46, PWM:29
22:28:17.802 -> Water_Temp.:96.46, PWM:29
22:28:18.598 -> Water_Temp.:96.46, PWM:28
22:28:19.361 -> Water_Temp.:96.46, PWM:28
22:28:20.152 -> Water_Temp.:96.46, PWM:29
22:28:20.911 -> Water_Temp.:96.46, PWM:29
22:28:21.704 -> Water_Temp.:96.46, PWM:29
22:28:22.496 -> Water_Temp.:96.46, PWM:29
22:28:23.256 -> Water_Temp.:96.46, PWM:29
22:28:24.050 -> Water_Temp.:96.46, PWM:29
22:28:24.815 -> Water_Temp.:96.46, PWM:28
22:28:25.612 -> Water_Temp.:96.46, PWM:29
22:28:26.368 -> Water_Temp.:96.46, PWM:28
22:28:27.166 -> Water_Temp.:96.46, PWM:28
22:28:27.928 -> Water_Temp.:96.46, PWM:28
22:28:28.719 -> Water_Temp.:96.46, PWM:29
22:28:29.513 -> Water_Temp.:96.46, PWM:29
22:28:30.269 -> Water_Temp.:96.46, PWM:29
22:28:31.064 -> Water_Temp.:96.46, PWM:29
22:28:31.857 -> Water_Temp.:96.46, PWM:29
22:28:32.616 -> Water_Temp.:96.46, PWM:29
22:28:33.413 -> Water_Temp.:96.46, PWM:29
22:28:34.173 -> Water_Temp.:96.46, PWM:28
22:28:35.755 -> Water_Temp.:96.46, PWM:28
22:28:36.514 -> Water_Temp.:96.46, PWM:29
22:28:37.310 -> Water_Temp.:96.46, PWM:29
22:28:38.070 -> Water_Temp.:96.57, PWM:29
22:28:38.863 -> Water_Temp.:96.46, PWM:28
22:28:39.655 -> Water_Temp.:96.46, PWM:29
22:28:40.417 -> Water_Temp.:96.57, PWM:29
22:28:41.213 -> Water_Temp.:96.57, PWM:29
22:28:41.970 -> Water_Temp.:96.57, PWM:28
22:28:42.769 -> Water_Temp.:96.57, PWM:29
22:28:43.530 -> Water_Temp.:96.46, PWM:29
22:28:44.327 -> Water_Temp.:96.57, PWM:29
22:28:45.087 -> Water_Temp.:96.57, PWM:29
22:28:45.883 -> Water_Temp.:96.57, PWM:29
22:28:46.645 -> Water_Temp.:96.46, PWM:29
22:28:47.435 -> Water_Temp.:96.46, PWM:29
22:28:48.991 -> Water_Temp.:96.46, PWM:29
22:28:50.546 -> Water_Temp.:96.46, PWM:29
22:28:51.342 -> Water_Temp.:96.46, PWM:29
22:28:52.101 -> Water_Temp.:96.46, PWM:28
22:28:52.897 -> Water_Temp.:96.57, PWM:28
22:28:53.692 -> Water_Temp.:96.57, PWM:28
22:28:54.456 -> Water_Temp.:96.57, PWM:29
22:28:55.252 -> Water_Temp.:96.57, PWM:29
22:28:56.012 -> Water_Temp.:96.57, PWM:28
22:28:56.809 -> Water_Temp.:96.57, PWM:29
22:28:58.326 -> Water_Temp.:96.57, PWM:28
22:28:59.122 -> Water_Temp.:96.57, PWM:29
22:28:59.878 -> Water_Temp.:96.57, PWM:29
22:29:00.671 -> Water_Temp.:96.57, PWM:29
22:29:01.465 -> Water_Temp.:96.57, PWM:29
22:29:02.222 -> Water_Temp.:96.57, PWM:28
22:29:03.022 -> Water_Temp.:96.57, PWM:28
22:29:03.783 -> Water_Temp.:96.57, PWM:29
22:29:05.339 -> Water_Temp.:96.57, PWM:28
22:29:06.889 -> Water_Temp.:96.57, PWM:28
22:29:07.652 -> Water_Temp.:96.57, PWM:28
22:29:08.447 -> Water_Temp.:96.57, PWM:28
22:29:09.206 -> Water_Temp.:96.57, PWM:28
22:29:10.000 -> Water_Temp.:96.57, PWM:29
22:29:10.762 -> Water_Temp.:96.57, PWM:29
22:29:11.555 -> Water_Temp.:96.57, PWM:29
22:29:12.346 -> Water_Temp.:96.57, PWM:28
22:29:13.107 -> Water_Temp.:96.57, PWM:29
22:29:13.895 -> Water_Temp.:96.57, PWM:28
22:29:14.689 -> Water_Temp.:96.57, PWM:28
22:29:15.449 -> Water_Temp.:96.57, PWM:29
22:29:16.245 -> Water_Temp.:96.57, PWM:29
22:29:17.004 -> Water_Temp.:96.57, PWM:29
22:29:17.799 -> Water_Temp.:96.57, PWM:29
22:29:18.561 -> Water_Temp.:96.57, PWM:28
22:29:19.357 -> Water_Temp.:96.57, PWM:29
22:29:20.911 -> Water_Temp.:96.57, PWM:28
22:29:21.704 -> Water_Temp.:96.57, PWM:29
22:29:22.462 -> Water_Temp.:96.57, PWM:29
22:29:23.256 -> Water_Temp.:96.57, PWM:29
22:29:24.017 -> Water_Temp.:96.69, PWM:29
22:29:24.816 -> Water_Temp.:96.69, PWM:29
22:29:25.578 -> Water_Temp.:96.69, PWM:28
22:29:26.375 -> Water_Temp.:96.69, PWM:29
22:29:27.167 -> Water_Temp.:96.69, PWM:29
22:29:27.926 -> Water_Temp.:96.69, PWM:28
22:29:28.720 -> Water_Temp.:96.69, PWM:29
22:29:30.246 -> Water_Temp.:96.69, PWM:29
22:29:31.046 -> Water_Temp.:96.69, PWM:29
22:29:32.601 -> Water_Temp.:96.69, PWM:29
22:29:33.361 -> Water_Temp.:96.69, PWM:28
22:29:34.157 -> Water_Temp.:96.69, PWM:28
22:29:34.916 -> Water_Temp.:96.69, PWM:29
22:29:35.709 -> Water_Temp.:96.69, PWM:29
22:29:36.500 -> Water_Temp.:96.69, PWM:28
22:29:37.259 -> Water_Temp.:96.69, PWM:29
22:29:38.055 -> Water_Temp.:96.69, PWM:28
22:29:38.842 -> Water_Temp.:96.69, PWM:28
22:29:39.603 -> Water_Temp.:96.69, PWM:28
22:29:40.396 -> Water_Temp.:96.69, PWM:29
22:29:41.159 -> Water_Temp.:96.69, PWM:29
22:29:41.954 -> Water_Temp.:96.69, PWM:29
22:29:42.749 -> Water_Temp.:96.69, PWM:28
22:29:43.508 -> Water_Temp.:96.69, PWM:29
22:29:45.068 -> Water_Temp.:96.80, PWM:29
22:29:45.866 -> Water_Temp.:96.69, PWM:29
22:29:46.629 -> Water_Temp.:96.69, PWM:29
22:29:47.426 -> Water_Temp.:96.69, PWM:29
22:29:48.187 -> Water_Temp.:96.69, PWM:29
22:29:49.741 -> Water_Temp.:96.80, PWM:28
22:29:50.536 -> Water_Temp.:96.80, PWM:28
22:29:51.299 -> Water_Temp.:96.80, PWM:28
22:29:52.098 -> Water_Temp.:96.80, PWM:29
22:29:52.862 -> Water_Temp.:96.80, PWM:29
22:29:53.658 -> Water_Temp.:96.69, PWM:28
22:29:55.181 -> Water_Temp.:96.69, PWM:29
22:29:55.975 -> Water_Temp.:96.69, PWM:29
22:29:56.768 -> Water_Temp.:96.69, PWM:28
22:29:57.529 -> Water_Temp.:96.69, PWM:29
22:29:58.320 -> Water_Temp.:96.69, PWM:29
22:29:59.081 -> Water_Temp.:96.69, PWM:29
22:29:59.878 -> Water_Temp.:96.69, PWM:28
22:30:00.636 -> Water_Temp.:96.69, PWM:29
22:30:01.436 -> Water_Temp.:96.69, PWM:29
22:30:02.227 -> Water_Temp.:96.69, PWM:28
22:30:02.986 -> Water_Temp.:96.69, PWM:28
22:30:03.781 -> Water_Temp.:96.69, PWM:29
22:30:04.540 -> Water_Temp.:96.69, PWM:28
22:30:05.335 -> Water_Temp.:96.69, PWM:29
22:30:06.123 -> Water_Temp.:96.69, PWM:29
22:30:06.880 -> Water_Temp.:96.69, PWM:28
22:30:07.676 -> Water_Temp.:96.69, PWM:29
22:30:08.435 -> Water_Temp.:96.69, PWM:29
22:30:09.230 -> Water_Temp.:96.69, PWM:28
22:30:10.024 -> Water_Temp.:96.69, PWM:28
22:30:10.786 -> Water_Temp.:96.69, PWM:28
22:30:11.582 -> Water_Temp.:96.69, PWM:28
22:30:12.342 -> Water_Temp.:96.69, PWM:28
22:30:13.139 -> Water_Temp.:96.69, PWM:28
22:30:13.899 -> Water_Temp.:96.69, PWM:28
22:30:14.700 -> Water_Temp.:96.69, PWM:29
22:30:15.464 -> Water_Temp.:96.69, PWM:29
22:30:16.256 -> Water_Temp.:96.69, PWM:29
22:30:17.016 -> Water_Temp.:96.69, PWM:28
22:30:17.812 -> Water_Temp.:96.69, PWM:29
22:30:18.607 -> Water_Temp.:96.69, PWM:29
22:30:19.367 -> Water_Temp.:96.69, PWM:29
22:30:20.163 -> Water_Temp.:96.69, PWM:28
22:30:20.923 -> Water_Temp.:96.69, PWM:28
22:30:21.718 -> Water_Temp.:96.80, PWM:29
22:30:22.478 -> Water_Temp.:96.80, PWM:28
22:30:24.062 -> Water_Temp.:96.80, PWM:29
22:30:24.823 -> Water_Temp.:96.80, PWM:29
22:30:25.615 -> Water_Temp.:96.80, PWM:29
22:30:26.373 -> Water_Temp.:96.80, PWM:29
22:30:27.165 -> Water_Temp.:96.80, PWM:29
22:30:27.961 -> Water_Temp.:96.80, PWM:28
22:30:28.718 -> Water_Temp.:96.80, PWM:29
22:30:29.513 -> Water_Temp.:96.80, PWM:28
22:30:30.275 -> Water_Temp.:96.80, PWM:29
22:30:31.829 -> Water_Temp.:96.80, PWM:29
22:30:32.622 -> Water_Temp.:96.80, PWM:28
22:30:33.382 -> Water_Temp.:96.80, PWM:29
22:30:34.176 -> Water_Temp.:96.80, PWM:29
22:30:34.969 -> Water_Temp.:96.80, PWM:29
22:30:35.728 -> Water_Temp.:96.80, PWM:29
22:30:36.528 -> Water_Temp.:96.80, PWM:29
22:30:37.287 -> Water_Temp.:96.80, PWM:28
22:30:38.080 -> Water_Temp.:96.80, PWM:28
22:30:38.842 -> Water_Temp.:96.80, PWM:29
22:30:39.636 -> Water_Temp.:96.80, PWM:29
22:30:41.159 -> Water_Temp.:96.80, PWM:29
22:30:41.950 -> Water_Temp.:96.91, PWM:28
22:30:42.711 -> Water_Temp.:96.91, PWM:29
22:30:43.507 -> Water_Temp.:96.91, PWM:29
22:30:44.301 -> Water_Temp.:96.91, PWM:28
22:30:45.066 -> Water_Temp.:96.91, PWM:28
22:30:45.859 -> Water_Temp.:96.91, PWM:29
22:30:46.620 -> Water_Temp.:96.91, PWM:29
22:30:47.414 -> Water_Temp.:96.91, PWM:28
22:30:48.172 -> Water_Temp.:96.91, PWM:28
22:30:48.967 -> Water_Temp.:96.91, PWM:29
22:30:50.524 -> Water_Temp.:96.91, PWM:29
22:30:51.321 -> Water_Temp.:96.91, PWM:29
22:30:52.083 -> Water_Temp.:96.91, PWM:29
22:30:52.872 -> Water_Temp.:96.91, PWM:29
22:30:53.629 -> Water_Temp.:96.91, PWM:28
22:30:54.421 -> Water_Temp.:96.91, PWM:29
22:30:55.215 -> Water_Temp.:96.91, PWM:29
22:30:55.978 -> Water_Temp.:96.91, PWM:28
22:30:56.770 -> Water_Temp.:96.91, PWM:29
22:30:57.558 -> Water_Temp.:96.91, PWM:29
22:30:58.318 -> Water_Temp.:96.91, PWM:28
22:30:59.113 -> Water_Temp.:96.91, PWM:28
22:30:59.874 -> Water_Temp.:96.91, PWM:28
22:31:00.672 -> Water_Temp.:96.91, PWM:29
22:31:01.431 -> Water_Temp.:96.91, PWM:29
22:31:02.228 -> Water_Temp.:96.91, PWM:28
22:31:03.020 -> Water_Temp.:96.91, PWM:29
22:31:03.780 -> Water_Temp.:96.91, PWM:28
22:31:04.577 -> Water_Temp.:96.91, PWM:28
22:31:05.339 -> Water_Temp.:96.91, PWM:29
22:31:06.133 -> Water_Temp.:96.91, PWM:29
22:31:06.896 -> Water_Temp.:96.91, PWM:29
22:31:07.694 -> Water_Temp.:96.91, PWM:29
22:31:08.457 -> Water_Temp.:97.02, PWM:29
22:31:09.252 -> Water_Temp.:96.91, PWM:29
22:31:10.013 -> Water_Temp.:97.02, PWM:28
22:31:10.808 -> Water_Temp.:97.02, PWM:28
22:31:11.572 -> Water_Temp.:97.02, PWM:29
22:31:12.369 -> Water_Temp.:96.91, PWM:28
22:31:13.161 -> Water_Temp.:96.91, PWM:29
22:31:13.925 -> Water_Temp.:97.02, PWM:29
22:31:14.722 -> Water_Temp.:97.02, PWM:29
22:31:15.485 -> Water_Temp.:96.91, PWM:29
22:31:16.279 -> Water_Temp.:97.02, PWM:28
22:31:17.036 -> Water_Temp.:97.02, PWM:29
22:31:17.833 -> Water_Temp.:97.02, PWM:29
22:31:18.594 -> Water_Temp.:97.02, PWM:28
22:31:19.393 -> Water_Temp.:97.02, PWM:29
22:31:20.154 -> Water_Temp.:97.02, PWM:28
22:31:20.952 -> Water_Temp.:97.02, PWM:28
22:31:21.709 -> Water_Temp.:97.02, PWM:28
22:31:22.502 -> Water_Temp.:97.02, PWM:29
22:31:23.295 -> Water_Temp.:97.02, PWM:28
22:31:24.060 -> Water_Temp.:97.02, PWM:29
22:31:24.856 -> Water_Temp.:97.02, PWM:29
22:31:25.618 -> Water_Temp.:97.02, PWM:29
22:31:27.175 -> Water_Temp.:97.02, PWM:29
22:31:27.937 -> Water_Temp.:97.02, PWM:29
22:31:28.703 -> Water_Temp.:97.02, PWM:29
22:31:29.494 -> Water_Temp.:97.14, PWM:29
22:31:31.044 -> Water_Temp.:97.02, PWM:28
22:31:31.839 -> Water_Temp.:97.14, PWM:29
22:31:32.602 -> Water_Temp.:97.14, PWM:29
22:31:33.398 -> Water_Temp.:97.14, PWM:29
22:31:34.195 -> Water_Temp.:97.14, PWM:29
22:31:34.958 -> Water_Temp.:97.14, PWM:29
22:31:35.718 -> Water_Temp.:97.14, PWM:28
22:31:36.515 -> Water_Temp.:97.14, PWM:29
22:31:37.312 -> Water_Temp.:97.14, PWM:29
22:31:38.072 -> Water_Temp.:97.14, PWM:29
22:31:38.865 -> Water_Temp.:97.14, PWM:28
22:31:39.626 -> Water_Temp.:97.14, PWM:28
22:31:40.423 -> Water_Temp.:97.14, PWM:29
22:31:41.184 -> Water_Temp.:97.14, PWM:29
22:31:41.977 -> Water_Temp.:97.14, PWM:28
22:31:42.740 -> Water_Temp.:97.14, PWM:29
22:31:43.530 -> Water_Temp.:97.14, PWM:29
22:31:44.322 -> Water_Temp.:97.14, PWM:29
22:31:45.081 -> Water_Temp.:97.14, PWM:29
22:31:46.670 -> Water_Temp.:97.14, PWM:29
22:31:47.429 -> Water_Temp.:97.14, PWM:28
22:31:48.227 -> Water_Temp.:97.14, PWM:29
22:31:48.989 -> Water_Temp.:97.14, PWM:29
22:31:49.786 -> Water_Temp.:97.14, PWM:29
22:31:50.543 -> Water_Temp.:97.14, PWM:29
22:31:51.337 -> Water_Temp.:97.14, PWM:29
22:31:52.102 -> Water_Temp.:97.14, PWM:29
22:31:52.895 -> Water_Temp.:97.14, PWM:28
22:31:53.691 -> Water_Temp.:97.14, PWM:28
22:31:54.448 -> Water_Temp.:97.14, PWM:29
22:31:56.008 -> Water_Temp.:97.14, PWM:29
22:31:57.565 -> Water_Temp.:97.14, PWM:28
22:31:58.327 -> Water_Temp.:97.14, PWM:28
22:31:59.124 -> Water_Temp.:97.14, PWM:28
22:31:59.887 -> Water_Temp.:97.14, PWM:28
22:32:00.651 -> Water_Temp.:97.14, PWM:28
22:32:01.445 -> Water_Temp.:97.14, PWM:28
22:32:02.236 -> Water_Temp.:97.14, PWM:28
22:32:02.997 -> Water_Temp.:97.14, PWM:29
22:32:03.794 -> Water_Temp.:97.14, PWM:29
22:32:04.557 -> Water_Temp.:97.14, PWM:29
22:32:05.350 -> Water_Temp.:97.14, PWM:28
22:32:06.111 -> Water_Temp.:97.14, PWM:29
22:32:06.903 -> Water_Temp.:97.14, PWM:29
22:32:07.692 -> Water_Temp.:97.14, PWM:28
22:32:08.457 -> Water_Temp.:97.25, PWM:29
22:32:09.250 -> Water_Temp.:97.25, PWM:29
22:32:10.045 -> Water_Temp.:97.25, PWM:28
22:32:10.808 -> Water_Temp.:97.25, PWM:29
22:32:11.598 -> Water_Temp.:97.25, PWM:29
22:32:13.152 -> Water_Temp.:97.25, PWM:29
22:32:13.912 -> Water_Temp.:97.25, PWM:29
22:32:14.711 -> Water_Temp.:97.25, PWM:29
22:32:15.477 -> Water_Temp.:97.25, PWM:29
22:32:16.269 -> Water_Temp.:97.25, PWM:29
22:32:17.029 -> Water_Temp.:97.25, PWM:29
22:32:17.825 -> Water_Temp.:97.25, PWM:28
22:32:18.621 -> Water_Temp.:97.25, PWM:28
22:32:19.382 -> Water_Temp.:97.25, PWM:28
22:32:20.177 -> Water_Temp.:97.25, PWM:29
22:32:20.941 -> Water_Temp.:97.25, PWM:29
22:32:21.736 -> Water_Temp.:97.25, PWM:29
22:32:22.496 -> Water_Temp.:97.25, PWM:29
22:32:23.290 -> Water_Temp.:97.25, PWM:29
22:32:24.054 -> Water_Temp.:97.25, PWM:29
22:32:24.850 -> Water_Temp.:97.25, PWM:29
22:32:25.611 -> Water_Temp.:97.25, PWM:29
22:32:26.407 -> Water_Temp.:97.25, PWM:29
22:32:27.171 -> Water_Temp.:97.25, PWM:28
22:32:28.757 -> Water_Temp.:97.25, PWM:29
22:32:29.519 -> Water_Temp.:97.25, PWM:29
22:32:30.310 -> Water_Temp.:97.25, PWM:28
22:32:31.074 -> Water_Temp.:97.25, PWM:28
22:32:31.865 -> Water_Temp.:97.25, PWM:29
22:32:32.656 -> Water_Temp.:97.25, PWM:28
22:32:33.417 -> Water_Temp.:97.25, PWM:28
22:32:34.212 -> Water_Temp.:97.25, PWM:29
22:32:34.975 -> Water_Temp.:97.36, PWM:28
22:32:35.770 -> Water_Temp.:97.36, PWM:29
22:32:37.323 -> Water_Temp.:97.36, PWM:29
22:32:38.115 -> Water_Temp.:97.36, PWM:28
22:32:38.876 -> Water_Temp.:97.36, PWM:29
22:32:39.668 -> Water_Temp.:97.25, PWM:28
22:32:40.428 -> Water_Temp.:97.36, PWM:28
22:32:41.221 -> Water_Temp.:97.36, PWM:29
22:32:42.013 -> Water_Temp.:97.36, PWM:29
22:32:42.770 -> Water_Temp.:97.36, PWM:29
22:32:43.564 -> Water_Temp.:97.36, PWM:28
22:32:44.360 -> Water_Temp.:97.36, PWM:29
22:32:45.122 -> Water_Temp.:97.36, PWM:29
22:32:45.884 -> Water_Temp.:97.36, PWM:29
22:32:46.682 -> Water_Temp.:97.36, PWM:28
22:32:47.444 -> Water_Temp.:97.36, PWM:29
22:32:48.241 -> Water_Temp.:97.36, PWM:29
22:32:49.003 -> Water_Temp.:97.36, PWM:28
22:32:50.593 -> Water_Temp.:97.36, PWM:29
22:32:51.354 -> Water_Temp.:97.36, PWM:29
22:32:52.150 -> Water_Temp.:97.36, PWM:29
22:32:52.908 -> Water_Temp.:97.36, PWM:28
22:32:53.701 -> Water_Temp.:97.36, PWM:28
22:32:54.496 -> Water_Temp.:97.36, PWM:28
22:32:55.256 -> Water_Temp.:97.36, PWM:28
22:32:56.047 -> Water_Temp.:97.36, PWM:29
22:32:57.568 -> Water_Temp.:97.36, PWM:28
22:32:58.356 -> Water_Temp.:97.36, PWM:28
22:32:59.152 -> Water_Temp.:97.47, PWM:29
22:32:59.909 -> Water_Temp.:97.47, PWM:28
22:33:01.469 -> Water_Temp.:97.47, PWM:28
22:33:02.230 -> Water_Temp.:97.47, PWM:29
22:33:03.783 -> Water_Temp.:97.47, PWM:29
22:33:04.573 -> Water_Temp.:97.47, PWM:29
22:33:05.330 -> Water_Temp.:97.47, PWM:29
22:33:06.128 -> Water_Temp.:97.47, PWM:28
22:33:07.676 -> Water_Temp.:97.47, PWM:29
22:33:08.437 -> Water_Temp.:97.47, PWM:29
22:33:09.235 -> Water_Temp.:97.47, PWM:29
22:33:10.028 -> Water_Temp.:97.47, PWM:29
22:33:10.786 -> Water_Temp.:97.47, PWM:29
22:33:11.581 -> Water_Temp.:97.47, PWM:29
22:33:12.336 -> Water_Temp.:97.47, PWM:28
22:33:13.131 -> Water_Temp.:97.47, PWM:29
22:33:13.924 -> Water_Temp.:97.47, PWM:28
22:33:14.683 -> Water_Temp.:97.47, PWM:29
22:33:15.482 -> Water_Temp.:97.47, PWM:29
22:33:17.002 -> Water_Temp.:97.47, PWM:29
22:33:17.797 -> Water_Temp.:97.47, PWM:29
22:33:18.560 -> Water_Temp.:97.47, PWM:29
22:33:19.356 -> Water_Temp.:97.47, PWM:29
22:33:20.118 -> Water_Temp.:97.47, PWM:29
22:33:20.912 -> Water_Temp.:97.47, PWM:28
22:33:21.670 -> Water_Temp.:97.47, PWM:28
22:33:22.462 -> Water_Temp.:97.47, PWM:29
22:33:23.258 -> Water_Temp.:97.47, PWM:28
22:33:24.020 -> Water_Temp.:97.47, PWM:29
22:33:24.816 -> Water_Temp.:97.36, PWM:28
22:33:25.574 -> Water_Temp.:97.36, PWM:29
22:33:26.373 -> Water_Temp.:97.36, PWM:28
22:33:27.135 -> Water_Temp.:97.47, PWM:28
22:33:27.925 -> Water_Temp.:99.05, PWM:29
22:33:28.719 -> Water_Temp.:101.07, PWM:28
22:33:29.479 -> Water_Temp.:98.71, PWM:28
22:33:30.274 -> Water_Temp.:98.26, PWM:28
22:33:31.039 -> Water_Temp.:98.15, PWM:29
22:33:31.832 -> Water_Temp.:97.92, PWM:29
22:33:32.592 -> Water_Temp.:97.92, PWM:29
22:33:33.390 -> Water_Temp.:97.81, PWM:28
22:33:34.148 -> Water_Temp.:97.81, PWM:28
22:33:34.947 -> Water_Temp.:97.70, PWM:29
22:33:35.706 -> Water_Temp.:97.70, PWM:28
22:33:36.502 -> Water_Temp.:97.70, PWM:28
22:33:37.290 -> Water_Temp.:97.59, PWM:28
22:33:38.047 -> Water_Temp.:97.59, PWM:29
22:33:38.840 -> Water_Temp.:97.59, PWM:29
22:33:39.634 -> Water_Temp.:97.59, PWM:29
22:33:40.392 -> Water_Temp.:97.59, PWM:28
22:33:41.184 -> Water_Temp.:97.59, PWM:29
22:33:41.977 -> Water_Temp.:97.59, PWM:28
22:33:42.738 -> Water_Temp.:97.59, PWM:29
22:33:44.287 -> Water_Temp.:97.47, PWM:29
22:33:45.811 -> Water_Temp.:97.47, PWM:29
22:33:46.604 -> Water_Temp.:97.47, PWM:28
22:33:47.400 -> Water_Temp.:97.47, PWM:28
22:33:48.162 -> Water_Temp.:97.47, PWM:28
22:33:48.955 -> Water_Temp.:97.47, PWM:29
22:33:49.716 -> Water_Temp.:97.47, PWM:28
22:33:50.509 -> Water_Temp.:97.47, PWM:29
22:33:51.270 -> Water_Temp.:97.47, PWM:28
22:33:52.065 -> Water_Temp.:97.47, PWM:29
22:33:52.862 -> Water_Temp.:97.47, PWM:29
22:33:53.622 -> Water_Temp.:97.47, PWM:29
22:33:54.413 -> Water_Temp.:97.47, PWM:29
22:33:55.204 -> Water_Temp.:97.47, PWM:28
22:33:55.965 -> Water_Temp.:97.47, PWM:28
22:33:56.761 -> Water_Temp.:97.47, PWM:28
22:33:57.517 -> Water_Temp.:97.47, PWM:29
22:33:58.312 -> Water_Temp.:97.47, PWM:29
22:33:59.077 -> Water_Temp.:97.47, PWM:29
22:33:59.869 -> Water_Temp.:97.47, PWM:29
22:34:00.631 -> Water_Temp.:97.47, PWM:28
22:34:01.429 -> Water_Temp.:97.47, PWM:29
22:34:02.219 -> Water_Temp.:97.47, PWM:29
22:34:02.983 -> Water_Temp.:97.47, PWM:29
22:34:03.777 -> Water_Temp.:97.47, PWM:29
22:34:05.334 -> Water_Temp.:97.47, PWM:29
22:34:06.095 -> Water_Temp.:97.47, PWM:29
22:34:06.889 -> Water_Temp.:97.47, PWM:28
22:34:07.649 -> Water_Temp.:97.47, PWM:29
22:34:09.200 -> Water_Temp.:97.47, PWM:28
22:34:09.994 -> Water_Temp.:97.47, PWM:28
22:34:10.761 -> Water_Temp.:97.47, PWM:29
22:34:11.555 -> Water_Temp.:97.47, PWM:28
22:34:12.315 -> Water_Temp.:97.47, PWM:29
22:34:13.107 -> Water_Temp.:97.47, PWM:29
22:34:13.902 -> Water_Temp.:97.59, PWM:28
22:34:14.663 -> Water_Temp.:97.59, PWM:29
22:34:15.460 -> Water_Temp.:97.59, PWM:29
22:34:16.219 -> Water_Temp.:97.59, PWM:29
22:34:17.015 -> Water_Temp.:97.59, PWM:29
22:34:17.775 -> Water_Temp.:97.59, PWM:29
22:34:18.572 -> Water_Temp.:97.59, PWM:28
22:34:20.121 -> Water_Temp.:97.59, PWM:29
22:34:20.886 -> Water_Temp.:97.59, PWM:28
22:34:21.676 -> Water_Temp.:97.59, PWM:29
22:34:22.437 -> Water_Temp.:97.59, PWM:28
22:34:23.234 -> Water_Temp.:97.59, PWM:28
22:34:23.993 -> Water_Temp.:97.59, PWM:29
22:34:24.788 -> Water_Temp.:97.59, PWM:29
22:34:25.579 -> Water_Temp.:97.59, PWM:29
22:34:26.341 -> Water_Temp.:97.59, PWM:28
22:34:27.131 -> Water_Temp.:97.59, PWM:29
22:34:27.922 -> Water_Temp.:97.70, PWM:28
22:34:28.685 -> Water_Temp.:97.70, PWM:29
22:34:29.476 -> Water_Temp.:97.70, PWM:29
22:34:30.236 -> Water_Temp.:97.70, PWM:29
22:34:31.032 -> Water_Temp.:97.70, PWM:28
22:34:31.824 -> Water_Temp.:97.70, PWM:28
22:34:32.582 -> Water_Temp.:97.70, PWM:29
22:34:33.377 -> Water_Temp.:97.70, PWM:29
22:34:34.136 -> Water_Temp.:97.70, PWM:28
22:34:34.931 -> Water_Temp.:97.70, PWM:28
22:34:35.725 -> Water_Temp.:97.70, PWM:29
22:34:36.485 -> Water_Temp.:97.70, PWM:29
22:34:37.280 -> Water_Temp.:97.70, PWM:28
22:34:38.040 -> Water_Temp.:97.70, PWM:29
22:34:38.835 -> Water_Temp.:97.70, PWM:28
22:34:39.594 -> Water_Temp.:97.70, PWM:29
22:34:40.386 -> Water_Temp.:97.70, PWM:29
22:34:41.183 -> Water_Temp.:97.70, PWM:29
22:34:41.945 -> Water_Temp.:97.70, PWM:29
22:34:42.744 -> Water_Temp.:97.70, PWM:29
22:34:43.505 -> Water_Temp.:97.70, PWM:28
22:34:45.063 -> Water_Temp.:97.70, PWM:29
22:34:45.818 -> Water_Temp.:97.70, PWM:28
22:34:46.607 -> Water_Temp.:97.70, PWM:29
22:34:47.404 -> Water_Temp.:97.70, PWM:29
22:34:48.958 -> Water_Temp.:97.70, PWM:29
22:34:49.716 -> Water_Temp.:97.70, PWM:29
22:34:50.511 -> Water_Temp.:97.70, PWM:28
22:34:52.036 -> Water_Temp.:97.70, PWM:28
22:34:52.836 -> Water_Temp.:97.70, PWM:28
22:34:53.626 -> Water_Temp.:97.70, PWM:29
22:34:54.387 -> Water_Temp.:97.59, PWM:28
22:34:55.183 -> Water_Temp.:97.70, PWM:29
22:34:55.943 -> Water_Temp.:97.70, PWM:28
22:34:56.733 -> Water_Temp.:97.59, PWM:29
22:34:57.527 -> Water_Temp.:97.59, PWM:29
22:34:59.843 -> Water_Temp.:97.59, PWM:28
22:35:00.603 -> Water_Temp.:97.59, PWM:29
22:35:01.398 -> Water_Temp.:97.59, PWM:29
22:35:02.190 -> Water_Temp.:97.59, PWM:29
22:35:03.710 -> Water_Temp.:97.70, PWM:28
22:35:04.503 -> Water_Temp.:97.70, PWM:29
22:35:05.295 -> Water_Temp.:97.70, PWM:29
22:35:06.054 -> Water_Temp.:97.70, PWM:29
22:35:06.852 -> Water_Temp.:97.70, PWM:29
22:35:07.610 -> Water_Temp.:97.70, PWM:29
22:35:08.401 -> Water_Temp.:97.70, PWM:28
22:35:09.161 -> Water_Temp.:97.70, PWM:29
22:35:10.718 -> Water_Temp.:97.70, PWM:28
22:35:11.510 -> Water_Temp.:97.70, PWM:28
22:35:12.300 -> Water_Temp.:97.70, PWM:28
22:35:13.062 -> Water_Temp.:97.70, PWM:29
22:35:13.856 -> Water_Temp.:97.70, PWM:28
22:35:14.649 -> Water_Temp.:97.70, PWM:28
22:35:15.412 -> Water_Temp.:97.70, PWM:28
22:35:16.208 -> Water_Temp.:97.70, PWM:29
22:35:16.970 -> Water_Temp.:97.70, PWM:29
22:35:17.765 -> Water_Temp.:97.70, PWM:29
22:35:18.521 -> Water_Temp.:97.70, PWM:28
22:35:19.317 -> Water_Temp.:97.70, PWM:28
22:35:20.076 -> Water_Temp.:97.70, PWM:29
22:35:20.869 -> Water_Temp.:97.70, PWM:28
22:35:21.662 -> Water_Temp.:97.70, PWM:29
22:35:23.214 -> Water_Temp.:97.70, PWM:28
22:35:24.009 -> Water_Temp.:97.70, PWM:28
22:35:24.764 -> Water_Temp.:97.70, PWM:29
22:35:25.558 -> Water_Temp.:97.81, PWM:29
22:35:26.350 -> Water_Temp.:97.81, PWM:29
22:35:27.109 -> Water_Temp.:97.81, PWM:29
22:35:27.899 -> Water_Temp.:97.81, PWM:29
22:35:28.660 -> Water_Temp.:97.81, PWM:29
22:35:29.454 -> Water_Temp.:97.81, PWM:29
22:35:30.250 -> Water_Temp.:97.81, PWM:29
22:35:31.011 -> Water_Temp.:97.81, PWM:29
22:35:31.803 -> Water_Temp.:97.81, PWM:28
22:35:32.562 -> Water_Temp.:97.81, PWM:29
22:35:33.358 -> Water_Temp.:97.81, PWM:29
22:35:34.118 -> Water_Temp.:97.81, PWM:29
22:35:34.912 -> Water_Temp.:97.81, PWM:29
22:35:35.703 -> Water_Temp.:97.81, PWM:29
22:35:36.462 -> Water_Temp.:97.81, PWM:29
22:35:37.256 -> Water_Temp.:97.81, PWM:28
22:35:38.016 -> Water_Temp.:97.81, PWM:29
22:35:38.804 -> Water_Temp.:97.81, PWM:29
22:35:39.598 -> Water_Temp.:97.81, PWM:28
22:35:40.389 -> Water_Temp.:97.81, PWM:29
22:35:41.150 -> Water_Temp.:97.81, PWM:28
22:35:41.942 -> Water_Temp.:97.81, PWM:29
22:35:42.702 -> Water_Temp.:97.81, PWM:28
22:35:43.492 -> Water_Temp.:97.81, PWM:29
22:35:45.049 -> Water_Temp.:97.81, PWM:29
22:35:45.810 -> Water_Temp.:97.81, PWM:29
22:35:46.606 -> Water_Temp.:97.81, PWM:28
22:35:47.366 -> Water_Temp.:97.81, PWM:29
22:35:48.160 -> Water_Temp.:97.92, PWM:28
22:35:48.920 -> Water_Temp.:97.92, PWM:28
22:35:49.717 -> Water_Temp.:97.81, PWM:29
22:35:50.477 -> Water_Temp.:97.92, PWM:29
22:35:51.271 -> Water_Temp.:97.92, PWM:29
22:35:52.066 -> Water_Temp.:97.92, PWM:28
22:35:52.826 -> Water_Temp.:97.92, PWM:28
22:35:53.620 -> Water_Temp.:97.92, PWM:28
22:35:54.382 -> Water_Temp.:97.92, PWM:28
22:35:55.178 -> Water_Temp.:97.92, PWM:29
22:35:55.934 -> Water_Temp.:97.92, PWM:29
22:35:56.730 -> Water_Temp.:97.92, PWM:29
22:35:57.524 -> Water_Temp.:97.92, PWM:29
22:35:58.287 -> Water_Temp.:97.92, PWM:29
22:35:59.083 -> Water_Temp.:97.92, PWM:29
22:36:00.642 -> Water_Temp.:97.92, PWM:29
22:36:02.202 -> Water_Temp.:97.92, PWM:29
22:36:02.962 -> Water_Temp.:97.92, PWM:29
22:36:03.761 -> Water_Temp.:97.92, PWM:29
22:36:04.520 -> Water_Temp.:97.92, PWM:28
22:36:05.313 -> Water_Temp.:97.92, PWM:28
22:36:06.073 -> Water_Temp.:97.92, PWM:28
22:36:06.868 -> Water_Temp.:97.92, PWM:28
22:36:07.660 -> Water_Temp.:97.92, PWM:29
22:36:08.421 -> Water_Temp.:97.92, PWM:29
22:36:09.218 -> Water_Temp.:97.92, PWM:29
22:36:09.976 -> Water_Temp.:97.92, PWM:29
22:36:10.769 -> Water_Temp.:97.92, PWM:29
22:36:13.082 -> Water_Temp.:97.92, PWM:29
22:36:13.844 -> Water_Temp.:97.92, PWM:29
22:36:14.637 -> Water_Temp.:97.92, PWM:28
22:36:15.430 -> Water_Temp.:97.92, PWM:29
22:36:16.191 -> Water_Temp.:97.92, PWM:29
22:36:16.980 -> Water_Temp.:97.92, PWM:28
22:36:17.742 -> Water_Temp.:97.92, PWM:29
22:36:18.541 -> Water_Temp.:97.92, PWM:28
22:36:19.297 -> Water_Temp.:97.92, PWM:29
22:36:20.098 -> Water_Temp.:97.92, PWM:29
22:36:20.890 -> Water_Temp.:97.92, PWM:28
22:36:21.656 -> Water_Temp.:97.92, PWM:28
22:36:22.452 -> Water_Temp.:97.92, PWM:29
22:36:23.211 -> Water_Temp.:97.92, PWM:28
22:36:24.006 -> Water_Temp.:97.92, PWM:28
22:36:24.766 -> Water_Temp.:97.92, PWM:29
22:36:25.565 -> Water_Temp.:97.92, PWM:28
22:36:26.324 -> Water_Temp.:97.92, PWM:28
22:36:27.118 -> Water_Temp.:97.92, PWM:29
22:36:27.880 -> Water_Temp.:97.92, PWM:29
22:36:28.677 -> Water_Temp.:97.92, PWM:29
22:36:29.439 -> Water_Temp.:97.92, PWM:28
22:36:31.025 -> Water_Temp.:99.50, PWM:29
22:36:31.782 -> Water_Temp.:101.52, PWM:28
22:36:32.577 -> Water_Temp.:99.16, PWM:28
22:36:33.341 -> Water_Temp.:98.82, PWM:28
22:36:34.136 -> Water_Temp.:98.60, PWM:29
22:36:34.929 -> Water_Temp.:98.49, PWM:29
22:36:35.689 -> Water_Temp.:98.37, PWM:29
22:36:36.483 -> Water_Temp.:98.26, PWM:28
22:36:37.992 -> Water_Temp.:98.15, PWM:28
22:36:38.784 -> Water_Temp.:98.15, PWM:29
22:36:39.580 -> Water_Temp.:98.15, PWM:29
22:36:40.342 -> Water_Temp.:98.15, PWM:28
22:36:41.138 -> Water_Temp.:98.15, PWM:29
22:36:41.899 -> Water_Temp.:98.15, PWM:29
22:36:42.692 -> Water_Temp.:98.04, PWM:28
22:36:43.452 -> Water_Temp.:98.04, PWM:28
22:36:44.250 -> Water_Temp.:98.04, PWM:29
22:36:45.805 -> Water_Temp.:98.04, PWM:28
22:36:46.595 -> Water_Temp.:98.04, PWM:28
22:36:47.356 -> Water_Temp.:98.04, PWM:29
22:36:48.152 -> Water_Temp.:98.04, PWM:29
22:36:48.913 -> Water_Temp.:98.04, PWM:28
22:36:49.707 -> Water_Temp.:98.04, PWM:28
22:36:50.467 -> Water_Temp.:98.04, PWM:29
22:36:51.266 -> Water_Temp.:98.04, PWM:29
22:36:52.026 -> Water_Temp.:98.04, PWM:28
22:36:52.822 -> Water_Temp.:97.92, PWM:28
22:36:53.616 -> Water_Temp.:97.92, PWM:28
22:36:54.378 -> Water_Temp.:97.92, PWM:29
22:36:55.172 -> Water_Temp.:97.92, PWM:29
22:36:55.932 -> Water_Temp.:97.92, PWM:29
22:36:57.489 -> Water_Temp.:97.92, PWM:29
22:36:58.287 -> Water_Temp.:97.92, PWM:28
22:36:59.051 -> Water_Temp.:97.92, PWM:28
22:36:59.841 -> Water_Temp.:97.92, PWM:28
22:37:00.635 -> Water_Temp.:97.92, PWM:29
22:37:01.399 -> Water_Temp.:97.92, PWM:28
22:37:02.189 -> Water_Temp.:97.92, PWM:29
22:37:02.949 -> Water_Temp.:97.92, PWM:28
22:37:03.741 -> Water_Temp.:97.92, PWM:28
22:37:05.294 -> Water_Temp.:97.92, PWM:29
22:37:06.094 -> Water_Temp.:97.92, PWM:29
22:37:06.848 -> Water_Temp.:97.92, PWM:29
22:37:07.645 -> Water_Temp.:97.92, PWM:29
22:37:08.440 -> Water_Temp.:98.04, PWM:29
22:37:09.198 -> Water_Temp.:98.04, PWM:29
22:37:09.994 -> Water_Temp.:97.92, PWM:29
22:37:10.753 -> Water_Temp.:98.04, PWM:29
22:37:12.309 -> Water_Temp.:98.04, PWM:29
22:37:13.103 -> Water_Temp.:98.04, PWM:29
22:37:13.864 -> Water_Temp.:98.04, PWM:28
22:37:14.662 -> Water_Temp.:98.04, PWM:28
22:37:15.423 -> Water_Temp.:98.04, PWM:29
22:37:16.221 -> Water_Temp.:98.04, PWM:29
22:37:16.986 -> Water_Temp.:98.04, PWM:29
22:37:17.778 -> Water_Temp.:98.04, PWM:29
22:37:18.573 -> Water_Temp.:98.04, PWM:28
22:37:19.334 -> Water_Temp.:98.04, PWM:29
22:37:20.131 -> Water_Temp.:98.04, PWM:28
22:37:20.898 -> Water_Temp.:98.04, PWM:29
22:37:22.452 -> Water_Temp.:98.04, PWM:28
22:37:23.246 -> Water_Temp.:98.04, PWM:28
22:37:24.009 -> Water_Temp.:98.04, PWM:28
22:37:24.810 -> Water_Temp.:98.04, PWM:28
22:37:25.573 -> Water_Temp.:98.04, PWM:28
22:37:27.135 -> Water_Temp.:98.04, PWM:28
22:37:27.894 -> Water_Temp.:97.92, PWM:29
22:37:28.694 -> Water_Temp.:97.92, PWM:29
22:37:29.454 -> Water_Temp.:97.92, PWM:29
22:37:30.980 -> Water_Temp.:98.04, PWM:28
22:37:31.777 -> Water_Temp.:98.04, PWM:28
22:37:34.056 -> Water_Temp.:98.04, PWM:28
22:37:34.845 -> Water_Temp.:97.92, PWM:28
22:37:35.641 -> Water_Temp.:98.04, PWM:29
22:37:36.395 -> Water_Temp.:98.04, PWM:29
22:37:37.951 -> Water_Temp.:98.04, PWM:29
22:37:38.745 -> Water_Temp.:98.04, PWM:28
22:37:39.508 -> Water_Temp.:97.92, PWM:29
22:37:41.068 -> Water_Temp.:98.04, PWM:29
22:37:41.826 -> Water_Temp.:98.04, PWM:29
22:37:42.622 -> Water_Temp.:98.04, PWM:28
22:37:43.414 -> Water_Temp.:98.04, PWM:29
22:37:44.176 -> Water_Temp.:98.04, PWM:29
22:37:44.973 -> Water_Temp.:98.04, PWM:29
22:37:45.734 -> Water_Temp.:98.04, PWM:29
22:37:46.527 -> Water_Temp.:98.04, PWM:29
22:37:47.284 -> Water_Temp.:98.04, PWM:29
22:37:48.078 -> Water_Temp.:98.04, PWM:28
22:37:48.876 -> Water_Temp.:98.04, PWM:28
22:37:49.634 -> Water_Temp.:98.04, PWM:29
22:37:50.430 -> Water_Temp.:98.04, PWM:29
22:37:51.942 -> Water_Temp.:98.04, PWM:29
22:37:52.740 -> Water_Temp.:98.04, PWM:29
22:37:53.501 -> Water_Temp.:98.04, PWM:29
22:37:54.301 -> Water_Temp.:98.04, PWM:29
22:37:55.066 -> Water_Temp.:98.15, PWM:28
22:37:55.858 -> Water_Temp.:98.15, PWM:29
22:37:56.650 -> Water_Temp.:98.04, PWM:28
22:37:57.413 -> Water_Temp.:98.15, PWM:28
22:37:58.965 -> Water_Temp.:98.15, PWM:29
22:37:59.760 -> Water_Temp.:98.15, PWM:28
22:38:00.522 -> Water_Temp.:98.15, PWM:29
22:38:01.320 -> Water_Temp.:98.15, PWM:29
22:38:02.082 -> Water_Temp.:98.15, PWM:29
22:38:02.876 -> Water_Temp.:98.15, PWM:29
22:38:03.637 -> Water_Temp.:98.15, PWM:29
22:38:04.434 -> Water_Temp.:98.15, PWM:28
22:38:05.230 -> Water_Temp.:98.15, PWM:29
22:38:05.990 -> Water_Temp.:98.15, PWM:29
22:38:06.787 -> Water_Temp.:98.15, PWM:28
22:38:07.548 -> Water_Temp.:98.15, PWM:28
22:38:08.346 -> Water_Temp.:98.15, PWM:29
22:38:09.109 -> Water_Temp.:98.15, PWM:29
22:38:09.904 -> Water_Temp.:98.15, PWM:28
22:38:10.667 -> Water_Temp.:98.15, PWM:28
22:38:11.464 -> Water_Temp.:98.15, PWM:29
22:38:12.222 -> Water_Temp.:98.15, PWM:28
22:38:13.014 -> Water_Temp.:98.15, PWM:28
22:38:13.805 -> Water_Temp.:98.15, PWM:29
22:38:14.567 -> Water_Temp.:98.15, PWM:28
22:38:15.365 -> Water_Temp.:98.15, PWM:29
22:38:16.124 -> Water_Temp.:98.15, PWM:29
22:38:16.920 -> Water_Temp.:98.15, PWM:29
22:38:17.674 -> Water_Temp.:98.15, PWM:28
22:38:18.471 -> Water_Temp.:98.15, PWM:28
22:38:19.267 -> Water_Temp.:98.15, PWM:28
22:38:20.026 -> Water_Temp.:98.15, PWM:29
22:38:20.825 -> Water_Temp.:98.15, PWM:28
22:38:21.588 -> Water_Temp.:98.15, PWM:29
22:38:22.379 -> Water_Temp.:98.15, PWM:28
22:38:23.138 -> Water_Temp.:98.15, PWM:29
22:38:23.935 -> Water_Temp.:98.15, PWM:29
22:38:24.698 -> Water_Temp.:98.15, PWM:28
22:38:25.489 -> Water_Temp.:98.15, PWM:28
22:38:26.280 -> Water_Temp.:98.15, PWM:28
22:38:27.037 -> Water_Temp.:98.15, PWM:29
22:38:27.833 -> Water_Temp.:98.15, PWM:29
22:38:28.595 -> Water_Temp.:98.15, PWM:28
22:38:29.390 -> Water_Temp.:98.15, PWM:29
22:38:30.184 -> Water_Temp.:98.15, PWM:28
22:38:31.705 -> Water_Temp.:98.15, PWM:29
22:38:33.260 -> Water_Temp.:98.15, PWM:28
22:38:34.053 -> Water_Temp.:98.15, PWM:28
22:38:34.814 -> Water_Temp.:98.15, PWM:28
22:38:35.612 -> Water_Temp.:98.15, PWM:29
22:38:36.368 -> Water_Temp.:98.15, PWM:28
22:38:37.164 -> Water_Temp.:98.15, PWM:29
22:38:37.928 -> Water_Temp.:98.15, PWM:28
22:38:38.724 -> Water_Temp.:98.15, PWM:28
22:38:39.521 -> Water_Temp.:98.15, PWM:28
22:38:40.282 -> Water_Temp.:98.15, PWM:29
22:38:41.080 -> Water_Temp.:98.15, PWM:29
22:38:41.838 -> Water_Temp.:98.15, PWM:28
22:38:42.632 -> Water_Temp.:98.15, PWM:28
22:38:43.390 -> Water_Temp.:98.15, PWM:28
22:38:44.181 -> Water_Temp.:98.15, PWM:28
22:38:44.977 -> Water_Temp.:98.15, PWM:29
22:38:45.743 -> Water_Temp.:98.15, PWM:28
22:38:46.536 -> Water_Temp.:98.15, PWM:29
22:38:48.085 -> Water_Temp.:98.15, PWM:29
22:38:48.849 -> Water_Temp.:98.15, PWM:28
22:38:49.645 -> Water_Temp.:98.15, PWM:29
22:38:50.406 -> Water_Temp.:98.15, PWM:29
22:38:51.202 -> Water_Temp.:98.15, PWM:29
22:38:51.962 -> Water_Temp.:98.15, PWM:29
22:38:52.758 -> Water_Temp.:98.15, PWM:29
22:38:53.517 -> Water_Temp.:98.15, PWM:28
22:38:54.319 -> Water_Temp.:98.15, PWM:29
22:38:55.110 -> Water_Temp.:98.15, PWM:28
22:38:55.873 -> Water_Temp.:98.15, PWM:29
22:38:56.666 -> Water_Temp.:98.15, PWM:29
22:38:57.427 -> Water_Temp.:98.15, PWM:28
22:38:58.221 -> Water_Temp.:98.15, PWM:29
22:38:58.981 -> Water_Temp.:98.15, PWM:29
22:38:59.774 -> Water_Temp.:98.15, PWM:28
22:39:00.532 -> Water_Temp.:98.15, PWM:28
22:39:01.326 -> Water_Temp.:98.15, PWM:28
22:39:02.121 -> Water_Temp.:98.15, PWM:29
22:39:02.878 -> Water_Temp.:98.15, PWM:28
22:39:04.435 -> Water_Temp.:98.15, PWM:29
22:39:05.235 -> Water_Temp.:98.15, PWM:28
22:39:05.996 -> Water_Temp.:98.15, PWM:29
22:39:06.791 -> Water_Temp.:98.15, PWM:28
22:39:08.344 -> Water_Temp.:98.15, PWM:28
22:39:09.142 -> Water_Temp.:98.15, PWM:29
22:39:09.904 -> Water_Temp.:98.15, PWM:28
22:39:10.698 -> Water_Temp.:98.15, PWM:28
22:39:11.459 -> Water_Temp.:98.26, PWM:29
22:39:12.257 -> Water_Temp.:98.26, PWM:29
22:39:13.015 -> Water_Temp.:98.26, PWM:28
22:39:13.808 -> Water_Temp.:98.15, PWM:29
22:39:14.574 -> Water_Temp.:98.15, PWM:28
22:39:15.366 -> Water_Temp.:98.26, PWM:29
22:39:16.161 -> Water_Temp.:98.26, PWM:29
22:39:16.919 -> Water_Temp.:98.26, PWM:29
22:39:17.716 -> Water_Temp.:98.26, PWM:28
22:39:18.478 -> Water_Temp.:98.26, PWM:29
22:39:20.032 -> Water_Temp.:98.26, PWM:28
22:39:20.796 -> Water_Temp.:98.26, PWM:29
22:39:21.591 -> Water_Temp.:98.15, PWM:28
22:39:22.351 -> Water_Temp.:98.15, PWM:28
22:39:23.145 -> Water_Temp.:98.15, PWM:28
22:39:23.904 -> Water_Temp.:98.15, PWM:29
22:39:24.697 -> Water_Temp.:98.15, PWM:28
22:39:25.493 -> Water_Temp.:98.15, PWM:29
22:39:26.255 -> Water_Temp.:98.15, PWM:28
22:39:27.048 -> Water_Temp.:98.15, PWM:28
22:39:27.811 -> Water_Temp.:99.84, PWM:28
22:39:28.607 -> Water_Temp.:101.86, PWM:28
22:39:29.364 -> Water_Temp.:99.50, PWM:29
22:39:30.162 -> Water_Temp.:99.05, PWM:29
22:39:31.725 -> Water_Temp.:98.71, PWM:28
22:39:32.486 -> Water_Temp.:98.60, PWM:28
22:39:33.275 -> Water_Temp.:98.60, PWM:28
22:39:34.793 -> Water_Temp.:98.49, PWM:28
22:39:35.588 -> Water_Temp.:98.49, PWM:29
22:39:36.352 -> Water_Temp.:98.37, PWM:29
22:39:37.148 -> Water_Temp.:98.37, PWM:29
22:39:37.910 -> Water_Temp.:98.37, PWM:29
22:39:40.229 -> Water_Temp.:98.37, PWM:28
22:39:40.992 -> Water_Temp.:98.37, PWM:29
22:39:41.790 -> Water_Temp.:98.37, PWM:29
22:39:42.551 -> Water_Temp.:98.37, PWM:29
22:39:43.344 -> Water_Temp.:98.37, PWM:29
22:39:44.101 -> Water_Temp.:98.37, PWM:28
22:39:44.904 -> Water_Temp.:98.37, PWM:29
22:39:45.668 -> Water_Temp.:98.37, PWM:29
22:39:46.460 -> Water_Temp.:98.37, PWM:28
22:39:47.225 -> Water_Temp.:98.37, PWM:28
22:39:48.022 -> Water_Temp.:98.37, PWM:28
22:39:48.781 -> Water_Temp.:98.37, PWM:28
22:39:49.574 -> Water_Temp.:98.37, PWM:29
22:39:50.369 -> Water_Temp.:98.26, PWM:28
22:39:51.132 -> Water_Temp.:98.37, PWM:29
22:39:51.921 -> Water_Temp.:98.26, PWM:29
22:39:52.680 -> Water_Temp.:98.37, PWM:29
22:39:53.475 -> Water_Temp.:98.26, PWM:29
22:39:54.236 -> Water_Temp.:98.26, PWM:28
22:39:55.035 -> Water_Temp.:98.26, PWM:28
22:39:55.829 -> Water_Temp.:98.26, PWM:29
22:39:56.589 -> Water_Temp.:98.26, PWM:29
22:39:57.384 -> Water_Temp.:98.26, PWM:28
22:39:58.146 -> Water_Temp.:98.26, PWM:29
22:39:58.941 -> Water_Temp.:98.26, PWM:29
22:39:59.698 -> Water_Temp.:98.26, PWM:28
22:40:00.494 -> Water_Temp.:98.26, PWM:28
22:40:02.016 -> Water_Temp.:98.26, PWM:29
22:40:02.809 -> Water_Temp.:98.26, PWM:28
22:40:03.571 -> Water_Temp.:98.26, PWM:28
22:40:04.360 -> Water_Temp.:98.26, PWM:28
22:40:05.155 -> Water_Temp.:98.26, PWM:28
22:40:05.913 -> Water_Temp.:98.26, PWM:28
22:40:06.705 -> Water_Temp.:98.26, PWM:29
22:40:07.497 -> Water_Temp.:98.26, PWM:29
22:40:08.253 -> Water_Temp.:98.26, PWM:29
22:40:09.049 -> Water_Temp.:98.26, PWM:29
22:40:09.845 -> Water_Temp.:98.26, PWM:28
22:40:10.607 -> Water_Temp.:98.26, PWM:29
22:40:11.372 -> Water_Temp.:98.26, PWM:28
22:40:12.164 -> Water_Temp.:98.26, PWM:28
22:40:12.957 -> Water_Temp.:98.26, PWM:29
22:40:13.720 -> Water_Temp.:98.26, PWM:29
22:40:14.513 -> Water_Temp.:98.26, PWM:29
22:40:15.273 -> Water_Temp.:98.26, PWM:29
22:40:16.068 -> Water_Temp.:98.37, PWM:28
22:40:16.858 -> Water_Temp.:98.37, PWM:29
22:40:17.618 -> Water_Temp.:98.37, PWM:29
22:40:18.410 -> Water_Temp.:98.37, PWM:28
22:40:19.172 -> Water_Temp.:98.37, PWM:28
22:40:20.726 -> Water_Temp.:98.37, PWM:29
22:40:21.486 -> Water_Temp.:98.37, PWM:29
22:40:22.280 -> Water_Temp.:98.37, PWM:29
22:40:23.042 -> Water_Temp.:98.37, PWM:28
22:40:23.836 -> Water_Temp.:98.37, PWM:29
22:40:24.633 -> Water_Temp.:98.37, PWM:28
22:40:25.395 -> Water_Temp.:98.37, PWM:29
22:40:26.950 -> Water_Temp.:98.37, PWM:28
22:40:27.744 -> Water_Temp.:98.37, PWM:29
22:40:28.502 -> Water_Temp.:98.37, PWM:29
22:40:29.298 -> Water_Temp.:98.37, PWM:29
22:40:30.096 -> Water_Temp.:98.37, PWM:29
22:40:30.862 -> Water_Temp.:98.37, PWM:29
22:40:31.623 -> Water_Temp.:98.37, PWM:28
22:40:32.417 -> Water_Temp.:98.37, PWM:29
22:40:33.211 -> Water_Temp.:98.37, PWM:29
22:40:33.970 -> Water_Temp.:98.37, PWM:28
22:40:34.765 -> Water_Temp.:98.37, PWM:28
22:40:35.527 -> Water_Temp.:98.37, PWM:29
22:40:36.320 -> Water_Temp.:98.37, PWM:28
22:40:37.114 -> Water_Temp.:98.37, PWM:28
22:40:37.878 -> Water_Temp.:98.37, PWM:29
22:40:38.673 -> Water_Temp.:98.37, PWM:28
22:40:39.433 -> Water_Temp.:98.37, PWM:28
22:40:40.226 -> Water_Temp.:98.37, PWM:28
22:40:40.989 -> Water_Temp.:98.37, PWM:29
22:40:41.785 -> Water_Temp.:98.37, PWM:29
22:40:42.550 -> Water_Temp.:98.37, PWM:29
22:40:43.352 -> Water_Temp.:98.37, PWM:29
22:40:44.116 -> Water_Temp.:98.37, PWM:28
22:40:44.911 -> Water_Temp.:98.37, PWM:28
22:40:45.674 -> Water_Temp.:98.37, PWM:29
22:40:46.464 -> Water_Temp.:98.37, PWM:29
22:40:47.224 -> Water_Temp.:98.37, PWM:28
22:40:48.015 -> Water_Temp.:98.37, PWM:28
22:40:48.812 -> Water_Temp.:98.37, PWM:28
22:40:49.566 -> Water_Temp.:98.37, PWM:29
22:40:50.361 -> Water_Temp.:98.37, PWM:29
22:40:51.122 -> Water_Temp.:98.37, PWM:28
22:40:52.681 -> Water_Temp.:98.37, PWM:28
22:40:53.443 -> Water_Temp.:98.37, PWM:28
22:40:54.238 -> Water_Temp.:98.37, PWM:29
22:40:55.005 -> Water_Temp.:98.37, PWM:29
22:40:55.800 -> Water_Temp.:98.37, PWM:28
22:40:56.592 -> Water_Temp.:98.37, PWM:28
22:40:57.351 -> Water_Temp.:98.37, PWM:29
22:40:58.146 -> Water_Temp.:98.37, PWM:29
22:40:58.905 -> Water_Temp.:98.37, PWM:28
22:40:59.694 -> Water_Temp.:98.37, PWM:28
22:41:00.491 -> Water_Temp.:98.37, PWM:29
22:41:01.255 -> Water_Temp.:98.37, PWM:29
22:41:02.046 -> Water_Temp.:98.37, PWM:29
22:41:02.804 -> Water_Temp.:98.37, PWM:28
22:41:03.598 -> Water_Temp.:98.37, PWM:29
22:41:04.393 -> Water_Temp.:98.37, PWM:29
22:41:05.155 -> Water_Temp.:98.37, PWM:29
22:41:05.951 -> Water_Temp.:98.37, PWM:28
22:41:06.712 -> Water_Temp.:98.37, PWM:29
22:41:07.511 -> Water_Temp.:98.37, PWM:29
22:41:08.273 -> Water_Temp.:98.37, PWM:29
22:41:09.071 -> Water_Temp.:98.37, PWM:29
22:41:09.829 -> Water_Temp.:98.37, PWM:29
22:41:10.624 -> Water_Temp.:98.49, PWM:29
22:41:11.383 -> Water_Temp.:98.37, PWM:29
22:41:12.180 -> Water_Temp.:98.37, PWM:29
22:41:12.941 -> Water_Temp.:98.37, PWM:29
22:41:13.734 -> Water_Temp.:98.37, PWM:29
22:41:14.528 -> Water_Temp.:98.37, PWM:28
22:41:15.289 -> Water_Temp.:98.37, PWM:29
22:41:16.082 -> Water_Temp.:98.37, PWM:29
22:41:16.842 -> Water_Temp.:98.37, PWM:28
22:41:17.630 -> Water_Temp.:98.37, PWM:28
22:41:18.427 -> Water_Temp.:98.37, PWM:29
22:41:19.981 -> Water_Temp.:98.37, PWM:29
22:41:20.744 -> Water_Temp.:98.37, PWM:29
22:41:21.539 -> Water_Temp.:98.37, PWM:28
22:41:22.300 -> Water_Temp.:98.37, PWM:28
22:41:23.094 -> Water_Temp.:98.37, PWM:28
22:41:23.888 -> Water_Temp.:98.37, PWM:29
22:41:24.654 -> Water_Temp.:98.37, PWM:29
22:41:26.209 -> Water_Temp.:98.37, PWM:29
22:41:27.005 -> Water_Temp.:98.37, PWM:29
22:41:27.765 -> Water_Temp.:98.37, PWM:29
22:41:28.559 -> Water_Temp.:98.37, PWM:28
22:41:29.323 -> Water_Temp.:98.37, PWM:28
22:41:30.118 -> Water_Temp.:98.37, PWM:29
22:41:30.881 -> Water_Temp.:98.37, PWM:28
22:41:31.673 -> Water_Temp.:98.37, PWM:28
22:41:32.467 -> Water_Temp.:98.37, PWM:29
22:41:33.224 -> Water_Temp.:98.37, PWM:28
22:41:34.018 -> Water_Temp.:98.37, PWM:29
22:41:34.780 -> Water_Temp.:98.37, PWM:28
22:41:35.572 -> Water_Temp.:98.37, PWM:29
22:41:36.367 -> Water_Temp.:98.37, PWM:29
22:41:37.129 -> Water_Temp.:98.37, PWM:29
22:41:37.924 -> Water_Temp.:98.37, PWM:28
22:41:38.683 -> Water_Temp.:98.37, PWM:28
22:41:39.479 -> Water_Temp.:98.37, PWM:31
22:41:40.241 -> Water_Temp.:98.37, PWM:31
22:41:41.041 -> Water_Temp.:98.37, PWM:31
22:41:41.802 -> Water_Temp.:98.37, PWM:30
22:41:42.594 -> Water_Temp.:98.37, PWM:30
22:41:43.382 -> Water_Temp.:98.37, PWM:30
22:41:44.145 -> Water_Temp.:98.49, PWM:30
22:41:44.944 -> Water_Temp.:98.49, PWM:31
22:41:45.703 -> Water_Temp.:98.37, PWM:30
22:41:46.499 -> Water_Temp.:98.49, PWM:30
22:41:47.257 -> Water_Temp.:98.49, PWM:30
22:41:48.051 -> Water_Temp.:98.49, PWM:31
22:41:48.845 -> Water_Temp.:98.49, PWM:30
22:41:49.607 -> Water_Temp.:98.49, PWM:30
22:41:50.402 -> Water_Temp.:98.49, PWM:31
22:41:51.926 -> Water_Temp.:98.49, PWM:31
22:41:52.718 -> Water_Temp.:98.49, PWM:31
22:41:53.481 -> Water_Temp.:98.49, PWM:30
22:41:54.278 -> Water_Temp.:98.49, PWM:31
22:41:55.041 -> Water_Temp.:98.49, PWM:31
22:41:55.832 -> Water_Temp.:98.49, PWM:31
22:41:56.619 -> Water_Temp.:98.49, PWM:30
22:41:57.383 -> Water_Temp.:98.49, PWM:30
22:41:58.174 -> Water_Temp.:98.49, PWM:31
22:41:58.970 -> Water_Temp.:98.49, PWM:31
22:41:59.728 -> Water_Temp.:98.49, PWM:31
22:42:00.522 -> Water_Temp.:98.49, PWM:31
22:42:01.282 -> Water_Temp.:98.49, PWM:31
22:42:02.074 -> Water_Temp.:98.49, PWM:31
22:42:02.864 -> Water_Temp.:98.49, PWM:31
22:42:03.624 -> Water_Temp.:98.49, PWM:31
22:42:04.416 -> Water_Temp.:98.49, PWM:31
22:42:05.177 -> Water_Temp.:98.49, PWM:31
22:42:05.973 -> Water_Temp.:98.49, PWM:31
22:42:06.768 -> Water_Temp.:98.49, PWM:31
22:42:07.527 -> Water_Temp.:98.49, PWM:31
22:42:08.320 -> Water_Temp.:98.49, PWM:31
22:42:09.082 -> Water_Temp.:98.49, PWM:31
22:42:09.875 -> Water_Temp.:98.49, PWM:31
22:42:10.638 -> Water_Temp.:98.49, PWM:31
22:42:11.429 -> Water_Temp.:98.49, PWM:31
22:42:12.221 -> Water_Temp.:98.49, PWM:31
22:42:12.982 -> Water_Temp.:98.49, PWM:31
22:42:13.776 -> Water_Temp.:98.49, PWM:31
22:42:14.539 -> Water_Temp.:98.49, PWM:31
22:42:15.336 -> Water_Temp.:98.49, PWM:31
22:42:16.100 -> Water_Temp.:98.49, PWM:31
22:42:16.893 -> Water_Temp.:98.49, PWM:31
22:42:17.687 -> Water_Temp.:98.49, PWM:31
22:42:18.448 -> Water_Temp.:98.49, PWM:31
22:42:19.245 -> Water_Temp.:98.49, PWM:31
22:42:20.004 -> Water_Temp.:98.49, PWM:31
22:42:20.800 -> Water_Temp.:98.49, PWM:31
22:42:21.561 -> Water_Temp.:98.49, PWM:31
22:42:22.360 -> Water_Temp.:98.49, PWM:31
22:42:23.121 -> Water_Temp.:98.49, PWM:31
22:42:23.917 -> Water_Temp.:100.06, PWM:31
22:42:24.680 -> Water_Temp.:102.20, PWM:31
22:42:25.476 -> Water_Temp.:99.84, PWM:31
22:42:26.234 -> Water_Temp.:99.39, PWM:31
22:42:27.029 -> Water_Temp.:99.16, PWM:31
22:42:27.819 -> Water_Temp.:99.05, PWM:31
22:42:28.579 -> Water_Temp.:98.94, PWM:31
22:42:29.375 -> Water_Temp.:98.94, PWM:31
22:42:30.139 -> Water_Temp.:98.82, PWM:30
22:42:30.936 -> Water_Temp.:98.82, PWM:31
22:42:31.727 -> Water_Temp.:98.71, PWM:31
22:42:32.490 -> Water_Temp.:98.71, PWM:31
22:42:33.285 -> Water_Temp.:98.71, PWM:31
22:42:34.044 -> Water_Temp.:98.71, PWM:31
22:42:34.841 -> Water_Temp.:98.71, PWM:31
22:42:35.603 -> Water_Temp.:98.71, PWM:31
22:42:36.396 -> Water_Temp.:98.71, PWM:31
22:42:37.155 -> Water_Temp.:98.71, PWM:31
22:42:37.949 -> Water_Temp.:98.71, PWM:31
22:42:39.501 -> Water_Temp.:98.71, PWM:31
22:42:40.266 -> Water_Temp.:98.71, PWM:31
22:42:42.584 -> Water_Temp.:98.71, PWM:31
22:42:43.379 -> Water_Temp.:98.71, PWM:31
22:42:44.138 -> Water_Temp.:98.71, PWM:31
22:42:44.938 -> Water_Temp.:98.71, PWM:31
22:42:45.701 -> Water_Temp.:98.71, PWM:31
22:42:46.499 -> Water_Temp.:98.71, PWM:31
22:42:48.052 -> Water_Temp.:98.71, PWM:31
22:42:48.814 -> Water_Temp.:98.71, PWM:31
22:42:49.609 -> Water_Temp.:98.60, PWM:31
22:42:50.398 -> Water_Temp.:98.60, PWM:31
22:42:51.159 -> Water_Temp.:98.60, PWM:31
22:42:51.948 -> Water_Temp.:98.60, PWM:31
22:42:52.740 -> Water_Temp.:98.60, PWM:31
22:42:53.501 -> Water_Temp.:98.60, PWM:31
22:42:54.295 -> Water_Temp.:98.60, PWM:31
22:42:55.090 -> Water_Temp.:98.60, PWM:31
22:42:55.851 -> Water_Temp.:98.60, PWM:31
22:42:56.642 -> Water_Temp.:98.60, PWM:31
22:42:57.402 -> Water_Temp.:98.60, PWM:31
22:42:58.195 -> Water_Temp.:98.60, PWM:31
22:42:58.959 -> Water_Temp.:98.60, PWM:31
22:42:59.759 -> Water_Temp.:98.60, PWM:31
22:43:00.516 -> Water_Temp.:98.60, PWM:31
22:43:01.314 -> Water_Temp.:98.60, PWM:31
22:43:02.107 -> Water_Temp.:98.60, PWM:31
22:43:02.863 -> Water_Temp.:98.60, PWM:31
22:43:03.657 -> Water_Temp.:98.60, PWM:31
22:43:04.419 -> Water_Temp.:98.60, PWM:31
22:43:05.215 -> Water_Temp.:98.60, PWM:31
22:43:05.977 -> Water_Temp.:98.60, PWM:31
22:43:06.771 -> Water_Temp.:98.60, PWM:31
22:43:07.565 -> Water_Temp.:98.71, PWM:31
22:43:08.325 -> Water_Temp.:98.60, PWM:31
22:43:09.119 -> Water_Temp.:98.71, PWM:31
22:43:09.884 -> Water_Temp.:100.29, PWM:30
22:43:10.681 -> Water_Temp.:102.31, PWM:30
22:43:11.444 -> Water_Temp.:99.95, PWM:31
22:43:12.235 -> Water_Temp.:99.50, PWM:31
22:43:13.028 -> Water_Temp.:99.39, PWM:31
22:43:13.788 -> Water_Temp.:99.27, PWM:31
22:43:14.583 -> Water_Temp.:99.16, PWM:31
22:43:15.340 -> Water_Temp.:99.05, PWM:31
22:43:17.656 -> Water_Temp.:98.94, PWM:31
22:43:18.409 -> Water_Temp.:98.94, PWM:31
22:43:19.202 -> Water_Temp.:98.94, PWM:31
22:43:19.996 -> Water_Temp.:98.94, PWM:31
22:43:20.753 -> Water_Temp.:98.82, PWM:31
22:43:21.550 -> Water_Temp.:98.82, PWM:31
22:43:22.309 -> Water_Temp.:98.82, PWM:31
22:43:23.101 -> Water_Temp.:98.82, PWM:31
22:43:23.862 -> Water_Temp.:98.82, PWM:31
22:43:24.661 -> Water_Temp.:98.82, PWM:31
22:43:25.450 -> Water_Temp.:98.82, PWM:31
22:43:26.209 -> Water_Temp.:98.82, PWM:31
22:43:27.006 -> Water_Temp.:98.82, PWM:31
22:43:27.766 -> Water_Temp.:98.82, PWM:31
22:43:28.560 -> Water_Temp.:98.82, PWM:31
22:43:29.353 -> Water_Temp.:98.82, PWM:31
22:43:30.113 -> Water_Temp.:98.82, PWM:31
22:43:30.910 -> Water_Temp.:98.82, PWM:31
22:43:31.667 -> Water_Temp.:98.82, PWM:30
22:43:32.461 -> Water_Temp.:98.82, PWM:31
22:43:33.221 -> Water_Temp.:98.82, PWM:31
22:43:34.014 -> Water_Temp.:98.82, PWM:31
22:43:34.806 -> Water_Temp.:98.82, PWM:31
22:43:35.566 -> Water_Temp.:98.82, PWM:31
22:43:36.362 -> Water_Temp.:98.82, PWM:31
22:43:37.124 -> Water_Temp.:98.82, PWM:31
22:43:37.916 -> Water_Temp.:98.82, PWM:31
22:43:38.712 -> Water_Temp.:98.82, PWM:31
22:43:39.473 -> Water_Temp.:98.82, PWM:31
22:43:40.270 -> Water_Temp.:98.82, PWM:31
22:43:41.035 -> Water_Temp.:98.82, PWM:31
22:43:43.351 -> Water_Temp.:98.82, PWM:31
22:43:44.143 -> Water_Temp.:98.82, PWM:31
22:43:44.906 -> Water_Temp.:98.82, PWM:31
22:43:45.703 -> Water_Temp.:98.82, PWM:31
22:43:46.491 -> Water_Temp.:98.82, PWM:31
22:43:47.249 -> Water_Temp.:98.82, PWM:31
22:43:48.045 -> Water_Temp.:98.82, PWM:31
22:43:48.839 -> Water_Temp.:98.82, PWM:30
22:43:49.599 -> Water_Temp.:98.82, PWM:31
22:43:50.394 -> Water_Temp.:98.82, PWM:31
22:43:51.155 -> Water_Temp.:98.82, PWM:31
22:43:51.951 -> Water_Temp.:98.82, PWM:31
22:43:52.710 -> Water_Temp.:98.82, PWM:31
22:43:53.504 -> Water_Temp.:98.82, PWM:31
22:43:54.297 -> Water_Temp.:98.82, PWM:31
22:43:55.058 -> Water_Temp.:98.82, PWM:31
22:43:55.853 -> Water_Temp.:98.82, PWM:31
22:43:56.614 -> Water_Temp.:98.82, PWM:31
22:43:57.408 -> Water_Temp.:98.71, PWM:31
22:43:58.196 -> Water_Temp.:98.71, PWM:31
22:43:58.958 -> Water_Temp.:98.71, PWM:31
22:43:59.752 -> Water_Temp.:98.82, PWM:31
22:44:00.512 -> Water_Temp.:98.71, PWM:31
22:44:01.308 -> Water_Temp.:98.71, PWM:31
22:44:02.068 -> Water_Temp.:98.71, PWM:31
22:44:02.860 -> Water_Temp.:98.71, PWM:31
22:44:03.653 -> Water_Temp.:98.71, PWM:31
22:44:04.415 -> Water_Temp.:98.71, PWM:31
22:44:05.211 -> Water_Temp.:98.71, PWM:31
22:44:05.966 -> Water_Temp.:98.71, PWM:31
22:44:07.518 -> Water_Temp.:98.71, PWM:31
22:44:08.313 -> Water_Temp.:98.71, PWM:31
22:44:09.076 -> Water_Temp.:98.71, PWM:31
22:44:09.871 -> Water_Temp.:98.71, PWM:31
22:44:10.634 -> Water_Temp.:98.71, PWM:31
22:44:11.435 -> Water_Temp.:98.71, PWM:31
22:44:12.192 -> Water_Temp.:98.71, PWM:30
22:44:12.984 -> Water_Temp.:98.71, PWM:31
22:44:13.745 -> Water_Temp.:98.71, PWM:31
22:44:14.542 -> Water_Temp.:98.71, PWM:31
22:44:15.338 -> Water_Temp.:98.71, PWM:31
22:44:16.098 -> Water_Temp.:98.82, PWM:31
22:44:16.891 -> Water_Temp.:98.82, PWM:31
22:44:17.653 -> Water_Temp.:98.82, PWM:31
22:44:18.445 -> Water_Temp.:98.82, PWM:31
22:44:19.207 -> Water_Temp.:98.82, PWM:31
22:44:20.003 -> Water_Temp.:98.82, PWM:31
22:44:20.767 -> Water_Temp.:98.82, PWM:31
22:44:21.567 -> Water_Temp.:98.82, PWM:31
22:44:22.325 -> Water_Temp.:98.82, PWM:31
22:44:23.118 -> Water_Temp.:98.82, PWM:31
22:44:23.914 -> Water_Temp.:98.82, PWM:31
22:44:24.677 -> Water_Temp.:98.82, PWM:31
22:44:25.471 -> Water_Temp.:98.82, PWM:31
22:44:26.228 -> Water_Temp.:98.82, PWM:30
22:44:27.787 -> Water_Temp.:98.82, PWM:31
22:44:28.544 -> Water_Temp.:98.82, PWM:31
22:44:29.333 -> Water_Temp.:98.82, PWM:31
22:44:30.127 -> Water_Temp.:98.82, PWM:31
22:44:30.892 -> Water_Temp.:98.82, PWM:31
22:44:31.685 -> Water_Temp.:98.82, PWM:31
22:44:32.446 -> Water_Temp.:98.82, PWM:31
22:44:33.238 -> Water_Temp.:98.82, PWM:31
22:44:34.031 -> Water_Temp.:98.82, PWM:31
22:44:34.791 -> Water_Temp.:98.82, PWM:31
22:44:36.340 -> Water_Temp.:98.82, PWM:31
22:44:37.103 -> Water_Temp.:98.82, PWM:31
22:44:37.899 -> Water_Temp.:98.82, PWM:31
22:44:38.660 -> Water_Temp.:98.82, PWM:31
22:44:39.453 -> Water_Temp.:98.82, PWM:31
22:44:40.249 -> Water_Temp.:98.94, PWM:31
22:44:41.014 -> Water_Temp.:98.82, PWM:31
22:44:41.805 -> Water_Temp.:98.94, PWM:31
22:44:42.563 -> Water_Temp.:98.94, PWM:31
22:44:43.354 -> Water_Temp.:98.82, PWM:31
22:44:44.115 -> Water_Temp.:98.82, PWM:31
22:44:44.910 -> Water_Temp.:98.94, PWM:31
22:44:45.704 -> Water_Temp.:98.82, PWM:31
22:44:46.462 -> Water_Temp.:98.82, PWM:31
22:44:47.251 -> Water_Temp.:98.82, PWM:31
22:44:48.041 -> Water_Temp.:98.82, PWM:31
22:44:48.805 -> Water_Temp.:98.82, PWM:31
22:44:50.364 -> Water_Temp.:98.82, PWM:31
22:44:51.920 -> Water_Temp.:98.82, PWM:31
22:44:52.711 -> Water_Temp.:98.82, PWM:31
22:44:53.474 -> Water_Temp.:98.82, PWM:31
22:44:54.269 -> Water_Temp.:98.82, PWM:31
22:44:55.036 -> Water_Temp.:98.82, PWM:31
22:44:55.828 -> Water_Temp.:98.82, PWM:31
22:44:56.624 -> Water_Temp.:98.82, PWM:31
22:44:57.377 -> Water_Temp.:98.82, PWM:31
22:44:58.170 -> Water_Temp.:98.82, PWM:31
22:44:58.965 -> Water_Temp.:98.82, PWM:31
22:44:59.722 -> Water_Temp.:98.82, PWM:31
22:45:00.519 -> Water_Temp.:98.82, PWM:31
22:45:01.284 -> Water_Temp.:98.82, PWM:31
22:45:02.077 -> Water_Temp.:98.82, PWM:31
22:45:02.837 -> Water_Temp.:98.82, PWM:31
22:45:04.395 -> Water_Temp.:98.82, PWM:31
22:45:05.194 -> Water_Temp.:98.82, PWM:31
22:45:05.954 -> Water_Temp.:98.82, PWM:31
22:45:07.505 -> Water_Temp.:98.82, PWM:31
22:45:08.291 -> Water_Temp.:98.82, PWM:31
22:45:09.078 -> Water_Temp.:98.82, PWM:31
22:45:09.865 -> Water_Temp.:98.82, PWM:31
22:45:10.622 -> Water_Temp.:98.82, PWM:31
22:45:11.413 -> Water_Temp.:98.82, PWM:31
22:45:12.210 -> Water_Temp.:98.94, PWM:30
22:45:12.976 -> Water_Temp.:98.82, PWM:31
22:45:13.769 -> Water_Temp.:98.94, PWM:31
22:45:14.534 -> Water_Temp.:98.94, PWM:31
22:45:15.329 -> Water_Temp.:98.94, PWM:31
22:45:16.086 -> Water_Temp.:98.94, PWM:31
22:45:16.882 -> Water_Temp.:98.94, PWM:31
22:45:17.643 -> Water_Temp.:98.94, PWM:31
22:45:18.436 -> Water_Temp.:98.94, PWM:31
22:45:19.199 -> Water_Temp.:98.94, PWM:30
22:45:19.993 -> Water_Temp.:98.94, PWM:31
22:45:20.755 -> Water_Temp.:98.94, PWM:31
22:45:21.552 -> Water_Temp.:98.94, PWM:31
22:45:22.349 -> Water_Temp.:98.94, PWM:31
22:45:23.898 -> Water_Temp.:98.94, PWM:31
22:45:25.455 -> Water_Temp.:98.94, PWM:31
22:45:26.248 -> Water_Temp.:98.94, PWM:31
22:45:27.009 -> Water_Temp.:98.94, PWM:31
22:45:27.803 -> Water_Temp.:98.94, PWM:31
22:45:28.564 -> Water_Temp.:98.94, PWM:31
22:45:29.357 -> Water_Temp.:98.94, PWM:31
22:45:30.120 -> Water_Temp.:98.94, PWM:31
22:45:30.910 -> Water_Temp.:98.94, PWM:31
22:45:31.708 -> Water_Temp.:98.94, PWM:31
22:45:32.467 -> Water_Temp.:99.05, PWM:31
22:45:33.260 -> Water_Temp.:99.05, PWM:31
22:45:34.021 -> Water_Temp.:99.05, PWM:31
22:45:34.817 -> Water_Temp.:99.05, PWM:31
22:45:35.579 -> Water_Temp.:99.05, PWM:31
22:45:36.376 -> Water_Temp.:99.05, PWM:31
22:45:37.136 -> Water_Temp.:99.05, PWM:31
22:45:37.927 -> Water_Temp.:99.05, PWM:31
22:45:38.718 -> Water_Temp.:99.05, PWM:31
22:45:39.476 -> Water_Temp.:99.05, PWM:31
22:45:40.275 -> Water_Temp.:99.05, PWM:31
22:45:41.034 -> Water_Temp.:99.05, PWM:31
22:45:41.827 -> Water_Temp.:99.05, PWM:31
22:45:42.619 -> Water_Temp.:99.05, PWM:31
22:45:43.380 -> Water_Temp.:99.05, PWM:31
22:45:44.171 -> Water_Temp.:99.05, PWM:31
22:45:44.967 -> Water_Temp.:99.05, PWM:31
22:45:45.730 -> Water_Temp.:99.05, PWM:31
22:45:46.527 -> Water_Temp.:99.05, PWM:31
22:45:47.282 -> Water_Temp.:99.05, PWM:31
22:45:48.076 -> Water_Temp.:99.05, PWM:31
22:45:48.835 -> Water_Temp.:99.05, PWM:31
22:45:49.626 -> Water_Temp.:99.05, PWM:31
22:45:50.422 -> Water_Temp.:99.05, PWM:31
22:45:51.183 -> Water_Temp.:99.05, PWM:31
22:45:51.976 -> Water_Temp.:99.05, PWM:31
22:45:52.735 -> Water_Temp.:99.05, PWM:31
22:45:53.526 -> Water_Temp.:99.05, PWM:31
22:45:54.322 -> Water_Temp.:99.05, PWM:31
22:45:55.089 -> Water_Temp.:99.05, PWM:31
22:45:55.880 -> Water_Temp.:99.05, PWM:31
22:45:56.639 -> Water_Temp.:99.05, PWM:31
22:45:57.430 -> Water_Temp.:99.05, PWM:31
22:45:58.191 -> Water_Temp.:99.05, PWM:31
22:45:58.983 -> Water_Temp.:99.05, PWM:31
22:45:59.777 -> Water_Temp.:99.05, PWM:31
22:46:00.540 -> Water_Temp.:99.05, PWM:31
22:46:01.339 -> Water_Temp.:99.05, PWM:31
22:46:02.095 -> Water_Temp.:99.05, PWM:31
22:46:02.895 -> Water_Temp.:99.05, PWM:31
22:46:03.656 -> Water_Temp.:99.05, PWM:31
22:46:04.451 -> Water_Temp.:99.05, PWM:31
22:46:05.218 -> Water_Temp.:99.05, PWM:31
22:46:06.010 -> Water_Temp.:99.05, PWM:31
22:46:06.805 -> Water_Temp.:99.05, PWM:31
22:46:07.569 -> Water_Temp.:99.05, PWM:31
22:46:08.363 -> Water_Temp.:99.05, PWM:31
22:46:09.125 -> Water_Temp.:99.05, PWM:31
22:46:09.918 -> Water_Temp.:99.05, PWM:31
22:46:10.679 -> Water_Temp.:99.05, PWM:31
22:46:11.473 -> Water_Temp.:99.05, PWM:31
22:46:12.235 -> Water_Temp.:99.05, PWM:31
22:46:13.820 -> Water_Temp.:99.05, PWM:31
22:46:14.578 -> Water_Temp.:99.05, PWM:31
22:46:15.376 -> Water_Temp.:99.05, PWM:31
22:46:16.134 -> Water_Temp.:99.05, PWM:31
22:46:16.929 -> Water_Temp.:99.05, PWM:31
22:46:17.722 -> Water_Temp.:99.05, PWM:31
22:46:18.479 -> Water_Temp.:99.05, PWM:31
22:46:19.274 -> Water_Temp.:99.05, PWM:31
22:46:20.038 -> Water_Temp.:99.05, PWM:31
22:46:20.834 -> Water_Temp.:99.05, PWM:31
22:46:21.594 -> Water_Temp.:99.05, PWM:31
22:46:22.389 -> Water_Temp.:99.05, PWM:31
22:46:23.183 -> Water_Temp.:99.05, PWM:31
22:46:23.944 -> Water_Temp.:99.05, PWM:31
22:46:24.739 -> Water_Temp.:99.05, PWM:31
22:46:25.497 -> Water_Temp.:99.05, PWM:31
22:46:26.287 -> Water_Temp.:99.05, PWM:31
22:46:27.081 -> Water_Temp.:99.05, PWM:31
22:46:27.836 -> Water_Temp.:99.05, PWM:31
22:46:28.626 -> Water_Temp.:99.05, PWM:31
22:46:29.421 -> Water_Temp.:99.05, PWM:31
22:46:30.177 -> Water_Temp.:99.05, PWM:30
22:46:30.974 -> Water_Temp.:99.05, PWM:31
22:46:31.730 -> Water_Temp.:99.05, PWM:31
22:46:32.522 -> Water_Temp.:99.05, PWM:31
22:46:33.315 -> Water_Temp.:99.05, PWM:31
22:46:34.077 -> Water_Temp.:99.05, PWM:31
22:46:34.876 -> Water_Temp.:99.05, PWM:31
22:46:35.636 -> Water_Temp.:99.05, PWM:31
22:46:36.432 -> Water_Temp.:99.05, PWM:31
22:46:37.192 -> Water_Temp.:99.05, PWM:31
22:46:37.989 -> Water_Temp.:99.05, PWM:31
22:46:38.748 -> Water_Temp.:99.05, PWM:31
22:46:39.544 -> Water_Temp.:99.05, PWM:31
22:46:40.340 -> Water_Temp.:99.16, PWM:31
22:46:41.102 -> Water_Temp.:99.05, PWM:31
22:46:41.896 -> Water_Temp.:99.16, PWM:31
22:46:42.656 -> Water_Temp.:99.16, PWM:31
22:46:43.450 -> Water_Temp.:99.16, PWM:34
22:46:44.212 -> Water_Temp.:99.16, PWM:34
22:46:45.013 -> Water_Temp.:99.16, PWM:34
22:46:45.771 -> Water_Temp.:99.16, PWM:34
22:46:46.569 -> Water_Temp.:99.16, PWM:34
22:46:47.361 -> Water_Temp.:99.16, PWM:34
22:46:48.122 -> Water_Temp.:99.16, PWM:34
22:46:48.921 -> Water_Temp.:99.16, PWM:34
22:46:49.679 -> Water_Temp.:99.16, PWM:34
22:46:50.475 -> Water_Temp.:99.16, PWM:34
22:46:51.243 -> Water_Temp.:99.16, PWM:34
22:46:52.032 -> Water_Temp.:99.16, PWM:34
22:46:52.794 -> Water_Temp.:99.16, PWM:34
22:46:53.589 -> Water_Temp.:99.16, PWM:34
22:46:54.349 -> Water_Temp.:99.16, PWM:34
22:46:55.149 -> Water_Temp.:99.16, PWM:34
22:46:55.942 -> Water_Temp.:99.16, PWM:34
22:46:56.702 -> Water_Temp.:99.16, PWM:34
22:46:57.498 -> Water_Temp.:99.16, PWM:34
22:46:58.259 -> Water_Temp.:99.16, PWM:34
22:46:59.054 -> Water_Temp.:99.16, PWM:34
22:46:59.814 -> Water_Temp.:99.16, PWM:34
22:47:01.374 -> Water_Temp.:99.16, PWM:34
22:47:02.133 -> Water_Temp.:99.16, PWM:34
22:47:03.682 -> Water_Temp.:99.16, PWM:34
22:47:04.480 -> Water_Temp.:99.16, PWM:34
22:47:05.246 -> Water_Temp.:99.16, PWM:34
22:47:06.038 -> Water_Temp.:99.16, PWM:34
22:47:06.796 -> Water_Temp.:99.16, PWM:34
22:47:07.591 -> Water_Temp.:99.16, PWM:34
22:47:08.383 -> Water_Temp.:99.16, PWM:34
22:47:09.144 -> Water_Temp.:99.16, PWM:34
22:47:09.938 -> Water_Temp.:99.16, PWM:34
22:47:10.697 -> Water_Temp.:99.16, PWM:34
22:47:12.283 -> Water_Temp.:99.16, PWM:34
22:47:13.041 -> Water_Temp.:99.16, PWM:34
22:47:13.834 -> Water_Temp.:99.16, PWM:34
22:47:14.598 -> Water_Temp.:99.16, PWM:34
22:47:16.152 -> Water_Temp.:99.16, PWM:34
22:47:16.911 -> Water_Temp.:99.16, PWM:34
22:47:17.704 -> Water_Temp.:99.16, PWM:34
22:47:18.502 -> Water_Temp.:99.16, PWM:34
22:47:19.259 -> Water_Temp.:99.16, PWM:34
22:47:20.049 -> Water_Temp.:99.16, PWM:34
22:47:20.815 -> Water_Temp.:99.16, PWM:34
22:47:21.611 -> Water_Temp.:99.27, PWM:34
22:47:22.372 -> Water_Temp.:99.27, PWM:34
22:47:23.964 -> Water_Temp.:99.16, PWM:34
22:47:25.522 -> Water_Temp.:99.16, PWM:34
22:47:26.280 -> Water_Temp.:99.27, PWM:34
22:47:27.077 -> Water_Temp.:99.27, PWM:34
22:47:27.835 -> Water_Temp.:99.16, PWM:34
22:47:28.635 -> Water_Temp.:99.27, PWM:34
22:47:30.151 -> Water_Temp.:99.27, PWM:34
22:47:30.950 -> Water_Temp.:99.27, PWM:34
22:47:31.714 -> Water_Temp.:99.27, PWM:34
22:47:32.509 -> Water_Temp.:99.27, PWM:34
22:47:33.271 -> Water_Temp.:99.27, PWM:34
22:47:34.826 -> Water_Temp.:99.27, PWM:34
22:47:35.622 -> Water_Temp.:99.27, PWM:34
22:47:36.379 -> Water_Temp.:99.27, PWM:34
22:47:37.172 -> Water_Temp.:100.85, PWM:34
22:47:37.932 -> Water_Temp.:102.87, PWM:34
22:47:38.727 -> Water_Temp.:100.51, PWM:34
22:47:39.520 -> Water_Temp.:100.06, PWM:34
22:47:40.283 -> Water_Temp.:99.95, PWM:34
22:47:41.082 -> Water_Temp.:99.84, PWM:34
22:47:41.843 -> Water_Temp.:99.72, PWM:34
22:47:42.641 -> Water_Temp.:99.61, PWM:34
22:47:43.403 -> Water_Temp.:99.61, PWM:34
22:47:45.721 -> Water_Temp.:99.50, PWM:34
22:47:46.515 -> Water_Temp.:99.39, PWM:34
22:47:47.275 -> Water_Temp.:99.39, PWM:34
22:47:48.075 -> Water_Temp.:99.39, PWM:34
22:47:48.838 -> Water_Temp.:99.39, PWM:34
22:47:49.637 -> Water_Temp.:99.39, PWM:34
22:47:50.398 -> Water_Temp.:99.39, PWM:34
22:47:51.161 -> Water_Temp.:99.27, PWM:34
22:47:51.950 -> Water_Temp.:99.27, PWM:34
22:47:52.738 -> Water_Temp.:99.27, PWM:34
22:47:53.537 -> Water_Temp.:99.27, PWM:34
22:47:54.302 -> Water_Temp.:99.27, PWM:34
22:47:55.067 -> Water_Temp.:99.27, PWM:34
22:47:55.856 -> Water_Temp.:99.27, PWM:34
22:47:56.620 -> Water_Temp.:99.27, PWM:34
22:47:57.422 -> Water_Temp.:99.27, PWM:34
22:47:58.180 -> Water_Temp.:99.27, PWM:34
22:47:58.975 -> Water_Temp.:99.27, PWM:34
22:47:59.774 -> Water_Temp.:99.27, PWM:34
22:48:00.538 -> Water_Temp.:99.27, PWM:34
22:48:01.302 -> Water_Temp.:99.27, PWM:34
22:48:02.099 -> Water_Temp.:99.27, PWM:34
22:48:02.863 -> Water_Temp.:99.27, PWM:34
22:48:03.659 -> Water_Temp.:99.27, PWM:34
22:48:04.456 -> Water_Temp.:99.27, PWM:34
22:48:05.220 -> Water_Temp.:99.27, PWM:34
22:48:06.013 -> Water_Temp.:99.27, PWM:34
22:48:06.773 -> Water_Temp.:99.27, PWM:34
22:48:07.570 -> Water_Temp.:99.27, PWM:34
22:48:08.333 -> Water_Temp.:99.27, PWM:34
22:48:09.128 -> Water_Temp.:99.39, PWM:34
22:48:09.891 -> Water_Temp.:99.39, PWM:34
22:48:11.447 -> Water_Temp.:99.39, PWM:34
22:48:12.206 -> Water_Temp.:99.39, PWM:34
22:48:13.760 -> Water_Temp.:99.39, PWM:34
22:48:14.556 -> Water_Temp.:99.39, PWM:34
22:48:16.076 -> Water_Temp.:99.39, PWM:34
22:48:16.868 -> Water_Temp.:99.39, PWM:34
22:48:17.630 -> Water_Temp.:99.39, PWM:34
22:48:18.426 -> Water_Temp.:99.39, PWM:34
22:48:19.187 -> Water_Temp.:99.39, PWM:34
22:48:19.984 -> Water_Temp.:99.39, PWM:34
22:48:20.746 -> Water_Temp.:99.39, PWM:34
22:48:21.540 -> Water_Temp.:99.39, PWM:34
22:48:22.333 -> Water_Temp.:99.27, PWM:34
22:48:23.090 -> Water_Temp.:99.39, PWM:34
22:48:23.883 -> Water_Temp.:99.27, PWM:34
22:48:24.647 -> Water_Temp.:99.27, PWM:34
22:48:25.449 -> Water_Temp.:99.39, PWM:34
22:48:26.209 -> Water_Temp.:99.27, PWM:34
22:48:27.004 -> Water_Temp.:99.27, PWM:34
22:48:27.797 -> Water_Temp.:99.39, PWM:34
22:48:28.561 -> Water_Temp.:99.39, PWM:34
22:48:29.356 -> Water_Temp.:99.27, PWM:34
22:48:30.120 -> Water_Temp.:99.39, PWM:34
22:48:30.917 -> Water_Temp.:99.39, PWM:34
22:48:31.671 -> Water_Temp.:99.39, PWM:34
22:48:32.466 -> Water_Temp.:99.39, PWM:34
22:48:33.229 -> Water_Temp.:99.39, PWM:34
22:48:34.023 -> Water_Temp.:99.39, PWM:34
22:48:34.786 -> Water_Temp.:99.39, PWM:34
22:48:35.582 -> Water_Temp.:99.39, PWM:34
22:48:36.375 -> Water_Temp.:99.39, PWM:34
22:48:37.134 -> Water_Temp.:99.39, PWM:34
22:48:37.932 -> Water_Temp.:99.39, PWM:34
22:48:38.693 -> Water_Temp.:99.39, PWM:34
22:48:40.220 -> Water_Temp.:99.39, PWM:34
22:48:41.014 -> Water_Temp.:99.39, PWM:34
22:48:41.810 -> Water_Temp.:99.39, PWM:34
22:48:42.569 -> Water_Temp.:99.39, PWM:34
22:48:43.366 -> Water_Temp.:99.39, PWM:34
22:48:44.130 -> Water_Temp.:99.39, PWM:34
22:48:44.925 -> Water_Temp.:99.39, PWM:34
22:48:45.689 -> Water_Temp.:99.39, PWM:34
22:48:46.484 -> Water_Temp.:99.39, PWM:34
22:48:47.245 -> Water_Temp.:99.39, PWM:34
22:48:48.042 -> Water_Temp.:99.39, PWM:34
22:48:48.802 -> Water_Temp.:99.39, PWM:34
22:48:49.594 -> Water_Temp.:99.39, PWM:34
22:48:50.389 -> Water_Temp.:99.39, PWM:34
22:48:51.149 -> Water_Temp.:99.39, PWM:34
22:48:51.943 -> Water_Temp.:99.39, PWM:34
22:48:52.699 -> Water_Temp.:99.39, PWM:34
22:48:53.494 -> Water_Temp.:99.39, PWM:34
22:48:54.256 -> Water_Temp.:99.39, PWM:34
22:48:55.057 -> Water_Temp.:99.39, PWM:34
22:48:55.849 -> Water_Temp.:99.39, PWM:34
22:48:56.611 -> Water_Temp.:99.39, PWM:34
22:48:57.398 -> Water_Temp.:99.39, PWM:34
22:48:58.937 -> Water_Temp.:99.39, PWM:34
22:48:59.731 -> Water_Temp.:99.39, PWM:34
22:49:00.524 -> Water_Temp.:99.39, PWM:34
22:49:01.286 -> Water_Temp.:99.39, PWM:34
22:49:02.077 -> Water_Temp.:99.39, PWM:34
22:49:03.625 -> Water_Temp.:99.39, PWM:34
22:49:04.415 -> Water_Temp.:99.39, PWM:34
22:49:05.181 -> Water_Temp.:99.50, PWM:34
22:49:05.970 -> Water_Temp.:99.50, PWM:34
22:49:07.492 -> Water_Temp.:99.50, PWM:34
22:49:08.284 -> Water_Temp.:99.50, PWM:34
22:49:09.075 -> Water_Temp.:99.50, PWM:34
22:49:09.834 -> Water_Temp.:99.50, PWM:34
22:49:10.630 -> Water_Temp.:99.50, PWM:34
22:49:11.390 -> Water_Temp.:99.50, PWM:34
22:49:12.183 -> Water_Temp.:99.50, PWM:34
22:49:12.978 -> Water_Temp.:99.50, PWM:34
22:49:13.738 -> Water_Temp.:99.50, PWM:34
22:49:14.534 -> Water_Temp.:99.50, PWM:34
22:49:16.817 -> Water_Temp.:99.50, PWM:34
22:49:17.616 -> Water_Temp.:99.50, PWM:34
22:49:18.377 -> Water_Temp.:99.50, PWM:34
22:49:19.171 -> Water_Temp.:99.50, PWM:34
22:49:19.933 -> Water_Temp.:99.50, PWM:34
22:49:20.731 -> Water_Temp.:99.50, PWM:34
22:49:21.487 -> Water_Temp.:99.50, PWM:34
22:49:22.280 -> Water_Temp.:99.39, PWM:34
22:49:23.075 -> Water_Temp.:99.39, PWM:34
22:49:23.836 -> Water_Temp.:99.39, PWM:34
22:49:24.601 -> Water_Temp.:99.39, PWM:34
22:49:25.399 -> Water_Temp.:99.39, PWM:34
22:49:26.194 -> Water_Temp.:99.39, PWM:34
22:49:26.954 -> Water_Temp.:99.39, PWM:34
22:49:27.752 -> Water_Temp.:99.39, PWM:34
22:49:28.514 -> Water_Temp.:99.39, PWM:34
22:49:29.309 -> Water_Temp.:99.39, PWM:34
22:49:30.071 -> Water_Temp.:99.39, PWM:34
22:49:30.870 -> Water_Temp.:99.39, PWM:34
22:49:31.633 -> Water_Temp.:99.39, PWM:34
22:49:32.432 -> Water_Temp.:99.39, PWM:34
22:49:33.193 -> Water_Temp.:99.39, PWM:34
22:49:33.988 -> Water_Temp.:99.39, PWM:34
22:49:34.749 -> Water_Temp.:99.39, PWM:34
22:49:35.546 -> Water_Temp.:99.39, PWM:34
22:49:36.307 -> Water_Temp.:101.07, PWM:34
22:49:37.103 -> Water_Temp.:103.10, PWM:34
22:49:37.863 -> Water_Temp.:100.74, PWM:34
22:49:38.660 -> Water_Temp.:100.29, PWM:34
22:49:39.452 -> Water_Temp.:100.17, PWM:34
22:49:40.215 -> Water_Temp.:99.95, PWM:34
22:49:41.014 -> Water_Temp.:99.95, PWM:34
22:49:41.775 -> Water_Temp.:99.84, PWM:34
22:49:43.329 -> Water_Temp.:99.72, PWM:34
22:49:44.091 -> Water_Temp.:99.61, PWM:34
22:49:44.889 -> Water_Temp.:99.61, PWM:34
22:49:45.652 -> Water_Temp.:99.61, PWM:34
22:49:46.444 -> Water_Temp.:99.61, PWM:34
22:49:48.755 -> Water_Temp.:99.50, PWM:34
22:49:49.515 -> Water_Temp.:99.50, PWM:34
22:49:50.315 -> Water_Temp.:99.50, PWM:34
22:49:51.074 -> Water_Temp.:99.50, PWM:34
22:49:51.867 -> Water_Temp.:99.50, PWM:34
22:49:53.432 -> Water_Temp.:99.50, PWM:34
22:49:54.200 -> Water_Temp.:99.50, PWM:34
22:49:55.001 -> Water_Temp.:99.50, PWM:34
22:49:55.763 -> Water_Temp.:99.50, PWM:34
22:49:56.556 -> Water_Temp.:99.50, PWM:34
22:49:57.348 -> Water_Temp.:99.50, PWM:34
22:49:58.105 -> Water_Temp.:99.50, PWM:34
22:49:58.901 -> Water_Temp.:99.50, PWM:34
22:50:00.454 -> Water_Temp.:99.61, PWM:34
22:50:01.217 -> Water_Temp.:99.61, PWM:34
22:50:02.008 -> Water_Temp.:99.61, PWM:34
22:50:02.802 -> Water_Temp.:99.61, PWM:34
22:50:03.558 -> Water_Temp.:99.61, PWM:34
22:50:04.351 -> Water_Temp.:99.61, PWM:34
22:50:05.116 -> Water_Temp.:99.61, PWM:34
22:50:05.912 -> Water_Temp.:99.61, PWM:34
22:50:06.670 -> Water_Temp.:99.61, PWM:34
22:50:07.466 -> Water_Temp.:99.61, PWM:34
22:50:08.260 -> Water_Temp.:99.61, PWM:34
22:50:09.018 -> Water_Temp.:99.61, PWM:34
22:50:09.815 -> Water_Temp.:99.61, PWM:34
22:50:10.573 -> Water_Temp.:99.61, PWM:34
22:50:11.373 -> Water_Temp.:99.61, PWM:34
22:50:12.130 -> Water_Temp.:99.61, PWM:34
22:50:12.925 -> Water_Temp.:99.61, PWM:34
22:50:13.718 -> Water_Temp.:99.61, PWM:34
22:50:14.474 -> Water_Temp.:99.61, PWM:34
22:50:15.275 -> Water_Temp.:99.61, PWM:34
22:50:16.036 -> Water_Temp.:99.61, PWM:34
22:50:16.829 -> Water_Temp.:99.61, PWM:34
22:50:17.589 -> Water_Temp.:99.61, PWM:34
22:50:18.382 -> Water_Temp.:99.72, PWM:34
22:50:19.173 -> Water_Temp.:99.61, PWM:34
22:50:20.725 -> Water_Temp.:99.72, PWM:34
22:50:21.490 -> Water_Temp.:99.72, PWM:34
22:50:22.285 -> Water_Temp.:99.72, PWM:34
22:50:23.078 -> Water_Temp.:99.72, PWM:34
22:50:24.601 -> Water_Temp.:99.72, PWM:34
22:50:25.397 -> Water_Temp.:99.72, PWM:34
22:50:26.158 -> Water_Temp.:99.72, PWM:34
22:50:26.954 -> Water_Temp.:99.72, PWM:34
22:50:28.505 -> Water_Temp.:99.72, PWM:34
22:50:29.298 -> Water_Temp.:99.72, PWM:34
22:50:30.061 -> Water_Temp.:99.72, PWM:34
22:50:30.856 -> Water_Temp.:99.72, PWM:34
22:50:31.650 -> Water_Temp.:99.72, PWM:34
22:50:32.412 -> Water_Temp.:99.72, PWM:34
22:50:33.204 -> Water_Temp.:99.72, PWM:34
22:50:33.965 -> Water_Temp.:99.72, PWM:34
22:50:34.764 -> Water_Temp.:99.72, PWM:34
22:50:36.291 -> Water_Temp.:99.72, PWM:34
22:50:37.087 -> Water_Temp.:99.72, PWM:34
22:50:37.846 -> Water_Temp.:99.72, PWM:34
22:50:38.641 -> Water_Temp.:99.72, PWM:34
22:50:39.401 -> Water_Temp.:99.72, PWM:34
22:50:40.195 -> Water_Temp.:99.72, PWM:34
22:50:40.994 -> Water_Temp.:99.72, PWM:34
22:50:41.749 -> Water_Temp.:99.72, PWM:34
22:50:42.545 -> Water_Temp.:99.72, PWM:34
22:50:43.306 -> Water_Temp.:99.72, PWM:34
22:50:44.099 -> Water_Temp.:99.72, PWM:34
22:50:44.858 -> Water_Temp.:99.72, PWM:34
22:50:45.656 -> Water_Temp.:99.72, PWM:34
22:50:46.452 -> Water_Temp.:99.72, PWM:34
22:50:47.210 -> Water_Temp.:99.72, PWM:34
22:50:48.003 -> Water_Temp.:99.72, PWM:34
22:50:49.522 -> Water_Temp.:99.72, PWM:34
22:50:50.319 -> Water_Temp.:99.61, PWM:34
22:50:51.080 -> Water_Temp.:99.61, PWM:34
22:50:51.878 -> Water_Temp.:99.61, PWM:34
22:50:52.637 -> Water_Temp.:99.61, PWM:34
22:50:54.189 -> Water_Temp.:99.61, PWM:34
22:50:54.983 -> Water_Temp.:99.61, PWM:34
22:50:55.743 -> Water_Temp.:99.61, PWM:34
22:50:56.530 -> Water_Temp.:99.61, PWM:34
22:50:57.327 -> Water_Temp.:99.61, PWM:34
22:50:58.087 -> Water_Temp.:99.61, PWM:34
22:50:58.879 -> Water_Temp.:99.61, PWM:34
22:50:59.639 -> Water_Temp.:99.61, PWM:34
22:51:01.196 -> Water_Temp.:99.61, PWM:34
22:51:01.959 -> Water_Temp.:99.61, PWM:34
22:51:02.755 -> Water_Temp.:99.61, PWM:34
22:51:03.517 -> Water_Temp.:99.61, PWM:34
22:51:04.309 -> Water_Temp.:99.61, PWM:34
22:51:05.108 -> Water_Temp.:99.61, PWM:34
22:51:05.865 -> Water_Temp.:99.61, PWM:34
22:51:06.658 -> Water_Temp.:99.61, PWM:34
22:51:07.451 -> Water_Temp.:99.61, PWM:34
22:51:08.211 -> Water_Temp.:99.61, PWM:34
22:51:09.006 -> Water_Temp.:99.61, PWM:34
22:51:09.770 -> Water_Temp.:99.61, PWM:34
22:51:10.567 -> Water_Temp.:99.61, PWM:34
22:51:12.122 -> Water_Temp.:99.61, PWM:34
22:51:12.878 -> Water_Temp.:99.61, PWM:34
22:51:13.673 -> Water_Temp.:99.61, PWM:34
22:51:14.435 -> Water_Temp.:99.61, PWM:34
22:51:15.233 -> Water_Temp.:99.72, PWM:34
22:51:16.025 -> Water_Temp.:99.72, PWM:34
22:51:16.783 -> Water_Temp.:99.72, PWM:34
22:51:17.575 -> Water_Temp.:99.72, PWM:34
22:51:18.362 -> Water_Temp.:99.72, PWM:34
22:51:19.121 -> Water_Temp.:99.72, PWM:34
22:51:19.917 -> Water_Temp.:99.72, PWM:34
22:51:20.680 -> Water_Temp.:99.72, PWM:34
22:51:21.473 -> Water_Temp.:99.72, PWM:34
22:51:22.235 -> Water_Temp.:99.72, PWM:34
22:51:23.035 -> Water_Temp.:99.72, PWM:34
22:51:23.798 -> Water_Temp.:99.72, PWM:34
22:51:24.594 -> Water_Temp.:99.72, PWM:34
22:51:25.358 -> Water_Temp.:99.72, PWM:34
22:51:26.152 -> Water_Temp.:99.72, PWM:34
22:51:26.946 -> Water_Temp.:99.72, PWM:34
22:51:27.706 -> Water_Temp.:99.72, PWM:34
22:51:28.504 -> Water_Temp.:99.72, PWM:34
22:51:29.266 -> Water_Temp.:99.72, PWM:34
22:51:30.059 -> Water_Temp.:99.72, PWM:34
22:51:30.816 -> Water_Temp.:99.72, PWM:34
22:51:31.607 -> Water_Temp.:99.72, PWM:34
22:51:32.400 -> Water_Temp.:99.72, PWM:34
22:51:33.162 -> Water_Temp.:99.72, PWM:34
22:51:33.954 -> Water_Temp.:99.72, PWM:34
22:51:34.714 -> Water_Temp.:99.72, PWM:34
22:51:35.510 -> Water_Temp.:99.72, PWM:34
22:51:36.301 -> Water_Temp.:99.72, PWM:34
22:51:37.056 -> Water_Temp.:99.72, PWM:34
22:51:37.851 -> Water_Temp.:99.72, PWM:34
22:51:38.643 -> Water_Temp.:99.72, PWM:34
22:51:39.404 -> Water_Temp.:99.72, PWM:34
22:51:40.195 -> Water_Temp.:99.72, PWM:34
22:51:40.961 -> Water_Temp.:99.72, PWM:34
22:51:41.755 -> Water_Temp.:99.72, PWM:34
22:51:42.548 -> Water_Temp.:99.72, PWM:34
22:51:44.068 -> Water_Temp.:99.72, PWM:34
22:51:44.864 -> Water_Temp.:99.72, PWM:34
22:51:45.630 -> Water_Temp.:99.72, PWM:34
22:51:46.425 -> Water_Temp.:99.72, PWM:35
22:51:47.185 -> Water_Temp.:99.72, PWM:34
22:51:47.983 -> Water_Temp.:99.72, PWM:34
22:51:48.741 -> Water_Temp.:99.61, PWM:34
22:51:49.534 -> Water_Temp.:99.61, PWM:34
22:51:50.326 -> Water_Temp.:99.61, PWM:34
22:51:51.087 -> Water_Temp.:99.61, PWM:34
22:51:51.882 -> Water_Temp.:99.61, PWM:34
22:51:53.406 -> Water_Temp.:99.61, PWM:34
22:51:54.194 -> Water_Temp.:99.61, PWM:34
22:51:54.958 -> Water_Temp.:99.61, PWM:34
22:51:55.754 -> Water_Temp.:99.61, PWM:34
22:51:56.546 -> Water_Temp.:99.61, PWM:34
22:51:57.305 -> Water_Temp.:99.61, PWM:34
22:51:58.095 -> Water_Temp.:99.72, PWM:34
22:51:58.886 -> Water_Temp.:99.61, PWM:34
22:51:59.644 -> Water_Temp.:99.61, PWM:34
22:52:00.439 -> Water_Temp.:99.72, PWM:34
22:52:01.204 -> Water_Temp.:99.61, PWM:34
22:52:01.998 -> Water_Temp.:99.72, PWM:34
22:52:04.315 -> Water_Temp.:99.72, PWM:34
22:52:05.116 -> Water_Temp.:99.72, PWM:34
22:52:05.880 -> Water_Temp.:99.72, PWM:34
22:52:06.675 -> Water_Temp.:99.72, PWM:34
22:52:07.467 -> Water_Temp.:99.72, PWM:34
22:52:08.225 -> Water_Temp.:99.72, PWM:34
22:52:09.021 -> Water_Temp.:99.72, PWM:34
22:52:09.782 -> Water_Temp.:99.72, PWM:34
22:52:10.578 -> Water_Temp.:99.72, PWM:34
22:52:11.337 -> Water_Temp.:99.72, PWM:34
22:52:12.133 -> Water_Temp.:99.72, PWM:34
22:52:12.894 -> Water_Temp.:99.72, PWM:34
22:52:13.691 -> Water_Temp.:99.72, PWM:34
22:52:14.484 -> Water_Temp.:99.72, PWM:34
22:52:15.244 -> Water_Temp.:99.72, PWM:34
22:52:16.040 -> Water_Temp.:99.72, PWM:34
22:52:16.799 -> Water_Temp.:99.72, PWM:34
22:52:17.594 -> Water_Temp.:99.72, PWM:34
22:52:18.354 -> Water_Temp.:99.72, PWM:34
22:52:19.940 -> Water_Temp.:99.72, PWM:34
22:52:20.702 -> Water_Temp.:99.72, PWM:34
22:52:21.498 -> Water_Temp.:99.72, PWM:34
22:52:23.815 -> Water_Temp.:99.72, PWM:34
22:52:26.131 -> Water_Temp.:99.72, PWM:34
22:52:26.922 -> Water_Temp.:99.84, PWM:34
22:52:27.716 -> Water_Temp.:99.84, PWM:35
22:52:28.477 -> Water_Temp.:99.84, PWM:34
22:52:29.270 -> Water_Temp.:99.84, PWM:34
22:52:30.030 -> Water_Temp.:99.84, PWM:34
22:52:30.829 -> Water_Temp.:99.84, PWM:34
22:52:32.353 -> Water_Temp.:99.84, PWM:34
22:52:33.149 -> Water_Temp.:99.84, PWM:34
22:52:33.911 -> Water_Temp.:99.84, PWM:34
22:52:34.710 -> Water_Temp.:99.84, PWM:34
22:52:35.472 -> Water_Temp.:99.84, PWM:34
22:52:37.057 -> Water_Temp.:99.84, PWM:34
22:52:37.820 -> Water_Temp.:99.84, PWM:34
22:52:38.611 -> Water_Temp.:99.84, PWM:34
22:52:39.368 -> Water_Temp.:99.84, PWM:34
22:52:40.159 -> Water_Temp.:99.84, PWM:34
22:52:40.954 -> Water_Temp.:99.84, PWM:34
22:52:41.708 -> Water_Temp.:99.84, PWM:34
22:52:42.502 -> Water_Temp.:99.84, PWM:34
22:52:43.294 -> Water_Temp.:99.84, PWM:34
22:52:44.055 -> Water_Temp.:99.84, PWM:34
22:52:44.852 -> Water_Temp.:99.84, PWM:34
22:52:45.617 -> Water_Temp.:99.84, PWM:34
22:52:46.409 -> Water_Temp.:99.84, PWM:34
22:52:47.200 -> Water_Temp.:99.84, PWM:34
22:52:47.959 -> Water_Temp.:99.84, PWM:34
22:52:48.751 -> Water_Temp.:99.84, PWM:34
22:52:49.515 -> Water_Temp.:99.84, PWM:34
22:52:50.310 -> Water_Temp.:99.84, PWM:34
22:52:51.073 -> Water_Temp.:99.84, PWM:34
22:52:51.868 -> Water_Temp.:99.84, PWM:34
22:52:52.629 -> Water_Temp.:99.84, PWM:34
22:52:53.421 -> Water_Temp.:99.84, PWM:34
22:52:54.214 -> Water_Temp.:99.84, PWM:34
22:52:54.971 -> Water_Temp.:99.84, PWM:34
22:52:56.554 -> Water_Temp.:99.84, PWM:34
22:52:57.315 -> Water_Temp.:99.84, PWM:34
22:52:58.111 -> Water_Temp.:99.84, PWM:34
22:52:58.868 -> Water_Temp.:99.84, PWM:34
22:52:59.662 -> Water_Temp.:99.84, PWM:34
22:53:00.457 -> Water_Temp.:99.84, PWM:34
22:53:01.220 -> Water_Temp.:99.84, PWM:34
22:53:02.015 -> Water_Temp.:99.84, PWM:34
22:53:03.576 -> Water_Temp.:99.84, PWM:34
22:53:04.338 -> Water_Temp.:99.84, PWM:34
22:53:05.105 -> Water_Temp.:99.84, PWM:34
22:53:05.900 -> Water_Temp.:99.84, PWM:34
22:53:06.696 -> Water_Temp.:99.84, PWM:34
22:53:07.456 -> Water_Temp.:99.84, PWM:34
22:53:08.250 -> Water_Temp.:99.84, PWM:34
22:53:09.014 -> Water_Temp.:99.84, PWM:34
22:53:09.811 -> Water_Temp.:99.84, PWM:34
22:53:10.574 -> Water_Temp.:99.84, PWM:34
22:53:11.371 -> Water_Temp.:99.84, PWM:34
22:53:12.130 -> Water_Temp.:99.84, PWM:34
22:53:12.924 -> Water_Temp.:99.84, PWM:34
22:53:13.683 -> Water_Temp.:99.84, PWM:34
22:53:14.477 -> Water_Temp.:99.84, PWM:34
22:53:15.274 -> Water_Temp.:99.84, PWM:34
22:53:16.034 -> Water_Temp.:99.84, PWM:34
22:53:18.352 -> Water_Temp.:99.84, PWM:34
22:53:19.110 -> Water_Temp.:99.84, PWM:34
22:53:19.902 -> Water_Temp.:99.84, PWM:34
22:53:20.695 -> Water_Temp.:99.84, PWM:34
22:53:23.006 -> Water_Temp.:99.95, PWM:34
22:53:23.766 -> Water_Temp.:99.95, PWM:34
22:53:25.321 -> Water_Temp.:99.95, PWM:34
22:53:26.118 -> Water_Temp.:99.95, PWM:34
22:53:26.907 -> Water_Temp.:99.95, PWM:34
22:53:27.667 -> Water_Temp.:99.95, PWM:34
22:53:28.465 -> Water_Temp.:99.95, PWM:34
22:53:29.220 -> Water_Temp.:99.95, PWM:34
22:53:30.015 -> Water_Temp.:99.95, PWM:34
22:53:30.813 -> Water_Temp.:100.06, PWM:34
22:53:31.574 -> Water_Temp.:100.06, PWM:34
22:53:32.370 -> Water_Temp.:100.06, PWM:34
22:53:33.132 -> Water_Temp.:100.06, PWM:34
22:53:33.930 -> Water_Temp.:100.06, PWM:34
22:53:34.693 -> Water_Temp.:100.06, PWM:34
22:53:36.251 -> Water_Temp.:100.06, PWM:34
22:53:37.011 -> Water_Temp.:100.06, PWM:34
22:53:37.802 -> Water_Temp.:100.06, PWM:34
22:53:38.563 -> Water_Temp.:100.06, PWM:34
22:53:39.355 -> Water_Temp.:100.06, PWM:34
22:53:40.153 -> Water_Temp.:100.06, PWM:34
22:53:40.915 -> Water_Temp.:100.06, PWM:34
22:53:41.710 -> Water_Temp.:100.06, PWM:34
22:53:42.469 -> Water_Temp.:100.06, PWM:34
22:53:43.262 -> Water_Temp.:100.06, PWM:34
22:53:44.054 -> Water_Temp.:100.06, PWM:34
22:53:44.814 -> Water_Temp.:100.06, PWM:34
22:53:45.611 -> Water_Temp.:100.06, PWM:34
22:53:46.367 -> Water_Temp.:100.06, PWM:34
22:53:47.159 -> Water_Temp.:100.06, PWM:34
22:53:47.952 -> Water_Temp.:100.06, PWM:34
22:53:48.710 -> Water_Temp.:100.06, PWM:34
22:53:49.509 -> Water_Temp.:100.06, PWM:34
22:53:50.272 -> Water_Temp.:100.06, PWM:34
22:53:51.071 -> Water_Temp.:100.06, PWM:34
22:53:51.829 -> Water_Temp.:100.06, PWM:34
22:53:52.626 -> Water_Temp.:100.06, PWM:34
22:53:53.384 -> Water_Temp.:100.06, PWM:34
22:53:54.179 -> Water_Temp.:100.06, PWM:34
22:53:54.943 -> Water_Temp.:100.06, PWM:34
22:53:55.740 -> Water_Temp.:100.06, PWM:34
22:53:56.531 -> Water_Temp.:100.06, PWM:34
22:53:57.294 -> Water_Temp.:100.06, PWM:34
22:53:58.087 -> Water_Temp.:100.06, PWM:34
22:53:58.851 -> Water_Temp.:100.06, PWM:34
22:53:59.643 -> Water_Temp.:100.06, PWM:34
22:54:00.400 -> Water_Temp.:100.06, PWM:34
22:54:01.196 -> Water_Temp.:100.06, PWM:34
22:54:01.988 -> Water_Temp.:100.06, PWM:34
22:54:03.543 -> Water_Temp.:103.66, PWM:34
22:54:04.300 -> Water_Temp.:101.30, PWM:34
22:54:05.098 -> Water_Temp.:100.96, PWM:34
22:54:05.860 -> Water_Temp.:100.74, PWM:34
22:54:06.656 -> Water_Temp.:100.62, PWM:34
22:54:07.450 -> Water_Temp.:100.51, PWM:34
22:54:08.214 -> Water_Temp.:100.40, PWM:34
22:54:09.006 -> Water_Temp.:100.40, PWM:34
22:54:09.765 -> Water_Temp.:100.40, PWM:34
22:54:10.559 -> Water_Temp.:100.40, PWM:34
22:54:12.120 -> Water_Temp.:100.29, PWM:34
22:54:12.911 -> Water_Temp.:100.29, PWM:34
22:54:13.669 -> Water_Temp.:100.29, PWM:34
22:54:14.465 -> Water_Temp.:100.29, PWM:34
22:54:15.228 -> Water_Temp.:100.29, PWM:34
22:54:16.026 -> Water_Temp.:100.29, PWM:34
22:54:16.783 -> Water_Temp.:100.29, PWM:34
22:54:17.580 -> Water_Temp.:100.29, PWM:34
22:54:18.340 -> Water_Temp.:100.29, PWM:34
22:54:19.133 -> Water_Temp.:100.29, PWM:34
22:54:19.929 -> Water_Temp.:100.29, PWM:34
22:54:20.688 -> Water_Temp.:100.29, PWM:34
22:54:21.486 -> Water_Temp.:100.29, PWM:34
22:54:22.247 -> Water_Temp.:100.29, PWM:34
22:54:23.039 -> Water_Temp.:100.29, PWM:34
22:54:23.798 -> Water_Temp.:100.29, PWM:34
22:54:24.590 -> Water_Temp.:100.29, PWM:34
22:54:26.152 -> Water_Temp.:100.29, PWM:34
22:54:26.916 -> Water_Temp.:100.17, PWM:34
22:54:27.711 -> Water_Temp.:100.29, PWM:34
22:54:28.473 -> Water_Temp.:100.29, PWM:34
22:54:29.273 -> Water_Temp.:100.29, PWM:34
22:54:30.032 -> Water_Temp.:100.29, PWM:34
22:54:30.826 -> Water_Temp.:100.29, PWM:34
22:54:31.584 -> Water_Temp.:100.29, PWM:34
22:54:32.376 -> Water_Temp.:100.29, PWM:34
22:54:33.169 -> Water_Temp.:100.29, PWM:34
22:54:33.926 -> Water_Temp.:100.29, PWM:34
22:54:34.722 -> Water_Temp.:100.29, PWM:34
22:54:35.483 -> Water_Temp.:100.29, PWM:34
22:54:36.277 -> Water_Temp.:100.29, PWM:34
22:54:37.071 -> Water_Temp.:100.29, PWM:34
22:54:37.830 -> Water_Temp.:100.29, PWM:34
22:54:38.625 -> Water_Temp.:100.29, PWM:34
22:54:39.384 -> Water_Temp.:100.29, PWM:34
22:54:40.180 -> Water_Temp.:100.29, PWM:34
22:54:40.942 -> Water_Temp.:100.40, PWM:34
22:54:41.735 -> Water_Temp.:100.29, PWM:34
22:54:42.530 -> Water_Temp.:100.29, PWM:34
22:54:43.290 -> Water_Temp.:100.29, PWM:34
22:54:44.086 -> Water_Temp.:100.29, PWM:34
22:54:44.847 -> Water_Temp.:100.29, PWM:34
22:54:45.644 -> Water_Temp.:100.29, PWM:34
22:54:46.399 -> Water_Temp.:100.29, PWM:34
22:54:47.197 -> Water_Temp.:100.29, PWM:34
22:54:47.987 -> Water_Temp.:100.29, PWM:34
22:54:48.748 -> Water_Temp.:100.29, PWM:34
22:54:49.543 -> Water_Temp.:101.86, PWM:34
22:54:50.303 -> Water_Temp.:103.89, PWM:34
22:54:51.097 -> Water_Temp.:101.52, PWM:34
22:54:51.890 -> Water_Temp.:101.19, PWM:34
22:54:52.648 -> Water_Temp.:100.96, PWM:34
22:54:53.442 -> Water_Temp.:100.85, PWM:34
22:54:54.202 -> Water_Temp.:100.74, PWM:34
22:54:55.000 -> Water_Temp.:100.74, PWM:34
22:54:55.766 -> Water_Temp.:100.62, PWM:34
22:54:56.559 -> Water_Temp.:100.62, PWM:34
22:54:57.322 -> Water_Temp.:100.51, PWM:34
22:54:58.118 -> Water_Temp.:100.51, PWM:34
22:54:58.911 -> Water_Temp.:100.51, PWM:34
22:55:00.431 -> Water_Temp.:100.51, PWM:34
22:55:01.196 -> Water_Temp.:100.51, PWM:34
22:55:01.994 -> Water_Temp.:100.40, PWM:34
22:55:02.754 -> Water_Temp.:100.40, PWM:34
22:55:03.553 -> Water_Temp.:100.40, PWM:34
22:55:04.348 -> Water_Temp.:100.40, PWM:34
22:55:05.111 -> Water_Temp.:100.40, PWM:34
22:55:05.907 -> Water_Temp.:100.40, PWM:34
22:55:06.664 -> Water_Temp.:100.40, PWM:34
22:55:07.458 -> Water_Temp.:100.40, PWM:34
22:55:08.216 -> Water_Temp.:100.40, PWM:34
22:55:09.009 -> Water_Temp.:100.40, PWM:34
22:55:09.802 -> Water_Temp.:100.40, PWM:34
22:55:10.559 -> Water_Temp.:100.40, PWM:34
22:55:11.348 -> Water_Temp.:100.40, PWM:34
22:55:12.139 -> Water_Temp.:100.40, PWM:34
22:55:12.897 -> Water_Temp.:100.51, PWM:34
22:55:13.692 -> Water_Temp.:100.40, PWM:34
22:55:14.486 -> Water_Temp.:100.51, PWM:34
22:55:15.247 -> Water_Temp.:100.40, PWM:34
22:55:16.045 -> Water_Temp.:100.51, PWM:34
22:55:16.807 -> Water_Temp.:100.51, PWM:34
22:55:17.603 -> Water_Temp.:100.51, PWM:34
22:55:18.358 -> Water_Temp.:100.51, PWM:34
22:55:19.155 -> Water_Temp.:100.51, PWM:34
22:55:19.948 -> Water_Temp.:100.51, PWM:34
22:55:20.709 -> Water_Temp.:100.51, PWM:34
22:55:21.502 -> Water_Temp.:100.51, PWM:34
22:55:22.257 -> Water_Temp.:100.51, PWM:34
22:55:23.047 -> Water_Temp.:100.51, PWM:34
22:55:23.837 -> Water_Temp.:100.51, PWM:34
22:55:24.601 -> Water_Temp.:100.51, PWM:34
22:55:25.398 -> Water_Temp.:100.51, PWM:34
22:55:26.160 -> Water_Temp.:102.09, PWM:34
22:55:26.956 -> Water_Temp.:104.00, PWM:34
22:55:27.714 -> Water_Temp.:101.75, PWM:34
22:55:28.515 -> Water_Temp.:101.30, PWM:34
22:55:29.275 -> Water_Temp.:101.07, PWM:34
22:55:30.071 -> Water_Temp.:100.96, PWM:34
22:55:30.862 -> Water_Temp.:100.85, PWM:34
22:55:31.622 -> Water_Temp.:100.85, PWM:34
22:55:32.418 -> Water_Temp.:100.74, PWM:34
22:55:33.178 -> Water_Temp.:100.74, PWM:34
22:55:33.978 -> Water_Temp.:100.74, PWM:34
22:55:34.742 -> Water_Temp.:100.62, PWM:34
22:55:35.537 -> Water_Temp.:100.62, PWM:34
22:55:36.297 -> Water_Temp.:100.62, PWM:34
22:55:37.889 -> Water_Temp.:100.62, PWM:34
22:55:38.646 -> Water_Temp.:100.62, PWM:34
22:55:39.440 -> Water_Temp.:100.51, PWM:34
22:55:40.202 -> Water_Temp.:100.51, PWM:34
22:55:40.998 -> Water_Temp.:100.51, PWM:34
22:55:41.754 -> Water_Temp.:100.51, PWM:34
22:55:42.548 -> Water_Temp.:100.51, PWM:34
22:55:43.341 -> Water_Temp.:100.51, PWM:34
22:55:44.859 -> Water_Temp.:100.51, PWM:34
22:55:45.657 -> Water_Temp.:100.51, PWM:34
22:55:46.417 -> Water_Temp.:100.51, PWM:34
22:55:47.209 -> Water_Temp.:100.51, PWM:34
22:55:48.001 -> Water_Temp.:100.51, PWM:34
22:55:48.762 -> Water_Temp.:100.51, PWM:34
22:55:49.554 -> Water_Temp.:100.51, PWM:34
22:55:50.314 -> Water_Temp.:100.51, PWM:34
22:55:51.105 -> Water_Temp.:100.51, PWM:34
22:55:51.895 -> Water_Temp.:100.51, PWM:34
22:55:52.657 -> Water_Temp.:100.51, PWM:34
22:55:53.453 -> Water_Temp.:100.51, PWM:34
22:55:55.016 -> Water_Temp.:100.51, PWM:34
22:55:55.781 -> Water_Temp.:100.51, PWM:34
22:55:56.578 -> Water_Temp.:100.51, PWM:34
22:55:57.335 -> Water_Temp.:100.51, PWM:34
22:55:58.127 -> Water_Temp.:100.51, PWM:34
22:55:58.921 -> Water_Temp.:100.51, PWM:34
22:55:59.681 -> Water_Temp.:100.51, PWM:34
22:56:00.473 -> Water_Temp.:100.51, PWM:34
22:56:01.235 -> Water_Temp.:100.62, PWM:34
22:56:02.028 -> Water_Temp.:100.62, PWM:34
22:56:02.823 -> Water_Temp.:100.62, PWM:34
22:56:03.584 -> Water_Temp.:100.62, PWM:34
22:56:04.375 -> Water_Temp.:100.62, PWM:34
22:56:05.138 -> Water_Temp.:100.62, PWM:34
22:56:05.932 -> Water_Temp.:100.62, PWM:34
22:56:06.724 -> Water_Temp.:100.62, PWM:34
22:56:08.241 -> Water_Temp.:100.62, PWM:34
22:56:09.034 -> Water_Temp.:100.62, PWM:34
22:56:10.550 -> Water_Temp.:100.62, PWM:34
22:56:11.349 -> Water_Temp.:100.62, PWM:34
22:56:12.110 -> Water_Temp.:100.62, PWM:34
22:56:12.902 -> Water_Temp.:100.51, PWM:34
22:56:13.697 -> Water_Temp.:100.62, PWM:34
22:56:14.457 -> Water_Temp.:100.62, PWM:34
22:56:15.254 -> Water_Temp.:100.51, PWM:34
22:56:16.015 -> Water_Temp.:100.62, PWM:34
22:56:16.807 -> Water_Temp.:100.62, PWM:34
22:56:17.572 -> Water_Temp.:100.51, PWM:34
22:56:18.366 -> Water_Temp.:100.62, PWM:34
22:56:19.160 -> Water_Temp.:100.62, PWM:34
22:56:19.917 -> Water_Temp.:100.62, PWM:34
22:56:20.714 -> Water_Temp.:100.62, PWM:34
22:56:21.473 -> Water_Temp.:100.62, PWM:34
22:56:22.266 -> Water_Temp.:100.62, PWM:34
22:56:23.027 -> Water_Temp.:100.62, PWM:34
22:56:23.820 -> Water_Temp.:100.62, PWM:34
22:56:24.619 -> Water_Temp.:100.62, PWM:34
22:56:25.375 -> Water_Temp.:100.62, PWM:34
22:56:26.172 -> Water_Temp.:100.62, PWM:34
22:56:26.930 -> Water_Temp.:100.62, PWM:34
22:56:28.487 -> Water_Temp.:100.62, PWM:34
22:56:29.280 -> Water_Temp.:100.62, PWM:34
22:56:30.039 -> Water_Temp.:100.62, PWM:34
22:56:30.836 -> Water_Temp.:100.62, PWM:34
22:56:31.596 -> Water_Temp.:100.62, PWM:34
22:56:32.391 -> Water_Temp.:100.62, PWM:34
22:56:33.150 -> Water_Temp.:100.62, PWM:34
22:56:34.708 -> Water_Temp.:100.62, PWM:34
22:56:35.506 -> Water_Temp.:100.62, PWM:34
22:56:36.301 -> Water_Temp.:100.62, PWM:34
22:56:37.061 -> Water_Temp.:100.62, PWM:34
22:56:37.858 -> Water_Temp.:100.62, PWM:34
22:56:38.610 -> Water_Temp.:100.62, PWM:34
22:56:39.404 -> Water_Temp.:100.62, PWM:34
22:56:40.200 -> Water_Temp.:100.62, PWM:34
22:56:40.961 -> Water_Temp.:100.62, PWM:34
22:56:41.757 -> Water_Temp.:100.62, PWM:34
22:56:42.516 -> Water_Temp.:100.62, PWM:34
22:56:43.313 -> Water_Temp.:100.62, PWM:34
22:56:44.072 -> Water_Temp.:100.62, PWM:34
22:56:45.632 -> Water_Temp.:100.62, PWM:34
22:56:46.428 -> Water_Temp.:100.62, PWM:34
22:56:47.190 -> Water_Temp.:100.62, PWM:34
22:56:47.986 -> Water_Temp.:100.62, PWM:34
22:56:48.776 -> Water_Temp.:100.62, PWM:34
22:56:49.536 -> Water_Temp.:100.62, PWM:34
22:56:50.331 -> Water_Temp.:100.62, PWM:34
22:56:51.095 -> Water_Temp.:100.62, PWM:34
22:56:51.888 -> Water_Temp.:100.62, PWM:34
22:56:52.647 -> Water_Temp.:100.62, PWM:34
22:56:53.440 -> Water_Temp.:100.62, PWM:34
22:56:54.237 -> Water_Temp.:100.62, PWM:34
22:56:55.000 -> Water_Temp.:100.62, PWM:34
22:56:55.793 -> Water_Temp.:100.62, PWM:34
22:56:56.554 -> Water_Temp.:100.62, PWM:34
22:56:57.350 -> Water_Temp.:100.62, PWM:34
22:56:58.111 -> Water_Temp.:100.62, PWM:34
22:56:59.699 -> Water_Temp.:100.62, PWM:35
22:57:00.458 -> Water_Temp.:100.62, PWM:39
22:57:01.253 -> Water_Temp.:100.62, PWM:39
22:57:02.014 -> Water_Temp.:100.62, PWM:39
22:57:02.807 -> Water_Temp.:100.62, PWM:36
22:57:03.570 -> Water_Temp.:100.62, PWM:36
22:57:04.361 -> Water_Temp.:100.62, PWM:36
22:57:05.155 -> Water_Temp.:100.62, PWM:36
22:57:05.913 -> Water_Temp.:100.62, PWM:36
22:57:06.707 -> Water_Temp.:100.62, PWM:36
22:57:07.468 -> Water_Temp.:100.62, PWM:36
22:57:08.261 -> Water_Temp.:100.62, PWM:37
22:57:09.056 -> Water_Temp.:100.62, PWM:36
22:57:09.818 -> Water_Temp.:100.62, PWM:37
22:57:10.613 -> Water_Temp.:100.62, PWM:37
22:57:11.375 -> Water_Temp.:100.62, PWM:37
22:57:12.167 -> Water_Temp.:100.62, PWM:37
22:57:12.959 -> Water_Temp.:100.62, PWM:37
22:57:13.719 -> Water_Temp.:100.62, PWM:37
22:57:14.515 -> Water_Temp.:100.62, PWM:37
22:57:16.067 -> Water_Temp.:100.62, PWM:37
22:57:16.828 -> Water_Temp.:100.62, PWM:37
22:57:17.620 -> Water_Temp.:100.62, PWM:37
22:57:18.378 -> Water_Temp.:100.74, PWM:37
22:57:19.177 -> Water_Temp.:100.74, PWM:37
22:57:19.938 -> Water_Temp.:100.74, PWM:37
22:57:20.728 -> Water_Temp.:100.74, PWM:37
22:57:21.516 -> Water_Temp.:100.74, PWM:37
22:57:22.274 -> Water_Temp.:100.74, PWM:37
22:57:23.070 -> Water_Temp.:100.74, PWM:37
22:57:24.594 -> Water_Temp.:100.74, PWM:37
22:57:25.393 -> Water_Temp.:100.74, PWM:37
22:57:26.153 -> Water_Temp.:100.74, PWM:37
22:57:26.948 -> Water_Temp.:100.74, PWM:37
22:57:27.742 -> Water_Temp.:100.74, PWM:37
22:57:28.500 -> Water_Temp.:100.74, PWM:37
22:57:30.819 -> Water_Temp.:100.74, PWM:37
22:57:31.614 -> Water_Temp.:100.74, PWM:37
22:57:32.375 -> Water_Temp.:100.74, PWM:37
22:57:33.168 -> Water_Temp.:100.74, PWM:37
22:57:33.927 -> Water_Temp.:100.74, PWM:37
22:57:34.726 -> Water_Temp.:100.74, PWM:37
22:57:35.489 -> Water_Temp.:100.74, PWM:37
22:57:36.283 -> Water_Temp.:100.74, PWM:37
22:57:37.076 -> Water_Temp.:100.74, PWM:37
22:57:37.837 -> Water_Temp.:100.74, PWM:37
22:57:38.630 -> Water_Temp.:100.74, PWM:37
22:57:39.391 -> Water_Temp.:100.74, PWM:37
22:57:40.182 -> Water_Temp.:100.74, PWM:37
22:57:40.978 -> Water_Temp.:100.74, PWM:37
22:57:41.734 -> Water_Temp.:100.74, PWM:37
22:57:42.527 -> Water_Temp.:100.74, PWM:37
22:57:43.286 -> Water_Temp.:100.74, PWM:37
22:57:44.077 -> Water_Temp.:100.74, PWM:37
22:57:44.872 -> Water_Temp.:100.74, PWM:37
22:57:45.633 -> Water_Temp.:100.74, PWM:37
22:57:46.427 -> Water_Temp.:100.74, PWM:37
22:57:47.187 -> Water_Temp.:100.74, PWM:37
22:57:47.984 -> Water_Temp.:100.74, PWM:37
22:57:48.745 -> Water_Temp.:100.74, PWM:37
22:57:49.542 -> Water_Temp.:100.74, PWM:37
22:57:50.337 -> Water_Temp.:100.74, PWM:37
22:57:51.098 -> Water_Temp.:100.74, PWM:37
22:57:51.896 -> Water_Temp.:100.74, PWM:37
22:57:52.658 -> Water_Temp.:100.74, PWM:37
22:57:53.455 -> Water_Temp.:100.74, PWM:37
22:57:54.217 -> Water_Temp.:100.74, PWM:37
22:57:55.777 -> Water_Temp.:100.74, PWM:37
22:57:56.536 -> Water_Temp.:100.74, PWM:37
22:57:57.328 -> Water_Temp.:100.74, PWM:37
22:57:58.090 -> Water_Temp.:100.74, PWM:37
22:57:59.644 -> Water_Temp.:100.74, PWM:37
22:58:00.403 -> Water_Temp.:100.74, PWM:37
22:58:01.205 -> Water_Temp.:100.74, PWM:37
22:58:01.969 -> Water_Temp.:100.74, PWM:37
22:58:02.760 -> Water_Temp.:100.74, PWM:37
22:58:03.517 -> Water_Temp.:100.74, PWM:37
22:58:04.309 -> Water_Temp.:100.74, PWM:37
22:58:05.104 -> Water_Temp.:100.74, PWM:37
22:58:05.860 -> Water_Temp.:100.74, PWM:37
22:58:06.656 -> Water_Temp.:100.74, PWM:37
22:58:07.450 -> Water_Temp.:100.74, PWM:37
22:58:08.207 -> Water_Temp.:100.74, PWM:37
22:58:09.002 -> Water_Temp.:100.74, PWM:37
22:58:09.759 -> Water_Temp.:100.74, PWM:37
22:58:10.556 -> Water_Temp.:100.74, PWM:37
22:58:11.320 -> Water_Temp.:100.74, PWM:37
22:58:12.872 -> Water_Temp.:100.74, PWM:37
22:58:13.668 -> Water_Temp.:100.74, PWM:37
22:58:14.428 -> Water_Temp.:100.74, PWM:37
22:58:15.228 -> Water_Temp.:100.74, PWM:37
22:58:15.988 -> Water_Temp.:100.74, PWM:37
22:58:16.784 -> Water_Temp.:100.74, PWM:37
22:58:17.547 -> Water_Temp.:100.74, PWM:37
22:58:19.130 -> Water_Temp.:100.74, PWM:37
22:58:19.891 -> Water_Temp.:100.74, PWM:37
22:58:20.685 -> Water_Temp.:100.74, PWM:37
22:58:21.448 -> Water_Temp.:100.74, PWM:37
22:58:22.245 -> Water_Temp.:100.74, PWM:37
22:58:23.003 -> Water_Temp.:100.74, PWM:37
22:58:23.799 -> Water_Temp.:100.74, PWM:37
22:58:24.564 -> Water_Temp.:100.85, PWM:37
22:58:25.365 -> Water_Temp.:100.74, PWM:37
22:58:26.123 -> Water_Temp.:100.85, PWM:37
22:58:26.920 -> Water_Temp.:100.74, PWM:37
22:58:27.683 -> Water_Temp.:100.85, PWM:37
22:58:28.478 -> Water_Temp.:100.85, PWM:37
22:58:29.240 -> Water_Temp.:100.85, PWM:37
22:58:30.035 -> Water_Temp.:100.85, PWM:37
22:58:30.831 -> Water_Temp.:100.85, PWM:37
22:58:31.592 -> Water_Temp.:100.85, PWM:37
22:58:32.384 -> Water_Temp.:100.85, PWM:37
22:58:33.146 -> Water_Temp.:100.85, PWM:37
22:58:33.938 -> Water_Temp.:100.85, PWM:37
22:58:34.729 -> Water_Temp.:100.85, PWM:37
22:58:35.491 -> Water_Temp.:100.85, PWM:37
22:58:36.287 -> Water_Temp.:100.85, PWM:37
22:58:37.047 -> Water_Temp.:100.85, PWM:37
22:58:37.837 -> Water_Temp.:100.85, PWM:37
22:58:38.597 -> Water_Temp.:100.85, PWM:37
22:58:39.390 -> Water_Temp.:100.85, PWM:37
22:58:40.189 -> Water_Temp.:100.85, PWM:37
22:58:40.954 -> Water_Temp.:100.85, PWM:37
22:58:41.745 -> Water_Temp.:100.85, PWM:37
22:58:42.501 -> Water_Temp.:100.85, PWM:37
22:58:43.293 -> Water_Temp.:100.96, PWM:37
22:58:44.087 -> Water_Temp.:100.96, PWM:37
22:58:44.850 -> Water_Temp.:100.96, PWM:37
22:58:45.645 -> Water_Temp.:100.96, PWM:37
22:58:46.432 -> Water_Temp.:100.96, PWM:37
22:58:47.194 -> Water_Temp.:100.96, PWM:37
22:58:47.987 -> Water_Temp.:100.96, PWM:37
22:58:48.749 -> Water_Temp.:100.96, PWM:37
22:58:49.549 -> Water_Temp.:100.96, PWM:37
22:58:50.306 -> Water_Temp.:100.96, PWM:37
22:58:51.862 -> Water_Temp.:100.96, PWM:37
22:58:52.657 -> Water_Temp.:100.96, PWM:37
22:58:53.417 -> Water_Temp.:100.96, PWM:37
22:58:54.212 -> Water_Temp.:100.96, PWM:37
22:58:55.004 -> Water_Temp.:100.96, PWM:37
22:58:55.763 -> Water_Temp.:100.96, PWM:37
22:58:57.315 -> Water_Temp.:100.96, PWM:37
22:58:58.105 -> Water_Temp.:100.96, PWM:37
22:58:58.898 -> Water_Temp.:100.96, PWM:37
22:58:59.660 -> Water_Temp.:100.96, PWM:37
22:59:00.455 -> Water_Temp.:100.96, PWM:37
22:59:01.219 -> Water_Temp.:100.96, PWM:37
22:59:02.010 -> Water_Temp.:100.96, PWM:37
22:59:02.805 -> Water_Temp.:100.96, PWM:37
22:59:03.564 -> Water_Temp.:100.96, PWM:37
22:59:04.357 -> Water_Temp.:100.96, PWM:37
22:59:05.120 -> Water_Temp.:100.96, PWM:37
22:59:06.674 -> Water_Temp.:100.96, PWM:37
22:59:07.471 -> Water_Temp.:100.96, PWM:37
22:59:08.267 -> Water_Temp.:100.96, PWM:37
22:59:09.026 -> Water_Temp.:101.07, PWM:37
22:59:09.817 -> Water_Temp.:100.96, PWM:37
22:59:10.611 -> Water_Temp.:100.96, PWM:37
22:59:11.372 -> Water_Temp.:101.07, PWM:37
22:59:12.135 -> Water_Temp.:101.07, PWM:37
22:59:12.931 -> Water_Temp.:101.07, PWM:37
22:59:13.694 -> Water_Temp.:101.07, PWM:37
22:59:14.491 -> Water_Temp.:101.07, PWM:37
22:59:15.289 -> Water_Temp.:101.07, PWM:37
22:59:16.809 -> Water_Temp.:101.07, PWM:37
22:59:17.604 -> Water_Temp.:101.07, PWM:37
22:59:18.395 -> Water_Temp.:101.07, PWM:37
22:59:19.156 -> Water_Temp.:101.07, PWM:37
22:59:20.713 -> Water_Temp.:101.07, PWM:37
22:59:21.473 -> Water_Temp.:101.07, PWM:37
22:59:22.271 -> Water_Temp.:101.07, PWM:37
22:59:23.028 -> Water_Temp.:101.07, PWM:37
22:59:23.822 -> Water_Temp.:101.07, PWM:37
22:59:24.584 -> Water_Temp.:101.07, PWM:37
22:59:25.386 -> Water_Temp.:101.07, PWM:37
22:59:26.146 -> Water_Temp.:101.07, PWM:37
22:59:26.937 -> Water_Temp.:101.07, PWM:37
22:59:27.732 -> Water_Temp.:101.07, PWM:37
22:59:29.277 -> Water_Temp.:101.07, PWM:37
22:59:30.073 -> Water_Temp.:101.07, PWM:37
22:59:30.834 -> Water_Temp.:101.07, PWM:37
22:59:31.627 -> Water_Temp.:101.07, PWM:37
22:59:32.384 -> Water_Temp.:101.07, PWM:37
22:59:33.182 -> Water_Temp.:101.07, PWM:37
22:59:33.975 -> Water_Temp.:101.07, PWM:37
22:59:34.739 -> Water_Temp.:101.07, PWM:37
22:59:36.295 -> Water_Temp.:101.07, PWM:37
22:59:37.055 -> Water_Temp.:101.07, PWM:37
22:59:37.849 -> Water_Temp.:101.07, PWM:37
22:59:38.642 -> Water_Temp.:101.07, PWM:37
22:59:39.398 -> Water_Temp.:101.07, PWM:37
22:59:40.188 -> Water_Temp.:101.07, PWM:37
22:59:40.982 -> Water_Temp.:101.07, PWM:37
22:59:41.744 -> Water_Temp.:101.07, PWM:37
22:59:42.538 -> Water_Temp.:101.07, PWM:37
22:59:43.298 -> Water_Temp.:101.07, PWM:37
22:59:44.088 -> Water_Temp.:101.07, PWM:37
22:59:44.882 -> Water_Temp.:101.19, PWM:37
22:59:45.640 -> Water_Temp.:101.19, PWM:37
22:59:46.433 -> Water_Temp.:101.19, PWM:37
22:59:47.192 -> Water_Temp.:101.19, PWM:37
22:59:47.986 -> Water_Temp.:101.19, PWM:37
22:59:48.779 -> Water_Temp.:101.19, PWM:37
22:59:49.539 -> Water_Temp.:101.19, PWM:37
22:59:50.331 -> Water_Temp.:101.19, PWM:37
22:59:51.094 -> Water_Temp.:101.19, PWM:37
22:59:51.885 -> Water_Temp.:101.19, PWM:37
22:59:52.679 -> Water_Temp.:101.19, PWM:37
22:59:54.237 -> Water_Temp.:101.19, PWM:37
22:59:54.999 -> Water_Temp.:101.19, PWM:37
22:59:55.792 -> Water_Temp.:101.19, PWM:37
22:59:56.551 -> Water_Temp.:101.19, PWM:37
22:59:57.346 -> Water_Temp.:101.19, PWM:37
22:59:58.139 -> Water_Temp.:101.19, PWM:37
22:59:58.900 -> Water_Temp.:101.19, PWM:37
22:59:59.695 -> Water_Temp.:101.19, PWM:37
23:00:00.454 -> Water_Temp.:101.19, PWM:37
23:00:01.247 -> Water_Temp.:101.19, PWM:37
23:00:02.041 -> Water_Temp.:101.19, PWM:37
23:00:02.802 -> Water_Temp.:101.19, PWM:37
23:00:03.597 -> Water_Temp.:101.19, PWM:37
23:00:04.357 -> Water_Temp.:101.19, PWM:37
23:00:05.156 -> Water_Temp.:101.19, PWM:37
23:00:05.921 -> Water_Temp.:101.19, PWM:37
23:00:06.718 -> Water_Temp.:101.19, PWM:37
23:00:07.481 -> Water_Temp.:101.19, PWM:37
23:00:08.276 -> Water_Temp.:101.19, PWM:37
23:00:09.041 -> Water_Temp.:101.19, PWM:37
23:00:10.564 -> Water_Temp.:101.19, PWM:37
23:00:11.359 -> Water_Temp.:101.19, PWM:37
23:00:12.154 -> Water_Temp.:101.19, PWM:37
23:00:12.915 -> Water_Temp.:101.19, PWM:37
23:00:13.713 -> Water_Temp.:101.19, PWM:37
23:00:15.234 -> Water_Temp.:101.19, PWM:37
23:00:16.026 -> Water_Temp.:101.19, PWM:37
23:00:18.335 -> Water_Temp.:101.19, PWM:37
23:00:19.126 -> Water_Temp.:101.30, PWM:37
23:00:19.887 -> Water_Temp.:101.30, PWM:37
23:00:20.681 -> Water_Temp.:101.30, PWM:37
23:00:21.475 -> Water_Temp.:101.30, PWM:37
23:00:22.233 -> Water_Temp.:101.30, PWM:37
23:00:23.023 -> Water_Temp.:101.30, PWM:37
23:00:23.783 -> Water_Temp.:101.30, PWM:37
23:00:24.579 -> Water_Temp.:101.30, PWM:37
23:00:25.341 -> Water_Temp.:101.30, PWM:37
23:00:26.140 -> Water_Temp.:101.30, PWM:37
23:00:26.902 -> Water_Temp.:101.30, PWM:37
23:00:27.694 -> Water_Temp.:101.30, PWM:37
23:00:28.490 -> Water_Temp.:101.30, PWM:37
23:00:29.249 -> Water_Temp.:101.30, PWM:37
23:00:30.045 -> Water_Temp.:101.30, PWM:37
23:00:30.805 -> Water_Temp.:101.30, PWM:37
23:00:31.600 -> Water_Temp.:101.30, PWM:37
23:00:32.393 -> Water_Temp.:101.30, PWM:37
23:00:33.153 -> Water_Temp.:101.30, PWM:37
23:00:33.946 -> Water_Temp.:101.30, PWM:37
23:00:34.709 -> Water_Temp.:101.30, PWM:37
23:00:36.265 -> Water_Temp.:101.30, PWM:37
23:00:37.061 -> Water_Temp.:101.30, PWM:37
23:00:37.855 -> Water_Temp.:101.30, PWM:37
23:00:38.613 -> Water_Temp.:101.30, PWM:37
23:00:39.407 -> Water_Temp.:101.41, PWM:37
23:00:40.167 -> Water_Temp.:101.41, PWM:37
23:00:40.967 -> Water_Temp.:101.41, PWM:37
23:00:42.521 -> Water_Temp.:101.41, PWM:37
23:00:43.313 -> Water_Temp.:101.41, PWM:37
23:00:44.076 -> Water_Temp.:101.41, PWM:37
23:00:45.631 -> Water_Temp.:101.41, PWM:37
23:00:46.423 -> Water_Temp.:101.41, PWM:37
23:00:47.186 -> Water_Temp.:101.41, PWM:37
23:00:47.976 -> Water_Temp.:101.41, PWM:37
23:00:48.769 -> Water_Temp.:101.41, PWM:37
23:00:49.528 -> Water_Temp.:101.41, PWM:37
23:00:50.324 -> Water_Temp.:101.41, PWM:37
23:00:51.087 -> Water_Temp.:101.41, PWM:37
23:00:51.879 -> Water_Temp.:101.41, PWM:37
23:00:52.640 -> Water_Temp.:101.41, PWM:37
23:00:53.434 -> Water_Temp.:101.41, PWM:37
23:00:54.225 -> Water_Temp.:101.41, PWM:37
23:00:54.988 -> Water_Temp.:101.41, PWM:37
23:00:55.781 -> Water_Temp.:101.41, PWM:37
23:00:57.296 -> Water_Temp.:101.30, PWM:37
23:00:58.093 -> Water_Temp.:101.30, PWM:37
23:00:58.887 -> Water_Temp.:101.30, PWM:37
23:00:59.650 -> Water_Temp.:101.30, PWM:37
23:01:00.445 -> Water_Temp.:101.30, PWM:37
23:01:01.208 -> Water_Temp.:101.30, PWM:37
23:01:02.002 -> Water_Temp.:101.30, PWM:37
23:01:02.762 -> Water_Temp.:101.30, PWM:37
23:01:03.554 -> Water_Temp.:101.30, PWM:37
23:01:04.316 -> Water_Temp.:101.30, PWM:37
23:01:05.114 -> Water_Temp.:101.30, PWM:37
23:01:05.905 -> Water_Temp.:101.30, PWM:37
23:01:06.666 -> Water_Temp.:101.30, PWM:37
23:01:07.462 -> Water_Temp.:101.30, PWM:37
23:01:08.222 -> Water_Temp.:101.30, PWM:37
23:01:09.016 -> Water_Temp.:101.30, PWM:37
23:01:09.777 -> Water_Temp.:101.30, PWM:37
23:01:10.571 -> Water_Temp.:101.30, PWM:37
23:01:11.334 -> Water_Temp.:101.30, PWM:37
23:01:12.127 -> Water_Temp.:101.30, PWM:37
23:01:12.918 -> Water_Temp.:101.30, PWM:37
23:01:13.678 -> Water_Temp.:101.30, PWM:37
23:01:14.473 -> Water_Temp.:101.30, PWM:37
23:01:15.238 -> Water_Temp.:101.30, PWM:37
23:01:16.034 -> Water_Temp.:101.30, PWM:37
23:01:16.827 -> Water_Temp.:101.30, PWM:37
23:01:17.585 -> Water_Temp.:101.30, PWM:37
23:01:18.383 -> Water_Temp.:101.30, PWM:37
23:01:19.144 -> Water_Temp.:101.30, PWM:37
23:01:19.942 -> Water_Temp.:101.30, PWM:37
23:01:20.705 -> Water_Temp.:101.30, PWM:37
23:01:21.501 -> Water_Temp.:101.30, PWM:37
23:01:22.262 -> Water_Temp.:101.30, PWM:37
23:01:23.059 -> Water_Temp.:101.41, PWM:37
23:01:23.821 -> Water_Temp.:101.41, PWM:37
23:01:24.617 -> Water_Temp.:101.41, PWM:37
23:01:26.175 -> Water_Temp.:101.41, PWM:37
23:01:26.934 -> Water_Temp.:101.41, PWM:37
23:01:27.724 -> Water_Temp.:101.41, PWM:37
23:01:29.272 -> Water_Temp.:101.41, PWM:37
23:01:30.068 -> Water_Temp.:101.41, PWM:37
23:01:30.830 -> Water_Temp.:101.41, PWM:37
23:01:31.626 -> Water_Temp.:101.41, PWM:37
23:01:32.420 -> Water_Temp.:101.41, PWM:37
23:01:33.178 -> Water_Temp.:101.41, PWM:37
23:01:33.975 -> Water_Temp.:101.41, PWM:37
23:01:34.732 -> Water_Temp.:101.41, PWM:37
23:01:35.529 -> Water_Temp.:101.41, PWM:37
23:01:36.288 -> Water_Temp.:101.41, PWM:37
23:01:37.086 -> Water_Temp.:101.41, PWM:37
23:01:37.876 -> Water_Temp.:101.52, PWM:37
23:01:38.635 -> Water_Temp.:101.52, PWM:37
23:01:39.429 -> Water_Temp.:101.52, PWM:37
23:01:40.224 -> Water_Temp.:101.52, PWM:37
23:01:40.989 -> Water_Temp.:101.52, PWM:37
23:01:41.780 -> Water_Temp.:101.52, PWM:37
23:01:42.533 -> Water_Temp.:101.52, PWM:37
23:01:43.329 -> Water_Temp.:101.52, PWM:37
23:01:44.092 -> Water_Temp.:101.52, PWM:37
23:01:45.680 -> Water_Temp.:101.52, PWM:37
23:01:46.434 -> Water_Temp.:101.52, PWM:37
23:01:47.228 -> Water_Temp.:101.52, PWM:37
23:01:48.021 -> Water_Temp.:101.52, PWM:37
23:01:48.781 -> Water_Temp.:101.52, PWM:37
23:01:49.571 -> Water_Temp.:101.52, PWM:37
23:01:50.330 -> Water_Temp.:101.52, PWM:37
23:01:51.129 -> Water_Temp.:101.52, PWM:37
23:01:51.891 -> Water_Temp.:101.52, PWM:37
23:01:52.688 -> Water_Temp.:101.52, PWM:37
23:01:53.480 -> Water_Temp.:101.52, PWM:37
23:01:54.239 -> Water_Temp.:101.52, PWM:37
23:01:55.038 -> Water_Temp.:101.52, PWM:37
23:01:55.797 -> Water_Temp.:101.52, PWM:37
23:01:57.346 -> Water_Temp.:101.52, PWM:37
23:01:58.106 -> Water_Temp.:101.52, PWM:37
23:01:58.900 -> Water_Temp.:101.52, PWM:37
23:01:59.693 -> Water_Temp.:101.52, PWM:37
23:02:01.246 -> Water_Temp.:101.52, PWM:37
23:02:02.003 -> Water_Temp.:101.52, PWM:37
23:02:02.794 -> Water_Temp.:101.52, PWM:37
23:02:03.588 -> Water_Temp.:101.52, PWM:37
23:02:04.349 -> Water_Temp.:101.52, PWM:37
23:02:05.146 -> Water_Temp.:101.52, PWM:37
23:02:05.903 -> Water_Temp.:101.52, PWM:37
23:02:06.697 -> Water_Temp.:101.52, PWM:37
23:02:07.487 -> Water_Temp.:101.52, PWM:37
23:02:08.250 -> Water_Temp.:101.52, PWM:37
23:02:09.047 -> Water_Temp.:101.52, PWM:37
23:02:09.808 -> Water_Temp.:101.52, PWM:37
23:02:10.604 -> Water_Temp.:101.52, PWM:37
23:02:11.370 -> Water_Temp.:101.52, PWM:37
23:02:12.164 -> Water_Temp.:101.52, PWM:37
23:02:12.925 -> Water_Temp.:101.52, PWM:37
23:02:13.717 -> Water_Temp.:103.10, PWM:37
23:02:14.512 -> Water_Temp.:105.12, PWM:37
23:02:15.275 -> Water_Temp.:102.76, PWM:37
23:02:16.072 -> Water_Temp.:102.42, PWM:37
23:02:16.828 -> Water_Temp.:102.20, PWM:37
23:02:17.622 -> Water_Temp.:102.09, PWM:37
23:02:18.382 -> Water_Temp.:101.97, PWM:37
23:02:19.179 -> Water_Temp.:101.97, PWM:37
23:02:20.742 -> Water_Temp.:101.86, PWM:37
23:02:21.505 -> Water_Temp.:101.75, PWM:37
23:02:23.058 -> Water_Temp.:101.75, PWM:37
23:02:24.614 -> Water_Temp.:101.75, PWM:37
23:02:25.378 -> Water_Temp.:101.75, PWM:37
23:02:26.172 -> Water_Temp.:101.75, PWM:37
23:02:27.692 -> Water_Temp.:101.64, PWM:37
23:02:28.489 -> Water_Temp.:101.64, PWM:37
23:02:29.284 -> Water_Temp.:101.64, PWM:37
23:02:30.044 -> Water_Temp.:101.64, PWM:37
23:02:30.843 -> Water_Temp.:101.64, PWM:37
23:02:31.603 -> Water_Temp.:101.64, PWM:37
23:02:32.403 -> Water_Temp.:101.64, PWM:37
23:02:33.160 -> Water_Temp.:101.64, PWM:37
23:02:34.718 -> Water_Temp.:101.64, PWM:37
23:02:35.514 -> Water_Temp.:101.64, PWM:37
23:02:36.271 -> Water_Temp.:101.64, PWM:37
23:02:37.066 -> Water_Temp.:101.64, PWM:37
23:02:37.862 -> Water_Temp.:101.64, PWM:37
23:02:38.623 -> Water_Temp.:101.64, PWM:37
23:02:39.422 -> Water_Temp.:101.64, PWM:37
23:02:40.183 -> Water_Temp.:101.64, PWM:37
23:02:40.979 -> Water_Temp.:101.64, PWM:37
23:02:41.738 -> Water_Temp.:101.64, PWM:37
23:02:42.530 -> Water_Temp.:101.64, PWM:37
23:02:43.292 -> Water_Temp.:101.64, PWM:37
23:02:44.088 -> Water_Temp.:101.64, PWM:37
23:02:44.849 -> Water_Temp.:101.64, PWM:37
23:02:45.647 -> Water_Temp.:101.64, PWM:37
23:02:46.439 -> Water_Temp.:101.64, PWM:37
23:02:48.722 -> Water_Temp.:101.64, PWM:37
23:02:49.480 -> Water_Temp.:101.64, PWM:37
23:02:50.271 -> Water_Temp.:101.64, PWM:37
23:02:51.068 -> Water_Temp.:101.64, PWM:37
23:02:51.824 -> Water_Temp.:101.64, PWM:37
23:02:52.620 -> Water_Temp.:101.64, PWM:37
23:02:53.381 -> Water_Temp.:101.64, PWM:37
23:02:54.937 -> Water_Temp.:101.64, PWM:37
23:02:55.700 -> Water_Temp.:101.64, PWM:37
23:02:56.494 -> Water_Temp.:101.64, PWM:37
23:02:57.286 -> Water_Temp.:101.64, PWM:37
23:02:58.047 -> Water_Temp.:101.64, PWM:37
23:02:58.842 -> Water_Temp.:101.64, PWM:37
23:02:59.601 -> Water_Temp.:101.64, PWM:37
23:03:00.395 -> Water_Temp.:101.64, PWM:37
23:03:01.921 -> Water_Temp.:101.64, PWM:37
23:03:02.715 -> Water_Temp.:101.64, PWM:37
23:03:03.477 -> Water_Temp.:101.64, PWM:37
23:03:04.275 -> Water_Temp.:101.64, PWM:37
23:03:05.037 -> Water_Temp.:101.64, PWM:37
23:03:05.834 -> Water_Temp.:101.75, PWM:37
23:03:06.594 -> Water_Temp.:101.64, PWM:37
23:03:07.389 -> Water_Temp.:101.64, PWM:37
23:03:08.151 -> Water_Temp.:101.64, PWM:37
23:03:08.950 -> Water_Temp.:101.75, PWM:37
23:03:09.712 -> Water_Temp.:101.75, PWM:37
23:03:10.506 -> Water_Temp.:101.75, PWM:37
23:03:11.272 -> Water_Temp.:101.75, PWM:37
23:03:12.068 -> Water_Temp.:101.75, PWM:37
23:03:12.863 -> Water_Temp.:101.75, PWM:37
23:03:13.623 -> Water_Temp.:101.75, PWM:37
23:03:14.414 -> Water_Temp.:101.75, PWM:37
23:03:15.171 -> Water_Temp.:101.75, PWM:37
23:03:15.970 -> Water_Temp.:101.75, PWM:37
23:03:16.730 -> Water_Temp.:101.75, PWM:37
23:03:17.527 -> Water_Temp.:101.75, PWM:37
23:03:18.323 -> Water_Temp.:101.75, PWM:37
23:03:19.083 -> Water_Temp.:101.75, PWM:37
23:03:19.877 -> Water_Temp.:101.75, PWM:37
23:03:20.635 -> Water_Temp.:101.75, PWM:37
23:03:21.428 -> Water_Temp.:101.75, PWM:37
23:03:22.191 -> Water_Temp.:101.75, PWM:37
23:03:23.745 -> Water_Temp.:101.75, PWM:37
23:03:24.508 -> Water_Temp.:101.75, PWM:37
23:03:25.305 -> Water_Temp.:101.75, PWM:37
23:03:26.066 -> Water_Temp.:101.75, PWM:37
23:03:26.864 -> Water_Temp.:101.75, PWM:37
23:03:27.625 -> Water_Temp.:101.75, PWM:37
23:03:28.422 -> Water_Temp.:101.75, PWM:37
23:03:29.181 -> Water_Temp.:101.75, PWM:37
23:03:29.975 -> Water_Temp.:101.75, PWM:37
23:03:30.767 -> Water_Temp.:101.75, PWM:37
23:03:31.529 -> Water_Temp.:101.75, PWM:37
23:03:32.322 -> Water_Temp.:101.75, PWM:37
23:03:33.084 -> Water_Temp.:101.75, PWM:37
23:03:33.878 -> Water_Temp.:101.75, PWM:37
23:03:34.643 -> Water_Temp.:101.75, PWM:37
23:03:35.437 -> Water_Temp.:101.75, PWM:37
23:03:36.200 -> Water_Temp.:101.75, PWM:37
23:03:36.988 -> Water_Temp.:101.75, PWM:37
23:03:37.781 -> Water_Temp.:101.75, PWM:37
23:03:38.573 -> Water_Temp.:101.75, PWM:37
23:03:39.333 -> Water_Temp.:101.75, PWM:37
23:03:40.129 -> Water_Temp.:101.75, PWM:37
23:03:40.892 -> Water_Temp.:101.75, PWM:37
23:03:41.689 -> Water_Temp.:101.75, PWM:37
23:03:42.448 -> Water_Temp.:101.75, PWM:37
23:03:43.243 -> Water_Temp.:101.75, PWM:37
23:03:44.005 -> Water_Temp.:101.75, PWM:37
23:03:44.803 -> Water_Temp.:101.75, PWM:37
23:03:45.565 -> Water_Temp.:101.75, PWM:37
23:03:46.353 -> Water_Temp.:101.75, PWM:37
23:03:47.909 -> Water_Temp.:101.75, PWM:37
23:03:48.669 -> Water_Temp.:101.75, PWM:37
23:03:50.254 -> Water_Temp.:101.75, PWM:37
23:03:51.014 -> Water_Temp.:101.75, PWM:37
23:03:51.805 -> Water_Temp.:101.75, PWM:37
23:03:52.599 -> Water_Temp.:101.75, PWM:37
23:03:53.357 -> Water_Temp.:101.75, PWM:37
23:03:54.152 -> Water_Temp.:101.75, PWM:37
23:03:54.911 -> Water_Temp.:101.75, PWM:37
23:03:55.707 -> Water_Temp.:101.75, PWM:37
23:03:56.499 -> Water_Temp.:101.75, PWM:37
23:03:57.259 -> Water_Temp.:101.86, PWM:37
23:03:58.052 -> Water_Temp.:101.75, PWM:37
23:03:58.811 -> Water_Temp.:101.75, PWM:37
23:03:59.605 -> Water_Temp.:101.75, PWM:37
23:04:00.366 -> Water_Temp.:101.75, PWM:37
23:04:01.162 -> Water_Temp.:101.75, PWM:37
23:04:01.953 -> Water_Temp.:101.86, PWM:37
23:04:02.713 -> Water_Temp.:101.75, PWM:37
23:04:03.507 -> Water_Temp.:101.75, PWM:37
23:04:04.266 -> Water_Temp.:101.75, PWM:37
23:04:05.067 -> Water_Temp.:101.75, PWM:37
23:04:05.829 -> Water_Temp.:101.75, PWM:37
23:04:06.629 -> Water_Temp.:101.75, PWM:37
23:04:08.190 -> Water_Temp.:101.75, PWM:37
23:04:08.948 -> Water_Temp.:101.75, PWM:37
23:04:09.738 -> Water_Temp.:101.75, PWM:37
23:04:10.535 -> Water_Temp.:101.75, PWM:37
23:04:11.299 -> Water_Temp.:101.75, PWM:37
23:04:12.088 -> Water_Temp.:101.86, PWM:37
23:04:13.638 -> Water_Temp.:101.86, PWM:37
23:04:14.429 -> Water_Temp.:101.86, PWM:37
23:04:15.190 -> Water_Temp.:101.86, PWM:37
23:04:15.988 -> Water_Temp.:101.86, PWM:37
23:04:16.746 -> Water_Temp.:101.86, PWM:37
23:04:17.542 -> Water_Temp.:101.86, PWM:37
23:04:18.339 -> Water_Temp.:101.86, PWM:37
23:04:19.100 -> Water_Temp.:101.86, PWM:37
23:04:19.892 -> Water_Temp.:101.86, PWM:37
23:04:20.654 -> Water_Temp.:101.86, PWM:37
23:04:21.452 -> Water_Temp.:101.86, PWM:37
23:04:22.210 -> Water_Temp.:101.86, PWM:37
23:04:23.004 -> Water_Temp.:101.86, PWM:37
23:04:23.768 -> Water_Temp.:101.86, PWM:37
23:04:24.563 -> Water_Temp.:101.86, PWM:37
23:04:25.324 -> Water_Temp.:101.86, PWM:37
23:04:26.120 -> Water_Temp.:101.86, PWM:37
23:04:26.914 -> Water_Temp.:101.86, PWM:37
23:04:27.676 -> Water_Temp.:101.86, PWM:37
23:04:28.471 -> Water_Temp.:101.86, PWM:37
23:04:29.231 -> Water_Temp.:101.86, PWM:37
23:04:30.027 -> Water_Temp.:101.86, PWM:37
23:04:30.792 -> Water_Temp.:101.86, PWM:37
23:04:31.587 -> Water_Temp.:101.86, PWM:37
23:04:33.141 -> Water_Temp.:101.86, PWM:37
23:04:33.905 -> Water_Temp.:101.86, PWM:37
23:04:34.696 -> Water_Temp.:101.86, PWM:37
23:04:35.496 -> Water_Temp.:101.86, PWM:37
23:04:36.253 -> Water_Temp.:101.86, PWM:37
23:04:37.045 -> Water_Temp.:101.86, PWM:37
23:04:37.805 -> Water_Temp.:101.86, PWM:37
23:04:38.597 -> Water_Temp.:101.86, PWM:37
23:04:39.390 -> Water_Temp.:101.86, PWM:37
23:04:40.149 -> Water_Temp.:101.86, PWM:37
23:04:40.942 -> Water_Temp.:101.86, PWM:37
23:04:41.701 -> Water_Temp.:101.86, PWM:37
23:04:42.499 -> Water_Temp.:101.86, PWM:37
23:04:43.293 -> Water_Temp.:101.86, PWM:37
23:04:44.052 -> Water_Temp.:101.86, PWM:37
23:04:44.851 -> Water_Temp.:101.86, PWM:37
23:04:45.612 -> Water_Temp.:101.86, PWM:37
23:04:46.407 -> Water_Temp.:101.86, PWM:37
23:04:47.166 -> Water_Temp.:101.86, PWM:37
23:04:47.960 -> Water_Temp.:101.86, PWM:37
23:04:50.306 -> Water_Temp.:101.86, PWM:37
23:04:51.068 -> Water_Temp.:101.86, PWM:37
23:04:51.863 -> Water_Temp.:101.86, PWM:37
23:04:52.620 -> Water_Temp.:101.86, PWM:37
23:04:53.415 -> Water_Temp.:101.86, PWM:37
23:04:54.210 -> Water_Temp.:101.86, PWM:37
23:04:56.520 -> Water_Temp.:101.86, PWM:37
23:04:57.311 -> Water_Temp.:101.86, PWM:37
23:04:58.066 -> Water_Temp.:101.86, PWM:37
23:04:58.860 -> Water_Temp.:101.86, PWM:37
23:04:59.619 -> Water_Temp.:101.86, PWM:37
23:05:00.412 -> Water_Temp.:101.86, PWM:37
23:05:01.211 -> Water_Temp.:101.86, PWM:37
23:05:01.971 -> Water_Temp.:101.86, PWM:37
23:05:02.766 -> Water_Temp.:101.86, PWM:37
23:05:03.524 -> Water_Temp.:101.86, PWM:37
23:05:05.085 -> Water_Temp.:101.86, PWM:37
23:05:05.879 -> Water_Temp.:101.86, PWM:37
23:05:06.640 -> Water_Temp.:101.86, PWM:37
23:05:07.436 -> Water_Temp.:101.86, PWM:37
23:05:08.230 -> Water_Temp.:101.86, PWM:37
23:05:08.989 -> Water_Temp.:101.86, PWM:37
23:05:09.783 -> Water_Temp.:101.86, PWM:37
23:05:10.542 -> Water_Temp.:101.97, PWM:37
23:05:11.340 -> Water_Temp.:101.97, PWM:37
23:05:12.099 -> Water_Temp.:101.97, PWM:37
23:05:12.892 -> Water_Temp.:101.97, PWM:37
23:05:13.690 -> Water_Temp.:101.97, PWM:37
23:05:14.448 -> Water_Temp.:101.97, PWM:37
23:05:15.248 -> Water_Temp.:101.97, PWM:37
23:05:16.012 -> Water_Temp.:101.97, PWM:37
23:05:16.802 -> Water_Temp.:101.97, PWM:37
23:05:17.562 -> Water_Temp.:101.97, PWM:37
23:05:18.358 -> Water_Temp.:101.97, PWM:37
23:05:19.151 -> Water_Temp.:101.97, PWM:37
23:05:19.909 -> Water_Temp.:101.97, PWM:37
23:05:20.706 -> Water_Temp.:101.97, PWM:37
23:05:21.471 -> Water_Temp.:101.97, PWM:37
23:05:22.266 -> Water_Temp.:101.97, PWM:37
23:05:23.022 -> Water_Temp.:101.97, PWM:37
23:05:23.818 -> Water_Temp.:101.97, PWM:37
23:05:24.584 -> Water_Temp.:101.97, PWM:37
23:05:25.379 -> Water_Temp.:101.97, PWM:37
23:05:26.140 -> Water_Temp.:101.97, PWM:37
23:05:26.931 -> Water_Temp.:101.97, PWM:37
23:05:27.729 -> Water_Temp.:101.97, PWM:37
23:05:28.489 -> Water_Temp.:101.97, PWM:37
23:05:29.283 -> Water_Temp.:101.97, PWM:37
23:05:30.047 -> Water_Temp.:101.97, PWM:37
23:05:30.846 -> Water_Temp.:101.97, PWM:37
23:05:31.604 -> Water_Temp.:101.97, PWM:37
23:05:32.400 -> Water_Temp.:101.97, PWM:37
23:05:33.192 -> Water_Temp.:101.97, PWM:37
23:05:34.714 -> Water_Temp.:101.97, PWM:37
23:05:35.476 -> Water_Temp.:101.97, PWM:37
23:05:36.268 -> Water_Temp.:102.09, PWM:37
23:05:37.063 -> Water_Temp.:101.97, PWM:37
23:05:37.822 -> Water_Temp.:102.09, PWM:37
23:05:38.616 -> Water_Temp.:102.09, PWM:37
23:05:39.374 -> Water_Temp.:102.09, PWM:37
23:05:40.171 -> Water_Temp.:102.09, PWM:37
23:05:40.967 -> Water_Temp.:102.09, PWM:37
23:05:41.723 -> Water_Temp.:102.09, PWM:37
23:05:42.517 -> Water_Temp.:102.09, PWM:37
23:05:43.276 -> Water_Temp.:102.09, PWM:37
23:05:44.069 -> Water_Temp.:102.09, PWM:37
23:05:44.865 -> Water_Temp.:102.09, PWM:37
23:05:45.629 -> Water_Temp.:102.09, PWM:37
23:05:46.421 -> Water_Temp.:102.09, PWM:37
23:05:47.181 -> Water_Temp.:102.09, PWM:37
23:05:47.977 -> Water_Temp.:102.09, PWM:37
23:05:49.499 -> Water_Temp.:102.09, PWM:37
23:05:50.295 -> Water_Temp.:102.09, PWM:37
23:05:51.063 -> Water_Temp.:102.09, PWM:37
23:05:51.862 -> Water_Temp.:102.09, PWM:37
23:05:52.620 -> Water_Temp.:102.09, PWM:37
23:05:53.414 -> Water_Temp.:102.09, PWM:37
23:05:54.175 -> Water_Temp.:102.09, PWM:37
23:05:54.970 -> Water_Temp.:102.09, PWM:37
23:05:55.765 -> Water_Temp.:102.09, PWM:37
23:05:56.524 -> Water_Temp.:102.09, PWM:37
23:05:57.320 -> Water_Temp.:102.09, PWM:37
23:05:58.079 -> Water_Temp.:102.09, PWM:37
23:05:58.870 -> Water_Temp.:102.09, PWM:37
23:05:59.662 -> Water_Temp.:102.09, PWM:37
23:06:00.427 -> Water_Temp.:102.09, PWM:37
23:06:01.988 -> Water_Temp.:102.09, PWM:37
23:06:02.787 -> Water_Temp.:102.09, PWM:37
23:06:03.547 -> Water_Temp.:102.09, PWM:37
23:06:04.343 -> Water_Temp.:102.09, PWM:37
23:06:05.110 -> Water_Temp.:102.09, PWM:37
23:06:05.901 -> Water_Temp.:102.09, PWM:37
23:06:06.661 -> Water_Temp.:102.09, PWM:37
23:06:07.455 -> Water_Temp.:102.09, PWM:37
23:06:08.216 -> Water_Temp.:101.97, PWM:37
23:06:09.012 -> Water_Temp.:101.97, PWM:37
23:06:09.810 -> Water_Temp.:101.97, PWM:37
23:06:10.565 -> Water_Temp.:101.97, PWM:37
23:06:11.362 -> Water_Temp.:101.97, PWM:37
23:06:13.649 -> Water_Temp.:101.97, PWM:37
23:06:14.444 -> Water_Temp.:101.97, PWM:37
23:06:15.208 -> Water_Temp.:101.97, PWM:37
23:06:16.007 -> Water_Temp.:101.97, PWM:37
23:06:16.763 -> Water_Temp.:101.97, PWM:37
23:06:17.556 -> Water_Temp.:101.97, PWM:37
23:06:18.349 -> Water_Temp.:101.97, PWM:37
23:06:19.109 -> Water_Temp.:101.97, PWM:37
23:06:19.905 -> Water_Temp.:101.97, PWM:37
23:06:20.699 -> Water_Temp.:101.97, PWM:37
23:06:21.463 -> Water_Temp.:101.97, PWM:37
23:06:22.257 -> Water_Temp.:101.97, PWM:37
23:06:23.015 -> Water_Temp.:101.97, PWM:37
23:06:23.808 -> Water_Temp.:101.97, PWM:37
23:06:24.567 -> Water_Temp.:101.97, PWM:37
23:06:26.891 -> Water_Temp.:101.97, PWM:37
23:06:27.685 -> Water_Temp.:101.97, PWM:37
23:06:28.443 -> Water_Temp.:101.97, PWM:37
23:06:29.238 -> Water_Temp.:101.97, PWM:37
23:06:30.001 -> Water_Temp.:101.97, PWM:37
23:06:30.799 -> Water_Temp.:101.97, PWM:37
23:06:31.559 -> Water_Temp.:101.97, PWM:37
23:06:32.355 -> Water_Temp.:102.09, PWM:37
23:06:33.149 -> Water_Temp.:101.97, PWM:37
23:06:33.906 -> Water_Temp.:101.97, PWM:37
23:06:34.702 -> Water_Temp.:101.97, PWM:37
23:06:35.465 -> Water_Temp.:101.97, PWM:37
23:06:36.256 -> Water_Temp.:101.97, PWM:37
23:06:37.049 -> Water_Temp.:101.97, PWM:37
23:06:37.810 -> Water_Temp.:102.09, PWM:37
23:06:38.606 -> Water_Temp.:102.09, PWM:37
23:06:39.368 -> Water_Temp.:101.97, PWM:37
23:06:40.162 -> Water_Temp.:102.09, PWM:37
23:06:40.927 -> Water_Temp.:102.09, PWM:37
23:06:41.722 -> Water_Temp.:102.09, PWM:37
23:06:42.485 -> Water_Temp.:102.09, PWM:37
23:06:43.283 -> Water_Temp.:102.09, PWM:37
23:06:44.045 -> Water_Temp.:102.09, PWM:37
23:06:44.842 -> Water_Temp.:102.09, PWM:37
23:06:45.607 -> Water_Temp.:101.97, PWM:37
23:06:46.406 -> Water_Temp.:102.09, PWM:37
23:06:47.167 -> Water_Temp.:101.97, PWM:37
23:06:47.961 -> Water_Temp.:102.09, PWM:37
23:06:48.719 -> Water_Temp.:101.97, PWM:37
23:06:49.513 -> Water_Temp.:102.09, PWM:37
23:06:50.308 -> Water_Temp.:102.09, PWM:37
23:06:51.069 -> Water_Temp.:102.09, PWM:37
23:06:51.861 -> Water_Temp.:101.97, PWM:37
23:06:52.619 -> Water_Temp.:101.97, PWM:37
23:06:53.413 -> Water_Temp.:103.66, PWM:37
23:06:54.210 -> Water_Temp.:105.69, PWM:37
23:06:54.973 -> Water_Temp.:103.21, PWM:37
23:06:55.769 -> Water_Temp.:102.87, PWM:37
23:06:57.284 -> Water_Temp.:102.54, PWM:37
23:06:58.076 -> Water_Temp.:102.42, PWM:37
23:06:58.836 -> Water_Temp.:102.31, PWM:37
23:07:00.395 -> Water_Temp.:102.20, PWM:37
23:07:01.158 -> Water_Temp.:102.20, PWM:37
23:07:01.950 -> Water_Temp.:102.20, PWM:37
23:07:03.470 -> Water_Temp.:102.09, PWM:37
23:07:04.262 -> Water_Temp.:102.09, PWM:37
23:07:05.055 -> Water_Temp.:102.09, PWM:37
23:07:05.816 -> Water_Temp.:102.09, PWM:37
23:07:06.609 -> Water_Temp.:102.09, PWM:37
23:07:07.364 -> Water_Temp.:102.09, PWM:37
23:07:08.160 -> Water_Temp.:102.09, PWM:37
23:07:08.953 -> Water_Temp.:102.09, PWM:37
23:07:09.712 -> Water_Temp.:102.09, PWM:37
23:07:10.508 -> Water_Temp.:102.09, PWM:37
23:07:11.269 -> Water_Temp.:102.09, PWM:37
23:07:12.064 -> Water_Temp.:102.09, PWM:37
23:07:12.828 -> Water_Temp.:102.09, PWM:37
23:07:13.622 -> Water_Temp.:102.09, PWM:37
23:07:14.416 -> Water_Temp.:102.09, PWM:37
23:07:15.937 -> Water_Temp.:102.09, PWM:37
23:07:16.725 -> Water_Temp.:102.09, PWM:37
23:07:17.487 -> Water_Temp.:102.09, PWM:37
23:07:18.283 -> Water_Temp.:102.09, PWM:37
23:07:19.077 -> Water_Temp.:102.09, PWM:37
23:07:19.837 -> Water_Temp.:102.09, PWM:37
23:07:21.391 -> Water_Temp.:102.09, PWM:37
23:07:22.181 -> Water_Temp.:102.09, PWM:37
23:07:22.975 -> Water_Temp.:102.09, PWM:37
23:07:23.733 -> Water_Temp.:102.09, PWM:37
23:07:24.526 -> Water_Temp.:102.09, PWM:37
23:07:25.323 -> Water_Temp.:102.09, PWM:37
23:07:26.081 -> Water_Temp.:102.09, PWM:37
23:07:26.847 -> Water_Temp.:102.09, PWM:37
23:07:27.641 -> Water_Temp.:101.97, PWM:37
23:07:29.173 -> Water_Temp.:101.97, PWM:37
23:07:29.969 -> Water_Temp.:101.97, PWM:37
23:07:30.733 -> Water_Temp.:101.97, PWM:37
23:07:31.532 -> Water_Temp.:101.97, PWM:37
23:07:32.323 -> Water_Temp.:101.97, PWM:37
23:07:33.082 -> Water_Temp.:101.97, PWM:37
23:07:33.875 -> Water_Temp.:101.97, PWM:37
23:07:34.638 -> Water_Temp.:101.97, PWM:37
23:07:35.431 -> Water_Temp.:101.97, PWM:37
23:07:36.192 -> Water_Temp.:101.97, PWM:37
23:07:36.982 -> Water_Temp.:101.97, PWM:37
23:07:37.773 -> Water_Temp.:101.97, PWM:37
23:07:39.319 -> Water_Temp.:101.97, PWM:37
23:07:40.877 -> Water_Temp.:101.97, PWM:37
23:07:41.632 -> Water_Temp.:101.97, PWM:37
23:07:42.427 -> Water_Temp.:101.97, PWM:37
23:07:43.190 -> Water_Temp.:101.97, PWM:37
23:07:43.982 -> Water_Temp.:101.97, PWM:37
23:07:44.779 -> Water_Temp.:101.97, PWM:37
23:07:45.541 -> Water_Temp.:101.97, PWM:37
23:07:46.335 -> Water_Temp.:101.97, PWM:37
23:07:47.852 -> Water_Temp.:101.97, PWM:37
23:07:48.647 -> Water_Temp.:101.97, PWM:37
23:07:49.443 -> Water_Temp.:101.97, PWM:37
23:07:50.199 -> Water_Temp.:101.97, PWM:37
23:07:50.995 -> Water_Temp.:101.97, PWM:37
23:07:52.548 -> Water_Temp.:101.97, PWM:37
23:07:53.338 -> Water_Temp.:101.97, PWM:37
23:07:54.097 -> Water_Temp.:101.97, PWM:37
23:07:54.898 -> Water_Temp.:101.97, PWM:37
23:07:55.659 -> Water_Temp.:101.97, PWM:37
23:07:57.213 -> Water_Temp.:101.97, PWM:37
23:07:57.972 -> Water_Temp.:101.97, PWM:37
23:07:58.763 -> Water_Temp.:101.97, PWM:37
23:07:59.524 -> Water_Temp.:101.97, PWM:37
23:08:00.321 -> Water_Temp.:101.97, PWM:37
23:08:01.085 -> Water_Temp.:101.97, PWM:37
23:08:01.876 -> Water_Temp.:101.97, PWM:37
23:08:02.668 -> Water_Temp.:101.97, PWM:37
23:08:03.427 -> Water_Temp.:101.97, PWM:37
23:08:04.224 -> Water_Temp.:101.97, PWM:37
23:08:04.987 -> Water_Temp.:101.97, PWM:37
23:08:06.574 -> Water_Temp.:101.97, PWM:37
23:08:07.331 -> Water_Temp.:101.97, PWM:37
23:08:08.127 -> Water_Temp.:101.97, PWM:37
23:08:08.882 -> Water_Temp.:101.97, PWM:37
23:08:09.677 -> Water_Temp.:101.97, PWM:37
23:08:10.471 -> Water_Temp.:101.97, PWM:37
23:08:11.233 -> Water_Temp.:101.97, PWM:37
23:08:12.024 -> Water_Temp.:101.97, PWM:37
23:08:12.783 -> Water_Temp.:101.97, PWM:37
23:08:13.580 -> Water_Temp.:101.97, PWM:37
23:08:14.341 -> Water_Temp.:101.97, PWM:37
23:08:15.139 -> Water_Temp.:101.97, PWM:40
23:08:15.932 -> Water_Temp.:101.97, PWM:41
23:08:16.691 -> Water_Temp.:102.09, PWM:41
23:08:17.483 -> Water_Temp.:101.97, PWM:41
23:08:18.241 -> Water_Temp.:102.09, PWM:41
23:08:19.036 -> Water_Temp.:102.09, PWM:41
23:08:20.593 -> Water_Temp.:102.09, PWM:41
23:08:21.354 -> Water_Temp.:102.09, PWM:41
23:08:22.148 -> Water_Temp.:102.09, PWM:41
23:08:22.909 -> Water_Temp.:102.09, PWM:41
23:08:23.701 -> Water_Temp.:102.09, PWM:41
23:08:24.491 -> Water_Temp.:102.09, PWM:41
23:08:25.258 -> Water_Temp.:102.09, PWM:41
23:08:26.053 -> Water_Temp.:102.09, PWM:41
23:08:26.812 -> Water_Temp.:102.09, PWM:41
23:08:27.604 -> Water_Temp.:102.09, PWM:41
23:08:28.364 -> Water_Temp.:102.09, PWM:41
23:08:29.155 -> Water_Temp.:102.09, PWM:41
23:08:29.949 -> Water_Temp.:102.09, PWM:41
23:08:30.711 -> Water_Temp.:102.09, PWM:41
23:08:31.508 -> Water_Temp.:102.09, PWM:41
23:08:32.264 -> Water_Temp.:102.09, PWM:41
23:08:33.059 -> Water_Temp.:102.09, PWM:41
23:08:33.820 -> Water_Temp.:102.09, PWM:41
23:08:34.612 -> Water_Temp.:102.09, PWM:41
23:08:35.407 -> Water_Temp.:101.97, PWM:41
23:08:36.167 -> Water_Temp.:102.09, PWM:41
23:08:36.956 -> Water_Temp.:102.09, PWM:41
23:08:37.750 -> Water_Temp.:102.09, PWM:41
23:08:38.507 -> Water_Temp.:101.97, PWM:41
23:08:39.299 -> Water_Temp.:101.97, PWM:41
23:08:40.095 -> Water_Temp.:101.97, PWM:41
23:08:40.856 -> Water_Temp.:101.97, PWM:41
23:08:41.652 -> Water_Temp.:102.09, PWM:41
23:08:42.413 -> Water_Temp.:102.09, PWM:41
23:08:43.210 -> Water_Temp.:102.09, PWM:41
23:08:43.967 -> Water_Temp.:102.09, PWM:41
23:08:44.760 -> Water_Temp.:102.09, PWM:41
23:08:45.522 -> Water_Temp.:102.09, PWM:41
23:08:46.313 -> Water_Temp.:102.09, PWM:41
23:08:47.105 -> Water_Temp.:103.66, PWM:41
23:08:47.864 -> Water_Temp.:105.69, PWM:41
23:08:48.654 -> Water_Temp.:103.32, PWM:41
23:08:49.450 -> Water_Temp.:102.99, PWM:41
23:08:50.210 -> Water_Temp.:102.76, PWM:41
23:08:51.009 -> Water_Temp.:102.65, PWM:41
23:08:51.771 -> Water_Temp.:102.54, PWM:41
23:08:52.564 -> Water_Temp.:102.42, PWM:41
23:08:54.087 -> Water_Temp.:102.31, PWM:41
23:08:54.883 -> Water_Temp.:102.31, PWM:41
23:08:55.677 -> Water_Temp.:102.31, PWM:41
23:08:57.231 -> Water_Temp.:102.20, PWM:41
23:08:57.990 -> Water_Temp.:102.20, PWM:41
23:08:58.787 -> Water_Temp.:102.20, PWM:41
23:08:59.579 -> Water_Temp.:102.20, PWM:41
23:09:00.340 -> Water_Temp.:102.20, PWM:41
23:09:01.137 -> Water_Temp.:102.20, PWM:41
23:09:01.897 -> Water_Temp.:102.20, PWM:41
23:09:02.695 -> Water_Temp.:102.20, PWM:41
23:09:03.450 -> Water_Temp.:102.20, PWM:41
23:09:04.247 -> Water_Temp.:102.20, PWM:41
23:09:05.006 -> Water_Temp.:102.20, PWM:41
23:09:05.800 -> Water_Temp.:102.09, PWM:41
23:09:06.593 -> Water_Temp.:102.20, PWM:41
23:09:07.358 -> Water_Temp.:102.09, PWM:41
23:09:08.155 -> Water_Temp.:102.09, PWM:41
23:09:08.915 -> Water_Temp.:102.09, PWM:41
23:09:09.706 -> Water_Temp.:102.09, PWM:41
23:09:10.468 -> Water_Temp.:102.20, PWM:41
23:09:11.265 -> Water_Temp.:102.20, PWM:41
23:09:12.785 -> Water_Temp.:102.09, PWM:41
23:09:13.582 -> Water_Temp.:102.09, PWM:41
23:09:14.378 -> Water_Temp.:102.09, PWM:41
23:09:15.139 -> Water_Temp.:102.09, PWM:41
23:09:15.933 -> Water_Temp.:102.09, PWM:41
23:09:16.693 -> Water_Temp.:102.09, PWM:41
23:09:17.487 -> Water_Temp.:102.09, PWM:41
23:09:18.249 -> Water_Temp.:102.09, PWM:41
23:09:19.793 -> Water_Temp.:102.09, PWM:41
23:09:20.587 -> Water_Temp.:102.09, PWM:41
23:09:21.351 -> Water_Temp.:102.09, PWM:41
23:09:22.139 -> Water_Temp.:102.09, PWM:41
23:09:22.930 -> Water_Temp.:102.09, PWM:41
23:09:23.692 -> Water_Temp.:102.09, PWM:41
23:09:24.492 -> Water_Temp.:102.09, PWM:41
23:09:25.257 -> Water_Temp.:102.09, PWM:41
23:09:26.051 -> Water_Temp.:102.20, PWM:41
23:09:26.806 -> Water_Temp.:102.09, PWM:41
23:09:27.602 -> Water_Temp.:102.20, PWM:41
23:09:28.394 -> Water_Temp.:102.20, PWM:41
23:09:29.950 -> Water_Temp.:102.20, PWM:41
23:09:30.708 -> Water_Temp.:102.09, PWM:41
23:09:31.505 -> Water_Temp.:102.20, PWM:41
23:09:32.259 -> Water_Temp.:102.20, PWM:41
23:09:33.056 -> Water_Temp.:102.20, PWM:41
23:09:33.848 -> Water_Temp.:102.20, PWM:41
23:09:34.610 -> Water_Temp.:102.20, PWM:41
23:09:35.407 -> Water_Temp.:102.20, PWM:41
23:09:36.163 -> Water_Temp.:102.20, PWM:41
23:09:36.957 -> Water_Temp.:102.20, PWM:41
23:09:37.750 -> Water_Temp.:102.20, PWM:41
23:09:38.509 -> Water_Temp.:102.20, PWM:41
23:09:39.302 -> Water_Temp.:102.20, PWM:41
23:09:40.058 -> Water_Temp.:102.20, PWM:41
23:09:40.857 -> Water_Temp.:102.20, PWM:41
23:09:41.619 -> Water_Temp.:102.20, PWM:41
23:09:42.414 -> Water_Temp.:102.20, PWM:41
23:09:43.212 -> Water_Temp.:102.20, PWM:41
23:09:43.968 -> Water_Temp.:102.20, PWM:41
23:09:44.758 -> Water_Temp.:102.20, PWM:41
23:09:45.521 -> Water_Temp.:102.20, PWM:41
23:09:46.317 -> Water_Temp.:102.20, PWM:41
23:09:47.109 -> Water_Temp.:102.20, PWM:41
23:09:48.628 -> Water_Temp.:102.20, PWM:41
23:09:49.423 -> Water_Temp.:102.20, PWM:41
23:09:50.183 -> Water_Temp.:102.20, PWM:41
23:09:50.978 -> Water_Temp.:102.20, PWM:41
23:09:51.771 -> Water_Temp.:102.20, PWM:41
23:09:52.530 -> Water_Temp.:102.20, PWM:41
23:09:53.326 -> Water_Temp.:102.20, PWM:41
23:09:54.088 -> Water_Temp.:102.20, PWM:41
23:09:54.884 -> Water_Temp.:102.20, PWM:41
23:09:56.436 -> Water_Temp.:102.31, PWM:41
23:09:57.231 -> Water_Temp.:102.31, PWM:41
23:09:57.993 -> Water_Temp.:102.31, PWM:41
23:09:58.786 -> Water_Temp.:102.31, PWM:41
23:10:00.308 -> Water_Temp.:102.31, PWM:41
23:10:01.108 -> Water_Temp.:102.31, PWM:41
23:10:01.870 -> Water_Temp.:102.31, PWM:41
23:10:02.662 -> Water_Temp.:102.31, PWM:41
23:10:03.460 -> Water_Temp.:102.31, PWM:41
23:10:04.217 -> Water_Temp.:102.31, PWM:41
23:10:05.011 -> Water_Temp.:102.31, PWM:41
23:10:06.560 -> Water_Temp.:102.31, PWM:41
23:10:07.320 -> Water_Temp.:102.31, PWM:41
23:10:08.111 -> Water_Temp.:102.31, PWM:41
23:10:08.906 -> Water_Temp.:102.31, PWM:41
23:10:09.666 -> Water_Temp.:102.31, PWM:41
23:10:10.460 -> Water_Temp.:102.31, PWM:41
23:10:11.219 -> Water_Temp.:102.31, PWM:41
23:10:12.014 -> Water_Temp.:102.31, PWM:41
23:10:13.568 -> Water_Temp.:102.31, PWM:41
23:10:14.324 -> Water_Temp.:102.31, PWM:41
23:10:15.119 -> Water_Temp.:102.31, PWM:41
23:10:15.911 -> Water_Temp.:102.31, PWM:41
23:10:16.674 -> Water_Temp.:102.31, PWM:41
23:10:17.470 -> Water_Temp.:102.31, PWM:41
23:10:18.229 -> Water_Temp.:102.31, PWM:41
23:10:19.022 -> Water_Temp.:102.31, PWM:41
23:10:19.785 -> Water_Temp.:102.31, PWM:41
23:10:20.582 -> Water_Temp.:102.31, PWM:41
23:10:21.345 -> Water_Temp.:102.31, PWM:41
23:10:22.137 -> Water_Temp.:102.31, PWM:41
23:10:22.932 -> Water_Temp.:102.31, PWM:41
23:10:23.692 -> Water_Temp.:102.31, PWM:41
23:10:24.483 -> Water_Temp.:102.31, PWM:41
23:10:25.245 -> Water_Temp.:102.31, PWM:41
23:10:26.040 -> Water_Temp.:102.31, PWM:41
23:10:26.800 -> Water_Temp.:102.31, PWM:41
23:10:27.589 -> Water_Temp.:102.31, PWM:41
23:10:28.383 -> Water_Temp.:102.31, PWM:41
23:10:29.143 -> Water_Temp.:102.31, PWM:41
23:10:29.939 -> Water_Temp.:102.42, PWM:41
23:10:30.734 -> Water_Temp.:102.31, PWM:41
23:10:31.495 -> Water_Temp.:102.31, PWM:41
23:10:32.292 -> Water_Temp.:102.31, PWM:41
23:10:33.050 -> Water_Temp.:102.31, PWM:41
23:10:33.844 -> Water_Temp.:102.31, PWM:41
23:10:34.603 -> Water_Temp.:102.31, PWM:41
23:10:35.397 -> Water_Temp.:102.31, PWM:41
23:10:36.950 -> Water_Temp.:102.31, PWM:41
23:10:37.738 -> Water_Temp.:102.31, PWM:41
23:10:38.499 -> Water_Temp.:102.31, PWM:41
23:10:39.293 -> Water_Temp.:102.31, PWM:41
23:10:40.050 -> Water_Temp.:102.42, PWM:41
23:10:40.845 -> Water_Temp.:102.42, PWM:41
23:10:41.607 -> Water_Temp.:102.31, PWM:41
23:10:43.162 -> Water_Temp.:102.42, PWM:41
23:10:43.923 -> Water_Temp.:102.42, PWM:41
23:10:44.720 -> Water_Temp.:102.42, PWM:41
23:10:45.483 -> Water_Temp.:102.42, PWM:41
23:10:46.278 -> Water_Temp.:102.42, PWM:41
23:10:47.037 -> Water_Temp.:102.42, PWM:41
23:10:47.828 -> Water_Temp.:102.42, PWM:41
23:10:48.627 -> Water_Temp.:102.42, PWM:41
23:10:49.388 -> Water_Temp.:102.42, PWM:41
23:10:50.182 -> Water_Temp.:102.42, PWM:41
23:10:50.943 -> Water_Temp.:102.42, PWM:41
23:10:51.737 -> Water_Temp.:102.42, PWM:41
23:10:52.497 -> Water_Temp.:102.42, PWM:41
23:10:53.295 -> Water_Temp.:102.42, PWM:41
23:10:54.087 -> Water_Temp.:102.42, PWM:41
23:10:54.848 -> Water_Temp.:102.42, PWM:41
23:10:55.639 -> Water_Temp.:102.42, PWM:41
23:10:56.400 -> Water_Temp.:102.42, PWM:41
23:10:57.193 -> Water_Temp.:102.42, PWM:41
23:10:57.990 -> Water_Temp.:102.42, PWM:41
23:10:58.749 -> Water_Temp.:102.42, PWM:41
23:10:59.547 -> Water_Temp.:102.42, PWM:41
23:11:00.311 -> Water_Temp.:102.42, PWM:41
23:11:01.109 -> Water_Temp.:102.42, PWM:41
23:11:01.870 -> Water_Temp.:102.42, PWM:41
23:11:02.667 -> Water_Temp.:102.42, PWM:41
23:11:03.428 -> Water_Temp.:102.42, PWM:41
23:11:04.223 -> Water_Temp.:102.42, PWM:41
23:11:04.985 -> Water_Temp.:102.42, PWM:41
23:11:05.780 -> Water_Temp.:102.42, PWM:41
23:11:06.539 -> Water_Temp.:102.42, PWM:41
23:11:08.093 -> Water_Temp.:102.42, PWM:41
23:11:08.855 -> Water_Temp.:102.42, PWM:41
23:11:09.650 -> Water_Temp.:102.42, PWM:41
23:11:10.443 -> Water_Temp.:102.54, PWM:41
23:11:11.205 -> Water_Temp.:102.42, PWM:41
23:11:12.005 -> Water_Temp.:102.54, PWM:41
23:11:12.767 -> Water_Temp.:102.54, PWM:41
23:11:14.323 -> Water_Temp.:102.54, PWM:41
23:11:15.122 -> Water_Temp.:102.54, PWM:41
23:11:15.881 -> Water_Temp.:102.54, PWM:41
23:11:16.677 -> Water_Temp.:102.54, PWM:41
23:11:17.436 -> Water_Temp.:102.54, PWM:41
23:11:18.230 -> Water_Temp.:102.54, PWM:41
23:11:19.022 -> Water_Temp.:102.54, PWM:41
23:11:19.776 -> Water_Temp.:102.54, PWM:41
23:11:20.575 -> Water_Temp.:102.54, PWM:41
23:11:21.367 -> Water_Temp.:102.54, PWM:41
23:11:22.126 -> Water_Temp.:102.54, PWM:41
23:11:22.921 -> Water_Temp.:102.54, PWM:41
23:11:23.680 -> Water_Temp.:102.54, PWM:41
23:11:24.476 -> Water_Temp.:102.54, PWM:41
23:11:25.244 -> Water_Temp.:102.54, PWM:41
23:11:26.041 -> Water_Temp.:102.54, PWM:41
23:11:26.799 -> Water_Temp.:102.54, PWM:41
23:11:27.591 -> Water_Temp.:102.54, PWM:41
23:11:28.383 -> Water_Temp.:102.54, PWM:41
23:11:29.144 -> Water_Temp.:102.54, PWM:41
23:11:29.936 -> Water_Temp.:102.54, PWM:41
23:11:30.696 -> Water_Temp.:102.54, PWM:41
23:11:31.492 -> Water_Temp.:102.54, PWM:41
23:11:32.254 -> Water_Temp.:102.54, PWM:41
23:11:33.051 -> Water_Temp.:102.54, PWM:41
23:11:33.845 -> Water_Temp.:102.54, PWM:41
23:11:34.604 -> Water_Temp.:102.54, PWM:41
23:11:35.398 -> Water_Temp.:102.54, PWM:41
23:11:36.157 -> Water_Temp.:102.54, PWM:41
23:11:38.477 -> Water_Temp.:102.54, PWM:41
23:11:40.035 -> Water_Temp.:102.54, PWM:41
23:11:40.830 -> Water_Temp.:102.54, PWM:41
23:11:41.589 -> Water_Temp.:102.54, PWM:41
23:11:42.383 -> Water_Temp.:102.54, PWM:41
23:11:43.145 -> Water_Temp.:102.65, PWM:41
23:11:43.938 -> Water_Temp.:104.22, PWM:41
23:11:44.733 -> Water_Temp.:106.25, PWM:41
23:11:45.488 -> Water_Temp.:103.89, PWM:41
23:11:46.289 -> Water_Temp.:103.44, PWM:41
23:11:47.807 -> Water_Temp.:103.10, PWM:41
23:11:48.603 -> Water_Temp.:103.10, PWM:41
23:11:49.364 -> Water_Temp.:102.99, PWM:41
23:11:50.157 -> Water_Temp.:102.87, PWM:41
23:11:50.918 -> Water_Temp.:102.87, PWM:41
23:11:51.709 -> Water_Temp.:102.87, PWM:41
23:11:52.501 -> Water_Temp.:102.87, PWM:41
23:11:53.261 -> Water_Temp.:102.76, PWM:41
23:11:54.058 -> Water_Temp.:102.76, PWM:41
23:11:54.820 -> Water_Temp.:102.76, PWM:41
23:11:55.616 -> Water_Temp.:102.76, PWM:41
23:11:56.409 -> Water_Temp.:102.76, PWM:41
23:11:57.928 -> Water_Temp.:102.76, PWM:41
23:11:58.723 -> Water_Temp.:102.76, PWM:41
23:11:59.483 -> Water_Temp.:102.76, PWM:41
23:12:00.277 -> Water_Temp.:102.76, PWM:41
23:12:01.077 -> Water_Temp.:102.76, PWM:41
23:12:01.837 -> Water_Temp.:102.76, PWM:41
23:12:02.632 -> Water_Temp.:102.65, PWM:41
23:12:03.392 -> Water_Temp.:102.65, PWM:41
23:12:04.186 -> Water_Temp.:102.65, PWM:41
23:12:04.947 -> Water_Temp.:102.65, PWM:41
23:12:05.746 -> Water_Temp.:102.65, PWM:41
23:12:06.503 -> Water_Temp.:102.65, PWM:41
23:12:07.297 -> Water_Temp.:102.65, PWM:41
23:12:08.092 -> Water_Temp.:102.65, PWM:41
23:12:08.852 -> Water_Temp.:102.65, PWM:41
23:12:09.649 -> Water_Temp.:102.65, PWM:41
23:12:10.410 -> Water_Temp.:102.65, PWM:41
23:12:11.205 -> Water_Temp.:102.65, PWM:41
23:12:11.965 -> Water_Temp.:102.65, PWM:41
23:12:12.760 -> Water_Temp.:102.76, PWM:41
23:12:13.522 -> Water_Temp.:102.65, PWM:41
23:12:14.318 -> Water_Temp.:102.76, PWM:41
23:12:15.111 -> Water_Temp.:102.76, PWM:41
23:12:15.875 -> Water_Temp.:102.76, PWM:41
23:12:16.667 -> Water_Temp.:102.76, PWM:41
23:12:17.429 -> Water_Temp.:102.76, PWM:41
23:12:18.223 -> Water_Temp.:102.76, PWM:41
23:12:18.984 -> Water_Temp.:102.76, PWM:41
23:12:19.779 -> Water_Temp.:102.76, PWM:41
23:12:20.576 -> Water_Temp.:102.76, PWM:42
23:12:21.341 -> Water_Temp.:102.76, PWM:46
23:12:22.133 -> Water_Temp.:102.76, PWM:46
23:12:22.889 -> Water_Temp.:102.65, PWM:46
23:12:23.682 -> Water_Temp.:102.76, PWM:43
23:12:24.475 -> Water_Temp.:102.65, PWM:43
23:12:25.234 -> Water_Temp.:102.65, PWM:43
23:12:26.025 -> Water_Temp.:102.65, PWM:43
23:12:26.786 -> Water_Temp.:102.65, PWM:43
23:12:28.335 -> Water_Temp.:102.65, PWM:43
23:12:29.126 -> Water_Temp.:102.65, PWM:43
23:12:29.881 -> Water_Temp.:102.65, PWM:43
23:12:30.677 -> Water_Temp.:102.65, PWM:43
23:12:31.439 -> Water_Temp.:102.65, PWM:43
23:12:32.236 -> Water_Temp.:102.65, PWM:43
23:12:32.994 -> Water_Temp.:102.65, PWM:43
23:12:33.788 -> Water_Temp.:102.65, PWM:43
23:12:34.583 -> Water_Temp.:102.65, PWM:43
23:12:35.346 -> Water_Temp.:102.65, PWM:43
23:12:36.139 -> Water_Temp.:102.65, PWM:43
23:12:36.930 -> Water_Temp.:102.65, PWM:43
23:12:37.684 -> Water_Temp.:102.65, PWM:43
23:12:38.471 -> Water_Temp.:102.65, PWM:43
23:12:39.265 -> Water_Temp.:102.65, PWM:43
23:12:40.027 -> Water_Temp.:102.65, PWM:43
23:12:42.342 -> Water_Temp.:102.65, PWM:43
23:12:43.135 -> Water_Temp.:102.65, PWM:43
23:12:43.892 -> Water_Temp.:102.65, PWM:43
23:12:44.684 -> Water_Temp.:102.65, PWM:43
23:12:45.446 -> Water_Temp.:102.65, PWM:43
23:12:46.241 -> Water_Temp.:102.65, PWM:43
23:12:47.032 -> Water_Temp.:102.65, PWM:43
23:12:48.552 -> Water_Temp.:102.65, PWM:43
23:12:49.345 -> Water_Temp.:102.65, PWM:43
23:12:50.139 -> Water_Temp.:102.65, PWM:43
23:12:50.896 -> Water_Temp.:102.65, PWM:43
23:12:51.688 -> Water_Temp.:102.65, PWM:43
23:12:52.482 -> Water_Temp.:102.65, PWM:43
23:12:53.243 -> Water_Temp.:102.65, PWM:43
23:12:54.035 -> Water_Temp.:102.65, PWM:43
23:12:54.799 -> Water_Temp.:102.65, PWM:43
23:12:55.592 -> Water_Temp.:102.65, PWM:43
23:12:56.352 -> Water_Temp.:102.65, PWM:43
23:12:57.150 -> Water_Temp.:102.65, PWM:43
23:12:57.942 -> Water_Temp.:102.65, PWM:43
23:12:58.705 -> Water_Temp.:102.65, PWM:43
23:12:59.467 -> Water_Temp.:102.65, PWM:43
23:13:00.263 -> Water_Temp.:102.65, PWM:43
23:13:01.025 -> Water_Temp.:102.65, PWM:43
23:13:01.814 -> Water_Temp.:102.65, PWM:43
23:13:02.604 -> Water_Temp.:102.65, PWM:43
23:13:03.396 -> Water_Temp.:102.65, PWM:43
23:13:04.155 -> Water_Temp.:102.65, PWM:43
23:13:04.953 -> Water_Temp.:102.65, PWM:43
23:13:05.714 -> Water_Temp.:102.65, PWM:43
23:13:06.508 -> Water_Temp.:102.65, PWM:43
23:13:07.300 -> Water_Temp.:102.65, PWM:43
23:13:08.062 -> Water_Temp.:102.65, PWM:43
23:13:09.617 -> Water_Temp.:102.65, PWM:43
23:13:10.374 -> Water_Temp.:102.65, PWM:43
23:13:11.170 -> Water_Temp.:102.65, PWM:43
23:13:11.929 -> Water_Temp.:102.65, PWM:43
23:13:12.725 -> Water_Temp.:102.65, PWM:43
23:13:13.488 -> Water_Temp.:102.65, PWM:43
23:13:14.283 -> Water_Temp.:102.65, PWM:43
23:13:15.078 -> Water_Temp.:102.65, PWM:43
23:13:15.841 -> Water_Temp.:102.65, PWM:43
23:13:16.635 -> Water_Temp.:102.65, PWM:43
23:13:17.396 -> Water_Temp.:102.65, PWM:43
23:13:18.191 -> Water_Temp.:102.65, PWM:43
23:13:18.953 -> Water_Temp.:102.65, PWM:43
23:13:19.746 -> Water_Temp.:102.65, PWM:43
23:13:20.537 -> Water_Temp.:102.65, PWM:43
23:13:21.299 -> Water_Temp.:102.65, PWM:43
23:13:22.089 -> Water_Temp.:102.65, PWM:43
23:13:22.848 -> Water_Temp.:102.65, PWM:43
23:13:23.644 -> Water_Temp.:102.65, PWM:43
23:13:24.442 -> Water_Temp.:102.65, PWM:43
23:13:25.205 -> Water_Temp.:102.76, PWM:43
23:13:25.994 -> Water_Temp.:102.65, PWM:43
23:13:26.752 -> Water_Temp.:102.76, PWM:43
23:13:27.542 -> Water_Temp.:102.76, PWM:43
23:13:28.338 -> Water_Temp.:102.76, PWM:43
23:13:31.389 -> Water_Temp.:102.76, PWM:43
23:13:32.180 -> Water_Temp.:102.76, PWM:43
23:13:33.733 -> Water_Temp.:102.76, PWM:43
23:13:34.525 -> Water_Temp.:102.65, PWM:43
23:13:35.290 -> Water_Temp.:102.65, PWM:43
23:13:36.848 -> Water_Temp.:102.76, PWM:43
23:13:37.609 -> Water_Temp.:102.76, PWM:43
23:13:38.401 -> Water_Temp.:102.76, PWM:43
23:13:39.160 -> Water_Temp.:102.76, PWM:43
23:13:39.951 -> Water_Temp.:102.76, PWM:43
23:13:40.716 -> Water_Temp.:102.76, PWM:43
23:13:41.510 -> Water_Temp.:102.76, PWM:43
23:13:42.303 -> Water_Temp.:102.76, PWM:43
23:13:43.067 -> Water_Temp.:102.76, PWM:43
23:13:43.858 -> Water_Temp.:102.76, PWM:43
23:13:44.618 -> Water_Temp.:102.76, PWM:43
23:13:45.413 -> Water_Temp.:102.76, PWM:43
23:13:46.175 -> Water_Temp.:102.76, PWM:43
23:13:46.972 -> Water_Temp.:102.76, PWM:43
23:13:47.766 -> Water_Temp.:102.76, PWM:43
23:13:49.322 -> Water_Temp.:102.76, PWM:43
23:13:50.084 -> Water_Temp.:102.76, PWM:43
23:13:50.880 -> Water_Temp.:102.76, PWM:43
23:13:51.639 -> Water_Temp.:102.76, PWM:43
23:13:52.433 -> Water_Temp.:102.76, PWM:43
23:13:53.194 -> Water_Temp.:102.76, PWM:43
23:13:53.984 -> Water_Temp.:102.76, PWM:43
23:13:54.782 -> Water_Temp.:102.76, PWM:43
23:13:55.546 -> Water_Temp.:102.76, PWM:43
23:13:56.339 -> Water_Temp.:102.76, PWM:43
23:13:57.095 -> Water_Temp.:102.87, PWM:43
23:13:57.894 -> Water_Temp.:102.76, PWM:43
23:13:58.685 -> Water_Temp.:102.76, PWM:43
23:13:59.447 -> Water_Temp.:102.87, PWM:43
23:14:00.243 -> Water_Temp.:102.87, PWM:43
23:14:01.005 -> Water_Temp.:102.87, PWM:43
23:14:01.801 -> Water_Temp.:102.87, PWM:43
23:14:02.562 -> Water_Temp.:102.87, PWM:43
23:14:03.351 -> Water_Temp.:102.87, PWM:43
23:14:04.144 -> Water_Temp.:102.87, PWM:43
23:14:04.904 -> Water_Temp.:102.87, PWM:43
23:14:05.699 -> Water_Temp.:102.87, PWM:43
23:14:07.258 -> Water_Temp.:106.47, PWM:43
23:14:08.018 -> Water_Temp.:104.11, PWM:43
23:14:08.812 -> Water_Temp.:103.77, PWM:43
23:14:09.606 -> Water_Temp.:103.55, PWM:43
23:14:10.366 -> Water_Temp.:103.44, PWM:43
23:14:11.164 -> Water_Temp.:103.32, PWM:43
23:14:11.923 -> Water_Temp.:103.21, PWM:43
23:14:12.715 -> Water_Temp.:103.21, PWM:43
23:14:13.505 -> Water_Temp.:103.10, PWM:43
23:14:14.263 -> Water_Temp.:103.10, PWM:43
23:14:15.064 -> Water_Temp.:103.10, PWM:43
23:14:15.825 -> Water_Temp.:103.10, PWM:43
23:14:17.373 -> Water_Temp.:102.99, PWM:43
23:14:18.170 -> Water_Temp.:102.99, PWM:43
23:14:18.929 -> Water_Temp.:102.99, PWM:43
23:14:20.480 -> Water_Temp.:102.99, PWM:43
23:14:21.277 -> Water_Temp.:102.99, PWM:43
23:14:22.038 -> Water_Temp.:102.99, PWM:43
23:14:22.832 -> Water_Temp.:102.99, PWM:43
23:14:23.626 -> Water_Temp.:102.99, PWM:43
23:14:24.387 -> Water_Temp.:102.99, PWM:43
23:14:25.178 -> Water_Temp.:102.99, PWM:43
23:14:25.936 -> Water_Temp.:102.99, PWM:43
23:14:26.729 -> Water_Temp.:102.99, PWM:43
23:14:27.521 -> Water_Temp.:102.99, PWM:43
23:14:28.277 -> Water_Temp.:102.99, PWM:43
23:14:29.071 -> Water_Temp.:102.99, PWM:42
23:14:29.866 -> Water_Temp.:102.99, PWM:42
23:14:30.628 -> Water_Temp.:102.99, PWM:43
23:14:31.422 -> Water_Temp.:102.99, PWM:43
23:14:32.182 -> Water_Temp.:102.99, PWM:43
23:14:32.975 -> Water_Temp.:102.99, PWM:43
23:14:33.734 -> Water_Temp.:102.99, PWM:43
23:14:34.525 -> Water_Temp.:102.99, PWM:43
23:14:35.324 -> Water_Temp.:102.99, PWM:43
23:14:36.084 -> Water_Temp.:102.99, PWM:43
23:14:36.876 -> Water_Temp.:102.99, PWM:43
23:14:37.668 -> Water_Temp.:102.99, PWM:43
23:14:38.424 -> Water_Temp.:102.87, PWM:43
23:14:39.218 -> Water_Temp.:102.99, PWM:43
23:14:39.976 -> Water_Temp.:102.87, PWM:43
23:14:40.772 -> Water_Temp.:102.87, PWM:43
23:14:41.561 -> Water_Temp.:102.87, PWM:43
23:14:42.323 -> Water_Temp.:102.87, PWM:43
23:14:43.116 -> Water_Temp.:102.87, PWM:43
23:14:43.908 -> Water_Temp.:102.87, PWM:43
23:14:44.666 -> Water_Temp.:102.87, PWM:43
23:14:45.468 -> Water_Temp.:102.87, PWM:43
23:14:46.229 -> Water_Temp.:102.87, PWM:43
23:14:47.025 -> Water_Temp.:102.87, PWM:43
23:14:47.781 -> Water_Temp.:102.87, PWM:43
23:14:48.576 -> Water_Temp.:102.87, PWM:43
23:14:49.336 -> Water_Temp.:102.87, PWM:43
23:14:50.128 -> Water_Temp.:102.87, PWM:43
23:14:50.923 -> Water_Temp.:102.87, PWM:43
23:14:51.684 -> Water_Temp.:102.87, PWM:43
23:14:52.478 -> Water_Temp.:102.87, PWM:43
23:14:53.237 -> Water_Temp.:102.87, PWM:43
23:14:54.029 -> Water_Temp.:102.87, PWM:43
23:14:54.826 -> Water_Temp.:102.87, PWM:43
23:14:55.587 -> Water_Temp.:102.99, PWM:43
23:14:56.382 -> Water_Temp.:102.99, PWM:43
23:14:57.143 -> Water_Temp.:102.99, PWM:43
23:14:57.936 -> Water_Temp.:102.99, PWM:43
23:14:58.698 -> Water_Temp.:102.99, PWM:43
23:14:59.493 -> Water_Temp.:102.99, PWM:43
23:15:01.815 -> Water_Temp.:102.99, PWM:43
23:15:02.576 -> Water_Temp.:102.99, PWM:43
23:15:03.368 -> Water_Temp.:102.99, PWM:43
23:15:04.128 -> Water_Temp.:102.99, PWM:43
23:15:05.688 -> Water_Temp.:102.99, PWM:43
23:15:06.481 -> Water_Temp.:102.99, PWM:43
23:15:07.242 -> Water_Temp.:102.99, PWM:43
23:15:08.037 -> Water_Temp.:102.99, PWM:43
23:15:08.834 -> Water_Temp.:102.99, PWM:43
23:15:09.594 -> Water_Temp.:102.99, PWM:43
23:15:10.386 -> Water_Temp.:102.99, PWM:43
23:15:11.150 -> Water_Temp.:102.99, PWM:43
23:15:11.944 -> Water_Temp.:102.99, PWM:43
23:15:12.735 -> Water_Temp.:102.99, PWM:43
23:15:13.496 -> Water_Temp.:102.99, PWM:43
23:15:15.052 -> Water_Temp.:102.99, PWM:43
23:15:15.813 -> Water_Temp.:102.99, PWM:43
23:15:16.610 -> Water_Temp.:102.99, PWM:43
23:15:17.371 -> Water_Temp.:102.99, PWM:43
23:15:18.165 -> Water_Temp.:102.99, PWM:43
23:15:18.926 -> Water_Temp.:102.99, PWM:43
23:15:19.716 -> Water_Temp.:102.99, PWM:43
23:15:20.513 -> Water_Temp.:102.99, PWM:43
23:15:21.278 -> Water_Temp.:102.99, PWM:43
23:15:22.831 -> Water_Temp.:102.99, PWM:43
23:15:23.591 -> Water_Temp.:102.99, PWM:43
23:15:24.389 -> Water_Temp.:102.99, PWM:43
23:15:25.153 -> Water_Temp.:102.99, PWM:43
23:15:25.943 -> Water_Temp.:102.99, PWM:43
23:15:26.739 -> Water_Temp.:102.99, PWM:43
23:15:29.053 -> Water_Temp.:102.99, PWM:43
23:15:29.814 -> Water_Temp.:102.99, PWM:43
23:15:30.611 -> Water_Temp.:102.99, PWM:43
23:15:31.376 -> Water_Temp.:102.99, PWM:43
23:15:32.171 -> Water_Temp.:102.99, PWM:43
23:15:32.934 -> Water_Temp.:102.99, PWM:43
23:15:33.729 -> Water_Temp.:102.99, PWM:43
23:15:34.491 -> Water_Temp.:102.99, PWM:43
23:15:35.284 -> Water_Temp.:102.99, PWM:43
23:15:36.076 -> Water_Temp.:102.99, PWM:43
23:15:36.834 -> Water_Temp.:102.99, PWM:43
23:15:38.387 -> Water_Temp.:103.10, PWM:43
23:15:39.180 -> Water_Temp.:103.10, PWM:43
23:15:39.940 -> Water_Temp.:103.10, PWM:43
23:15:40.735 -> Water_Temp.:103.10, PWM:43
23:15:41.493 -> Water_Temp.:103.10, PWM:43
23:15:42.289 -> Water_Temp.:103.10, PWM:43
23:15:43.048 -> Water_Temp.:103.10, PWM:43
23:15:44.602 -> Water_Temp.:103.10, PWM:43
23:15:45.393 -> Water_Temp.:103.10, PWM:43
23:15:46.155 -> Water_Temp.:103.10, PWM:43
23:15:46.946 -> Water_Temp.:103.10, PWM:43
23:15:48.500 -> Water_Temp.:103.10, PWM:43
23:15:49.291 -> Water_Temp.:103.10, PWM:43
23:15:50.054 -> Water_Temp.:103.10, PWM:43
23:15:50.851 -> Water_Temp.:103.10, PWM:43
23:15:51.608 -> Water_Temp.:103.10, PWM:43
23:15:52.404 -> Water_Temp.:103.10, PWM:43
23:15:53.164 -> Water_Temp.:103.10, PWM:43
23:15:53.958 -> Water_Temp.:103.10, PWM:43
23:15:54.720 -> Water_Temp.:103.10, PWM:43
23:15:55.513 -> Water_Temp.:103.10, PWM:43
23:15:56.303 -> Water_Temp.:103.10, PWM:43
23:15:57.092 -> Water_Temp.:103.10, PWM:43
23:15:57.854 -> Water_Temp.:103.10, PWM:43
23:15:58.648 -> Water_Temp.:103.10, PWM:43
23:15:59.405 -> Water_Temp.:103.10, PWM:43
23:16:00.201 -> Water_Temp.:103.10, PWM:43
23:16:00.994 -> Water_Temp.:103.10, PWM:43
23:16:01.754 -> Water_Temp.:103.10, PWM:43
23:16:02.547 -> Water_Temp.:103.10, PWM:43
23:16:03.305 -> Water_Temp.:103.10, PWM:43
23:16:04.102 -> Water_Temp.:103.10, PWM:43
23:16:04.863 -> Water_Temp.:103.10, PWM:43
23:16:05.658 -> Water_Temp.:103.10, PWM:43
23:16:06.454 -> Water_Temp.:103.10, PWM:43
23:16:07.211 -> Water_Temp.:103.10, PWM:43
23:16:08.004 -> Water_Temp.:103.10, PWM:43
23:16:08.763 -> Water_Temp.:103.10, PWM:43
23:16:09.554 -> Water_Temp.:103.10, PWM:43
23:16:10.351 -> Water_Temp.:103.10, PWM:43
23:16:11.115 -> Water_Temp.:103.10, PWM:43
23:16:11.907 -> Water_Temp.:103.10, PWM:43
23:16:13.456 -> Water_Temp.:103.10, PWM:43
23:16:14.249 -> Water_Temp.:103.10, PWM:43
23:16:15.010 -> Water_Temp.:103.10, PWM:43
23:16:15.806 -> Water_Temp.:103.10, PWM:43
23:16:16.563 -> Water_Temp.:103.10, PWM:43
23:16:17.359 -> Water_Temp.:103.10, PWM:43
23:16:18.119 -> Water_Temp.:103.10, PWM:43
23:16:18.911 -> Water_Temp.:103.10, PWM:43
23:16:19.702 -> Water_Temp.:103.10, PWM:43
23:16:20.493 -> Water_Temp.:103.10, PWM:43
23:16:22.011 -> Water_Temp.:103.21, PWM:43
23:16:22.806 -> Water_Temp.:103.21, PWM:43
23:16:23.563 -> Water_Temp.:103.21, PWM:43
23:16:24.361 -> Water_Temp.:103.21, PWM:43
23:16:25.123 -> Water_Temp.:103.21, PWM:43
23:16:25.912 -> Water_Temp.:103.21, PWM:43
23:16:26.701 -> Water_Temp.:103.21, PWM:43
23:16:27.460 -> Water_Temp.:103.21, PWM:43
23:16:28.254 -> Water_Temp.:103.21, PWM:43
23:16:29.047 -> Water_Temp.:103.21, PWM:43
23:16:29.809 -> Water_Temp.:103.21, PWM:43
23:16:30.604 -> Water_Temp.:103.21, PWM:43
23:16:31.366 -> Water_Temp.:103.21, PWM:43
23:16:32.159 -> Water_Temp.:103.21, PWM:43
23:16:32.917 -> Water_Temp.:103.21, PWM:43
23:16:33.708 -> Water_Temp.:103.21, PWM:43
23:16:34.503 -> Water_Temp.:103.21, PWM:43
23:16:35.267 -> Water_Temp.:103.21, PWM:43
23:16:36.058 -> Water_Temp.:103.21, PWM:43
23:16:36.853 -> Water_Temp.:103.21, PWM:43
23:16:37.611 -> Water_Temp.:103.21, PWM:43
23:16:38.404 -> Water_Temp.:103.21, PWM:43
23:16:39.160 -> Water_Temp.:103.21, PWM:43
23:16:39.951 -> Water_Temp.:103.21, PWM:43
23:16:40.748 -> Water_Temp.:103.21, PWM:43
23:16:41.508 -> Water_Temp.:103.32, PWM:43
23:16:42.300 -> Water_Temp.:103.21, PWM:43
23:16:43.059 -> Water_Temp.:103.32, PWM:43
23:16:43.853 -> Water_Temp.:103.32, PWM:43
23:16:44.651 -> Water_Temp.:103.32, PWM:43
23:16:45.415 -> Water_Temp.:103.32, PWM:43
23:16:46.180 -> Water_Temp.:103.32, PWM:43
23:16:46.973 -> Water_Temp.:103.21, PWM:43
23:16:47.770 -> Water_Temp.:103.32, PWM:43
23:16:48.530 -> Water_Temp.:103.32, PWM:43
23:16:50.086 -> Water_Temp.:103.32, PWM:43
23:16:51.640 -> Water_Temp.:103.32, PWM:43
23:16:52.432 -> Water_Temp.:103.32, PWM:43
23:16:53.227 -> Water_Temp.:103.32, PWM:43
23:16:53.988 -> Water_Temp.:103.32, PWM:43
23:16:54.784 -> Water_Temp.:103.32, PWM:43
23:16:55.549 -> Water_Temp.:103.32, PWM:43
23:16:56.346 -> Water_Temp.:103.32, PWM:43
23:16:57.103 -> Water_Temp.:103.32, PWM:43
23:16:57.899 -> Water_Temp.:103.32, PWM:43
23:16:58.659 -> Water_Temp.:103.32, PWM:43
23:16:59.451 -> Water_Temp.:103.32, PWM:43
23:17:00.216 -> Water_Temp.:103.32, PWM:43
23:17:01.012 -> Water_Temp.:103.32, PWM:43
23:17:02.533 -> Water_Temp.:103.32, PWM:43
23:17:03.331 -> Water_Temp.:103.32, PWM:43
23:17:04.089 -> Water_Temp.:103.32, PWM:43
23:17:04.886 -> Water_Temp.:103.32, PWM:43
23:17:06.444 -> Water_Temp.:103.32, PWM:43
23:17:07.237 -> Water_Temp.:103.32, PWM:43
23:17:07.999 -> Water_Temp.:103.32, PWM:43
23:17:08.786 -> Water_Temp.:103.32, PWM:43
23:17:09.578 -> Water_Temp.:103.32, PWM:43
23:17:10.339 -> Water_Temp.:103.32, PWM:43
23:17:11.133 -> Water_Temp.:103.32, PWM:43
23:17:11.891 -> Water_Temp.:103.32, PWM:43
23:17:12.686 -> Water_Temp.:103.44, PWM:43
23:17:13.480 -> Water_Temp.:103.32, PWM:43
23:17:14.241 -> Water_Temp.:103.44, PWM:43
23:17:15.033 -> Water_Temp.:103.44, PWM:43
23:17:16.558 -> Water_Temp.:103.44, PWM:43
23:17:17.354 -> Water_Temp.:103.44, PWM:43
23:17:18.115 -> Water_Temp.:103.44, PWM:43
23:17:18.915 -> Water_Temp.:103.44, PWM:43
23:17:19.677 -> Water_Temp.:103.44, PWM:43
23:17:20.473 -> Water_Temp.:103.44, PWM:43
23:17:21.239 -> Water_Temp.:103.44, PWM:43
23:17:22.031 -> Water_Temp.:103.44, PWM:43
23:17:24.346 -> Water_Temp.:103.44, PWM:43
23:17:25.141 -> Water_Temp.:103.44, PWM:43
23:17:25.900 -> Water_Temp.:103.44, PWM:43
23:17:27.459 -> Water_Temp.:103.44, PWM:43
23:17:28.250 -> Water_Temp.:103.44, PWM:43
23:17:29.010 -> Water_Temp.:103.44, PWM:43
23:17:29.801 -> Water_Temp.:103.44, PWM:43
23:17:30.592 -> Water_Temp.:103.44, PWM:43
23:17:31.351 -> Water_Temp.:103.44, PWM:43
23:17:32.147 -> Water_Temp.:103.44, PWM:43
23:17:32.904 -> Water_Temp.:103.44, PWM:43
23:17:33.699 -> Water_Temp.:103.44, PWM:43
23:17:34.492 -> Water_Temp.:103.44, PWM:43
23:17:35.255 -> Water_Temp.:103.44, PWM:43
23:17:36.051 -> Water_Temp.:103.44, PWM:43
23:17:36.811 -> Water_Temp.:103.44, PWM:43
23:17:37.602 -> Water_Temp.:103.44, PWM:43
23:17:38.395 -> Water_Temp.:103.44, PWM:43
23:17:39.155 -> Water_Temp.:103.44, PWM:43
23:17:39.948 -> Water_Temp.:103.44, PWM:43
23:17:40.711 -> Water_Temp.:103.44, PWM:43
23:17:41.507 -> Water_Temp.:103.44, PWM:43
23:17:42.267 -> Water_Temp.:103.44, PWM:43
23:17:43.062 -> Water_Temp.:103.44, PWM:43
23:17:43.823 -> Water_Temp.:103.44, PWM:43
23:17:44.619 -> Water_Temp.:103.44, PWM:43
23:17:45.414 -> Water_Temp.:103.44, PWM:43
23:17:46.174 -> Water_Temp.:103.44, PWM:43
23:17:46.965 -> Water_Temp.:103.44, PWM:43
23:17:48.522 -> Water_Temp.:103.44, PWM:43
23:17:49.289 -> Water_Temp.:103.44, PWM:43
23:17:50.841 -> Water_Temp.:103.44, PWM:43
23:17:51.639 -> Water_Temp.:103.44, PWM:43
23:17:52.399 -> Water_Temp.:103.44, PWM:43
23:17:53.195 -> Water_Temp.:103.44, PWM:43
23:17:54.724 -> Water_Temp.:103.55, PWM:43
23:17:55.492 -> Water_Temp.:103.44, PWM:43
23:17:56.258 -> Water_Temp.:103.44, PWM:43
23:17:57.051 -> Water_Temp.:103.44, PWM:43
23:17:57.845 -> Water_Temp.:103.44, PWM:43
23:17:58.608 -> Water_Temp.:103.44, PWM:43
23:17:59.403 -> Water_Temp.:103.44, PWM:43
23:18:00.158 -> Water_Temp.:103.44, PWM:43
23:18:00.952 -> Water_Temp.:103.44, PWM:43
23:18:01.745 -> Water_Temp.:103.44, PWM:43
23:18:02.504 -> Water_Temp.:103.44, PWM:43
23:18:04.060 -> Water_Temp.:103.44, PWM:43
23:18:05.618 -> Water_Temp.:103.44, PWM:43
23:18:06.380 -> Water_Temp.:103.44, PWM:43
23:18:07.173 -> Water_Temp.:103.44, PWM:43
23:18:07.968 -> Water_Temp.:103.44, PWM:43
23:18:08.728 -> Water_Temp.:103.44, PWM:43
23:18:09.522 -> Water_Temp.:103.44, PWM:43
23:18:10.280 -> Water_Temp.:103.44, PWM:43
23:18:11.076 -> Water_Temp.:103.44, PWM:43
23:18:11.839 -> Water_Temp.:103.44, PWM:43
23:18:12.632 -> Water_Temp.:103.44, PWM:43
23:18:13.428 -> Water_Temp.:103.44, PWM:43
23:18:14.189 -> Water_Temp.:103.44, PWM:43
23:18:14.989 -> Water_Temp.:103.44, PWM:43
23:18:15.746 -> Water_Temp.:103.44, PWM:43
23:18:16.543 -> Water_Temp.:103.44, PWM:43
23:18:17.302 -> Water_Temp.:103.44, PWM:43
23:18:18.099 -> Water_Temp.:103.44, PWM:43
23:18:18.890 -> Water_Temp.:103.44, PWM:43
23:18:19.651 -> Water_Temp.:103.55, PWM:43
23:18:20.444 -> Water_Temp.:103.44, PWM:43
23:18:21.206 -> Water_Temp.:103.55, PWM:43
23:18:22.005 -> Water_Temp.:103.44, PWM:43
23:18:22.767 -> Water_Temp.:103.44, PWM:43
23:18:23.564 -> Water_Temp.:103.55, PWM:43
23:18:24.319 -> Water_Temp.:103.55, PWM:43
23:18:25.111 -> Water_Temp.:103.55, PWM:43
23:18:25.906 -> Water_Temp.:103.44, PWM:43
23:18:26.662 -> Water_Temp.:103.44, PWM:47
23:18:27.461 -> Water_Temp.:103.44, PWM:47
23:18:29.021 -> Water_Temp.:103.44, PWM:47
23:18:30.575 -> Water_Temp.:103.55, PWM:46
23:18:31.338 -> Water_Temp.:103.44, PWM:46
23:18:32.130 -> Water_Temp.:103.44, PWM:46
23:18:32.924 -> Water_Temp.:103.44, PWM:47
23:18:33.683 -> Water_Temp.:103.44, PWM:47
23:18:34.475 -> Water_Temp.:103.44, PWM:47
23:18:35.270 -> Water_Temp.:103.44, PWM:47
23:18:36.032 -> Water_Temp.:103.44, PWM:47
23:18:36.823 -> Water_Temp.:103.44, PWM:47
23:18:37.586 -> Water_Temp.:103.44, PWM:47
23:18:38.381 -> Water_Temp.:103.44, PWM:46
23:18:39.139 -> Water_Temp.:103.44, PWM:46
23:18:39.932 -> Water_Temp.:103.44, PWM:47
23:18:41.490 -> Water_Temp.:103.44, PWM:47
23:18:42.247 -> Water_Temp.:103.44, PWM:47
23:18:43.042 -> Water_Temp.:103.44, PWM:46
23:18:43.806 -> Water_Temp.:103.44, PWM:47
23:18:44.604 -> Water_Temp.:103.44, PWM:47
23:18:45.368 -> Water_Temp.:103.44, PWM:47
23:18:46.164 -> Water_Temp.:103.44, PWM:47
23:18:47.684 -> Water_Temp.:103.44, PWM:47
23:18:48.482 -> Water_Temp.:103.44, PWM:47
23:18:49.273 -> Water_Temp.:103.44, PWM:47
23:18:50.037 -> Water_Temp.:103.44, PWM:47
23:18:50.801 -> Water_Temp.:103.44, PWM:47
23:18:51.598 -> Water_Temp.:103.44, PWM:47
23:18:52.395 -> Water_Temp.:103.44, PWM:46
23:18:53.153 -> Water_Temp.:103.44, PWM:47
23:18:53.946 -> Water_Temp.:103.44, PWM:47
23:18:54.735 -> Water_Temp.:103.44, PWM:47
23:18:55.493 -> Water_Temp.:103.44, PWM:47
23:18:56.292 -> Water_Temp.:103.55, PWM:47
23:18:57.055 -> Water_Temp.:103.55, PWM:46
23:18:57.852 -> Water_Temp.:103.55, PWM:47
23:18:58.606 -> Water_Temp.:103.55, PWM:47
23:18:59.395 -> Water_Temp.:103.55, PWM:47
23:19:00.186 -> Water_Temp.:103.55, PWM:47
23:19:00.950 -> Water_Temp.:103.55, PWM:47
23:19:01.746 -> Water_Temp.:103.55, PWM:46
23:19:02.506 -> Water_Temp.:103.55, PWM:47
23:19:03.301 -> Water_Temp.:103.55, PWM:47
23:19:04.094 -> Water_Temp.:103.55, PWM:46
23:19:04.856 -> Water_Temp.:103.55, PWM:47
23:19:05.651 -> Water_Temp.:103.55, PWM:47
23:19:06.409 -> Water_Temp.:103.55, PWM:47
23:19:07.199 -> Water_Temp.:103.55, PWM:47
23:19:07.995 -> Water_Temp.:103.55, PWM:46
23:19:08.754 -> Water_Temp.:103.55, PWM:47
23:19:09.547 -> Water_Temp.:103.55, PWM:47
23:19:10.302 -> Water_Temp.:103.55, PWM:47
23:19:11.097 -> Water_Temp.:103.55, PWM:47
23:19:11.890 -> Water_Temp.:103.55, PWM:47
23:19:12.648 -> Water_Temp.:103.55, PWM:47
23:19:14.209 -> Water_Temp.:103.55, PWM:47
23:19:14.973 -> Water_Temp.:103.55, PWM:47
23:19:15.772 -> Water_Temp.:103.55, PWM:47
23:19:16.529 -> Water_Temp.:103.55, PWM:47
23:19:17.325 -> Water_Temp.:103.55, PWM:47
23:19:18.085 -> Water_Temp.:103.55, PWM:47
23:19:18.879 -> Water_Temp.:103.55, PWM:47
23:19:20.437 -> Water_Temp.:103.55, PWM:47
23:19:21.198 -> Water_Temp.:103.55, PWM:47
23:19:21.991 -> Water_Temp.:103.55, PWM:47
23:19:23.543 -> Water_Temp.:103.55, PWM:47
23:19:24.305 -> Water_Temp.:103.66, PWM:47
23:19:25.103 -> Water_Temp.:103.55, PWM:47
23:19:26.655 -> Water_Temp.:103.55, PWM:47
23:19:27.450 -> Water_Temp.:103.55, PWM:47
23:19:28.208 -> Water_Temp.:103.55, PWM:47
23:19:29.793 -> Water_Temp.:103.55, PWM:47
23:19:30.551 -> Water_Temp.:103.55, PWM:47
23:19:31.349 -> Water_Temp.:103.55, PWM:47
23:19:32.902 -> Water_Temp.:103.55, PWM:47
23:19:33.664 -> Water_Temp.:103.55, PWM:47
23:19:34.458 -> Water_Temp.:103.55, PWM:46
23:19:35.219 -> Water_Temp.:103.55, PWM:47
23:19:36.013 -> Water_Temp.:103.55, PWM:47
23:19:36.802 -> Water_Temp.:103.55, PWM:47
23:19:37.568 -> Water_Temp.:103.55, PWM:47
23:19:39.128 -> Water_Temp.:103.55, PWM:47
23:19:39.886 -> Water_Temp.:103.66, PWM:47
23:19:40.652 -> Water_Temp.:103.66, PWM:47
23:19:41.452 -> Water_Temp.:103.55, PWM:46
23:19:42.212 -> Water_Temp.:103.66, PWM:47
23:19:43.008 -> Water_Temp.:103.66, PWM:47
23:19:43.801 -> Water_Temp.:103.66, PWM:47
23:19:44.564 -> Water_Temp.:103.66, PWM:47
23:19:45.359 -> Water_Temp.:103.66, PWM:47
23:19:46.876 -> Water_Temp.:103.66, PWM:47
23:19:47.667 -> Water_Temp.:103.66, PWM:47
23:19:48.461 -> Water_Temp.:103.66, PWM:47
23:19:49.223 -> Water_Temp.:103.66, PWM:47
23:19:50.018 -> Water_Temp.:103.66, PWM:47
23:19:50.780 -> Water_Temp.:103.66, PWM:46
23:19:51.578 -> Water_Temp.:103.66, PWM:47
23:19:52.334 -> Water_Temp.:103.66, PWM:47
23:19:53.127 -> Water_Temp.:103.66, PWM:47
23:19:53.922 -> Water_Temp.:103.66, PWM:47
23:19:54.686 -> Water_Temp.:103.66, PWM:47
23:19:55.447 -> Water_Temp.:103.66, PWM:47
23:19:56.245 -> Water_Temp.:103.66, PWM:47
23:19:57.006 -> Water_Temp.:103.66, PWM:47
23:19:57.801 -> Water_Temp.:103.66, PWM:46
23:19:58.598 -> Water_Temp.:103.66, PWM:47
23:19:59.355 -> Water_Temp.:103.66, PWM:47
23:20:00.147 -> Water_Temp.:103.66, PWM:47
23:20:00.909 -> Water_Temp.:103.66, PWM:47
23:20:01.698 -> Water_Temp.:103.66, PWM:47
23:20:02.493 -> Water_Temp.:103.77, PWM:47
23:20:03.255 -> Water_Temp.:103.66, PWM:47
23:20:04.052 -> Water_Temp.:103.77, PWM:47
23:20:04.812 -> Water_Temp.:103.77, PWM:47
23:20:05.607 -> Water_Temp.:103.77, PWM:47
23:20:06.398 -> Water_Temp.:103.77, PWM:47
23:20:07.159 -> Water_Temp.:103.66, PWM:46
23:20:07.952 -> Water_Temp.:103.77, PWM:47
23:20:08.711 -> Water_Temp.:103.77, PWM:47
23:20:10.299 -> Water_Temp.:103.66, PWM:47
23:20:11.063 -> Water_Temp.:103.66, PWM:47
23:20:12.580 -> Water_Temp.:103.66, PWM:47
23:20:13.376 -> Water_Temp.:103.66, PWM:47
23:20:14.138 -> Water_Temp.:103.66, PWM:47
23:20:14.937 -> Water_Temp.:103.66, PWM:47
23:20:15.699 -> Water_Temp.:103.66, PWM:47
23:20:16.494 -> Water_Temp.:103.66, PWM:47
23:20:17.288 -> Water_Temp.:103.66, PWM:47
23:20:18.049 -> Water_Temp.:103.66, PWM:47
23:20:18.842 -> Water_Temp.:103.66, PWM:46
23:20:19.606 -> Water_Temp.:103.66, PWM:47
23:20:20.400 -> Water_Temp.:103.66, PWM:47
23:20:21.161 -> Water_Temp.:103.66, PWM:47
23:20:21.950 -> Water_Temp.:103.66, PWM:46
23:20:22.738 -> Water_Temp.:103.66, PWM:47
23:20:23.534 -> Water_Temp.:103.66, PWM:47
23:20:24.294 -> Water_Temp.:103.66, PWM:47
23:20:25.091 -> Water_Temp.:103.66, PWM:47
23:20:25.851 -> Water_Temp.:103.66, PWM:47
23:20:26.641 -> Water_Temp.:103.66, PWM:47
23:20:27.401 -> Water_Temp.:105.24, PWM:47
23:20:28.194 -> Water_Temp.:107.26, PWM:47
23:20:28.987 -> Water_Temp.:104.90, PWM:47
23:20:29.746 -> Water_Temp.:104.56, PWM:47
23:20:30.535 -> Water_Temp.:104.34, PWM:47
23:20:31.333 -> Water_Temp.:104.22, PWM:47
23:20:32.092 -> Water_Temp.:104.11, PWM:47
23:20:32.884 -> Water_Temp.:104.00, PWM:47
23:20:33.643 -> Water_Temp.:104.00, PWM:47
23:20:34.436 -> Water_Temp.:104.00, PWM:47
23:20:35.231 -> Water_Temp.:104.00, PWM:47
23:20:35.989 -> Water_Temp.:103.89, PWM:47
23:20:36.785 -> Water_Temp.:103.89, PWM:47
23:20:37.539 -> Water_Temp.:103.89, PWM:47
23:20:38.335 -> Water_Temp.:103.89, PWM:47
23:20:39.129 -> Water_Temp.:103.89, PWM:47
23:20:40.688 -> Water_Temp.:103.89, PWM:47
23:20:41.449 -> Water_Temp.:103.89, PWM:47
23:20:42.246 -> Water_Temp.:103.89, PWM:47
23:20:43.004 -> Water_Temp.:103.89, PWM:47
23:20:43.800 -> Water_Temp.:103.89, PWM:47
23:20:44.557 -> Water_Temp.:103.89, PWM:47
23:20:45.354 -> Water_Temp.:103.89, PWM:47
23:20:46.148 -> Water_Temp.:103.89, PWM:47
23:20:47.668 -> Water_Temp.:103.89, PWM:46
23:20:48.462 -> Water_Temp.:103.89, PWM:47
23:20:49.221 -> Water_Temp.:103.89, PWM:47
23:20:50.015 -> Water_Temp.:103.89, PWM:47
23:20:50.775 -> Water_Temp.:103.89, PWM:47
23:20:51.571 -> Water_Temp.:103.89, PWM:47
23:20:52.359 -> Water_Temp.:104.00, PWM:47
23:20:53.119 -> Water_Temp.:104.00, PWM:47
23:20:53.917 -> Water_Temp.:104.00, PWM:47
23:20:54.679 -> Water_Temp.:104.00, PWM:46
23:20:55.476 -> Water_Temp.:104.00, PWM:47
23:20:56.267 -> Water_Temp.:104.00, PWM:46
23:20:57.822 -> Water_Temp.:104.00, PWM:47
23:20:58.579 -> Water_Temp.:104.00, PWM:47
23:20:59.373 -> Water_Temp.:104.00, PWM:47
23:21:00.167 -> Water_Temp.:104.00, PWM:47
23:21:00.931 -> Water_Temp.:104.00, PWM:47
23:21:01.727 -> Water_Temp.:104.00, PWM:47
23:21:02.487 -> Water_Temp.:104.00, PWM:47
23:21:04.012 -> Water_Temp.:104.00, PWM:47
23:21:04.810 -> Water_Temp.:104.00, PWM:47
23:21:05.605 -> Water_Temp.:104.00, PWM:47
23:21:06.364 -> Water_Temp.:104.00, PWM:47
23:21:07.157 -> Water_Temp.:104.00, PWM:46
23:21:07.913 -> Water_Temp.:104.00, PWM:47
23:21:08.708 -> Water_Temp.:104.00, PWM:47
23:21:09.503 -> Water_Temp.:104.00, PWM:47
23:21:10.262 -> Water_Temp.:104.00, PWM:47
23:21:11.061 -> Water_Temp.:104.00, PWM:47
23:21:11.819 -> Water_Temp.:104.00, PWM:47
23:21:12.606 -> Water_Temp.:104.00, PWM:46
23:21:13.403 -> Water_Temp.:104.00, PWM:47
23:21:14.161 -> Water_Temp.:104.00, PWM:47
23:21:14.959 -> Water_Temp.:104.00, PWM:47
23:21:15.719 -> Water_Temp.:104.00, PWM:47
23:21:16.511 -> Water_Temp.:104.00, PWM:47
23:21:17.303 -> Water_Temp.:104.00, PWM:46
23:21:18.854 -> Water_Temp.:104.00, PWM:47
23:21:19.642 -> Water_Temp.:104.00, PWM:47
23:21:20.400 -> Water_Temp.:104.00, PWM:47
23:21:21.199 -> Water_Temp.:104.00, PWM:47
23:21:21.959 -> Water_Temp.:104.00, PWM:47
23:21:22.753 -> Water_Temp.:104.00, PWM:47
23:21:23.544 -> Water_Temp.:104.00, PWM:47
23:21:24.304 -> Water_Temp.:104.00, PWM:47
23:21:25.100 -> Water_Temp.:104.00, PWM:47
23:21:25.859 -> Water_Temp.:104.00, PWM:47
23:21:26.651 -> Water_Temp.:104.00, PWM:47
23:21:27.411 -> Water_Temp.:104.00, PWM:47
23:21:28.207 -> Water_Temp.:104.00, PWM:47
23:21:29.000 -> Water_Temp.:104.00, PWM:47
23:21:29.757 -> Water_Temp.:104.00, PWM:47
23:21:31.312 -> Water_Temp.:104.00, PWM:46
23:21:32.071 -> Water_Temp.:104.00, PWM:47
23:21:32.861 -> Water_Temp.:104.00, PWM:47
23:21:33.656 -> Water_Temp.:104.00, PWM:47
23:21:35.183 -> Water_Temp.:104.00, PWM:47
23:21:35.976 -> Water_Temp.:104.11, PWM:47
23:21:37.530 -> Water_Temp.:104.11, PWM:47
23:21:38.290 -> Water_Temp.:104.11, PWM:47
23:21:39.085 -> Water_Temp.:104.11, PWM:47
23:21:39.875 -> Water_Temp.:104.11, PWM:47
23:21:40.638 -> Water_Temp.:104.11, PWM:47
23:21:41.436 -> Water_Temp.:104.11, PWM:46
23:21:42.198 -> Water_Temp.:104.11, PWM:46
23:21:42.996 -> Water_Temp.:104.11, PWM:47
23:21:43.757 -> Water_Temp.:104.11, PWM:47
23:21:44.552 -> Water_Temp.:104.11, PWM:47
23:21:45.316 -> Water_Temp.:104.11, PWM:47
23:21:46.109 -> Water_Temp.:104.11, PWM:47
23:21:46.869 -> Water_Temp.:104.11, PWM:47
23:21:47.662 -> Water_Temp.:104.11, PWM:47
23:21:48.455 -> Water_Temp.:104.11, PWM:47
23:21:49.213 -> Water_Temp.:104.11, PWM:47
23:21:50.009 -> Water_Temp.:104.11, PWM:47
23:21:50.772 -> Water_Temp.:104.11, PWM:47
23:21:52.326 -> Water_Temp.:104.11, PWM:47
23:21:53.082 -> Water_Temp.:104.11, PWM:47
23:21:53.875 -> Water_Temp.:104.11, PWM:47
23:21:54.671 -> Water_Temp.:104.11, PWM:47
23:21:55.433 -> Water_Temp.:104.11, PWM:47
23:21:56.228 -> Water_Temp.:104.11, PWM:47
23:21:56.991 -> Water_Temp.:104.11, PWM:47
23:21:57.785 -> Water_Temp.:104.11, PWM:47
23:21:58.543 -> Water_Temp.:104.11, PWM:47
23:21:59.339 -> Water_Temp.:104.11, PWM:47
23:22:00.132 -> Water_Temp.:104.11, PWM:47
23:22:00.893 -> Water_Temp.:104.11, PWM:47
23:22:01.689 -> Water_Temp.:104.11, PWM:47
23:22:02.451 -> Water_Temp.:104.11, PWM:47
23:22:03.252 -> Water_Temp.:104.11, PWM:47
23:22:04.015 -> Water_Temp.:104.11, PWM:47
23:22:04.777 -> Water_Temp.:105.69, PWM:47
23:22:05.574 -> Water_Temp.:107.82, PWM:47
23:22:07.125 -> Water_Temp.:105.01, PWM:47
23:22:07.917 -> Water_Temp.:104.79, PWM:47
23:22:08.677 -> Water_Temp.:104.67, PWM:46
23:22:09.471 -> Water_Temp.:104.56, PWM:47
23:22:10.264 -> Water_Temp.:104.56, PWM:47
23:22:11.825 -> Water_Temp.:104.45, PWM:47
23:22:13.342 -> Water_Temp.:104.34, PWM:47
23:22:14.137 -> Water_Temp.:104.34, PWM:47
23:22:14.935 -> Water_Temp.:104.34, PWM:47
23:22:15.699 -> Water_Temp.:104.34, PWM:47
23:22:16.491 -> Water_Temp.:104.34, PWM:47
23:22:17.254 -> Water_Temp.:104.34, PWM:47
23:22:18.048 -> Water_Temp.:104.34, PWM:47
23:22:18.807 -> Water_Temp.:104.34, PWM:47
23:22:19.603 -> Water_Temp.:104.34, PWM:47
23:22:20.362 -> Water_Temp.:104.34, PWM:47
23:22:21.158 -> Water_Temp.:104.34, PWM:47
23:22:22.710 -> Water_Temp.:104.34, PWM:47
23:22:23.473 -> Water_Temp.:104.34, PWM:47
23:22:24.267 -> Water_Temp.:104.22, PWM:47
23:22:25.032 -> Water_Temp.:104.22, PWM:47
23:22:25.827 -> Water_Temp.:104.22, PWM:47
23:22:26.585 -> Water_Temp.:104.22, PWM:47
23:22:27.380 -> Water_Temp.:104.22, PWM:47
23:22:28.140 -> Water_Temp.:104.22, PWM:47
23:22:29.694 -> Water_Temp.:104.22, PWM:46
23:22:30.490 -> Water_Temp.:104.22, PWM:47
23:22:31.255 -> Water_Temp.:104.22, PWM:47
23:22:32.047 -> Water_Temp.:104.22, PWM:47
23:22:32.805 -> Water_Temp.:104.22, PWM:47
23:22:33.601 -> Water_Temp.:104.22, PWM:47
23:22:34.393 -> Water_Temp.:104.22, PWM:46
23:22:35.919 -> Water_Temp.:104.22, PWM:47
23:22:36.710 -> Water_Temp.:104.22, PWM:47
23:22:37.473 -> Water_Temp.:104.22, PWM:47
23:22:38.264 -> Water_Temp.:104.22, PWM:47
23:22:39.055 -> Water_Temp.:104.22, PWM:47
23:22:39.814 -> Water_Temp.:104.22, PWM:47
23:22:40.608 -> Water_Temp.:104.22, PWM:47
23:22:41.366 -> Water_Temp.:104.22, PWM:47
23:22:42.160 -> Water_Temp.:104.22, PWM:47
23:22:42.951 -> Water_Temp.:104.22, PWM:47
23:22:43.711 -> Water_Temp.:104.22, PWM:47
23:22:44.508 -> Water_Temp.:104.22, PWM:47
23:22:46.064 -> Water_Temp.:104.34, PWM:47
23:22:46.857 -> Water_Temp.:104.34, PWM:47
23:22:47.617 -> Water_Temp.:104.34, PWM:47
23:22:48.416 -> Water_Temp.:104.34, PWM:47
23:22:49.173 -> Water_Temp.:104.34, PWM:47
23:22:49.966 -> Water_Temp.:104.34, PWM:47
23:22:50.724 -> Water_Temp.:104.34, PWM:47
23:22:51.518 -> Water_Temp.:104.34, PWM:47
23:22:52.309 -> Water_Temp.:104.34, PWM:47
23:22:53.070 -> Water_Temp.:104.34, PWM:47
23:22:53.861 -> Water_Temp.:104.34, PWM:47
23:22:54.657 -> Water_Temp.:104.34, PWM:47
23:22:55.420 -> Water_Temp.:104.34, PWM:47
23:22:56.214 -> Water_Temp.:104.45, PWM:47
23:22:56.971 -> Water_Temp.:104.45, PWM:47
23:22:57.760 -> Water_Temp.:104.45, PWM:47
23:22:58.555 -> Water_Temp.:104.45, PWM:47
23:22:59.318 -> Water_Temp.:104.45, PWM:47
23:23:00.112 -> Water_Temp.:104.45, PWM:47
23:23:00.874 -> Water_Temp.:104.45, PWM:47
23:23:01.669 -> Water_Temp.:104.45, PWM:47
23:23:02.428 -> Water_Temp.:104.45, PWM:47
23:23:03.222 -> Water_Temp.:104.45, PWM:47
23:23:04.776 -> Water_Temp.:106.02, PWM:47
23:23:05.572 -> Water_Temp.:108.05, PWM:47
23:23:06.333 -> Water_Temp.:105.69, PWM:47
23:23:07.126 -> Water_Temp.:105.24, PWM:47
23:23:08.645 -> Water_Temp.:104.90, PWM:47
23:23:09.441 -> Water_Temp.:104.79, PWM:47
23:23:10.200 -> Water_Temp.:104.67, PWM:47
23:23:10.998 -> Water_Temp.:104.67, PWM:47
23:23:11.761 -> Water_Temp.:104.56, PWM:47
23:23:12.554 -> Water_Temp.:104.56, PWM:47
23:23:13.345 -> Water_Temp.:104.56, PWM:47
23:23:14.102 -> Water_Temp.:104.56, PWM:47
23:23:14.899 -> Water_Temp.:104.56, PWM:47
23:23:15.662 -> Water_Temp.:104.45, PWM:47
23:23:17.248 -> Water_Temp.:104.45, PWM:47
23:23:18.008 -> Water_Temp.:104.45, PWM:47
23:23:18.802 -> Water_Temp.:104.45, PWM:47
23:23:19.562 -> Water_Temp.:104.45, PWM:47
23:23:20.354 -> Water_Temp.:104.45, PWM:47
23:23:21.117 -> Water_Temp.:104.45, PWM:47
23:23:21.907 -> Water_Temp.:104.45, PWM:47
23:23:22.699 -> Water_Temp.:104.45, PWM:47
23:23:24.252 -> Water_Temp.:104.45, PWM:47
23:23:25.014 -> Water_Temp.:104.45, PWM:47
23:23:25.807 -> Water_Temp.:104.45, PWM:47
23:23:26.599 -> Water_Temp.:104.45, PWM:47
23:23:27.359 -> Water_Temp.:104.45, PWM:47
23:23:28.153 -> Water_Temp.:104.45, PWM:47
23:23:28.946 -> Water_Temp.:106.02, PWM:47
23:23:29.703 -> Water_Temp.:108.05, PWM:47
23:23:30.498 -> Water_Temp.:105.69, PWM:47
23:23:31.260 -> Water_Temp.:105.35, PWM:47
23:23:32.056 -> Water_Temp.:105.12, PWM:47
23:23:32.816 -> Water_Temp.:105.01, PWM:47
23:23:33.610 -> Water_Temp.:104.90, PWM:47
23:23:34.400 -> Water_Temp.:104.90, PWM:47
23:23:35.165 -> Water_Temp.:104.79, PWM:47
23:23:35.957 -> Water_Temp.:104.79, PWM:47
23:23:36.717 -> Water_Temp.:104.67, PWM:47
23:23:37.507 -> Water_Temp.:104.67, PWM:47
23:23:38.300 -> Water_Temp.:104.67, PWM:47
23:23:39.059 -> Water_Temp.:104.67, PWM:47
23:23:39.855 -> Water_Temp.:104.67, PWM:47
23:23:40.619 -> Water_Temp.:104.56, PWM:47
23:23:41.415 -> Water_Temp.:104.56, PWM:47
23:23:42.208 -> Water_Temp.:106.25, PWM:47
23:23:42.968 -> Water_Temp.:108.16, PWM:47
23:23:43.763 -> Water_Temp.:105.80, PWM:47
23:23:44.519 -> Water_Temp.:105.46, PWM:47
23:23:45.318 -> Water_Temp.:105.24, PWM:47
23:23:46.079 -> Water_Temp.:105.01, PWM:47
23:23:46.870 -> Water_Temp.:105.01, PWM:47
23:23:47.665 -> Water_Temp.:104.90, PWM:47
23:23:49.224 -> Water_Temp.:104.79, PWM:47
23:23:49.986 -> Water_Temp.:104.79, PWM:47
23:23:50.783 -> Water_Temp.:104.67, PWM:47
23:23:51.547 -> Water_Temp.:104.67, PWM:47
23:23:52.342 -> Water_Temp.:104.67, PWM:47
23:23:54.661 -> Water_Temp.:104.67, PWM:47
23:23:55.456 -> Water_Temp.:104.67, PWM:47
23:23:56.214 -> Water_Temp.:104.56, PWM:47
23:23:57.008 -> Water_Temp.:104.56, PWM:47
23:23:57.771 -> Water_Temp.:104.56, PWM:47
23:23:58.568 -> Water_Temp.:104.56, PWM:47
23:23:59.366 -> Water_Temp.:104.56, PWM:47
23:24:00.126 -> Water_Temp.:104.56, PWM:47
23:24:00.921 -> Water_Temp.:104.56, PWM:47
23:24:01.678 -> Water_Temp.:104.56, PWM:47
23:24:02.474 -> Water_Temp.:104.56, PWM:47
23:24:03.232 -> Water_Temp.:104.56, PWM:47
23:24:04.029 -> Water_Temp.:104.56, PWM:47
23:24:04.825 -> Water_Temp.:104.56, PWM:47
23:24:05.587 -> Water_Temp.:104.56, PWM:47
23:24:06.385 -> Water_Temp.:104.56, PWM:47
23:24:07.144 -> Water_Temp.:104.56, PWM:47
23:24:07.936 -> Water_Temp.:104.56, PWM:47
23:24:08.695 -> Water_Temp.:104.56, PWM:47
23:24:09.493 -> Water_Temp.:104.45, PWM:47
23:24:10.251 -> Water_Temp.:104.45, PWM:47
23:24:11.051 -> Water_Temp.:104.45, PWM:47
23:24:11.814 -> Water_Temp.:104.45, PWM:47
23:24:12.611 -> Water_Temp.:104.45, PWM:47
23:24:13.403 -> Water_Temp.:104.45, PWM:47
23:24:14.165 -> Water_Temp.:104.45, PWM:47
23:24:14.965 -> Water_Temp.:104.45, PWM:47
23:24:16.519 -> Water_Temp.:104.45, PWM:47
23:24:17.273 -> Water_Temp.:104.45, PWM:47
23:24:18.067 -> Water_Temp.:104.45, PWM:47
23:24:18.861 -> Water_Temp.:104.45, PWM:47
23:24:19.622 -> Water_Temp.:104.45, PWM:47
23:24:20.416 -> Water_Temp.:104.45, PWM:47
23:24:21.177 -> Water_Temp.:104.45, PWM:47
23:24:21.971 -> Water_Temp.:104.45, PWM:47
23:24:22.731 -> Water_Temp.:104.45, PWM:47
23:24:23.526 -> Water_Temp.:104.45, PWM:47
23:24:24.321 -> Water_Temp.:104.45, PWM:47
23:24:25.085 -> Water_Temp.:104.45, PWM:47
23:24:25.877 -> Water_Temp.:104.45, PWM:47
23:24:26.632 -> Water_Temp.:104.56, PWM:47
23:24:27.428 -> Water_Temp.:104.56, PWM:47
23:24:28.190 -> Water_Temp.:104.56, PWM:47
23:24:28.984 -> Water_Temp.:104.56, PWM:47
23:24:29.779 -> Water_Temp.:104.56, PWM:47
23:24:30.538 -> Water_Temp.:104.56, PWM:47
23:24:31.338 -> Water_Temp.:104.56, PWM:47
23:24:32.098 -> Water_Temp.:104.56, PWM:47
23:24:32.892 -> Water_Temp.:104.56, PWM:47
23:24:34.413 -> Water_Temp.:104.56, PWM:47
23:24:35.208 -> Water_Temp.:104.56, PWM:47
23:24:36.003 -> Water_Temp.:104.56, PWM:47
23:24:36.758 -> Water_Temp.:104.56, PWM:47
23:24:37.552 -> Water_Temp.:104.56, PWM:47
23:24:38.345 -> Water_Temp.:104.56, PWM:47
23:24:39.103 -> Water_Temp.:104.56, PWM:47
23:24:39.895 -> Water_Temp.:104.56, PWM:47
23:24:40.657 -> Water_Temp.:104.56, PWM:47
23:24:41.453 -> Water_Temp.:104.56, PWM:47
23:24:42.243 -> Water_Temp.:104.56, PWM:47
23:24:43.001 -> Water_Temp.:104.56, PWM:47
23:24:43.793 -> Water_Temp.:104.45, PWM:47
23:24:44.554 -> Water_Temp.:104.45, PWM:47
23:24:45.351 -> Water_Temp.:104.45, PWM:47
23:24:46.145 -> Water_Temp.:104.45, PWM:47
23:24:46.908 -> Water_Temp.:104.45, PWM:47
23:24:47.702 -> Water_Temp.:104.45, PWM:47
23:24:48.459 -> Water_Temp.:104.45, PWM:47
23:24:49.255 -> Water_Temp.:104.45, PWM:47
23:24:50.014 -> Water_Temp.:104.45, PWM:47
23:24:50.806 -> Water_Temp.:104.45, PWM:47
23:24:51.606 -> Water_Temp.:104.45, PWM:47
23:24:52.365 -> Water_Temp.:104.56, PWM:47
23:24:53.913 -> Water_Temp.:104.56, PWM:47
23:24:54.709 -> Water_Temp.:104.56, PWM:47
23:24:55.470 -> Water_Temp.:104.56, PWM:47
23:24:57.021 -> Water_Temp.:104.56, PWM:47
23:24:57.819 -> Water_Temp.:104.56, PWM:47
23:24:58.574 -> Water_Temp.:104.56, PWM:47
23:25:00.130 -> Water_Temp.:104.56, PWM:47
23:25:00.893 -> Water_Temp.:104.56, PWM:47
23:25:01.688 -> Water_Temp.:104.67, PWM:47
23:25:02.449 -> Water_Temp.:104.67, PWM:47
23:25:04.004 -> Water_Temp.:104.67, PWM:47
23:25:04.766 -> Water_Temp.:104.67, PWM:47
23:25:05.563 -> Water_Temp.:104.67, PWM:47
23:25:06.359 -> Water_Temp.:104.67, PWM:47
23:25:07.912 -> Water_Temp.:104.67, PWM:47
23:25:08.670 -> Water_Temp.:104.67, PWM:47
23:25:09.467 -> Water_Temp.:104.67, PWM:47
23:25:10.228 -> Water_Temp.:104.67, PWM:47
23:25:11.022 -> Water_Temp.:104.67, PWM:47
23:25:11.815 -> Water_Temp.:104.67, PWM:47
23:25:12.578 -> Water_Temp.:104.67, PWM:47
23:25:13.371 -> Water_Temp.:104.67, PWM:47
23:25:14.134 -> Water_Temp.:104.67, PWM:47
23:25:14.932 -> Water_Temp.:104.67, PWM:47
23:25:15.693 -> Water_Temp.:104.67, PWM:47
23:25:16.486 -> Water_Temp.:104.67, PWM:47
23:25:17.247 -> Water_Temp.:104.67, PWM:47
23:25:18.043 -> Water_Temp.:104.67, PWM:47
23:25:18.839 -> Water_Temp.:104.67, PWM:47
23:25:20.394 -> Water_Temp.:104.67, PWM:47
23:25:21.156 -> Water_Temp.:104.67, PWM:47
23:25:21.953 -> Water_Temp.:104.67, PWM:47
23:25:22.711 -> Water_Temp.:104.67, PWM:47
23:25:24.267 -> Water_Temp.:104.67, PWM:47
23:25:25.063 -> Water_Temp.:104.67, PWM:47
23:25:25.854 -> Water_Temp.:104.67, PWM:47
23:25:26.613 -> Water_Temp.:104.67, PWM:47
23:25:27.407 -> Water_Temp.:104.67, PWM:47
23:25:28.170 -> Water_Temp.:104.67, PWM:47
23:25:29.727 -> Water_Temp.:104.67, PWM:47
23:25:30.521 -> Water_Temp.:104.67, PWM:47
23:25:31.284 -> Water_Temp.:104.67, PWM:47
23:25:32.079 -> Water_Temp.:104.79, PWM:47
23:25:32.840 -> Water_Temp.:104.79, PWM:47
23:25:34.393 -> Water_Temp.:104.79, PWM:47
23:25:35.156 -> Water_Temp.:104.79, PWM:47
23:25:35.952 -> Water_Temp.:104.79, PWM:47
23:25:36.709 -> Water_Temp.:104.79, PWM:47
23:25:37.504 -> Water_Temp.:104.79, PWM:47
23:25:38.297 -> Water_Temp.:104.79, PWM:47
23:25:39.059 -> Water_Temp.:106.36, PWM:47
23:25:39.857 -> Water_Temp.:108.39, PWM:47
23:25:40.617 -> Water_Temp.:106.02, PWM:47
23:25:41.414 -> Water_Temp.:105.57, PWM:47
23:25:42.176 -> Water_Temp.:105.35, PWM:47
23:25:42.971 -> Water_Temp.:105.24, PWM:47
23:25:43.733 -> Water_Temp.:105.24, PWM:47
23:25:45.287 -> Water_Temp.:105.01, PWM:47
23:25:46.083 -> Water_Temp.:105.01, PWM:47
23:25:46.844 -> Water_Temp.:105.01, PWM:47
23:25:47.640 -> Water_Temp.:104.90, PWM:47
23:25:48.402 -> Water_Temp.:104.90, PWM:47
23:25:49.202 -> Water_Temp.:104.90, PWM:47
23:25:49.959 -> Water_Temp.:104.90, PWM:47
23:25:50.756 -> Water_Temp.:104.90, PWM:47
23:25:51.517 -> Water_Temp.:104.90, PWM:47
23:25:53.074 -> Water_Temp.:104.79, PWM:47
23:25:53.831 -> Water_Temp.:104.79, PWM:47
23:25:54.628 -> Water_Temp.:104.79, PWM:47
23:25:55.391 -> Water_Temp.:104.79, PWM:47
23:25:56.189 -> Water_Temp.:104.79, PWM:47
23:25:56.979 -> Water_Temp.:104.79, PWM:47
23:25:57.739 -> Water_Temp.:104.79, PWM:47
23:25:58.534 -> Water_Temp.:104.79, PWM:47
23:25:59.294 -> Water_Temp.:104.79, PWM:47
23:26:00.089 -> Water_Temp.:104.79, PWM:47
23:26:00.851 -> Water_Temp.:104.79, PWM:47
23:26:02.405 -> Water_Temp.:104.79, PWM:47
23:26:03.171 -> Water_Temp.:104.79, PWM:47
23:26:03.964 -> Water_Temp.:104.79, PWM:47
23:26:04.756 -> Water_Temp.:104.79, PWM:47
23:26:05.521 -> Water_Temp.:104.79, PWM:47
23:26:06.315 -> Water_Temp.:104.79, PWM:47
23:26:07.074 -> Water_Temp.:104.79, PWM:47
23:26:07.867 -> Water_Temp.:104.79, PWM:47
23:26:08.661 -> Water_Temp.:104.79, PWM:47
23:26:09.422 -> Water_Temp.:104.79, PWM:47
23:26:10.215 -> Water_Temp.:104.79, PWM:47
23:26:10.976 -> Water_Temp.:104.79, PWM:47
23:26:11.770 -> Water_Temp.:104.79, PWM:47
23:26:12.528 -> Water_Temp.:104.79, PWM:47
23:26:13.321 -> Water_Temp.:104.79, PWM:47
23:26:14.115 -> Water_Temp.:104.79, PWM:47
23:26:14.875 -> Water_Temp.:104.90, PWM:47
23:26:15.671 -> Water_Temp.:104.79, PWM:47
23:26:16.432 -> Water_Temp.:104.79, PWM:47
23:26:17.228 -> Water_Temp.:104.90, PWM:47
23:26:17.991 -> Water_Temp.:104.90, PWM:47
23:26:18.786 -> Water_Temp.:104.90, PWM:47
23:26:19.547 -> Water_Temp.:104.79, PWM:47
23:26:20.336 -> Water_Temp.:104.90, PWM:47
23:26:21.135 -> Water_Temp.:104.90, PWM:47
23:26:21.893 -> Water_Temp.:104.90, PWM:47
23:26:22.689 -> Water_Temp.:104.90, PWM:47
23:26:23.449 -> Water_Temp.:104.79, PWM:47
23:26:24.246 -> Water_Temp.:104.79, PWM:47
23:26:25.009 -> Water_Temp.:104.79, PWM:47
23:26:26.592 -> Water_Temp.:104.79, PWM:47
23:26:27.352 -> Water_Temp.:104.79, PWM:47
23:26:28.148 -> Water_Temp.:104.79, PWM:47
23:26:28.906 -> Water_Temp.:104.79, PWM:47
23:26:29.702 -> Water_Temp.:104.79, PWM:47
23:26:30.498 -> Water_Temp.:104.79, PWM:47
23:26:31.263 -> Water_Temp.:104.79, PWM:47
23:26:32.054 -> Water_Temp.:104.79, PWM:47
23:26:32.816 -> Water_Temp.:104.79, PWM:47
23:26:33.610 -> Water_Temp.:104.79, PWM:47
23:26:34.369 -> Water_Temp.:104.79, PWM:47
23:26:35.169 -> Water_Temp.:104.79, PWM:47
23:26:35.930 -> Water_Temp.:104.79, PWM:47
23:26:36.724 -> Water_Temp.:104.79, PWM:47
23:26:37.519 -> Water_Temp.:104.79, PWM:47
23:26:38.278 -> Water_Temp.:104.79, PWM:47
23:26:39.072 -> Water_Temp.:104.79, PWM:47
23:26:39.829 -> Water_Temp.:104.79, PWM:47
23:26:40.625 -> Water_Temp.:104.79, PWM:47
23:26:42.143 -> Water_Temp.:104.90, PWM:47
23:26:42.941 -> Water_Temp.:104.79, PWM:47
23:26:43.737 -> Water_Temp.:104.90, PWM:47
23:26:44.500 -> Water_Temp.:104.79, PWM:47
23:26:45.261 -> Water_Temp.:104.79, PWM:47
23:26:46.058 -> Water_Temp.:104.79, PWM:47
23:26:46.853 -> Water_Temp.:104.79, PWM:47
23:26:47.612 -> Water_Temp.:104.79, PWM:47
23:26:48.405 -> Water_Temp.:104.79, PWM:47
23:26:49.198 -> Water_Temp.:104.79, PWM:47
23:26:49.961 -> Water_Temp.:104.79, PWM:47
23:26:50.723 -> Water_Temp.:104.79, PWM:47
23:26:51.517 -> Water_Temp.:104.79, PWM:47
23:26:52.310 -> Water_Temp.:104.79, PWM:47
23:26:53.067 -> Water_Temp.:104.79, PWM:47
23:26:53.854 -> Water_Temp.:104.79, PWM:47
23:26:54.646 -> Water_Temp.:104.79, PWM:47
23:26:55.407 -> Water_Temp.:104.79, PWM:47
23:26:56.201 -> Water_Temp.:104.90, PWM:47
23:26:56.963 -> Water_Temp.:104.79, PWM:47
23:26:57.755 -> Water_Temp.:104.90, PWM:47
23:26:58.548 -> Water_Temp.:104.90, PWM:47
23:26:59.309 -> Water_Temp.:104.90, PWM:52
23:27:00.102 -> Water_Temp.:104.90, PWM:51
23:27:00.865 -> Water_Temp.:104.79, PWM:51
23:27:01.663 -> Water_Temp.:104.90, PWM:51
23:27:02.422 -> Water_Temp.:104.90, PWM:49
23:27:03.216 -> Water_Temp.:104.79, PWM:50
23:27:04.014 -> Water_Temp.:104.79, PWM:50
23:27:04.777 -> Water_Temp.:104.79, PWM:50
23:27:05.573 -> Water_Temp.:104.79, PWM:50
23:27:06.333 -> Water_Temp.:104.79, PWM:50
23:27:07.128 -> Water_Temp.:104.79, PWM:50
23:27:07.892 -> Water_Temp.:104.79, PWM:50
23:27:08.689 -> Water_Temp.:104.79, PWM:50
23:27:09.450 -> Water_Temp.:104.79, PWM:50
23:27:10.247 -> Water_Temp.:104.79, PWM:50
23:27:11.007 -> Water_Temp.:104.79, PWM:49
23:27:11.796 -> Water_Temp.:104.90, PWM:50
23:27:12.591 -> Water_Temp.:104.90, PWM:49
23:27:13.353 -> Water_Temp.:104.90, PWM:50
23:27:14.150 -> Water_Temp.:104.90, PWM:50
23:27:14.914 -> Water_Temp.:104.90, PWM:49
23:27:15.711 -> Water_Temp.:104.90, PWM:49
23:27:16.467 -> Water_Temp.:104.90, PWM:50
23:27:17.263 -> Water_Temp.:104.90, PWM:50
23:27:18.023 -> Water_Temp.:104.90, PWM:49
23:27:18.820 -> Water_Temp.:104.90, PWM:50
23:27:19.582 -> Water_Temp.:104.90, PWM:50
23:27:20.371 -> Water_Temp.:104.90, PWM:49
23:27:21.168 -> Water_Temp.:104.90, PWM:49
23:27:21.926 -> Water_Temp.:104.90, PWM:50
23:27:22.719 -> Water_Temp.:104.90, PWM:50
23:27:23.510 -> Water_Temp.:104.90, PWM:49
23:27:24.268 -> Water_Temp.:104.90, PWM:50
23:27:25.060 -> Water_Temp.:104.90, PWM:50
23:27:25.822 -> Water_Temp.:104.90, PWM:50
23:27:26.618 -> Water_Temp.:105.01, PWM:50
23:27:27.415 -> Water_Temp.:104.90, PWM:49
23:27:28.173 -> Water_Temp.:105.01, PWM:49
23:27:29.725 -> Water_Temp.:104.90, PWM:49
23:27:30.520 -> Water_Temp.:104.90, PWM:50
23:27:31.314 -> Water_Temp.:104.90, PWM:50
23:27:32.076 -> Water_Temp.:104.90, PWM:50
23:27:32.871 -> Water_Temp.:104.90, PWM:50
23:27:33.631 -> Water_Temp.:104.90, PWM:50
23:27:34.428 -> Water_Temp.:104.90, PWM:49
23:27:35.191 -> Water_Temp.:104.90, PWM:50
23:27:35.981 -> Water_Temp.:106.47, PWM:49
23:27:36.773 -> Water_Temp.:108.50, PWM:50
23:27:37.534 -> Water_Temp.:106.14, PWM:50
23:27:38.329 -> Water_Temp.:105.69, PWM:49
23:27:39.087 -> Water_Temp.:105.46, PWM:50
23:27:39.885 -> Water_Temp.:105.35, PWM:50
23:27:40.650 -> Water_Temp.:105.24, PWM:50
23:27:41.445 -> Water_Temp.:105.12, PWM:50
23:27:42.206 -> Water_Temp.:105.12, PWM:50
23:27:43.000 -> Water_Temp.:105.01, PWM:50
23:27:43.760 -> Water_Temp.:105.01, PWM:50
23:27:44.560 -> Water_Temp.:105.01, PWM:50
23:27:45.322 -> Water_Temp.:104.90, PWM:50
23:27:46.116 -> Water_Temp.:104.90, PWM:50
23:27:47.662 -> Water_Temp.:104.90, PWM:50
23:27:48.458 -> Water_Temp.:104.90, PWM:50
23:27:49.219 -> Water_Temp.:104.90, PWM:50
23:27:50.014 -> Water_Temp.:104.90, PWM:50
23:27:50.772 -> Water_Temp.:104.90, PWM:50
23:27:51.567 -> Water_Temp.:104.90, PWM:50
23:27:52.359 -> Water_Temp.:104.90, PWM:49
23:27:53.914 -> Water_Temp.:104.90, PWM:50
23:27:54.674 -> Water_Temp.:104.90, PWM:50
23:27:55.470 -> Water_Temp.:104.90, PWM:49
23:27:56.229 -> Water_Temp.:104.90, PWM:49
23:27:57.023 -> Water_Temp.:104.90, PWM:50
23:27:57.818 -> Water_Temp.:104.79, PWM:50
23:27:58.581 -> Water_Temp.:104.79, PWM:50
23:27:59.376 -> Water_Temp.:104.79, PWM:50
23:28:00.133 -> Water_Temp.:104.79, PWM:50
23:28:00.928 -> Water_Temp.:104.90, PWM:50
23:28:01.689 -> Water_Temp.:104.90, PWM:50
23:28:02.480 -> Water_Temp.:104.90, PWM:50
23:28:03.270 -> Water_Temp.:104.90, PWM:50
23:28:04.825 -> Water_Temp.:104.90, PWM:50
23:28:05.584 -> Water_Temp.:104.90, PWM:50
23:28:06.378 -> Water_Temp.:104.90, PWM:50
23:28:07.135 -> Water_Temp.:104.90, PWM:50
23:28:07.931 -> Water_Temp.:104.90, PWM:50
23:28:08.727 -> Water_Temp.:104.90, PWM:49
23:28:09.488 -> Water_Temp.:104.90, PWM:50
23:28:10.281 -> Water_Temp.:104.90, PWM:49
23:28:11.045 -> Water_Temp.:104.90, PWM:50
23:28:11.835 -> Water_Temp.:104.90, PWM:49
23:28:12.627 -> Water_Temp.:104.90, PWM:50
23:28:13.382 -> Water_Temp.:104.90, PWM:49
23:28:14.173 -> Water_Temp.:104.90, PWM:50
23:28:14.937 -> Water_Temp.:104.90, PWM:50
23:28:15.736 -> Water_Temp.:104.90, PWM:49
23:28:16.494 -> Water_Temp.:105.01, PWM:50
23:28:17.289 -> Water_Temp.:105.01, PWM:49
23:28:18.086 -> Water_Temp.:104.90, PWM:50
23:28:18.849 -> Water_Temp.:105.01, PWM:49
23:28:19.642 -> Water_Temp.:105.01, PWM:49
23:28:20.398 -> Water_Temp.:105.01, PWM:50
23:28:21.192 -> Water_Temp.:105.01, PWM:49
23:28:22.740 -> Water_Temp.:105.01, PWM:50
23:28:24.260 -> Water_Temp.:105.01, PWM:49
23:28:25.055 -> Water_Temp.:105.01, PWM:50
23:28:25.814 -> Water_Temp.:105.01, PWM:50
23:28:26.607 -> Water_Temp.:105.01, PWM:50
23:28:27.398 -> Water_Temp.:105.01, PWM:50
23:28:28.154 -> Water_Temp.:105.01, PWM:49
23:28:28.949 -> Water_Temp.:105.01, PWM:50
23:28:29.712 -> Water_Temp.:105.01, PWM:49
23:28:30.507 -> Water_Temp.:105.01, PWM:49
23:28:31.268 -> Water_Temp.:105.01, PWM:50
23:28:32.061 -> Water_Temp.:105.01, PWM:50
23:28:32.856 -> Water_Temp.:105.01, PWM:50
23:28:33.615 -> Water_Temp.:105.01, PWM:50
23:28:34.408 -> Water_Temp.:105.01, PWM:50
23:28:35.171 -> Water_Temp.:105.01, PWM:50
23:28:36.759 -> Water_Temp.:105.01, PWM:50
23:28:38.309 -> Water_Temp.:105.12, PWM:49
23:28:39.071 -> Water_Temp.:105.12, PWM:50
23:28:39.862 -> Water_Temp.:105.12, PWM:50
23:28:40.656 -> Water_Temp.:105.12, PWM:49
23:28:41.418 -> Water_Temp.:105.12, PWM:50
23:28:42.212 -> Water_Temp.:105.12, PWM:49
23:28:42.974 -> Water_Temp.:105.12, PWM:49
23:28:43.770 -> Water_Temp.:105.12, PWM:50
23:28:44.530 -> Water_Temp.:105.12, PWM:49
23:28:45.324 -> Water_Temp.:105.12, PWM:50
23:28:46.118 -> Water_Temp.:105.12, PWM:50
23:28:46.875 -> Water_Temp.:105.12, PWM:50
23:28:47.668 -> Water_Temp.:105.12, PWM:49
23:28:48.426 -> Water_Temp.:105.12, PWM:50
23:28:49.217 -> Water_Temp.:105.12, PWM:50
23:28:50.009 -> Water_Temp.:105.12, PWM:50
23:28:50.770 -> Water_Temp.:105.12, PWM:50
23:28:51.567 -> Water_Temp.:105.12, PWM:50
23:28:52.329 -> Water_Temp.:105.12, PWM:50
23:28:53.120 -> Water_Temp.:105.12, PWM:50
23:28:53.917 -> Water_Temp.:105.12, PWM:49
23:28:54.678 -> Water_Temp.:105.12, PWM:50
23:28:55.473 -> Water_Temp.:105.12, PWM:50
23:28:56.228 -> Water_Temp.:105.12, PWM:50
23:28:57.779 -> Water_Temp.:105.12, PWM:50
23:28:58.575 -> Water_Temp.:105.24, PWM:50
23:29:00.133 -> Water_Temp.:105.24, PWM:50
23:29:00.892 -> Water_Temp.:105.24, PWM:49
23:29:01.686 -> Water_Temp.:105.24, PWM:49
23:29:02.446 -> Water_Temp.:105.24, PWM:50
23:29:03.241 -> Water_Temp.:105.24, PWM:50
23:29:04.800 -> Water_Temp.:105.24, PWM:49
23:29:06.359 -> Water_Temp.:105.24, PWM:50
23:29:07.152 -> Water_Temp.:105.24, PWM:50
23:29:07.912 -> Water_Temp.:105.24, PWM:50
23:29:08.705 -> Water_Temp.:105.24, PWM:49
23:29:09.464 -> Water_Temp.:105.24, PWM:50
23:29:10.260 -> Water_Temp.:105.24, PWM:49
23:29:11.022 -> Water_Temp.:105.35, PWM:50
23:29:11.812 -> Water_Temp.:105.35, PWM:49
23:29:12.606 -> Water_Temp.:105.24, PWM:50
23:29:13.370 -> Water_Temp.:105.35, PWM:50
23:29:14.165 -> Water_Temp.:105.24, PWM:50
23:29:14.929 -> Water_Temp.:105.35, PWM:50
23:29:15.726 -> Water_Temp.:105.24, PWM:49
23:29:16.486 -> Water_Temp.:105.24, PWM:50
23:29:17.279 -> Water_Temp.:105.24, PWM:50
23:29:18.075 -> Water_Temp.:105.24, PWM:50
23:29:18.837 -> Water_Temp.:105.24, PWM:50
23:29:19.631 -> Water_Temp.:105.24, PWM:50
23:29:20.387 -> Water_Temp.:105.35, PWM:50
23:29:21.181 -> Water_Temp.:105.35, PWM:50
23:29:21.942 -> Water_Temp.:105.35, PWM:50
23:29:22.738 -> Water_Temp.:105.35, PWM:49
23:29:23.502 -> Water_Temp.:105.35, PWM:50
23:29:24.298 -> Water_Temp.:105.35, PWM:50
23:29:25.095 -> Water_Temp.:105.35, PWM:50
23:29:25.856 -> Water_Temp.:105.35, PWM:50
23:29:26.647 -> Water_Temp.:105.35, PWM:50
23:29:27.408 -> Water_Temp.:105.35, PWM:50
23:29:28.203 -> Water_Temp.:105.35, PWM:50
23:29:28.966 -> Water_Temp.:105.35, PWM:50
23:29:29.758 -> Water_Temp.:105.35, PWM:49
23:29:30.553 -> Water_Temp.:105.35, PWM:49
23:29:31.314 -> Water_Temp.:105.24, PWM:49
23:29:32.105 -> Water_Temp.:105.35, PWM:50
23:29:32.870 -> Water_Temp.:105.35, PWM:49
23:29:33.663 -> Water_Temp.:105.24, PWM:49
23:29:34.425 -> Water_Temp.:105.35, PWM:50
23:29:35.224 -> Water_Temp.:105.35, PWM:50
23:29:35.984 -> Water_Temp.:105.35, PWM:50
23:29:36.774 -> Water_Temp.:105.24, PWM:50
23:29:37.565 -> Water_Temp.:105.24, PWM:50
23:29:38.320 -> Water_Temp.:105.35, PWM:49
23:29:39.115 -> Water_Temp.:105.24, PWM:50
23:29:39.904 -> Water_Temp.:105.24, PWM:50
23:29:40.666 -> Water_Temp.:105.24, PWM:50
23:29:41.460 -> Water_Temp.:105.24, PWM:50
23:29:42.256 -> Water_Temp.:105.24, PWM:50
23:29:43.016 -> Water_Temp.:105.35, PWM:49
23:29:43.812 -> Water_Temp.:105.24, PWM:50
23:29:44.575 -> Water_Temp.:105.24, PWM:50
23:29:45.372 -> Water_Temp.:105.24, PWM:50
23:29:46.135 -> Water_Temp.:105.35, PWM:50
23:29:46.921 -> Water_Temp.:105.35, PWM:50
23:29:47.716 -> Water_Temp.:105.35, PWM:50
23:29:49.264 -> Water_Temp.:105.35, PWM:50
23:29:50.025 -> Water_Temp.:105.35, PWM:49
23:29:50.819 -> Water_Temp.:105.35, PWM:50
23:29:51.582 -> Water_Temp.:105.35, PWM:50
23:29:52.374 -> Water_Temp.:105.35, PWM:50
23:29:53.170 -> Water_Temp.:105.35, PWM:50
23:29:53.936 -> Water_Temp.:105.35, PWM:50
23:29:54.730 -> Water_Temp.:105.35, PWM:50
23:29:55.490 -> Water_Temp.:105.35, PWM:50
23:29:56.284 -> Water_Temp.:105.35, PWM:50
23:29:57.072 -> Water_Temp.:105.35, PWM:50
23:29:57.831 -> Water_Temp.:105.35, PWM:50
23:29:58.625 -> Water_Temp.:105.35, PWM:50
23:29:59.382 -> Water_Temp.:105.35, PWM:50
23:30:00.178 -> Water_Temp.:105.35, PWM:50
23:30:00.972 -> Water_Temp.:105.35, PWM:49
23:30:01.733 -> Water_Temp.:105.35, PWM:49
23:30:02.528 -> Water_Temp.:105.35, PWM:50
23:30:03.289 -> Water_Temp.:105.35, PWM:50
23:30:04.080 -> Water_Temp.:105.35, PWM:50
23:30:04.840 -> Water_Temp.:105.35, PWM:50
23:30:05.633 -> Water_Temp.:105.35, PWM:49
23:30:07.181 -> Water_Temp.:105.35, PWM:50
23:30:07.977 -> Water_Temp.:105.35, PWM:50
23:30:08.739 -> Water_Temp.:105.35, PWM:50
23:30:09.534 -> Water_Temp.:105.35, PWM:50
23:30:10.294 -> Water_Temp.:105.35, PWM:49
23:30:11.094 -> Water_Temp.:105.35, PWM:49
23:30:11.852 -> Water_Temp.:105.35, PWM:50
23:30:13.405 -> Water_Temp.:105.35, PWM:49
23:30:14.195 -> Water_Temp.:105.35, PWM:50
23:30:14.957 -> Water_Temp.:105.35, PWM:50
23:30:15.751 -> Water_Temp.:105.35, PWM:49
23:30:16.539 -> Water_Temp.:105.35, PWM:49
23:30:17.301 -> Water_Temp.:105.35, PWM:50
23:30:18.093 -> Water_Temp.:105.35, PWM:50
23:30:18.854 -> Water_Temp.:105.35, PWM:50
23:30:19.645 -> Water_Temp.:105.35, PWM:50
23:30:20.434 -> Water_Temp.:105.35, PWM:49
23:30:21.197 -> Water_Temp.:105.35, PWM:49
23:30:21.992 -> Water_Temp.:105.24, PWM:50
23:30:22.753 -> Water_Temp.:105.35, PWM:49
23:30:23.549 -> Water_Temp.:105.35, PWM:50
23:30:24.309 -> Water_Temp.:105.24, PWM:50
23:30:25.107 -> Water_Temp.:105.24, PWM:49
23:30:25.901 -> Water_Temp.:105.24, PWM:50
23:30:26.657 -> Water_Temp.:105.24, PWM:50
23:30:27.451 -> Water_Temp.:105.35, PWM:50
23:30:28.214 -> Water_Temp.:105.35, PWM:50
23:30:29.012 -> Water_Temp.:105.35, PWM:50
23:30:29.774 -> Water_Temp.:105.35, PWM:49
23:30:30.571 -> Water_Temp.:105.35, PWM:50
23:30:31.334 -> Water_Temp.:105.35, PWM:50
23:30:32.891 -> Water_Temp.:105.35, PWM:50
23:30:33.684 -> Water_Temp.:105.35, PWM:50
23:30:34.446 -> Water_Temp.:105.35, PWM:49
23:30:35.239 -> Water_Temp.:105.46, PWM:50
23:30:35.996 -> Water_Temp.:105.35, PWM:50
23:30:36.786 -> Water_Temp.:105.46, PWM:50
23:30:37.584 -> Water_Temp.:105.46, PWM:50
23:30:38.343 -> Water_Temp.:105.46, PWM:50
23:30:39.135 -> Water_Temp.:105.46, PWM:49
23:30:40.687 -> Water_Temp.:105.46, PWM:50
23:30:41.447 -> Water_Temp.:105.46, PWM:50
23:30:42.242 -> Water_Temp.:105.46, PWM:49
23:30:43.000 -> Water_Temp.:105.46, PWM:50
23:30:43.793 -> Water_Temp.:105.46, PWM:50
23:30:44.588 -> Water_Temp.:105.46, PWM:49
23:30:45.344 -> Water_Temp.:105.46, PWM:50
23:30:46.137 -> Water_Temp.:105.46, PWM:50
23:30:46.898 -> Water_Temp.:105.46, PWM:49
23:30:47.694 -> Water_Temp.:105.46, PWM:50
23:30:48.487 -> Water_Temp.:105.46, PWM:50
23:30:49.250 -> Water_Temp.:105.46, PWM:50
23:30:50.048 -> Water_Temp.:105.46, PWM:50
23:30:50.810 -> Water_Temp.:105.46, PWM:49
23:30:51.608 -> Water_Temp.:105.46, PWM:50
23:30:52.368 -> Water_Temp.:105.46, PWM:49
23:30:53.159 -> Water_Temp.:105.46, PWM:49
23:30:53.916 -> Water_Temp.:105.46, PWM:50
23:30:54.707 -> Water_Temp.:105.46, PWM:49
23:30:55.501 -> Water_Temp.:105.46, PWM:49
23:30:56.263 -> Water_Temp.:105.57, PWM:49
23:30:57.817 -> Water_Temp.:105.57, PWM:50
23:30:59.394 -> Water_Temp.:105.57, PWM:50
23:31:00.154 -> Water_Temp.:105.57, PWM:49
23:31:00.954 -> Water_Temp.:105.57, PWM:50
23:31:01.718 -> Water_Temp.:105.46, PWM:50
23:31:02.509 -> Water_Temp.:105.46, PWM:49
23:31:03.271 -> Water_Temp.:105.46, PWM:49
23:31:04.064 -> Water_Temp.:105.57, PWM:50
23:31:04.822 -> Water_Temp.:105.57, PWM:50
23:31:05.618 -> Water_Temp.:105.57, PWM:50
23:31:06.409 -> Water_Temp.:105.57, PWM:50
23:31:07.170 -> Water_Temp.:105.57, PWM:50
23:31:07.962 -> Water_Temp.:105.57, PWM:50
23:31:08.725 -> Water_Temp.:105.57, PWM:49
23:31:09.518 -> Water_Temp.:105.57, PWM:50
23:31:10.314 -> Water_Temp.:105.57, PWM:50
23:31:11.074 -> Water_Temp.:105.57, PWM:50
23:31:11.869 -> Water_Temp.:105.57, PWM:49
23:31:12.626 -> Water_Temp.:105.57, PWM:50
23:31:13.417 -> Water_Temp.:105.57, PWM:50
23:31:14.209 -> Water_Temp.:105.57, PWM:50
23:31:14.970 -> Water_Temp.:105.57, PWM:50
23:31:15.768 -> Water_Temp.:105.57, PWM:49
23:31:16.525 -> Water_Temp.:105.57, PWM:49
23:31:17.318 -> Water_Temp.:105.57, PWM:49
23:31:18.869 -> Water_Temp.:105.57, PWM:50
23:31:19.630 -> Water_Temp.:105.57, PWM:50
23:31:20.424 -> Water_Temp.:105.57, PWM:50
23:31:21.188 -> Water_Temp.:105.57, PWM:50
23:31:21.984 -> Water_Temp.:105.57, PWM:50
23:31:22.774 -> Water_Temp.:105.57, PWM:49
23:31:23.535 -> Water_Temp.:105.57, PWM:49
23:31:24.330 -> Water_Temp.:105.57, PWM:50
23:31:25.090 -> Water_Temp.:105.57, PWM:49
23:31:25.883 -> Water_Temp.:105.57, PWM:50
23:31:27.430 -> Water_Temp.:105.57, PWM:50
23:31:28.983 -> Water_Temp.:105.57, PWM:50
23:31:29.777 -> Water_Temp.:105.57, PWM:50
23:31:30.538 -> Water_Temp.:105.57, PWM:49
23:31:31.338 -> Water_Temp.:105.57, PWM:49
23:31:32.098 -> Water_Temp.:105.57, PWM:50
23:31:32.890 -> Water_Temp.:105.57, PWM:50
23:31:33.682 -> Water_Temp.:105.57, PWM:49
23:31:34.440 -> Water_Temp.:105.57, PWM:50
23:31:35.239 -> Water_Temp.:105.57, PWM:50
23:31:36.000 -> Water_Temp.:105.57, PWM:50
23:31:36.792 -> Water_Temp.:105.57, PWM:50
23:31:37.552 -> Water_Temp.:105.57, PWM:50
23:31:38.343 -> Water_Temp.:105.57, PWM:50
23:31:39.140 -> Water_Temp.:105.57, PWM:50
23:31:39.901 -> Water_Temp.:105.57, PWM:50
23:31:40.696 -> Water_Temp.:105.57, PWM:50
23:31:41.459 -> Water_Temp.:105.57, PWM:50
23:31:42.253 -> Water_Temp.:105.57, PWM:49
23:31:43.013 -> Water_Temp.:105.57, PWM:50
23:31:43.808 -> Water_Temp.:105.57, PWM:50
23:31:44.601 -> Water_Temp.:105.57, PWM:50
23:31:45.360 -> Water_Temp.:105.57, PWM:50
23:31:46.154 -> Water_Temp.:105.57, PWM:49
23:31:46.916 -> Water_Temp.:105.57, PWM:50
23:31:48.468 -> Water_Temp.:105.57, PWM:50
23:31:49.258 -> Water_Temp.:105.57, PWM:49
23:31:50.023 -> Water_Temp.:105.57, PWM:50
23:31:50.818 -> Water_Temp.:105.57, PWM:50
23:31:51.579 -> Water_Temp.:105.57, PWM:50
23:31:52.374 -> Water_Temp.:105.57, PWM:49
23:31:53.135 -> Water_Temp.:105.57, PWM:50
23:31:53.934 -> Water_Temp.:105.57, PWM:50
23:31:54.694 -> Water_Temp.:105.57, PWM:49
23:31:55.492 -> Water_Temp.:105.57, PWM:50
23:31:56.250 -> Water_Temp.:105.57, PWM:50
23:31:57.042 -> Water_Temp.:105.57, PWM:50
23:31:57.834 -> Water_Temp.:105.57, PWM:50
23:31:58.597 -> Water_Temp.:105.57, PWM:50
23:31:59.390 -> Water_Temp.:105.57, PWM:49
23:32:00.148 -> Water_Temp.:105.57, PWM:50
23:32:00.945 -> Water_Temp.:105.57, PWM:49
23:32:01.740 -> Water_Temp.:105.57, PWM:49
23:32:02.499 -> Water_Temp.:105.57, PWM:50
23:32:03.296 -> Water_Temp.:105.57, PWM:50
23:32:04.058 -> Water_Temp.:105.57, PWM:50
23:32:04.850 -> Water_Temp.:105.57, PWM:50
23:32:05.610 -> Water_Temp.:105.57, PWM:50
23:32:06.405 -> Water_Temp.:105.57, PWM:50
23:32:07.196 -> Water_Temp.:105.57, PWM:50
23:32:07.958 -> Water_Temp.:105.57, PWM:50
23:32:08.756 -> Water_Temp.:105.57, PWM:49
23:32:09.516 -> Water_Temp.:105.57, PWM:50
23:32:11.068 -> Water_Temp.:105.69, PWM:50
23:32:11.861 -> Water_Temp.:105.69, PWM:50
23:32:12.622 -> Water_Temp.:105.69, PWM:49
23:32:13.417 -> Water_Temp.:105.69, PWM:50
23:32:14.180 -> Water_Temp.:105.69, PWM:50
23:32:14.977 -> Water_Temp.:105.69, PWM:50
23:32:15.740 -> Water_Temp.:105.69, PWM:49
23:32:16.530 -> Water_Temp.:105.69, PWM:50
23:32:17.325 -> Water_Temp.:105.69, PWM:50
23:32:18.087 -> Water_Temp.:105.69, PWM:50
23:32:18.882 -> Water_Temp.:105.69, PWM:50
23:32:19.641 -> Water_Temp.:105.69, PWM:50
23:32:20.435 -> Water_Temp.:105.69, PWM:50
23:32:21.201 -> Water_Temp.:105.69, PWM:50
23:32:21.995 -> Water_Temp.:105.69, PWM:50
23:32:22.753 -> Water_Temp.:105.69, PWM:50
23:32:23.541 -> Water_Temp.:105.69, PWM:50
23:32:24.338 -> Water_Temp.:105.69, PWM:50
23:32:25.101 -> Water_Temp.:105.69, PWM:50
23:32:25.895 -> Water_Temp.:105.69, PWM:49
23:32:26.686 -> Water_Temp.:105.69, PWM:49
23:32:27.445 -> Water_Temp.:105.69, PWM:49
23:32:28.237 -> Water_Temp.:105.69, PWM:50
23:32:28.996 -> Water_Temp.:105.69, PWM:50
23:32:29.789 -> Water_Temp.:105.69, PWM:50
23:32:30.585 -> Water_Temp.:105.69, PWM:50
23:32:32.106 -> Water_Temp.:105.69, PWM:50
23:32:32.902 -> Water_Temp.:105.69, PWM:50
23:32:33.663 -> Water_Temp.:105.69, PWM:50
23:32:34.454 -> Water_Temp.:105.69, PWM:49
23:32:35.213 -> Water_Temp.:105.69, PWM:50
23:32:36.008 -> Water_Temp.:105.69, PWM:50
23:32:36.797 -> Water_Temp.:105.69, PWM:49
23:32:37.555 -> Water_Temp.:105.69, PWM:50
23:32:38.351 -> Water_Temp.:105.69, PWM:50
23:32:39.112 -> Water_Temp.:105.69, PWM:50
23:32:39.903 -> Water_Temp.:105.69, PWM:50
23:32:40.700 -> Water_Temp.:105.69, PWM:50
23:32:41.462 -> Water_Temp.:105.69, PWM:50
23:32:42.254 -> Water_Temp.:105.69, PWM:50
23:32:43.048 -> Water_Temp.:105.69, PWM:49
23:32:43.807 -> Water_Temp.:105.69, PWM:49
23:32:45.366 -> Water_Temp.:105.69, PWM:50
23:32:46.160 -> Water_Temp.:105.69, PWM:50
23:32:46.914 -> Water_Temp.:105.69, PWM:50
23:32:47.708 -> Water_Temp.:105.69, PWM:50
23:32:48.501 -> Water_Temp.:105.69, PWM:50
23:32:49.259 -> Water_Temp.:105.69, PWM:50
23:32:50.056 -> Water_Temp.:105.69, PWM:50
23:32:50.815 -> Water_Temp.:105.69, PWM:50
23:32:51.609 -> Water_Temp.:105.69, PWM:49
23:32:52.398 -> Water_Temp.:105.69, PWM:50
23:32:53.157 -> Water_Temp.:105.69, PWM:49
23:32:53.949 -> Water_Temp.:105.69, PWM:50
23:32:54.745 -> Water_Temp.:105.69, PWM:50
23:32:55.503 -> Water_Temp.:105.69, PWM:50
23:32:56.295 -> Water_Temp.:105.69, PWM:49
23:32:57.055 -> Water_Temp.:105.69, PWM:50
23:32:57.849 -> Water_Temp.:105.69, PWM:50
23:32:58.641 -> Water_Temp.:105.69, PWM:50
23:32:59.404 -> Water_Temp.:105.69, PWM:50
23:33:00.198 -> Water_Temp.:105.69, PWM:49
23:33:00.962 -> Water_Temp.:105.69, PWM:49
23:33:01.759 -> Water_Temp.:105.69, PWM:50
23:33:02.522 -> Water_Temp.:105.69, PWM:50
23:33:03.319 -> Water_Temp.:105.69, PWM:50
23:33:04.085 -> Water_Temp.:105.69, PWM:50
23:33:04.880 -> Water_Temp.:105.69, PWM:50
23:33:05.643 -> Water_Temp.:105.69, PWM:50
23:33:07.201 -> Water_Temp.:105.69, PWM:50
23:33:07.997 -> Water_Temp.:105.69, PWM:50
23:33:08.757 -> Water_Temp.:105.69, PWM:50
23:33:09.552 -> Water_Temp.:105.69, PWM:50
23:33:11.109 -> Water_Temp.:105.69, PWM:49
23:33:12.629 -> Water_Temp.:105.69, PWM:50
23:33:13.424 -> Water_Temp.:105.69, PWM:49
23:33:14.216 -> Water_Temp.:105.69, PWM:50
23:33:14.981 -> Water_Temp.:105.69, PWM:50
23:33:15.744 -> Water_Temp.:105.69, PWM:50
23:33:16.538 -> Water_Temp.:105.80, PWM:49
23:33:17.329 -> Water_Temp.:105.69, PWM:50
23:33:18.089 -> Water_Temp.:105.69, PWM:50
23:33:18.886 -> Water_Temp.:105.69, PWM:50
23:33:19.680 -> Water_Temp.:105.80, PWM:49
23:33:20.441 -> Water_Temp.:105.69, PWM:50
23:33:21.237 -> Water_Temp.:105.69, PWM:50
23:33:21.990 -> Water_Temp.:105.69, PWM:50
23:33:22.785 -> Water_Temp.:105.69, PWM:50
23:33:23.577 -> Water_Temp.:105.69, PWM:50
23:33:24.334 -> Water_Temp.:105.69, PWM:50
23:33:25.132 -> Water_Temp.:105.69, PWM:49
23:33:25.891 -> Water_Temp.:105.69, PWM:50
23:33:26.685 -> Water_Temp.:105.69, PWM:50
23:33:27.477 -> Water_Temp.:105.69, PWM:50
23:33:28.233 -> Water_Temp.:105.69, PWM:51
23:33:29.030 -> Water_Temp.:105.69, PWM:53
23:33:29.790 -> Water_Temp.:105.69, PWM:53
23:33:30.583 -> Water_Temp.:105.69, PWM:53
23:33:31.377 -> Water_Temp.:105.69, PWM:53
23:33:32.134 -> Water_Temp.:105.80, PWM:53
23:33:32.930 -> Water_Temp.:105.69, PWM:53
23:33:33.693 -> Water_Temp.:105.80, PWM:53
23:33:34.487 -> Water_Temp.:105.80, PWM:53
23:33:35.248 -> Water_Temp.:105.80, PWM:53
23:33:36.044 -> Water_Temp.:105.80, PWM:53
23:33:36.806 -> Water_Temp.:105.80, PWM:53
23:33:37.598 -> Water_Temp.:105.80, PWM:53
23:33:38.395 -> Water_Temp.:105.80, PWM:53
23:33:39.918 -> Water_Temp.:105.80, PWM:53
23:33:40.709 -> Water_Temp.:105.80, PWM:53
23:33:41.506 -> Water_Temp.:105.80, PWM:53
23:33:42.263 -> Water_Temp.:105.80, PWM:53
23:33:43.054 -> Water_Temp.:105.80, PWM:53
23:33:43.813 -> Water_Temp.:105.80, PWM:53
23:33:44.608 -> Water_Temp.:105.80, PWM:53
23:33:45.405 -> Water_Temp.:105.80, PWM:53
23:33:46.166 -> Water_Temp.:105.80, PWM:53
23:33:46.959 -> Water_Temp.:105.80, PWM:53
23:33:47.719 -> Water_Temp.:105.80, PWM:53
23:33:48.514 -> Water_Temp.:105.80, PWM:53
23:33:49.275 -> Water_Temp.:105.80, PWM:53
23:33:50.829 -> Water_Temp.:105.80, PWM:53
23:33:51.628 -> Water_Temp.:105.80, PWM:53
23:33:52.389 -> Water_Temp.:105.80, PWM:53
23:33:53.183 -> Water_Temp.:105.80, PWM:53
23:33:53.945 -> Water_Temp.:105.80, PWM:53
23:33:54.741 -> Water_Temp.:105.80, PWM:53
23:33:55.505 -> Water_Temp.:105.91, PWM:53
23:33:56.302 -> Water_Temp.:105.80, PWM:53
23:33:57.092 -> Water_Temp.:105.80, PWM:53
23:33:57.849 -> Water_Temp.:105.80, PWM:53
23:33:58.644 -> Water_Temp.:105.91, PWM:53
23:33:59.406 -> Water_Temp.:105.91, PWM:53
23:34:00.206 -> Water_Temp.:105.80, PWM:53
23:34:00.970 -> Water_Temp.:105.80, PWM:53
23:34:01.764 -> Water_Temp.:105.80, PWM:53
23:34:02.553 -> Water_Temp.:105.80, PWM:53
23:34:03.309 -> Water_Temp.:105.80, PWM:53
23:34:04.104 -> Water_Temp.:105.80, PWM:53
23:34:04.866 -> Water_Temp.:105.80, PWM:53
23:34:05.664 -> Water_Temp.:105.80, PWM:53
23:34:06.422 -> Water_Temp.:105.80, PWM:53
23:34:07.219 -> Water_Temp.:105.80, PWM:53
23:34:08.013 -> Water_Temp.:105.80, PWM:53
23:34:08.779 -> Water_Temp.:105.80, PWM:53
23:34:09.576 -> Water_Temp.:105.80, PWM:53
23:34:10.336 -> Water_Temp.:105.80, PWM:53
23:34:11.133 -> Water_Temp.:105.80, PWM:53
23:34:11.892 -> Water_Temp.:105.80, PWM:53
23:34:12.684 -> Water_Temp.:105.80, PWM:53
23:34:13.444 -> Water_Temp.:105.80, PWM:53
23:34:14.240 -> Water_Temp.:105.80, PWM:53
23:34:15.002 -> Water_Temp.:105.80, PWM:53
23:34:16.592 -> Water_Temp.:105.80, PWM:53
23:34:17.349 -> Water_Temp.:105.80, PWM:53
23:34:18.142 -> Water_Temp.:105.80, PWM:53
23:34:18.903 -> Water_Temp.:105.80, PWM:53
23:34:19.696 -> Water_Temp.:105.80, PWM:53
23:34:20.487 -> Water_Temp.:105.80, PWM:53
23:34:21.247 -> Water_Temp.:105.80, PWM:53
23:34:22.040 -> Water_Temp.:105.80, PWM:53
23:34:22.805 -> Water_Temp.:105.80, PWM:53
23:34:23.598 -> Water_Temp.:105.80, PWM:53
23:34:24.389 -> Water_Temp.:105.80, PWM:53
23:34:25.154 -> Water_Temp.:105.80, PWM:53
23:34:25.949 -> Water_Temp.:105.80, PWM:53
23:34:26.711 -> Water_Temp.:105.91, PWM:53
23:34:27.507 -> Water_Temp.:105.91, PWM:53
23:34:28.263 -> Water_Temp.:105.91, PWM:53
23:34:29.057 -> Water_Temp.:105.91, PWM:53
23:34:29.852 -> Water_Temp.:105.91, PWM:53
23:34:30.609 -> Water_Temp.:105.91, PWM:53
23:34:31.406 -> Water_Temp.:105.91, PWM:53
23:34:32.167 -> Water_Temp.:105.91, PWM:53
23:34:32.960 -> Water_Temp.:105.91, PWM:53
23:34:33.721 -> Water_Temp.:105.91, PWM:53
23:34:34.515 -> Water_Temp.:105.91, PWM:53
23:34:35.310 -> Water_Temp.:105.91, PWM:53
23:34:36.067 -> Water_Temp.:105.91, PWM:53
23:34:36.865 -> Water_Temp.:105.91, PWM:53
23:34:37.627 -> Water_Temp.:105.91, PWM:53
23:34:38.424 -> Water_Temp.:105.91, PWM:53
23:34:39.947 -> Water_Temp.:105.91, PWM:53
23:34:40.741 -> Water_Temp.:105.91, PWM:53
23:34:41.504 -> Water_Temp.:105.91, PWM:53
23:34:42.302 -> Water_Temp.:105.91, PWM:53
23:34:43.064 -> Water_Temp.:105.91, PWM:53
23:34:43.860 -> Water_Temp.:105.91, PWM:53
23:34:44.618 -> Water_Temp.:105.91, PWM:53
23:34:45.414 -> Water_Temp.:105.91, PWM:53
23:34:46.175 -> Water_Temp.:106.02, PWM:53
23:34:47.727 -> Water_Temp.:106.02, PWM:53
23:34:48.527 -> Water_Temp.:106.02, PWM:53
23:34:49.285 -> Water_Temp.:106.02, PWM:53
23:34:50.077 -> Water_Temp.:106.02, PWM:53
23:34:50.839 -> Water_Temp.:106.02, PWM:53
23:34:51.638 -> Water_Temp.:106.02, PWM:53
23:34:52.400 -> Water_Temp.:106.02, PWM:53
23:34:53.958 -> Water_Temp.:106.02, PWM:53
23:34:54.717 -> Water_Temp.:106.02, PWM:53
23:34:55.512 -> Water_Temp.:106.02, PWM:53
23:34:56.304 -> Water_Temp.:106.14, PWM:53
23:34:57.857 -> Water_Temp.:106.14, PWM:53
23:34:58.617 -> Water_Temp.:106.14, PWM:53
23:34:59.407 -> Water_Temp.:106.14, PWM:53
23:35:00.198 -> Water_Temp.:106.14, PWM:53
23:35:00.963 -> Water_Temp.:106.14, PWM:53
23:35:01.759 -> Water_Temp.:106.14, PWM:53
23:35:02.520 -> Water_Temp.:106.14, PWM:53
23:35:03.314 -> Water_Temp.:106.14, PWM:53
23:35:04.074 -> Water_Temp.:106.14, PWM:53
23:35:04.870 -> Water_Temp.:106.14, PWM:53
23:35:05.665 -> Water_Temp.:106.14, PWM:53
23:35:06.420 -> Water_Temp.:106.14, PWM:53
23:35:07.213 -> Water_Temp.:106.14, PWM:53
23:35:08.004 -> Water_Temp.:106.14, PWM:53
23:35:08.763 -> Water_Temp.:106.14, PWM:53
23:35:09.560 -> Water_Temp.:106.14, PWM:53
23:35:10.321 -> Water_Temp.:106.02, PWM:53
23:35:11.111 -> Water_Temp.:106.02, PWM:53
23:35:11.905 -> Water_Temp.:106.02, PWM:53
23:35:12.666 -> Water_Temp.:106.14, PWM:53
23:35:13.460 -> Water_Temp.:106.14, PWM:53
23:35:14.987 -> Water_Temp.:106.14, PWM:53
23:35:15.786 -> Water_Temp.:106.14, PWM:53
23:35:16.547 -> Water_Temp.:106.14, PWM:53
23:35:17.339 -> Water_Temp.:106.14, PWM:53
23:35:18.135 -> Water_Temp.:106.14, PWM:53
23:35:18.897 -> Water_Temp.:106.14, PWM:53
23:35:19.689 -> Water_Temp.:106.14, PWM:53
23:35:20.451 -> Water_Temp.:106.14, PWM:53
23:35:21.250 -> Water_Temp.:106.14, PWM:53
23:35:22.010 -> Water_Temp.:106.14, PWM:53
23:35:22.801 -> Water_Temp.:106.14, PWM:53
23:35:23.562 -> Water_Temp.:106.14, PWM:53
23:35:24.354 -> Water_Temp.:106.14, PWM:53
23:35:25.152 -> Water_Temp.:106.14, PWM:53
23:35:25.914 -> Water_Temp.:106.14, PWM:53
23:35:26.709 -> Water_Temp.:106.14, PWM:53
23:35:27.466 -> Water_Temp.:106.14, PWM:53
23:35:28.260 -> Water_Temp.:106.14, PWM:53
23:35:29.053 -> Water_Temp.:106.14, PWM:53
23:35:29.815 -> Water_Temp.:106.14, PWM:53
23:35:30.608 -> Water_Temp.:106.14, PWM:53
23:35:31.366 -> Water_Temp.:106.14, PWM:53
23:35:32.159 -> Water_Temp.:106.14, PWM:53
23:35:32.922 -> Water_Temp.:106.14, PWM:53
23:35:33.718 -> Water_Temp.:106.14, PWM:53
23:35:34.481 -> Water_Temp.:106.25, PWM:53
23:35:35.279 -> Water_Temp.:106.25, PWM:53
23:35:36.071 -> Water_Temp.:106.25, PWM:53
23:35:36.831 -> Water_Temp.:106.25, PWM:53
23:35:39.141 -> Water_Temp.:106.25, PWM:53
23:35:39.903 -> Water_Temp.:106.25, PWM:53
23:35:40.700 -> Water_Temp.:106.25, PWM:53
23:35:41.495 -> Water_Temp.:106.25, PWM:53
23:35:42.253 -> Water_Temp.:106.25, PWM:53
23:35:43.043 -> Water_Temp.:106.25, PWM:53
23:35:43.839 -> Water_Temp.:106.25, PWM:53
23:35:44.597 -> Water_Temp.:106.25, PWM:53
23:35:45.395 -> Water_Temp.:106.25, PWM:53
23:35:46.156 -> Water_Temp.:106.25, PWM:53
23:35:46.947 -> Water_Temp.:106.25, PWM:53
23:35:47.739 -> Water_Temp.:106.25, PWM:53
23:35:48.500 -> Water_Temp.:106.25, PWM:53
23:35:49.297 -> Water_Temp.:106.25, PWM:53
23:35:50.058 -> Water_Temp.:106.25, PWM:53
23:35:50.849 -> Water_Temp.:106.25, PWM:53
23:35:51.608 -> Water_Temp.:106.25, PWM:53
23:35:52.405 -> Water_Temp.:106.25, PWM:53
23:35:53.956 -> Water_Temp.:106.25, PWM:53
23:35:54.716 -> Water_Temp.:106.25, PWM:53
23:35:55.510 -> Water_Temp.:106.25, PWM:53
23:35:56.273 -> Water_Temp.:106.25, PWM:53
23:35:57.066 -> Water_Temp.:106.25, PWM:53
23:35:57.827 -> Water_Temp.:106.25, PWM:53
23:35:58.621 -> Water_Temp.:106.14, PWM:53
23:35:59.414 -> Water_Temp.:106.14, PWM:53
23:36:00.174 -> Water_Temp.:106.14, PWM:53
23:36:00.970 -> Water_Temp.:106.14, PWM:53
23:36:01.731 -> Water_Temp.:106.14, PWM:53
23:36:02.523 -> Water_Temp.:106.14, PWM:53
23:36:03.314 -> Water_Temp.:106.14, PWM:53
23:36:04.073 -> Water_Temp.:106.14, PWM:53
23:36:04.871 -> Water_Temp.:106.14, PWM:53
23:36:05.632 -> Water_Temp.:106.14, PWM:53
23:36:06.424 -> Water_Temp.:106.14, PWM:53
23:36:07.188 -> Water_Temp.:106.14, PWM:53
23:36:07.981 -> Water_Temp.:106.14, PWM:53
23:36:08.774 -> Water_Temp.:106.14, PWM:53
23:36:09.532 -> Water_Temp.:106.14, PWM:53
23:36:10.325 -> Water_Temp.:106.14, PWM:53
23:36:11.088 -> Water_Temp.:106.14, PWM:53
23:36:11.885 -> Water_Temp.:106.14, PWM:53
23:36:12.647 -> Water_Temp.:106.14, PWM:53
23:36:13.440 -> Water_Temp.:106.14, PWM:53
23:36:14.235 -> Water_Temp.:106.14, PWM:53
23:36:14.999 -> Water_Temp.:106.14, PWM:53
23:36:15.798 -> Water_Temp.:106.14, PWM:53
23:36:16.558 -> Water_Temp.:106.14, PWM:53
23:36:17.353 -> Water_Temp.:106.14, PWM:53
23:36:18.113 -> Water_Temp.:106.14, PWM:53
23:36:18.904 -> Water_Temp.:106.14, PWM:53
23:36:19.696 -> Water_Temp.:106.14, PWM:53
23:36:20.459 -> Water_Temp.:106.14, PWM:53
23:36:21.254 -> Water_Temp.:106.14, PWM:53
23:36:22.011 -> Water_Temp.:106.14, PWM:53
23:36:22.811 -> Water_Temp.:106.14, PWM:53
23:36:23.568 -> Water_Temp.:106.14, PWM:53
23:36:24.359 -> Water_Temp.:106.14, PWM:53
23:36:25.125 -> Water_Temp.:106.14, PWM:53
23:36:27.477 -> Water_Temp.:106.14, PWM:53
23:36:28.240 -> Water_Temp.:106.25, PWM:53
23:36:29.036 -> Water_Temp.:106.25, PWM:53
23:36:29.795 -> Water_Temp.:106.14, PWM:53
23:36:30.590 -> Water_Temp.:106.25, PWM:53
23:36:31.353 -> Water_Temp.:106.25, PWM:53
23:36:32.149 -> Water_Temp.:106.25, PWM:53
23:36:32.940 -> Water_Temp.:106.25, PWM:53
23:36:33.697 -> Water_Temp.:106.25, PWM:53
23:36:34.495 -> Water_Temp.:106.25, PWM:53
23:36:35.257 -> Water_Temp.:106.25, PWM:53
23:36:36.051 -> Water_Temp.:106.25, PWM:53
23:36:36.844 -> Water_Temp.:106.25, PWM:53
23:36:37.607 -> Water_Temp.:106.25, PWM:53
23:36:38.401 -> Water_Temp.:106.25, PWM:53
23:36:39.159 -> Water_Temp.:106.25, PWM:53
23:36:39.955 -> Water_Temp.:106.25, PWM:53
23:36:40.715 -> Water_Temp.:106.25, PWM:53
23:36:41.509 -> Water_Temp.:106.25, PWM:53
23:36:42.269 -> Water_Temp.:106.25, PWM:53
23:36:43.060 -> Water_Temp.:106.25, PWM:53
23:36:43.856 -> Water_Temp.:106.25, PWM:53
23:36:44.618 -> Water_Temp.:106.25, PWM:53
23:36:45.417 -> Water_Temp.:106.25, PWM:53
23:36:46.178 -> Water_Temp.:106.25, PWM:53
23:36:46.975 -> Water_Temp.:106.25, PWM:53
23:36:47.734 -> Water_Temp.:106.25, PWM:53
23:36:48.531 -> Water_Temp.:106.25, PWM:53
23:36:49.324 -> Water_Temp.:106.25, PWM:53
23:36:50.845 -> Water_Temp.:106.25, PWM:53
23:36:51.637 -> Water_Temp.:106.25, PWM:53
23:36:52.393 -> Water_Temp.:106.25, PWM:53
23:36:53.186 -> Water_Temp.:106.25, PWM:53
23:36:53.982 -> Water_Temp.:106.25, PWM:53
23:36:55.514 -> Water_Temp.:106.25, PWM:53
23:36:56.308 -> Water_Temp.:106.25, PWM:53
23:36:57.103 -> Water_Temp.:106.25, PWM:53
23:36:57.862 -> Water_Temp.:106.25, PWM:53
23:36:58.653 -> Water_Temp.:106.25, PWM:56
23:36:59.414 -> Water_Temp.:106.25, PWM:56
23:37:00.212 -> Water_Temp.:106.25, PWM:56
23:37:00.976 -> Water_Temp.:106.25, PWM:56
23:37:02.530 -> Water_Temp.:106.25, PWM:56
23:37:03.286 -> Water_Temp.:106.25, PWM:56
23:37:04.081 -> Water_Temp.:106.25, PWM:56
23:37:04.875 -> Water_Temp.:106.25, PWM:56
23:37:05.638 -> Water_Temp.:106.25, PWM:56
23:37:06.431 -> Water_Temp.:106.25, PWM:56
23:37:07.193 -> Water_Temp.:106.25, PWM:56
23:37:07.989 -> Water_Temp.:106.25, PWM:56
23:37:08.745 -> Water_Temp.:106.25, PWM:56
23:37:09.536 -> Water_Temp.:106.25, PWM:56
23:37:10.331 -> Water_Temp.:106.25, PWM:56
23:37:11.098 -> Water_Temp.:106.36, PWM:56
23:37:11.895 -> Water_Temp.:106.36, PWM:56
23:37:12.654 -> Water_Temp.:106.36, PWM:56
23:37:14.208 -> Water_Temp.:106.25, PWM:56
23:37:14.971 -> Water_Temp.:106.25, PWM:56
23:37:15.768 -> Water_Temp.:106.36, PWM:56
23:37:16.527 -> Water_Temp.:106.36, PWM:56
23:37:17.322 -> Water_Temp.:106.36, PWM:56
23:37:18.117 -> Water_Temp.:106.36, PWM:56
23:37:18.877 -> Water_Temp.:106.36, PWM:56
23:37:19.670 -> Water_Temp.:106.36, PWM:56
23:37:20.432 -> Water_Temp.:106.36, PWM:56
23:37:21.233 -> Water_Temp.:106.36, PWM:56
23:37:21.993 -> Water_Temp.:106.36, PWM:56
23:37:22.789 -> Water_Temp.:106.36, PWM:56
23:37:23.548 -> Water_Temp.:106.36, PWM:56
23:37:24.341 -> Water_Temp.:106.36, PWM:56
23:37:25.106 -> Water_Temp.:106.36, PWM:56
23:37:25.895 -> Water_Temp.:106.36, PWM:56
23:37:26.692 -> Water_Temp.:106.36, PWM:56
23:37:27.452 -> Water_Temp.:106.36, PWM:56
23:37:28.247 -> Water_Temp.:106.36, PWM:56
23:37:29.003 -> Water_Temp.:106.36, PWM:56
23:37:29.797 -> Water_Temp.:106.36, PWM:56
23:37:30.593 -> Water_Temp.:106.36, PWM:56
23:37:31.357 -> Water_Temp.:106.36, PWM:56
23:37:32.157 -> Water_Temp.:106.36, PWM:56
23:37:32.916 -> Water_Temp.:106.36, PWM:56
23:37:33.703 -> Water_Temp.:106.36, PWM:56
23:37:34.494 -> Water_Temp.:106.36, PWM:56
23:37:35.255 -> Water_Temp.:106.36, PWM:56
23:37:36.048 -> Water_Temp.:106.36, PWM:56
23:37:36.805 -> Water_Temp.:106.36, PWM:56
23:37:37.602 -> Water_Temp.:106.36, PWM:56
23:37:38.393 -> Water_Temp.:106.36, PWM:56
23:37:39.156 -> Water_Temp.:106.36, PWM:56
23:37:39.948 -> Water_Temp.:106.25, PWM:56
23:37:40.711 -> Water_Temp.:106.36, PWM:56
23:37:41.501 -> Water_Temp.:106.25, PWM:56
23:37:42.294 -> Water_Temp.:106.25, PWM:56
23:37:43.053 -> Water_Temp.:106.25, PWM:56
23:37:43.848 -> Water_Temp.:106.25, PWM:56
23:37:44.609 -> Water_Temp.:106.25, PWM:56
23:37:45.408 -> Water_Temp.:106.25, PWM:56
23:37:46.169 -> Water_Temp.:106.25, PWM:56
23:37:48.485 -> Water_Temp.:106.36, PWM:56
23:37:49.279 -> Water_Temp.:106.36, PWM:56
23:37:50.835 -> Water_Temp.:106.36, PWM:56
23:37:51.599 -> Water_Temp.:106.36, PWM:56
23:37:52.387 -> Water_Temp.:106.36, PWM:56
23:37:53.183 -> Water_Temp.:106.36, PWM:56
23:37:53.943 -> Water_Temp.:106.36, PWM:56
23:37:54.738 -> Water_Temp.:106.36, PWM:56
23:37:55.500 -> Water_Temp.:106.36, PWM:56
23:37:56.293 -> Water_Temp.:106.47, PWM:56
23:37:57.089 -> Water_Temp.:106.36, PWM:56
23:37:58.609 -> Water_Temp.:106.36, PWM:56
23:37:59.405 -> Water_Temp.:106.36, PWM:56
23:38:00.164 -> Water_Temp.:106.36, PWM:56
23:38:00.962 -> Water_Temp.:106.36, PWM:56
23:38:01.727 -> Water_Temp.:106.36, PWM:56
23:38:02.517 -> Water_Temp.:106.36, PWM:56
23:38:03.277 -> Water_Temp.:106.36, PWM:56
23:38:04.067 -> Water_Temp.:106.36, PWM:56
23:38:05.621 -> Water_Temp.:106.36, PWM:56
23:38:06.413 -> Water_Temp.:106.36, PWM:56
23:38:07.176 -> Water_Temp.:106.36, PWM:56
23:38:07.972 -> Water_Temp.:106.36, PWM:56
23:38:08.761 -> Water_Temp.:106.36, PWM:56
23:38:09.522 -> Water_Temp.:106.36, PWM:56
23:38:10.315 -> Water_Temp.:106.36, PWM:56
23:38:11.080 -> Water_Temp.:106.36, PWM:56
23:38:11.871 -> Water_Temp.:106.47, PWM:56
23:38:12.663 -> Water_Temp.:106.47, PWM:0
23:38:13.425 -> Water_Temp.:106.47, PWM:0
23:38:14.218 -> Water_Temp.:106.47, PWM:0
23:38:14.983 -> Water_Temp.:106.47, PWM:0
23:38:15.780 -> Water_Temp.:106.47, PWM:0
23:38:16.536 -> Water_Temp.:106.47, PWM:0
23:38:17.326 -> Water_Temp.:106.47, PWM:0
23:38:18.121 -> Water_Temp.:106.47, PWM:0
23:38:18.882 -> Water_Temp.:106.47, PWM:0
23:38:19.675 -> Water_Temp.:106.47, PWM:0
23:38:20.437 -> Water_Temp.:106.47, PWM:0
23:38:21.238 -> Water_Temp.:106.47, PWM:0
23:38:21.995 -> Water_Temp.:106.47, PWM:0
23:38:22.790 -> Water_Temp.:106.47, PWM:0
23:38:23.580 -> Water_Temp.:106.47, PWM:0
23:38:24.337 -> Water_Temp.:106.47, PWM:0
23:38:25.133 -> Water_Temp.:106.47, PWM:0
23:38:25.894 -> Water_Temp.:106.47, PWM:0
23:38:26.686 -> Water_Temp.:106.47, PWM:0
23:38:27.478 -> Water_Temp.:106.47, PWM:0
23:38:28.236 -> Water_Temp.:106.47, PWM:0
23:38:29.029 -> Water_Temp.:106.47, PWM:0
23:38:29.821 -> Water_Temp.:106.47, PWM:0
23:38:30.582 -> Water_Temp.:106.47, PWM:0
23:38:31.376 -> Water_Temp.:106.47, PWM:0
23:38:32.137 -> Water_Temp.:106.47, PWM:0
23:38:32.930 -> Water_Temp.:106.47, PWM:0
23:38:33.693 -> Water_Temp.:106.47, PWM:0
23:38:34.483 -> Water_Temp.:106.47, PWM:0
23:38:35.280 -> Water_Temp.:106.47, PWM:0
23:38:36.039 -> Water_Temp.:106.47, PWM:0
23:38:36.831 -> Water_Temp.:106.47, PWM:0
23:38:37.622 -> Water_Temp.:106.47, PWM:0
23:38:38.384 -> Water_Temp.:106.47, PWM:0
23:38:39.184 -> Water_Temp.:106.47, PWM:0
23:38:39.944 -> Water_Temp.:106.47, PWM:0
23:38:40.734 -> Water_Temp.:106.47, PWM:0
23:38:41.494 -> Water_Temp.:106.36, PWM:0
23:38:42.293 -> Water_Temp.:106.36, PWM:0
23:38:43.055 -> Water_Temp.:106.36, PWM:0
23:38:43.851 -> Water_Temp.:106.36, PWM:0
23:38:44.639 -> Water_Temp.:106.36, PWM:0
23:38:45.404 -> Water_Temp.:106.36, PWM:0
23:38:46.201 -> Water_Temp.:106.36, PWM:0
23:38:46.962 -> Water_Temp.:106.36, PWM:0
23:38:47.755 -> Water_Temp.:106.36, PWM:0
23:38:48.516 -> Water_Temp.:106.36, PWM:0
23:38:49.309 -> Water_Temp.:106.36, PWM:0
23:38:50.101 -> Water_Temp.:106.36, PWM:0
23:38:50.860 -> Water_Temp.:106.36, PWM:0
23:38:51.660 -> Water_Temp.:106.36, PWM:0
23:38:52.418 -> Water_Temp.:106.25, PWM:0
23:38:53.212 -> Water_Temp.:106.25, PWM:0
23:38:54.001 -> Water_Temp.:106.25, PWM:0
23:38:54.761 -> Water_Temp.:106.25, PWM:0
23:38:55.559 -> Water_Temp.:106.25, PWM:0
23:38:56.318 -> Water_Temp.:106.25, PWM:0
23:38:57.114 -> Water_Temp.:106.25, PWM:0
23:38:57.870 -> Water_Temp.:106.25, PWM:0
23:38:58.664 -> Water_Temp.:106.25, PWM:0
23:38:59.458 -> Water_Temp.:106.25, PWM:0
23:39:00.216 -> Water_Temp.:106.25, PWM:0
23:39:01.005 -> Water_Temp.:106.25, PWM:0
23:39:01.796 -> Water_Temp.:106.25, PWM:0
23:39:02.556 -> Water_Temp.:106.25, PWM:0
23:39:03.352 -> Water_Temp.:106.25, PWM:0
23:39:04.112 -> Water_Temp.:106.25, PWM:0
23:39:04.908 -> Water_Temp.:106.25, PWM:0
23:39:05.671 -> Water_Temp.:106.25, PWM:0
23:39:06.464 -> Water_Temp.:106.25, PWM:0
23:39:07.255 -> Water_Temp.:106.25, PWM:0
23:39:08.014 -> Water_Temp.:106.14, PWM:0
23:39:08.807 -> Water_Temp.:106.14, PWM:0
23:39:09.598 -> Water_Temp.:106.14, PWM:0
23:39:10.363 -> Water_Temp.:106.14, PWM:0
23:39:11.160 -> Water_Temp.:106.14, PWM:0
23:39:11.921 -> Water_Temp.:106.14, PWM:0
23:39:12.713 -> Water_Temp.:106.14, PWM:0
23:39:13.471 -> Water_Temp.:106.14, PWM:0
23:39:14.266 -> Water_Temp.:106.14, PWM:0
23:39:15.027 -> Water_Temp.:106.14, PWM:0
23:39:15.825 -> Water_Temp.:106.14, PWM:0
23:39:16.619 -> Water_Temp.:106.14, PWM:0
23:39:17.376 -> Water_Temp.:106.14, PWM:0
23:39:18.164 -> Water_Temp.:106.02, PWM:0
23:39:18.955 -> Water_Temp.:106.14, PWM:0
23:39:19.719 -> Water_Temp.:106.02, PWM:0
23:39:20.503 -> Water_Temp.:106.02, PWM:0
23:39:21.291 -> Water_Temp.:106.02, PWM:0
23:39:22.047 -> Water_Temp.:106.02, PWM:0
23:39:22.843 -> Water_Temp.:106.02, PWM:0
23:39:23.629 -> Water_Temp.:106.02, PWM:0
23:39:24.422 -> Water_Temp.:106.02, PWM:0
23:39:25.180 -> Water_Temp.:106.02, PWM:0

## RAW_CAL_MLX-6
22:06:55.604 -> Average Cal sensor Temp.: 86.36 Ambient Temp.: 93.83 F
22:06:57.178 -> Average Cal sensor Temp.: 86.46 Ambient Temp.: 93.83 F
22:06:58.722 -> Average Cal sensor Temp.: 86.30 Ambient Temp.: 93.83 F
22:07:00.293 -> Average Cal sensor Temp.: 86.32 Ambient Temp.: 93.83 F
22:07:01.857 -> Average Cal sensor Temp.: 86.32 Ambient Temp.: 93.90 F
22:07:03.390 -> Average Cal sensor Temp.: 86.26 Ambient Temp.: 93.83 F
22:07:04.951 -> Average Cal sensor Temp.: 86.36 Ambient Temp.: 93.87 F
22:07:06.519 -> Average Cal sensor Temp.: 86.26 Ambient Temp.: 93.90 F
22:07:08.089 -> Average Cal sensor Temp.: 86.30 Ambient Temp.: 93.87 F
22:07:09.631 -> Average Cal sensor Temp.: 86.25 Ambient Temp.: 93.83 F
22:07:11.201 -> Average Cal sensor Temp.: 86.29 Ambient Temp.: 93.87 F
22:07:12.740 -> Average Cal sensor Temp.: 86.41 Ambient Temp.: 93.83 F
22:07:14.310 -> Average Cal sensor Temp.: 86.50 Ambient Temp.: 93.83 F
22:07:15.873 -> Average Cal sensor Temp.: 86.37 Ambient Temp.: 93.90 F
22:07:17.409 -> Average Cal sensor Temp.: 86.37 Ambient Temp.: 93.87 F
22:07:18.980 -> Average Cal sensor Temp.: 86.41 Ambient Temp.: 93.87 F
22:07:20.517 -> Average Cal sensor Temp.: 86.55 Ambient Temp.: 93.87 F
22:07:22.081 -> Average Cal sensor Temp.: 86.50 Ambient Temp.: 93.90 F
22:07:23.646 -> Average Cal sensor Temp.: 86.48 Ambient Temp.: 93.90 F
22:07:25.211 -> Average Cal sensor Temp.: 86.45 Ambient Temp.: 93.90 F
22:07:26.742 -> Average Cal sensor Temp.: 86.54 Ambient Temp.: 93.90 F
22:07:28.316 -> Average Cal sensor Temp.: 86.49 Ambient Temp.: 93.90 F
22:07:29.880 -> Average Cal sensor Temp.: 86.52 Ambient Temp.: 93.90 F
22:07:31.413 -> Average Cal sensor Temp.: 86.53 Ambient Temp.: 93.87 F
22:07:32.979 -> Average Cal sensor Temp.: 86.70 Ambient Temp.: 93.90 F
22:07:34.545 -> Average Cal sensor Temp.: 86.52 Ambient Temp.: 93.94 F
22:07:36.081 -> Average Cal sensor Temp.: 86.43 Ambient Temp.: 93.87 F
22:07:37.646 -> Average Cal sensor Temp.: 86.64 Ambient Temp.: 93.94 F
22:07:39.212 -> Average Cal sensor Temp.: 86.53 Ambient Temp.: 93.90 F
22:07:40.773 -> Average Cal sensor Temp.: 86.47 Ambient Temp.: 93.90 F
22:07:42.335 -> Average Cal sensor Temp.: 86.55 Ambient Temp.: 93.94 F
22:07:43.873 -> Average Cal sensor Temp.: 86.48 Ambient Temp.: 93.97 F
22:07:45.430 -> Average Cal sensor Temp.: 86.50 Ambient Temp.: 93.94 F
22:07:46.996 -> Average Cal sensor Temp.: 86.58 Ambient Temp.: 93.90 F
22:07:48.564 -> Average Cal sensor Temp.: 86.65 Ambient Temp.: 93.97 F
22:07:50.127 -> Average Cal sensor Temp.: 86.77 Ambient Temp.: 93.94 F
22:07:51.680 -> Average Cal sensor Temp.: 86.73 Ambient Temp.: 93.94 F
22:07:53.226 -> Average Cal sensor Temp.: 86.75 Ambient Temp.: 93.90 F
22:07:54.775 -> Average Cal sensor Temp.: 86.78 Ambient Temp.: 93.94 F
22:07:56.334 -> Average Cal sensor Temp.: 86.73 Ambient Temp.: 93.94 F
22:07:57.886 -> Average Cal sensor Temp.: 86.72 Ambient Temp.: 93.97 F
22:07:59.439 -> Average Cal sensor Temp.: 86.70 Ambient Temp.: 93.97 F
22:08:00.996 -> Average Cal sensor Temp.: 86.74 Ambient Temp.: 93.97 F
22:08:02.548 -> Average Cal sensor Temp.: 86.82 Ambient Temp.: 93.94 F
22:08:04.138 -> Average Cal sensor Temp.: 86.76 Ambient Temp.: 93.97 F
22:08:05.687 -> Average Cal sensor Temp.: 86.78 Ambient Temp.: 93.97 F
22:08:07.239 -> Average Cal sensor Temp.: 86.84 Ambient Temp.: 93.94 F
22:08:08.795 -> Average Cal sensor Temp.: 86.90 Ambient Temp.: 93.97 F
22:08:10.353 -> Average Cal sensor Temp.: 86.92 Ambient Temp.: 94.01 F
22:08:11.904 -> Average Cal sensor Temp.: 86.88 Ambient Temp.: 93.94 F
22:08:13.457 -> Average Cal sensor Temp.: 86.77 Ambient Temp.: 94.08 F
22:08:15.014 -> Average Cal sensor Temp.: 86.66 Ambient Temp.: 93.94 F
22:08:16.568 -> Average Cal sensor Temp.: 86.77 Ambient Temp.: 94.01 F
22:08:18.155 -> Average Cal sensor Temp.: 86.68 Ambient Temp.: 93.94 F
22:08:19.708 -> Average Cal sensor Temp.: 86.61 Ambient Temp.: 93.94 F
22:08:21.267 -> Average Cal sensor Temp.: 86.76 Ambient Temp.: 93.97 F
22:08:22.818 -> Average Cal sensor Temp.: 86.65 Ambient Temp.: 93.97 F
22:08:24.370 -> Average Cal sensor Temp.: 86.77 Ambient Temp.: 93.97 F
22:08:25.927 -> Average Cal sensor Temp.: 86.72 Ambient Temp.: 93.97 F
22:08:27.477 -> Average Cal sensor Temp.: 86.75 Ambient Temp.: 93.97 F
22:08:29.024 -> Average Cal sensor Temp.: 86.79 Ambient Temp.: 93.97 F
22:08:30.584 -> Average Cal sensor Temp.: 86.77 Ambient Temp.: 94.08 F
22:08:32.132 -> Average Cal sensor Temp.: 86.87 Ambient Temp.: 93.97 F
22:08:33.718 -> Average Cal sensor Temp.: 86.90 Ambient Temp.: 94.01 F
22:08:35.270 -> Average Cal sensor Temp.: 87.02 Ambient Temp.: 94.01 F
22:08:36.826 -> Average Cal sensor Temp.: 87.12 Ambient Temp.: 94.01 F
22:08:38.381 -> Average Cal sensor Temp.: 87.05 Ambient Temp.: 94.01 F
22:08:39.931 -> Average Cal sensor Temp.: 87.09 Ambient Temp.: 94.08 F
22:08:41.490 -> Average Cal sensor Temp.: 87.08 Ambient Temp.: 94.08 F
22:08:43.045 -> Average Cal sensor Temp.: 87.15 Ambient Temp.: 94.01 F
22:08:44.599 -> Average Cal sensor Temp.: 87.08 Ambient Temp.: 94.01 F
22:08:46.155 -> Average Cal sensor Temp.: 87.17 Ambient Temp.: 94.01 F
22:08:47.711 -> Average Cal sensor Temp.: 87.12 Ambient Temp.: 94.08 F
22:08:49.262 -> Average Cal sensor Temp.: 87.20 Ambient Temp.: 94.08 F
22:08:50.812 -> Average Cal sensor Temp.: 87.11 Ambient Temp.: 94.08 F
22:08:52.367 -> Average Cal sensor Temp.: 86.98 Ambient Temp.: 94.12 F
22:08:53.922 -> Average Cal sensor Temp.: 86.93 Ambient Temp.: 94.08 F
22:08:55.509 -> Average Cal sensor Temp.: 86.97 Ambient Temp.: 94.08 F
22:08:57.063 -> Average Cal sensor Temp.: 87.10 Ambient Temp.: 94.12 F
22:08:58.617 -> Average Cal sensor Temp.: 87.15 Ambient Temp.: 94.08 F
22:09:00.168 -> Average Cal sensor Temp.: 87.06 Ambient Temp.: 94.15 F
22:09:01.725 -> Average Cal sensor Temp.: 87.17 Ambient Temp.: 94.08 F
22:09:03.279 -> Average Cal sensor Temp.: 87.28 Ambient Temp.: 94.08 F
22:09:04.837 -> Average Cal sensor Temp.: 87.30 Ambient Temp.: 94.12 F
22:09:06.394 -> Average Cal sensor Temp.: 87.17 Ambient Temp.: 94.12 F
22:09:07.954 -> Average Cal sensor Temp.: 87.30 Ambient Temp.: 94.15 F
22:09:09.505 -> Average Cal sensor Temp.: 87.30 Ambient Temp.: 94.08 F
22:09:11.054 -> Average Cal sensor Temp.: 87.36 Ambient Temp.: 94.12 F
22:09:12.607 -> Average Cal sensor Temp.: 87.43 Ambient Temp.: 94.19 F
22:09:14.157 -> Average Cal sensor Temp.: 87.39 Ambient Temp.: 94.12 F
22:09:15.712 -> Average Cal sensor Temp.: 87.31 Ambient Temp.: 94.12 F
22:09:17.266 -> Average Cal sensor Temp.: 87.25 Ambient Temp.: 94.15 F
22:09:18.820 -> Average Cal sensor Temp.: 87.32 Ambient Temp.: 94.19 F
22:09:20.373 -> Average Cal sensor Temp.: 87.25 Ambient Temp.: 94.12 F
22:09:21.925 -> Average Cal sensor Temp.: 87.22 Ambient Temp.: 94.15 F
22:09:23.480 -> Average Cal sensor Temp.: 87.35 Ambient Temp.: 94.19 F
22:09:25.036 -> Average Cal sensor Temp.: 87.16 Ambient Temp.: 94.12 F
22:09:26.592 -> Average Cal sensor Temp.: 87.35 Ambient Temp.: 94.19 F
22:09:28.142 -> Average Cal sensor Temp.: 87.24 Ambient Temp.: 94.15 F
22:09:29.728 -> Average Cal sensor Temp.: 87.28 Ambient Temp.: 94.12 F
22:09:31.280 -> Average Cal sensor Temp.: 87.34 Ambient Temp.: 94.19 F
22:09:32.833 -> Average Cal sensor Temp.: 87.37 Ambient Temp.: 94.15 F
22:09:34.391 -> Average Cal sensor Temp.: 87.43 Ambient Temp.: 94.23 F
22:09:35.946 -> Average Cal sensor Temp.: 87.48 Ambient Temp.: 94.19 F
22:09:37.505 -> Average Cal sensor Temp.: 87.43 Ambient Temp.: 94.26 F
22:09:39.059 -> Average Cal sensor Temp.: 87.45 Ambient Temp.: 94.15 F
22:09:40.614 -> Average Cal sensor Temp.: 87.31 Ambient Temp.: 94.23 F
22:09:42.158 -> Average Cal sensor Temp.: 87.35 Ambient Temp.: 94.19 F
22:09:43.712 -> Average Cal sensor Temp.: 87.52 Ambient Temp.: 94.23 F
22:09:45.272 -> Average Cal sensor Temp.: 87.53 Ambient Temp.: 94.19 F
22:09:46.823 -> Average Cal sensor Temp.: 87.51 Ambient Temp.: 94.26 F
22:09:48.381 -> Average Cal sensor Temp.: 87.48 Ambient Temp.: 94.19 F
22:09:49.932 -> Average Cal sensor Temp.: 87.43 Ambient Temp.: 94.15 F
22:09:51.484 -> Average Cal sensor Temp.: 87.44 Ambient Temp.: 94.23 F
22:09:53.042 -> Average Cal sensor Temp.: 87.37 Ambient Temp.: 94.23 F
22:09:54.595 -> Average Cal sensor Temp.: 87.41 Ambient Temp.: 94.26 F
22:09:56.150 -> Average Cal sensor Temp.: 87.42 Ambient Temp.: 94.19 F
22:09:57.700 -> Average Cal sensor Temp.: 87.50 Ambient Temp.: 94.23 F
22:09:59.284 -> Average Cal sensor Temp.: 87.61 Ambient Temp.: 94.19 F
22:10:00.836 -> Average Cal sensor Temp.: 87.66 Ambient Temp.: 94.19 F
22:10:02.389 -> Average Cal sensor Temp.: 87.46 Ambient Temp.: 94.19 F
22:10:03.942 -> Average Cal sensor Temp.: 87.57 Ambient Temp.: 94.23 F
22:10:05.503 -> Average Cal sensor Temp.: 87.62 Ambient Temp.: 94.23 F
22:10:07.053 -> Average Cal sensor Temp.: 87.48 Ambient Temp.: 94.19 F
22:10:08.610 -> Average Cal sensor Temp.: 87.52 Ambient Temp.: 94.23 F
22:10:10.161 -> Average Cal sensor Temp.: 87.62 Ambient Temp.: 94.19 F
22:10:11.710 -> Average Cal sensor Temp.: 87.61 Ambient Temp.: 94.23 F
22:10:13.268 -> Average Cal sensor Temp.: 87.57 Ambient Temp.: 94.23 F
22:10:14.821 -> Average Cal sensor Temp.: 87.66 Ambient Temp.: 94.26 F
22:10:16.371 -> Average Cal sensor Temp.: 87.70 Ambient Temp.: 94.26 F
22:10:17.929 -> Average Cal sensor Temp.: 87.76 Ambient Temp.: 94.26 F
22:10:19.482 -> Average Cal sensor Temp.: 87.77 Ambient Temp.: 94.26 F
22:10:21.035 -> Average Cal sensor Temp.: 87.75 Ambient Temp.: 94.26 F
22:10:22.617 -> Average Cal sensor Temp.: 87.69 Ambient Temp.: 94.26 F
22:10:24.167 -> Average Cal sensor Temp.: 87.80 Ambient Temp.: 94.30 F
22:10:25.720 -> Average Cal sensor Temp.: 87.75 Ambient Temp.: 94.26 F
22:10:27.273 -> Average Cal sensor Temp.: 87.78 Ambient Temp.: 94.30 F
22:10:28.819 -> Average Cal sensor Temp.: 87.76 Ambient Temp.: 94.26 F
22:10:30.376 -> Average Cal sensor Temp.: 87.84 Ambient Temp.: 94.26 F
22:10:31.933 -> Average Cal sensor Temp.: 87.84 Ambient Temp.: 94.30 F
22:10:33.520 -> Average Cal sensor Temp.: 87.77 Ambient Temp.: 94.26 F
22:10:35.075 -> Average Cal sensor Temp.: 87.84 Ambient Temp.: 94.26 F
22:10:36.627 -> Average Cal sensor Temp.: 87.95 Ambient Temp.: 94.30 F
22:10:38.182 -> Average Cal sensor Temp.: 87.90 Ambient Temp.: 94.33 F
22:10:39.735 -> Average Cal sensor Temp.: 87.92 Ambient Temp.: 94.33 F
22:10:41.291 -> Average Cal sensor Temp.: 87.99 Ambient Temp.: 94.26 F
22:10:42.836 -> Average Cal sensor Temp.: 87.93 Ambient Temp.: 94.30 F
22:10:44.387 -> Average Cal sensor Temp.: 87.85 Ambient Temp.: 94.37 F
22:10:45.943 -> Average Cal sensor Temp.: 87.83 Ambient Temp.: 94.30 F
22:10:47.499 -> Average Cal sensor Temp.: 87.85 Ambient Temp.: 94.33 F
22:10:49.055 -> Average Cal sensor Temp.: 87.97 Ambient Temp.: 94.30 F
22:10:50.610 -> Average Cal sensor Temp.: 87.93 Ambient Temp.: 94.33 F
22:10:52.170 -> Average Cal sensor Temp.: 87.77 Ambient Temp.: 94.37 F
22:10:53.729 -> Average Cal sensor Temp.: 87.84 Ambient Temp.: 94.37 F
22:10:55.284 -> Average Cal sensor Temp.: 87.84 Ambient Temp.: 94.33 F
22:10:56.838 -> Average Cal sensor Temp.: 87.87 Ambient Temp.: 94.33 F
22:10:58.392 -> Average Cal sensor Temp.: 88.13 Ambient Temp.: 94.37 F
22:10:59.948 -> Average Cal sensor Temp.: 87.99 Ambient Temp.: 94.33 F
22:11:01.506 -> Average Cal sensor Temp.: 87.93 Ambient Temp.: 94.33 F
22:11:03.060 -> Average Cal sensor Temp.: 87.98 Ambient Temp.: 94.37 F
22:11:04.617 -> Average Cal sensor Temp.: 88.03 Ambient Temp.: 94.33 F
22:11:06.171 -> Average Cal sensor Temp.: 88.07 Ambient Temp.: 94.37 F
22:11:07.725 -> Average Cal sensor Temp.: 88.01 Ambient Temp.: 94.41 F
22:11:09.313 -> Average Cal sensor Temp.: 88.06 Ambient Temp.: 94.37 F
22:11:10.840 -> Average Cal sensor Temp.: 88.08 Ambient Temp.: 94.41 F
22:11:12.399 -> Average Cal sensor Temp.: 88.14 Ambient Temp.: 94.37 F
22:11:13.955 -> Average Cal sensor Temp.: 88.07 Ambient Temp.: 94.37 F
22:11:15.510 -> Average Cal sensor Temp.: 88.19 Ambient Temp.: 94.37 F
22:11:17.066 -> Average Cal sensor Temp.: 88.32 Ambient Temp.: 94.41 F
22:11:18.652 -> Average Cal sensor Temp.: 88.28 Ambient Temp.: 94.37 F
22:11:20.178 -> Average Cal sensor Temp.: 88.29 Ambient Temp.: 94.41 F
22:11:21.762 -> Average Cal sensor Temp.: 88.33 Ambient Temp.: 94.37 F
22:11:23.312 -> Average Cal sensor Temp.: 88.25 Ambient Temp.: 94.41 F
22:11:24.869 -> Average Cal sensor Temp.: 88.26 Ambient Temp.: 94.41 F
22:11:26.421 -> Average Cal sensor Temp.: 88.13 Ambient Temp.: 94.48 F
22:11:27.974 -> Average Cal sensor Temp.: 88.08 Ambient Temp.: 94.48 F
22:11:29.534 -> Average Cal sensor Temp.: 88.08 Ambient Temp.: 94.37 F
22:11:31.091 -> Average Cal sensor Temp.: 88.14 Ambient Temp.: 94.41 F
22:11:32.639 -> Average Cal sensor Temp.: 88.11 Ambient Temp.: 94.41 F
22:11:34.192 -> Average Cal sensor Temp.: 88.21 Ambient Temp.: 94.48 F
22:11:35.780 -> Average Cal sensor Temp.: 88.16 Ambient Temp.: 94.51 F
22:11:37.304 -> Average Cal sensor Temp.: 88.17 Ambient Temp.: 94.51 F
22:11:38.890 -> Average Cal sensor Temp.: 88.43 Ambient Temp.: 94.48 F
22:11:40.447 -> Average Cal sensor Temp.: 88.55 Ambient Temp.: 94.41 F
22:11:41.995 -> Average Cal sensor Temp.: 88.62 Ambient Temp.: 94.51 F
22:11:43.553 -> Average Cal sensor Temp.: 88.53 Ambient Temp.: 94.55 F
22:11:45.107 -> Average Cal sensor Temp.: 88.49 Ambient Temp.: 94.51 F
22:11:46.666 -> Average Cal sensor Temp.: 88.54 Ambient Temp.: 94.51 F
22:11:48.214 -> Average Cal sensor Temp.: 88.50 Ambient Temp.: 94.51 F
22:11:49.761 -> Average Cal sensor Temp.: 88.52 Ambient Temp.: 94.51 F
22:11:51.316 -> Average Cal sensor Temp.: 88.53 Ambient Temp.: 94.55 F
22:11:52.868 -> Average Cal sensor Temp.: 88.56 Ambient Temp.: 94.51 F
22:11:54.423 -> Average Cal sensor Temp.: 88.58 Ambient Temp.: 94.51 F
22:11:55.980 -> Average Cal sensor Temp.: 88.55 Ambient Temp.: 94.55 F
22:11:57.565 -> Average Cal sensor Temp.: 88.55 Ambient Temp.: 94.51 F
22:11:59.120 -> Average Cal sensor Temp.: 88.50 Ambient Temp.: 94.55 F
22:12:00.673 -> Average Cal sensor Temp.: 88.56 Ambient Temp.: 94.55 F
22:12:02.228 -> Average Cal sensor Temp.: 88.44 Ambient Temp.: 94.48 F
22:12:03.780 -> Average Cal sensor Temp.: 88.47 Ambient Temp.: 94.55 F
22:12:05.334 -> Average Cal sensor Temp.: 88.42 Ambient Temp.: 94.55 F
22:12:06.885 -> Average Cal sensor Temp.: 88.39 Ambient Temp.: 94.55 F
22:12:08.433 -> Average Cal sensor Temp.: 88.60 Ambient Temp.: 94.55 F
22:12:10.017 -> Average Cal sensor Temp.: 88.51 Ambient Temp.: 94.55 F
22:12:11.574 -> Average Cal sensor Temp.: 88.44 Ambient Temp.: 94.55 F
22:12:13.122 -> Average Cal sensor Temp.: 88.59 Ambient Temp.: 94.55 F
22:12:14.678 -> Average Cal sensor Temp.: 88.65 Ambient Temp.: 94.55 F
22:12:16.237 -> Average Cal sensor Temp.: 88.69 Ambient Temp.: 94.59 F
22:12:17.786 -> Average Cal sensor Temp.: 88.56 Ambient Temp.: 94.62 F
22:12:19.338 -> Average Cal sensor Temp.: 88.60 Ambient Temp.: 94.62 F
22:12:20.899 -> Average Cal sensor Temp.: 88.63 Ambient Temp.: 94.55 F
22:12:22.452 -> Average Cal sensor Temp.: 88.62 Ambient Temp.: 94.62 F
22:12:24.039 -> Average Cal sensor Temp.: 88.57 Ambient Temp.: 94.55 F
22:12:25.561 -> Average Cal sensor Temp.: 88.72 Ambient Temp.: 94.59 F
22:12:27.146 -> Average Cal sensor Temp.: 88.68 Ambient Temp.: 94.59 F
22:12:28.695 -> Average Cal sensor Temp.: 88.83 Ambient Temp.: 94.59 F
22:12:30.252 -> Average Cal sensor Temp.: 88.72 Ambient Temp.: 94.62 F
22:12:31.804 -> Average Cal sensor Temp.: 88.79 Ambient Temp.: 94.59 F
22:12:33.355 -> Average Cal sensor Temp.: 88.92 Ambient Temp.: 94.62 F
22:12:34.906 -> Average Cal sensor Temp.: 88.75 Ambient Temp.: 94.59 F
22:12:36.493 -> Average Cal sensor Temp.: 89.02 Ambient Temp.: 94.62 F
22:12:38.046 -> Average Cal sensor Temp.: 89.03 Ambient Temp.: 94.59 F
22:12:39.602 -> Average Cal sensor Temp.: 88.75 Ambient Temp.: 94.55 F
22:12:41.157 -> Average Cal sensor Temp.: 88.62 Ambient Temp.: 94.66 F
22:12:42.704 -> Average Cal sensor Temp.: 88.88 Ambient Temp.: 94.66 F
22:12:44.259 -> Average Cal sensor Temp.: 88.88 Ambient Temp.: 94.62 F
22:12:45.817 -> Average Cal sensor Temp.: 88.81 Ambient Temp.: 94.62 F
22:12:47.367 -> Average Cal sensor Temp.: 88.86 Ambient Temp.: 94.62 F
22:12:48.958 -> Average Cal sensor Temp.: 88.75 Ambient Temp.: 94.62 F
22:12:50.513 -> Average Cal sensor Temp.: 89.05 Ambient Temp.: 94.62 F
22:12:52.061 -> Average Cal sensor Temp.: 89.02 Ambient Temp.: 94.62 F
22:12:53.610 -> Average Cal sensor Temp.: 88.95 Ambient Temp.: 94.69 F
22:12:55.163 -> Average Cal sensor Temp.: 89.00 Ambient Temp.: 94.66 F
22:12:56.715 -> Average Cal sensor Temp.: 88.96 Ambient Temp.: 94.62 F
22:12:58.305 -> Average Cal sensor Temp.: 89.10 Ambient Temp.: 94.62 F
22:12:59.860 -> Average Cal sensor Temp.: 89.09 Ambient Temp.: 94.69 F
22:13:01.415 -> Average Cal sensor Temp.: 89.11 Ambient Temp.: 94.66 F
22:13:02.968 -> Average Cal sensor Temp.: 89.06 Ambient Temp.: 94.69 F
22:13:04.516 -> Average Cal sensor Temp.: 89.10 Ambient Temp.: 94.69 F
22:13:06.072 -> Average Cal sensor Temp.: 89.15 Ambient Temp.: 94.69 F
22:13:07.629 -> Average Cal sensor Temp.: 89.24 Ambient Temp.: 94.69 F
22:13:09.184 -> Average Cal sensor Temp.: 89.11 Ambient Temp.: 94.66 F
22:13:10.741 -> Average Cal sensor Temp.: 89.05 Ambient Temp.: 94.69 F
22:13:12.291 -> Average Cal sensor Temp.: 89.02 Ambient Temp.: 94.73 F
22:13:13.876 -> Average Cal sensor Temp.: 89.02 Ambient Temp.: 94.69 F
22:13:15.404 -> Average Cal sensor Temp.: 89.04 Ambient Temp.: 94.77 F
22:13:16.988 -> Average Cal sensor Temp.: 88.96 Ambient Temp.: 94.69 F
22:13:18.542 -> Average Cal sensor Temp.: 88.95 Ambient Temp.: 94.73 F
22:13:20.094 -> Average Cal sensor Temp.: 89.10 Ambient Temp.: 94.69 F
22:13:21.651 -> Average Cal sensor Temp.: 88.99 Ambient Temp.: 94.69 F
22:13:23.200 -> Average Cal sensor Temp.: 88.91 Ambient Temp.: 94.73 F
22:13:24.754 -> Average Cal sensor Temp.: 88.99 Ambient Temp.: 94.69 F
22:13:26.309 -> Average Cal sensor Temp.: 89.05 Ambient Temp.: 94.69 F
22:13:27.897 -> Average Cal sensor Temp.: 89.10 Ambient Temp.: 94.73 F
22:13:29.451 -> Average Cal sensor Temp.: 89.04 Ambient Temp.: 94.77 F
22:13:31.004 -> Average Cal sensor Temp.: 89.04 Ambient Temp.: 94.73 F
22:13:32.553 -> Average Cal sensor Temp.: 89.11 Ambient Temp.: 94.73 F
22:13:34.102 -> Average Cal sensor Temp.: 89.18 Ambient Temp.: 94.66 F
22:13:35.661 -> Average Cal sensor Temp.: 89.16 Ambient Temp.: 94.73 F
22:13:37.216 -> Average Cal sensor Temp.: 89.09 Ambient Temp.: 94.69 F
22:13:38.766 -> Average Cal sensor Temp.: 89.31 Ambient Temp.: 94.77 F
22:13:40.323 -> Average Cal sensor Temp.: 89.30 Ambient Temp.: 94.73 F
22:13:41.910 -> Average Cal sensor Temp.: 89.32 Ambient Temp.: 94.77 F
22:13:43.458 -> Average Cal sensor Temp.: 89.35 Ambient Temp.: 94.77 F
22:13:45.013 -> Average Cal sensor Temp.: 89.43 Ambient Temp.: 94.77 F
22:13:46.568 -> Average Cal sensor Temp.: 89.43 Ambient Temp.: 94.69 F
22:13:48.119 -> Average Cal sensor Temp.: 89.54 Ambient Temp.: 94.73 F
22:13:49.669 -> Average Cal sensor Temp.: 89.60 Ambient Temp.: 94.73 F
22:13:51.227 -> Average Cal sensor Temp.: 89.60 Ambient Temp.: 94.73 F
22:13:52.776 -> Average Cal sensor Temp.: 89.61 Ambient Temp.: 94.80 F
22:13:54.365 -> Average Cal sensor Temp.: 89.60 Ambient Temp.: 94.77 F
22:13:55.917 -> Average Cal sensor Temp.: 89.69 Ambient Temp.: 94.73 F
22:13:57.470 -> Average Cal sensor Temp.: 89.50 Ambient Temp.: 94.80 F
22:13:59.021 -> Average Cal sensor Temp.: 89.58 Ambient Temp.: 94.77 F
22:14:00.575 -> Average Cal sensor Temp.: 89.45 Ambient Temp.: 94.80 F
22:14:02.129 -> Average Cal sensor Temp.: 89.47 Ambient Temp.: 94.80 F
22:14:03.682 -> Average Cal sensor Temp.: 89.47 Ambient Temp.: 94.87 F
22:14:05.239 -> Average Cal sensor Temp.: 89.43 Ambient Temp.: 94.80 F
22:14:06.820 -> Average Cal sensor Temp.: 89.44 Ambient Temp.: 94.80 F
22:14:08.373 -> Average Cal sensor Temp.: 89.31 Ambient Temp.: 94.87 F
22:14:09.929 -> Average Cal sensor Temp.: 89.22 Ambient Temp.: 94.80 F
22:14:11.484 -> Average Cal sensor Temp.: 89.40 Ambient Temp.: 94.77 F
22:14:13.040 -> Average Cal sensor Temp.: 89.42 Ambient Temp.: 94.80 F
22:14:14.592 -> Average Cal sensor Temp.: 89.60 Ambient Temp.: 94.77 F
22:14:16.145 -> Average Cal sensor Temp.: 89.59 Ambient Temp.: 94.87 F
22:14:17.727 -> Average Cal sensor Temp.: 89.58 Ambient Temp.: 94.80 F
22:14:19.281 -> Average Cal sensor Temp.: 89.57 Ambient Temp.: 94.80 F
22:14:20.836 -> Average Cal sensor Temp.: 89.47 Ambient Temp.: 94.87 F
22:14:22.392 -> Average Cal sensor Temp.: 89.46 Ambient Temp.: 94.87 F
22:14:23.949 -> Average Cal sensor Temp.: 89.66 Ambient Temp.: 94.87 F
22:14:25.506 -> Average Cal sensor Temp.: 89.63 Ambient Temp.: 94.87 F
22:14:27.059 -> Average Cal sensor Temp.: 89.59 Ambient Temp.: 94.80 F
22:14:28.606 -> Average Cal sensor Temp.: 89.48 Ambient Temp.: 94.87 F
22:14:30.161 -> Average Cal sensor Temp.: 89.56 Ambient Temp.: 94.91 F
22:14:31.717 -> Average Cal sensor Temp.: 89.55 Ambient Temp.: 94.87 F
22:14:33.305 -> Average Cal sensor Temp.: 89.62 Ambient Temp.: 94.87 F
22:14:34.857 -> Average Cal sensor Temp.: 89.67 Ambient Temp.: 94.80 F
22:14:36.406 -> Average Cal sensor Temp.: 89.76 Ambient Temp.: 94.80 F
22:14:37.965 -> Average Cal sensor Temp.: 89.90 Ambient Temp.: 94.87 F
22:14:39.516 -> Average Cal sensor Temp.: 89.84 Ambient Temp.: 94.87 F
22:14:41.070 -> Average Cal sensor Temp.: 89.90 Ambient Temp.: 94.87 F
22:14:42.624 -> Average Cal sensor Temp.: 89.80 Ambient Temp.: 94.87 F
22:14:44.212 -> Average Cal sensor Temp.: 89.82 Ambient Temp.: 94.87 F
22:14:45.764 -> Average Cal sensor Temp.: 89.92 Ambient Temp.: 94.80 F
22:14:47.319 -> Average Cal sensor Temp.: 89.90 Ambient Temp.: 94.87 F
22:14:48.873 -> Average Cal sensor Temp.: 89.70 Ambient Temp.: 94.80 F
22:14:50.430 -> Average Cal sensor Temp.: 89.91 Ambient Temp.: 94.87 F
22:14:51.982 -> Average Cal sensor Temp.: 90.00 Ambient Temp.: 94.87 F
22:14:53.531 -> Average Cal sensor Temp.: 90.08 Ambient Temp.: 94.80 F
22:14:55.093 -> Average Cal sensor Temp.: 89.88 Ambient Temp.: 94.80 F
22:14:56.675 -> Average Cal sensor Temp.: 89.87 Ambient Temp.: 94.87 F
22:14:58.230 -> Average Cal sensor Temp.: 89.69 Ambient Temp.: 94.87 F
22:14:59.780 -> Average Cal sensor Temp.: 89.72 Ambient Temp.: 94.87 F
22:15:01.332 -> Average Cal sensor Temp.: 89.49 Ambient Temp.: 94.80 F
22:15:02.886 -> Average Cal sensor Temp.: 89.56 Ambient Temp.: 94.87 F
22:15:04.437 -> Average Cal sensor Temp.: 89.59 Ambient Temp.: 94.91 F
22:15:05.991 -> Average Cal sensor Temp.: 89.53 Ambient Temp.: 94.87 F
22:15:07.579 -> Average Cal sensor Temp.: 89.73 Ambient Temp.: 94.87 F
22:15:09.132 -> Average Cal sensor Temp.: 89.73 Ambient Temp.: 94.87 F
22:15:10.688 -> Average Cal sensor Temp.: 89.84 Ambient Temp.: 94.80 F
22:15:12.240 -> Average Cal sensor Temp.: 89.99 Ambient Temp.: 94.77 F
22:15:13.794 -> Average Cal sensor Temp.: 90.03 Ambient Temp.: 94.87 F
22:15:15.348 -> Average Cal sensor Temp.: 89.92 Ambient Temp.: 94.91 F
22:15:16.904 -> Average Cal sensor Temp.: 89.98 Ambient Temp.: 94.87 F
22:15:18.457 -> Average Cal sensor Temp.: 89.91 Ambient Temp.: 94.80 F
22:15:20.007 -> Average Cal sensor Temp.: 89.89 Ambient Temp.: 94.87 F
22:15:21.590 -> Average Cal sensor Temp.: 90.09 Ambient Temp.: 94.87 F
22:15:23.144 -> Average Cal sensor Temp.: 90.14 Ambient Temp.: 94.87 F
22:15:24.700 -> Average Cal sensor Temp.: 90.14 Ambient Temp.: 94.91 F
22:15:26.256 -> Average Cal sensor Temp.: 90.06 Ambient Temp.: 94.91 F
22:15:27.810 -> Average Cal sensor Temp.: 90.32 Ambient Temp.: 94.87 F
22:15:29.368 -> Average Cal sensor Temp.: 90.23 Ambient Temp.: 94.87 F
22:15:30.926 -> Average Cal sensor Temp.: 90.18 Ambient Temp.: 94.80 F
22:15:32.484 -> Average Cal sensor Temp.: 90.14 Ambient Temp.: 94.80 F
22:15:34.038 -> Average Cal sensor Temp.: 90.26 Ambient Temp.: 94.91 F
22:15:35.599 -> Average Cal sensor Temp.: 90.04 Ambient Temp.: 94.91 F
22:15:37.155 -> Average Cal sensor Temp.: 90.09 Ambient Temp.: 94.91 F
22:15:38.707 -> Average Cal sensor Temp.: 90.11 Ambient Temp.: 94.91 F
22:15:40.259 -> Average Cal sensor Temp.: 90.14 Ambient Temp.: 94.91 F
22:15:41.812 -> Average Cal sensor Temp.: 90.17 Ambient Temp.: 94.87 F
22:15:43.366 -> Average Cal sensor Temp.: 90.26 Ambient Temp.: 94.91 F
22:15:44.920 -> Average Cal sensor Temp.: 90.32 Ambient Temp.: 94.87 F
22:15:46.475 -> Average Cal sensor Temp.: 90.39 Ambient Temp.: 94.91 F
22:15:48.032 -> Average Cal sensor Temp.: 90.35 Ambient Temp.: 94.95 F
22:15:49.589 -> Average Cal sensor Temp.: 90.25 Ambient Temp.: 94.87 F
22:15:51.148 -> Average Cal sensor Temp.: 90.30 Ambient Temp.: 94.91 F
22:15:52.701 -> Average Cal sensor Temp.: 90.35 Ambient Temp.: 94.91 F
22:15:54.262 -> Average Cal sensor Temp.: 90.20 Ambient Temp.: 94.87 F
22:15:55.815 -> Average Cal sensor Temp.: 90.12 Ambient Temp.: 94.91 F
22:15:57.402 -> Average Cal sensor Temp.: 90.13 Ambient Temp.: 94.91 F
22:15:58.954 -> Average Cal sensor Temp.: 90.27 Ambient Temp.: 94.91 F
22:16:00.504 -> Average Cal sensor Temp.: 90.39 Ambient Temp.: 94.95 F
22:16:02.055 -> Average Cal sensor Temp.: 90.29 Ambient Temp.: 94.91 F
22:16:03.614 -> Average Cal sensor Temp.: 90.21 Ambient Temp.: 94.91 F
22:16:05.165 -> Average Cal sensor Temp.: 90.50 Ambient Temp.: 94.98 F
22:16:06.718 -> Average Cal sensor Temp.: 90.51 Ambient Temp.: 94.95 F
22:16:08.268 -> Average Cal sensor Temp.: 90.33 Ambient Temp.: 94.98 F
22:16:09.821 -> Average Cal sensor Temp.: 90.34 Ambient Temp.: 95.02 F
22:16:11.379 -> Average Cal sensor Temp.: 90.50 Ambient Temp.: 94.95 F
22:16:12.934 -> Average Cal sensor Temp.: 90.44 Ambient Temp.: 94.98 F
22:16:14.525 -> Average Cal sensor Temp.: 90.52 Ambient Temp.: 94.95 F
22:16:16.080 -> Average Cal sensor Temp.: 90.49 Ambient Temp.: 94.95 F
22:16:17.631 -> Average Cal sensor Temp.: 90.46 Ambient Temp.: 94.98 F
22:16:19.182 -> Average Cal sensor Temp.: 90.47 Ambient Temp.: 94.95 F
22:16:20.739 -> Average Cal sensor Temp.: 90.55 Ambient Temp.: 94.98 F
22:16:22.289 -> Average Cal sensor Temp.: 90.56 Ambient Temp.: 94.98 F
22:16:23.840 -> Average Cal sensor Temp.: 90.45 Ambient Temp.: 94.95 F
22:16:25.428 -> Average Cal sensor Temp.: 90.48 Ambient Temp.: 94.98 F
22:16:26.980 -> Average Cal sensor Temp.: 90.43 Ambient Temp.: 94.95 F
22:16:28.533 -> Average Cal sensor Temp.: 90.45 Ambient Temp.: 94.98 F
22:16:30.090 -> Average Cal sensor Temp.: 90.40 Ambient Temp.: 94.98 F
22:16:31.648 -> Average Cal sensor Temp.: 90.54 Ambient Temp.: 95.02 F
22:16:33.203 -> Average Cal sensor Temp.: 90.46 Ambient Temp.: 94.98 F
22:16:34.758 -> Average Cal sensor Temp.: 90.55 Ambient Temp.: 95.02 F
22:16:36.315 -> Average Cal sensor Temp.: 90.55 Ambient Temp.: 95.02 F
22:16:37.871 -> Average Cal sensor Temp.: 90.67 Ambient Temp.: 95.05 F
22:16:39.424 -> Average Cal sensor Temp.: 90.68 Ambient Temp.: 95.02 F
22:16:40.981 -> Average Cal sensor Temp.: 90.65 Ambient Temp.: 94.98 F
22:16:42.531 -> Average Cal sensor Temp.: 90.61 Ambient Temp.: 94.98 F
22:16:44.117 -> Average Cal sensor Temp.: 90.65 Ambient Temp.: 95.05 F
22:16:45.642 -> Average Cal sensor Temp.: 90.67 Ambient Temp.: 95.02 F
22:16:47.229 -> Average Cal sensor Temp.: 90.58 Ambient Temp.: 95.05 F
22:16:48.754 -> Average Cal sensor Temp.: 90.65 Ambient Temp.: 95.05 F
22:16:50.343 -> Average Cal sensor Temp.: 90.72 Ambient Temp.: 95.05 F
22:16:51.895 -> Average Cal sensor Temp.: 90.63 Ambient Temp.: 95.02 F
22:16:53.445 -> Average Cal sensor Temp.: 90.65 Ambient Temp.: 95.05 F
22:16:55.003 -> Average Cal sensor Temp.: 90.66 Ambient Temp.: 95.16 F
22:16:56.550 -> Average Cal sensor Temp.: 90.85 Ambient Temp.: 95.05 F
22:16:58.103 -> Average Cal sensor Temp.: 90.79 Ambient Temp.: 95.09 F
22:16:59.660 -> Average Cal sensor Temp.: 90.80 Ambient Temp.: 95.05 F
22:17:01.215 -> Average Cal sensor Temp.: 90.96 Ambient Temp.: 95.02 F
22:17:02.775 -> Average Cal sensor Temp.: 90.94 Ambient Temp.: 95.05 F
22:17:04.333 -> Average Cal sensor Temp.: 91.00 Ambient Temp.: 95.02 F
22:17:05.885 -> Average Cal sensor Temp.: 90.97 Ambient Temp.: 95.05 F
22:17:07.444 -> Average Cal sensor Temp.: 90.96 Ambient Temp.: 95.09 F
22:17:09.001 -> Average Cal sensor Temp.: 90.91 Ambient Temp.: 95.09 F
22:17:10.560 -> Average Cal sensor Temp.: 90.94 Ambient Temp.: 95.05 F
22:17:12.114 -> Average Cal sensor Temp.: 91.03 Ambient Temp.: 95.09 F
22:17:13.663 -> Average Cal sensor Temp.: 90.91 Ambient Temp.: 95.05 F
22:17:15.221 -> Average Cal sensor Temp.: 90.91 Ambient Temp.: 95.05 F
22:17:16.806 -> Average Cal sensor Temp.: 90.87 Ambient Temp.: 95.13 F
22:17:18.363 -> Average Cal sensor Temp.: 90.90 Ambient Temp.: 95.16 F
22:17:19.917 -> Average Cal sensor Temp.: 90.95 Ambient Temp.: 95.13 F
22:17:21.470 -> Average Cal sensor Temp.: 90.89 Ambient Temp.: 95.13 F
22:17:23.026 -> Average Cal sensor Temp.: 90.75 Ambient Temp.: 95.05 F
22:17:24.578 -> Average Cal sensor Temp.: 90.85 Ambient Temp.: 95.09 F
22:17:26.136 -> Average Cal sensor Temp.: 90.73 Ambient Temp.: 95.16 F
22:17:27.686 -> Average Cal sensor Temp.: 90.87 Ambient Temp.: 95.09 F
22:17:29.241 -> Average Cal sensor Temp.: 91.06 Ambient Temp.: 95.13 F
22:17:30.796 -> Average Cal sensor Temp.: 91.06 Ambient Temp.: 95.16 F
22:17:32.347 -> Average Cal sensor Temp.: 91.08 Ambient Temp.: 95.16 F
22:17:33.900 -> Average Cal sensor Temp.: 90.96 Ambient Temp.: 95.13 F
22:17:35.454 -> Average Cal sensor Temp.: 91.00 Ambient Temp.: 95.13 F
22:17:37.038 -> Average Cal sensor Temp.: 90.95 Ambient Temp.: 95.16 F
22:17:38.592 -> Average Cal sensor Temp.: 91.07 Ambient Temp.: 95.13 F
22:17:40.116 -> Average Cal sensor Temp.: 91.18 Ambient Temp.: 95.16 F
22:17:41.674 -> Average Cal sensor Temp.: 91.16 Ambient Temp.: 95.20 F
22:17:43.228 -> Average Cal sensor Temp.: 91.07 Ambient Temp.: 95.16 F
22:17:44.778 -> Average Cal sensor Temp.: 90.95 Ambient Temp.: 95.20 F
22:17:46.337 -> Average Cal sensor Temp.: 90.95 Ambient Temp.: 95.27 F
22:17:47.896 -> Average Cal sensor Temp.: 91.08 Ambient Temp.: 95.20 F
22:17:49.481 -> Average Cal sensor Temp.: 91.08 Ambient Temp.: 95.20 F
22:17:51.037 -> Average Cal sensor Temp.: 91.13 Ambient Temp.: 95.16 F
22:17:52.589 -> Average Cal sensor Temp.: 91.35 Ambient Temp.: 95.27 F
22:17:54.137 -> Average Cal sensor Temp.: 91.29 Ambient Temp.: 95.20 F
22:17:55.694 -> Average Cal sensor Temp.: 91.39 Ambient Temp.: 95.16 F
22:17:57.247 -> Average Cal sensor Temp.: 92.13 Ambient Temp.: 95.27 F
22:17:58.800 -> Average Cal sensor Temp.: 92.38 Ambient Temp.: 95.20 F
22:18:00.356 -> Average Cal sensor Temp.: 92.39 Ambient Temp.: 95.20 F
22:18:01.910 -> Average Cal sensor Temp.: 92.21 Ambient Temp.: 95.16 F
22:18:03.466 -> Average Cal sensor Temp.: 92.07 Ambient Temp.: 95.20 F
22:18:05.026 -> Average Cal sensor Temp.: 91.94 Ambient Temp.: 95.16 F
22:18:06.583 -> Average Cal sensor Temp.: 91.82 Ambient Temp.: 95.20 F
22:18:08.133 -> Average Cal sensor Temp.: 91.80 Ambient Temp.: 95.27 F
22:18:09.666 -> Average Cal sensor Temp.: 91.69 Ambient Temp.: 95.20 F
22:18:11.226 -> Average Cal sensor Temp.: 91.79 Ambient Temp.: 95.31 F
22:18:12.784 -> Average Cal sensor Temp.: 91.59 Ambient Temp.: 95.27 F
22:18:14.336 -> Average Cal sensor Temp.: 91.62 Ambient Temp.: 95.27 F
22:18:15.892 -> Average Cal sensor Temp.: 91.68 Ambient Temp.: 95.20 F
22:18:17.445 -> Average Cal sensor Temp.: 91.82 Ambient Temp.: 95.20 F
22:18:19.005 -> Average Cal sensor Temp.: 91.65 Ambient Temp.: 95.31 F
22:18:20.564 -> Average Cal sensor Temp.: 91.57 Ambient Temp.: 95.27 F
22:18:22.119 -> Average Cal sensor Temp.: 91.49 Ambient Temp.: 95.27 F
22:18:22.707 -> Ambient Temp.: 95.27 F
22:18:24.267 -> Average Cal sensor Temp.: 91.89 Ambient Temp.: 95.31 F
22:18:25.824 -> Average Cal sensor Temp.: 91.77 Ambient Temp.: 95.31 F
22:18:27.377 -> Average Cal sensor Temp.: 91.36 Ambient Temp.: 95.20 F
22:18:28.932 -> Average Cal sensor Temp.: 91.58 Ambient Temp.: 95.27 F
22:18:30.488 -> Average Cal sensor Temp.: 91.34 Ambient Temp.: 95.27 F
22:18:32.014 -> Average Cal sensor Temp.: 91.44 Ambient Temp.: 95.20 F
22:18:33.595 -> Average Cal sensor Temp.: 91.51 Ambient Temp.: 95.31 F
22:18:35.151 -> Average Cal sensor Temp.: 91.43 Ambient Temp.: 95.31 F
22:18:36.708 -> Average Cal sensor Temp.: 91.49 Ambient Temp.: 95.27 F
22:18:38.262 -> Average Cal sensor Temp.: 91.17 Ambient Temp.: 95.31 F
22:18:39.816 -> Average Cal sensor Temp.: 91.01 Ambient Temp.: 95.31 F
22:18:41.338 -> Average Cal sensor Temp.: 90.89 Ambient Temp.: 95.31 F
22:18:42.927 -> Average Cal sensor Temp.: 90.90 Ambient Temp.: 95.31 F
22:18:44.481 -> Average Cal sensor Temp.: 91.20 Ambient Temp.: 95.27 F
22:18:46.035 -> Average Cal sensor Temp.: 91.34 Ambient Temp.: 95.27 F
22:18:47.586 -> Average Cal sensor Temp.: 91.25 Ambient Temp.: 95.31 F
22:18:49.137 -> Average Cal sensor Temp.: 91.39 Ambient Temp.: 95.31 F
22:18:50.694 -> Average Cal sensor Temp.: 91.09 Ambient Temp.: 95.31 F
22:18:52.248 -> Average Cal sensor Temp.: 91.27 Ambient Temp.: 95.31 F
22:18:53.803 -> Average Cal sensor Temp.: 91.29 Ambient Temp.: 95.27 F
22:18:55.358 -> Average Cal sensor Temp.: 91.31 Ambient Temp.: 95.31 F
22:18:56.914 -> Average Cal sensor Temp.: 91.21 Ambient Temp.: 95.31 F
22:18:58.464 -> Average Cal sensor Temp.: 91.27 Ambient Temp.: 95.27 F
22:19:00.016 -> Average Cal sensor Temp.: 91.34 Ambient Temp.: 95.27 F
22:19:01.574 -> Average Cal sensor Temp.: 91.33 Ambient Temp.: 95.20 F
22:19:03.131 -> Average Cal sensor Temp.: 91.50 Ambient Temp.: 95.27 F
22:19:04.655 -> Average Cal sensor Temp.: 91.49 Ambient Temp.: 95.31 F
22:19:05.245 -> Ambient Temp.: 95.31 F
22:19:06.806 -> Average Cal sensor Temp.: 91.39 Ambient Temp.: 95.20 F
22:19:07.396 -> Ambient Temp.: 95.20 F
22:19:07.987 -> Ambient Temp.: 95.20 F
22:19:08.539 -> Ambient Temp.: 95.20 F
22:19:10.093 -> Average Cal sensor Temp.: 91.80 Ambient Temp.: 95.27 F
22:19:11.648 -> Average Cal sensor Temp.: 91.72 Ambient Temp.: 95.31 F
22:19:13.208 -> Average Cal sensor Temp.: 91.79 Ambient Temp.: 95.31 F
22:19:14.769 -> Average Cal sensor Temp.: 91.56 Ambient Temp.: 95.27 F
22:19:16.326 -> Average Cal sensor Temp.: 91.71 Ambient Temp.: 95.31 F
22:19:17.880 -> Average Cal sensor Temp.: 91.88 Ambient Temp.: 95.31 F
22:19:18.472 -> Ambient Temp.: 95.31 F
22:19:19.025 -> Ambient Temp.: 95.31 F
22:19:20.585 -> Average Cal sensor Temp.: 91.76 Ambient Temp.: 95.27 F
22:19:22.141 -> Average Cal sensor Temp.: 91.88 Ambient Temp.: 95.31 F
22:19:23.694 -> Average Cal sensor Temp.: 91.83 Ambient Temp.: 95.27 F
22:19:25.249 -> Average Cal sensor Temp.: 91.84 Ambient Temp.: 95.34 F
22:19:26.802 -> Average Cal sensor Temp.: 91.78 Ambient Temp.: 95.31 F
22:19:28.354 -> Average Cal sensor Temp.: 91.66 Ambient Temp.: 95.34 F
22:19:29.910 -> Average Cal sensor Temp.: 91.80 Ambient Temp.: 95.31 F
22:19:31.467 -> Average Cal sensor Temp.: 91.75 Ambient Temp.: 95.31 F
22:19:33.022 -> Average Cal sensor Temp.: 91.56 Ambient Temp.: 95.27 F
22:19:34.572 -> Average Cal sensor Temp.: 91.67 Ambient Temp.: 95.34 F
22:19:35.879 -> Ambient Temp.: 95.34 F
22:19:36.468 -> Ambient Temp.: 95.34 F
22:19:38.017 -> Average Cal sensor Temp.: 91.81 Ambient Temp.: 95.34 F
22:19:39.577 -> Average Cal sensor Temp.: 91.89 Ambient Temp.: 95.31 F
22:19:41.133 -> Average Cal sensor Temp.: 91.92 Ambient Temp.: 95.34 F
22:19:42.684 -> Average Cal sensor Temp.: 91.94 Ambient Temp.: 95.31 F
22:19:44.236 -> Average Cal sensor Temp.: 91.66 Ambient Temp.: 95.34 F
22:19:45.790 -> Average Cal sensor Temp.: 91.83 Ambient Temp.: 95.38 F
22:19:47.348 -> Average Cal sensor Temp.: 91.75 Ambient Temp.: 95.31 F
22:19:48.896 -> Average Cal sensor Temp.: 91.89 Ambient Temp.: 95.34 F
22:19:50.481 -> Average Cal sensor Temp.: 91.94 Ambient Temp.: 95.34 F
22:19:52.036 -> Average Cal sensor Temp.: 92.17 Ambient Temp.: 95.34 F
22:19:53.562 -> Average Cal sensor Temp.: 92.11 Ambient Temp.: 95.34 F
22:19:55.119 -> Average Cal sensor Temp.: 92.03 Ambient Temp.: 95.34 F
22:19:56.673 -> Average Cal sensor Temp.: 92.14 Ambient Temp.: 95.38 F
22:19:58.259 -> Average Cal sensor Temp.: 92.01 Ambient Temp.: 95.34 F
22:19:59.806 -> Average Cal sensor Temp.: 92.15 Ambient Temp.: 95.34 F
22:20:01.366 -> Average Cal sensor Temp.: 92.04 Ambient Temp.: 95.34 F
22:20:02.918 -> Average Cal sensor Temp.: 91.86 Ambient Temp.: 95.38 F
22:20:04.472 -> Average Cal sensor Temp.: 92.00 Ambient Temp.: 95.38 F
22:20:06.025 -> Average Cal sensor Temp.: 92.04 Ambient Temp.: 95.41 F
22:20:07.570 -> Average Cal sensor Temp.: 91.93 Ambient Temp.: 95.38 F
22:20:09.124 -> Average Cal sensor Temp.: 91.99 Ambient Temp.: 95.38 F
22:20:10.683 -> Average Cal sensor Temp.: 92.04 Ambient Temp.: 95.41 F
22:20:12.231 -> Average Cal sensor Temp.: 92.00 Ambient Temp.: 95.41 F
22:20:13.814 -> Average Cal sensor Temp.: 91.87 Ambient Temp.: 95.38 F
22:20:15.375 -> Average Cal sensor Temp.: 91.99 Ambient Temp.: 95.38 F
22:20:16.897 -> Average Cal sensor Temp.: 92.00 Ambient Temp.: 95.34 F
22:20:18.486 -> Average Cal sensor Temp.: 92.06 Ambient Temp.: 95.34 F
22:20:20.042 -> Average Cal sensor Temp.: 92.04 Ambient Temp.: 95.38 F
22:20:21.599 -> Average Cal sensor Temp.: 92.16 Ambient Temp.: 95.38 F
22:20:23.151 -> Average Cal sensor Temp.: 91.97 Ambient Temp.: 95.38 F
22:20:24.709 -> Average Cal sensor Temp.: 92.17 Ambient Temp.: 95.38 F
22:20:26.257 -> Average Cal sensor Temp.: 92.36 Ambient Temp.: 95.34 F
22:20:27.816 -> Average Cal sensor Temp.: 91.94 Ambient Temp.: 95.34 F
22:20:29.374 -> Average Cal sensor Temp.: 92.10 Ambient Temp.: 95.41 F
22:20:30.933 -> Average Cal sensor Temp.: 92.12 Ambient Temp.: 95.38 F
22:20:32.455 -> Average Cal sensor Temp.: 92.19 Ambient Temp.: 95.38 F
22:20:34.011 -> Average Cal sensor Temp.: 92.30 Ambient Temp.: 95.38 F
22:20:35.599 -> Average Cal sensor Temp.: 92.24 Ambient Temp.: 95.45 F
22:20:37.150 -> Average Cal sensor Temp.: 92.04 Ambient Temp.: 95.41 F
22:20:38.706 -> Average Cal sensor Temp.: 92.03 Ambient Temp.: 95.41 F
22:20:40.262 -> Average Cal sensor Temp.: 92.31 Ambient Temp.: 95.38 F
22:20:41.817 -> Average Cal sensor Temp.: 92.53 Ambient Temp.: 95.41 F
22:20:43.371 -> Average Cal sensor Temp.: 92.21 Ambient Temp.: 95.45 F
22:20:44.927 -> Average Cal sensor Temp.: 92.16 Ambient Temp.: 95.45 F
22:20:46.481 -> Average Cal sensor Temp.: 92.23 Ambient Temp.: 95.41 F
22:20:48.035 -> Average Cal sensor Temp.: 92.30 Ambient Temp.: 95.41 F
22:20:49.585 -> Average Cal sensor Temp.: 92.26 Ambient Temp.: 95.41 F
22:20:51.142 -> Average Cal sensor Temp.: 92.11 Ambient Temp.: 95.41 F
22:20:52.701 -> Average Cal sensor Temp.: 92.20 Ambient Temp.: 95.41 F
22:20:54.252 -> Average Cal sensor Temp.: 92.33 Ambient Temp.: 95.45 F
22:20:55.807 -> Average Cal sensor Temp.: 92.12 Ambient Temp.: 95.41 F
22:20:57.359 -> Average Cal sensor Temp.: 92.06 Ambient Temp.: 95.45 F
22:20:58.909 -> Average Cal sensor Temp.: 92.17 Ambient Temp.: 95.41 F
22:21:00.463 -> Average Cal sensor Temp.: 92.03 Ambient Temp.: 95.38 F
22:21:02.016 -> Average Cal sensor Temp.: 92.17 Ambient Temp.: 95.41 F
22:21:03.565 -> Average Cal sensor Temp.: 92.55 Ambient Temp.: 95.41 F
22:21:05.120 -> Average Cal sensor Temp.: 92.65 Ambient Temp.: 95.41 F
22:21:06.679 -> Average Cal sensor Temp.: 92.38 Ambient Temp.: 95.41 F
22:21:08.238 -> Average Cal sensor Temp.: 92.41 Ambient Temp.: 95.41 F
22:21:09.788 -> Average Cal sensor Temp.: 92.02 Ambient Temp.: 95.41 F
22:21:11.346 -> Average Cal sensor Temp.: 92.12 Ambient Temp.: 95.41 F
22:21:12.903 -> Average Cal sensor Temp.: 92.22 Ambient Temp.: 95.45 F
22:21:14.454 -> Average Cal sensor Temp.: 92.41 Ambient Temp.: 95.41 F
22:21:16.011 -> Average Cal sensor Temp.: 92.12 Ambient Temp.: 95.45 F
22:21:17.565 -> Average Cal sensor Temp.: 92.41 Ambient Temp.: 95.41 F
22:21:19.121 -> Average Cal sensor Temp.: 92.20 Ambient Temp.: 95.41 F
22:21:20.670 -> Average Cal sensor Temp.: 92.48 Ambient Temp.: 95.41 F
22:21:22.222 -> Average Cal sensor Temp.: 92.70 Ambient Temp.: 95.45 F
22:21:23.772 -> Average Cal sensor Temp.: 92.93 Ambient Temp.: 95.49 F
22:21:25.331 -> Average Cal sensor Temp.: 92.80 Ambient Temp.: 95.45 F
22:21:26.915 -> Average Cal sensor Temp.: 92.94 Ambient Temp.: 95.41 F
22:21:28.469 -> Average Cal sensor Temp.: 92.62 Ambient Temp.: 95.45 F
22:21:30.027 -> Average Cal sensor Temp.: 92.24 Ambient Temp.: 95.45 F
22:21:31.583 -> Average Cal sensor Temp.: 92.52 Ambient Temp.: 95.49 F
22:21:33.132 -> Average Cal sensor Temp.: 92.72 Ambient Temp.: 95.45 F
22:21:34.686 -> Average Cal sensor Temp.: 92.87 Ambient Temp.: 95.45 F
22:21:36.244 -> Average Cal sensor Temp.: 92.88 Ambient Temp.: 95.49 F
22:21:37.797 -> Average Cal sensor Temp.: 92.97 Ambient Temp.: 95.49 F
22:21:39.356 -> Average Cal sensor Temp.: 92.80 Ambient Temp.: 95.41 F
22:21:40.907 -> Average Cal sensor Temp.: 92.89 Ambient Temp.: 95.49 F
22:21:42.460 -> Average Cal sensor Temp.: 92.90 Ambient Temp.: 95.49 F
22:21:44.009 -> Average Cal sensor Temp.: 92.91 Ambient Temp.: 95.45 F
22:21:45.569 -> Average Cal sensor Temp.: 92.99 Ambient Temp.: 95.52 F
22:21:47.122 -> Average Cal sensor Temp.: 92.97 Ambient Temp.: 95.45 F
22:21:48.677 -> Average Cal sensor Temp.: 92.99 Ambient Temp.: 95.45 F
22:21:50.263 -> Average Cal sensor Temp.: 92.93 Ambient Temp.: 95.49 F
22:21:50.817 -> Ambient Temp.: 95.49 F
22:21:52.024 -> Ambient Temp.: 95.41 F
22:21:53.581 -> Average Cal sensor Temp.: 92.62 Ambient Temp.: 95.49 F
22:21:55.139 -> Average Cal sensor Temp.: 92.74 Ambient Temp.: 95.52 F
22:21:56.690 -> Average Cal sensor Temp.: 92.87 Ambient Temp.: 95.49 F
22:21:58.245 -> Average Cal sensor Temp.: 92.97 Ambient Temp.: 95.49 F
22:21:59.801 -> Average Cal sensor Temp.: 92.93 Ambient Temp.: 95.52 F
22:22:01.223 -> Ambient Temp.: 95.56 F
22:22:02.773 -> Average Cal sensor Temp.: 92.78 Ambient Temp.: 95.52 F
22:22:04.327 -> Average Cal sensor Temp.: 92.58 Ambient Temp.: 95.49 F
22:22:05.883 -> Average Cal sensor Temp.: 92.64 Ambient Temp.: 95.49 F
22:22:06.470 -> Ambient Temp.: 95.49 F
22:22:08.020 -> Average Cal sensor Temp.: 92.74 Ambient Temp.: 95.52 F
22:22:09.578 -> Average Cal sensor Temp.: 92.96 Ambient Temp.: 95.52 F
22:22:11.131 -> Average Cal sensor Temp.: 92.85 Ambient Temp.: 95.49 F
22:22:12.683 -> Average Cal sensor Temp.: 92.69 Ambient Temp.: 95.49 F
22:22:14.270 -> Average Cal sensor Temp.: 92.67 Ambient Temp.: 95.52 F
22:22:15.825 -> Average Cal sensor Temp.: 92.68 Ambient Temp.: 95.52 F
22:22:17.382 -> Average Cal sensor Temp.: 92.75 Ambient Temp.: 95.49 F
22:22:18.937 -> Average Cal sensor Temp.: 93.02 Ambient Temp.: 95.52 F
22:22:20.492 -> Average Cal sensor Temp.: 93.17 Ambient Temp.: 95.56 F
22:22:22.048 -> Average Cal sensor Temp.: 93.15 Ambient Temp.: 95.49 F
22:22:23.595 -> Average Cal sensor Temp.: 93.15 Ambient Temp.: 95.52 F
22:22:25.152 -> Average Cal sensor Temp.: 93.22 Ambient Temp.: 95.52 F
22:22:25.743 -> Ambient Temp.: 95.52 F
22:22:26.295 -> Ambient Temp.: 95.52 F
22:22:26.880 -> Ambient Temp.: 95.52 F
22:22:27.463 -> Ambient Temp.: 95.52 F
22:22:28.046 -> Ambient Temp.: 95.52 F
22:22:28.630 -> Ambient Temp.: 95.52 F
22:22:29.217 -> Ambient Temp.: 95.52 F
22:22:29.907 -> Ambient Temp.: 95.56 F
22:22:30.732 -> Ambient Temp.: 95.56 F
22:22:31.323 -> Ambient Temp.: 95.56 F
22:22:32.871 -> Average Cal sensor Temp.: 92.99 Ambient Temp.: 95.52 F
22:22:34.425 -> Average Cal sensor Temp.: 93.40 Ambient Temp.: 95.52 F
22:22:35.980 -> Average Cal sensor Temp.: 93.74 Ambient Temp.: 95.49 F
22:22:36.565 -> Ambient Temp.: 95.49 F
22:22:37.149 -> Ambient Temp.: 95.49 F
22:22:38.703 -> Average Cal sensor Temp.: 93.53 Ambient Temp.: 95.52 F
22:22:40.259 -> Average Cal sensor Temp.: 93.34 Ambient Temp.: 95.52 F
22:22:41.785 -> Average Cal sensor Temp.: 93.14 Ambient Temp.: 95.52 F
22:22:43.368 -> Average Cal sensor Temp.: 93.17 Ambient Temp.: 95.56 F
22:22:44.924 -> Average Cal sensor Temp.: 93.27 Ambient Temp.: 95.59 F
22:22:46.447 -> Average Cal sensor Temp.: 93.17 Ambient Temp.: 95.59 F
22:22:48.029 -> Average Cal sensor Temp.: 93.05 Ambient Temp.: 95.56 F
22:22:49.585 -> Average Cal sensor Temp.: 93.11 Ambient Temp.: 95.52 F
22:22:51.142 -> Average Cal sensor Temp.: 93.14 Ambient Temp.: 95.56 F
22:22:52.698 -> Average Cal sensor Temp.: 93.39 Ambient Temp.: 95.56 F
22:22:54.225 -> Average Cal sensor Temp.: 93.46 Ambient Temp.: 95.59 F
22:22:55.813 -> Average Cal sensor Temp.: 93.08 Ambient Temp.: 95.56 F
22:22:57.368 -> Average Cal sensor Temp.: 93.20 Ambient Temp.: 95.59 F
22:22:58.919 -> Average Cal sensor Temp.: 93.20 Ambient Temp.: 95.52 F
22:23:00.474 -> Average Cal sensor Temp.: 93.40 Ambient Temp.: 95.56 F
22:23:02.032 -> Average Cal sensor Temp.: 93.47 Ambient Temp.: 95.56 F
22:23:03.559 -> Average Cal sensor Temp.: 93.28 Ambient Temp.: 95.56 F
22:23:05.121 -> Average Cal sensor Temp.: 92.85 Ambient Temp.: 95.56 F
22:23:06.673 -> Average Cal sensor Temp.: 93.08 Ambient Temp.: 95.59 F
22:23:08.228 -> Average Cal sensor Temp.: 93.16 Ambient Temp.: 95.56 F
22:23:09.788 -> Average Cal sensor Temp.: 93.13 Ambient Temp.: 95.59 F
22:23:11.343 -> Average Cal sensor Temp.: 93.23 Ambient Temp.: 95.59 F
22:23:12.894 -> Average Cal sensor Temp.: 93.03 Ambient Temp.: 95.59 F
22:23:14.482 -> Average Cal sensor Temp.: 93.56 Ambient Temp.: 95.56 F
22:23:16.034 -> Average Cal sensor Temp.: 94.70 Ambient Temp.: 95.56 F
22:23:17.590 -> Average Cal sensor Temp.: 94.47 Ambient Temp.: 95.59 F
22:23:19.140 -> Average Cal sensor Temp.: 94.17 Ambient Temp.: 95.59 F
22:23:20.695 -> Average Cal sensor Temp.: 93.84 Ambient Temp.: 95.67 F
22:23:22.249 -> Average Cal sensor Temp.: 93.52 Ambient Temp.: 95.59 F
22:23:23.804 -> Average Cal sensor Temp.: 93.40 Ambient Temp.: 95.59 F
22:23:25.356 -> Average Cal sensor Temp.: 93.05 Ambient Temp.: 95.59 F
22:23:26.904 -> Average Cal sensor Temp.: 93.26 Ambient Temp.: 95.59 F
22:23:28.452 -> Average Cal sensor Temp.: 93.20 Ambient Temp.: 95.67 F
22:23:30.036 -> Average Cal sensor Temp.: 93.24 Ambient Temp.: 95.56 F
22:23:31.565 -> Average Cal sensor Temp.: 92.94 Ambient Temp.: 95.59 F
22:23:33.152 -> Average Cal sensor Temp.: 92.97 Ambient Temp.: 95.59 F
22:23:34.709 -> Average Cal sensor Temp.: 93.14 Ambient Temp.: 95.67 F
22:23:36.259 -> Average Cal sensor Temp.: 93.13 Ambient Temp.: 95.59 F
22:23:37.809 -> Average Cal sensor Temp.: 93.14 Ambient Temp.: 95.59 F
22:23:39.360 -> Average Cal sensor Temp.: 93.20 Ambient Temp.: 95.59 F
22:23:40.923 -> Average Cal sensor Temp.: 93.42 Ambient Temp.: 95.67 F
22:23:42.481 -> Average Cal sensor Temp.: 93.35 Ambient Temp.: 95.59 F
22:23:44.037 -> Average Cal sensor Temp.: 93.46 Ambient Temp.: 95.70 F
22:23:45.594 -> Average Cal sensor Temp.: 93.58 Ambient Temp.: 95.59 F
22:23:47.148 -> Average Cal sensor Temp.: 93.47 Ambient Temp.: 95.67 F
22:23:48.703 -> Average Cal sensor Temp.: 93.42 Ambient Temp.: 95.59 F
22:23:50.257 -> Average Cal sensor Temp.: 93.60 Ambient Temp.: 95.67 F
22:23:51.814 -> Average Cal sensor Temp.: 93.47 Ambient Temp.: 95.67 F
22:23:53.370 -> Average Cal sensor Temp.: 93.48 Ambient Temp.: 95.67 F
22:23:54.930 -> Average Cal sensor Temp.: 93.47 Ambient Temp.: 95.59 F
22:23:56.482 -> Average Cal sensor Temp.: 93.39 Ambient Temp.: 95.67 F
22:23:58.041 -> Average Cal sensor Temp.: 93.55 Ambient Temp.: 95.67 F
22:23:59.594 -> Average Cal sensor Temp.: 93.43 Ambient Temp.: 95.59 F
22:24:01.151 -> Average Cal sensor Temp.: 93.51 Ambient Temp.: 95.70 F
22:24:02.704 -> Average Cal sensor Temp.: 93.47 Ambient Temp.: 95.70 F
22:24:04.255 -> Average Cal sensor Temp.: 93.51 Ambient Temp.: 95.67 F
22:24:05.811 -> Average Cal sensor Temp.: 93.52 Ambient Temp.: 95.67 F
22:24:07.369 -> Average Cal sensor Temp.: 93.44 Ambient Temp.: 95.67 F
22:24:08.921 -> Average Cal sensor Temp.: 93.41 Ambient Temp.: 95.70 F
22:24:10.474 -> Average Cal sensor Temp.: 93.49 Ambient Temp.: 95.67 F
22:24:12.027 -> Average Cal sensor Temp.: 93.48 Ambient Temp.: 95.67 F
22:24:13.581 -> Average Cal sensor Temp.: 93.24 Ambient Temp.: 95.70 F
22:24:15.138 -> Average Cal sensor Temp.: 93.17 Ambient Temp.: 95.59 F
22:24:16.690 -> Average Cal sensor Temp.: 93.29 Ambient Temp.: 95.67 F
22:24:18.239 -> Average Cal sensor Temp.: 93.19 Ambient Temp.: 95.67 F
22:24:19.787 -> Average Cal sensor Temp.: 93.22 Ambient Temp.: 95.67 F
22:24:21.343 -> Average Cal sensor Temp.: 93.30 Ambient Temp.: 95.67 F
22:24:22.930 -> Average Cal sensor Temp.: 93.41 Ambient Temp.: 95.74 F
22:24:24.487 -> Average Cal sensor Temp.: 93.30 Ambient Temp.: 95.67 F
22:24:26.037 -> Average Cal sensor Temp.: 93.14 Ambient Temp.: 95.70 F
22:24:27.595 -> Average Cal sensor Temp.: 93.20 Ambient Temp.: 95.74 F
22:24:29.119 -> Average Cal sensor Temp.: 93.09 Ambient Temp.: 95.67 F
22:24:30.676 -> Average Cal sensor Temp.: 93.19 Ambient Temp.: 95.67 F
22:24:32.232 -> Average Cal sensor Temp.: 93.23 Ambient Temp.: 95.74 F
22:24:33.820 -> Average Cal sensor Temp.: 93.20 Ambient Temp.: 95.70 F
22:24:35.346 -> Average Cal sensor Temp.: 93.28 Ambient Temp.: 95.70 F
22:24:36.905 -> Average Cal sensor Temp.: 93.26 Ambient Temp.: 95.70 F
22:24:38.461 -> Average Cal sensor Temp.: 93.33 Ambient Temp.: 95.70 F
22:24:40.012 -> Average Cal sensor Temp.: 93.34 Ambient Temp.: 95.67 F
22:24:41.567 -> Average Cal sensor Temp.: 93.56 Ambient Temp.: 95.67 F
22:24:43.122 -> Average Cal sensor Temp.: 93.65 Ambient Temp.: 95.74 F
22:24:44.679 -> Average Cal sensor Temp.: 93.68 Ambient Temp.: 95.67 F
22:24:46.240 -> Average Cal sensor Temp.: 93.72 Ambient Temp.: 95.70 F
22:24:47.789 -> Average Cal sensor Temp.: 93.71 Ambient Temp.: 95.59 F
22:24:49.374 -> Average Cal sensor Temp.: 93.95 Ambient Temp.: 95.67 F
22:24:50.928 -> Average Cal sensor Temp.: 93.74 Ambient Temp.: 95.67 F
22:24:52.484 -> Average Cal sensor Temp.: 93.71 Ambient Temp.: 95.67 F
22:24:54.038 -> Average Cal sensor Temp.: 93.48 Ambient Temp.: 95.74 F
22:24:55.598 -> Average Cal sensor Temp.: 93.43 Ambient Temp.: 95.70 F
22:24:57.152 -> Average Cal sensor Temp.: 93.56 Ambient Temp.: 95.70 F
22:24:58.707 -> Average Cal sensor Temp.: 93.69 Ambient Temp.: 95.70 F
22:25:00.261 -> Average Cal sensor Temp.: 93.61 Ambient Temp.: 95.70 F
22:25:01.819 -> Average Cal sensor Temp.: 93.61 Ambient Temp.: 95.77 F
22:25:03.375 -> Average Cal sensor Temp.: 93.72 Ambient Temp.: 95.74 F
22:25:04.929 -> Average Cal sensor Temp.: 93.74 Ambient Temp.: 95.70 F
22:25:06.479 -> Average Cal sensor Temp.: 93.76 Ambient Temp.: 95.74 F
22:25:08.029 -> Average Cal sensor Temp.: 93.48 Ambient Temp.: 95.74 F
22:25:09.585 -> Average Cal sensor Temp.: 93.48 Ambient Temp.: 95.74 F
22:25:11.139 -> Average Cal sensor Temp.: 93.43 Ambient Temp.: 95.74 F
22:25:12.696 -> Average Cal sensor Temp.: 93.40 Ambient Temp.: 95.70 F
22:25:14.251 -> Average Cal sensor Temp.: 93.76 Ambient Temp.: 95.70 F
22:25:15.811 -> Average Cal sensor Temp.: 93.85 Ambient Temp.: 95.74 F
22:25:17.366 -> Average Cal sensor Temp.: 93.82 Ambient Temp.: 95.74 F
22:25:18.925 -> Average Cal sensor Temp.: 93.96 Ambient Temp.: 95.77 F
22:25:20.475 -> Average Cal sensor Temp.: 93.79 Ambient Temp.: 95.74 F
22:25:22.031 -> Average Cal sensor Temp.: 93.99 Ambient Temp.: 95.77 F
22:25:23.580 -> Average Cal sensor Temp.: 93.93 Ambient Temp.: 95.77 F
22:25:25.136 -> Average Cal sensor Temp.: 93.74 Ambient Temp.: 95.77 F
22:25:26.694 -> Average Cal sensor Temp.: 93.79 Ambient Temp.: 95.74 F
22:25:28.250 -> Average Cal sensor Temp.: 93.63 Ambient Temp.: 95.70 F
22:25:29.801 -> Average Cal sensor Temp.: 93.36 Ambient Temp.: 95.74 F
22:25:31.355 -> Average Cal sensor Temp.: 93.31 Ambient Temp.: 95.74 F
22:25:32.910 -> Average Cal sensor Temp.: 93.60 Ambient Temp.: 95.77 F
22:25:34.466 -> Average Cal sensor Temp.: 93.56 Ambient Temp.: 95.74 F
22:25:36.024 -> Average Cal sensor Temp.: 93.53 Ambient Temp.: 95.74 F
22:25:37.576 -> Average Cal sensor Temp.: 93.35 Ambient Temp.: 95.77 F
22:25:39.129 -> Average Cal sensor Temp.: 93.28 Ambient Temp.: 95.74 F
22:25:40.691 -> Average Cal sensor Temp.: 93.45 Ambient Temp.: 95.74 F
22:25:42.241 -> Average Cal sensor Temp.: 93.52 Ambient Temp.: 95.74 F
22:25:43.796 -> Average Cal sensor Temp.: 93.56 Ambient Temp.: 95.77 F
22:25:45.350 -> Average Cal sensor Temp.: 93.74 Ambient Temp.: 95.77 F
22:25:46.906 -> Average Cal sensor Temp.: 93.83 Ambient Temp.: 95.74 F
22:25:48.465 -> Average Cal sensor Temp.: 93.87 Ambient Temp.: 95.74 F
22:25:50.018 -> Average Cal sensor Temp.: 94.16 Ambient Temp.: 95.74 F
22:25:51.569 -> Average Cal sensor Temp.: 93.88 Ambient Temp.: 95.70 F
22:25:53.125 -> Average Cal sensor Temp.: 94.20 Ambient Temp.: 95.74 F
22:25:54.686 -> Average Cal sensor Temp.: 94.31 Ambient Temp.: 95.81 F
22:25:56.239 -> Average Cal sensor Temp.: 94.15 Ambient Temp.: 95.77 F
22:25:57.789 -> Average Cal sensor Temp.: 94.00 Ambient Temp.: 95.85 F
22:25:59.342 -> Average Cal sensor Temp.: 93.77 Ambient Temp.: 95.77 F
22:26:00.928 -> Average Cal sensor Temp.: 93.73 Ambient Temp.: 95.81 F
22:26:02.481 -> Average Cal sensor Temp.: 94.04 Ambient Temp.: 95.74 F
22:26:04.032 -> Average Cal sensor Temp.: 93.96 Ambient Temp.: 95.81 F
22:26:05.588 -> Average Cal sensor Temp.: 93.89 Ambient Temp.: 95.74 F
22:26:07.144 -> Average Cal sensor Temp.: 93.87 Ambient Temp.: 95.74 F
22:26:08.704 -> Average Cal sensor Temp.: 93.71 Ambient Temp.: 95.77 F
22:26:10.255 -> Average Cal sensor Temp.: 93.66 Ambient Temp.: 95.74 F
22:26:11.807 -> Average Cal sensor Temp.: 93.93 Ambient Temp.: 95.74 F
22:26:13.359 -> Average Cal sensor Temp.: 93.87 Ambient Temp.: 95.77 F
22:26:14.915 -> Average Cal sensor Temp.: 93.69 Ambient Temp.: 95.81 F
22:26:16.469 -> Average Cal sensor Temp.: 93.66 Ambient Temp.: 95.74 F
22:26:18.026 -> Average Cal sensor Temp.: 93.99 Ambient Temp.: 95.74 F
22:26:19.575 -> Average Cal sensor Temp.: 94.06 Ambient Temp.: 95.81 F
22:26:21.132 -> Average Cal sensor Temp.: 93.96 Ambient Temp.: 95.77 F
22:26:22.690 -> Average Cal sensor Temp.: 94.10 Ambient Temp.: 95.77 F
22:26:24.247 -> Average Cal sensor Temp.: 94.21 Ambient Temp.: 95.81 F
22:26:25.802 -> Average Cal sensor Temp.: 94.10 Ambient Temp.: 95.77 F
22:26:27.352 -> Average Cal sensor Temp.: 93.89 Ambient Temp.: 95.77 F
22:26:28.909 -> Average Cal sensor Temp.: 93.82 Ambient Temp.: 95.81 F
22:26:30.466 -> Average Cal sensor Temp.: 93.93 Ambient Temp.: 95.81 F
22:26:32.022 -> Average Cal sensor Temp.: 94.17 Ambient Temp.: 95.77 F
22:26:33.573 -> Average Cal sensor Temp.: 94.20 Ambient Temp.: 95.85 F
22:26:35.133 -> Average Cal sensor Temp.: 94.29 Ambient Temp.: 95.85 F
22:26:36.691 -> Average Cal sensor Temp.: 93.96 Ambient Temp.: 95.77 F
22:26:38.240 -> Average Cal sensor Temp.: 94.17 Ambient Temp.: 95.81 F
22:26:39.826 -> Average Cal sensor Temp.: 94.24 Ambient Temp.: 95.81 F
22:26:41.353 -> Average Cal sensor Temp.: 94.24 Ambient Temp.: 95.81 F
22:26:42.908 -> Average Cal sensor Temp.: 93.61 Ambient Temp.: 95.81 F
22:26:44.490 -> Average Cal sensor Temp.: 93.76 Ambient Temp.: 95.81 F
22:26:46.046 -> Average Cal sensor Temp.: 94.15 Ambient Temp.: 95.81 F
22:26:47.598 -> Average Cal sensor Temp.: 94.25 Ambient Temp.: 95.81 F
22:26:49.148 -> Average Cal sensor Temp.: 94.34 Ambient Temp.: 95.81 F
22:26:50.701 -> Average Cal sensor Temp.: 94.20 Ambient Temp.: 95.81 F
22:26:52.258 -> Average Cal sensor Temp.: 94.24 Ambient Temp.: 95.77 F
22:26:53.807 -> Average Cal sensor Temp.: 94.26 Ambient Temp.: 95.85 F
22:26:55.365 -> Average Cal sensor Temp.: 94.39 Ambient Temp.: 95.85 F
22:26:56.915 -> Average Cal sensor Temp.: 94.41 Ambient Temp.: 95.77 F
22:26:58.467 -> Average Cal sensor Temp.: 94.30 Ambient Temp.: 95.77 F
22:27:00.023 -> Average Cal sensor Temp.: 94.46 Ambient Temp.: 95.85 F
22:27:01.579 -> Average Cal sensor Temp.: 94.21 Ambient Temp.: 95.81 F
22:27:03.163 -> Average Cal sensor Temp.: 93.97 Ambient Temp.: 95.81 F
22:27:04.709 -> Average Cal sensor Temp.: 94.26 Ambient Temp.: 95.81 F
22:27:06.266 -> Average Cal sensor Temp.: 94.32 Ambient Temp.: 95.81 F
22:27:07.813 -> Average Cal sensor Temp.: 94.13 Ambient Temp.: 95.85 F
22:27:09.364 -> Average Cal sensor Temp.: 94.42 Ambient Temp.: 95.81 F
22:27:10.917 -> Average Cal sensor Temp.: 94.46 Ambient Temp.: 95.81 F
22:27:12.475 -> Average Cal sensor Temp.: 94.22 Ambient Temp.: 95.81 F
22:27:14.031 -> Average Cal sensor Temp.: 94.01 Ambient Temp.: 95.88 F
22:27:15.589 -> Average Cal sensor Temp.: 93.83 Ambient Temp.: 95.81 F
22:27:17.144 -> Average Cal sensor Temp.: 94.14 Ambient Temp.: 95.81 F
22:27:18.695 -> Average Cal sensor Temp.: 94.13 Ambient Temp.: 95.81 F
22:27:20.246 -> Average Cal sensor Temp.: 94.25 Ambient Temp.: 95.85 F
22:27:21.803 -> Average Cal sensor Temp.: 94.27 Ambient Temp.: 95.85 F
22:27:23.390 -> Average Cal sensor Temp.: 94.15 Ambient Temp.: 95.81 F
22:27:24.943 -> Average Cal sensor Temp.: 94.46 Ambient Temp.: 95.81 F
22:27:26.496 -> Average Cal sensor Temp.: 94.43 Ambient Temp.: 95.85 F
22:27:28.045 -> Average Cal sensor Temp.: 94.42 Ambient Temp.: 95.88 F
22:27:29.605 -> Average Cal sensor Temp.: 94.28 Ambient Temp.: 95.85 F
22:27:31.162 -> Average Cal sensor Temp.: 94.25 Ambient Temp.: 95.81 F
22:27:32.715 -> Average Cal sensor Temp.: 94.10 Ambient Temp.: 95.88 F
22:27:34.275 -> Average Cal sensor Temp.: 94.44 Ambient Temp.: 95.85 F
22:27:35.830 -> Average Cal sensor Temp.: 94.65 Ambient Temp.: 95.88 F
22:27:37.384 -> Average Cal sensor Temp.: 94.75 Ambient Temp.: 95.88 F
22:27:38.936 -> Average Cal sensor Temp.: 94.68 Ambient Temp.: 95.81 F
22:27:40.491 -> Average Cal sensor Temp.: 94.61 Ambient Temp.: 95.81 F
22:27:42.047 -> Average Cal sensor Temp.: 94.65 Ambient Temp.: 95.85 F
22:27:43.609 -> Average Cal sensor Temp.: 94.06 Ambient Temp.: 95.88 F
22:27:45.166 -> Average Cal sensor Temp.: 94.28 Ambient Temp.: 95.88 F
22:27:46.715 -> Average Cal sensor Temp.: 94.55 Ambient Temp.: 95.85 F
22:27:48.267 -> Average Cal sensor Temp.: 94.27 Ambient Temp.: 95.88 F
22:27:49.852 -> Average Cal sensor Temp.: 94.11 Ambient Temp.: 95.92 F
22:27:51.405 -> Average Cal sensor Temp.: 94.30 Ambient Temp.: 95.85 F
22:27:52.955 -> Average Cal sensor Temp.: 94.52 Ambient Temp.: 95.85 F
22:27:54.507 -> Average Cal sensor Temp.: 94.74 Ambient Temp.: 95.88 F
22:27:56.064 -> Average Cal sensor Temp.: 94.55 Ambient Temp.: 95.88 F
22:27:57.621 -> Average Cal sensor Temp.: 94.24 Ambient Temp.: 95.88 F
22:27:59.173 -> Average Cal sensor Temp.: 94.43 Ambient Temp.: 95.81 F
22:28:00.727 -> Average Cal sensor Temp.: 94.30 Ambient Temp.: 95.88 F
22:28:02.286 -> Average Cal sensor Temp.: 94.11 Ambient Temp.: 95.85 F
22:28:03.843 -> Average Cal sensor Temp.: 94.46 Ambient Temp.: 95.88 F
22:28:05.397 -> Average Cal sensor Temp.: 94.64 Ambient Temp.: 95.92 F
22:28:06.952 -> Average Cal sensor Temp.: 94.50 Ambient Temp.: 95.92 F
22:28:08.508 -> Average Cal sensor Temp.: 94.34 Ambient Temp.: 95.92 F
22:28:10.064 -> Average Cal sensor Temp.: 94.01 Ambient Temp.: 95.92 F
22:28:11.618 -> Average Cal sensor Temp.: 94.32 Ambient Temp.: 95.92 F
22:28:13.172 -> Average Cal sensor Temp.: 94.44 Ambient Temp.: 95.88 F
22:28:14.725 -> Average Cal sensor Temp.: 94.12 Ambient Temp.: 95.92 F
22:28:16.282 -> Average Cal sensor Temp.: 94.45 Ambient Temp.: 95.88 F
22:28:17.869 -> Average Cal sensor Temp.: 94.57 Ambient Temp.: 95.92 F
22:28:19.420 -> Average Cal sensor Temp.: 94.61 Ambient Temp.: 95.85 F
22:28:20.977 -> Average Cal sensor Temp.: 94.73 Ambient Temp.: 95.92 F
22:28:22.531 -> Average Cal sensor Temp.: 94.58 Ambient Temp.: 95.95 F
22:28:24.085 -> Average Cal sensor Temp.: 94.58 Ambient Temp.: 95.88 F
22:28:25.647 -> Average Cal sensor Temp.: 94.73 Ambient Temp.: 95.92 F
22:28:27.200 -> Average Cal sensor Temp.: 94.50 Ambient Temp.: 95.92 F
22:28:28.754 -> Average Cal sensor Temp.: 94.10 Ambient Temp.: 95.88 F
22:28:30.302 -> Average Cal sensor Temp.: 93.96 Ambient Temp.: 95.88 F
22:28:31.857 -> Average Cal sensor Temp.: 94.38 Ambient Temp.: 95.92 F
22:28:33.413 -> Average Cal sensor Temp.: 94.30 Ambient Temp.: 95.92 F
22:28:34.968 -> Average Cal sensor Temp.: 94.17 Ambient Temp.: 95.92 F
22:28:36.519 -> Average Cal sensor Temp.: 94.42 Ambient Temp.: 95.95 F
22:28:38.075 -> Average Cal sensor Temp.: 94.27 Ambient Temp.: 95.92 F
22:28:39.626 -> Average Cal sensor Temp.: 94.06 Ambient Temp.: 95.92 F
22:28:41.184 -> Average Cal sensor Temp.: 94.51 Ambient Temp.: 95.95 F
22:28:42.740 -> Average Cal sensor Temp.: 94.48 Ambient Temp.: 95.92 F
22:28:44.329 -> Average Cal sensor Temp.: 94.57 Ambient Temp.: 95.92 F
22:28:45.883 -> Average Cal sensor Temp.: 94.36 Ambient Temp.: 95.88 F
22:28:47.435 -> Average Cal sensor Temp.: 94.31 Ambient Temp.: 95.92 F
22:28:48.989 -> Average Cal sensor Temp.: 94.05 Ambient Temp.: 95.88 F
22:28:50.546 -> Average Cal sensor Temp.: 94.57 Ambient Temp.: 95.92 F
22:28:52.101 -> Average Cal sensor Temp.: 94.47 Ambient Temp.: 95.92 F
22:28:53.657 -> Average Cal sensor Temp.: 94.27 Ambient Temp.: 95.88 F
22:28:55.217 -> Average Cal sensor Temp.: 93.89 Ambient Temp.: 95.92 F
22:28:56.771 -> Average Cal sensor Temp.: 93.52 Ambient Temp.: 95.92 F
22:28:58.322 -> Average Cal sensor Temp.: 93.97 Ambient Temp.: 95.95 F
22:28:59.872 -> Average Cal sensor Temp.: 94.45 Ambient Temp.: 95.92 F
22:29:01.427 -> Average Cal sensor Temp.: 94.21 Ambient Temp.: 95.88 F
22:29:02.983 -> Average Cal sensor Temp.: 94.09 Ambient Temp.: 95.88 F
22:29:04.538 -> Average Cal sensor Temp.: 94.50 Ambient Temp.: 95.92 F
22:29:06.091 -> Average Cal sensor Temp.: 94.55 Ambient Temp.: 95.88 F
22:29:07.652 -> Average Cal sensor Temp.: 94.38 Ambient Temp.: 95.95 F
22:29:09.206 -> Average Cal sensor Temp.: 94.33 Ambient Temp.: 95.92 F
22:29:10.762 -> Average Cal sensor Temp.: 94.50 Ambient Temp.: 95.92 F
22:29:12.311 -> Average Cal sensor Temp.: 94.59 Ambient Temp.: 95.88 F
22:29:13.895 -> Average Cal sensor Temp.: 94.85 Ambient Temp.: 95.92 F
22:29:15.449 -> Average Cal sensor Temp.: 95.05 Ambient Temp.: 95.95 F
22:29:16.996 -> Average Cal sensor Temp.: 95.03 Ambient Temp.: 95.92 F
22:29:18.547 -> Average Cal sensor Temp.: 94.99 Ambient Temp.: 95.88 F
22:29:20.110 -> Average Cal sensor Temp.: 94.54 Ambient Temp.: 95.92 F
22:29:21.665 -> Average Cal sensor Temp.: 94.41 Ambient Temp.: 95.88 F
22:29:23.216 -> Average Cal sensor Temp.: 94.59 Ambient Temp.: 95.92 F
22:29:24.769 -> Average Cal sensor Temp.: 94.69 Ambient Temp.: 95.95 F
22:29:26.325 -> Average Cal sensor Temp.: 94.42 Ambient Temp.: 95.95 F
22:29:27.910 -> Average Cal sensor Temp.: 94.53 Ambient Temp.: 95.92 F
22:29:29.461 -> Average Cal sensor Temp.: 94.16 Ambient Temp.: 95.92 F
22:29:31.017 -> Average Cal sensor Temp.: 94.40 Ambient Temp.: 95.88 F
22:29:32.574 -> Average Cal sensor Temp.: 94.39 Ambient Temp.: 95.92 F
22:29:34.130 -> Average Cal sensor Temp.: 94.32 Ambient Temp.: 95.95 F
22:29:35.682 -> Average Cal sensor Temp.: 94.51 Ambient Temp.: 95.92 F
22:29:37.229 -> Average Cal sensor Temp.: 94.50 Ambient Temp.: 95.92 F
22:29:38.780 -> Average Cal sensor Temp.: 94.74 Ambient Temp.: 95.92 F
22:29:40.329 -> Average Cal sensor Temp.: 94.78 Ambient Temp.: 95.92 F
22:29:41.919 -> Average Cal sensor Temp.: 94.48 Ambient Temp.: 95.92 F
22:29:43.473 -> Average Cal sensor Temp.: 94.33 Ambient Temp.: 95.99 F
22:29:45.000 -> Average Cal sensor Temp.: 94.76 Ambient Temp.: 95.92 F
22:29:46.559 -> Average Cal sensor Temp.: 94.42 Ambient Temp.: 95.95 F
22:29:48.118 -> Average Cal sensor Temp.: 94.69 Ambient Temp.: 95.95 F
22:29:49.672 -> Average Cal sensor Temp.: 95.00 Ambient Temp.: 95.88 F
22:29:51.228 -> Average Cal sensor Temp.: 94.91 Ambient Temp.: 95.95 F
22:29:52.815 -> Average Cal sensor Temp.: 95.08 Ambient Temp.: 95.95 F
22:29:54.369 -> Average Cal sensor Temp.: 95.19 Ambient Temp.: 95.99 F
22:29:55.924 -> Average Cal sensor Temp.: 94.81 Ambient Temp.: 95.92 F
22:29:57.469 -> Average Cal sensor Temp.: 94.91 Ambient Temp.: 95.92 F
22:29:59.024 -> Average Cal sensor Temp.: 94.85 Ambient Temp.: 95.95 F
22:30:00.578 -> Average Cal sensor Temp.: 94.73 Ambient Temp.: 95.92 F
22:30:02.134 -> Average Cal sensor Temp.: 94.79 Ambient Temp.: 95.92 F
22:30:03.689 -> Average Cal sensor Temp.: 94.87 Ambient Temp.: 95.95 F
22:30:05.241 -> Average Cal sensor Temp.: 94.73 Ambient Temp.: 95.95 F
22:30:06.796 -> Average Cal sensor Temp.: 94.67 Ambient Temp.: 95.92 F
22:30:08.348 -> Average Cal sensor Temp.: 94.46 Ambient Temp.: 95.95 F
22:30:09.936 -> Average Cal sensor Temp.: 94.58 Ambient Temp.: 95.92 F
22:30:11.492 -> Average Cal sensor Temp.: 94.71 Ambient Temp.: 95.99 F
22:30:13.046 -> Average Cal sensor Temp.: 94.95 Ambient Temp.: 95.95 F
22:30:14.599 -> Average Cal sensor Temp.: 94.39 Ambient Temp.: 95.88 F
22:30:16.162 -> Average Cal sensor Temp.: 94.64 Ambient Temp.: 95.95 F
22:30:17.718 -> Average Cal sensor Temp.: 94.60 Ambient Temp.: 95.95 F
22:30:19.271 -> Average Cal sensor Temp.: 94.61 Ambient Temp.: 95.95 F
22:30:20.829 -> Average Cal sensor Temp.: 94.84 Ambient Temp.: 95.92 F
22:30:22.379 -> Average Cal sensor Temp.: 95.00 Ambient Temp.: 95.92 F
22:30:23.935 -> Average Cal sensor Temp.: 94.71 Ambient Temp.: 95.99 F
22:30:25.486 -> Average Cal sensor Temp.: 94.73 Ambient Temp.: 95.95 F
22:30:27.036 -> Average Cal sensor Temp.: 94.73 Ambient Temp.: 95.99 F
22:30:28.583 -> Average Cal sensor Temp.: 94.95 Ambient Temp.: 95.95 F
22:30:30.140 -> Average Cal sensor Temp.: 95.01 Ambient Temp.: 95.95 F
22:30:31.699 -> Average Cal sensor Temp.: 94.59 Ambient Temp.: 95.95 F
22:30:33.254 -> Average Cal sensor Temp.: 94.84 Ambient Temp.: 95.99 F
22:30:34.813 -> Average Cal sensor Temp.: 94.96 Ambient Temp.: 95.99 F
22:30:36.367 -> Average Cal sensor Temp.: 94.48 Ambient Temp.: 95.95 F
22:30:37.956 -> Average Cal sensor Temp.: 94.85 Ambient Temp.: 95.99 F
22:30:39.512 -> Average Cal sensor Temp.: 94.98 Ambient Temp.: 95.95 F
22:30:41.071 -> Average Cal sensor Temp.: 95.10 Ambient Temp.: 95.95 F
22:30:42.595 -> Average Cal sensor Temp.: 94.68 Ambient Temp.: 95.99 F
22:30:44.184 -> Average Cal sensor Temp.: 95.16 Ambient Temp.: 95.99 F
22:30:45.709 -> Average Cal sensor Temp.: 94.97 Ambient Temp.: 95.95 F
22:30:47.295 -> Average Cal sensor Temp.: 95.02 Ambient Temp.: 95.99 F
22:30:48.849 -> Average Cal sensor Temp.: 95.27 Ambient Temp.: 95.99 F
22:30:50.406 -> Average Cal sensor Temp.: 95.13 Ambient Temp.: 96.10 F
22:30:51.964 -> Average Cal sensor Temp.: 95.22 Ambient Temp.: 95.99 F
22:30:53.514 -> Average Cal sensor Temp.: 95.25 Ambient Temp.: 95.99 F
22:30:55.069 -> Average Cal sensor Temp.: 95.42 Ambient Temp.: 95.99 F
22:30:56.628 -> Average Cal sensor Temp.: 95.13 Ambient Temp.: 95.99 F
22:30:58.181 -> Average Cal sensor Temp.: 94.90 Ambient Temp.: 96.03 F
22:30:59.736 -> Average Cal sensor Temp.: 94.86 Ambient Temp.: 96.03 F
22:31:01.293 -> Average Cal sensor Temp.: 94.96 Ambient Temp.: 95.95 F
22:31:02.847 -> Average Cal sensor Temp.: 94.91 Ambient Temp.: 96.03 F
22:31:04.403 -> Average Cal sensor Temp.: 94.72 Ambient Temp.: 95.99 F
22:31:05.960 -> Average Cal sensor Temp.: 94.88 Ambient Temp.: 95.99 F
22:31:07.519 -> Average Cal sensor Temp.: 95.10 Ambient Temp.: 96.03 F
22:31:09.082 -> Average Cal sensor Temp.: 95.08 Ambient Temp.: 96.03 F
22:31:10.635 -> Average Cal sensor Temp.: 94.96 Ambient Temp.: 96.03 F
22:31:12.192 -> Average Cal sensor Temp.: 94.91 Ambient Temp.: 95.99 F
22:31:13.745 -> Average Cal sensor Temp.: 95.09 Ambient Temp.: 96.03 F
22:31:15.302 -> Average Cal sensor Temp.: 95.26 Ambient Temp.: 95.99 F
22:31:16.858 -> Average Cal sensor Temp.: 95.43 Ambient Temp.: 95.99 F
22:31:18.409 -> Average Cal sensor Temp.: 95.49 Ambient Temp.: 96.03 F
22:31:19.964 -> Average Cal sensor Temp.: 95.37 Ambient Temp.: 95.99 F
22:31:21.523 -> Average Cal sensor Temp.: 95.21 Ambient Temp.: 96.03 F
22:31:23.081 -> Average Cal sensor Temp.: 95.24 Ambient Temp.: 95.99 F
22:31:24.634 -> Average Cal sensor Temp.: 94.87 Ambient Temp.: 96.10 F
22:31:26.196 -> Average Cal sensor Temp.: 95.24 Ambient Temp.: 95.99 F
22:31:27.753 -> Average Cal sensor Temp.: 95.26 Ambient Temp.: 96.03 F
22:31:29.310 -> Average Cal sensor Temp.: 95.40 Ambient Temp.: 96.03 F
22:31:30.863 -> Average Cal sensor Temp.: 95.30 Ambient Temp.: 96.10 F
22:31:32.419 -> Average Cal sensor Temp.: 95.30 Ambient Temp.: 96.10 F
22:31:33.977 -> Average Cal sensor Temp.: 95.17 Ambient Temp.: 96.03 F
22:31:35.532 -> Average Cal sensor Temp.: 95.23 Ambient Temp.: 96.03 F
22:31:37.084 -> Average Cal sensor Temp.: 95.17 Ambient Temp.: 96.03 F
22:31:38.636 -> Average Cal sensor Temp.: 95.54 Ambient Temp.: 96.10 F
22:31:40.193 -> Average Cal sensor Temp.: 95.50 Ambient Temp.: 96.03 F
22:31:41.751 -> Average Cal sensor Temp.: 95.26 Ambient Temp.: 96.03 F
22:31:43.305 -> Average Cal sensor Temp.: 95.27 Ambient Temp.: 96.03 F
22:31:44.861 -> Average Cal sensor Temp.: 95.34 Ambient Temp.: 96.10 F
22:31:46.417 -> Average Cal sensor Temp.: 95.10 Ambient Temp.: 96.10 F
22:31:47.977 -> Average Cal sensor Temp.: 95.27 Ambient Temp.: 96.10 F
22:31:49.534 -> Average Cal sensor Temp.: 95.40 Ambient Temp.: 96.10 F
22:31:51.090 -> Average Cal sensor Temp.: 95.49 Ambient Temp.: 96.10 F
22:31:52.637 -> Average Cal sensor Temp.: 94.88 Ambient Temp.: 96.03 F
22:31:54.194 -> Average Cal sensor Temp.: 95.27 Ambient Temp.: 96.10 F
22:31:55.749 -> Average Cal sensor Temp.: 95.06 Ambient Temp.: 96.03 F
22:31:57.307 -> Average Cal sensor Temp.: 95.11 Ambient Temp.: 96.10 F
22:31:58.862 -> Average Cal sensor Temp.: 95.34 Ambient Temp.: 96.17 F
22:32:00.413 -> Average Cal sensor Temp.: 95.19 Ambient Temp.: 96.03 F
22:32:01.968 -> Average Cal sensor Temp.: 95.16 Ambient Temp.: 95.99 F
22:32:03.525 -> Average Cal sensor Temp.: 95.34 Ambient Temp.: 96.13 F
22:32:05.081 -> Average Cal sensor Temp.: 95.37 Ambient Temp.: 96.10 F
22:32:06.632 -> Average Cal sensor Temp.: 95.17 Ambient Temp.: 96.10 F
22:32:08.214 -> Average Cal sensor Temp.: 95.18 Ambient Temp.: 96.10 F
22:32:09.769 -> Average Cal sensor Temp.: 95.23 Ambient Temp.: 96.10 F
22:32:11.324 -> Average Cal sensor Temp.: 95.21 Ambient Temp.: 96.13 F
22:32:12.876 -> Average Cal sensor Temp.: 95.35 Ambient Temp.: 96.17 F
22:32:14.433 -> Average Cal sensor Temp.: 95.27 Ambient Temp.: 96.10 F
22:32:15.993 -> Average Cal sensor Temp.: 95.58 Ambient Temp.: 96.13 F
22:32:17.548 -> Average Cal sensor Temp.: 95.38 Ambient Temp.: 96.17 F
22:32:19.105 -> Average Cal sensor Temp.: 95.44 Ambient Temp.: 96.10 F
22:32:20.628 -> Average Cal sensor Temp.: 95.41 Ambient Temp.: 96.13 F
22:32:22.184 -> Average Cal sensor Temp.: 95.51 Ambient Temp.: 96.13 F
22:32:23.741 -> Average Cal sensor Temp.: 95.44 Ambient Temp.: 96.13 F
22:32:25.297 -> Average Cal sensor Temp.: 94.88 Ambient Temp.: 96.10 F
22:32:26.883 -> Average Cal sensor Temp.: 95.26 Ambient Temp.: 96.10 F
22:32:28.437 -> Average Cal sensor Temp.: 95.48 Ambient Temp.: 96.13 F
22:32:29.985 -> Average Cal sensor Temp.: 95.30 Ambient Temp.: 96.17 F
22:32:31.538 -> Average Cal sensor Temp.: 95.13 Ambient Temp.: 96.17 F
22:32:33.092 -> Average Cal sensor Temp.: 95.22 Ambient Temp.: 96.17 F
22:32:34.651 -> Average Cal sensor Temp.: 95.59 Ambient Temp.: 96.17 F
22:32:36.205 -> Average Cal sensor Temp.: 95.70 Ambient Temp.: 96.17 F
22:32:37.757 -> Average Cal sensor Temp.: 95.69 Ambient Temp.: 96.17 F
22:32:39.315 -> Average Cal sensor Temp.: 95.54 Ambient Temp.: 96.21 F
22:32:40.872 -> Average Cal sensor Temp.: 95.03 Ambient Temp.: 96.17 F
22:32:42.427 -> Average Cal sensor Temp.: 95.16 Ambient Temp.: 96.13 F
22:32:43.979 -> Average Cal sensor Temp.: 95.15 Ambient Temp.: 96.13 F
22:32:45.538 -> Average Cal sensor Temp.: 95.08 Ambient Temp.: 96.17 F
22:32:47.099 -> Average Cal sensor Temp.: 95.34 Ambient Temp.: 96.17 F
22:32:48.655 -> Average Cal sensor Temp.: 95.46 Ambient Temp.: 96.13 F
22:32:50.212 -> Average Cal sensor Temp.: 95.50 Ambient Temp.: 96.17 F
22:32:51.768 -> Average Cal sensor Temp.: 95.54 Ambient Temp.: 96.13 F
22:32:53.319 -> Average Cal sensor Temp.: 95.02 Ambient Temp.: 96.13 F
22:32:54.876 -> Average Cal sensor Temp.: 95.34 Ambient Temp.: 96.21 F
22:32:56.428 -> Average Cal sensor Temp.: 95.19 Ambient Temp.: 96.13 F
22:32:57.981 -> Average Cal sensor Temp.: 95.36 Ambient Temp.: 96.21 F
22:32:59.532 -> Average Cal sensor Temp.: 95.18 Ambient Temp.: 96.13 F
22:33:01.083 -> Average Cal sensor Temp.: 95.14 Ambient Temp.: 96.13 F
22:33:02.637 -> Average Cal sensor Temp.: 95.43 Ambient Temp.: 96.13 F
22:33:04.190 -> Average Cal sensor Temp.: 95.69 Ambient Temp.: 96.21 F
22:33:05.748 -> Average Cal sensor Temp.: 95.76 Ambient Temp.: 96.17 F
22:33:07.298 -> Average Cal sensor Temp.: 95.56 Ambient Temp.: 96.13 F
22:33:08.854 -> Average Cal sensor Temp.: 95.62 Ambient Temp.: 96.17 F
22:33:10.407 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 96.13 F
22:33:11.963 -> Average Cal sensor Temp.: 95.70 Ambient Temp.: 96.17 F
22:33:13.544 -> Average Cal sensor Temp.: 95.75 Ambient Temp.: 96.17 F
22:33:15.066 -> Average Cal sensor Temp.: 95.69 Ambient Temp.: 96.13 F
22:33:16.622 -> Average Cal sensor Temp.: 95.28 Ambient Temp.: 96.17 F
22:33:18.181 -> Average Cal sensor Temp.: 95.47 Ambient Temp.: 96.24 F
22:33:19.738 -> Average Cal sensor Temp.: 95.47 Ambient Temp.: 96.17 F
22:33:21.293 -> Average Cal sensor Temp.: 95.39 Ambient Temp.: 96.17 F
22:33:22.877 -> Average Cal sensor Temp.: 95.59 Ambient Temp.: 96.17 F
22:33:24.399 -> Average Cal sensor Temp.: 95.37 Ambient Temp.: 96.17 F
22:33:25.956 -> Average Cal sensor Temp.: 95.02 Ambient Temp.: 96.17 F
22:33:27.541 -> Average Cal sensor Temp.: 95.42 Ambient Temp.: 96.21 F
22:33:29.097 -> Average Cal sensor Temp.: 95.43 Ambient Temp.: 96.17 F
22:33:30.652 -> Average Cal sensor Temp.: 95.39 Ambient Temp.: 96.17 F
22:33:32.203 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 96.17 F
22:33:33.757 -> Average Cal sensor Temp.: 95.37 Ambient Temp.: 96.13 F
22:33:35.316 -> Average Cal sensor Temp.: 95.08 Ambient Temp.: 96.17 F
22:33:36.867 -> Average Cal sensor Temp.: 95.22 Ambient Temp.: 96.10 F
22:33:38.418 -> Average Cal sensor Temp.: 95.60 Ambient Temp.: 96.13 F
22:33:39.966 -> Average Cal sensor Temp.: 95.36 Ambient Temp.: 96.10 F
22:33:41.518 -> Average Cal sensor Temp.: 95.37 Ambient Temp.: 96.13 F
22:33:43.073 -> Average Cal sensor Temp.: 95.48 Ambient Temp.: 96.13 F
22:33:44.627 -> Average Cal sensor Temp.: 95.62 Ambient Temp.: 96.17 F
22:33:46.182 -> Average Cal sensor Temp.: 95.72 Ambient Temp.: 96.17 F
22:33:47.736 -> Average Cal sensor Temp.: 95.39 Ambient Temp.: 96.13 F
22:33:49.288 -> Average Cal sensor Temp.: 95.62 Ambient Temp.: 96.13 F
22:33:50.843 -> Average Cal sensor Temp.: 95.54 Ambient Temp.: 96.13 F
22:33:52.404 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 96.17 F
22:33:53.965 -> Average Cal sensor Temp.: 95.60 Ambient Temp.: 96.13 F
22:33:55.515 -> Average Cal sensor Temp.: 95.17 Ambient Temp.: 96.13 F
22:33:57.072 -> Average Cal sensor Temp.: 95.57 Ambient Temp.: 96.17 F
22:33:58.627 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 96.17 F
22:34:00.181 -> Average Cal sensor Temp.: 95.58 Ambient Temp.: 96.17 F
22:34:01.739 -> Average Cal sensor Temp.: 95.61 Ambient Temp.: 96.17 F
22:34:03.293 -> Average Cal sensor Temp.: 95.75 Ambient Temp.: 96.13 F
22:34:04.851 -> Average Cal sensor Temp.: 95.54 Ambient Temp.: 96.17 F
22:34:06.405 -> Average Cal sensor Temp.: 95.47 Ambient Temp.: 96.13 F
22:34:07.960 -> Average Cal sensor Temp.: 95.64 Ambient Temp.: 96.13 F
22:34:09.510 -> Average Cal sensor Temp.: 95.67 Ambient Temp.: 96.21 F
22:34:11.072 -> Average Cal sensor Temp.: 95.31 Ambient Temp.: 96.13 F
22:34:12.626 -> Average Cal sensor Temp.: 95.50 Ambient Temp.: 96.17 F
22:34:14.178 -> Average Cal sensor Temp.: 95.39 Ambient Temp.: 96.17 F
22:34:15.737 -> Average Cal sensor Temp.: 95.65 Ambient Temp.: 96.17 F
22:34:17.293 -> Average Cal sensor Temp.: 95.55 Ambient Temp.: 96.17 F
22:34:18.848 -> Average Cal sensor Temp.: 95.69 Ambient Temp.: 96.17 F
22:34:20.399 -> Average Cal sensor Temp.: 95.72 Ambient Temp.: 96.13 F
22:34:21.952 -> Average Cal sensor Temp.: 95.75 Ambient Temp.: 96.17 F
22:34:23.510 -> Average Cal sensor Temp.: 95.56 Ambient Temp.: 96.21 F
22:34:25.065 -> Average Cal sensor Temp.: 95.59 Ambient Temp.: 96.17 F
22:34:26.614 -> Average Cal sensor Temp.: 95.78 Ambient Temp.: 96.17 F
22:34:28.165 -> Average Cal sensor Temp.: 95.84 Ambient Temp.: 96.17 F
22:34:29.752 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 96.13 F
22:34:31.275 -> Average Cal sensor Temp.: 95.89 Ambient Temp.: 96.24 F
22:34:32.859 -> Average Cal sensor Temp.: 95.66 Ambient Temp.: 96.24 F
22:34:34.410 -> Average Cal sensor Temp.: 95.77 Ambient Temp.: 96.17 F
22:34:35.967 -> Average Cal sensor Temp.: 95.85 Ambient Temp.: 96.21 F
22:34:37.522 -> Average Cal sensor Temp.: 95.98 Ambient Temp.: 96.17 F
22:34:39.076 -> Average Cal sensor Temp.: 95.82 Ambient Temp.: 96.21 F
22:34:40.629 -> Average Cal sensor Temp.: 96.03 Ambient Temp.: 96.21 F
22:34:42.189 -> Average Cal sensor Temp.: 95.57 Ambient Temp.: 96.17 F
22:34:43.748 -> Average Cal sensor Temp.: 95.77 Ambient Temp.: 96.21 F
22:34:45.304 -> Average Cal sensor Temp.: 95.54 Ambient Temp.: 96.21 F
22:34:46.851 -> Average Cal sensor Temp.: 95.33 Ambient Temp.: 96.17 F
22:34:48.408 -> Average Cal sensor Temp.: 95.74 Ambient Temp.: 96.24 F
22:34:49.958 -> Average Cal sensor Temp.: 95.84 Ambient Temp.: 96.24 F
22:34:51.513 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 96.24 F
22:34:53.068 -> Average Cal sensor Temp.: 95.85 Ambient Temp.: 96.21 F
22:34:54.624 -> Average Cal sensor Temp.: 95.80 Ambient Temp.: 96.17 F
22:34:56.176 -> Average Cal sensor Temp.: 95.49 Ambient Temp.: 96.21 F
22:34:57.726 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 96.24 F
22:34:59.315 -> Average Cal sensor Temp.: 95.62 Ambient Temp.: 96.24 F
22:35:00.869 -> Average Cal sensor Temp.: 95.64 Ambient Temp.: 96.21 F
22:35:02.418 -> Average Cal sensor Temp.: 95.79 Ambient Temp.: 96.21 F
22:35:03.975 -> Average Cal sensor Temp.: 95.72 Ambient Temp.: 96.21 F
22:35:05.532 -> Average Cal sensor Temp.: 95.97 Ambient Temp.: 96.28 F
22:35:07.087 -> Average Cal sensor Temp.: 95.85 Ambient Temp.: 96.24 F
22:35:08.643 -> Average Cal sensor Temp.: 95.88 Ambient Temp.: 96.21 F
22:35:10.203 -> Average Cal sensor Temp.: 95.76 Ambient Temp.: 96.24 F
22:35:11.757 -> Average Cal sensor Temp.: 95.45 Ambient Temp.: 96.21 F
22:35:13.308 -> Average Cal sensor Temp.: 95.50 Ambient Temp.: 96.24 F
22:35:14.871 -> Average Cal sensor Temp.: 95.80 Ambient Temp.: 96.21 F
22:35:16.428 -> Average Cal sensor Temp.: 95.93 Ambient Temp.: 96.28 F
22:35:17.985 -> Average Cal sensor Temp.: 95.91 Ambient Temp.: 96.21 F
22:35:19.537 -> Average Cal sensor Temp.: 95.83 Ambient Temp.: 96.21 F
22:35:21.093 -> Average Cal sensor Temp.: 95.90 Ambient Temp.: 96.21 F
22:35:22.649 -> Average Cal sensor Temp.: 95.90 Ambient Temp.: 96.24 F
22:35:24.202 -> Average Cal sensor Temp.: 95.78 Ambient Temp.: 96.24 F
22:35:25.751 -> Average Cal sensor Temp.: 95.80 Ambient Temp.: 96.21 F
22:35:27.302 -> Average Cal sensor Temp.: 95.86 Ambient Temp.: 96.24 F
22:35:28.852 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 96.24 F
22:35:30.406 -> Average Cal sensor Temp.: 95.72 Ambient Temp.: 96.24 F
22:35:31.960 -> Average Cal sensor Temp.: 95.42 Ambient Temp.: 96.24 F
22:35:33.517 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 96.21 F
22:35:35.076 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 96.24 F
22:35:36.659 -> Average Cal sensor Temp.: 95.84 Ambient Temp.: 96.21 F
22:35:38.214 -> Average Cal sensor Temp.: 95.82 Ambient Temp.: 96.24 F
22:35:39.769 -> Average Cal sensor Temp.: 95.71 Ambient Temp.: 96.21 F
22:35:41.324 -> Average Cal sensor Temp.: 95.43 Ambient Temp.: 96.28 F
22:35:42.875 -> Average Cal sensor Temp.: 95.68 Ambient Temp.: 96.24 F
22:35:44.423 -> Average Cal sensor Temp.: 95.79 Ambient Temp.: 96.24 F
22:35:45.977 -> Average Cal sensor Temp.: 95.69 Ambient Temp.: 96.24 F
22:35:47.531 -> Average Cal sensor Temp.: 95.22 Ambient Temp.: 96.28 F
22:35:49.088 -> Average Cal sensor Temp.: 95.58 Ambient Temp.: 96.24 F
22:35:50.639 -> Average Cal sensor Temp.: 95.75 Ambient Temp.: 96.24 F
22:35:52.191 -> Average Cal sensor Temp.: 95.95 Ambient Temp.: 96.21 F
22:35:53.747 -> Average Cal sensor Temp.: 96.10 Ambient Temp.: 96.24 F
22:35:55.303 -> Average Cal sensor Temp.: 96.22 Ambient Temp.: 96.24 F
22:35:56.890 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 96.24 F
22:35:58.441 -> Average Cal sensor Temp.: 95.69 Ambient Temp.: 96.24 F
22:36:00.000 -> Average Cal sensor Temp.: 96.00 Ambient Temp.: 96.24 F
22:36:01.553 -> Average Cal sensor Temp.: 96.05 Ambient Temp.: 96.24 F
22:36:03.106 -> Average Cal sensor Temp.: 95.93 Ambient Temp.: 96.21 F
22:36:04.664 -> Average Cal sensor Temp.: 96.11 Ambient Temp.: 96.24 F
22:36:06.218 -> Average Cal sensor Temp.: 96.13 Ambient Temp.: 96.28 F
22:36:07.776 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 96.24 F
22:36:09.328 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 96.28 F
22:36:10.885 -> Average Cal sensor Temp.: 95.62 Ambient Temp.: 96.24 F
22:36:12.438 -> Average Cal sensor Temp.: 95.78 Ambient Temp.: 96.28 F
22:36:13.990 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 96.24 F
22:36:15.537 -> Average Cal sensor Temp.: 95.83 Ambient Temp.: 96.24 F
22:36:17.096 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 96.28 F
22:36:18.649 -> Average Cal sensor Temp.: 96.01 Ambient Temp.: 96.24 F
22:36:20.207 -> Average Cal sensor Temp.: 95.83 Ambient Temp.: 96.28 F
22:36:21.796 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 96.24 F
22:36:23.350 -> Average Cal sensor Temp.: 96.01 Ambient Temp.: 96.24 F
22:36:24.902 -> Average Cal sensor Temp.: 96.14 Ambient Temp.: 96.24 F
22:36:26.462 -> Average Cal sensor Temp.: 95.93 Ambient Temp.: 96.24 F
22:36:28.019 -> Average Cal sensor Temp.: 95.84 Ambient Temp.: 96.28 F
22:36:29.576 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 96.24 F
22:36:31.128 -> Average Cal sensor Temp.: 96.03 Ambient Temp.: 96.24 F
22:36:32.680 -> Average Cal sensor Temp.: 96.23 Ambient Temp.: 96.24 F
22:36:34.238 -> Average Cal sensor Temp.: 96.20 Ambient Temp.: 96.24 F
22:36:35.792 -> Average Cal sensor Temp.: 96.14 Ambient Temp.: 96.24 F
22:36:37.341 -> Average Cal sensor Temp.: 96.12 Ambient Temp.: 96.24 F
22:36:38.890 -> Average Cal sensor Temp.: 96.23 Ambient Temp.: 96.31 F
22:36:40.446 -> Average Cal sensor Temp.: 95.86 Ambient Temp.: 96.28 F
22:36:42.003 -> Average Cal sensor Temp.: 96.07 Ambient Temp.: 96.28 F
22:36:43.557 -> Average Cal sensor Temp.: 95.89 Ambient Temp.: 96.24 F
22:36:45.116 -> Average Cal sensor Temp.: 95.93 Ambient Temp.: 96.28 F
22:36:46.699 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 96.28 F
22:36:48.256 -> Average Cal sensor Temp.: 95.86 Ambient Temp.: 96.28 F
22:36:49.810 -> Average Cal sensor Temp.: 96.05 Ambient Temp.: 96.24 F
22:36:51.369 -> Average Cal sensor Temp.: 95.86 Ambient Temp.: 96.24 F
22:36:52.926 -> Average Cal sensor Temp.: 95.93 Ambient Temp.: 96.24 F
22:36:54.481 -> Average Cal sensor Temp.: 95.82 Ambient Temp.: 96.28 F
22:36:56.035 -> Average Cal sensor Temp.: 95.99 Ambient Temp.: 96.28 F
22:36:57.592 -> Average Cal sensor Temp.: 95.94 Ambient Temp.: 96.24 F
22:36:59.154 -> Average Cal sensor Temp.: 95.90 Ambient Temp.: 96.24 F
22:37:00.708 -> Average Cal sensor Temp.: 96.12 Ambient Temp.: 96.24 F
22:37:02.258 -> Average Cal sensor Temp.: 95.90 Ambient Temp.: 96.28 F
22:37:03.809 -> Average Cal sensor Temp.: 95.93 Ambient Temp.: 96.31 F
22:37:05.363 -> Average Cal sensor Temp.: 95.95 Ambient Temp.: 96.24 F
22:37:06.923 -> Average Cal sensor Temp.: 95.95 Ambient Temp.: 96.28 F
22:37:08.481 -> Average Cal sensor Temp.: 95.86 Ambient Temp.: 96.28 F
22:37:10.033 -> Average Cal sensor Temp.: 95.90 Ambient Temp.: 96.21 F
22:37:11.589 -> Average Cal sensor Temp.: 95.86 Ambient Temp.: 96.24 F
22:37:13.173 -> Average Cal sensor Temp.: 96.01 Ambient Temp.: 96.28 F
22:37:14.697 -> Average Cal sensor Temp.: 95.95 Ambient Temp.: 96.24 F
22:37:16.256 -> Average Cal sensor Temp.: 95.99 Ambient Temp.: 96.24 F
22:37:17.812 -> Average Cal sensor Temp.: 95.93 Ambient Temp.: 96.28 F
22:37:19.369 -> Average Cal sensor Temp.: 96.02 Ambient Temp.: 96.24 F
22:37:20.933 -> Average Cal sensor Temp.: 96.05 Ambient Temp.: 96.28 F
22:37:22.485 -> Average Cal sensor Temp.: 96.15 Ambient Temp.: 96.24 F
22:37:24.044 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 96.24 F
22:37:25.601 -> Average Cal sensor Temp.: 96.18 Ambient Temp.: 96.24 F
22:37:27.162 -> Average Cal sensor Temp.: 96.19 Ambient Temp.: 96.24 F
22:37:28.719 -> Average Cal sensor Temp.: 96.11 Ambient Temp.: 96.24 F
22:37:30.279 -> Average Cal sensor Temp.: 96.21 Ambient Temp.: 96.24 F
22:37:31.836 -> Average Cal sensor Temp.: 96.21 Ambient Temp.: 96.24 F
22:37:33.391 -> Average Cal sensor Temp.: 96.07 Ambient Temp.: 96.28 F
22:37:34.943 -> Average Cal sensor Temp.: 95.61 Ambient Temp.: 96.24 F
22:37:36.493 -> Average Cal sensor Temp.: 95.60 Ambient Temp.: 96.24 F
22:37:38.052 -> Average Cal sensor Temp.: 95.94 Ambient Temp.: 96.24 F
22:37:39.605 -> Average Cal sensor Temp.: 95.88 Ambient Temp.: 96.24 F
22:37:41.166 -> Average Cal sensor Temp.: 95.72 Ambient Temp.: 96.24 F
22:37:42.717 -> Average Cal sensor Temp.: 95.46 Ambient Temp.: 96.24 F
22:37:44.276 -> Average Cal sensor Temp.: 95.58 Ambient Temp.: 96.24 F
22:37:45.835 -> Average Cal sensor Temp.: 95.79 Ambient Temp.: 96.21 F
22:37:47.387 -> Average Cal sensor Temp.: 96.13 Ambient Temp.: 96.24 F
22:37:48.944 -> Average Cal sensor Temp.: 96.13 Ambient Temp.: 96.24 F
22:37:50.499 -> Average Cal sensor Temp.: 96.26 Ambient Temp.: 96.28 F
22:37:52.081 -> Average Cal sensor Temp.: 96.42 Ambient Temp.: 96.24 F
22:37:53.639 -> Average Cal sensor Temp.: 96.31 Ambient Temp.: 96.24 F
22:37:55.169 -> Average Cal sensor Temp.: 96.27 Ambient Temp.: 96.28 F
22:37:56.755 -> Average Cal sensor Temp.: 96.28 Ambient Temp.: 96.28 F
22:37:58.312 -> Average Cal sensor Temp.: 96.23 Ambient Temp.: 96.24 F
22:37:59.865 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 96.28 F
22:38:01.423 -> Average Cal sensor Temp.: 95.88 Ambient Temp.: 96.24 F
22:38:02.980 -> Average Cal sensor Temp.: 96.06 Ambient Temp.: 96.28 F
22:38:04.533 -> Average Cal sensor Temp.: 96.22 Ambient Temp.: 96.28 F
22:38:06.092 -> Average Cal sensor Temp.: 95.90 Ambient Temp.: 96.31 F
22:38:07.647 -> Average Cal sensor Temp.: 95.90 Ambient Temp.: 96.24 F
22:38:09.200 -> Average Cal sensor Temp.: 95.87 Ambient Temp.: 96.24 F
22:38:10.758 -> Average Cal sensor Temp.: 95.97 Ambient Temp.: 96.28 F
22:38:12.314 -> Average Cal sensor Temp.: 96.00 Ambient Temp.: 96.24 F
22:38:13.864 -> Average Cal sensor Temp.: 96.18 Ambient Temp.: 96.28 F
22:38:15.421 -> Average Cal sensor Temp.: 96.16 Ambient Temp.: 96.24 F
22:38:16.968 -> Average Cal sensor Temp.: 95.86 Ambient Temp.: 96.28 F
22:38:18.527 -> Average Cal sensor Temp.: 95.90 Ambient Temp.: 96.21 F
22:38:20.110 -> Average Cal sensor Temp.: 95.82 Ambient Temp.: 96.24 F
22:38:21.662 -> Average Cal sensor Temp.: 95.98 Ambient Temp.: 96.28 F
22:38:23.214 -> Average Cal sensor Temp.: 95.80 Ambient Temp.: 96.28 F
22:38:24.768 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 96.24 F
22:38:26.315 -> Average Cal sensor Temp.: 96.10 Ambient Temp.: 96.28 F
22:38:27.868 -> Average Cal sensor Temp.: 96.05 Ambient Temp.: 96.31 F
22:38:29.425 -> Average Cal sensor Temp.: 96.23 Ambient Temp.: 96.28 F
22:38:31.014 -> Average Cal sensor Temp.: 96.34 Ambient Temp.: 96.24 F
22:38:32.569 -> Average Cal sensor Temp.: 96.19 Ambient Temp.: 96.28 F
22:38:34.120 -> Average Cal sensor Temp.: 96.16 Ambient Temp.: 96.24 F
22:38:35.677 -> Average Cal sensor Temp.: 96.19 Ambient Temp.: 96.31 F
22:38:37.225 -> Average Cal sensor Temp.: 96.26 Ambient Temp.: 96.28 F
22:38:38.780 -> Average Cal sensor Temp.: 95.94 Ambient Temp.: 96.31 F
22:38:40.335 -> Average Cal sensor Temp.: 96.07 Ambient Temp.: 96.28 F
22:38:41.890 -> Average Cal sensor Temp.: 95.97 Ambient Temp.: 96.24 F
22:38:43.445 -> Average Cal sensor Temp.: 95.93 Ambient Temp.: 96.28 F
22:38:45.002 -> Average Cal sensor Temp.: 96.09 Ambient Temp.: 96.31 F
22:38:46.554 -> Average Cal sensor Temp.: 96.03 Ambient Temp.: 96.24 F
22:38:48.141 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 96.28 F
22:38:49.687 -> Average Cal sensor Temp.: 96.12 Ambient Temp.: 96.35 F
22:38:51.239 -> Average Cal sensor Temp.: 95.73 Ambient Temp.: 96.31 F
22:38:52.797 -> Average Cal sensor Temp.: 96.10 Ambient Temp.: 96.35 F
22:38:54.355 -> Average Cal sensor Temp.: 96.03 Ambient Temp.: 96.35 F
22:38:55.907 -> Average Cal sensor Temp.: 96.19 Ambient Temp.: 96.28 F
22:38:57.454 -> Average Cal sensor Temp.: 96.32 Ambient Temp.: 96.31 F
22:38:59.011 -> Average Cal sensor Temp.: 96.34 Ambient Temp.: 96.31 F
22:39:00.560 -> Average Cal sensor Temp.: 96.53 Ambient Temp.: 96.35 F
22:39:02.147 -> Average Cal sensor Temp.: 96.47 Ambient Temp.: 96.35 F
22:39:03.704 -> Average Cal sensor Temp.: 96.44 Ambient Temp.: 96.35 F
22:39:05.258 -> Average Cal sensor Temp.: 96.46 Ambient Temp.: 96.35 F
22:39:06.808 -> Average Cal sensor Temp.: 96.32 Ambient Temp.: 96.35 F
22:39:08.360 -> Average Cal sensor Temp.: 96.33 Ambient Temp.: 96.31 F
22:39:09.916 -> Average Cal sensor Temp.: 96.49 Ambient Temp.: 96.35 F
22:39:11.473 -> Average Cal sensor Temp.: 96.21 Ambient Temp.: 96.35 F
22:39:13.027 -> Average Cal sensor Temp.: 95.89 Ambient Temp.: 96.31 F
22:39:14.579 -> Average Cal sensor Temp.: 95.87 Ambient Temp.: 96.35 F
22:39:16.137 -> Average Cal sensor Temp.: 96.15 Ambient Temp.: 96.35 F
22:39:17.687 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 96.39 F
22:39:19.247 -> Average Cal sensor Temp.: 96.29 Ambient Temp.: 96.35 F
22:39:20.809 -> Average Cal sensor Temp.: 96.13 Ambient Temp.: 96.31 F
22:39:22.363 -> Average Cal sensor Temp.: 96.26 Ambient Temp.: 96.35 F
22:39:23.917 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 96.31 F
22:39:25.471 -> Average Cal sensor Temp.: 96.45 Ambient Temp.: 96.35 F
22:39:27.028 -> Average Cal sensor Temp.: 96.42 Ambient Temp.: 96.39 F
22:39:28.585 -> Average Cal sensor Temp.: 96.51 Ambient Temp.: 96.35 F
22:39:30.141 -> Average Cal sensor Temp.: 96.66 Ambient Temp.: 96.39 F
22:39:31.703 -> Average Cal sensor Temp.: 96.48 Ambient Temp.: 96.35 F
22:39:33.246 -> Average Cal sensor Temp.: 96.18 Ambient Temp.: 96.35 F
22:39:34.808 -> Average Cal sensor Temp.: 96.33 Ambient Temp.: 96.39 F
22:39:36.366 -> Average Cal sensor Temp.: 96.47 Ambient Temp.: 96.31 F
22:39:37.918 -> Average Cal sensor Temp.: 96.37 Ambient Temp.: 96.39 F
22:39:39.476 -> Average Cal sensor Temp.: 96.36 Ambient Temp.: 96.35 F
22:39:41.035 -> Average Cal sensor Temp.: 96.24 Ambient Temp.: 96.35 F
22:39:42.593 -> Average Cal sensor Temp.: 96.21 Ambient Temp.: 96.39 F
22:39:44.143 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 96.39 F
22:39:45.705 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 96.39 F
22:39:47.259 -> Average Cal sensor Temp.: 96.33 Ambient Temp.: 96.35 F
22:39:48.815 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 96.39 F
22:39:50.369 -> Average Cal sensor Temp.: 96.21 Ambient Temp.: 96.39 F
22:39:51.956 -> Average Cal sensor Temp.: 96.15 Ambient Temp.: 96.35 F
22:39:53.509 -> Average Cal sensor Temp.: 95.86 Ambient Temp.: 96.31 F
22:39:55.070 -> Average Cal sensor Temp.: 95.98 Ambient Temp.: 96.35 F
22:39:56.624 -> Average Cal sensor Temp.: 96.20 Ambient Temp.: 96.35 F
22:39:58.180 -> Average Cal sensor Temp.: 96.03 Ambient Temp.: 96.39 F
22:39:59.734 -> Average Cal sensor Temp.: 96.21 Ambient Temp.: 96.31 F
22:40:01.290 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 96.35 F
22:40:02.844 -> Average Cal sensor Temp.: 96.04 Ambient Temp.: 96.35 F
22:40:04.393 -> Average Cal sensor Temp.: 96.16 Ambient Temp.: 96.35 F
22:40:05.948 -> Average Cal sensor Temp.: 96.39 Ambient Temp.: 96.39 F
22:40:07.497 -> Average Cal sensor Temp.: 96.53 Ambient Temp.: 96.35 F
22:40:09.049 -> Average Cal sensor Temp.: 96.67 Ambient Temp.: 96.42 F
22:40:10.607 -> Average Cal sensor Temp.: 96.71 Ambient Temp.: 96.39 F
22:40:12.164 -> Average Cal sensor Temp.: 96.42 Ambient Temp.: 96.42 F
22:40:13.720 -> Average Cal sensor Temp.: 96.50 Ambient Temp.: 96.39 F
22:40:15.307 -> Average Cal sensor Temp.: 96.63 Ambient Temp.: 96.35 F
22:40:16.858 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 96.39 F
22:40:18.410 -> Average Cal sensor Temp.: 96.68 Ambient Temp.: 96.31 F
22:40:19.969 -> Average Cal sensor Temp.: 96.29 Ambient Temp.: 96.42 F
22:40:21.522 -> Average Cal sensor Temp.: 96.01 Ambient Temp.: 96.35 F
22:40:23.077 -> Average Cal sensor Temp.: 96.15 Ambient Temp.: 96.35 F
22:40:24.633 -> Average Cal sensor Temp.: 96.27 Ambient Temp.: 96.42 F
22:40:26.191 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 96.39 F
22:40:27.744 -> Average Cal sensor Temp.: 96.40 Ambient Temp.: 96.42 F
22:40:29.298 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 96.39 F
22:40:30.862 -> Average Cal sensor Temp.: 96.23 Ambient Temp.: 96.42 F
22:40:32.417 -> Average Cal sensor Temp.: 96.28 Ambient Temp.: 96.42 F
22:40:33.970 -> Average Cal sensor Temp.: 96.19 Ambient Temp.: 96.35 F
22:40:35.522 -> Average Cal sensor Temp.: 96.24 Ambient Temp.: 96.35 F
22:40:37.075 -> Average Cal sensor Temp.: 96.19 Ambient Temp.: 96.39 F
22:40:38.635 -> Average Cal sensor Temp.: 96.28 Ambient Temp.: 96.42 F
22:40:40.188 -> Average Cal sensor Temp.: 96.14 Ambient Temp.: 96.39 F
22:40:41.742 -> Average Cal sensor Temp.: 96.28 Ambient Temp.: 96.39 F
22:40:43.310 -> Average Cal sensor Temp.: 96.34 Ambient Temp.: 96.39 F
22:40:44.869 -> Average Cal sensor Temp.: 96.42 Ambient Temp.: 96.39 F
22:40:46.416 -> Average Cal sensor Temp.: 96.49 Ambient Temp.: 96.39 F
22:40:47.971 -> Average Cal sensor Temp.: 96.66 Ambient Temp.: 96.42 F
22:40:49.561 -> Average Cal sensor Temp.: 96.40 Ambient Temp.: 96.39 F
22:40:51.114 -> Average Cal sensor Temp.: 96.33 Ambient Temp.: 96.39 F
22:40:52.664 -> Average Cal sensor Temp.: 95.99 Ambient Temp.: 96.42 F
22:40:54.221 -> Average Cal sensor Temp.: 96.22 Ambient Temp.: 96.39 F
22:40:55.777 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 96.39 F
22:40:57.335 -> Average Cal sensor Temp.: 96.28 Ambient Temp.: 96.39 F
22:40:58.888 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 96.42 F
22:41:00.440 -> Average Cal sensor Temp.: 95.97 Ambient Temp.: 96.49 F
22:41:01.996 -> Average Cal sensor Temp.: 96.33 Ambient Temp.: 96.39 F
22:41:03.547 -> Average Cal sensor Temp.: 96.39 Ambient Temp.: 96.39 F
22:41:05.103 -> Average Cal sensor Temp.: 96.40 Ambient Temp.: 96.39 F
22:41:06.660 -> Average Cal sensor Temp.: 96.54 Ambient Temp.: 96.39 F
22:41:08.248 -> Average Cal sensor Temp.: 96.21 Ambient Temp.: 96.39 F
22:41:09.804 -> Average Cal sensor Temp.: 96.36 Ambient Temp.: 96.35 F
22:41:11.327 -> Average Cal sensor Temp.: 96.24 Ambient Temp.: 96.39 F
22:41:12.912 -> Average Cal sensor Temp.: 96.11 Ambient Temp.: 96.39 F
22:41:14.466 -> Average Cal sensor Temp.: 95.99 Ambient Temp.: 96.35 F
22:41:16.023 -> Average Cal sensor Temp.: 96.23 Ambient Temp.: 96.39 F
22:41:17.574 -> Average Cal sensor Temp.: 96.49 Ambient Temp.: 96.39 F
22:41:19.131 -> Average Cal sensor Temp.: 96.55 Ambient Temp.: 96.35 F
22:41:20.688 -> Average Cal sensor Temp.: 96.63 Ambient Temp.: 96.39 F
22:41:22.240 -> Average Cal sensor Temp.: 96.62 Ambient Temp.: 96.39 F
22:41:23.798 -> Average Cal sensor Temp.: 96.57 Ambient Temp.: 96.42 F
22:41:25.351 -> Average Cal sensor Temp.: 96.58 Ambient Temp.: 96.49 F
22:41:26.908 -> Average Cal sensor Temp.: 96.71 Ambient Temp.: 96.42 F
22:41:28.466 -> Average Cal sensor Temp.: 96.47 Ambient Temp.: 96.39 F
22:41:30.021 -> Average Cal sensor Temp.: 96.32 Ambient Temp.: 96.35 F
22:41:31.577 -> Average Cal sensor Temp.: 96.24 Ambient Temp.: 96.35 F
22:41:33.132 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 96.39 F
22:41:34.690 -> Average Cal sensor Temp.: 96.41 Ambient Temp.: 96.35 F
22:41:36.242 -> Average Cal sensor Temp.: 96.40 Ambient Temp.: 96.39 F
22:41:37.797 -> Average Cal sensor Temp.: 96.52 Ambient Temp.: 96.35 F
22:41:39.350 -> Average Cal sensor Temp.: 96.57 Ambient Temp.: 96.39 F
22:41:40.937 -> Average Cal sensor Temp.: 96.77 Ambient Temp.: 96.39 F
22:41:42.493 -> Average Cal sensor Temp.: 96.56 Ambient Temp.: 96.39 F
22:41:44.040 -> Average Cal sensor Temp.: 96.39 Ambient Temp.: 96.42 F
22:41:45.605 -> Average Cal sensor Temp.: 96.33 Ambient Temp.: 96.35 F
22:41:47.156 -> Average Cal sensor Temp.: 96.41 Ambient Temp.: 96.39 F
22:41:48.707 -> Average Cal sensor Temp.: 96.22 Ambient Temp.: 96.39 F
22:41:50.261 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 96.39 F
22:41:51.816 -> Average Cal sensor Temp.: 96.38 Ambient Temp.: 96.49 F
22:41:53.369 -> Average Cal sensor Temp.: 96.43 Ambient Temp.: 96.42 F
22:41:54.927 -> Average Cal sensor Temp.: 96.32 Ambient Temp.: 96.39 F
22:41:56.477 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 96.35 F
22:41:58.034 -> Average Cal sensor Temp.: 96.35 Ambient Temp.: 96.39 F
22:41:59.624 -> Average Cal sensor Temp.: 96.26 Ambient Temp.: 96.35 F
22:42:01.151 -> Average Cal sensor Temp.: 95.68 Ambient Temp.: 96.39 F
22:42:02.737 -> Average Cal sensor Temp.: 96.07 Ambient Temp.: 96.35 F
22:42:04.293 -> Average Cal sensor Temp.: 96.04 Ambient Temp.: 96.35 F
22:42:05.848 -> Average Cal sensor Temp.: 96.18 Ambient Temp.: 96.39 F
22:42:07.404 -> Average Cal sensor Temp.: 96.21 Ambient Temp.: 96.39 F
22:42:08.959 -> Average Cal sensor Temp.: 96.48 Ambient Temp.: 96.35 F
22:42:10.515 -> Average Cal sensor Temp.: 96.40 Ambient Temp.: 96.39 F
22:42:12.067 -> Average Cal sensor Temp.: 96.61 Ambient Temp.: 96.39 F
22:42:13.623 -> Average Cal sensor Temp.: 96.55 Ambient Temp.: 96.39 F
22:42:15.181 -> Average Cal sensor Temp.: 96.52 Ambient Temp.: 96.39 F
22:42:16.737 -> Average Cal sensor Temp.: 96.28 Ambient Temp.: 96.39 F
22:42:18.295 -> Average Cal sensor Temp.: 96.15 Ambient Temp.: 96.35 F
22:42:19.848 -> Average Cal sensor Temp.: 95.96 Ambient Temp.: 96.39 F
22:42:21.410 -> Average Cal sensor Temp.: 96.20 Ambient Temp.: 96.39 F
22:42:22.969 -> Average Cal sensor Temp.: 96.39 Ambient Temp.: 96.31 F
22:42:24.517 -> Average Cal sensor Temp.: 96.34 Ambient Temp.: 96.39 F
22:42:26.072 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 96.39 F
22:42:27.629 -> Average Cal sensor Temp.: 96.35 Ambient Temp.: 96.39 F
22:42:29.179 -> Average Cal sensor Temp.: 96.41 Ambient Temp.: 96.35 F
22:42:30.741 -> Average Cal sensor Temp.: 96.35 Ambient Temp.: 96.35 F
22:42:32.297 -> Average Cal sensor Temp.: 96.49 Ambient Temp.: 96.39 F
22:42:33.846 -> Average Cal sensor Temp.: 96.50 Ambient Temp.: 96.35 F
22:42:35.410 -> Average Cal sensor Temp.: 96.53 Ambient Temp.: 96.42 F
22:42:36.972 -> Average Cal sensor Temp.: 96.77 Ambient Temp.: 96.39 F
22:42:38.526 -> Average Cal sensor Temp.: 96.62 Ambient Temp.: 96.35 F
22:42:40.081 -> Average Cal sensor Temp.: 96.43 Ambient Temp.: 96.35 F
22:42:41.632 -> Average Cal sensor Temp.: 96.63 Ambient Temp.: 96.35 F
22:42:43.188 -> Average Cal sensor Temp.: 96.64 Ambient Temp.: 96.39 F
22:42:44.741 -> Average Cal sensor Temp.: 96.68 Ambient Temp.: 96.35 F
22:42:46.333 -> Average Cal sensor Temp.: 96.69 Ambient Temp.: 96.39 F
22:42:47.888 -> Average Cal sensor Temp.: 96.54 Ambient Temp.: 96.39 F
22:42:49.442 -> Average Cal sensor Temp.: 96.48 Ambient Temp.: 96.35 F
22:42:50.995 -> Average Cal sensor Temp.: 96.53 Ambient Temp.: 96.35 F
22:42:52.554 -> Average Cal sensor Temp.: 96.59 Ambient Temp.: 96.42 F
22:42:54.108 -> Average Cal sensor Temp.: 96.63 Ambient Temp.: 96.39 F
22:42:55.670 -> Average Cal sensor Temp.: 96.73 Ambient Temp.: 96.39 F
22:42:57.219 -> Average Cal sensor Temp.: 96.89 Ambient Temp.: 96.42 F
22:42:58.773 -> Average Cal sensor Temp.: 96.79 Ambient Temp.: 96.39 F
22:43:00.332 -> Average Cal sensor Temp.: 96.67 Ambient Temp.: 96.42 F
22:43:01.885 -> Average Cal sensor Temp.: 96.95 Ambient Temp.: 96.49 F
22:43:03.435 -> Average Cal sensor Temp.: 96.84 Ambient Temp.: 96.39 F
22:43:04.992 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 96.39 F
22:43:06.545 -> Average Cal sensor Temp.: 96.65 Ambient Temp.: 96.49 F
22:43:08.135 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 96.39 F
22:43:09.688 -> Average Cal sensor Temp.: 96.87 Ambient Temp.: 96.42 F
22:43:11.248 -> Average Cal sensor Temp.: 96.83 Ambient Temp.: 96.42 F
22:43:12.799 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 96.35 F
22:43:14.357 -> Average Cal sensor Temp.: 96.53 Ambient Temp.: 96.42 F
22:43:15.915 -> Average Cal sensor Temp.: 96.74 Ambient Temp.: 96.42 F
22:43:17.466 -> Average Cal sensor Temp.: 96.81 Ambient Temp.: 96.42 F
22:43:19.016 -> Average Cal sensor Temp.: 96.38 Ambient Temp.: 96.39 F
22:43:20.572 -> Average Cal sensor Temp.: 96.54 Ambient Temp.: 96.42 F
22:43:22.132 -> Average Cal sensor Temp.: 96.65 Ambient Temp.: 96.49 F
22:43:23.686 -> Average Cal sensor Temp.: 96.52 Ambient Temp.: 96.42 F
22:43:25.241 -> Average Cal sensor Temp.: 96.82 Ambient Temp.: 96.49 F
22:43:26.798 -> Average Cal sensor Temp.: 96.49 Ambient Temp.: 96.42 F
22:43:28.354 -> Average Cal sensor Temp.: 96.35 Ambient Temp.: 96.42 F
22:43:29.907 -> Average Cal sensor Temp.: 96.45 Ambient Temp.: 96.49 F
22:43:31.463 -> Average Cal sensor Temp.: 96.66 Ambient Temp.: 96.42 F
22:43:33.047 -> Average Cal sensor Temp.: 96.65 Ambient Temp.: 96.42 F
22:43:34.600 -> Average Cal sensor Temp.: 96.70 Ambient Temp.: 96.42 F
22:43:36.154 -> Average Cal sensor Temp.: 96.50 Ambient Temp.: 96.49 F
22:43:37.709 -> Average Cal sensor Temp.: 96.84 Ambient Temp.: 96.49 F
22:43:39.267 -> Average Cal sensor Temp.: 96.96 Ambient Temp.: 96.49 F
22:43:40.826 -> Average Cal sensor Temp.: 96.86 Ambient Temp.: 96.42 F
22:43:42.380 -> Average Cal sensor Temp.: 96.85 Ambient Temp.: 96.53 F
22:43:43.936 -> Average Cal sensor Temp.: 96.80 Ambient Temp.: 96.42 F
22:43:45.497 -> Average Cal sensor Temp.: 96.70 Ambient Temp.: 96.42 F
22:43:47.044 -> Average Cal sensor Temp.: 96.81 Ambient Temp.: 96.39 F
22:43:48.599 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 96.49 F
22:43:50.187 -> Average Cal sensor Temp.: 96.43 Ambient Temp.: 96.42 F
22:43:51.742 -> Average Cal sensor Temp.: 96.44 Ambient Temp.: 96.42 F
22:43:53.297 -> Average Cal sensor Temp.: 96.45 Ambient Temp.: 96.42 F
22:43:54.851 -> Average Cal sensor Temp.: 96.51 Ambient Temp.: 96.42 F
22:43:56.408 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 96.42 F
22:43:57.956 -> Average Cal sensor Temp.: 96.48 Ambient Temp.: 96.49 F
22:43:59.510 -> Average Cal sensor Temp.: 96.39 Ambient Temp.: 96.42 F
22:44:01.064 -> Average Cal sensor Temp.: 96.53 Ambient Temp.: 96.49 F
22:44:02.619 -> Average Cal sensor Temp.: 96.56 Ambient Temp.: 96.42 F
22:44:04.174 -> Average Cal sensor Temp.: 96.57 Ambient Temp.: 96.42 F
22:44:05.726 -> Average Cal sensor Temp.: 96.70 Ambient Temp.: 96.42 F
22:44:07.311 -> Average Cal sensor Temp.: 96.93 Ambient Temp.: 96.42 F
22:44:08.867 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 96.42 F
22:44:10.426 -> Average Cal sensor Temp.: 96.96 Ambient Temp.: 96.42 F
22:44:11.954 -> Average Cal sensor Temp.: 96.35 Ambient Temp.: 96.42 F
22:44:13.542 -> Average Cal sensor Temp.: 96.65 Ambient Temp.: 96.42 F
22:44:15.095 -> Average Cal sensor Temp.: 96.55 Ambient Temp.: 96.49 F
22:44:16.648 -> Average Cal sensor Temp.: 96.85 Ambient Temp.: 96.42 F
22:44:18.203 -> Average Cal sensor Temp.: 97.00 Ambient Temp.: 96.42 F
22:44:19.760 -> Average Cal sensor Temp.: 97.00 Ambient Temp.: 96.42 F
22:44:21.324 -> Average Cal sensor Temp.: 96.98 Ambient Temp.: 96.49 F
22:44:22.877 -> Average Cal sensor Temp.: 96.91 Ambient Temp.: 96.42 F
22:44:24.436 -> Average Cal sensor Temp.: 96.73 Ambient Temp.: 96.42 F
22:44:25.986 -> Average Cal sensor Temp.: 96.72 Ambient Temp.: 96.49 F
22:44:27.544 -> Average Cal sensor Temp.: 96.43 Ambient Temp.: 96.49 F
22:44:29.094 -> Average Cal sensor Temp.: 96.71 Ambient Temp.: 96.42 F
22:44:30.649 -> Average Cal sensor Temp.: 96.61 Ambient Temp.: 96.49 F
22:44:32.210 -> Average Cal sensor Temp.: 96.63 Ambient Temp.: 96.49 F
22:44:33.765 -> Average Cal sensor Temp.: 96.47 Ambient Temp.: 96.49 F
22:44:35.317 -> Average Cal sensor Temp.: 96.55 Ambient Temp.: 96.49 F
22:44:36.900 -> Average Cal sensor Temp.: 96.67 Ambient Temp.: 96.49 F
22:44:38.453 -> Average Cal sensor Temp.: 96.79 Ambient Temp.: 96.42 F
22:44:40.007 -> Average Cal sensor Temp.: 96.69 Ambient Temp.: 96.42 F
22:44:41.564 -> Average Cal sensor Temp.: 96.93 Ambient Temp.: 96.49 F
22:44:43.117 -> Average Cal sensor Temp.: 96.95 Ambient Temp.: 96.49 F
22:44:44.668 -> Average Cal sensor Temp.: 96.90 Ambient Temp.: 96.42 F
22:44:46.222 -> Average Cal sensor Temp.: 96.88 Ambient Temp.: 96.42 F
22:44:47.777 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 96.49 F
22:44:49.364 -> Average Cal sensor Temp.: 96.86 Ambient Temp.: 96.42 F
22:44:50.922 -> Average Cal sensor Temp.: 96.90 Ambient Temp.: 96.42 F
22:44:52.471 -> Average Cal sensor Temp.: 96.43 Ambient Temp.: 96.53 F
22:44:54.028 -> Average Cal sensor Temp.: 96.59 Ambient Temp.: 96.49 F
22:44:55.588 -> Average Cal sensor Temp.: 96.80 Ambient Temp.: 96.53 F
22:44:57.144 -> Average Cal sensor Temp.: 96.74 Ambient Temp.: 96.49 F
22:44:58.699 -> Average Cal sensor Temp.: 96.79 Ambient Temp.: 96.53 F
22:45:00.259 -> Average Cal sensor Temp.: 96.85 Ambient Temp.: 96.53 F
22:45:01.825 -> Average Cal sensor Temp.: 96.76 Ambient Temp.: 96.42 F
22:45:03.378 -> Average Cal sensor Temp.: 96.73 Ambient Temp.: 96.49 F
22:45:04.932 -> Average Cal sensor Temp.: 96.70 Ambient Temp.: 96.49 F
22:45:06.483 -> Average Cal sensor Temp.: 96.61 Ambient Temp.: 96.49 F
22:45:08.031 -> Average Cal sensor Temp.: 96.88 Ambient Temp.: 96.53 F
22:45:09.596 -> Average Cal sensor Temp.: 96.68 Ambient Temp.: 96.49 F
22:45:11.140 -> Average Cal sensor Temp.: 96.62 Ambient Temp.: 96.53 F
22:45:12.697 -> Average Cal sensor Temp.: 96.51 Ambient Temp.: 96.53 F
22:45:14.256 -> Average Cal sensor Temp.: 96.59 Ambient Temp.: 96.53 F
22:45:15.812 -> Average Cal sensor Temp.: 96.62 Ambient Temp.: 96.53 F
22:45:17.366 -> Average Cal sensor Temp.: 96.77 Ambient Temp.: 96.60 F
22:45:18.923 -> Average Cal sensor Temp.: 96.73 Ambient Temp.: 96.53 F
22:45:20.480 -> Average Cal sensor Temp.: 96.60 Ambient Temp.: 96.53 F
22:45:22.036 -> Average Cal sensor Temp.: 96.77 Ambient Temp.: 96.53 F
22:45:23.622 -> Average Cal sensor Temp.: 96.87 Ambient Temp.: 96.53 F
22:45:25.179 -> Average Cal sensor Temp.: 96.57 Ambient Temp.: 96.42 F
22:45:26.732 -> Average Cal sensor Temp.: 96.74 Ambient Temp.: 96.49 F
22:45:28.286 -> Average Cal sensor Temp.: 96.90 Ambient Temp.: 96.49 F
22:45:29.841 -> Average Cal sensor Temp.: 96.76 Ambient Temp.: 96.53 F
22:45:31.396 -> Average Cal sensor Temp.: 96.68 Ambient Temp.: 96.53 F
22:45:32.950 -> Average Cal sensor Temp.: 96.40 Ambient Temp.: 96.49 F
22:45:34.505 -> Average Cal sensor Temp.: 96.39 Ambient Temp.: 96.53 F
22:45:36.064 -> Average Cal sensor Temp.: 96.76 Ambient Temp.: 96.49 F
22:45:37.618 -> Average Cal sensor Temp.: 96.56 Ambient Temp.: 96.42 F
22:45:39.167 -> Average Cal sensor Temp.: 96.60 Ambient Temp.: 96.49 F
22:45:40.724 -> Average Cal sensor Temp.: 96.85 Ambient Temp.: 96.49 F
22:45:42.308 -> Average Cal sensor Temp.: 96.88 Ambient Temp.: 96.53 F
22:45:43.857 -> Average Cal sensor Temp.: 96.93 Ambient Temp.: 96.49 F
22:45:45.415 -> Average Cal sensor Temp.: 96.83 Ambient Temp.: 96.53 F
22:45:46.968 -> Average Cal sensor Temp.: 96.65 Ambient Temp.: 96.53 F
22:45:48.521 -> Average Cal sensor Temp.: 96.55 Ambient Temp.: 96.57 F
22:45:50.073 -> Average Cal sensor Temp.: 96.58 Ambient Temp.: 96.53 F
22:45:51.656 -> Average Cal sensor Temp.: 96.52 Ambient Temp.: 96.53 F
22:45:53.179 -> Average Cal sensor Temp.: 96.62 Ambient Temp.: 96.53 F
22:45:54.740 -> Average Cal sensor Temp.: 96.92 Ambient Temp.: 96.49 F
22:45:56.294 -> Average Cal sensor Temp.: 96.68 Ambient Temp.: 96.53 F
22:45:57.879 -> Average Cal sensor Temp.: 96.94 Ambient Temp.: 96.53 F
22:45:59.431 -> Average Cal sensor Temp.: 96.99 Ambient Temp.: 96.49 F
22:46:00.994 -> Average Cal sensor Temp.: 96.94 Ambient Temp.: 96.53 F
22:46:02.547 -> Average Cal sensor Temp.: 96.95 Ambient Temp.: 96.53 F
22:46:04.105 -> Average Cal sensor Temp.: 97.05 Ambient Temp.: 96.49 F
22:46:05.666 -> Average Cal sensor Temp.: 96.84 Ambient Temp.: 96.49 F
22:46:07.219 -> Average Cal sensor Temp.: 96.85 Ambient Temp.: 96.60 F
22:46:08.774 -> Average Cal sensor Temp.: 96.85 Ambient Temp.: 96.49 F
22:46:10.331 -> Average Cal sensor Temp.: 96.60 Ambient Temp.: 96.49 F
22:46:11.889 -> Average Cal sensor Temp.: 96.35 Ambient Temp.: 96.53 F
22:46:13.441 -> Average Cal sensor Temp.: 96.57 Ambient Temp.: 96.53 F
22:46:14.995 -> Average Cal sensor Temp.: 96.77 Ambient Temp.: 96.53 F
22:46:16.557 -> Average Cal sensor Temp.: 96.83 Ambient Temp.: 96.53 F
22:46:18.114 -> Average Cal sensor Temp.: 96.89 Ambient Temp.: 96.53 F
22:46:19.672 -> Average Cal sensor Temp.: 96.93 Ambient Temp.: 96.53 F
22:46:21.229 -> Average Cal sensor Temp.: 97.05 Ambient Temp.: 96.49 F
22:46:22.783 -> Average Cal sensor Temp.: 97.04 Ambient Temp.: 96.49 F
22:46:24.342 -> Average Cal sensor Temp.: 97.14 Ambient Temp.: 96.53 F
22:46:25.899 -> Average Cal sensor Temp.: 96.94 Ambient Temp.: 96.53 F
22:46:27.455 -> Average Cal sensor Temp.: 96.98 Ambient Temp.: 96.42 F
22:46:29.005 -> Average Cal sensor Temp.: 97.03 Ambient Temp.: 96.49 F
22:46:30.556 -> Average Cal sensor Temp.: 96.98 Ambient Temp.: 96.53 F
22:46:32.110 -> Average Cal sensor Temp.: 96.93 Ambient Temp.: 96.49 F
22:46:33.695 -> Average Cal sensor Temp.: 96.63 Ambient Temp.: 96.53 F
22:46:35.257 -> Average Cal sensor Temp.: 96.71 Ambient Temp.: 96.53 F
22:46:36.813 -> Average Cal sensor Temp.: 96.53 Ambient Temp.: 96.57 F
22:46:38.369 -> Average Cal sensor Temp.: 96.50 Ambient Temp.: 96.57 F
22:46:39.925 -> Average Cal sensor Temp.: 96.49 Ambient Temp.: 96.49 F
22:46:41.482 -> Average Cal sensor Temp.: 96.80 Ambient Temp.: 96.49 F
22:46:43.037 -> Average Cal sensor Temp.: 96.93 Ambient Temp.: 96.53 F
22:46:44.595 -> Average Cal sensor Temp.: 96.84 Ambient Temp.: 96.49 F
22:46:46.155 -> Average Cal sensor Temp.: 96.99 Ambient Temp.: 96.49 F
22:46:47.710 -> Average Cal sensor Temp.: 97.02 Ambient Temp.: 96.53 F
22:46:49.264 -> Average Cal sensor Temp.: 96.83 Ambient Temp.: 96.53 F
22:46:50.824 -> Average Cal sensor Temp.: 96.98 Ambient Temp.: 96.49 F
22:46:52.378 -> Average Cal sensor Temp.: 96.92 Ambient Temp.: 96.49 F
22:46:53.935 -> Average Cal sensor Temp.: 96.87 Ambient Temp.: 96.49 F
22:46:55.493 -> Average Cal sensor Temp.: 96.93 Ambient Temp.: 96.49 F
22:46:57.049 -> Average Cal sensor Temp.: 97.00 Ambient Temp.: 96.49 F
22:46:58.605 -> Average Cal sensor Temp.: 96.85 Ambient Temp.: 96.53 F
22:47:00.159 -> Average Cal sensor Temp.: 96.84 Ambient Temp.: 96.53 F
22:47:01.717 -> Average Cal sensor Temp.: 96.93 Ambient Temp.: 96.49 F
22:47:03.267 -> Average Cal sensor Temp.: 97.05 Ambient Temp.: 96.53 F
22:47:04.823 -> Average Cal sensor Temp.: 97.05 Ambient Temp.: 96.57 F
22:47:06.380 -> Average Cal sensor Temp.: 96.93 Ambient Temp.: 96.49 F
22:47:07.933 -> Average Cal sensor Temp.: 96.89 Ambient Temp.: 96.53 F
22:47:09.489 -> Average Cal sensor Temp.: 96.84 Ambient Temp.: 96.53 F
22:47:11.047 -> Average Cal sensor Temp.: 96.71 Ambient Temp.: 96.49 F
22:47:12.629 -> Average Cal sensor Temp.: 96.94 Ambient Temp.: 96.53 F
22:47:14.180 -> Average Cal sensor Temp.: 97.06 Ambient Temp.: 96.57 F
22:47:15.741 -> Average Cal sensor Temp.: 97.05 Ambient Temp.: 96.57 F
22:47:17.294 -> Average Cal sensor Temp.: 96.98 Ambient Temp.: 96.53 F
22:47:18.847 -> Average Cal sensor Temp.: 96.95 Ambient Temp.: 96.57 F
22:47:20.392 -> Average Cal sensor Temp.: 96.95 Ambient Temp.: 96.53 F
22:47:21.945 -> Average Cal sensor Temp.: 97.12 Ambient Temp.: 96.53 F
22:47:23.529 -> Average Cal sensor Temp.: 96.96 Ambient Temp.: 96.53 F
22:47:25.055 -> Average Cal sensor Temp.: 96.98 Ambient Temp.: 96.57 F
22:47:26.643 -> Average Cal sensor Temp.: 96.96 Ambient Temp.: 96.53 F
22:47:28.199 -> Average Cal sensor Temp.: 97.00 Ambient Temp.: 96.53 F
22:47:29.755 -> Average Cal sensor Temp.: 97.07 Ambient Temp.: 96.57 F
22:47:31.285 -> Average Cal sensor Temp.: 97.10 Ambient Temp.: 96.57 F
22:47:32.842 -> Average Cal sensor Temp.: 97.04 Ambient Temp.: 96.57 F
22:47:34.394 -> Average Cal sensor Temp.: 96.91 Ambient Temp.: 96.57 F
22:47:35.953 -> Average Cal sensor Temp.: 97.09 Ambient Temp.: 96.53 F
22:47:37.511 -> Average Cal sensor Temp.: 97.23 Ambient Temp.: 96.53 F
22:47:39.066 -> Average Cal sensor Temp.: 96.80 Ambient Temp.: 96.53 F
22:47:40.621 -> Average Cal sensor Temp.: 97.00 Ambient Temp.: 96.53 F
22:47:42.176 -> Average Cal sensor Temp.: 97.05 Ambient Temp.: 96.49 F
22:47:43.758 -> Average Cal sensor Temp.: 96.72 Ambient Temp.: 96.53 F
22:47:45.319 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 96.53 F
22:47:46.875 -> Average Cal sensor Temp.: 96.85 Ambient Temp.: 96.57 F
22:47:48.432 -> Average Cal sensor Temp.: 96.57 Ambient Temp.: 96.53 F
22:47:49.986 -> Average Cal sensor Temp.: 96.52 Ambient Temp.: 96.57 F
22:47:51.541 -> Average Cal sensor Temp.: 96.87 Ambient Temp.: 96.53 F
22:47:53.092 -> Average Cal sensor Temp.: 97.13 Ambient Temp.: 96.53 F
22:47:54.651 -> Average Cal sensor Temp.: 97.09 Ambient Temp.: 96.57 F
22:47:56.206 -> Average Cal sensor Temp.: 97.30 Ambient Temp.: 96.60 F
22:47:57.767 -> Average Cal sensor Temp.: 97.26 Ambient Temp.: 96.64 F
22:47:59.322 -> Average Cal sensor Temp.: 97.35 Ambient Temp.: 96.57 F
22:48:00.851 -> Average Cal sensor Temp.: 97.15 Ambient Temp.: 96.53 F
22:48:02.412 -> Average Cal sensor Temp.: 96.64 Ambient Temp.: 96.49 F
22:48:03.971 -> Average Cal sensor Temp.: 96.86 Ambient Temp.: 96.53 F
22:48:05.531 -> Average Cal sensor Temp.: 96.96 Ambient Temp.: 96.57 F
22:48:07.087 -> Average Cal sensor Temp.: 96.70 Ambient Temp.: 96.53 F
22:48:08.643 -> Average Cal sensor Temp.: 96.53 Ambient Temp.: 96.53 F
22:48:10.203 -> Average Cal sensor Temp.: 96.96 Ambient Temp.: 96.53 F
22:48:11.757 -> Average Cal sensor Temp.: 97.00 Ambient Temp.: 96.53 F
22:48:13.312 -> Average Cal sensor Temp.: 97.12 Ambient Temp.: 96.57 F
22:48:14.867 -> Average Cal sensor Temp.: 96.88 Ambient Temp.: 96.49 F
22:48:16.422 -> Average Cal sensor Temp.: 96.95 Ambient Temp.: 96.57 F
22:48:17.974 -> Average Cal sensor Temp.: 97.01 Ambient Temp.: 96.53 F
22:48:19.535 -> Average Cal sensor Temp.: 97.00 Ambient Temp.: 96.53 F
22:48:21.093 -> Average Cal sensor Temp.: 97.08 Ambient Temp.: 96.53 F
22:48:22.642 -> Average Cal sensor Temp.: 97.16 Ambient Temp.: 96.57 F
22:48:24.198 -> Average Cal sensor Temp.: 97.48 Ambient Temp.: 96.53 F
22:48:25.760 -> Average Cal sensor Temp.: 97.25 Ambient Temp.: 96.60 F
22:48:27.314 -> Average Cal sensor Temp.: 97.53 Ambient Temp.: 96.57 F
22:48:28.871 -> Average Cal sensor Temp.: 97.58 Ambient Temp.: 96.57 F
22:48:30.433 -> Average Cal sensor Temp.: 97.47 Ambient Temp.: 96.60 F
22:48:31.981 -> Average Cal sensor Temp.: 97.20 Ambient Temp.: 96.60 F
22:48:33.539 -> Average Cal sensor Temp.: 97.36 Ambient Temp.: 96.57 F
22:48:35.090 -> Average Cal sensor Temp.: 97.07 Ambient Temp.: 96.57 F
22:48:36.679 -> Average Cal sensor Temp.: 97.18 Ambient Temp.: 96.60 F
22:48:38.209 -> Average Cal sensor Temp.: 97.32 Ambient Temp.: 96.60 F
22:48:39.762 -> Average Cal sensor Temp.: 97.33 Ambient Temp.: 96.60 F
22:48:41.320 -> Average Cal sensor Temp.: 97.33 Ambient Temp.: 96.53 F
22:48:42.872 -> Average Cal sensor Temp.: 97.27 Ambient Temp.: 96.57 F
22:48:44.424 -> Average Cal sensor Temp.: 97.07 Ambient Temp.: 96.60 F
22:48:46.014 -> Average Cal sensor Temp.: 97.23 Ambient Temp.: 96.57 F
22:48:47.563 -> Average Cal sensor Temp.: 97.25 Ambient Temp.: 96.57 F
22:48:49.111 -> Average Cal sensor Temp.: 97.34 Ambient Temp.: 96.60 F
22:48:50.664 -> Average Cal sensor Temp.: 97.49 Ambient Temp.: 96.60 F
22:48:52.219 -> Average Cal sensor Temp.: 97.50 Ambient Temp.: 96.64 F
22:48:53.770 -> Average Cal sensor Temp.: 97.45 Ambient Temp.: 96.53 F
22:48:55.331 -> Average Cal sensor Temp.: 97.66 Ambient Temp.: 96.60 F
22:48:56.891 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 96.57 F
22:48:58.443 -> Average Cal sensor Temp.: 97.23 Ambient Temp.: 96.64 F
22:48:59.996 -> Average Cal sensor Temp.: 97.16 Ambient Temp.: 96.60 F
22:49:01.550 -> Average Cal sensor Temp.: 97.33 Ambient Temp.: 96.64 F
22:49:03.101 -> Average Cal sensor Temp.: 97.50 Ambient Temp.: 96.64 F
22:49:04.659 -> Average Cal sensor Temp.: 97.37 Ambient Temp.: 96.60 F
22:49:06.213 -> Average Cal sensor Temp.: 97.66 Ambient Temp.: 96.60 F
22:49:07.770 -> Average Cal sensor Temp.: 97.56 Ambient Temp.: 96.64 F
22:49:09.351 -> Average Cal sensor Temp.: 97.17 Ambient Temp.: 96.64 F
22:49:10.907 -> Average Cal sensor Temp.: 96.95 Ambient Temp.: 96.60 F
22:49:12.460 -> Average Cal sensor Temp.: 97.11 Ambient Temp.: 96.64 F
22:49:14.014 -> Average Cal sensor Temp.: 97.06 Ambient Temp.: 96.60 F
22:49:15.575 -> Average Cal sensor Temp.: 96.87 Ambient Temp.: 96.64 F
22:49:17.127 -> Average Cal sensor Temp.: 97.28 Ambient Temp.: 96.57 F
22:49:18.681 -> Average Cal sensor Temp.: 97.09 Ambient Temp.: 96.60 F
22:49:20.237 -> Average Cal sensor Temp.: 97.05 Ambient Temp.: 96.60 F
22:49:21.794 -> Average Cal sensor Temp.: 97.24 Ambient Temp.: 96.60 F
22:49:23.345 -> Average Cal sensor Temp.: 97.19 Ambient Temp.: 96.64 F
22:49:24.903 -> Average Cal sensor Temp.: 97.27 Ambient Temp.: 96.64 F
22:49:26.460 -> Average Cal sensor Temp.: 97.45 Ambient Temp.: 96.60 F
22:49:28.014 -> Average Cal sensor Temp.: 97.30 Ambient Temp.: 96.60 F
22:49:29.574 -> Average Cal sensor Temp.: 97.41 Ambient Temp.: 96.60 F
22:49:31.141 -> Average Cal sensor Temp.: 96.93 Ambient Temp.: 96.64 F
22:49:32.696 -> Average Cal sensor Temp.: 97.11 Ambient Temp.: 96.57 F
22:49:34.252 -> Average Cal sensor Temp.: 97.29 Ambient Temp.: 96.57 F
22:49:35.815 -> Average Cal sensor Temp.: 97.48 Ambient Temp.: 96.64 F
22:49:37.364 -> Average Cal sensor Temp.: 97.33 Ambient Temp.: 96.64 F
22:49:38.922 -> Average Cal sensor Temp.: 96.90 Ambient Temp.: 96.64 F
22:49:40.483 -> Average Cal sensor Temp.: 97.24 Ambient Temp.: 96.57 F
22:49:42.044 -> Average Cal sensor Temp.: 97.43 Ambient Temp.: 96.64 F
22:49:43.598 -> Average Cal sensor Temp.: 97.42 Ambient Temp.: 96.60 F
22:49:45.153 -> Average Cal sensor Temp.: 97.69 Ambient Temp.: 96.64 F
22:49:46.705 -> Average Cal sensor Temp.: 97.50 Ambient Temp.: 96.60 F
22:49:48.255 -> Average Cal sensor Temp.: 97.60 Ambient Temp.: 96.64 F
22:49:49.806 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 96.60 F
22:49:51.367 -> Average Cal sensor Temp.: 97.57 Ambient Temp.: 96.53 F
22:49:52.932 -> Average Cal sensor Temp.: 97.55 Ambient Temp.: 96.60 F
22:49:54.500 -> Average Cal sensor Temp.: 97.42 Ambient Temp.: 96.64 F
22:49:56.028 -> Average Cal sensor Temp.: 97.50 Ambient Temp.: 96.64 F
22:49:57.582 -> Average Cal sensor Temp.: 97.73 Ambient Temp.: 96.64 F
22:49:59.166 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 96.57 F
22:50:00.717 -> Average Cal sensor Temp.: 97.99 Ambient Temp.: 96.60 F
22:50:02.269 -> Average Cal sensor Temp.: 97.95 Ambient Temp.: 96.64 F
22:50:03.824 -> Average Cal sensor Temp.: 97.97 Ambient Temp.: 96.64 F
22:50:05.385 -> Average Cal sensor Temp.: 98.03 Ambient Temp.: 96.64 F
22:50:06.941 -> Average Cal sensor Temp.: 97.83 Ambient Temp.: 96.64 F
22:50:08.495 -> Average Cal sensor Temp.: 97.86 Ambient Temp.: 96.60 F
22:50:10.054 -> Average Cal sensor Temp.: 97.70 Ambient Temp.: 96.67 F
22:50:11.615 -> Average Cal sensor Temp.: 97.77 Ambient Temp.: 96.64 F
22:50:13.166 -> Average Cal sensor Temp.: 98.05 Ambient Temp.: 96.71 F
22:50:14.720 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 96.71 F
22:50:16.277 -> Average Cal sensor Temp.: 97.94 Ambient Temp.: 96.67 F
22:50:17.832 -> Average Cal sensor Temp.: 97.80 Ambient Temp.: 96.64 F
22:50:19.378 -> Average Cal sensor Temp.: 97.68 Ambient Temp.: 96.64 F
22:50:20.938 -> Average Cal sensor Temp.: 97.91 Ambient Temp.: 96.64 F
22:50:22.492 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 96.71 F
22:50:24.079 -> Average Cal sensor Temp.: 97.59 Ambient Temp.: 96.67 F
22:50:25.605 -> Average Cal sensor Temp.: 97.26 Ambient Temp.: 96.71 F
22:50:27.161 -> Average Cal sensor Temp.: 97.30 Ambient Temp.: 96.60 F
22:50:28.747 -> Average Cal sensor Temp.: 97.42 Ambient Temp.: 96.64 F
22:50:30.302 -> Average Cal sensor Temp.: 97.41 Ambient Temp.: 96.67 F
22:50:31.857 -> Average Cal sensor Temp.: 97.30 Ambient Temp.: 96.71 F
22:50:33.411 -> Average Cal sensor Temp.: 97.57 Ambient Temp.: 96.64 F
22:50:34.972 -> Average Cal sensor Temp.: 97.29 Ambient Temp.: 96.67 F
22:50:36.496 -> Average Cal sensor Temp.: 97.23 Ambient Temp.: 96.67 F
22:50:38.053 -> Average Cal sensor Temp.: 97.38 Ambient Temp.: 96.71 F
22:50:39.639 -> Average Cal sensor Temp.: 97.29 Ambient Temp.: 96.71 F
22:50:41.167 -> Average Cal sensor Temp.: 97.37 Ambient Temp.: 96.67 F
22:50:42.718 -> Average Cal sensor Temp.: 97.53 Ambient Temp.: 96.64 F
22:50:44.306 -> Average Cal sensor Temp.: 97.70 Ambient Temp.: 96.67 F
22:50:45.861 -> Average Cal sensor Temp.: 97.25 Ambient Temp.: 96.64 F
22:50:47.416 -> Average Cal sensor Temp.: 97.27 Ambient Temp.: 96.64 F
22:50:48.970 -> Average Cal sensor Temp.: 97.29 Ambient Temp.: 96.64 F
22:50:50.527 -> Average Cal sensor Temp.: 97.62 Ambient Temp.: 96.67 F
22:50:52.085 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 96.67 F
22:50:53.637 -> Average Cal sensor Temp.: 97.70 Ambient Temp.: 96.64 F
22:50:55.192 -> Average Cal sensor Temp.: 97.74 Ambient Temp.: 96.64 F
22:50:56.736 -> Average Cal sensor Temp.: 97.77 Ambient Temp.: 96.60 F
22:50:58.292 -> Average Cal sensor Temp.: 97.43 Ambient Temp.: 96.67 F
22:50:59.846 -> Average Cal sensor Temp.: 97.50 Ambient Temp.: 96.64 F
22:51:01.404 -> Average Cal sensor Temp.: 97.52 Ambient Temp.: 96.64 F
22:51:02.963 -> Average Cal sensor Temp.: 97.57 Ambient Temp.: 96.64 F
22:51:04.516 -> Average Cal sensor Temp.: 97.70 Ambient Temp.: 96.67 F
22:51:06.075 -> Average Cal sensor Temp.: 97.69 Ambient Temp.: 96.64 F
22:51:07.634 -> Average Cal sensor Temp.: 97.35 Ambient Temp.: 96.64 F
22:51:09.185 -> Average Cal sensor Temp.: 97.21 Ambient Temp.: 96.67 F
22:51:10.739 -> Average Cal sensor Temp.: 97.56 Ambient Temp.: 96.64 F
22:51:12.294 -> Average Cal sensor Temp.: 97.74 Ambient Temp.: 96.64 F
22:51:13.847 -> Average Cal sensor Temp.: 97.83 Ambient Temp.: 96.64 F
22:51:15.406 -> Average Cal sensor Temp.: 97.74 Ambient Temp.: 96.64 F
22:51:16.957 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 96.64 F
22:51:18.540 -> Average Cal sensor Temp.: 97.64 Ambient Temp.: 96.64 F
22:51:20.096 -> Average Cal sensor Temp.: 97.70 Ambient Temp.: 96.64 F
22:51:21.656 -> Average Cal sensor Temp.: 97.64 Ambient Temp.: 96.64 F
22:51:23.215 -> Average Cal sensor Temp.: 97.82 Ambient Temp.: 96.60 F
22:51:24.773 -> Average Cal sensor Temp.: 97.19 Ambient Temp.: 96.67 F
22:51:26.327 -> Average Cal sensor Temp.: 97.21 Ambient Temp.: 96.67 F
22:51:27.880 -> Average Cal sensor Temp.: 97.40 Ambient Temp.: 96.64 F
22:51:29.439 -> Average Cal sensor Temp.: 97.52 Ambient Temp.: 96.67 F
22:51:30.989 -> Average Cal sensor Temp.: 97.36 Ambient Temp.: 96.64 F
22:51:32.538 -> Average Cal sensor Temp.: 97.07 Ambient Temp.: 96.67 F
22:51:34.091 -> Average Cal sensor Temp.: 97.02 Ambient Temp.: 96.64 F
22:51:35.648 -> Average Cal sensor Temp.: 97.19 Ambient Temp.: 96.60 F
22:51:37.197 -> Average Cal sensor Temp.: 97.46 Ambient Temp.: 96.60 F
22:51:38.784 -> Average Cal sensor Temp.: 97.43 Ambient Temp.: 96.64 F
22:51:40.339 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 96.60 F
22:51:41.866 -> Average Cal sensor Temp.: 97.59 Ambient Temp.: 96.60 F
22:51:43.453 -> Average Cal sensor Temp.: 97.65 Ambient Temp.: 96.60 F
22:51:45.009 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 96.64 F
22:51:46.561 -> Average Cal sensor Temp.: 97.89 Ambient Temp.: 96.64 F
22:51:48.119 -> Average Cal sensor Temp.: 97.82 Ambient Temp.: 96.67 F
22:51:49.670 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 96.67 F
22:51:51.226 -> Average Cal sensor Temp.: 97.81 Ambient Temp.: 96.64 F
22:51:52.784 -> Average Cal sensor Temp.: 97.69 Ambient Temp.: 96.71 F
22:51:54.337 -> Average Cal sensor Temp.: 97.66 Ambient Temp.: 96.67 F
22:51:55.894 -> Average Cal sensor Temp.: 97.54 Ambient Temp.: 96.67 F
22:51:57.451 -> Average Cal sensor Temp.: 97.33 Ambient Temp.: 96.67 F
22:51:59.004 -> Average Cal sensor Temp.: 96.97 Ambient Temp.: 96.64 F
22:52:00.558 -> Average Cal sensor Temp.: 97.15 Ambient Temp.: 96.67 F
22:52:02.114 -> Average Cal sensor Temp.: 97.37 Ambient Temp.: 96.60 F
22:52:03.663 -> Average Cal sensor Temp.: 97.31 Ambient Temp.: 96.64 F
22:52:05.219 -> Average Cal sensor Temp.: 97.47 Ambient Temp.: 96.64 F
22:52:06.774 -> Average Cal sensor Temp.: 97.49 Ambient Temp.: 96.64 F
22:52:08.360 -> Average Cal sensor Temp.: 97.73 Ambient Temp.: 96.64 F
22:52:09.916 -> Average Cal sensor Temp.: 97.71 Ambient Temp.: 96.60 F
22:52:11.471 -> Average Cal sensor Temp.: 97.83 Ambient Temp.: 96.67 F
22:52:13.023 -> Average Cal sensor Temp.: 97.89 Ambient Temp.: 96.60 F
22:52:14.585 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 96.67 F
22:52:16.143 -> Average Cal sensor Temp.: 97.72 Ambient Temp.: 96.60 F
22:52:17.699 -> Average Cal sensor Temp.: 97.60 Ambient Temp.: 96.64 F
22:52:19.252 -> Average Cal sensor Temp.: 97.74 Ambient Temp.: 96.67 F
22:52:20.807 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 96.71 F
22:52:22.361 -> Average Cal sensor Temp.: 97.80 Ambient Temp.: 96.67 F
22:52:23.915 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 96.67 F
22:52:25.470 -> Average Cal sensor Temp.: 97.56 Ambient Temp.: 96.64 F
22:52:27.024 -> Average Cal sensor Temp.: 97.64 Ambient Temp.: 96.64 F
22:52:28.580 -> Average Cal sensor Temp.: 97.56 Ambient Temp.: 96.71 F
22:52:30.134 -> Average Cal sensor Temp.: 97.80 Ambient Temp.: 96.67 F
22:52:31.692 -> Average Cal sensor Temp.: 97.81 Ambient Temp.: 96.71 F
22:52:33.254 -> Average Cal sensor Temp.: 97.97 Ambient Temp.: 96.71 F
22:52:34.815 -> Average Cal sensor Temp.: 97.84 Ambient Temp.: 96.64 F
22:52:36.369 -> Average Cal sensor Temp.: 97.99 Ambient Temp.: 96.67 F
22:52:37.922 -> Average Cal sensor Temp.: 97.71 Ambient Temp.: 96.67 F
22:52:39.471 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 96.71 F
22:52:41.022 -> Average Cal sensor Temp.: 97.98 Ambient Temp.: 96.67 F
22:52:42.605 -> Average Cal sensor Temp.: 98.02 Ambient Temp.: 96.71 F
22:52:44.159 -> Average Cal sensor Temp.: 97.87 Ambient Temp.: 96.67 F
22:52:45.686 -> Average Cal sensor Temp.: 97.94 Ambient Temp.: 96.71 F
22:52:47.269 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 96.75 F
22:52:48.821 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 96.71 F
22:52:50.378 -> Average Cal sensor Temp.: 97.91 Ambient Temp.: 96.71 F
22:52:51.936 -> Average Cal sensor Temp.: 97.19 Ambient Temp.: 96.75 F
22:52:53.493 -> Average Cal sensor Temp.: 97.45 Ambient Temp.: 96.67 F
22:52:55.045 -> Average Cal sensor Temp.: 97.78 Ambient Temp.: 96.67 F
22:52:56.603 -> Average Cal sensor Temp.: 97.83 Ambient Temp.: 96.75 F
22:52:58.158 -> Average Cal sensor Temp.: 97.83 Ambient Temp.: 96.67 F
22:52:59.715 -> Average Cal sensor Temp.: 97.94 Ambient Temp.: 96.71 F
22:53:01.268 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 96.71 F
22:53:02.826 -> Average Cal sensor Temp.: 97.86 Ambient Temp.: 96.64 F
22:53:04.379 -> Average Cal sensor Temp.: 97.74 Ambient Temp.: 96.67 F
22:53:05.934 -> Average Cal sensor Temp.: 97.63 Ambient Temp.: 96.71 F
22:53:07.484 -> Average Cal sensor Temp.: 97.59 Ambient Temp.: 96.75 F
22:53:09.034 -> Average Cal sensor Temp.: 97.28 Ambient Temp.: 96.71 F
22:53:10.584 -> Average Cal sensor Temp.: 97.54 Ambient Temp.: 96.67 F
22:53:12.139 -> Average Cal sensor Temp.: 97.65 Ambient Temp.: 96.71 F
22:53:13.692 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 96.71 F
22:53:15.246 -> Average Cal sensor Temp.: 97.85 Ambient Temp.: 96.67 F
22:53:16.836 -> Average Cal sensor Temp.: 97.93 Ambient Temp.: 96.71 F
22:53:18.390 -> Average Cal sensor Temp.: 97.74 Ambient Temp.: 96.75 F
22:53:19.940 -> Average Cal sensor Temp.: 96.91 Ambient Temp.: 96.71 F
22:53:21.498 -> Average Cal sensor Temp.: 97.84 Ambient Temp.: 96.67 F
22:53:23.048 -> Average Cal sensor Temp.: 98.17 Ambient Temp.: 96.71 F
22:53:24.606 -> Average Cal sensor Temp.: 97.84 Ambient Temp.: 96.71 F
22:53:26.164 -> Average Cal sensor Temp.: 97.70 Ambient Temp.: 96.75 F
22:53:27.718 -> Average Cal sensor Temp.: 97.89 Ambient Temp.: 96.71 F
22:53:29.273 -> Average Cal sensor Temp.: 97.83 Ambient Temp.: 96.75 F
22:53:30.829 -> Average Cal sensor Temp.: 97.47 Ambient Temp.: 96.67 F
22:53:32.382 -> Average Cal sensor Temp.: 97.54 Ambient Temp.: 96.67 F
22:53:33.939 -> Average Cal sensor Temp.: 97.62 Ambient Temp.: 96.75 F
22:53:35.498 -> Average Cal sensor Temp.: 97.76 Ambient Temp.: 96.67 F
22:53:37.055 -> Average Cal sensor Temp.: 97.71 Ambient Temp.: 96.67 F
22:53:38.606 -> Average Cal sensor Temp.: 97.80 Ambient Temp.: 96.71 F
22:53:40.160 -> Average Cal sensor Temp.: 97.78 Ambient Temp.: 96.75 F
22:53:41.721 -> Average Cal sensor Temp.: 98.13 Ambient Temp.: 96.71 F
22:53:43.274 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 96.71 F
22:53:44.832 -> Average Cal sensor Temp.: 97.96 Ambient Temp.: 96.75 F
22:53:46.391 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 96.71 F
22:53:47.945 -> Average Cal sensor Temp.: 97.64 Ambient Temp.: 96.71 F
22:53:49.502 -> Average Cal sensor Temp.: 97.84 Ambient Temp.: 96.78 F
22:53:51.058 -> Average Cal sensor Temp.: 97.66 Ambient Temp.: 96.71 F
22:53:52.605 -> Average Cal sensor Temp.: 97.77 Ambient Temp.: 96.67 F
22:53:54.154 -> Average Cal sensor Temp.: 97.86 Ambient Temp.: 96.75 F
22:53:55.714 -> Average Cal sensor Temp.: 97.78 Ambient Temp.: 96.75 F
22:53:57.269 -> Average Cal sensor Temp.: 97.75 Ambient Temp.: 96.71 F
22:53:58.821 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 96.71 F
22:54:00.376 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 96.78 F
22:54:01.929 -> Average Cal sensor Temp.: 97.62 Ambient Temp.: 96.75 F
22:54:03.486 -> Average Cal sensor Temp.: 97.57 Ambient Temp.: 96.71 F
22:54:05.044 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 96.78 F
22:54:06.604 -> Average Cal sensor Temp.: 98.13 Ambient Temp.: 96.75 F
22:54:08.162 -> Average Cal sensor Temp.: 98.32 Ambient Temp.: 96.75 F
22:54:09.715 -> Average Cal sensor Temp.: 98.21 Ambient Temp.: 96.75 F
22:54:11.276 -> Average Cal sensor Temp.: 97.97 Ambient Temp.: 96.75 F
22:54:12.827 -> Average Cal sensor Temp.: 97.91 Ambient Temp.: 96.71 F
22:54:14.378 -> Average Cal sensor Temp.: 97.64 Ambient Temp.: 96.71 F
22:54:15.934 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 96.75 F
22:54:17.483 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 96.75 F
22:54:19.068 -> Average Cal sensor Temp.: 98.28 Ambient Temp.: 96.78 F
22:54:20.624 -> Average Cal sensor Temp.: 98.10 Ambient Temp.: 96.75 F
22:54:22.149 -> Average Cal sensor Temp.: 97.91 Ambient Temp.: 96.75 F
22:54:23.737 -> Average Cal sensor Temp.: 97.75 Ambient Temp.: 96.78 F
22:54:25.292 -> Average Cal sensor Temp.: 97.75 Ambient Temp.: 96.75 F
22:54:26.846 -> Average Cal sensor Temp.: 97.62 Ambient Temp.: 96.75 F
22:54:28.370 -> Average Cal sensor Temp.: 97.82 Ambient Temp.: 96.71 F
22:54:29.930 -> Average Cal sensor Temp.: 97.99 Ambient Temp.: 96.75 F
22:54:31.515 -> Average Cal sensor Temp.: 98.18 Ambient Temp.: 96.75 F
22:54:33.039 -> Average Cal sensor Temp.: 98.18 Ambient Temp.: 96.78 F
22:54:34.629 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 96.75 F
22:54:36.185 -> Average Cal sensor Temp.: 98.16 Ambient Temp.: 96.75 F
22:54:37.710 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 96.78 F
22:54:39.298 -> Average Cal sensor Temp.: 98.37 Ambient Temp.: 96.78 F
22:54:40.852 -> Average Cal sensor Temp.: 98.24 Ambient Temp.: 96.78 F
22:54:42.406 -> Average Cal sensor Temp.: 98.21 Ambient Temp.: 96.78 F
22:54:43.955 -> Average Cal sensor Temp.: 98.25 Ambient Temp.: 96.78 F
22:54:45.511 -> Average Cal sensor Temp.: 97.85 Ambient Temp.: 96.75 F
22:54:47.065 -> Average Cal sensor Temp.: 97.69 Ambient Temp.: 96.78 F
22:54:48.625 -> Average Cal sensor Temp.: 98.08 Ambient Temp.: 96.78 F
22:54:50.183 -> Average Cal sensor Temp.: 98.07 Ambient Temp.: 96.75 F
22:54:51.737 -> Average Cal sensor Temp.: 97.70 Ambient Temp.: 96.82 F
22:54:53.294 -> Average Cal sensor Temp.: 97.71 Ambient Temp.: 96.82 F
22:54:54.850 -> Average Cal sensor Temp.: 97.36 Ambient Temp.: 96.78 F
22:54:56.408 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 96.75 F
22:54:57.960 -> Average Cal sensor Temp.: 97.97 Ambient Temp.: 96.75 F
22:54:59.508 -> Average Cal sensor Temp.: 97.68 Ambient Temp.: 96.78 F
22:55:01.064 -> Average Cal sensor Temp.: 97.85 Ambient Temp.: 96.71 F
22:55:02.616 -> Average Cal sensor Temp.: 98.23 Ambient Temp.: 96.78 F
22:55:04.174 -> Average Cal sensor Temp.: 98.29 Ambient Temp.: 96.75 F
22:55:05.735 -> Average Cal sensor Temp.: 98.10 Ambient Temp.: 96.78 F
22:55:07.284 -> Average Cal sensor Temp.: 98.18 Ambient Temp.: 96.78 F
22:55:08.860 -> Average Cal sensor Temp.: 98.36 Ambient Temp.: 96.78 F
22:55:10.406 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 96.78 F
22:55:11.959 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 96.82 F
22:55:13.511 -> Average Cal sensor Temp.: 98.08 Ambient Temp.: 96.75 F
22:55:15.096 -> Average Cal sensor Temp.: 98.26 Ambient Temp.: 96.82 F
22:55:16.651 -> Average Cal sensor Temp.: 98.21 Ambient Temp.: 96.78 F
22:55:18.209 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 96.71 F
22:55:19.766 -> Average Cal sensor Temp.: 98.23 Ambient Temp.: 96.71 F
22:55:21.322 -> Average Cal sensor Temp.: 98.07 Ambient Temp.: 96.67 F
22:55:22.878 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 96.78 F
22:55:24.436 -> Average Cal sensor Temp.: 97.77 Ambient Temp.: 96.82 F
22:55:25.989 -> Average Cal sensor Temp.: 97.99 Ambient Temp.: 96.78 F
22:55:27.542 -> Average Cal sensor Temp.: 98.04 Ambient Temp.: 96.71 F
22:55:29.100 -> Average Cal sensor Temp.: 98.14 Ambient Temp.: 96.75 F
22:55:30.656 -> Average Cal sensor Temp.: 97.86 Ambient Temp.: 96.75 F
22:55:32.208 -> Average Cal sensor Temp.: 98.07 Ambient Temp.: 96.78 F
22:55:33.770 -> Average Cal sensor Temp.: 98.13 Ambient Temp.: 96.78 F
22:55:35.332 -> Average Cal sensor Temp.: 98.32 Ambient Temp.: 96.78 F
22:55:36.886 -> Average Cal sensor Temp.: 98.41 Ambient Temp.: 96.78 F
22:55:38.440 -> Average Cal sensor Temp.: 98.10 Ambient Temp.: 96.75 F
22:55:39.997 -> Average Cal sensor Temp.: 98.37 Ambient Temp.: 96.75 F
22:55:41.556 -> Average Cal sensor Temp.: 98.26 Ambient Temp.: 96.78 F
22:55:43.109 -> Average Cal sensor Temp.: 98.53 Ambient Temp.: 96.78 F
22:55:44.665 -> Average Cal sensor Temp.: 98.55 Ambient Temp.: 96.82 F
22:55:46.223 -> Average Cal sensor Temp.: 98.71 Ambient Temp.: 96.82 F
22:55:47.777 -> Average Cal sensor Temp.: 98.52 Ambient Temp.: 96.82 F
22:55:49.333 -> Average Cal sensor Temp.: 98.54 Ambient Temp.: 96.78 F
22:55:50.892 -> Average Cal sensor Temp.: 98.71 Ambient Temp.: 96.82 F
22:55:52.449 -> Average Cal sensor Temp.: 98.68 Ambient Temp.: 96.89 F
22:55:54.008 -> Average Cal sensor Temp.: 98.24 Ambient Temp.: 96.82 F
22:55:55.573 -> Average Cal sensor Temp.: 98.44 Ambient Temp.: 96.78 F
22:55:57.128 -> Average Cal sensor Temp.: 98.53 Ambient Temp.: 96.82 F
22:55:58.678 -> Average Cal sensor Temp.: 98.34 Ambient Temp.: 96.89 F
22:56:00.233 -> Average Cal sensor Temp.: 98.32 Ambient Temp.: 96.89 F
22:56:01.790 -> Average Cal sensor Temp.: 97.36 Ambient Temp.: 96.78 F
22:56:03.343 -> Average Cal sensor Temp.: 97.74 Ambient Temp.: 96.78 F
22:56:04.895 -> Average Cal sensor Temp.: 97.73 Ambient Temp.: 96.82 F
22:56:06.449 -> Average Cal sensor Temp.: 98.02 Ambient Temp.: 96.78 F
22:56:08.008 -> Average Cal sensor Temp.: 98.18 Ambient Temp.: 96.82 F
22:56:09.565 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 96.78 F
22:56:11.124 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 96.82 F
22:56:12.678 -> Average Cal sensor Temp.: 98.18 Ambient Temp.: 96.75 F
22:56:14.232 -> Average Cal sensor Temp.: 98.17 Ambient Temp.: 96.75 F
22:56:15.788 -> Average Cal sensor Temp.: 97.99 Ambient Temp.: 96.78 F
22:56:17.341 -> Average Cal sensor Temp.: 97.96 Ambient Temp.: 96.78 F
22:56:18.928 -> Average Cal sensor Temp.: 98.07 Ambient Temp.: 96.71 F
22:56:20.478 -> Average Cal sensor Temp.: 98.24 Ambient Temp.: 96.67 F
22:56:22.037 -> Average Cal sensor Temp.: 98.28 Ambient Temp.: 96.78 F
22:56:23.597 -> Average Cal sensor Temp.: 98.21 Ambient Temp.: 96.78 F
22:56:25.155 -> Average Cal sensor Temp.: 98.66 Ambient Temp.: 96.71 F
22:56:26.707 -> Average Cal sensor Temp.: 98.32 Ambient Temp.: 96.75 F
22:56:28.261 -> Average Cal sensor Temp.: 98.37 Ambient Temp.: 96.75 F
22:56:29.822 -> Average Cal sensor Temp.: 98.64 Ambient Temp.: 96.75 F
22:56:31.378 -> Average Cal sensor Temp.: 98.78 Ambient Temp.: 96.78 F
22:56:32.927 -> Average Cal sensor Temp.: 98.86 Ambient Temp.: 96.78 F
22:56:34.484 -> Average Cal sensor Temp.: 98.37 Ambient Temp.: 96.75 F
22:56:36.039 -> Average Cal sensor Temp.: 98.49 Ambient Temp.: 96.71 F
22:56:37.592 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 96.71 F
22:56:39.143 -> Average Cal sensor Temp.: 98.21 Ambient Temp.: 96.71 F
22:56:40.698 -> Average Cal sensor Temp.: 97.99 Ambient Temp.: 96.75 F
22:56:42.260 -> Average Cal sensor Temp.: 98.42 Ambient Temp.: 96.78 F
22:56:43.812 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 96.71 F
22:56:45.369 -> Average Cal sensor Temp.: 98.19 Ambient Temp.: 96.75 F
22:56:46.921 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 96.71 F
22:56:48.484 -> Average Cal sensor Temp.: 98.16 Ambient Temp.: 96.75 F
22:56:50.037 -> Average Cal sensor Temp.: 98.42 Ambient Temp.: 96.78 F
22:56:51.589 -> Average Cal sensor Temp.: 97.97 Ambient Temp.: 96.75 F
22:56:53.148 -> Average Cal sensor Temp.: 97.65 Ambient Temp.: 96.75 F
22:56:54.705 -> Average Cal sensor Temp.: 97.68 Ambient Temp.: 96.75 F
22:56:56.265 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 96.78 F
22:56:57.819 -> Average Cal sensor Temp.: 98.47 Ambient Temp.: 96.67 F
22:56:59.371 -> Average Cal sensor Temp.: 98.14 Ambient Temp.: 96.71 F
22:57:00.926 -> Average Cal sensor Temp.: 98.32 Ambient Temp.: 96.71 F
22:57:02.483 -> Average Cal sensor Temp.: 98.44 Ambient Temp.: 96.71 F
22:57:04.046 -> Average Cal sensor Temp.: 98.55 Ambient Temp.: 96.78 F
22:57:05.604 -> Average Cal sensor Temp.: 98.63 Ambient Temp.: 96.75 F
22:57:07.155 -> Average Cal sensor Temp.: 98.67 Ambient Temp.: 96.75 F
22:57:08.710 -> Average Cal sensor Temp.: 98.78 Ambient Temp.: 96.75 F
22:57:10.267 -> Average Cal sensor Temp.: 98.61 Ambient Temp.: 96.75 F
22:57:11.853 -> Average Cal sensor Temp.: 98.62 Ambient Temp.: 96.78 F
22:57:13.409 -> Average Cal sensor Temp.: 98.57 Ambient Temp.: 96.75 F
22:57:14.964 -> Average Cal sensor Temp.: 98.69 Ambient Temp.: 96.75 F
22:57:16.517 -> Average Cal sensor Temp.: 98.27 Ambient Temp.: 96.75 F
22:57:18.073 -> Average Cal sensor Temp.: 98.39 Ambient Temp.: 96.75 F
22:57:19.628 -> Average Cal sensor Temp.: 98.24 Ambient Temp.: 96.78 F
22:57:21.174 -> Average Cal sensor Temp.: 98.21 Ambient Temp.: 96.78 F
22:57:22.724 -> Average Cal sensor Temp.: 98.46 Ambient Temp.: 96.75 F
22:57:24.283 -> Average Cal sensor Temp.: 98.32 Ambient Temp.: 96.78 F
22:57:25.842 -> Average Cal sensor Temp.: 98.61 Ambient Temp.: 96.75 F
22:57:27.398 -> Average Cal sensor Temp.: 98.78 Ambient Temp.: 96.78 F
22:57:28.950 -> Average Cal sensor Temp.: 98.24 Ambient Temp.: 96.75 F
22:57:30.540 -> Average Cal sensor Temp.: 98.61 Ambient Temp.: 96.78 F
22:57:32.096 -> Average Cal sensor Temp.: 98.18 Ambient Temp.: 96.75 F
22:57:33.652 -> Average Cal sensor Temp.: 98.23 Ambient Temp.: 96.75 F
22:57:35.177 -> Average Cal sensor Temp.: 98.62 Ambient Temp.: 96.75 F
22:57:36.763 -> Average Cal sensor Temp.: 98.74 Ambient Temp.: 96.75 F
22:57:38.315 -> Average Cal sensor Temp.: 98.51 Ambient Temp.: 96.78 F
22:57:39.870 -> Average Cal sensor Temp.: 98.74 Ambient Temp.: 96.78 F
22:57:41.425 -> Average Cal sensor Temp.: 98.46 Ambient Temp.: 96.78 F
22:57:42.973 -> Average Cal sensor Temp.: 98.66 Ambient Temp.: 96.78 F
22:57:44.528 -> Average Cal sensor Temp.: 98.79 Ambient Temp.: 96.78 F
22:57:46.082 -> Average Cal sensor Temp.: 98.65 Ambient Temp.: 96.78 F
22:57:47.638 -> Average Cal sensor Temp.: 98.68 Ambient Temp.: 96.78 F
22:57:49.195 -> Average Cal sensor Temp.: 98.90 Ambient Temp.: 96.78 F
22:57:50.780 -> Average Cal sensor Temp.: 98.50 Ambient Temp.: 96.78 F
22:57:52.337 -> Average Cal sensor Temp.: 98.65 Ambient Temp.: 96.75 F
22:57:53.894 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 96.71 F
22:57:55.454 -> Average Cal sensor Temp.: 98.19 Ambient Temp.: 96.78 F
22:57:57.009 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 96.78 F
22:57:58.560 -> Average Cal sensor Temp.: 98.31 Ambient Temp.: 96.78 F
22:58:00.109 -> Average Cal sensor Temp.: 98.47 Ambient Temp.: 96.78 F
22:58:01.668 -> Average Cal sensor Temp.: 98.53 Ambient Temp.: 96.71 F
22:58:03.219 -> Average Cal sensor Temp.: 98.21 Ambient Temp.: 96.75 F
22:58:04.777 -> Average Cal sensor Temp.: 98.57 Ambient Temp.: 96.78 F
22:58:06.330 -> Average Cal sensor Temp.: 98.77 Ambient Temp.: 96.75 F
22:58:07.884 -> Average Cal sensor Temp.: 98.75 Ambient Temp.: 96.78 F
22:58:09.441 -> Average Cal sensor Temp.: 98.80 Ambient Temp.: 96.78 F
22:58:11.001 -> Average Cal sensor Temp.: 98.73 Ambient Temp.: 96.82 F
22:58:12.552 -> Average Cal sensor Temp.: 98.87 Ambient Temp.: 96.82 F
22:58:14.105 -> Average Cal sensor Temp.: 98.37 Ambient Temp.: 96.78 F
22:58:15.668 -> Average Cal sensor Temp.: 98.70 Ambient Temp.: 96.78 F
22:58:17.225 -> Average Cal sensor Temp.: 98.67 Ambient Temp.: 96.78 F
22:58:18.776 -> Average Cal sensor Temp.: 98.70 Ambient Temp.: 96.78 F
22:58:20.334 -> Average Cal sensor Temp.: 98.64 Ambient Temp.: 96.78 F
22:58:21.923 -> Average Cal sensor Temp.: 98.62 Ambient Temp.: 96.82 F
22:58:23.476 -> Average Cal sensor Temp.: 98.79 Ambient Temp.: 96.82 F
22:58:25.037 -> Average Cal sensor Temp.: 98.68 Ambient Temp.: 96.82 F
22:58:26.591 -> Average Cal sensor Temp.: 98.73 Ambient Temp.: 96.82 F
22:58:28.140 -> Average Cal sensor Temp.: 98.78 Ambient Temp.: 96.78 F
22:58:29.696 -> Average Cal sensor Temp.: 98.79 Ambient Temp.: 96.89 F
22:58:31.256 -> Average Cal sensor Temp.: 98.66 Ambient Temp.: 96.82 F
22:58:32.806 -> Average Cal sensor Temp.: 99.00 Ambient Temp.: 96.78 F
22:58:34.390 -> Average Cal sensor Temp.: 98.96 Ambient Temp.: 96.82 F
22:58:35.941 -> Average Cal sensor Temp.: 98.53 Ambient Temp.: 96.89 F
22:58:37.492 -> Average Cal sensor Temp.: 98.74 Ambient Temp.: 96.82 F
22:58:39.044 -> Average Cal sensor Temp.: 98.99 Ambient Temp.: 96.82 F
22:58:40.606 -> Average Cal sensor Temp.: 98.82 Ambient Temp.: 96.78 F
22:58:42.157 -> Average Cal sensor Temp.: 99.01 Ambient Temp.: 96.82 F
22:58:43.708 -> Average Cal sensor Temp.: 98.83 Ambient Temp.: 96.82 F
22:58:45.293 -> Average Cal sensor Temp.: 99.13 Ambient Temp.: 96.82 F
22:58:46.846 -> Average Cal sensor Temp.: 99.11 Ambient Temp.: 96.82 F
22:58:48.402 -> Average Cal sensor Temp.: 98.81 Ambient Temp.: 96.89 F
22:58:49.961 -> Average Cal sensor Temp.: 98.52 Ambient Temp.: 96.78 F
22:58:51.517 -> Average Cal sensor Temp.: 98.62 Ambient Temp.: 96.78 F
22:58:53.072 -> Average Cal sensor Temp.: 98.75 Ambient Temp.: 96.78 F
22:58:54.623 -> Average Cal sensor Temp.: 98.47 Ambient Temp.: 96.89 F
22:58:56.177 -> Average Cal sensor Temp.: 98.34 Ambient Temp.: 96.82 F
22:58:57.726 -> Average Cal sensor Temp.: 97.62 Ambient Temp.: 96.82 F
22:58:59.279 -> Average Cal sensor Temp.: 98.25 Ambient Temp.: 96.82 F
22:59:00.838 -> Average Cal sensor Temp.: 98.33 Ambient Temp.: 96.78 F
22:59:02.393 -> Average Cal sensor Temp.: 98.65 Ambient Temp.: 96.78 F
22:59:03.979 -> Average Cal sensor Temp.: 98.57 Ambient Temp.: 96.78 F
22:59:05.535 -> Average Cal sensor Temp.: 98.87 Ambient Temp.: 96.82 F
22:59:07.087 -> Average Cal sensor Temp.: 99.07 Ambient Temp.: 96.78 F
22:59:08.639 -> Average Cal sensor Temp.: 99.00 Ambient Temp.: 96.82 F
22:59:10.188 -> Average Cal sensor Temp.: 99.19 Ambient Temp.: 96.82 F
22:59:11.742 -> Average Cal sensor Temp.: 98.81 Ambient Temp.: 96.78 F
22:59:13.300 -> Average Cal sensor Temp.: 98.79 Ambient Temp.: 96.82 F
22:59:14.849 -> Average Cal sensor Temp.: 99.01 Ambient Temp.: 96.82 F
22:59:16.439 -> Average Cal sensor Temp.: 98.82 Ambient Temp.: 96.89 F
22:59:17.997 -> Average Cal sensor Temp.: 99.13 Ambient Temp.: 96.78 F
22:59:19.550 -> Average Cal sensor Temp.: 99.18 Ambient Temp.: 96.93 F
22:59:21.107 -> Average Cal sensor Temp.: 99.34 Ambient Temp.: 96.89 F
22:59:22.657 -> Average Cal sensor Temp.: 99.09 Ambient Temp.: 96.78 F
22:59:24.210 -> Average Cal sensor Temp.: 98.89 Ambient Temp.: 96.89 F
22:59:25.767 -> Average Cal sensor Temp.: 98.81 Ambient Temp.: 96.82 F
22:59:27.315 -> Average Cal sensor Temp.: 98.87 Ambient Temp.: 96.82 F
22:59:28.899 -> Average Cal sensor Temp.: 99.19 Ambient Temp.: 96.93 F
22:59:30.452 -> Average Cal sensor Temp.: 98.73 Ambient Temp.: 96.82 F
22:59:32.006 -> Average Cal sensor Temp.: 98.72 Ambient Temp.: 96.89 F
22:59:33.559 -> Average Cal sensor Temp.: 98.93 Ambient Temp.: 96.89 F
22:59:35.118 -> Average Cal sensor Temp.: 99.00 Ambient Temp.: 96.89 F
22:59:36.667 -> Average Cal sensor Temp.: 98.77 Ambient Temp.: 96.89 F
22:59:38.220 -> Average Cal sensor Temp.: 98.53 Ambient Temp.: 96.89 F
22:59:39.777 -> Average Cal sensor Temp.: 98.86 Ambient Temp.: 96.89 F
22:59:41.363 -> Average Cal sensor Temp.: 98.90 Ambient Temp.: 96.89 F
22:59:42.914 -> Average Cal sensor Temp.: 98.96 Ambient Temp.: 96.93 F
22:59:44.466 -> Average Cal sensor Temp.: 98.68 Ambient Temp.: 96.89 F
22:59:46.019 -> Average Cal sensor Temp.: 99.13 Ambient Temp.: 96.82 F
22:59:47.568 -> Average Cal sensor Temp.: 99.21 Ambient Temp.: 96.89 F
22:59:49.155 -> Average Cal sensor Temp.: 99.23 Ambient Temp.: 96.82 F
22:59:50.712 -> Average Cal sensor Temp.: 98.81 Ambient Temp.: 96.82 F
22:59:52.260 -> Average Cal sensor Temp.: 99.08 Ambient Temp.: 96.89 F
22:59:53.813 -> Average Cal sensor Temp.: 99.33 Ambient Temp.: 96.93 F
22:59:55.365 -> Average Cal sensor Temp.: 99.31 Ambient Temp.: 96.89 F
22:59:56.914 -> Average Cal sensor Temp.: 99.17 Ambient Temp.: 96.89 F
22:59:58.469 -> Average Cal sensor Temp.: 99.21 Ambient Temp.: 96.89 F
23:00:00.056 -> Average Cal sensor Temp.: 99.03 Ambient Temp.: 96.89 F
23:00:01.612 -> Average Cal sensor Temp.: 99.01 Ambient Temp.: 96.82 F
23:00:03.170 -> Average Cal sensor Temp.: 99.01 Ambient Temp.: 96.93 F
23:00:04.723 -> Average Cal sensor Temp.: 98.88 Ambient Temp.: 96.93 F
23:00:06.277 -> Average Cal sensor Temp.: 99.09 Ambient Temp.: 96.82 F
23:00:07.832 -> Average Cal sensor Temp.: 98.90 Ambient Temp.: 96.89 F
23:00:09.392 -> Average Cal sensor Temp.: 99.06 Ambient Temp.: 96.93 F
23:00:10.944 -> Average Cal sensor Temp.: 99.18 Ambient Temp.: 96.93 F
23:00:12.500 -> Average Cal sensor Temp.: 99.02 Ambient Temp.: 96.93 F
23:00:14.053 -> Average Cal sensor Temp.: 98.85 Ambient Temp.: 96.89 F
23:00:15.613 -> Average Cal sensor Temp.: 99.18 Ambient Temp.: 96.89 F
23:00:17.169 -> Average Cal sensor Temp.: 98.98 Ambient Temp.: 96.89 F
23:00:18.725 -> Average Cal sensor Temp.: 98.83 Ambient Temp.: 96.89 F
23:00:20.281 -> Average Cal sensor Temp.: 98.90 Ambient Temp.: 96.93 F
23:00:21.839 -> Average Cal sensor Temp.: 99.12 Ambient Temp.: 96.96 F
23:00:23.396 -> Average Cal sensor Temp.: 99.41 Ambient Temp.: 96.93 F
23:00:24.956 -> Average Cal sensor Temp.: 99.44 Ambient Temp.: 96.89 F
23:00:26.540 -> Average Cal sensor Temp.: 99.62 Ambient Temp.: 96.96 F
23:00:28.091 -> Average Cal sensor Temp.: 99.44 Ambient Temp.: 96.93 F
23:00:29.642 -> Average Cal sensor Temp.: 99.14 Ambient Temp.: 96.96 F
23:00:31.198 -> Average Cal sensor Temp.: 99.03 Ambient Temp.: 96.89 F
23:00:32.776 -> Average Cal sensor Temp.: 99.06 Ambient Temp.: 96.89 F
23:00:34.326 -> Average Cal sensor Temp.: 99.18 Ambient Temp.: 96.96 F
23:00:35.883 -> Average Cal sensor Temp.: 99.36 Ambient Temp.: 96.93 F
23:00:37.441 -> Average Cal sensor Temp.: 99.21 Ambient Temp.: 96.96 F
23:00:38.993 -> Average Cal sensor Temp.: 98.88 Ambient Temp.: 96.96 F
23:00:40.549 -> Average Cal sensor Temp.: 99.22 Ambient Temp.: 96.93 F
23:00:42.107 -> Average Cal sensor Temp.: 98.98 Ambient Temp.: 96.96 F
23:00:43.657 -> Average Cal sensor Temp.: 98.62 Ambient Temp.: 96.96 F
23:00:45.247 -> Average Cal sensor Temp.: 98.66 Ambient Temp.: 96.96 F
23:00:46.799 -> Average Cal sensor Temp.: 98.14 Ambient Temp.: 97.00 F
23:00:48.351 -> Average Cal sensor Temp.: 98.19 Ambient Temp.: 96.93 F
23:00:49.909 -> Average Cal sensor Temp.: 98.67 Ambient Temp.: 96.93 F
23:00:51.466 -> Average Cal sensor Temp.: 98.66 Ambient Temp.: 96.93 F
23:00:53.020 -> Average Cal sensor Temp.: 98.57 Ambient Temp.: 96.89 F
23:00:54.571 -> Average Cal sensor Temp.: 99.07 Ambient Temp.: 96.93 F
23:00:56.124 -> Average Cal sensor Temp.: 99.13 Ambient Temp.: 96.93 F
23:00:57.678 -> Average Cal sensor Temp.: 99.03 Ambient Temp.: 96.93 F
23:00:59.235 -> Average Cal sensor Temp.: 99.31 Ambient Temp.: 96.89 F
23:01:00.792 -> Average Cal sensor Temp.: 99.33 Ambient Temp.: 96.89 F
23:01:02.349 -> Average Cal sensor Temp.: 99.28 Ambient Temp.: 96.89 F
23:01:03.935 -> Average Cal sensor Temp.: 99.36 Ambient Temp.: 96.93 F
23:01:05.493 -> Average Cal sensor Temp.: 98.93 Ambient Temp.: 96.93 F
23:01:07.046 -> Average Cal sensor Temp.: 99.38 Ambient Temp.: 96.93 F
23:01:08.601 -> Average Cal sensor Temp.: 99.50 Ambient Temp.: 96.93 F
23:01:10.156 -> Average Cal sensor Temp.: 99.65 Ambient Temp.: 96.93 F
23:01:11.713 -> Average Cal sensor Temp.: 99.51 Ambient Temp.: 96.93 F
23:01:13.268 -> Average Cal sensor Temp.: 99.16 Ambient Temp.: 96.96 F
23:01:14.822 -> Average Cal sensor Temp.: 99.14 Ambient Temp.: 97.00 F
23:01:16.379 -> Average Cal sensor Temp.: 99.19 Ambient Temp.: 96.96 F
23:01:17.966 -> Average Cal sensor Temp.: 99.36 Ambient Temp.: 96.93 F
23:01:19.525 -> Average Cal sensor Temp.: 99.27 Ambient Temp.: 96.96 F
23:01:21.054 -> Average Cal sensor Temp.: 99.29 Ambient Temp.: 96.96 F
23:01:22.608 -> Average Cal sensor Temp.: 99.38 Ambient Temp.: 96.96 F
23:01:24.167 -> Average Cal sensor Temp.: 99.04 Ambient Temp.: 96.96 F
23:01:25.759 -> Average Cal sensor Temp.: 99.02 Ambient Temp.: 96.96 F
23:01:27.310 -> Average Cal sensor Temp.: 98.94 Ambient Temp.: 96.96 F
23:01:28.859 -> Average Cal sensor Temp.: 99.18 Ambient Temp.: 96.93 F
23:01:30.414 -> Average Cal sensor Temp.: 98.58 Ambient Temp.: 96.93 F
23:01:31.972 -> Average Cal sensor Temp.: 99.28 Ambient Temp.: 96.96 F
23:01:33.524 -> Average Cal sensor Temp.: 99.22 Ambient Temp.: 96.96 F
23:01:35.081 -> Average Cal sensor Temp.: 99.19 Ambient Temp.: 97.00 F
23:01:36.634 -> Average Cal sensor Temp.: 99.18 Ambient Temp.: 96.96 F
23:01:38.219 -> Average Cal sensor Temp.: 99.34 Ambient Temp.: 96.93 F
23:01:39.774 -> Average Cal sensor Temp.: 99.47 Ambient Temp.: 96.93 F
23:01:41.334 -> Average Cal sensor Temp.: 99.64 Ambient Temp.: 96.96 F
23:01:42.880 -> Average Cal sensor Temp.: 99.64 Ambient Temp.: 96.96 F
23:01:44.435 -> Average Cal sensor Temp.: 99.46 Ambient Temp.: 96.96 F
23:01:46.022 -> Average Cal sensor Temp.: 99.62 Ambient Temp.: 97.00 F
23:01:47.574 -> Average Cal sensor Temp.: 99.01 Ambient Temp.: 97.00 F
23:01:49.125 -> Average Cal sensor Temp.: 99.14 Ambient Temp.: 97.00 F
23:01:50.683 -> Average Cal sensor Temp.: 99.38 Ambient Temp.: 96.96 F
23:01:52.238 -> Average Cal sensor Temp.: 99.35 Ambient Temp.: 97.00 F
23:01:53.792 -> Average Cal sensor Temp.: 99.19 Ambient Temp.: 96.96 F
23:01:55.350 -> Average Cal sensor Temp.: 99.30 Ambient Temp.: 96.96 F
23:01:56.899 -> Average Cal sensor Temp.: 98.99 Ambient Temp.: 97.03 F
23:01:58.487 -> Average Cal sensor Temp.: 99.03 Ambient Temp.: 96.96 F
23:02:00.040 -> Average Cal sensor Temp.: 99.21 Ambient Temp.: 96.96 F
23:02:01.601 -> Average Cal sensor Temp.: 99.39 Ambient Temp.: 96.96 F
23:02:03.157 -> Average Cal sensor Temp.: 99.67 Ambient Temp.: 96.96 F
23:02:04.713 -> Average Cal sensor Temp.: 99.65 Ambient Temp.: 97.00 F
23:02:06.265 -> Average Cal sensor Temp.: 99.33 Ambient Temp.: 96.96 F
23:02:07.820 -> Average Cal sensor Temp.: 99.30 Ambient Temp.: 97.00 F
23:02:09.379 -> Average Cal sensor Temp.: 99.45 Ambient Temp.: 96.96 F
23:02:10.937 -> Average Cal sensor Temp.: 99.48 Ambient Temp.: 96.93 F
23:02:12.495 -> Average Cal sensor Temp.: 99.65 Ambient Temp.: 96.96 F
23:02:14.051 -> Average Cal sensor Temp.: 99.45 Ambient Temp.: 96.96 F
23:02:15.607 -> Average Cal sensor Temp.: 99.41 Ambient Temp.: 97.00 F
23:02:17.161 -> Average Cal sensor Temp.: 99.37 Ambient Temp.: 96.96 F
23:02:18.721 -> Average Cal sensor Temp.: 99.40 Ambient Temp.: 96.96 F
23:02:20.277 -> Average Cal sensor Temp.: 99.54 Ambient Temp.: 97.00 F
23:02:21.872 -> Average Cal sensor Temp.: 99.12 Ambient Temp.: 97.03 F
23:02:23.395 -> Average Cal sensor Temp.: 99.32 Ambient Temp.: 96.96 F
23:02:24.985 -> Average Cal sensor Temp.: 99.06 Ambient Temp.: 97.00 F
23:02:26.541 -> Average Cal sensor Temp.: 99.00 Ambient Temp.: 96.96 F
23:02:28.090 -> Average Cal sensor Temp.: 99.20 Ambient Temp.: 96.96 F
23:02:29.643 -> Average Cal sensor Temp.: 99.37 Ambient Temp.: 96.96 F
23:02:31.206 -> Average Cal sensor Temp.: 99.59 Ambient Temp.: 97.00 F
23:02:32.761 -> Average Cal sensor Temp.: 99.40 Ambient Temp.: 97.00 F
23:02:34.313 -> Average Cal sensor Temp.: 99.39 Ambient Temp.: 96.96 F
23:02:35.872 -> Average Cal sensor Temp.: 99.54 Ambient Temp.: 96.96 F
23:02:37.456 -> Average Cal sensor Temp.: 99.44 Ambient Temp.: 96.96 F
23:02:39.008 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 97.00 F
23:02:40.565 -> Average Cal sensor Temp.: 99.33 Ambient Temp.: 97.00 F
23:02:42.116 -> Average Cal sensor Temp.: 99.30 Ambient Temp.: 97.00 F
23:02:43.673 -> Average Cal sensor Temp.: 99.34 Ambient Temp.: 96.96 F
23:02:45.228 -> Average Cal sensor Temp.: 99.30 Ambient Temp.: 96.96 F
23:02:46.783 -> Average Cal sensor Temp.: 99.34 Ambient Temp.: 96.96 F
23:02:48.341 -> Average Cal sensor Temp.: 99.16 Ambient Temp.: 97.00 F
23:02:49.897 -> Average Cal sensor Temp.: 99.43 Ambient Temp.: 96.96 F
23:02:51.454 -> Average Cal sensor Temp.: 99.04 Ambient Temp.: 97.00 F
23:02:53.008 -> Average Cal sensor Temp.: 99.23 Ambient Temp.: 97.03 F
23:02:54.566 -> Average Cal sensor Temp.: 99.13 Ambient Temp.: 96.96 F
23:02:56.127 -> Average Cal sensor Temp.: 98.91 Ambient Temp.: 97.07 F
23:02:57.678 -> Average Cal sensor Temp.: 99.28 Ambient Temp.: 97.03 F
23:02:59.267 -> Average Cal sensor Temp.: 99.49 Ambient Temp.: 97.03 F
23:03:00.825 -> Average Cal sensor Temp.: 99.76 Ambient Temp.: 96.96 F
23:03:02.380 -> Average Cal sensor Temp.: 99.60 Ambient Temp.: 97.07 F
23:03:03.935 -> Average Cal sensor Temp.: 99.45 Ambient Temp.: 97.07 F
23:03:05.490 -> Average Cal sensor Temp.: 99.30 Ambient Temp.: 97.00 F
23:03:07.043 -> Average Cal sensor Temp.: 99.58 Ambient Temp.: 97.03 F
23:03:08.602 -> Average Cal sensor Temp.: 99.35 Ambient Temp.: 96.96 F
23:03:10.159 -> Average Cal sensor Temp.: 99.40 Ambient Temp.: 97.03 F
23:03:11.721 -> Average Cal sensor Temp.: 99.67 Ambient Temp.: 97.03 F
23:03:13.277 -> Average Cal sensor Temp.: 99.77 Ambient Temp.: 96.96 F
23:03:14.861 -> Average Cal sensor Temp.: 99.75 Ambient Temp.: 97.07 F
23:03:16.419 -> Average Cal sensor Temp.: 99.64 Ambient Temp.: 97.03 F
23:03:17.976 -> Average Cal sensor Temp.: 99.82 Ambient Temp.: 97.03 F
23:03:19.537 -> Average Cal sensor Temp.: 99.56 Ambient Temp.: 97.07 F
23:03:21.094 -> Average Cal sensor Temp.: 99.79 Ambient Temp.: 97.03 F
23:03:22.651 -> Average Cal sensor Temp.: 99.65 Ambient Temp.: 97.07 F
23:03:24.204 -> Average Cal sensor Temp.: 99.55 Ambient Temp.: 97.07 F
23:03:25.796 -> Average Cal sensor Temp.: 99.31 Ambient Temp.: 97.03 F
23:03:27.352 -> Average Cal sensor Temp.: 99.36 Ambient Temp.: 97.07 F
23:03:28.906 -> Average Cal sensor Temp.: 98.78 Ambient Temp.: 97.07 F
23:03:30.456 -> Average Cal sensor Temp.: 98.92 Ambient Temp.: 97.14 F
23:03:32.012 -> Average Cal sensor Temp.: 99.33 Ambient Temp.: 97.07 F
23:03:33.566 -> Average Cal sensor Temp.: 99.07 Ambient Temp.: 97.07 F
23:03:35.124 -> Average Cal sensor Temp.: 99.49 Ambient Temp.: 97.00 F
23:03:36.672 -> Average Cal sensor Temp.: 99.55 Ambient Temp.: 97.07 F
23:03:38.230 -> Average Cal sensor Temp.: 99.35 Ambient Temp.: 97.07 F
23:03:39.784 -> Average Cal sensor Temp.: 99.36 Ambient Temp.: 97.07 F
23:03:41.343 -> Average Cal sensor Temp.: 99.36 Ambient Temp.: 97.03 F
23:03:42.892 -> Average Cal sensor Temp.: 99.45 Ambient Temp.: 97.07 F
23:03:44.478 -> Average Cal sensor Temp.: 99.59 Ambient Temp.: 97.07 F
23:03:46.039 -> Average Cal sensor Temp.: 99.64 Ambient Temp.: 97.07 F
23:03:47.596 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 97.07 F
23:03:49.150 -> Average Cal sensor Temp.: 98.44 Ambient Temp.: 97.07 F
23:03:50.702 -> Average Cal sensor Temp.: 99.01 Ambient Temp.: 97.03 F
23:03:52.251 -> Average Cal sensor Temp.: 99.13 Ambient Temp.: 97.00 F
23:03:53.806 -> Average Cal sensor Temp.: 99.34 Ambient Temp.: 97.07 F
23:03:55.361 -> Average Cal sensor Temp.: 99.37 Ambient Temp.: 97.07 F
23:03:56.919 -> Average Cal sensor Temp.: 99.18 Ambient Temp.: 97.00 F
23:03:58.475 -> Average Cal sensor Temp.: 99.50 Ambient Temp.: 97.03 F
23:04:00.031 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 97.03 F
23:04:01.622 -> Average Cal sensor Temp.: 99.58 Ambient Temp.: 97.07 F
23:04:03.181 -> Average Cal sensor Temp.: 99.53 Ambient Temp.: 97.03 F
23:04:04.738 -> Average Cal sensor Temp.: 99.66 Ambient Temp.: 97.03 F
23:04:06.294 -> Average Cal sensor Temp.: 99.87 Ambient Temp.: 97.03 F
23:04:07.850 -> Average Cal sensor Temp.: 99.42 Ambient Temp.: 97.07 F
23:04:09.408 -> Average Cal sensor Temp.: 99.05 Ambient Temp.: 97.03 F
23:04:10.967 -> Average Cal sensor Temp.: 99.22 Ambient Temp.: 97.07 F
23:04:12.510 -> Average Cal sensor Temp.: 99.18 Ambient Temp.: 97.00 F
23:04:14.068 -> Average Cal sensor Temp.: 99.38 Ambient Temp.: 97.03 F
23:04:15.627 -> Average Cal sensor Temp.: 99.52 Ambient Temp.: 97.03 F
23:04:17.185 -> Average Cal sensor Temp.: 99.23 Ambient Temp.: 97.11 F
23:04:18.742 -> Average Cal sensor Temp.: 99.19 Ambient Temp.: 97.00 F
23:04:20.299 -> Average Cal sensor Temp.: 99.25 Ambient Temp.: 97.03 F
23:04:21.862 -> Average Cal sensor Temp.: 99.44 Ambient Temp.: 97.00 F
23:04:23.415 -> Average Cal sensor Temp.: 99.87 Ambient Temp.: 97.07 F
23:04:24.971 -> Average Cal sensor Temp.: 99.91 Ambient Temp.: 97.03 F
23:04:26.555 -> Average Cal sensor Temp.: 99.94 Ambient Temp.: 97.07 F
23:04:28.114 -> Average Cal sensor Temp.: 99.25 Ambient Temp.: 97.03 F
23:04:29.668 -> Average Cal sensor Temp.: 99.57 Ambient Temp.: 97.03 F
23:04:31.224 -> Average Cal sensor Temp.: 99.71 Ambient Temp.: 97.03 F
23:04:32.774 -> Average Cal sensor Temp.: 99.64 Ambient Temp.: 97.03 F
23:04:34.359 -> Average Cal sensor Temp.: 99.92 Ambient Temp.: 97.03 F
23:04:35.917 -> Average Cal sensor Temp.: 99.79 Ambient Temp.: 97.00 F
23:04:37.465 -> Average Cal sensor Temp.: 99.77 Ambient Temp.: 97.03 F
23:04:39.044 -> Average Cal sensor Temp.: 99.58 Ambient Temp.: 97.07 F
23:04:40.598 -> Average Cal sensor Temp.: 99.84 Ambient Temp.: 97.07 F
23:04:42.155 -> Average Cal sensor Temp.: 99.59 Ambient Temp.: 97.07 F
23:04:43.706 -> Average Cal sensor Temp.: 99.77 Ambient Temp.: 97.03 F
23:04:45.265 -> Average Cal sensor Temp.: 99.68 Ambient Temp.: 97.03 F
23:04:46.855 -> Average Cal sensor Temp.: 99.65 Ambient Temp.: 97.07 F
23:04:48.410 -> Average Cal sensor Temp.: 99.81 Ambient Temp.: 97.03 F
23:04:49.961 -> Average Cal sensor Temp.: 99.40 Ambient Temp.: 97.11 F
23:04:51.517 -> Average Cal sensor Temp.: 99.30 Ambient Temp.: 97.03 F
23:04:53.070 -> Average Cal sensor Temp.: 99.46 Ambient Temp.: 97.03 F
23:04:54.626 -> Average Cal sensor Temp.: 99.53 Ambient Temp.: 97.03 F
23:04:56.213 -> Average Cal sensor Temp.: 99.71 Ambient Temp.: 97.00 F
23:04:57.769 -> Average Cal sensor Temp.: 99.52 Ambient Temp.: 97.03 F
23:04:59.322 -> Average Cal sensor Temp.: 99.58 Ambient Temp.: 97.03 F
23:05:00.885 -> Average Cal sensor Temp.: 99.80 Ambient Temp.: 97.03 F
23:05:02.443 -> Average Cal sensor Temp.: 99.86 Ambient Temp.: 97.07 F
23:05:04.002 -> Average Cal sensor Temp.: 99.95 Ambient Temp.: 97.11 F
23:05:05.562 -> Average Cal sensor Temp.: 100.10 Ambient Temp.: 97.11 F
23:05:07.150 -> Average Cal sensor Temp.: 99.68 Ambient Temp.: 97.03 F
23:05:08.704 -> Average Cal sensor Temp.: 99.99 Ambient Temp.: 97.11 F
23:05:10.266 -> Average Cal sensor Temp.: 100.04 Ambient Temp.: 97.07 F
23:05:11.824 -> Average Cal sensor Temp.: 99.83 Ambient Temp.: 97.11 F
23:05:13.379 -> Average Cal sensor Temp.: 99.84 Ambient Temp.: 97.03 F
23:05:14.936 -> Average Cal sensor Temp.: 100.02 Ambient Temp.: 97.07 F
23:05:16.529 -> Average Cal sensor Temp.: 99.86 Ambient Temp.: 97.11 F
23:05:18.080 -> Average Cal sensor Temp.: 100.08 Ambient Temp.: 97.14 F
23:05:19.633 -> Average Cal sensor Temp.: 100.11 Ambient Temp.: 97.11 F
23:05:21.193 -> Average Cal sensor Temp.: 100.17 Ambient Temp.: 97.07 F
23:05:22.781 -> Average Cal sensor Temp.: 100.22 Ambient Temp.: 97.11 F
23:05:24.339 -> Average Cal sensor Temp.: 99.68 Ambient Temp.: 97.14 F
23:05:25.899 -> Average Cal sensor Temp.: 99.57 Ambient Temp.: 97.14 F
23:05:27.451 -> Average Cal sensor Temp.: 99.80 Ambient Temp.: 97.11 F
23:05:29.007 -> Average Cal sensor Temp.: 99.83 Ambient Temp.: 97.11 F
23:05:30.563 -> Average Cal sensor Temp.: 99.97 Ambient Temp.: 97.11 F
23:05:32.124 -> Average Cal sensor Temp.: 100.17 Ambient Temp.: 97.11 F
23:05:33.710 -> Average Cal sensor Temp.: 99.98 Ambient Temp.: 97.14 F
23:05:35.268 -> Average Cal sensor Temp.: 100.14 Ambient Temp.: 97.11 F
23:05:36.822 -> Average Cal sensor Temp.: 100.05 Ambient Temp.: 97.11 F
23:05:38.374 -> Average Cal sensor Temp.: 99.97 Ambient Temp.: 97.11 F
23:05:39.965 -> Average Cal sensor Temp.: 99.91 Ambient Temp.: 97.11 F
23:05:41.516 -> Average Cal sensor Temp.: 99.60 Ambient Temp.: 97.11 F
23:05:43.068 -> Average Cal sensor Temp.: 99.74 Ambient Temp.: 97.14 F
23:05:44.621 -> Average Cal sensor Temp.: 99.90 Ambient Temp.: 97.11 F
23:05:46.184 -> Average Cal sensor Temp.: 99.72 Ambient Temp.: 97.11 F
23:05:47.768 -> Average Cal sensor Temp.: 99.79 Ambient Temp.: 97.14 F
23:05:49.326 -> Average Cal sensor Temp.: 99.82 Ambient Temp.: 97.14 F
23:05:50.888 -> Average Cal sensor Temp.: 99.77 Ambient Temp.: 97.18 F
23:05:52.448 -> Average Cal sensor Temp.: 99.66 Ambient Temp.: 97.14 F
23:05:54.002 -> Average Cal sensor Temp.: 100.00 Ambient Temp.: 97.14 F
23:05:55.557 -> Average Cal sensor Temp.: 100.00 Ambient Temp.: 97.11 F
23:05:57.147 -> Average Cal sensor Temp.: 100.05 Ambient Temp.: 97.14 F
23:05:58.697 -> Average Cal sensor Temp.: 99.75 Ambient Temp.: 97.11 F
23:06:00.253 -> Average Cal sensor Temp.: 99.85 Ambient Temp.: 97.14 F
23:06:01.814 -> Average Cal sensor Temp.: 99.43 Ambient Temp.: 97.07 F
23:06:03.374 -> Average Cal sensor Temp.: 99.73 Ambient Temp.: 97.11 F
23:06:04.935 -> Average Cal sensor Temp.: 99.80 Ambient Temp.: 97.11 F
23:06:06.483 -> Average Cal sensor Temp.: 99.60 Ambient Temp.: 97.11 F
23:06:08.068 -> Average Cal sensor Temp.: 99.76 Ambient Temp.: 97.03 F
23:06:09.620 -> Average Cal sensor Temp.: 99.95 Ambient Temp.: 97.14 F
23:06:11.169 -> Average Cal sensor Temp.: 99.91 Ambient Temp.: 97.11 F
23:06:12.723 -> Average Cal sensor Temp.: 99.52 Ambient Temp.: 97.14 F
23:06:14.313 -> Average Cal sensor Temp.: 99.70 Ambient Temp.: 97.14 F
23:06:15.872 -> Average Cal sensor Temp.: 99.68 Ambient Temp.: 97.11 F
23:06:17.426 -> Average Cal sensor Temp.: 99.25 Ambient Temp.: 97.07 F
23:06:18.978 -> Average Cal sensor Temp.: 99.32 Ambient Temp.: 97.11 F
23:06:20.534 -> Average Cal sensor Temp.: 99.51 Ambient Temp.: 97.07 F
23:06:22.096 -> Average Cal sensor Temp.: 99.87 Ambient Temp.: 97.14 F
23:06:23.650 -> Average Cal sensor Temp.: 100.09 Ambient Temp.: 97.11 F
23:06:25.214 -> Average Cal sensor Temp.: 99.90 Ambient Temp.: 97.11 F
23:06:26.767 -> Average Cal sensor Temp.: 99.94 Ambient Temp.: 97.11 F
23:06:28.355 -> Average Cal sensor Temp.: 99.98 Ambient Temp.: 97.14 F
23:06:29.909 -> Average Cal sensor Temp.: 99.89 Ambient Temp.: 97.18 F
23:06:31.467 -> Average Cal sensor Temp.: 99.70 Ambient Temp.: 97.14 F
23:06:33.022 -> Average Cal sensor Temp.: 99.76 Ambient Temp.: 97.14 F
23:06:34.604 -> Average Cal sensor Temp.: 99.84 Ambient Temp.: 97.11 F
23:06:36.164 -> Average Cal sensor Temp.: 100.09 Ambient Temp.: 97.18 F
23:06:37.720 -> Average Cal sensor Temp.: 100.14 Ambient Temp.: 97.14 F
23:06:39.278 -> Average Cal sensor Temp.: 99.82 Ambient Temp.: 97.18 F
23:06:40.834 -> Average Cal sensor Temp.: 99.46 Ambient Temp.: 97.14 F
23:06:42.387 -> Average Cal sensor Temp.: 99.73 Ambient Temp.: 97.11 F
23:06:43.975 -> Average Cal sensor Temp.: 99.53 Ambient Temp.: 97.14 F
23:06:45.537 -> Average Cal sensor Temp.: 99.59 Ambient Temp.: 97.11 F
23:06:47.088 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 97.14 F
23:06:48.633 -> Average Cal sensor Temp.: 99.56 Ambient Temp.: 97.11 F
23:06:50.222 -> Average Cal sensor Temp.: 99.51 Ambient Temp.: 97.14 F
23:06:51.773 -> Average Cal sensor Temp.: 99.62 Ambient Temp.: 97.14 F
23:06:53.327 -> Average Cal sensor Temp.: 99.77 Ambient Temp.: 97.14 F
23:06:54.887 -> Average Cal sensor Temp.: 99.75 Ambient Temp.: 97.14 F
23:06:56.434 -> Average Cal sensor Temp.: 99.78 Ambient Temp.: 97.11 F
23:06:58.025 -> Average Cal sensor Temp.: 99.96 Ambient Temp.: 97.11 F
23:06:59.584 -> Average Cal sensor Temp.: 100.02 Ambient Temp.: 97.11 F
23:07:01.139 -> Average Cal sensor Temp.: 99.70 Ambient Temp.: 97.11 F
23:07:02.688 -> Average Cal sensor Temp.: 99.81 Ambient Temp.: 97.11 F
23:07:04.270 -> Average Cal sensor Temp.: 99.92 Ambient Temp.: 97.07 F
23:07:05.823 -> Average Cal sensor Temp.: 99.93 Ambient Temp.: 97.14 F
23:07:07.374 -> Average Cal sensor Temp.: 99.80 Ambient Temp.: 97.14 F
23:07:08.928 -> Average Cal sensor Temp.: 99.96 Ambient Temp.: 97.07 F
23:07:10.518 -> Average Cal sensor Temp.: 99.77 Ambient Temp.: 97.11 F
23:07:12.068 -> Average Cal sensor Temp.: 99.77 Ambient Temp.: 97.11 F
23:07:13.622 -> Average Cal sensor Temp.: 99.95 Ambient Temp.: 97.11 F
23:07:15.177 -> Average Cal sensor Temp.: 100.00 Ambient Temp.: 97.07 F
23:07:16.759 -> Average Cal sensor Temp.: 99.81 Ambient Temp.: 97.14 F
23:07:18.313 -> Average Cal sensor Temp.: 99.89 Ambient Temp.: 97.14 F
23:07:19.865 -> Average Cal sensor Temp.: 99.66 Ambient Temp.: 97.11 F
23:07:21.421 -> Average Cal sensor Temp.: 99.60 Ambient Temp.: 97.14 F
23:07:23.010 -> Average Cal sensor Temp.: 99.47 Ambient Temp.: 97.11 F
23:07:24.559 -> Average Cal sensor Temp.: 99.87 Ambient Temp.: 97.14 F
23:07:26.110 -> Average Cal sensor Temp.: 100.00 Ambient Temp.: 97.11 F
23:07:27.670 -> Average Cal sensor Temp.: 100.08 Ambient Temp.: 97.14 F
23:07:29.237 -> Average Cal sensor Temp.: 100.13 Ambient Temp.: 97.14 F
23:07:30.794 -> Average Cal sensor Temp.: 100.03 Ambient Temp.: 97.14 F
23:07:32.385 -> Average Cal sensor Temp.: 99.89 Ambient Temp.: 97.14 F
23:07:33.940 -> Average Cal sensor Temp.: 100.01 Ambient Temp.: 97.11 F
23:07:35.499 -> Average Cal sensor Temp.: 100.29 Ambient Temp.: 97.14 F
23:07:37.051 -> Average Cal sensor Temp.: 99.94 Ambient Temp.: 97.14 F
23:07:38.628 -> Average Cal sensor Temp.: 99.74 Ambient Temp.: 97.14 F
23:07:40.182 -> Average Cal sensor Temp.: 99.85 Ambient Temp.: 97.14 F
23:07:41.736 -> Average Cal sensor Temp.: 99.73 Ambient Temp.: 97.18 F
23:07:43.328 -> Average Cal sensor Temp.: 99.66 Ambient Temp.: 97.14 F
23:07:44.883 -> Average Cal sensor Temp.: 99.65 Ambient Temp.: 97.14 F
23:07:46.438 -> Average Cal sensor Temp.: 99.40 Ambient Temp.: 97.14 F
23:07:47.988 -> Average Cal sensor Temp.: 99.69 Ambient Temp.: 97.14 F
23:07:49.545 -> Average Cal sensor Temp.: 99.77 Ambient Temp.: 97.14 F
23:07:51.132 -> Average Cal sensor Temp.: 99.48 Ambient Temp.: 97.11 F
23:07:52.684 -> Average Cal sensor Temp.: 99.51 Ambient Temp.: 97.11 F
23:07:54.236 -> Average Cal sensor Temp.: 99.48 Ambient Temp.: 97.07 F
23:07:55.797 -> Average Cal sensor Temp.: 99.43 Ambient Temp.: 97.14 F
23:07:57.351 -> Average Cal sensor Temp.: 99.27 Ambient Temp.: 97.14 F
23:07:58.903 -> Average Cal sensor Temp.: 99.57 Ambient Temp.: 97.11 F
23:08:00.461 -> Average Cal sensor Temp.: 99.61 Ambient Temp.: 97.11 F
23:08:02.014 -> Average Cal sensor Temp.: 99.59 Ambient Temp.: 97.14 F
23:08:03.568 -> Average Cal sensor Temp.: 99.76 Ambient Temp.: 97.14 F
23:08:05.129 -> Average Cal sensor Temp.: 99.79 Ambient Temp.: 97.11 F
23:08:06.721 -> Average Cal sensor Temp.: 99.90 Ambient Temp.: 97.11 F
23:08:08.277 -> Average Cal sensor Temp.: 99.91 Ambient Temp.: 97.14 F
23:08:09.830 -> Average Cal sensor Temp.: 99.79 Ambient Temp.: 97.07 F
23:08:11.388 -> Average Cal sensor Temp.: 99.96 Ambient Temp.: 97.11 F
23:08:12.943 -> Average Cal sensor Temp.: 99.81 Ambient Temp.: 97.14 F
23:08:14.499 -> Average Cal sensor Temp.: 100.08 Ambient Temp.: 97.14 F
23:08:16.092 -> Average Cal sensor Temp.: 99.93 Ambient Temp.: 97.14 F
23:08:17.647 -> Average Cal sensor Temp.: 100.04 Ambient Temp.: 97.14 F
23:08:19.203 -> Average Cal sensor Temp.: 100.00 Ambient Temp.: 97.11 F
23:08:20.756 -> Average Cal sensor Temp.: 100.08 Ambient Temp.: 97.14 F
23:08:22.346 -> Average Cal sensor Temp.: 99.55 Ambient Temp.: 97.11 F
23:08:23.901 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 97.11 F
23:08:25.464 -> Average Cal sensor Temp.: 99.66 Ambient Temp.: 97.07 F
23:08:27.010 -> Average Cal sensor Temp.: 99.60 Ambient Temp.: 97.14 F
23:08:28.593 -> Average Cal sensor Temp.: 99.68 Ambient Temp.: 97.11 F
23:08:30.149 -> Average Cal sensor Temp.: 99.99 Ambient Temp.: 97.11 F
23:08:31.707 -> Average Cal sensor Temp.: 100.12 Ambient Temp.: 97.14 F
23:08:33.262 -> Average Cal sensor Temp.: 100.10 Ambient Temp.: 97.11 F
23:08:34.815 -> Average Cal sensor Temp.: 100.00 Ambient Temp.: 97.18 F
23:08:36.402 -> Average Cal sensor Temp.: 99.91 Ambient Temp.: 97.11 F
23:08:37.956 -> Average Cal sensor Temp.: 100.01 Ambient Temp.: 97.14 F
23:08:39.507 -> Average Cal sensor Temp.: 99.83 Ambient Temp.: 97.18 F
23:08:41.099 -> Average Cal sensor Temp.: 100.02 Ambient Temp.: 97.18 F
23:08:42.655 -> Average Cal sensor Temp.: 100.15 Ambient Temp.: 97.14 F
23:08:44.210 -> Average Cal sensor Temp.: 100.26 Ambient Temp.: 97.11 F
23:08:45.763 -> Average Cal sensor Temp.: 100.28 Ambient Temp.: 97.11 F
23:08:47.346 -> Average Cal sensor Temp.: 100.28 Ambient Temp.: 97.14 F
23:08:48.897 -> Average Cal sensor Temp.: 100.22 Ambient Temp.: 97.11 F
23:08:50.452 -> Average Cal sensor Temp.: 100.36 Ambient Temp.: 97.14 F
23:08:52.047 -> Average Cal sensor Temp.: 99.93 Ambient Temp.: 97.14 F
23:08:53.601 -> Average Cal sensor Temp.: 99.87 Ambient Temp.: 97.11 F
23:08:55.157 -> Average Cal sensor Temp.: 100.12 Ambient Temp.: 97.14 F
23:08:56.713 -> Average Cal sensor Temp.: 100.08 Ambient Temp.: 97.11 F
23:08:58.267 -> Average Cal sensor Temp.: 100.15 Ambient Temp.: 97.11 F
23:08:59.855 -> Average Cal sensor Temp.: 99.98 Ambient Temp.: 97.11 F
23:09:01.413 -> Average Cal sensor Temp.: 100.09 Ambient Temp.: 97.14 F
23:09:02.969 -> Average Cal sensor Temp.: 100.07 Ambient Temp.: 97.14 F
23:09:04.554 -> Average Cal sensor Temp.: 100.17 Ambient Temp.: 97.14 F
23:09:06.110 -> Average Cal sensor Temp.: 100.22 Ambient Temp.: 97.18 F
23:09:07.670 -> Average Cal sensor Temp.: 100.18 Ambient Temp.: 97.14 F
23:09:09.224 -> Average Cal sensor Temp.: 99.29 Ambient Temp.: 97.14 F
23:09:10.780 -> Average Cal sensor Temp.: 99.20 Ambient Temp.: 97.14 F
23:09:12.336 -> Average Cal sensor Temp.: 99.50 Ambient Temp.: 97.14 F
23:09:13.894 -> Average Cal sensor Temp.: 99.97 Ambient Temp.: 97.14 F
23:09:15.450 -> Average Cal sensor Temp.: 100.23 Ambient Temp.: 97.11 F
23:09:17.037 -> Average Cal sensor Temp.: 100.39 Ambient Temp.: 97.11 F
23:09:18.592 -> Average Cal sensor Temp.: 100.46 Ambient Temp.: 97.14 F
23:09:20.172 -> Average Cal sensor Temp.: 100.19 Ambient Temp.: 97.07 F
23:09:21.725 -> Average Cal sensor Temp.: 100.40 Ambient Temp.: 97.14 F
23:09:23.277 -> Average Cal sensor Temp.: 100.07 Ambient Temp.: 97.11 F
23:09:24.841 -> Average Cal sensor Temp.: 100.14 Ambient Temp.: 97.11 F
23:09:26.429 -> Average Cal sensor Temp.: 100.17 Ambient Temp.: 97.11 F
23:09:27.977 -> Average Cal sensor Temp.: 100.09 Ambient Temp.: 97.14 F
23:09:29.534 -> Average Cal sensor Temp.: 100.02 Ambient Temp.: 97.07 F
23:09:31.125 -> Average Cal sensor Temp.: 100.30 Ambient Temp.: 97.11 F
23:09:32.674 -> Average Cal sensor Temp.: 100.51 Ambient Temp.: 97.07 F
23:09:34.228 -> Average Cal sensor Temp.: 100.08 Ambient Temp.: 97.11 F
23:09:35.821 -> Average Cal sensor Temp.: 100.16 Ambient Temp.: 97.11 F
23:09:37.372 -> Average Cal sensor Temp.: 100.08 Ambient Temp.: 97.11 F
23:09:38.918 -> Average Cal sensor Temp.: 100.31 Ambient Temp.: 97.14 F
23:09:40.510 -> Average Cal sensor Temp.: 100.28 Ambient Temp.: 97.14 F
23:09:42.069 -> Average Cal sensor Temp.: 100.36 Ambient Temp.: 97.07 F
23:09:43.624 -> Average Cal sensor Temp.: 100.40 Ambient Temp.: 97.14 F
23:09:45.208 -> Average Cal sensor Temp.: 100.12 Ambient Temp.: 97.14 F
23:09:46.761 -> Average Cal sensor Temp.: 100.04 Ambient Temp.: 97.14 F
23:09:48.315 -> Average Cal sensor Temp.: 99.59 Ambient Temp.: 97.11 F
23:09:49.864 -> Average Cal sensor Temp.: 99.81 Ambient Temp.: 97.14 F
23:09:51.454 -> Average Cal sensor Temp.: 100.09 Ambient Temp.: 97.14 F
23:09:53.015 -> Average Cal sensor Temp.: 100.24 Ambient Temp.: 97.14 F
23:09:54.570 -> Average Cal sensor Temp.: 100.35 Ambient Temp.: 97.14 F
23:09:56.127 -> Average Cal sensor Temp.: 100.49 Ambient Temp.: 97.14 F
23:09:57.717 -> Average Cal sensor Temp.: 100.52 Ambient Temp.: 97.11 F
23:09:59.270 -> Average Cal sensor Temp.: 100.66 Ambient Temp.: 97.18 F
23:10:00.831 -> Average Cal sensor Temp.: 100.50 Ambient Temp.: 97.14 F
23:10:02.386 -> Average Cal sensor Temp.: 100.33 Ambient Temp.: 97.14 F
23:10:03.975 -> Average Cal sensor Temp.: 100.42 Ambient Temp.: 97.18 F
23:10:05.527 -> Average Cal sensor Temp.: 100.16 Ambient Temp.: 97.14 F
23:10:07.080 -> Average Cal sensor Temp.: 99.99 Ambient Temp.: 97.14 F
23:10:08.664 -> Average Cal sensor Temp.: 100.02 Ambient Temp.: 97.07 F
23:10:10.219 -> Average Cal sensor Temp.: 99.93 Ambient Temp.: 97.14 F
23:10:11.771 -> Average Cal sensor Temp.: 99.88 Ambient Temp.: 97.18 F
23:10:13.328 -> Average Cal sensor Temp.: 100.03 Ambient Temp.: 97.11 F
23:10:14.911 -> Average Cal sensor Temp.: 99.90 Ambient Temp.: 97.11 F
23:10:16.468 -> Average Cal sensor Temp.: 100.22 Ambient Temp.: 97.18 F
23:10:18.022 -> Average Cal sensor Temp.: 99.70 Ambient Temp.: 97.11 F
23:10:19.610 -> Average Cal sensor Temp.: 100.13 Ambient Temp.: 97.11 F
23:10:21.172 -> Average Cal sensor Temp.: 100.04 Ambient Temp.: 97.14 F
23:10:22.724 -> Average Cal sensor Temp.: 100.09 Ambient Temp.: 97.14 F
23:10:23.931 -> Ambient Temp.: 97.11 F
23:10:24.516 -> Ambient Temp.: 97.11 F
23:10:26.040 -> Average Cal sensor Temp.: 99.96 Ambient Temp.: 97.18 F
23:10:27.630 -> Average Cal sensor Temp.: 100.54 Ambient Temp.: 97.14 F
23:10:29.183 -> Average Cal sensor Temp.: 100.36 Ambient Temp.: 97.14 F
23:10:30.734 -> Average Cal sensor Temp.: 100.21 Ambient Temp.: 97.07 F
23:10:32.325 -> Average Cal sensor Temp.: 100.04 Ambient Temp.: 97.18 F
23:10:33.877 -> Average Cal sensor Temp.: 100.12 Ambient Temp.: 97.18 F
23:10:35.432 -> Average Cal sensor Temp.: 99.50 Ambient Temp.: 97.14 F
23:10:36.984 -> Average Cal sensor Temp.: 99.58 Ambient Temp.: 97.14 F
23:10:38.566 -> Average Cal sensor Temp.: 99.97 Ambient Temp.: 97.18 F
23:10:40.120 -> Average Cal sensor Temp.: 99.76 Ambient Temp.: 97.18 F
23:10:41.676 -> Average Cal sensor Temp.: 99.55 Ambient Temp.: 97.18 F
23:10:43.231 -> Average Cal sensor Temp.: 100.02 Ambient Temp.: 97.18 F
23:10:44.825 -> Average Cal sensor Temp.: 99.27 Ambient Temp.: 97.11 F
23:10:46.382 -> Average Cal sensor Temp.: 99.97 Ambient Temp.: 97.11 F
23:10:47.932 -> Average Cal sensor Temp.: 99.86 Ambient Temp.: 97.14 F
23:10:49.492 -> Average Cal sensor Temp.: 100.21 Ambient Temp.: 97.14 F
23:10:51.048 -> Average Cal sensor Temp.: 100.42 Ambient Temp.: 97.11 F
23:10:52.599 -> Average Cal sensor Temp.: 100.42 Ambient Temp.: 97.14 F
23:10:54.192 -> Average Cal sensor Temp.: 99.85 Ambient Temp.: 97.11 F
23:10:55.743 -> Average Cal sensor Temp.: 99.87 Ambient Temp.: 97.11 F
23:10:57.298 -> Average Cal sensor Temp.: 100.48 Ambient Temp.: 97.18 F
23:10:58.854 -> Average Cal sensor Temp.: 100.29 Ambient Temp.: 97.11 F
23:11:00.450 -> Average Cal sensor Temp.: 100.68 Ambient Temp.: 97.11 F
23:11:02.010 -> Average Cal sensor Temp.: 100.50 Ambient Temp.: 97.14 F
23:11:03.571 -> Average Cal sensor Temp.: 100.44 Ambient Temp.: 97.03 F
23:11:05.127 -> Average Cal sensor Temp.: 100.57 Ambient Temp.: 97.14 F
23:11:06.677 -> Average Cal sensor Temp.: 100.47 Ambient Temp.: 97.18 F
23:11:08.267 -> Average Cal sensor Temp.: 100.60 Ambient Temp.: 97.14 F
23:11:09.823 -> Average Cal sensor Temp.: 100.60 Ambient Temp.: 97.18 F
23:11:11.379 -> Average Cal sensor Temp.: 100.43 Ambient Temp.: 97.14 F
23:11:12.941 -> Average Cal sensor Temp.: 100.63 Ambient Temp.: 97.11 F
23:11:14.532 -> Average Cal sensor Temp.: 100.40 Ambient Temp.: 97.11 F
23:11:16.090 -> Average Cal sensor Temp.: 100.60 Ambient Temp.: 97.11 F
23:11:17.642 -> Average Cal sensor Temp.: 100.55 Ambient Temp.: 97.14 F
23:11:19.194 -> Average Cal sensor Temp.: 100.33 Ambient Temp.: 97.14 F
23:11:20.783 -> Average Cal sensor Temp.: 100.44 Ambient Temp.: 97.18 F
23:11:22.333 -> Average Cal sensor Temp.: 100.49 Ambient Temp.: 97.14 F
23:11:23.922 -> Average Cal sensor Temp.: 100.55 Ambient Temp.: 97.21 F
23:11:25.486 -> Average Cal sensor Temp.: 100.46 Ambient Temp.: 97.14 F
23:11:27.039 -> Average Cal sensor Temp.: 99.89 Ambient Temp.: 97.18 F
23:11:28.591 -> Average Cal sensor Temp.: 99.71 Ambient Temp.: 97.14 F
23:11:30.143 -> Average Cal sensor Temp.: 100.03 Ambient Temp.: 97.18 F
23:11:31.735 -> Average Cal sensor Temp.: 100.44 Ambient Temp.: 97.18 F
23:11:33.295 -> Average Cal sensor Temp.: 100.57 Ambient Temp.: 97.18 F
23:11:34.855 -> Average Cal sensor Temp.: 100.85 Ambient Temp.: 97.18 F
23:11:36.410 -> Average Cal sensor Temp.: 100.65 Ambient Temp.: 97.14 F
23:11:37.996 -> Average Cal sensor Temp.: 100.65 Ambient Temp.: 97.14 F
23:11:39.549 -> Average Cal sensor Temp.: 100.63 Ambient Temp.: 97.21 F
23:11:41.107 -> Average Cal sensor Temp.: 100.77 Ambient Temp.: 97.18 F
23:11:42.659 -> Average Cal sensor Temp.: 99.88 Ambient Temp.: 97.14 F
23:11:44.249 -> Average Cal sensor Temp.: 100.17 Ambient Temp.: 97.21 F
23:11:45.801 -> Average Cal sensor Temp.: 99.96 Ambient Temp.: 97.18 F
23:11:47.358 -> Average Cal sensor Temp.: 99.98 Ambient Temp.: 97.14 F
23:11:48.915 -> Average Cal sensor Temp.: 100.11 Ambient Temp.: 97.18 F
23:11:50.503 -> Average Cal sensor Temp.: 100.15 Ambient Temp.: 97.21 F
23:11:52.053 -> Average Cal sensor Temp.: 100.01 Ambient Temp.: 97.18 F
23:11:53.606 -> Average Cal sensor Temp.: 100.30 Ambient Temp.: 97.21 F
23:11:55.166 -> Average Cal sensor Temp.: 100.28 Ambient Temp.: 97.14 F
23:11:56.754 -> Average Cal sensor Temp.: 100.48 Ambient Temp.: 97.18 F
23:11:58.310 -> Average Cal sensor Temp.: 100.20 Ambient Temp.: 97.21 F
23:11:59.862 -> Average Cal sensor Temp.: 99.69 Ambient Temp.: 97.18 F
23:12:01.457 -> Average Cal sensor Temp.: 100.01 Ambient Temp.: 97.18 F
23:12:03.012 -> Average Cal sensor Temp.: 99.99 Ambient Temp.: 97.21 F
23:12:04.567 -> Average Cal sensor Temp.: 99.56 Ambient Temp.: 97.18 F
23:12:06.127 -> Average Cal sensor Temp.: 99.95 Ambient Temp.: 97.18 F
23:12:07.685 -> Average Cal sensor Temp.: 99.78 Ambient Temp.: 97.14 F
23:12:09.244 -> Average Cal sensor Temp.: 99.88 Ambient Temp.: 97.21 F
23:12:10.831 -> Average Cal sensor Temp.: 100.34 Ambient Temp.: 97.18 F
23:12:12.379 -> Average Cal sensor Temp.: 100.33 Ambient Temp.: 97.14 F
23:12:13.934 -> Average Cal sensor Temp.: 100.23 Ambient Temp.: 97.21 F
23:12:15.493 -> Average Cal sensor Temp.: 99.95 Ambient Temp.: 97.18 F
23:12:17.082 -> Average Cal sensor Temp.: 100.49 Ambient Temp.: 97.14 F
23:12:18.637 -> Average Cal sensor Temp.: 100.62 Ambient Temp.: 97.14 F
23:12:20.195 -> Average Cal sensor Temp.: 99.96 Ambient Temp.: 97.14 F
23:12:21.755 -> Average Cal sensor Temp.: 99.54 Ambient Temp.: 97.14 F
23:12:23.336 -> Average Cal sensor Temp.: 99.94 Ambient Temp.: 97.14 F
23:12:24.890 -> Average Cal sensor Temp.: 100.09 Ambient Temp.: 97.18 F
23:12:26.438 -> Average Cal sensor Temp.: 99.52 Ambient Temp.: 97.18 F
23:12:28.027 -> Average Cal sensor Temp.: 99.82 Ambient Temp.: 97.14 F
23:12:29.572 -> Average Cal sensor Temp.: 100.03 Ambient Temp.: 97.18 F
23:12:31.130 -> Average Cal sensor Temp.: 100.13 Ambient Temp.: 97.14 F
23:12:32.719 -> Average Cal sensor Temp.: 99.72 Ambient Temp.: 97.14 F
23:12:34.272 -> Average Cal sensor Temp.: 100.13 Ambient Temp.: 97.18 F
23:12:35.830 -> Average Cal sensor Temp.: 100.00 Ambient Temp.: 97.14 F
23:12:37.408 -> Average Cal sensor Temp.: 99.92 Ambient Temp.: 97.11 F
23:12:38.956 -> Average Cal sensor Temp.: 100.14 Ambient Temp.: 97.11 F
23:12:40.512 -> Average Cal sensor Temp.: 100.25 Ambient Temp.: 97.18 F
23:12:42.100 -> Average Cal sensor Temp.: 100.10 Ambient Temp.: 97.18 F
23:12:43.650 -> Average Cal sensor Temp.: 100.18 Ambient Temp.: 97.14 F
23:12:45.205 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 97.07 F
23:12:46.789 -> Average Cal sensor Temp.: 99.64 Ambient Temp.: 97.14 F
23:12:48.343 -> Average Cal sensor Temp.: 100.23 Ambient Temp.: 97.07 F
23:12:49.896 -> Average Cal sensor Temp.: 100.35 Ambient Temp.: 97.11 F
23:12:51.447 -> Average Cal sensor Temp.: 100.31 Ambient Temp.: 97.11 F
23:12:53.034 -> Average Cal sensor Temp.: 100.44 Ambient Temp.: 97.11 F
23:12:54.590 -> Average Cal sensor Temp.: 100.23 Ambient Temp.: 97.11 F
23:12:56.146 -> Average Cal sensor Temp.: 100.36 Ambient Temp.: 97.14 F
23:12:57.735 -> Average Cal sensor Temp.: 100.27 Ambient Temp.: 97.11 F
23:12:59.295 -> Average Cal sensor Temp.: 99.88 Ambient Temp.: 97.11 F
23:13:00.850 -> Average Cal sensor Temp.: 100.12 Ambient Temp.: 97.11 F
23:13:02.399 -> Average Cal sensor Temp.: 100.27 Ambient Temp.: 97.14 F
23:13:03.982 -> Average Cal sensor Temp.: 100.66 Ambient Temp.: 97.11 F
23:13:05.543 -> Average Cal sensor Temp.: 100.79 Ambient Temp.: 97.14 F
23:13:07.093 -> Average Cal sensor Temp.: 100.62 Ambient Temp.: 97.11 F
23:13:08.649 -> Average Cal sensor Temp.: 100.61 Ambient Temp.: 97.11 F
23:13:10.235 -> Average Cal sensor Temp.: 100.85 Ambient Temp.: 97.14 F
23:13:11.789 -> Average Cal sensor Temp.: 100.18 Ambient Temp.: 97.14 F
23:13:13.350 -> Average Cal sensor Temp.: 100.15 Ambient Temp.: 97.14 F
23:13:14.905 -> Average Cal sensor Temp.: 100.49 Ambient Temp.: 97.11 F
23:13:16.497 -> Average Cal sensor Temp.: 100.51 Ambient Temp.: 97.14 F
23:13:18.055 -> Average Cal sensor Temp.: 100.70 Ambient Temp.: 97.18 F
23:13:19.609 -> Average Cal sensor Temp.: 100.58 Ambient Temp.: 97.14 F
23:13:21.165 -> Average Cal sensor Temp.: 100.71 Ambient Temp.: 97.11 F
23:13:22.748 -> Average Cal sensor Temp.: 100.49 Ambient Temp.: 97.14 F
23:13:24.302 -> Average Cal sensor Temp.: 100.51 Ambient Temp.: 97.14 F
23:13:25.857 -> Average Cal sensor Temp.: 100.55 Ambient Temp.: 97.18 F
23:13:27.439 -> Average Cal sensor Temp.: 100.53 Ambient Temp.: 97.14 F
23:13:28.995 -> Average Cal sensor Temp.: 100.81 Ambient Temp.: 97.18 F
23:13:30.556 -> Average Cal sensor Temp.: 100.89 Ambient Temp.: 97.18 F
23:13:32.145 -> Average Cal sensor Temp.: 100.62 Ambient Temp.: 97.29 F
23:13:33.698 -> Average Cal sensor Temp.: 100.61 Ambient Temp.: 97.14 F
23:13:35.251 -> Average Cal sensor Temp.: 100.30 Ambient Temp.: 97.29 F
23:13:36.839 -> Average Cal sensor Temp.: 100.52 Ambient Temp.: 97.18 F
23:13:38.391 -> Average Cal sensor Temp.: 100.54 Ambient Temp.: 97.18 F
23:13:39.948 -> Average Cal sensor Temp.: 100.73 Ambient Temp.: 97.18 F
23:13:41.503 -> Average Cal sensor Temp.: 100.93 Ambient Temp.: 97.29 F
23:13:43.090 -> Average Cal sensor Temp.: 100.51 Ambient Temp.: 97.21 F
23:13:44.640 -> Average Cal sensor Temp.: 100.59 Ambient Temp.: 97.21 F
23:13:46.199 -> Average Cal sensor Temp.: 100.75 Ambient Temp.: 97.21 F
23:13:47.785 -> Average Cal sensor Temp.: 100.83 Ambient Temp.: 97.29 F
23:13:49.340 -> Average Cal sensor Temp.: 100.92 Ambient Temp.: 97.29 F
23:13:50.897 -> Average Cal sensor Temp.: 100.51 Ambient Temp.: 97.29 F
23:13:52.488 -> Average Cal sensor Temp.: 100.67 Ambient Temp.: 97.21 F
23:13:54.040 -> Average Cal sensor Temp.: 100.60 Ambient Temp.: 97.21 F
23:13:55.597 -> Average Cal sensor Temp.: 100.68 Ambient Temp.: 97.21 F
23:13:57.152 -> Average Cal sensor Temp.: 100.85 Ambient Temp.: 97.29 F
23:13:58.736 -> Average Cal sensor Temp.: 100.71 Ambient Temp.: 97.29 F
23:14:00.290 -> Average Cal sensor Temp.: 101.00 Ambient Temp.: 97.29 F
23:14:01.877 -> Average Cal sensor Temp.: 100.25 Ambient Temp.: 97.29 F
23:14:03.432 -> Average Cal sensor Temp.: 100.53 Ambient Temp.: 100.49 F
23:14:04.986 -> Average Cal sensor Temp.: 100.36 Ambient Temp.: 97.29 F
23:14:06.573 -> Average Cal sensor Temp.: 100.44 Ambient Temp.: 97.32 F
23:14:08.130 -> Average Cal sensor Temp.: 100.26 Ambient Temp.: 97.32 F
23:14:09.677 -> Average Cal sensor Temp.: 100.20 Ambient Temp.: 97.29 F
23:14:11.269 -> Average Cal sensor Temp.: 99.73 Ambient Temp.: 97.32 F
23:14:12.818 -> Average Cal sensor Temp.: 100.31 Ambient Temp.: 97.32 F
23:14:14.368 -> Average Cal sensor Temp.: 100.53 Ambient Temp.: 97.32 F
23:14:15.928 -> Average Cal sensor Temp.: 100.40 Ambient Temp.: 97.29 F
23:14:17.510 -> Average Cal sensor Temp.: 100.34 Ambient Temp.: 97.21 F
23:14:19.068 -> Average Cal sensor Temp.: 100.60 Ambient Temp.: 97.21 F
23:14:20.653 -> Average Cal sensor Temp.: 100.59 Ambient Temp.: 97.21 F
23:14:22.211 -> Average Cal sensor Temp.: 100.61 Ambient Temp.: 97.21 F
23:14:23.766 -> Average Cal sensor Temp.: 100.09 Ambient Temp.: 97.29 F
23:14:25.350 -> Average Cal sensor Temp.: 99.83 Ambient Temp.: 97.21 F
23:14:26.903 -> Average Cal sensor Temp.: 100.18 Ambient Temp.: 97.21 F
23:14:28.451 -> Average Cal sensor Temp.: 100.45 Ambient Temp.: 97.21 F
23:14:30.040 -> Average Cal sensor Temp.: 99.76 Ambient Temp.: 97.18 F
23:14:31.596 -> Average Cal sensor Temp.: 100.05 Ambient Temp.: 97.14 F
23:14:33.147 -> Average Cal sensor Temp.: 100.52 Ambient Temp.: 97.21 F
23:14:34.699 -> Average Cal sensor Temp.: 100.78 Ambient Temp.: 97.18 F
23:14:36.292 -> Average Cal sensor Temp.: 100.94 Ambient Temp.: 97.21 F
23:14:37.840 -> Average Cal sensor Temp.: 101.08 Ambient Temp.: 97.21 F
23:14:39.391 -> Average Cal sensor Temp.: 101.02 Ambient Temp.: 97.29 F
23:14:40.979 -> Average Cal sensor Temp.: 100.65 Ambient Temp.: 97.29 F
23:14:42.529 -> Average Cal sensor Temp.: 100.89 Ambient Temp.: 97.29 F
23:14:44.114 -> Average Cal sensor Temp.: 101.01 Ambient Temp.: 97.18 F
23:14:45.675 -> Average Cal sensor Temp.: 101.16 Ambient Temp.: 97.21 F
23:14:47.230 -> Average Cal sensor Temp.: 100.78 Ambient Temp.: 97.18 F
23:14:48.785 -> Average Cal sensor Temp.: 100.93 Ambient Temp.: 97.29 F
23:14:50.371 -> Average Cal sensor Temp.: 101.25 Ambient Temp.: 97.21 F
23:14:51.919 -> Average Cal sensor Temp.: 100.28 Ambient Temp.: 97.21 F
23:14:53.470 -> Average Cal sensor Temp.: 100.57 Ambient Temp.: 97.29 F
23:14:55.062 -> Average Cal sensor Temp.: 100.54 Ambient Temp.: 97.29 F
23:14:56.617 -> Average Cal sensor Temp.: 100.63 Ambient Temp.: 97.29 F
23:14:58.164 -> Average Cal sensor Temp.: 100.31 Ambient Temp.: 97.29 F
23:14:59.754 -> Average Cal sensor Temp.: 100.80 Ambient Temp.: 97.21 F
23:15:01.315 -> Average Cal sensor Temp.: 100.75 Ambient Temp.: 97.29 F
23:15:02.864 -> Average Cal sensor Temp.: 100.59 Ambient Temp.: 97.29 F
23:15:04.452 -> Average Cal sensor Temp.: 100.94 Ambient Temp.: 97.32 F
23:15:06.014 -> Average Cal sensor Temp.: 101.16 Ambient Temp.: 97.29 F
23:15:07.567 -> Average Cal sensor Temp.: 100.26 Ambient Temp.: 97.29 F
23:15:09.153 -> Average Cal sensor Temp.: 100.74 Ambient Temp.: 97.29 F
23:15:10.708 -> Average Cal sensor Temp.: 100.45 Ambient Temp.: 97.21 F
23:15:12.261 -> Average Cal sensor Temp.: 100.60 Ambient Temp.: 97.29 F
23:15:13.844 -> Average Cal sensor Temp.: 100.68 Ambient Temp.: 97.29 F
23:15:15.403 -> Average Cal sensor Temp.: 100.19 Ambient Temp.: 97.29 F
23:15:16.960 -> Average Cal sensor Temp.: 100.44 Ambient Temp.: 97.21 F
23:15:18.546 -> Average Cal sensor Temp.: 100.84 Ambient Temp.: 97.21 F
23:15:20.097 -> Average Cal sensor Temp.: 100.25 Ambient Temp.: 97.21 F
23:15:21.654 -> Average Cal sensor Temp.: 100.11 Ambient Temp.: 97.29 F
23:15:23.210 -> Average Cal sensor Temp.: 100.53 Ambient Temp.: 97.21 F
23:15:24.770 -> Average Cal sensor Temp.: 100.58 Ambient Temp.: 97.21 F
23:15:26.358 -> Average Cal sensor Temp.: 100.64 Ambient Temp.: 97.21 F
23:15:27.912 -> Average Cal sensor Temp.: 100.88 Ambient Temp.: 97.18 F
23:15:29.467 -> Average Cal sensor Temp.: 101.08 Ambient Temp.: 97.21 F
23:15:31.028 -> Average Cal sensor Temp.: 100.99 Ambient Temp.: 97.21 F
23:15:32.620 -> Average Cal sensor Temp.: 100.94 Ambient Temp.: 97.14 F
23:15:34.180 -> Average Cal sensor Temp.: 101.03 Ambient Temp.: 97.21 F
23:15:35.731 -> Average Cal sensor Temp.: 100.99 Ambient Temp.: 97.18 F
23:15:37.318 -> Average Cal sensor Temp.: 101.21 Ambient Temp.: 97.21 F
23:15:38.869 -> Average Cal sensor Temp.: 101.25 Ambient Temp.: 97.21 F
23:15:40.424 -> Average Cal sensor Temp.: 101.02 Ambient Temp.: 97.21 F
23:15:41.979 -> Average Cal sensor Temp.: 101.31 Ambient Temp.: 97.21 F
23:15:43.567 -> Average Cal sensor Temp.: 101.31 Ambient Temp.: 97.21 F
23:15:45.120 -> Average Cal sensor Temp.: 101.16 Ambient Temp.: 97.21 F
23:15:46.670 -> Average Cal sensor Temp.: 100.97 Ambient Temp.: 97.29 F
23:15:48.225 -> Average Cal sensor Temp.: 100.71 Ambient Temp.: 97.29 F
23:15:49.810 -> Average Cal sensor Temp.: 100.58 Ambient Temp.: 97.29 F
23:15:51.366 -> Average Cal sensor Temp.: 100.73 Ambient Temp.: 97.32 F
23:15:52.922 -> Average Cal sensor Temp.: 101.03 Ambient Temp.: 97.29 F
23:15:54.513 -> Average Cal sensor Temp.: 101.00 Ambient Temp.: 97.21 F
23:15:56.065 -> Average Cal sensor Temp.: 100.98 Ambient Temp.: 97.21 F
23:15:57.616 -> Average Cal sensor Temp.: 101.44 Ambient Temp.: 97.21 F
23:15:59.205 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 97.29 F
23:16:00.757 -> Average Cal sensor Temp.: 101.13 Ambient Temp.: 97.21 F
23:16:02.314 -> Average Cal sensor Temp.: 101.28 Ambient Temp.: 97.29 F
23:16:03.861 -> Average Cal sensor Temp.: 101.22 Ambient Temp.: 97.29 F
23:16:05.452 -> Average Cal sensor Temp.: 101.03 Ambient Temp.: 97.36 F
23:16:07.004 -> Average Cal sensor Temp.: 100.94 Ambient Temp.: 97.29 F
23:16:08.556 -> Average Cal sensor Temp.: 100.84 Ambient Temp.: 97.29 F
23:16:10.143 -> Average Cal sensor Temp.: 100.80 Ambient Temp.: 97.36 F
23:16:11.699 -> Average Cal sensor Temp.: 101.04 Ambient Temp.: 97.29 F
23:16:13.284 -> Average Cal sensor Temp.: 101.08 Ambient Temp.: 97.32 F
23:16:14.838 -> Average Cal sensor Temp.: 101.22 Ambient Temp.: 97.32 F
23:16:16.391 -> Average Cal sensor Temp.: 101.20 Ambient Temp.: 97.32 F
23:16:17.947 -> Average Cal sensor Temp.: 100.98 Ambient Temp.: 97.29 F
23:16:19.530 -> Average Cal sensor Temp.: 100.18 Ambient Temp.: 97.29 F
23:16:21.085 -> Average Cal sensor Temp.: 100.63 Ambient Temp.: 97.32 F
23:16:22.639 -> Average Cal sensor Temp.: 100.98 Ambient Temp.: 97.29 F
23:16:24.225 -> Average Cal sensor Temp.: 101.12 Ambient Temp.: 97.36 F
23:16:25.781 -> Average Cal sensor Temp.: 100.83 Ambient Temp.: 97.32 F
23:16:27.334 -> Average Cal sensor Temp.: 100.81 Ambient Temp.: 97.36 F
23:16:28.915 -> Average Cal sensor Temp.: 100.59 Ambient Temp.: 97.36 F
23:16:30.469 -> Average Cal sensor Temp.: 100.63 Ambient Temp.: 97.32 F
23:16:32.055 -> Average Cal sensor Temp.: 100.70 Ambient Temp.: 97.32 F
23:16:33.606 -> Average Cal sensor Temp.: 101.17 Ambient Temp.: 97.32 F
23:16:35.164 -> Average Cal sensor Temp.: 101.32 Ambient Temp.: 97.36 F
23:16:36.749 -> Average Cal sensor Temp.: 101.35 Ambient Temp.: 97.29 F
23:16:38.297 -> Average Cal sensor Temp.: 101.39 Ambient Temp.: 97.32 F
23:16:39.846 -> Average Cal sensor Temp.: 101.21 Ambient Temp.: 97.32 F
23:16:41.403 -> Average Cal sensor Temp.: 101.12 Ambient Temp.: 97.32 F
23:16:42.990 -> Average Cal sensor Temp.: 100.83 Ambient Temp.: 97.36 F
23:16:44.546 -> Average Cal sensor Temp.: 100.65 Ambient Temp.: 97.32 F
23:16:46.110 -> Average Cal sensor Temp.: 99.72 Ambient Temp.: 97.29 F
23:16:47.666 -> Average Cal sensor Temp.: 100.53 Ambient Temp.: 97.29 F
23:16:49.223 -> Average Cal sensor Temp.: 101.18 Ambient Temp.: 97.32 F
23:16:50.811 -> Average Cal sensor Temp.: 101.15 Ambient Temp.: 97.21 F
23:16:52.363 -> Average Cal sensor Temp.: 101.16 Ambient Temp.: 97.29 F
23:16:53.919 -> Average Cal sensor Temp.: 101.12 Ambient Temp.: 97.32 F
23:16:55.480 -> Average Cal sensor Temp.: 100.75 Ambient Temp.: 97.32 F
23:16:57.033 -> Average Cal sensor Temp.: 101.14 Ambient Temp.: 97.29 F
23:16:58.624 -> Average Cal sensor Temp.: 101.25 Ambient Temp.: 97.32 F
23:17:00.181 -> Average Cal sensor Temp.: 101.40 Ambient Temp.: 97.32 F
23:17:01.739 -> Average Cal sensor Temp.: 101.26 Ambient Temp.: 97.36 F
23:17:03.296 -> Average Cal sensor Temp.: 101.15 Ambient Temp.: 97.21 F
23:17:04.851 -> Average Cal sensor Temp.: 101.16 Ambient Temp.: 97.36 F
23:17:06.444 -> Average Cal sensor Temp.: 101.18 Ambient Temp.: 97.29 F
23:17:07.999 -> Average Cal sensor Temp.: 101.34 Ambient Temp.: 97.39 F
23:17:09.544 -> Average Cal sensor Temp.: 100.94 Ambient Temp.: 97.36 F
23:17:11.133 -> Average Cal sensor Temp.: 100.56 Ambient Temp.: 97.32 F
23:17:12.686 -> Average Cal sensor Temp.: 99.70 Ambient Temp.: 97.32 F
23:17:14.241 -> Average Cal sensor Temp.: 100.60 Ambient Temp.: 97.32 F
23:17:15.796 -> Average Cal sensor Temp.: 100.42 Ambient Temp.: 97.29 F
23:17:17.389 -> Average Cal sensor Temp.: 100.44 Ambient Temp.: 97.29 F
23:17:18.949 -> Average Cal sensor Temp.: 100.51 Ambient Temp.: 97.32 F
23:17:20.508 -> Average Cal sensor Temp.: 100.59 Ambient Temp.: 97.32 F
23:17:22.066 -> Average Cal sensor Temp.: 101.04 Ambient Temp.: 97.32 F
23:17:23.622 -> Average Cal sensor Temp.: 101.26 Ambient Temp.: 97.21 F
23:17:25.210 -> Average Cal sensor Temp.: 101.27 Ambient Temp.: 97.29 F
23:17:26.758 -> Average Cal sensor Temp.: 101.52 Ambient Temp.: 97.32 F
23:17:28.317 -> Average Cal sensor Temp.: 101.54 Ambient Temp.: 97.32 F
23:17:29.870 -> Average Cal sensor Temp.: 101.18 Ambient Temp.: 97.36 F
23:17:31.455 -> Average Cal sensor Temp.: 101.32 Ambient Temp.: 97.32 F
23:17:33.008 -> Average Cal sensor Temp.: 101.16 Ambient Temp.: 97.32 F
23:17:34.561 -> Average Cal sensor Temp.: 101.25 Ambient Temp.: 97.32 F
23:17:36.120 -> Average Cal sensor Temp.: 101.24 Ambient Temp.: 97.32 F
23:17:37.706 -> Average Cal sensor Temp.: 101.59 Ambient Temp.: 97.29 F
23:17:39.258 -> Average Cal sensor Temp.: 101.65 Ambient Temp.: 97.36 F
23:17:40.816 -> Average Cal sensor Temp.: 101.21 Ambient Temp.: 97.29 F
23:17:42.405 -> Average Cal sensor Temp.: 100.90 Ambient Temp.: 97.36 F
23:17:43.962 -> Average Cal sensor Temp.: 101.20 Ambient Temp.: 97.36 F
23:17:45.519 -> Average Cal sensor Temp.: 100.72 Ambient Temp.: 97.39 F
23:17:47.070 -> Average Cal sensor Temp.: 100.72 Ambient Temp.: 97.39 F
23:17:48.662 -> Average Cal sensor Temp.: 100.91 Ambient Temp.: 97.39 F
23:17:50.219 -> Average Cal sensor Temp.: 101.13 Ambient Temp.: 97.36 F
23:17:51.776 -> Average Cal sensor Temp.: 101.16 Ambient Temp.: 97.32 F
23:17:53.333 -> Average Cal sensor Temp.: 101.23 Ambient Temp.: 97.39 F
23:17:54.899 -> Average Cal sensor Temp.: 100.80 Ambient Temp.: 97.39 F
23:17:56.463 -> Average Cal sensor Temp.: 101.18 Ambient Temp.: 97.39 F
23:17:58.018 -> Average Cal sensor Temp.: 101.10 Ambient Temp.: 97.36 F
23:17:59.608 -> Average Cal sensor Temp.: 101.14 Ambient Temp.: 97.36 F
23:18:01.161 -> Average Cal sensor Temp.: 101.11 Ambient Temp.: 97.36 F
23:18:02.713 -> Average Cal sensor Temp.: 101.34 Ambient Temp.: 97.36 F
23:18:04.269 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 97.36 F
23:18:05.860 -> Average Cal sensor Temp.: 101.16 Ambient Temp.: 97.39 F
23:18:07.417 -> Average Cal sensor Temp.: 101.04 Ambient Temp.: 97.39 F
23:18:08.968 -> Average Cal sensor Temp.: 101.21 Ambient Temp.: 97.36 F
23:18:10.523 -> Average Cal sensor Temp.: 101.06 Ambient Temp.: 97.39 F
23:18:12.082 -> Average Cal sensor Temp.: 100.88 Ambient Temp.: 97.39 F
23:18:13.672 -> Average Cal sensor Temp.: 101.20 Ambient Temp.: 97.39 F
23:18:15.232 -> Average Cal sensor Temp.: 101.25 Ambient Temp.: 97.36 F
23:18:16.785 -> Average Cal sensor Temp.: 101.01 Ambient Temp.: 97.36 F
23:18:18.338 -> Average Cal sensor Temp.: 101.11 Ambient Temp.: 97.39 F
23:18:19.927 -> Average Cal sensor Temp.: 101.06 Ambient Temp.: 97.39 F
23:18:21.484 -> Average Cal sensor Temp.: 100.77 Ambient Temp.: 97.39 F
23:18:23.045 -> Average Cal sensor Temp.: 101.11 Ambient Temp.: 97.43 F
23:18:24.598 -> Average Cal sensor Temp.: 100.87 Ambient Temp.: 97.36 F
23:18:26.155 -> Average Cal sensor Temp.: 101.46 Ambient Temp.: 97.39 F
23:18:27.742 -> Average Cal sensor Temp.: 101.71 Ambient Temp.: 97.43 F
23:18:29.299 -> Average Cal sensor Temp.: 101.42 Ambient Temp.: 97.39 F
23:18:30.854 -> Average Cal sensor Temp.: 101.41 Ambient Temp.: 97.39 F
23:18:32.437 -> Average Cal sensor Temp.: 101.48 Ambient Temp.: 97.36 F
23:18:33.994 -> Average Cal sensor Temp.: 101.28 Ambient Temp.: 97.39 F
23:18:35.546 -> Average Cal sensor Temp.: 101.68 Ambient Temp.: 97.39 F
23:18:37.101 -> Average Cal sensor Temp.: 101.66 Ambient Temp.: 97.36 F
23:18:38.693 -> Average Cal sensor Temp.: 101.66 Ambient Temp.: 97.43 F
23:18:40.241 -> Average Cal sensor Temp.: 101.69 Ambient Temp.: 97.43 F
23:18:41.799 -> Average Cal sensor Temp.: 101.48 Ambient Temp.: 97.43 F
23:18:43.389 -> Average Cal sensor Temp.: 100.34 Ambient Temp.: 97.39 F
23:18:44.951 -> Average Cal sensor Temp.: 100.71 Ambient Temp.: 97.39 F
23:18:46.509 -> Average Cal sensor Temp.: 100.46 Ambient Temp.: 97.32 F
23:18:48.068 -> Average Cal sensor Temp.: 100.73 Ambient Temp.: 97.36 F
23:18:49.620 -> Average Cal sensor Temp.: 101.10 Ambient Temp.: 97.36 F
23:18:51.209 -> Average Cal sensor Temp.: 101.00 Ambient Temp.: 97.32 F
23:18:52.764 -> Average Cal sensor Temp.: 101.34 Ambient Temp.: 97.36 F
23:18:54.317 -> Average Cal sensor Temp.: 101.17 Ambient Temp.: 97.39 F
23:18:55.879 -> Average Cal sensor Temp.: 101.36 Ambient Temp.: 97.39 F
23:18:57.436 -> Average Cal sensor Temp.: 101.21 Ambient Temp.: 97.39 F
23:18:59.025 -> Average Cal sensor Temp.: 100.97 Ambient Temp.: 97.32 F
23:19:00.577 -> Average Cal sensor Temp.: 101.43 Ambient Temp.: 97.36 F
23:19:02.135 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 97.36 F
23:19:03.721 -> Average Cal sensor Temp.: 100.64 Ambient Temp.: 97.32 F
23:19:05.278 -> Average Cal sensor Temp.: 100.94 Ambient Temp.: 97.39 F
23:19:06.830 -> Average Cal sensor Temp.: 101.14 Ambient Temp.: 97.36 F
23:19:08.385 -> Average Cal sensor Temp.: 100.95 Ambient Temp.: 97.32 F
23:19:09.974 -> Average Cal sensor Temp.: 101.09 Ambient Temp.: 97.32 F
23:19:11.528 -> Average Cal sensor Temp.: 100.83 Ambient Temp.: 97.36 F
23:19:13.081 -> Average Cal sensor Temp.: 100.97 Ambient Temp.: 97.36 F
23:19:14.676 -> Average Cal sensor Temp.: 100.58 Ambient Temp.: 97.36 F
23:19:16.241 -> Average Cal sensor Temp.: 100.98 Ambient Temp.: 97.32 F
23:19:17.796 -> Average Cal sensor Temp.: 101.41 Ambient Temp.: 97.32 F
23:19:19.352 -> Average Cal sensor Temp.: 101.35 Ambient Temp.: 97.39 F
23:19:20.901 -> Average Cal sensor Temp.: 101.63 Ambient Temp.: 97.36 F
23:19:22.485 -> Average Cal sensor Temp.: 101.64 Ambient Temp.: 97.36 F
23:19:24.044 -> Average Cal sensor Temp.: 101.68 Ambient Temp.: 97.36 F
23:19:25.601 -> Average Cal sensor Temp.: 101.55 Ambient Temp.: 97.36 F
23:19:27.156 -> Average Cal sensor Temp.: 101.21 Ambient Temp.: 97.36 F
23:19:28.749 -> Average Cal sensor Temp.: 101.52 Ambient Temp.: 97.36 F
23:19:30.303 -> Average Cal sensor Temp.: 101.74 Ambient Temp.: 97.39 F
23:19:31.861 -> Average Cal sensor Temp.: 100.95 Ambient Temp.: 97.32 F
23:19:33.420 -> Average Cal sensor Temp.: 100.88 Ambient Temp.: 97.32 F
23:19:34.975 -> Average Cal sensor Temp.: 101.27 Ambient Temp.: 97.32 F
23:19:36.561 -> Average Cal sensor Temp.: 101.28 Ambient Temp.: 97.36 F
23:19:38.125 -> Average Cal sensor Temp.: 100.62 Ambient Temp.: 97.39 F
23:19:39.677 -> Average Cal sensor Temp.: 101.26 Ambient Temp.: 97.36 F
23:19:41.234 -> Average Cal sensor Temp.: 101.45 Ambient Temp.: 97.36 F
23:19:42.826 -> Average Cal sensor Temp.: 101.38 Ambient Temp.: 97.36 F
23:19:44.379 -> Average Cal sensor Temp.: 101.51 Ambient Temp.: 97.39 F
23:19:45.937 -> Average Cal sensor Temp.: 101.58 Ambient Temp.: 97.36 F
23:19:47.490 -> Average Cal sensor Temp.: 101.61 Ambient Temp.: 97.36 F
23:19:49.048 -> Average Cal sensor Temp.: 100.90 Ambient Temp.: 97.32 F
23:19:50.640 -> Average Cal sensor Temp.: 101.48 Ambient Temp.: 97.32 F
23:19:52.197 -> Average Cal sensor Temp.: 101.66 Ambient Temp.: 97.32 F
23:19:53.748 -> Average Cal sensor Temp.: 101.62 Ambient Temp.: 97.36 F
23:19:55.306 -> Average Cal sensor Temp.: 101.51 Ambient Temp.: 97.36 F
23:19:56.895 -> Average Cal sensor Temp.: 101.84 Ambient Temp.: 97.32 F
23:19:58.446 -> Average Cal sensor Temp.: 101.44 Ambient Temp.: 97.36 F
23:20:00.002 -> Average Cal sensor Temp.: 101.37 Ambient Temp.: 97.36 F
23:20:01.556 -> Average Cal sensor Temp.: 101.68 Ambient Temp.: 97.39 F
23:20:03.148 -> Average Cal sensor Temp.: 101.39 Ambient Temp.: 97.39 F
23:20:04.700 -> Average Cal sensor Temp.: 101.25 Ambient Temp.: 97.32 F
23:20:06.254 -> Average Cal sensor Temp.: 101.36 Ambient Temp.: 97.36 F
23:20:07.811 -> Average Cal sensor Temp.: 101.60 Ambient Temp.: 97.36 F
23:20:09.400 -> Average Cal sensor Temp.: 101.48 Ambient Temp.: 97.39 F
23:20:10.958 -> Average Cal sensor Temp.: 101.62 Ambient Temp.: 97.39 F
23:20:12.511 -> Average Cal sensor Temp.: 101.62 Ambient Temp.: 97.39 F
23:20:14.069 -> Average Cal sensor Temp.: 101.71 Ambient Temp.: 97.39 F
23:20:15.631 -> Average Cal sensor Temp.: 101.66 Ambient Temp.: 97.39 F
23:20:17.219 -> Average Cal sensor Temp.: 101.16 Ambient Temp.: 97.39 F
23:20:18.772 -> Average Cal sensor Temp.: 101.17 Ambient Temp.: 97.39 F
23:20:20.330 -> Average Cal sensor Temp.: 101.58 Ambient Temp.: 97.39 F
23:20:21.881 -> Average Cal sensor Temp.: 101.49 Ambient Temp.: 97.39 F
23:20:23.465 -> Average Cal sensor Temp.: 101.43 Ambient Temp.: 97.39 F
23:20:25.021 -> Average Cal sensor Temp.: 101.84 Ambient Temp.: 97.43 F
23:20:26.573 -> Average Cal sensor Temp.: 101.90 Ambient Temp.: 97.39 F
23:20:28.160 -> Average Cal sensor Temp.: 101.87 Ambient Temp.: 97.43 F
23:20:29.712 -> Average Cal sensor Temp.: 101.74 Ambient Temp.: 97.43 F
23:20:31.264 -> Average Cal sensor Temp.: 101.57 Ambient Temp.: 97.36 F
23:20:32.848 -> Average Cal sensor Temp.: 101.79 Ambient Temp.: 97.43 F
23:20:34.401 -> Average Cal sensor Temp.: 101.86 Ambient Temp.: 97.47 F
23:20:35.954 -> Average Cal sensor Temp.: 101.75 Ambient Temp.: 97.39 F
23:20:37.539 -> Average Cal sensor Temp.: 101.78 Ambient Temp.: 97.39 F
23:20:39.095 -> Average Cal sensor Temp.: 101.88 Ambient Temp.: 97.43 F
23:20:40.654 -> Average Cal sensor Temp.: 102.01 Ambient Temp.: 97.50 F
23:20:42.211 -> Average Cal sensor Temp.: 101.88 Ambient Temp.: 97.47 F
23:20:43.800 -> Average Cal sensor Temp.: 101.98 Ambient Temp.: 97.47 F
23:20:45.354 -> Average Cal sensor Temp.: 101.80 Ambient Temp.: 97.43 F
23:20:46.912 -> Average Cal sensor Temp.: 101.68 Ambient Temp.: 97.39 F
23:20:48.470 -> Average Cal sensor Temp.: 101.56 Ambient Temp.: 97.50 F
23:20:50.060 -> Average Cal sensor Temp.: 101.29 Ambient Temp.: 97.43 F
23:20:51.617 -> Average Cal sensor Temp.: 101.34 Ambient Temp.: 97.39 F
23:20:53.163 -> Average Cal sensor Temp.: 100.62 Ambient Temp.: 97.39 F
23:20:54.749 -> Average Cal sensor Temp.: 101.35 Ambient Temp.: 97.39 F
23:20:56.302 -> Average Cal sensor Temp.: 101.71 Ambient Temp.: 97.39 F
23:20:57.856 -> Average Cal sensor Temp.: 101.88 Ambient Temp.: 97.39 F
23:20:59.442 -> Average Cal sensor Temp.: 102.12 Ambient Temp.: 97.43 F
23:21:01.001 -> Average Cal sensor Temp.: 101.93 Ambient Temp.: 97.47 F
23:21:02.557 -> Average Cal sensor Temp.: 101.66 Ambient Temp.: 97.39 F
23:21:04.116 -> Average Cal sensor Temp.: 101.98 Ambient Temp.: 97.47 F
23:21:05.674 -> Average Cal sensor Temp.: 101.27 Ambient Temp.: 97.50 F
23:21:07.256 -> Average Cal sensor Temp.: 101.04 Ambient Temp.: 97.43 F
23:21:08.811 -> Average Cal sensor Temp.: 101.27 Ambient Temp.: 97.43 F
23:21:10.365 -> Average Cal sensor Temp.: 100.94 Ambient Temp.: 97.36 F
23:21:11.954 -> Average Cal sensor Temp.: 101.29 Ambient Temp.: 97.43 F
23:21:13.507 -> Average Cal sensor Temp.: 101.70 Ambient Temp.: 97.39 F
23:21:15.062 -> Average Cal sensor Temp.: 101.46 Ambient Temp.: 97.43 F
23:21:16.621 -> Average Cal sensor Temp.: 101.66 Ambient Temp.: 97.36 F
23:21:18.210 -> Average Cal sensor Temp.: 101.67 Ambient Temp.: 97.43 F
23:21:19.764 -> Average Cal sensor Temp.: 101.84 Ambient Temp.: 97.43 F
23:21:21.321 -> Average Cal sensor Temp.: 102.03 Ambient Temp.: 97.43 F
23:21:22.903 -> Average Cal sensor Temp.: 101.98 Ambient Temp.: 97.39 F
23:21:24.460 -> Average Cal sensor Temp.: 102.16 Ambient Temp.: 97.39 F
23:21:26.018 -> Average Cal sensor Temp.: 101.84 Ambient Temp.: 97.43 F
23:21:27.570 -> Average Cal sensor Temp.: 101.91 Ambient Temp.: 97.43 F
23:21:29.156 -> Average Cal sensor Temp.: 101.67 Ambient Temp.: 97.43 F
23:21:30.706 -> Average Cal sensor Temp.: 101.68 Ambient Temp.: 97.39 F
23:21:32.265 -> Average Cal sensor Temp.: 101.74 Ambient Temp.: 97.43 F
23:21:33.817 -> Average Cal sensor Temp.: 102.14 Ambient Temp.: 97.43 F
23:21:35.410 -> Average Cal sensor Temp.: 102.12 Ambient Temp.: 97.39 F
23:21:36.964 -> Average Cal sensor Temp.: 101.98 Ambient Temp.: 97.39 F
23:21:38.517 -> Average Cal sensor Temp.: 101.88 Ambient Temp.: 97.43 F
23:21:40.072 -> Average Cal sensor Temp.: 101.95 Ambient Temp.: 97.39 F
23:21:41.663 -> Average Cal sensor Temp.: 101.85 Ambient Temp.: 97.39 F
23:21:43.220 -> Average Cal sensor Temp.: 101.69 Ambient Temp.: 97.39 F
23:21:44.780 -> Average Cal sensor Temp.: 102.00 Ambient Temp.: 97.39 F
23:21:46.334 -> Average Cal sensor Temp.: 101.89 Ambient Temp.: 97.39 F
23:21:47.914 -> Average Cal sensor Temp.: 102.13 Ambient Temp.: 97.43 F
23:21:49.466 -> Average Cal sensor Temp.: 102.39 Ambient Temp.: 97.47 F
23:21:51.020 -> Average Cal sensor Temp.: 102.07 Ambient Temp.: 97.43 F
23:21:52.606 -> Average Cal sensor Temp.: 102.19 Ambient Temp.: 97.43 F
23:21:54.161 -> Average Cal sensor Temp.: 102.43 Ambient Temp.: 97.47 F
23:21:55.716 -> Average Cal sensor Temp.: 102.05 Ambient Temp.: 97.47 F
23:21:57.302 -> Average Cal sensor Temp.: 101.95 Ambient Temp.: 97.43 F
23:21:58.854 -> Average Cal sensor Temp.: 102.14 Ambient Temp.: 97.39 F
23:22:00.408 -> Average Cal sensor Temp.: 101.99 Ambient Temp.: 97.43 F
23:22:02.000 -> Average Cal sensor Temp.: 102.16 Ambient Temp.: 97.43 F
23:22:03.555 -> Average Cal sensor Temp.: 101.81 Ambient Temp.: 97.39 F
23:22:05.111 -> Average Cal sensor Temp.: 101.83 Ambient Temp.: 97.43 F
23:22:06.663 -> Average Cal sensor Temp.: 102.15 Ambient Temp.: 97.43 F
23:22:08.221 -> Average Cal sensor Temp.: 101.78 Ambient Temp.: 97.43 F
23:22:09.808 -> Average Cal sensor Temp.: 101.94 Ambient Temp.: 97.47 F
23:22:11.366 -> Average Cal sensor Temp.: 101.68 Ambient Temp.: 97.39 F
23:22:12.912 -> Average Cal sensor Temp.: 101.82 Ambient Temp.: 97.43 F
23:22:14.500 -> Average Cal sensor Temp.: 101.80 Ambient Temp.: 97.47 F
23:22:16.054 -> Average Cal sensor Temp.: 101.87 Ambient Temp.: 97.43 F
23:22:17.603 -> Average Cal sensor Temp.: 101.65 Ambient Temp.: 97.47 F
23:22:19.197 -> Average Cal sensor Temp.: 101.62 Ambient Temp.: 97.47 F
23:22:20.752 -> Average Cal sensor Temp.: 101.66 Ambient Temp.: 97.47 F
23:22:22.305 -> Average Cal sensor Temp.: 102.01 Ambient Temp.: 97.47 F
23:22:23.865 -> Average Cal sensor Temp.: 102.10 Ambient Temp.: 97.47 F
23:22:25.426 -> Average Cal sensor Temp.: 102.26 Ambient Temp.: 97.50 F
23:22:27.010 -> Average Cal sensor Temp.: 102.35 Ambient Temp.: 97.47 F
23:22:28.571 -> Average Cal sensor Temp.: 102.60 Ambient Temp.: 97.47 F
23:22:30.125 -> Average Cal sensor Temp.: 102.71 Ambient Temp.: 97.50 F
23:22:31.680 -> Average Cal sensor Temp.: 102.22 Ambient Temp.: 97.50 F
23:22:33.272 -> Average Cal sensor Temp.: 102.25 Ambient Temp.: 97.47 F
23:22:34.832 -> Average Cal sensor Temp.: 102.21 Ambient Temp.: 97.47 F
23:22:36.387 -> Average Cal sensor Temp.: 101.79 Ambient Temp.: 97.47 F
23:22:37.938 -> Average Cal sensor Temp.: 101.66 Ambient Temp.: 97.50 F
23:22:39.493 -> Average Cal sensor Temp.: 101.81 Ambient Temp.: 97.43 F
23:22:41.085 -> Average Cal sensor Temp.: 101.75 Ambient Temp.: 97.50 F
23:22:42.639 -> Average Cal sensor Temp.: 102.09 Ambient Temp.: 97.50 F
23:22:44.190 -> Average Cal sensor Temp.: 102.11 Ambient Temp.: 97.50 F
23:22:45.750 -> Average Cal sensor Temp.: 102.38 Ambient Temp.: 97.50 F
23:22:47.335 -> Average Cal sensor Temp.: 102.11 Ambient Temp.: 97.47 F
23:22:48.897 -> Average Cal sensor Temp.: 101.87 Ambient Temp.: 97.50 F
23:22:50.448 -> Average Cal sensor Temp.: 102.12 Ambient Temp.: 97.50 F
23:22:52.002 -> Average Cal sensor Temp.: 102.04 Ambient Temp.: 97.47 F
23:22:53.583 -> Average Cal sensor Temp.: 102.27 Ambient Temp.: 97.50 F
23:22:55.143 -> Average Cal sensor Temp.: 102.37 Ambient Temp.: 97.50 F
23:22:56.695 -> Average Cal sensor Temp.: 101.89 Ambient Temp.: 97.50 F
23:22:58.275 -> Average Cal sensor Temp.: 102.03 Ambient Temp.: 97.50 F
23:22:59.833 -> Average Cal sensor Temp.: 102.24 Ambient Temp.: 97.50 F
23:23:01.391 -> Average Cal sensor Temp.: 101.68 Ambient Temp.: 97.50 F
23:23:02.946 -> Average Cal sensor Temp.: 101.98 Ambient Temp.: 97.50 F
23:23:04.532 -> Average Cal sensor Temp.: 101.94 Ambient Temp.: 97.50 F
23:23:06.092 -> Average Cal sensor Temp.: 102.38 Ambient Temp.: 97.54 F
23:23:07.644 -> Average Cal sensor Temp.: 102.38 Ambient Temp.: 97.50 F
23:23:09.233 -> Average Cal sensor Temp.: 102.42 Ambient Temp.: 97.47 F
23:23:10.790 -> Average Cal sensor Temp.: 101.81 Ambient Temp.: 97.47 F
23:23:12.349 -> Average Cal sensor Temp.: 101.52 Ambient Temp.: 97.50 F
23:23:13.902 -> Average Cal sensor Temp.: 101.82 Ambient Temp.: 97.50 F
23:23:15.461 -> Average Cal sensor Temp.: 102.30 Ambient Temp.: 97.50 F
23:23:17.051 -> Average Cal sensor Temp.: 102.39 Ambient Temp.: 97.50 F
23:23:18.608 -> Average Cal sensor Temp.: 101.89 Ambient Temp.: 97.47 F
23:23:20.160 -> Average Cal sensor Temp.: 102.29 Ambient Temp.: 97.47 F
23:23:21.744 -> Average Cal sensor Temp.: 102.34 Ambient Temp.: 97.47 F
23:23:23.301 -> Average Cal sensor Temp.: 102.50 Ambient Temp.: 97.50 F
23:23:24.859 -> Average Cal sensor Temp.: 102.57 Ambient Temp.: 97.47 F
23:23:26.409 -> Average Cal sensor Temp.: 102.26 Ambient Temp.: 97.50 F
23:23:27.993 -> Average Cal sensor Temp.: 102.44 Ambient Temp.: 97.50 F
23:23:29.546 -> Average Cal sensor Temp.: 102.29 Ambient Temp.: 97.50 F
23:23:31.102 -> Average Cal sensor Temp.: 102.67 Ambient Temp.: 97.50 F
23:23:32.692 -> Average Cal sensor Temp.: 102.64 Ambient Temp.: 97.47 F
23:23:34.243 -> Average Cal sensor Temp.: 101.85 Ambient Temp.: 97.43 F
23:23:35.797 -> Average Cal sensor Temp.: 101.46 Ambient Temp.: 97.47 F
23:23:37.381 -> Average Cal sensor Temp.: 101.78 Ambient Temp.: 97.43 F
23:23:38.941 -> Average Cal sensor Temp.: 101.62 Ambient Temp.: 97.50 F
23:23:40.493 -> Average Cal sensor Temp.: 101.74 Ambient Temp.: 97.47 F
23:23:42.048 -> Average Cal sensor Temp.: 102.26 Ambient Temp.: 97.43 F
23:23:43.634 -> Average Cal sensor Temp.: 102.02 Ambient Temp.: 97.47 F
23:23:45.191 -> Average Cal sensor Temp.: 102.07 Ambient Temp.: 97.47 F
23:23:46.751 -> Average Cal sensor Temp.: 102.29 Ambient Temp.: 97.50 F
23:23:48.309 -> Average Cal sensor Temp.: 101.63 Ambient Temp.: 97.47 F
23:23:49.865 -> Average Cal sensor Temp.: 102.12 Ambient Temp.: 97.43 F
23:23:51.427 -> Average Cal sensor Temp.: 102.17 Ambient Temp.: 97.47 F
23:23:53.020 -> Average Cal sensor Temp.: 101.97 Ambient Temp.: 97.54 F
23:23:54.574 -> Average Cal sensor Temp.: 102.02 Ambient Temp.: 97.50 F
23:23:56.128 -> Average Cal sensor Temp.: 102.00 Ambient Temp.: 97.47 F
23:23:57.713 -> Average Cal sensor Temp.: 101.80 Ambient Temp.: 97.43 F
23:23:59.267 -> Average Cal sensor Temp.: 101.86 Ambient Temp.: 97.47 F
23:24:00.827 -> Average Cal sensor Temp.: 101.88 Ambient Temp.: 97.47 F
23:24:02.382 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 97.47 F
23:24:03.937 -> Average Cal sensor Temp.: 99.98 Ambient Temp.: 97.47 F
23:24:05.520 -> Average Cal sensor Temp.: 99.86 Ambient Temp.: 97.47 F
23:24:07.074 -> Average Cal sensor Temp.: 103.05 Ambient Temp.: 97.47 F
23:24:08.626 -> Average Cal sensor Temp.: 102.63 Ambient Temp.: 97.43 F
23:24:10.216 -> Average Cal sensor Temp.: 102.63 Ambient Temp.: 97.47 F
23:24:11.780 -> Average Cal sensor Temp.: 101.75 Ambient Temp.: 97.43 F
23:24:13.333 -> Average Cal sensor Temp.: 101.28 Ambient Temp.: 97.47 F
23:24:14.891 -> Average Cal sensor Temp.: 101.86 Ambient Temp.: 97.43 F
23:24:16.447 -> Average Cal sensor Temp.: 101.71 Ambient Temp.: 97.43 F
23:24:18.033 -> Average Cal sensor Temp.: 101.84 Ambient Temp.: 97.39 F
23:24:19.588 -> Average Cal sensor Temp.: 102.14 Ambient Temp.: 97.47 F
23:24:21.142 -> Average Cal sensor Temp.: 102.26 Ambient Temp.: 97.47 F
23:24:22.731 -> Average Cal sensor Temp.: 102.63 Ambient Temp.: 97.50 F
23:24:24.286 -> Average Cal sensor Temp.: 102.69 Ambient Temp.: 97.43 F
23:24:25.844 -> Average Cal sensor Temp.: 102.44 Ambient Temp.: 97.47 F
23:24:27.394 -> Average Cal sensor Temp.: 102.65 Ambient Temp.: 97.47 F
23:24:28.984 -> Average Cal sensor Temp.: 102.42 Ambient Temp.: 97.47 F
23:24:30.538 -> Average Cal sensor Temp.: 102.29 Ambient Temp.: 97.54 F
23:24:32.098 -> Average Cal sensor Temp.: 102.32 Ambient Temp.: 97.50 F
23:24:33.656 -> Average Cal sensor Temp.: 102.48 Ambient Temp.: 97.50 F
23:24:35.243 -> Average Cal sensor Temp.: 102.71 Ambient Temp.: 97.47 F
23:24:36.797 -> Average Cal sensor Temp.: 102.85 Ambient Temp.: 97.54 F
23:24:38.347 -> Average Cal sensor Temp.: 102.63 Ambient Temp.: 97.47 F
23:24:39.904 -> Average Cal sensor Temp.: 102.47 Ambient Temp.: 97.57 F
23:24:41.495 -> Average Cal sensor Temp.: 102.58 Ambient Temp.: 97.54 F
23:24:43.046 -> Average Cal sensor Temp.: 102.24 Ambient Temp.: 97.54 F
23:24:44.602 -> Average Cal sensor Temp.: 102.49 Ambient Temp.: 97.54 F
23:24:46.189 -> Average Cal sensor Temp.: 102.38 Ambient Temp.: 97.57 F
23:24:47.745 -> Average Cal sensor Temp.: 102.06 Ambient Temp.: 97.54 F
23:24:49.298 -> Average Cal sensor Temp.: 101.72 Ambient Temp.: 97.54 F
23:24:50.855 -> Average Cal sensor Temp.: 101.53 Ambient Temp.: 97.54 F
23:24:52.412 -> Average Cal sensor Temp.: 102.13 Ambient Temp.: 97.54 F
23:24:53.998 -> Average Cal sensor Temp.: 102.23 Ambient Temp.: 97.54 F
23:24:55.552 -> Average Cal sensor Temp.: 102.42 Ambient Temp.: 97.54 F
23:24:57.102 -> Average Cal sensor Temp.: 102.70 Ambient Temp.: 97.54 F
23:24:58.691 -> Average Cal sensor Temp.: 102.87 Ambient Temp.: 97.57 F
23:25:00.253 -> Average Cal sensor Temp.: 102.69 Ambient Temp.: 97.61 F
23:25:01.809 -> Average Cal sensor Temp.: 102.78 Ambient Temp.: 97.61 F
23:25:03.359 -> Average Cal sensor Temp.: 102.49 Ambient Temp.: 97.57 F
23:25:04.944 -> Average Cal sensor Temp.: 102.53 Ambient Temp.: 97.61 F
23:25:06.497 -> Average Cal sensor Temp.: 102.60 Ambient Temp.: 97.57 F
23:25:08.083 -> Average Cal sensor Temp.: 102.66 Ambient Temp.: 97.61 F
23:25:09.639 -> Average Cal sensor Temp.: 102.69 Ambient Temp.: 97.68 F
23:25:11.195 -> Average Cal sensor Temp.: 102.13 Ambient Temp.: 97.57 F
23:25:12.746 -> Average Cal sensor Temp.: 102.61 Ambient Temp.: 97.61 F
23:25:14.305 -> Average Cal sensor Temp.: 102.46 Ambient Temp.: 97.68 F
23:25:15.899 -> Average Cal sensor Temp.: 102.34 Ambient Temp.: 97.61 F
23:25:17.454 -> Average Cal sensor Temp.: 102.08 Ambient Temp.: 97.68 F
23:25:19.013 -> Average Cal sensor Temp.: 102.20 Ambient Temp.: 97.61 F
23:25:20.571 -> Average Cal sensor Temp.: 102.43 Ambient Temp.: 97.68 F
23:25:22.155 -> Average Cal sensor Temp.: 102.51 Ambient Temp.: 97.72 F
23:25:23.714 -> Average Cal sensor Temp.: 102.21 Ambient Temp.: 97.72 F
23:25:25.268 -> Average Cal sensor Temp.: 102.22 Ambient Temp.: 97.61 F
23:25:26.822 -> Average Cal sensor Temp.: 102.50 Ambient Temp.: 97.72 F
23:25:28.414 -> Average Cal sensor Temp.: 102.35 Ambient Temp.: 97.75 F
23:25:29.970 -> Average Cal sensor Temp.: 102.44 Ambient Temp.: 97.72 F
23:25:31.526 -> Average Cal sensor Temp.: 102.53 Ambient Temp.: 97.68 F
23:25:33.083 -> Average Cal sensor Temp.: 102.43 Ambient Temp.: 97.75 F
23:25:34.636 -> Average Cal sensor Temp.: 101.76 Ambient Temp.: 97.68 F
23:25:36.229 -> Average Cal sensor Temp.: 101.89 Ambient Temp.: 97.68 F
23:25:37.788 -> Average Cal sensor Temp.: 102.39 Ambient Temp.: 97.68 F
23:25:39.344 -> Average Cal sensor Temp.: 102.61 Ambient Temp.: 97.72 F
23:25:40.906 -> Average Cal sensor Temp.: 102.32 Ambient Temp.: 97.72 F
23:25:42.455 -> Average Cal sensor Temp.: 102.38 Ambient Temp.: 97.61 F
23:25:44.045 -> Average Cal sensor Temp.: 102.67 Ambient Temp.: 97.61 F
23:25:45.601 -> Average Cal sensor Temp.: 102.53 Ambient Temp.: 97.61 F
23:25:47.158 -> Average Cal sensor Temp.: 102.55 Ambient Temp.: 97.68 F
23:25:48.742 -> Average Cal sensor Temp.: 102.53 Ambient Temp.: 97.72 F
23:25:50.293 -> Average Cal sensor Temp.: 102.42 Ambient Temp.: 97.61 F
23:25:51.852 -> Average Cal sensor Temp.: 102.44 Ambient Temp.: 97.61 F
23:25:53.412 -> Average Cal sensor Temp.: 102.49 Ambient Temp.: 97.68 F
23:25:54.971 -> Average Cal sensor Temp.: 102.74 Ambient Temp.: 97.72 F
23:25:56.563 -> Average Cal sensor Temp.: 102.94 Ambient Temp.: 97.61 F
23:25:58.120 -> Average Cal sensor Temp.: 102.82 Ambient Temp.: 97.72 F
23:25:59.676 -> Average Cal sensor Temp.: 101.83 Ambient Temp.: 97.61 F
23:26:01.235 -> Average Cal sensor Temp.: 102.56 Ambient Temp.: 97.61 F
23:26:02.789 -> Average Cal sensor Temp.: 102.47 Ambient Temp.: 97.57 F
23:26:04.376 -> Average Cal sensor Temp.: 102.15 Ambient Temp.: 97.72 F
23:26:05.935 -> Average Cal sensor Temp.: 102.51 Ambient Temp.: 97.68 F
23:26:07.490 -> Average Cal sensor Temp.: 102.70 Ambient Temp.: 97.61 F
23:26:09.040 -> Average Cal sensor Temp.: 102.01 Ambient Temp.: 97.68 F
23:26:10.628 -> Average Cal sensor Temp.: 102.18 Ambient Temp.: 97.68 F
23:26:12.182 -> Average Cal sensor Temp.: 102.51 Ambient Temp.: 97.72 F
23:26:13.735 -> Average Cal sensor Temp.: 102.34 Ambient Temp.: 97.68 F
23:26:15.324 -> Average Cal sensor Temp.: 102.10 Ambient Temp.: 97.68 F
23:26:16.883 -> Average Cal sensor Temp.: 102.18 Ambient Temp.: 97.68 F
23:26:18.441 -> Average Cal sensor Temp.: 102.45 Ambient Temp.: 97.61 F
23:26:19.993 -> Average Cal sensor Temp.: 102.26 Ambient Temp.: 97.72 F
23:26:21.583 -> Average Cal sensor Temp.: 102.32 Ambient Temp.: 97.61 F
23:26:23.138 -> Average Cal sensor Temp.: 102.17 Ambient Temp.: 97.68 F
23:26:24.695 -> Average Cal sensor Temp.: 102.43 Ambient Temp.: 97.68 F
23:26:26.247 -> Average Cal sensor Temp.: 102.31 Ambient Temp.: 97.68 F
23:26:27.837 -> Average Cal sensor Temp.: 102.48 Ambient Temp.: 97.68 F
23:26:29.391 -> Average Cal sensor Temp.: 102.39 Ambient Temp.: 97.68 F
23:26:30.949 -> Average Cal sensor Temp.: 102.56 Ambient Temp.: 97.68 F
23:26:32.502 -> Average Cal sensor Temp.: 102.69 Ambient Temp.: 97.72 F
23:26:34.091 -> Average Cal sensor Temp.: 102.26 Ambient Temp.: 97.68 F
23:26:35.655 -> Average Cal sensor Temp.: 102.21 Ambient Temp.: 97.68 F
23:26:37.208 -> Average Cal sensor Temp.: 102.12 Ambient Temp.: 97.61 F
23:26:38.764 -> Average Cal sensor Temp.: 102.50 Ambient Temp.: 97.68 F
23:26:40.351 -> Average Cal sensor Temp.: 102.49 Ambient Temp.: 97.61 F
23:26:41.903 -> Average Cal sensor Temp.: 102.64 Ambient Temp.: 97.68 F
23:26:43.460 -> Average Cal sensor Temp.: 102.64 Ambient Temp.: 97.72 F
23:26:45.019 -> Average Cal sensor Temp.: 102.46 Ambient Temp.: 97.68 F
23:26:46.611 -> Average Cal sensor Temp.: 102.76 Ambient Temp.: 97.72 F
23:26:48.171 -> Average Cal sensor Temp.: 102.81 Ambient Temp.: 97.68 F
23:26:49.727 -> Average Cal sensor Temp.: 102.78 Ambient Temp.: 97.75 F
23:26:51.287 -> Average Cal sensor Temp.: 102.95 Ambient Temp.: 97.68 F
23:26:52.868 -> Average Cal sensor Temp.: 102.37 Ambient Temp.: 97.72 F
23:26:54.421 -> Average Cal sensor Temp.: 102.41 Ambient Temp.: 97.61 F
23:26:55.980 -> Average Cal sensor Temp.: 102.71 Ambient Temp.: 97.68 F
23:26:57.535 -> Average Cal sensor Temp.: 102.69 Ambient Temp.: 97.72 F
23:26:59.122 -> Average Cal sensor Temp.: 102.44 Ambient Temp.: 97.75 F
23:27:00.675 -> Average Cal sensor Temp.: 102.44 Ambient Temp.: 97.61 F
23:27:02.228 -> Average Cal sensor Temp.: 102.60 Ambient Temp.: 97.72 F
23:27:03.785 -> Average Cal sensor Temp.: 103.00 Ambient Temp.: 97.75 F
23:27:05.378 -> Average Cal sensor Temp.: 103.27 Ambient Temp.: 97.72 F
23:27:06.935 -> Average Cal sensor Temp.: 103.00 Ambient Temp.: 97.75 F
23:27:08.490 -> Average Cal sensor Temp.: 102.96 Ambient Temp.: 97.75 F
23:27:10.040 -> Average Cal sensor Temp.: 102.02 Ambient Temp.: 97.72 F
23:27:11.625 -> Average Cal sensor Temp.: 102.02 Ambient Temp.: 97.75 F
23:27:13.179 -> Average Cal sensor Temp.: 102.33 Ambient Temp.: 97.79 F
23:27:14.739 -> Average Cal sensor Temp.: 102.34 Ambient Temp.: 97.72 F
23:27:16.296 -> Average Cal sensor Temp.: 102.34 Ambient Temp.: 97.75 F
23:27:17.885 -> Average Cal sensor Temp.: 102.31 Ambient Temp.: 97.68 F
23:27:19.444 -> Average Cal sensor Temp.: 102.58 Ambient Temp.: 97.68 F
23:27:20.993 -> Average Cal sensor Temp.: 102.48 Ambient Temp.: 97.68 F
23:27:22.582 -> Average Cal sensor Temp.: 102.58 Ambient Temp.: 97.72 F
23:27:24.130 -> Average Cal sensor Temp.: 102.66 Ambient Temp.: 97.75 F
23:27:25.683 -> Average Cal sensor Temp.: 102.35 Ambient Temp.: 97.68 F
23:27:27.271 -> Average Cal sensor Temp.: 102.42 Ambient Temp.: 97.68 F
23:27:28.825 -> Average Cal sensor Temp.: 102.12 Ambient Temp.: 97.61 F
23:27:30.384 -> Average Cal sensor Temp.: 102.46 Ambient Temp.: 97.72 F
23:27:31.967 -> Average Cal sensor Temp.: 102.70 Ambient Temp.: 97.68 F
23:27:33.520 -> Average Cal sensor Temp.: 102.97 Ambient Temp.: 97.61 F
23:27:35.074 -> Average Cal sensor Temp.: 103.24 Ambient Temp.: 97.68 F
23:27:36.661 -> Average Cal sensor Temp.: 103.32 Ambient Temp.: 97.61 F
23:27:38.216 -> Average Cal sensor Temp.: 102.50 Ambient Temp.: 97.61 F
23:27:39.768 -> Average Cal sensor Temp.: 102.55 Ambient Temp.: 97.68 F
23:27:41.327 -> Average Cal sensor Temp.: 102.76 Ambient Temp.: 97.61 F
23:27:42.914 -> Average Cal sensor Temp.: 102.87 Ambient Temp.: 97.54 F
23:27:44.473 -> Average Cal sensor Temp.: 103.11 Ambient Temp.: 97.68 F
23:27:46.028 -> Average Cal sensor Temp.: 103.16 Ambient Temp.: 97.68 F
23:27:47.579 -> Average Cal sensor Temp.: 102.80 Ambient Temp.: 97.61 F
23:27:49.165 -> Average Cal sensor Temp.: 103.16 Ambient Temp.: 97.72 F
23:27:50.723 -> Average Cal sensor Temp.: 103.07 Ambient Temp.: 97.68 F
23:27:52.282 -> Average Cal sensor Temp.: 102.82 Ambient Temp.: 97.68 F
23:27:53.837 -> Average Cal sensor Temp.: 103.02 Ambient Temp.: 97.61 F
23:27:55.423 -> Average Cal sensor Temp.: 102.98 Ambient Temp.: 97.68 F
23:27:56.970 -> Average Cal sensor Temp.: 103.22 Ambient Temp.: 97.68 F
23:27:58.555 -> Average Cal sensor Temp.: 102.97 Ambient Temp.: 97.68 F
23:28:00.115 -> Average Cal sensor Temp.: 102.66 Ambient Temp.: 97.68 F
23:28:01.673 -> Average Cal sensor Temp.: 102.60 Ambient Temp.: 97.68 F
23:28:03.227 -> Average Cal sensor Temp.: 102.85 Ambient Temp.: 97.61 F
23:28:04.821 -> Average Cal sensor Temp.: 103.02 Ambient Temp.: 97.68 F
23:28:06.378 -> Average Cal sensor Temp.: 102.75 Ambient Temp.: 97.61 F
23:28:07.931 -> Average Cal sensor Temp.: 102.87 Ambient Temp.: 97.68 F
23:28:09.488 -> Average Cal sensor Temp.: 103.26 Ambient Temp.: 97.68 F
23:28:11.047 -> Average Cal sensor Temp.: 103.17 Ambient Temp.: 97.68 F
23:28:12.624 -> Average Cal sensor Temp.: 102.59 Ambient Temp.: 97.68 F
23:28:14.173 -> Average Cal sensor Temp.: 102.92 Ambient Temp.: 97.61 F
23:28:15.769 -> Average Cal sensor Temp.: 103.10 Ambient Temp.: 97.61 F
23:28:17.324 -> Average Cal sensor Temp.: 103.10 Ambient Temp.: 97.68 F
23:28:18.884 -> Average Cal sensor Temp.: 103.03 Ambient Temp.: 97.72 F
23:28:20.433 -> Average Cal sensor Temp.: 103.02 Ambient Temp.: 97.68 F
23:28:22.019 -> Average Cal sensor Temp.: 103.36 Ambient Temp.: 97.72 F
23:28:23.570 -> Average Cal sensor Temp.: 103.26 Ambient Temp.: 97.75 F
23:28:25.125 -> Average Cal sensor Temp.: 103.14 Ambient Temp.: 97.72 F
23:28:26.711 -> Average Cal sensor Temp.: 102.95 Ambient Temp.: 97.72 F
23:28:28.268 -> Average Cal sensor Temp.: 102.75 Ambient Temp.: 97.72 F
23:28:29.823 -> Average Cal sensor Temp.: 103.25 Ambient Temp.: 97.83 F
23:28:31.380 -> Average Cal sensor Temp.: 103.22 Ambient Temp.: 97.79 F
23:28:32.967 -> Average Cal sensor Temp.: 103.27 Ambient Temp.: 97.75 F
23:28:34.519 -> Average Cal sensor Temp.: 103.57 Ambient Temp.: 97.79 F
23:28:36.080 -> Average Cal sensor Temp.: 103.16 Ambient Temp.: 97.75 F
23:28:37.628 -> Average Cal sensor Temp.: 103.21 Ambient Temp.: 97.79 F
23:28:39.220 -> Average Cal sensor Temp.: 103.25 Ambient Temp.: 97.79 F
23:28:40.775 -> Average Cal sensor Temp.: 102.94 Ambient Temp.: 97.79 F
23:28:42.328 -> Average Cal sensor Temp.: 103.12 Ambient Temp.: 97.79 F
23:28:43.883 -> Average Cal sensor Temp.: 103.15 Ambient Temp.: 97.79 F
23:28:45.474 -> Average Cal sensor Temp.: 103.33 Ambient Temp.: 97.79 F
23:28:47.030 -> Average Cal sensor Temp.: 103.48 Ambient Temp.: 97.83 F
23:28:48.583 -> Average Cal sensor Temp.: 103.28 Ambient Temp.: 97.79 F
23:28:50.138 -> Average Cal sensor Temp.: 103.20 Ambient Temp.: 97.86 F
23:28:51.725 -> Average Cal sensor Temp.: 103.29 Ambient Temp.: 97.86 F
23:28:53.274 -> Average Cal sensor Temp.: 103.54 Ambient Temp.: 97.83 F
23:28:54.828 -> Average Cal sensor Temp.: 103.25 Ambient Temp.: 97.83 F
23:28:56.411 -> Average Cal sensor Temp.: 103.38 Ambient Temp.: 97.86 F
23:28:57.963 -> Average Cal sensor Temp.: 103.26 Ambient Temp.: 97.86 F
23:28:59.554 -> Average Cal sensor Temp.: 102.69 Ambient Temp.: 97.83 F
23:29:01.106 -> Average Cal sensor Temp.: 102.54 Ambient Temp.: 97.90 F
23:29:02.656 -> Average Cal sensor Temp.: 102.92 Ambient Temp.: 97.86 F
23:29:04.244 -> Average Cal sensor Temp.: 103.04 Ambient Temp.: 97.90 F
23:29:05.803 -> Average Cal sensor Temp.: 102.73 Ambient Temp.: 97.83 F
23:29:07.359 -> Average Cal sensor Temp.: 103.19 Ambient Temp.: 97.86 F
23:29:08.914 -> Average Cal sensor Temp.: 102.86 Ambient Temp.: 97.83 F
23:29:10.503 -> Average Cal sensor Temp.: 103.20 Ambient Temp.: 97.83 F
23:29:12.054 -> Average Cal sensor Temp.: 103.01 Ambient Temp.: 97.86 F
23:29:13.612 -> Average Cal sensor Temp.: 103.16 Ambient Temp.: 97.86 F
23:29:15.173 -> Average Cal sensor Temp.: 103.11 Ambient Temp.: 97.86 F
23:29:16.760 -> Average Cal sensor Temp.: 103.20 Ambient Temp.: 97.86 F
23:29:18.316 -> Average Cal sensor Temp.: 103.21 Ambient Temp.: 97.86 F
23:29:19.874 -> Average Cal sensor Temp.: 103.13 Ambient Temp.: 97.83 F
23:29:21.433 -> Average Cal sensor Temp.: 103.14 Ambient Temp.: 97.86 F
23:29:22.989 -> Average Cal sensor Temp.: 103.35 Ambient Temp.: 97.90 F
23:29:24.577 -> Average Cal sensor Temp.: 103.42 Ambient Temp.: 97.86 F
23:29:26.131 -> Average Cal sensor Temp.: 103.08 Ambient Temp.: 97.90 F
23:29:27.685 -> Average Cal sensor Temp.: 102.14 Ambient Temp.: 97.86 F
23:29:29.278 -> Average Cal sensor Temp.: 102.20 Ambient Temp.: 97.86 F
23:29:30.834 -> Average Cal sensor Temp.: 102.57 Ambient Temp.: 97.83 F
23:29:32.393 -> Average Cal sensor Temp.: 102.99 Ambient Temp.: 97.83 F
23:29:33.950 -> Average Cal sensor Temp.: 103.26 Ambient Temp.: 97.90 F
23:29:35.510 -> Average Cal sensor Temp.: 103.59 Ambient Temp.: 97.86 F
23:29:37.101 -> Average Cal sensor Temp.: 103.05 Ambient Temp.: 97.83 F
23:29:38.651 -> Average Cal sensor Temp.: 102.93 Ambient Temp.: 97.83 F
23:29:40.201 -> Average Cal sensor Temp.: 103.24 Ambient Temp.: 97.83 F
23:29:41.792 -> Average Cal sensor Temp.: 102.91 Ambient Temp.: 97.79 F
23:29:43.346 -> Average Cal sensor Temp.: 103.24 Ambient Temp.: 97.86 F
23:29:44.899 -> Average Cal sensor Temp.: 102.99 Ambient Temp.: 97.79 F
23:29:46.456 -> Average Cal sensor Temp.: 103.02 Ambient Temp.: 97.79 F
23:29:48.038 -> Average Cal sensor Temp.: 103.39 Ambient Temp.: 97.83 F
23:29:49.594 -> Average Cal sensor Temp.: 103.13 Ambient Temp.: 97.86 F
23:29:51.148 -> Average Cal sensor Temp.: 103.37 Ambient Temp.: 97.83 F
23:29:52.740 -> Average Cal sensor Temp.: 102.98 Ambient Temp.: 97.86 F
23:29:54.293 -> Average Cal sensor Temp.: 102.57 Ambient Temp.: 97.79 F
23:29:55.848 -> Average Cal sensor Temp.: 102.86 Ambient Temp.: 97.79 F
23:29:57.433 -> Average Cal sensor Temp.: 102.99 Ambient Temp.: 97.83 F
23:29:58.990 -> Average Cal sensor Temp.: 103.45 Ambient Temp.: 97.86 F
23:30:00.548 -> Average Cal sensor Temp.: 103.52 Ambient Temp.: 97.86 F
23:30:02.132 -> Average Cal sensor Temp.: 103.33 Ambient Temp.: 97.83 F
23:30:03.688 -> Average Cal sensor Temp.: 103.38 Ambient Temp.: 97.83 F
23:30:05.248 -> Average Cal sensor Temp.: 103.39 Ambient Temp.: 97.83 F
23:30:06.794 -> Average Cal sensor Temp.: 103.15 Ambient Temp.: 97.83 F
23:30:08.381 -> Average Cal sensor Temp.: 103.04 Ambient Temp.: 97.83 F
23:30:09.932 -> Average Cal sensor Temp.: 103.23 Ambient Temp.: 97.83 F
23:30:11.485 -> Average Cal sensor Temp.: 102.87 Ambient Temp.: 97.83 F
23:30:13.070 -> Average Cal sensor Temp.: 102.83 Ambient Temp.: 97.83 F
23:30:14.622 -> Average Cal sensor Temp.: 103.29 Ambient Temp.: 97.86 F
23:30:16.180 -> Average Cal sensor Temp.: 102.87 Ambient Temp.: 97.86 F
23:30:17.773 -> Average Cal sensor Temp.: 102.82 Ambient Temp.: 97.83 F
23:30:19.326 -> Average Cal sensor Temp.: 103.30 Ambient Temp.: 97.83 F
23:30:20.882 -> Average Cal sensor Temp.: 103.05 Ambient Temp.: 97.79 F
23:30:22.439 -> Average Cal sensor Temp.: 102.84 Ambient Temp.: 97.83 F
23:30:24.031 -> Average Cal sensor Temp.: 103.29 Ambient Temp.: 97.83 F
23:30:25.592 -> Average Cal sensor Temp.: 103.44 Ambient Temp.: 97.83 F
23:30:27.139 -> Average Cal sensor Temp.: 103.45 Ambient Temp.: 97.79 F
23:30:28.725 -> Average Cal sensor Temp.: 103.57 Ambient Temp.: 97.83 F
23:30:30.280 -> Average Cal sensor Temp.: 103.74 Ambient Temp.: 97.83 F
23:30:31.832 -> Average Cal sensor Temp.: 103.57 Ambient Temp.: 97.86 F
23:30:33.387 -> Average Cal sensor Temp.: 103.45 Ambient Temp.: 97.86 F
23:30:34.978 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 97.79 F
23:30:36.532 -> Average Cal sensor Temp.: 103.64 Ambient Temp.: 97.86 F
23:30:38.084 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 97.86 F
23:30:39.672 -> Average Cal sensor Temp.: 103.61 Ambient Temp.: 97.86 F
23:30:41.223 -> Average Cal sensor Temp.: 103.33 Ambient Temp.: 97.83 F
23:30:42.778 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 97.90 F
23:30:44.362 -> Average Cal sensor Temp.: 103.05 Ambient Temp.: 97.86 F
23:30:45.920 -> Average Cal sensor Temp.: 103.28 Ambient Temp.: 97.86 F
23:30:47.468 -> Average Cal sensor Temp.: 103.04 Ambient Temp.: 97.83 F
23:30:49.048 -> Average Cal sensor Temp.: 103.05 Ambient Temp.: 97.86 F
23:30:50.605 -> Average Cal sensor Temp.: 103.36 Ambient Temp.: 97.79 F
23:30:52.191 -> Average Cal sensor Temp.: 103.48 Ambient Temp.: 97.86 F
23:30:53.742 -> Average Cal sensor Temp.: 103.34 Ambient Temp.: 97.83 F
23:30:55.296 -> Average Cal sensor Temp.: 103.19 Ambient Temp.: 97.83 F
23:30:56.884 -> Average Cal sensor Temp.: 103.51 Ambient Temp.: 97.83 F
23:30:58.433 -> Average Cal sensor Temp.: 103.85 Ambient Temp.: 97.86 F
23:31:00.015 -> Average Cal sensor Temp.: 103.67 Ambient Temp.: 97.79 F
23:31:01.546 -> Average Cal sensor Temp.: 103.59 Ambient Temp.: 97.86 F
23:31:03.133 -> Average Cal sensor Temp.: 103.15 Ambient Temp.: 97.86 F
23:31:04.683 -> Average Cal sensor Temp.: 103.40 Ambient Temp.: 97.86 F
23:31:06.273 -> Average Cal sensor Temp.: 103.55 Ambient Temp.: 97.86 F
23:31:07.822 -> Average Cal sensor Temp.: 103.32 Ambient Temp.: 97.83 F
23:31:09.379 -> Average Cal sensor Temp.: 103.49 Ambient Temp.: 97.86 F
23:31:10.936 -> Average Cal sensor Temp.: 103.60 Ambient Temp.: 97.83 F
23:31:12.524 -> Average Cal sensor Temp.: 103.66 Ambient Temp.: 97.86 F
23:31:14.070 -> Average Cal sensor Temp.: 103.55 Ambient Temp.: 97.86 F
23:31:15.630 -> Average Cal sensor Temp.: 103.38 Ambient Temp.: 97.86 F
23:31:17.214 -> Average Cal sensor Temp.: 102.85 Ambient Temp.: 97.86 F
23:31:18.766 -> Average Cal sensor Temp.: 103.14 Ambient Temp.: 97.90 F
23:31:20.321 -> Average Cal sensor Temp.: 103.38 Ambient Temp.: 97.86 F
23:31:21.915 -> Average Cal sensor Temp.: 103.53 Ambient Temp.: 97.83 F
23:31:23.465 -> Average Cal sensor Temp.: 103.33 Ambient Temp.: 97.86 F
23:31:25.022 -> Average Cal sensor Temp.: 103.69 Ambient Temp.: 97.86 F
23:31:26.604 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 97.86 F
23:31:28.153 -> Average Cal sensor Temp.: 103.66 Ambient Temp.: 97.86 F
23:31:29.742 -> Average Cal sensor Temp.: 102.94 Ambient Temp.: 97.83 F
23:31:31.304 -> Average Cal sensor Temp.: 102.97 Ambient Temp.: 97.83 F
23:31:32.856 -> Average Cal sensor Temp.: 103.10 Ambient Temp.: 97.83 F
23:31:34.405 -> Average Cal sensor Temp.: 103.22 Ambient Temp.: 97.83 F
23:31:35.965 -> Average Cal sensor Temp.: 103.58 Ambient Temp.: 97.83 F
23:31:37.552 -> Average Cal sensor Temp.: 103.46 Ambient Temp.: 97.83 F
23:31:39.105 -> Average Cal sensor Temp.: 103.17 Ambient Temp.: 97.83 F
23:31:40.661 -> Average Cal sensor Temp.: 103.39 Ambient Temp.: 97.79 F
23:31:42.253 -> Average Cal sensor Temp.: 103.52 Ambient Temp.: 97.79 F
23:31:43.808 -> Average Cal sensor Temp.: 103.57 Ambient Temp.: 97.83 F
23:31:45.360 -> Average Cal sensor Temp.: 103.72 Ambient Temp.: 97.83 F
23:31:46.916 -> Average Cal sensor Temp.: 103.78 Ambient Temp.: 97.90 F
23:31:48.503 -> Average Cal sensor Temp.: 103.90 Ambient Temp.: 97.90 F
23:31:50.057 -> Average Cal sensor Temp.: 103.76 Ambient Temp.: 97.90 F
23:31:51.614 -> Average Cal sensor Temp.: 103.78 Ambient Temp.: 97.83 F
23:31:53.205 -> Average Cal sensor Temp.: 103.98 Ambient Temp.: 97.79 F
23:31:54.763 -> Average Cal sensor Temp.: 103.96 Ambient Temp.: 97.83 F
23:31:56.319 -> Average Cal sensor Temp.: 103.32 Ambient Temp.: 97.86 F
23:31:57.869 -> Average Cal sensor Temp.: 103.60 Ambient Temp.: 97.79 F
23:31:59.458 -> Average Cal sensor Temp.: 103.27 Ambient Temp.: 97.86 F
23:32:01.014 -> Average Cal sensor Temp.: 102.92 Ambient Temp.: 97.83 F
23:32:02.569 -> Average Cal sensor Temp.: 103.54 Ambient Temp.: 97.83 F
23:32:04.157 -> Average Cal sensor Temp.: 103.70 Ambient Temp.: 97.83 F
23:32:05.714 -> Average Cal sensor Temp.: 103.47 Ambient Temp.: 97.83 F
23:32:07.265 -> Average Cal sensor Temp.: 103.31 Ambient Temp.: 97.83 F
23:32:08.825 -> Average Cal sensor Temp.: 103.18 Ambient Temp.: 97.86 F
23:32:10.411 -> Average Cal sensor Temp.: 103.37 Ambient Temp.: 97.79 F
23:32:11.964 -> Average Cal sensor Temp.: 103.58 Ambient Temp.: 97.83 F
23:32:13.521 -> Average Cal sensor Temp.: 103.77 Ambient Temp.: 97.90 F
23:32:15.081 -> Average Cal sensor Temp.: 103.10 Ambient Temp.: 97.83 F
23:32:16.668 -> Average Cal sensor Temp.: 103.26 Ambient Temp.: 97.86 F
23:32:18.225 -> Average Cal sensor Temp.: 103.51 Ambient Temp.: 97.86 F
23:32:19.781 -> Average Cal sensor Temp.: 103.37 Ambient Temp.: 97.86 F
23:32:21.338 -> Average Cal sensor Temp.: 103.54 Ambient Temp.: 97.86 F
23:32:22.924 -> Average Cal sensor Temp.: 103.84 Ambient Temp.: 97.86 F
23:32:24.477 -> Average Cal sensor Temp.: 103.95 Ambient Temp.: 97.90 F
23:32:26.033 -> Average Cal sensor Temp.: 103.82 Ambient Temp.: 97.90 F
23:32:27.616 -> Average Cal sensor Temp.: 103.88 Ambient Temp.: 97.90 F
23:32:29.169 -> Average Cal sensor Temp.: 103.57 Ambient Temp.: 97.86 F
23:32:30.724 -> Average Cal sensor Temp.: 103.55 Ambient Temp.: 97.86 F
23:32:32.313 -> Average Cal sensor Temp.: 102.98 Ambient Temp.: 97.90 F
23:32:33.869 -> Average Cal sensor Temp.: 103.16 Ambient Temp.: 97.86 F
23:32:35.419 -> Average Cal sensor Temp.: 103.07 Ambient Temp.: 97.83 F
23:32:37.001 -> Average Cal sensor Temp.: 102.83 Ambient Temp.: 97.86 F
23:32:38.558 -> Average Cal sensor Temp.: 103.28 Ambient Temp.: 97.86 F
23:32:40.110 -> Average Cal sensor Temp.: 103.39 Ambient Temp.: 97.83 F
23:32:41.701 -> Average Cal sensor Temp.: 103.59 Ambient Temp.: 97.93 F
23:32:43.254 -> Average Cal sensor Temp.: 103.98 Ambient Temp.: 97.86 F
23:32:44.809 -> Average Cal sensor Temp.: 103.88 Ambient Temp.: 97.86 F
23:32:46.366 -> Average Cal sensor Temp.: 103.97 Ambient Temp.: 97.90 F
23:32:47.958 -> Average Cal sensor Temp.: 103.99 Ambient Temp.: 97.86 F
23:32:49.512 -> Average Cal sensor Temp.: 103.83 Ambient Temp.: 97.90 F
23:32:51.064 -> Average Cal sensor Temp.: 103.66 Ambient Temp.: 97.86 F
23:32:52.655 -> Average Cal sensor Temp.: 103.84 Ambient Temp.: 97.90 F
23:32:54.206 -> Average Cal sensor Temp.: 103.88 Ambient Temp.: 97.83 F
23:32:55.764 -> Average Cal sensor Temp.: 103.69 Ambient Temp.: 97.86 F
23:32:57.349 -> Average Cal sensor Temp.: 103.56 Ambient Temp.: 97.90 F
23:32:58.902 -> Average Cal sensor Temp.: 103.56 Ambient Temp.: 97.86 F
23:33:00.458 -> Average Cal sensor Temp.: 103.68 Ambient Temp.: 97.86 F
23:33:02.045 -> Average Cal sensor Temp.: 103.64 Ambient Temp.: 97.90 F
23:33:03.597 -> Average Cal sensor Temp.: 103.56 Ambient Temp.: 97.93 F
23:33:05.157 -> Average Cal sensor Temp.: 103.66 Ambient Temp.: 97.90 F
23:33:06.717 -> Average Cal sensor Temp.: 103.36 Ambient Temp.: 97.90 F
23:33:08.309 -> Average Cal sensor Temp.: 103.15 Ambient Temp.: 97.90 F
23:33:09.863 -> Average Cal sensor Temp.: 103.62 Ambient Temp.: 97.90 F
23:33:11.422 -> Average Cal sensor Temp.: 103.52 Ambient Temp.: 97.90 F
23:33:12.976 -> Average Cal sensor Temp.: 103.85 Ambient Temp.: 97.93 F
23:33:14.564 -> Average Cal sensor Temp.: 103.01 Ambient Temp.: 97.93 F
23:33:16.124 -> Average Cal sensor Temp.: 103.19 Ambient Temp.: 97.93 F
23:33:17.672 -> Average Cal sensor Temp.: 102.98 Ambient Temp.: 97.93 F
23:33:19.232 -> Average Cal sensor Temp.: 102.95 Ambient Temp.: 97.90 F
23:33:20.823 -> Average Cal sensor Temp.: 103.41 Ambient Temp.: 97.90 F
23:33:22.370 -> Average Cal sensor Temp.: 103.71 Ambient Temp.: 97.90 F
23:33:23.920 -> Average Cal sensor Temp.: 103.78 Ambient Temp.: 97.90 F
23:33:25.510 -> Average Cal sensor Temp.: 103.72 Ambient Temp.: 97.90 F
23:33:27.065 -> Average Cal sensor Temp.: 103.62 Ambient Temp.: 97.90 F
23:33:28.618 -> Average Cal sensor Temp.: 103.57 Ambient Temp.: 97.93 F
23:33:30.203 -> Average Cal sensor Temp.: 103.54 Ambient Temp.: 97.90 F
23:33:31.753 -> Average Cal sensor Temp.: 103.44 Ambient Temp.: 97.93 F
23:33:33.310 -> Average Cal sensor Temp.: 103.52 Ambient Temp.: 97.90 F
23:33:34.901 -> Average Cal sensor Temp.: 103.68 Ambient Temp.: 97.93 F
23:33:36.459 -> Average Cal sensor Temp.: 103.70 Ambient Temp.: 98.01 F
23:33:38.014 -> Average Cal sensor Temp.: 103.62 Ambient Temp.: 97.90 F
23:33:39.570 -> Average Cal sensor Temp.: 103.60 Ambient Temp.: 97.93 F
23:33:41.161 -> Average Cal sensor Temp.: 103.71 Ambient Temp.: 97.97 F
23:33:42.710 -> Average Cal sensor Temp.: 103.26 Ambient Temp.: 97.93 F
23:33:44.262 -> Average Cal sensor Temp.: 103.29 Ambient Temp.: 97.93 F
23:33:45.822 -> Average Cal sensor Temp.: 102.94 Ambient Temp.: 97.86 F
23:33:47.406 -> Average Cal sensor Temp.: 103.28 Ambient Temp.: 97.90 F
23:33:48.964 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 97.90 F
23:33:50.518 -> Average Cal sensor Temp.: 103.10 Ambient Temp.: 97.86 F
23:33:52.113 -> Average Cal sensor Temp.: 103.45 Ambient Temp.: 97.86 F
23:33:53.667 -> Average Cal sensor Temp.: 103.99 Ambient Temp.: 97.86 F
23:33:55.228 -> Average Cal sensor Temp.: 103.78 Ambient Temp.: 97.90 F
23:33:56.783 -> Average Cal sensor Temp.: 104.05 Ambient Temp.: 97.90 F
23:33:58.367 -> Average Cal sensor Temp.: 103.88 Ambient Temp.: 97.93 F
23:33:59.929 -> Average Cal sensor Temp.: 103.91 Ambient Temp.: 97.97 F
23:34:01.489 -> Average Cal sensor Temp.: 103.77 Ambient Temp.: 97.93 F
23:34:03.033 -> Average Cal sensor Temp.: 103.80 Ambient Temp.: 97.93 F
23:34:04.624 -> Average Cal sensor Temp.: 103.69 Ambient Temp.: 97.86 F
23:34:06.183 -> Average Cal sensor Temp.: 103.49 Ambient Temp.: 97.90 F
23:34:07.739 -> Average Cal sensor Temp.: 103.43 Ambient Temp.: 97.86 F
23:34:09.298 -> Average Cal sensor Temp.: 103.63 Ambient Temp.: 97.90 F
23:34:10.890 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 97.90 F
23:34:12.444 -> Average Cal sensor Temp.: 104.09 Ambient Temp.: 97.86 F
23:34:13.997 -> Average Cal sensor Temp.: 103.94 Ambient Temp.: 97.93 F
23:34:15.557 -> Average Cal sensor Temp.: 103.90 Ambient Temp.: 97.93 F
23:34:17.145 -> Average Cal sensor Temp.: 103.49 Ambient Temp.: 97.90 F
23:34:18.694 -> Average Cal sensor Temp.: 103.79 Ambient Temp.: 97.90 F
23:34:20.280 -> Average Cal sensor Temp.: 104.07 Ambient Temp.: 97.97 F
23:34:21.835 -> Average Cal sensor Temp.: 104.16 Ambient Temp.: 97.90 F
23:34:23.389 -> Average Cal sensor Temp.: 104.09 Ambient Temp.: 97.93 F
23:34:24.980 -> Average Cal sensor Temp.: 104.26 Ambient Temp.: 97.90 F
23:34:26.537 -> Average Cal sensor Temp.: 104.21 Ambient Temp.: 97.93 F
23:34:28.092 -> Average Cal sensor Temp.: 104.22 Ambient Temp.: 97.97 F
23:34:29.645 -> Average Cal sensor Temp.: 104.48 Ambient Temp.: 97.93 F
23:34:31.235 -> Average Cal sensor Temp.: 103.97 Ambient Temp.: 97.97 F
23:34:32.789 -> Average Cal sensor Temp.: 103.91 Ambient Temp.: 97.97 F
23:34:34.343 -> Average Cal sensor Temp.: 103.89 Ambient Temp.: 97.97 F
23:34:35.902 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 97.90 F
23:34:37.493 -> Average Cal sensor Temp.: 103.52 Ambient Temp.: 97.97 F
23:34:39.053 -> Average Cal sensor Temp.: 103.68 Ambient Temp.: 98.01 F
23:34:40.610 -> Average Cal sensor Temp.: 103.76 Ambient Temp.: 97.97 F
23:34:42.166 -> Average Cal sensor Temp.: 104.14 Ambient Temp.: 98.01 F
23:34:43.722 -> Average Cal sensor Temp.: 104.40 Ambient Temp.: 98.01 F
23:34:45.317 -> Average Cal sensor Temp.: 104.18 Ambient Temp.: 98.08 F
23:34:46.877 -> Average Cal sensor Temp.: 103.85 Ambient Temp.: 98.08 F
23:34:48.433 -> Average Cal sensor Temp.: 103.82 Ambient Temp.: 97.97 F
23:34:49.984 -> Average Cal sensor Temp.: 103.06 Ambient Temp.: 98.01 F
23:34:51.572 -> Average Cal sensor Temp.: 103.64 Ambient Temp.: 97.97 F
23:34:53.128 -> Average Cal sensor Temp.: 103.78 Ambient Temp.: 97.97 F
23:34:54.682 -> Average Cal sensor Temp.: 104.09 Ambient Temp.: 97.90 F
23:34:56.270 -> Average Cal sensor Temp.: 104.02 Ambient Temp.: 97.97 F
23:34:57.822 -> Average Cal sensor Temp.: 103.88 Ambient Temp.: 97.97 F
23:34:59.373 -> Average Cal sensor Temp.: 103.94 Ambient Temp.: 97.97 F
23:35:00.963 -> Average Cal sensor Temp.: 104.30 Ambient Temp.: 97.97 F
23:35:02.520 -> Average Cal sensor Temp.: 104.59 Ambient Temp.: 97.93 F
23:35:04.080 -> Average Cal sensor Temp.: 104.48 Ambient Temp.: 98.01 F
23:35:05.636 -> Average Cal sensor Temp.: 104.21 Ambient Temp.: 98.01 F
23:35:07.223 -> Average Cal sensor Temp.: 103.42 Ambient Temp.: 97.93 F
23:35:08.777 -> Average Cal sensor Temp.: 103.67 Ambient Temp.: 98.01 F
23:35:10.332 -> Average Cal sensor Temp.: 103.94 Ambient Temp.: 97.93 F
23:35:11.921 -> Average Cal sensor Temp.: 103.16 Ambient Temp.: 97.93 F
23:35:13.480 -> Average Cal sensor Temp.: 103.62 Ambient Temp.: 97.93 F
23:35:15.030 -> Average Cal sensor Temp.: 103.98 Ambient Temp.: 97.93 F
23:35:16.590 -> Average Cal sensor Temp.: 103.42 Ambient Temp.: 97.93 F
23:35:18.179 -> Average Cal sensor Temp.: 102.83 Ambient Temp.: 97.93 F
23:35:19.735 -> Average Cal sensor Temp.: 103.63 Ambient Temp.: 97.93 F
23:35:21.294 -> Average Cal sensor Temp.: 103.56 Ambient Temp.: 97.86 F
23:35:22.875 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 97.93 F
23:35:24.429 -> Average Cal sensor Temp.: 103.95 Ambient Temp.: 97.93 F
23:35:25.983 -> Average Cal sensor Temp.: 104.22 Ambient Temp.: 97.93 F
23:35:27.536 -> Average Cal sensor Temp.: 104.17 Ambient Temp.: 97.90 F
23:35:29.123 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 97.90 F
23:35:30.676 -> Average Cal sensor Temp.: 103.64 Ambient Temp.: 97.86 F
23:35:32.229 -> Average Cal sensor Temp.: 104.05 Ambient Temp.: 97.93 F
23:35:33.823 -> Average Cal sensor Temp.: 103.87 Ambient Temp.: 97.97 F
23:35:35.381 -> Average Cal sensor Temp.: 104.22 Ambient Temp.: 97.86 F
23:35:36.935 -> Average Cal sensor Temp.: 104.61 Ambient Temp.: 97.90 F
23:35:38.520 -> Average Cal sensor Temp.: 104.71 Ambient Temp.: 97.93 F
23:35:40.077 -> Average Cal sensor Temp.: 104.09 Ambient Temp.: 97.86 F
23:35:41.633 -> Average Cal sensor Temp.: 104.39 Ambient Temp.: 97.90 F
23:35:43.183 -> Average Cal sensor Temp.: 104.06 Ambient Temp.: 97.90 F
23:35:44.772 -> Average Cal sensor Temp.: 104.36 Ambient Temp.: 97.93 F
23:35:46.326 -> Average Cal sensor Temp.: 104.44 Ambient Temp.: 97.90 F
23:35:47.879 -> Average Cal sensor Temp.: 104.50 Ambient Temp.: 97.90 F
23:35:49.471 -> Average Cal sensor Temp.: 104.04 Ambient Temp.: 97.90 F
23:35:51.022 -> Average Cal sensor Temp.: 103.93 Ambient Temp.: 97.97 F
23:35:52.574 -> Average Cal sensor Temp.: 104.04 Ambient Temp.: 97.90 F
23:35:54.162 -> Average Cal sensor Temp.: 103.24 Ambient Temp.: 97.90 F
23:35:55.719 -> Average Cal sensor Temp.: 103.89 Ambient Temp.: 97.86 F
23:35:57.274 -> Average Cal sensor Temp.: 103.95 Ambient Temp.: 97.90 F
23:35:58.863 -> Average Cal sensor Temp.: 104.09 Ambient Temp.: 97.90 F
23:36:00.418 -> Average Cal sensor Temp.: 104.31 Ambient Temp.: 97.90 F
23:36:01.973 -> Average Cal sensor Temp.: 104.28 Ambient Temp.: 97.86 F
23:36:03.557 -> Average Cal sensor Temp.: 104.41 Ambient Temp.: 97.93 F
23:36:05.119 -> Average Cal sensor Temp.: 104.10 Ambient Temp.: 97.86 F
23:36:06.676 -> Average Cal sensor Temp.: 103.88 Ambient Temp.: 97.97 F
23:36:08.230 -> Average Cal sensor Temp.: 103.61 Ambient Temp.: 97.86 F
23:36:09.793 -> Average Cal sensor Temp.: 103.38 Ambient Temp.: 97.83 F
23:36:11.380 -> Average Cal sensor Temp.: 103.79 Ambient Temp.: 97.90 F
23:36:12.930 -> Average Cal sensor Temp.: 103.87 Ambient Temp.: 97.93 F
23:36:14.517 -> Average Cal sensor Temp.: 104.11 Ambient Temp.: 97.93 F
23:36:16.073 -> Average Cal sensor Temp.: 104.38 Ambient Temp.: 97.86 F
23:36:17.628 -> Average Cal sensor Temp.: 104.07 Ambient Temp.: 97.86 F
23:36:19.181 -> Average Cal sensor Temp.: 104.42 Ambient Temp.: 97.90 F
23:36:20.764 -> Average Cal sensor Temp.: 104.15 Ambient Temp.: 97.93 F
23:36:22.324 -> Average Cal sensor Temp.: 104.36 Ambient Temp.: 97.86 F
23:36:23.877 -> Average Cal sensor Temp.: 104.55 Ambient Temp.: 97.90 F
23:36:25.438 -> Average Cal sensor Temp.: 103.80 Ambient Temp.: 97.86 F
23:36:27.029 -> Average Cal sensor Temp.: 104.27 Ambient Temp.: 97.93 F
23:36:28.586 -> Average Cal sensor Temp.: 104.31 Ambient Temp.: 97.86 F
23:36:30.140 -> Average Cal sensor Temp.: 104.39 Ambient Temp.: 97.90 F
23:36:31.700 -> Average Cal sensor Temp.: 104.13 Ambient Temp.: 97.86 F
23:36:33.289 -> Average Cal sensor Temp.: 104.19 Ambient Temp.: 97.90 F
23:36:34.847 -> Average Cal sensor Temp.: 104.21 Ambient Temp.: 97.86 F
23:36:36.405 -> Average Cal sensor Temp.: 104.24 Ambient Temp.: 97.93 F
23:36:37.987 -> Average Cal sensor Temp.: 103.74 Ambient Temp.: 97.90 F
23:36:39.539 -> Average Cal sensor Temp.: 104.02 Ambient Temp.: 97.86 F
23:36:41.097 -> Average Cal sensor Temp.: 104.12 Ambient Temp.: 97.90 F
23:36:42.652 -> Average Cal sensor Temp.: 104.01 Ambient Temp.: 97.90 F
23:36:44.247 -> Average Cal sensor Temp.: 104.44 Ambient Temp.: 97.90 F
23:36:45.804 -> Average Cal sensor Temp.: 104.34 Ambient Temp.: 97.90 F
23:36:47.357 -> Average Cal sensor Temp.: 104.21 Ambient Temp.: 97.93 F
23:36:48.918 -> Average Cal sensor Temp.: 104.13 Ambient Temp.: 97.93 F
23:36:50.474 -> Average Cal sensor Temp.: 103.93 Ambient Temp.: 97.93 F
23:36:52.059 -> Average Cal sensor Temp.: 103.88 Ambient Temp.: 97.90 F
23:36:53.619 -> Average Cal sensor Temp.: 103.96 Ambient Temp.: 97.93 F
23:36:55.182 -> Average Cal sensor Temp.: 103.69 Ambient Temp.: 97.90 F
23:36:56.741 -> Average Cal sensor Temp.: 103.61 Ambient Temp.: 97.97 F
23:36:58.325 -> Average Cal sensor Temp.: 103.96 Ambient Temp.: 97.90 F
23:36:59.874 -> Average Cal sensor Temp.: 103.63 Ambient Temp.: 97.93 F
23:37:01.430 -> Average Cal sensor Temp.: 103.72 Ambient Temp.: 97.90 F
23:37:03.020 -> Average Cal sensor Temp.: 104.05 Ambient Temp.: 97.90 F
23:37:04.574 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 97.90 F
23:37:06.131 -> Average Cal sensor Temp.: 103.85 Ambient Temp.: 97.90 F
23:37:07.687 -> Average Cal sensor Temp.: 103.77 Ambient Temp.: 97.90 F
23:37:09.275 -> Average Cal sensor Temp.: 103.88 Ambient Temp.: 97.86 F
23:37:10.830 -> Average Cal sensor Temp.: 104.02 Ambient Temp.: 97.90 F
23:37:12.389 -> Average Cal sensor Temp.: 104.36 Ambient Temp.: 97.86 F
23:37:13.946 -> Average Cal sensor Temp.: 104.35 Ambient Temp.: 97.93 F
23:37:15.537 -> Average Cal sensor Temp.: 104.36 Ambient Temp.: 97.90 F
23:37:17.088 -> Average Cal sensor Temp.: 104.37 Ambient Temp.: 97.93 F
23:37:18.642 -> Average Cal sensor Temp.: 104.10 Ambient Temp.: 97.93 F
23:37:20.235 -> Average Cal sensor Temp.: 103.96 Ambient Temp.: 97.90 F
23:37:21.792 -> Average Cal sensor Temp.: 104.13 Ambient Temp.: 97.86 F
23:37:23.349 -> Average Cal sensor Temp.: 104.18 Ambient Temp.: 97.90 F
23:37:24.902 -> Average Cal sensor Temp.: 104.38 Ambient Temp.: 97.90 F
23:37:26.494 -> Average Cal sensor Temp.: 104.45 Ambient Temp.: 97.93 F
23:37:28.049 -> Average Cal sensor Temp.: 104.14 Ambient Temp.: 97.93 F
23:37:29.605 -> Average Cal sensor Temp.: 103.78 Ambient Temp.: 97.93 F
23:37:31.161 -> Average Cal sensor Temp.: 103.75 Ambient Temp.: 97.93 F
23:37:32.752 -> Average Cal sensor Temp.: 103.37 Ambient Temp.: 97.90 F
23:37:34.301 -> Average Cal sensor Temp.: 103.90 Ambient Temp.: 97.86 F
23:37:35.856 -> Average Cal sensor Temp.: 104.32 Ambient Temp.: 97.90 F
23:37:37.444 -> Average Cal sensor Temp.: 104.46 Ambient Temp.: 97.93 F
23:37:38.997 -> Average Cal sensor Temp.: 104.70 Ambient Temp.: 97.93 F
23:37:40.551 -> Average Cal sensor Temp.: 104.41 Ambient Temp.: 97.86 F
23:37:42.139 -> Average Cal sensor Temp.: 104.40 Ambient Temp.: 97.93 F
23:37:43.690 -> Average Cal sensor Temp.: 104.17 Ambient Temp.: 98.01 F
23:37:45.252 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 97.97 F
23:37:46.808 -> Average Cal sensor Temp.: 103.99 Ambient Temp.: 97.97 F
23:37:48.394 -> Average Cal sensor Temp.: 104.27 Ambient Temp.: 97.86 F
23:37:49.948 -> Average Cal sensor Temp.: 104.24 Ambient Temp.: 97.93 F
23:37:51.500 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 97.90 F
23:37:53.091 -> Average Cal sensor Temp.: 104.18 Ambient Temp.: 97.93 F
23:37:54.645 -> Average Cal sensor Temp.: 103.82 Ambient Temp.: 97.93 F
23:37:56.202 -> Average Cal sensor Temp.: 103.77 Ambient Temp.: 97.97 F
23:37:57.757 -> Average Cal sensor Temp.: 104.04 Ambient Temp.: 97.93 F
23:37:59.342 -> Average Cal sensor Temp.: 103.83 Ambient Temp.: 97.93 F
23:38:00.899 -> Average Cal sensor Temp.: 104.38 Ambient Temp.: 97.90 F
23:38:02.457 -> Average Cal sensor Temp.: 103.91 Ambient Temp.: 97.90 F
23:38:04.045 -> Average Cal sensor Temp.: 103.68 Ambient Temp.: 97.86 F
23:38:05.601 -> Average Cal sensor Temp.: 104.18 Ambient Temp.: 97.90 F
23:38:07.155 -> Average Cal sensor Temp.: 104.48 Ambient Temp.: 97.90 F
23:38:08.708 -> Average Cal sensor Temp.: 104.36 Ambient Temp.: 97.90 F
23:38:10.294 -> Average Cal sensor Temp.: 104.67 Ambient Temp.: 97.86 F
23:38:11.853 -> Average Cal sensor Temp.: 104.49 Ambient Temp.: 97.90 F
23:38:13.409 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 97.93 F
23:38:15.002 -> Average Cal sensor Temp.: 104.31 Ambient Temp.: 97.83 F
23:38:16.558 -> Average Cal sensor Temp.: 104.29 Ambient Temp.: 97.86 F
23:38:18.113 -> Average Cal sensor Temp.: 104.26 Ambient Temp.: 97.90 F
23:38:19.671 -> Average Cal sensor Temp.: 104.45 Ambient Temp.: 97.90 F
23:38:21.235 -> Average Cal sensor Temp.: 104.25 Ambient Temp.: 97.90 F
23:38:22.819 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 97.90 F
23:38:24.371 -> Average Cal sensor Temp.: 104.19 Ambient Temp.: 97.97 F
23:38:25.929 -> Average Cal sensor Temp.: 103.92 Ambient Temp.: 97.93 F
23:38:27.513 -> Average Cal sensor Temp.: 104.17 Ambient Temp.: 97.97 F
23:38:29.064 -> Average Cal sensor Temp.: 104.34 Ambient Temp.: 97.97 F
23:38:30.650 -> Average Cal sensor Temp.: 104.49 Ambient Temp.: 97.97 F
23:38:32.204 -> Average Cal sensor Temp.: 104.62 Ambient Temp.: 97.93 F
23:38:33.761 -> Average Cal sensor Temp.: 104.50 Ambient Temp.: 97.93 F
23:38:35.315 -> Average Cal sensor Temp.: 104.58 Ambient Temp.: 97.97 F
23:38:36.905 -> Average Cal sensor Temp.: 104.71 Ambient Temp.: 98.01 F
23:38:38.456 -> Average Cal sensor Temp.: 104.77 Ambient Temp.: 97.97 F
23:38:40.012 -> Average Cal sensor Temp.: 104.58 Ambient Temp.: 98.08 F
23:38:41.599 -> Average Cal sensor Temp.: 104.53 Ambient Temp.: 98.01 F
23:38:43.158 -> Average Cal sensor Temp.: 104.34 Ambient Temp.: 98.08 F
23:38:44.709 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 98.08 F
23:38:46.301 -> Average Cal sensor Temp.: 104.25 Ambient Temp.: 98.01 F
23:38:47.859 -> Average Cal sensor Temp.: 103.84 Ambient Temp.: 98.08 F
23:38:49.415 -> Average Cal sensor Temp.: 104.02 Ambient Temp.: 98.08 F
23:38:50.998 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 98.08 F
23:38:52.555 -> Average Cal sensor Temp.: 104.18 Ambient Temp.: 98.08 F
23:38:54.105 -> Average Cal sensor Temp.: 104.18 Ambient Temp.: 98.08 F
23:38:55.663 -> Average Cal sensor Temp.: 104.23 Ambient Temp.: 98.11 F
23:38:57.253 -> Average Cal sensor Temp.: 104.23 Ambient Temp.: 98.11 F
23:38:58.799 -> Average Cal sensor Temp.: 103.66 Ambient Temp.: 98.01 F
23:39:00.382 -> Average Cal sensor Temp.: 104.01 Ambient Temp.: 98.11 F
23:39:01.936 -> Average Cal sensor Temp.: 104.09 Ambient Temp.: 98.08 F
23:39:03.497 -> Average Cal sensor Temp.: 103.94 Ambient Temp.: 98.11 F
23:39:05.054 -> Average Cal sensor Temp.: 103.98 Ambient Temp.: 98.11 F
23:39:06.643 -> Average Cal sensor Temp.: 104.04 Ambient Temp.: 98.11 F
23:39:08.197 -> Average Cal sensor Temp.: 102.66 Ambient Temp.: 98.08 F
23:39:09.752 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 98.08 F
23:39:11.311 -> Average Cal sensor Temp.: 103.81 Ambient Temp.: 98.08 F
23:39:12.896 -> Average Cal sensor Temp.: 103.64 Ambient Temp.: 98.08 F
23:39:14.447 -> Average Cal sensor Temp.: 103.29 Ambient Temp.: 98.08 F
23:39:16.041 -> Average Cal sensor Temp.: 103.08 Ambient Temp.: 97.97 F
23:39:17.598 -> Average Cal sensor Temp.: 103.43 Ambient Temp.: 97.97 F
23:39:19.149 -> Average Cal sensor Temp.: 103.64 Ambient Temp.: 98.01 F
23:39:20.703 -> Average Cal sensor Temp.: 103.77 Ambient Temp.: 98.01 F
23:39:22.287 -> Average Cal sensor Temp.: 103.77 Ambient Temp.: 98.01 F
23:39:23.844 -> Average Cal sensor Temp.: 103.52 Ambient Temp.: 98.01 F
23:39:25.394 -> Average Cal sensor Temp.: 103.12 Ambient Temp.: 98.01 F