## SET-7
Reference Clinical Thermometer Readout | Reference DIGI Thermometer Readout | Reference Water Temperature (Sensor) | Calibrating contactless sensor Readout (2-3 cm) | Calibrating contactless sensor AMBIENT Readout | DIFF bw clinical and sensor | DIFF bw DIGI and sensor | DIFF BETWEEN CLINICAL AND DIGI
Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F) | Read at (deg F)
94 | 93.9 | 93.31 | 91.2 | 94.59 | 2.8 | 2.7 | 0.1
94.4 | 94.2 | 93.65 | 91.44 | 94.66 | 2.96 | 2.76 | 0.2
94.8 | 94.7 | 93.99 | 91.94 | 94.8 | 2.86 | 2.76 | 0.1
95 | 94.9 | 94.21 | 92.21 | 94.77 | 2.79 | 2.69 | 0.1
95.4 | 95.5 | 94.44 | 92.33 | 94.8 | 3.07 | 3.17 | -0.1
95.8 | 96 | 94.77 | 92.69 | 94.95 | 3.11 | 3.31 | -0.2
96 | 96 | 95 | 93.06 | 94.98 | 2.94 | 2.94 | 0
96.4 | 96.3 | 95.45 | 93.45 | 95.05 | 2.95 | 2.85 | 0.1
96.8 | 96.6 | 95.68 | 93.81 | 95.16 | 2.99 | 2.79 | 0.2
97 | 96.8 | 96.01 | 94.17 | 95.27 | 2.83 | 2.63 | 0.2
97.4 | 97.6 | 96.35 | 94.35 | 95.34 | 3.05 | 3.25 | -0.2
97.8 | 97.8 | 96.69 | 94.74 | 95.45 | 3.06 | 3.06 | 0
98 | 98.2 | 97.02 | 94.95 | 95.52 | 3.05 | 3.25 | -0.2
98.2 | 98.2 | 97.25 | 95.16 | 95.56 | 3.04 | 3.04 | 0
98.4 | 98.5 | 97.36 | 95.37 | 95.59 | 3.03 | 3.13 | -0.1
98.6 | 98.7 | 97.47 | 95.5 | 95.67 | 3.1 | 3.2 | -0.1
98.8 | 98.9 | 97.7 | 95.67 | 95.7 | 3.13 | 3.23 | -0.1
99 | 98.9 | 97.92 | 95.74 | 95.7 | 3.26 | 3.16 | 0.1
99.2 | 99.2 | 98.15 | 95.93 | 95.77 | 3.27 | 3.27 | 0
99.4 | 99.3 | 98.37 | 96.16 | 95.81 | 3.24 | 3.14 | 0.1
99.6 | 99.5 | 98.49 | 96.26 | 95.77 | 3.34 | 3.24 | 0.1
99.8 | 99.9 | 98.6 | 96.52 | 95.81 | 3.28 | 3.38 | -0.1
100 | 100.1 | 98.82 | 96.74 | 95.85 | 3.26 | 3.36 | -0.1
100.2 | 100.4 | 99.05 | 96.89 | 95.92 | 3.31 | 3.51 | -0.2
100.4 | 100.6 | 99.27 | 97.09 | 95.95 | 3.31 | 3.51 | -0.2
100.6 | 100.7 | 99.5 | 97.26 | 96.03 | 3.34 | 3.44 | -0.1
100.8 | 101 | 99.61 | 97.33 | 95.99 | 3.47 | 3.67 | -0.2
101 | 101 | 99.84 | 97.65 | 96.1 | 3.35 | 3.35 | 0
101.2 | 101.1 | 100.06 | 97.85 | 96.13 | 3.35 | 3.25 | 0.1
101.4 | 101.3 | 100.17 | 98.07 | 96.17 | 3.33 | 3.23 | 0.1
101.6 | 101.5 | 100.51 | 98.16 | 96.17 | 3.44 | 3.34 | 0.1
101.8 | 101.9 | 100.62 | 98.45 | 96.21 | 3.35 | 3.45 | -0.1
102 | 102.2 | 100.74 | 98.62 | 96.21 | 3.38 | 3.58 | -0.2
102.2 | 102.4 | 100.96 | 98.86 | 96.21 | 3.34 | 3.54 | -0.2
102.4 | 102.5 | 101.3 | 99.1 | 96.35 | 3.3 | 3.4 | -0.1
102.6 | 102.6 | 101.41 | 99.24 | 96.42 | 3.36 | 3.36 | 0
102.8 | 102.7 | 101.75 | 99.42 | 96.53 | 3.38 | 3.28 | 0.1
103 | 103.1 | 101.86 | 99.59 | 96.57 | 3.41 | 3.51 | -0.1
103.4 | 103.8 | 102.2 | 100.08 | 96.49 | 3.32 | 3.72 | -0.4
103.8 | 103.8 | 102.65 | 100.45 | 96.71 | 3.35 | 3.35 | 0
104 | 104.2 | 102.76 | 100.58 | 96.78 | 3.42 | 3.62 | -0.2
104.4 | 104.6 | 102.65 | 100.85 | 96.67 | 3.55 | 3.75 | -0.2
104.8 | 105.3 | 102.87 | 101.05 | 96.75 | 3.75 | 4.25 | -0.5
105 | 105.5 | 103.21 | 101.25 | 96.82 | 3.75 | 4.25 | -0.5
105.4 | 105.8 | 103.66 | 101.56 | 96.93 | 3.84 | 4.24 | -0.4
105.8 | 106 | 103.77 | 101.81 | 96.89 | 3.99 | 4.19 | -0.2
106 | 106.4 | 103.89 | 102.22 | 97 | 3.78 | 4.18 | -0.4
106.4 | 106.6 | 104.9 | 102.5 | 97.03 | 3.9 | 4.1 | -0.2
106.8 | 107.2 | 105.35 | 103.38 | 97.14 | 3.42 | 3.82 | -0.4
107 | 107.5 | 105.46 | 103.5 | 97.18 | 3.5 | 4 | -0.5
107.4 | 107.7 | 105.8 | 103.88 | 97.21 | 3.52 | 3.82 | -0.3
107.8 | 108.2 | 106.25 | 104.14 | 97.21 | 3.66 | 4.06 | -0.4
108 | 108.3 | 106.47 | 104.54 | 97.39 | 3.46 | 3.76 | -0.3

## RAW_HEATER-7
16:41:02.589 -> Water_Temp.:89.15, PWM:22
16:41:03.377 -> Water_Temp.:89.15, PWM:22
16:41:04.163 -> Water_Temp.:89.26, PWM:22
16:41:04.920 -> Water_Temp.:89.26, PWM:22
16:41:06.508 -> Water_Temp.:89.26, PWM:22
16:41:07.272 -> Water_Temp.:89.26, PWM:22
16:41:08.067 -> Water_Temp.:89.26, PWM:22
16:41:08.826 -> Water_Temp.:89.26, PWM:22
16:41:09.619 -> Water_Temp.:89.26, PWM:22
16:41:10.380 -> Water_Temp.:89.26, PWM:22
16:41:11.173 -> Water_Temp.:89.26, PWM:22
16:41:11.967 -> Water_Temp.:89.26, PWM:22
16:41:12.723 -> Water_Temp.:89.26, PWM:22
16:41:13.517 -> Water_Temp.:89.26, PWM:22
16:41:14.312 -> Water_Temp.:89.37, PWM:22
16:41:15.070 -> Water_Temp.:89.37, PWM:22
16:41:15.863 -> Water_Temp.:89.37, PWM:22
16:41:16.620 -> Water_Temp.:89.37, PWM:22
16:41:17.413 -> Water_Temp.:89.37, PWM:22
16:41:18.203 -> Water_Temp.:89.37, PWM:22
16:41:18.965 -> Water_Temp.:89.37, PWM:22
16:41:19.761 -> Water_Temp.:89.37, PWM:22
16:41:20.518 -> Water_Temp.:89.37, PWM:22
16:41:21.307 -> Water_Temp.:89.37, PWM:22
16:41:22.105 -> Water_Temp.:89.37, PWM:22
16:41:23.621 -> Water_Temp.:89.37, PWM:22
16:41:24.420 -> Water_Temp.:89.37, PWM:22
16:41:25.211 -> Water_Temp.:89.37, PWM:22
16:41:25.965 -> Water_Temp.:89.49, PWM:22
16:41:26.760 -> Water_Temp.:89.49, PWM:22
16:41:27.521 -> Water_Temp.:89.49, PWM:22
16:41:28.315 -> Water_Temp.:89.49, PWM:22
16:41:29.111 -> Water_Temp.:89.49, PWM:22
16:41:29.874 -> Water_Temp.:89.49, PWM:22
16:41:30.667 -> Water_Temp.:89.49, PWM:22
16:41:31.429 -> Water_Temp.:89.49, PWM:22
16:41:32.221 -> Water_Temp.:89.49, PWM:22
16:41:32.979 -> Water_Temp.:89.49, PWM:22
16:41:33.771 -> Water_Temp.:89.49, PWM:22
16:41:34.564 -> Water_Temp.:89.49, PWM:22
16:41:35.327 -> Water_Temp.:89.49, PWM:22
16:41:36.125 -> Water_Temp.:89.49, PWM:22
16:41:36.885 -> Water_Temp.:89.60, PWM:22
16:41:37.681 -> Water_Temp.:89.60, PWM:22
16:41:38.439 -> Water_Temp.:89.60, PWM:22
16:41:39.234 -> Water_Temp.:89.60, PWM:22
16:41:40.022 -> Water_Temp.:89.60, PWM:22
16:41:40.779 -> Water_Temp.:89.60, PWM:22
16:41:41.574 -> Water_Temp.:89.60, PWM:22
16:41:42.369 -> Water_Temp.:89.60, PWM:22
16:41:43.130 -> Water_Temp.:89.60, PWM:22
16:41:43.926 -> Water_Temp.:89.60, PWM:22
16:41:44.682 -> Water_Temp.:89.60, PWM:22
16:41:45.477 -> Water_Temp.:89.60, PWM:22
16:41:46.269 -> Water_Temp.:89.60, PWM:22
16:41:47.816 -> Water_Temp.:89.60, PWM:22
16:41:48.575 -> Water_Temp.:89.60, PWM:22
16:41:49.369 -> Water_Temp.:89.60, PWM:22
16:41:50.127 -> Water_Temp.:89.60, PWM:22
16:41:50.924 -> Water_Temp.:89.60, PWM:22
16:41:51.687 -> Water_Temp.:89.60, PWM:22
16:41:52.478 -> Water_Temp.:89.60, PWM:22
16:41:53.269 -> Water_Temp.:89.60, PWM:22
16:41:54.026 -> Water_Temp.:89.60, PWM:22
16:41:54.813 -> Water_Temp.:89.60, PWM:22
16:41:55.608 -> Water_Temp.:89.60, PWM:22
16:41:56.368 -> Water_Temp.:89.60, PWM:22
16:41:57.160 -> Water_Temp.:89.60, PWM:22
16:41:57.950 -> Water_Temp.:89.60, PWM:22
16:41:58.707 -> Water_Temp.:89.60, PWM:22
16:41:59.498 -> Water_Temp.:89.60, PWM:22
16:42:01.049 -> Water_Temp.:89.60, PWM:22
16:42:01.846 -> Water_Temp.:89.60, PWM:22
16:42:02.637 -> Water_Temp.:89.60, PWM:22
16:42:03.397 -> Water_Temp.:89.60, PWM:22
16:42:04.189 -> Water_Temp.:89.71, PWM:22
16:42:04.950 -> Water_Temp.:89.60, PWM:22
16:42:05.745 -> Water_Temp.:89.60, PWM:22
16:42:06.504 -> Water_Temp.:89.60, PWM:22
16:42:07.296 -> Water_Temp.:89.60, PWM:22
16:42:08.088 -> Water_Temp.:89.60, PWM:22
16:42:08.851 -> Water_Temp.:89.60, PWM:22
16:42:09.642 -> Water_Temp.:89.60, PWM:22
16:42:10.402 -> Water_Temp.:89.71, PWM:22
16:42:11.193 -> Water_Temp.:89.71, PWM:22
16:42:11.987 -> Water_Temp.:89.71, PWM:22
16:42:12.746 -> Water_Temp.:89.71, PWM:22
16:42:13.539 -> Water_Temp.:89.71, PWM:22
16:42:14.332 -> Water_Temp.:89.71, PWM:22
16:42:15.085 -> Water_Temp.:89.71, PWM:22
16:42:15.877 -> Water_Temp.:89.71, PWM:22
16:42:16.671 -> Water_Temp.:89.71, PWM:22
16:42:17.427 -> Water_Temp.:89.71, PWM:22
16:42:18.221 -> Water_Temp.:89.71, PWM:22
16:42:19.014 -> Water_Temp.:89.71, PWM:22
16:42:19.774 -> Water_Temp.:89.71, PWM:22
16:42:20.566 -> Water_Temp.:89.71, PWM:22
16:42:21.324 -> Water_Temp.:89.71, PWM:22
16:42:22.118 -> Water_Temp.:89.71, PWM:22
16:42:22.912 -> Water_Temp.:91.40, PWM:22
16:42:23.673 -> Water_Temp.:93.43, PWM:22
16:42:24.466 -> Water_Temp.:90.95, PWM:22
16:42:25.229 -> Water_Temp.:90.61, PWM:22
16:42:26.018 -> Water_Temp.:90.39, PWM:22
16:42:26.812 -> Water_Temp.:90.27, PWM:22
16:42:27.573 -> Water_Temp.:90.16, PWM:22
16:42:28.366 -> Water_Temp.:90.16, PWM:22
16:42:29.128 -> Water_Temp.:90.05, PWM:22
16:42:29.922 -> Water_Temp.:90.05, PWM:22
16:42:30.680 -> Water_Temp.:90.05, PWM:22
16:42:31.473 -> Water_Temp.:89.94, PWM:22
16:42:32.268 -> Water_Temp.:89.94, PWM:22
16:42:33.027 -> Water_Temp.:89.94, PWM:22
16:42:33.818 -> Water_Temp.:89.94, PWM:22
16:42:34.610 -> Water_Temp.:91.62, PWM:22
16:42:35.370 -> Water_Temp.:93.65, PWM:22
16:42:36.924 -> Water_Temp.:90.84, PWM:22
16:42:37.716 -> Water_Temp.:90.61, PWM:22
16:42:38.508 -> Water_Temp.:90.50, PWM:22
16:42:39.265 -> Water_Temp.:90.39, PWM:22
16:42:40.059 -> Water_Temp.:90.27, PWM:22
16:42:40.818 -> Water_Temp.:90.27, PWM:22
16:42:41.611 -> Water_Temp.:90.27, PWM:22
16:42:42.408 -> Water_Temp.:90.16, PWM:22
16:42:43.173 -> Water_Temp.:90.16, PWM:22
16:42:43.965 -> Water_Temp.:90.16, PWM:22
16:42:44.724 -> Water_Temp.:90.05, PWM:22
16:42:45.510 -> Water_Temp.:90.05, PWM:22
16:42:46.306 -> Water_Temp.:90.05, PWM:22
16:42:47.067 -> Water_Temp.:90.05, PWM:22
16:42:47.862 -> Water_Temp.:90.05, PWM:22
16:42:48.622 -> Water_Temp.:90.05, PWM:22
16:42:49.414 -> Water_Temp.:90.05, PWM:22
16:42:50.207 -> Water_Temp.:90.05, PWM:22
16:42:50.967 -> Water_Temp.:90.05, PWM:22
16:42:51.762 -> Water_Temp.:90.05, PWM:22
16:42:52.522 -> Water_Temp.:90.05, PWM:22
16:42:53.319 -> Water_Temp.:90.05, PWM:22
16:42:54.079 -> Water_Temp.:90.05, PWM:22
16:42:54.876 -> Water_Temp.:90.05, PWM:22
16:42:56.430 -> Water_Temp.:90.16, PWM:22
16:42:57.223 -> Water_Temp.:90.05, PWM:22
16:42:57.984 -> Water_Temp.:90.05, PWM:22
16:42:58.775 -> Water_Temp.:90.05, PWM:22
16:43:00.328 -> Water_Temp.:90.05, PWM:22
16:43:01.126 -> Water_Temp.:90.16, PWM:22
16:43:01.884 -> Water_Temp.:90.16, PWM:22
16:43:02.677 -> Water_Temp.:90.16, PWM:22
16:43:03.465 -> Water_Temp.:90.16, PWM:22
16:43:04.222 -> Water_Temp.:90.16, PWM:22
16:43:05.016 -> Water_Temp.:90.16, PWM:22
16:43:05.775 -> Water_Temp.:90.16, PWM:22
16:43:06.569 -> Water_Temp.:90.05, PWM:22
16:43:07.362 -> Water_Temp.:90.05, PWM:22
16:43:08.122 -> Water_Temp.:90.05, PWM:22
16:43:08.917 -> Water_Temp.:90.16, PWM:22
16:43:09.674 -> Water_Temp.:90.16, PWM:22
16:43:10.471 -> Water_Temp.:90.16, PWM:22
16:43:11.261 -> Water_Temp.:90.16, PWM:22
16:43:12.019 -> Water_Temp.:90.16, PWM:22
16:43:12.816 -> Water_Temp.:90.16, PWM:22
16:43:13.607 -> Water_Temp.:90.16, PWM:22
16:43:14.364 -> Water_Temp.:90.16, PWM:22
16:43:15.162 -> Water_Temp.:90.16, PWM:22
16:43:16.683 -> Water_Temp.:90.16, PWM:22
16:43:17.480 -> Water_Temp.:90.16, PWM:22
16:43:18.234 -> Water_Temp.:90.16, PWM:22
16:43:19.032 -> Water_Temp.:90.16, PWM:22
16:43:19.796 -> Water_Temp.:90.16, PWM:22
16:43:20.588 -> Water_Temp.:90.16, PWM:22
16:43:21.346 -> Water_Temp.:90.16, PWM:22
16:43:22.140 -> Water_Temp.:90.16, PWM:22
16:43:22.933 -> Water_Temp.:90.16, PWM:22
16:43:23.691 -> Water_Temp.:90.27, PWM:22
16:43:24.482 -> Water_Temp.:90.27, PWM:22
16:43:25.276 -> Water_Temp.:90.27, PWM:22
16:43:26.037 -> Water_Temp.:90.16, PWM:22
16:43:26.832 -> Water_Temp.:90.27, PWM:22
16:43:27.594 -> Water_Temp.:90.27, PWM:22
16:43:29.148 -> Water_Temp.:90.16, PWM:22
16:43:29.946 -> Water_Temp.:90.16, PWM:22
16:43:30.705 -> Water_Temp.:90.16, PWM:22
16:43:31.497 -> Water_Temp.:90.16, PWM:22
16:43:32.297 -> Water_Temp.:90.16, PWM:22
16:43:33.053 -> Water_Temp.:90.16, PWM:22
16:43:33.845 -> Water_Temp.:90.16, PWM:22
16:43:34.607 -> Water_Temp.:90.16, PWM:22
16:43:35.403 -> Water_Temp.:90.16, PWM:22
16:43:36.165 -> Water_Temp.:90.16, PWM:22
16:43:36.961 -> Water_Temp.:90.16, PWM:22
16:43:37.752 -> Water_Temp.:90.16, PWM:22
16:43:38.509 -> Water_Temp.:90.16, PWM:22
16:43:39.299 -> Water_Temp.:90.16, PWM:22
16:43:40.095 -> Water_Temp.:90.16, PWM:22
16:43:41.640 -> Water_Temp.:90.16, PWM:22
16:43:42.399 -> Water_Temp.:90.16, PWM:22
16:43:43.192 -> Water_Temp.:90.16, PWM:22
16:43:43.985 -> Water_Temp.:90.16, PWM:22
16:43:44.744 -> Water_Temp.:90.27, PWM:22
16:43:45.540 -> Water_Temp.:90.27, PWM:22
16:43:46.300 -> Water_Temp.:90.27, PWM:22
16:43:47.850 -> Water_Temp.:90.27, PWM:22
16:43:48.613 -> Water_Temp.:90.27, PWM:22
16:43:49.410 -> Water_Temp.:90.27, PWM:22
16:43:50.168 -> Water_Temp.:90.27, PWM:22
16:43:50.963 -> Water_Temp.:90.27, PWM:22
16:43:51.756 -> Water_Temp.:90.27, PWM:22
16:43:52.509 -> Water_Temp.:90.27, PWM:22
16:43:53.300 -> Water_Temp.:90.27, PWM:29
16:43:54.095 -> Water_Temp.:90.27, PWM:29
16:43:54.853 -> Water_Temp.:90.27, PWM:29
16:43:55.649 -> Water_Temp.:90.27, PWM:26
16:43:56.409 -> Water_Temp.:90.27, PWM:26
16:43:57.968 -> Water_Temp.:90.27, PWM:26
16:43:58.725 -> Water_Temp.:90.27, PWM:26
16:43:59.517 -> Water_Temp.:90.27, PWM:26
16:44:00.313 -> Water_Temp.:90.27, PWM:27
16:44:01.070 -> Water_Temp.:90.27, PWM:26
16:44:01.863 -> Water_Temp.:90.27, PWM:26
16:44:02.624 -> Water_Temp.:90.39, PWM:26
16:44:04.215 -> Water_Temp.:90.39, PWM:26
16:44:04.975 -> Water_Temp.:90.39, PWM:26
16:44:05.773 -> Water_Temp.:90.39, PWM:26
16:44:06.534 -> Water_Temp.:90.50, PWM:27
16:44:07.326 -> Water_Temp.:90.50, PWM:27
16:44:08.085 -> Water_Temp.:90.50, PWM:26
16:44:08.877 -> Water_Temp.:90.50, PWM:27
16:44:10.428 -> Water_Temp.:90.61, PWM:26
16:44:11.192 -> Water_Temp.:90.61, PWM:26
16:44:11.990 -> Water_Temp.:90.61, PWM:27
16:44:12.747 -> Water_Temp.:90.61, PWM:26
16:44:13.540 -> Water_Temp.:90.61, PWM:26
16:44:15.095 -> Water_Temp.:90.61, PWM:27
16:44:15.856 -> Water_Temp.:90.61, PWM:27
16:44:16.650 -> Water_Temp.:90.72, PWM:26
16:44:17.445 -> Water_Temp.:90.72, PWM:26
16:44:18.204 -> Water_Temp.:90.72, PWM:26
16:44:18.995 -> Water_Temp.:90.72, PWM:26
16:44:19.788 -> Water_Temp.:90.72, PWM:26
16:44:20.543 -> Water_Temp.:90.72, PWM:27
16:44:21.340 -> Water_Temp.:90.72, PWM:27
16:44:22.101 -> Water_Temp.:90.72, PWM:26
16:44:22.896 -> Water_Temp.:90.72, PWM:26
16:44:23.654 -> Water_Temp.:90.72, PWM:27
16:44:24.446 -> Water_Temp.:90.72, PWM:26
16:44:25.236 -> Water_Temp.:90.72, PWM:26
16:44:25.996 -> Water_Temp.:90.72, PWM:27
16:44:26.790 -> Water_Temp.:90.72, PWM:26
16:44:27.584 -> Water_Temp.:90.72, PWM:27
16:44:28.346 -> Water_Temp.:90.72, PWM:26
16:44:29.143 -> Water_Temp.:90.72, PWM:27
16:44:29.901 -> Water_Temp.:90.72, PWM:27
16:44:30.694 -> Water_Temp.:90.72, PWM:26
16:44:31.454 -> Water_Temp.:90.72, PWM:27
16:44:32.247 -> Water_Temp.:90.72, PWM:27
16:44:33.043 -> Water_Temp.:90.72, PWM:27
16:44:33.804 -> Water_Temp.:90.72, PWM:26
16:44:34.599 -> Water_Temp.:90.72, PWM:26
16:44:35.362 -> Water_Temp.:90.72, PWM:27
16:44:36.155 -> Water_Temp.:90.72, PWM:26
16:44:36.912 -> Water_Temp.:90.72, PWM:26
16:44:37.704 -> Water_Temp.:90.72, PWM:26
16:44:38.499 -> Water_Temp.:90.72, PWM:26
16:44:39.259 -> Water_Temp.:90.72, PWM:27
16:44:40.055 -> Water_Temp.:90.72, PWM:26
16:44:40.814 -> Water_Temp.:90.72, PWM:26
16:44:41.605 -> Water_Temp.:90.72, PWM:27
16:44:42.405 -> Water_Temp.:90.84, PWM:26
16:44:43.168 -> Water_Temp.:90.72, PWM:27
16:44:43.964 -> Water_Temp.:90.72, PWM:26
16:44:44.724 -> Water_Temp.:90.84, PWM:26
16:44:45.523 -> Water_Temp.:90.84, PWM:26
16:44:46.285 -> Water_Temp.:90.84, PWM:26
16:44:47.076 -> Water_Temp.:90.84, PWM:27
16:44:47.836 -> Water_Temp.:90.84, PWM:26
16:44:48.630 -> Water_Temp.:90.84, PWM:26
16:44:49.425 -> Water_Temp.:90.84, PWM:26
16:44:50.184 -> Water_Temp.:90.84, PWM:26
16:44:50.974 -> Water_Temp.:90.95, PWM:27
16:44:51.731 -> Water_Temp.:90.95, PWM:26
16:44:52.523 -> Water_Temp.:90.95, PWM:27
16:44:54.076 -> Water_Temp.:90.95, PWM:27
16:44:54.869 -> Water_Temp.:90.95, PWM:27
16:44:55.631 -> Water_Temp.:90.95, PWM:27
16:44:56.428 -> Water_Temp.:90.95, PWM:26
16:44:57.188 -> Water_Temp.:90.95, PWM:27
16:44:57.985 -> Water_Temp.:90.95, PWM:26
16:44:58.745 -> Water_Temp.:90.95, PWM:26
16:44:59.540 -> Water_Temp.:90.95, PWM:26
16:45:00.302 -> Water_Temp.:90.95, PWM:26
16:45:01.099 -> Water_Temp.:90.95, PWM:27
16:45:01.859 -> Water_Temp.:90.95, PWM:26
16:45:02.653 -> Water_Temp.:90.95, PWM:27
16:45:03.448 -> Water_Temp.:90.95, PWM:26
16:45:04.206 -> Water_Temp.:90.95, PWM:26
16:45:05.000 -> Water_Temp.:90.95, PWM:26
16:45:05.756 -> Water_Temp.:90.95, PWM:27
16:45:06.551 -> Water_Temp.:90.95, PWM:27
16:45:07.347 -> Water_Temp.:90.95, PWM:27
16:45:08.106 -> Water_Temp.:90.95, PWM:27
16:45:08.903 -> Water_Temp.:90.95, PWM:26
16:45:09.664 -> Water_Temp.:90.95, PWM:27
16:45:10.458 -> Water_Temp.:90.95, PWM:26
16:45:11.223 -> Water_Temp.:91.06, PWM:27
16:45:12.012 -> Water_Temp.:90.95, PWM:26
16:45:12.803 -> Water_Temp.:91.06, PWM:26
16:45:13.564 -> Water_Temp.:91.06, PWM:26
16:45:14.360 -> Water_Temp.:90.95, PWM:26
16:45:15.118 -> Water_Temp.:91.06, PWM:26
16:45:15.914 -> Water_Temp.:91.06, PWM:26
16:45:16.675 -> Water_Temp.:91.06, PWM:26
16:45:17.469 -> Water_Temp.:91.06, PWM:27
16:45:18.266 -> Water_Temp.:91.06, PWM:26
16:45:19.028 -> Water_Temp.:91.06, PWM:27
16:45:19.824 -> Water_Temp.:91.06, PWM:27
16:45:20.583 -> Water_Temp.:91.06, PWM:26
16:45:21.375 -> Water_Temp.:91.06, PWM:27
16:45:22.932 -> Water_Temp.:91.06, PWM:26
16:45:23.721 -> Water_Temp.:91.06, PWM:27
16:45:24.481 -> Water_Temp.:91.06, PWM:26
16:45:25.273 -> Water_Temp.:91.06, PWM:26
16:45:26.033 -> Water_Temp.:91.06, PWM:27
16:45:26.829 -> Water_Temp.:91.06, PWM:26
16:45:27.624 -> Water_Temp.:91.06, PWM:26
16:45:28.380 -> Water_Temp.:91.06, PWM:27
16:45:29.177 -> Water_Temp.:91.06, PWM:27
16:45:29.937 -> Water_Temp.:91.06, PWM:26
16:45:30.731 -> Water_Temp.:91.06, PWM:26
16:45:31.523 -> Water_Temp.:91.06, PWM:26
16:45:32.284 -> Water_Temp.:91.06, PWM:26
16:45:33.077 -> Water_Temp.:91.06, PWM:26
16:45:33.837 -> Water_Temp.:91.06, PWM:26
16:45:34.630 -> Water_Temp.:91.06, PWM:27
16:45:35.425 -> Water_Temp.:91.06, PWM:26
16:45:36.184 -> Water_Temp.:91.06, PWM:26
16:45:36.979 -> Water_Temp.:91.06, PWM:27
16:45:37.739 -> Water_Temp.:91.06, PWM:26
16:45:38.533 -> Water_Temp.:91.06, PWM:27
16:45:39.294 -> Water_Temp.:91.06, PWM:26
16:45:40.086 -> Water_Temp.:91.18, PWM:26
16:45:40.881 -> Water_Temp.:91.06, PWM:26
16:45:41.638 -> Water_Temp.:91.18, PWM:26
16:45:42.432 -> Water_Temp.:91.18, PWM:26
16:45:43.985 -> Water_Temp.:91.18, PWM:27
16:45:45.505 -> Water_Temp.:91.18, PWM:26
16:45:46.301 -> Water_Temp.:91.18, PWM:26
16:45:47.088 -> Water_Temp.:91.18, PWM:27
16:45:47.846 -> Water_Temp.:91.18, PWM:27
16:45:48.643 -> Water_Temp.:91.18, PWM:26
16:45:49.400 -> Water_Temp.:91.18, PWM:26
16:45:50.195 -> Water_Temp.:91.18, PWM:26
16:45:50.990 -> Water_Temp.:91.18, PWM:27
16:45:51.750 -> Water_Temp.:91.18, PWM:26
16:45:52.543 -> Water_Temp.:91.18, PWM:26
16:45:53.301 -> Water_Temp.:91.18, PWM:26
16:45:54.091 -> Water_Temp.:91.29, PWM:26
16:45:54.884 -> Water_Temp.:91.29, PWM:27
16:45:55.670 -> Water_Temp.:91.29, PWM:26
16:45:56.426 -> Water_Temp.:91.29, PWM:26
16:45:57.221 -> Water_Temp.:91.29, PWM:27
16:45:57.980 -> Water_Temp.:91.29, PWM:27
16:45:58.773 -> Water_Temp.:91.29, PWM:26
16:45:59.565 -> Water_Temp.:91.29, PWM:27
16:46:00.325 -> Water_Temp.:91.29, PWM:26
16:46:01.118 -> Water_Temp.:91.29, PWM:27
16:46:01.907 -> Water_Temp.:91.29, PWM:26
16:46:02.668 -> Water_Temp.:91.29, PWM:27
16:46:03.461 -> Water_Temp.:91.29, PWM:27
16:46:04.218 -> Water_Temp.:91.29, PWM:26
16:46:05.772 -> Water_Temp.:91.29, PWM:27
16:46:06.567 -> Water_Temp.:91.29, PWM:27
16:46:07.326 -> Water_Temp.:91.40, PWM:26
16:46:08.878 -> Water_Temp.:91.29, PWM:26
16:46:09.638 -> Water_Temp.:91.29, PWM:26
16:46:10.429 -> Water_Temp.:91.40, PWM:26
16:46:11.220 -> Water_Temp.:91.29, PWM:26
16:46:11.975 -> Water_Temp.:91.40, PWM:27
16:46:12.767 -> Water_Temp.:91.40, PWM:26
16:46:13.560 -> Water_Temp.:91.29, PWM:26
16:46:14.322 -> Water_Temp.:91.40, PWM:27
16:46:15.119 -> Water_Temp.:91.40, PWM:26
16:46:15.880 -> Water_Temp.:91.40, PWM:26
16:46:16.676 -> Water_Temp.:91.40, PWM:26
16:46:17.435 -> Water_Temp.:91.40, PWM:26
16:46:18.228 -> Water_Temp.:91.40, PWM:26
16:46:19.024 -> Water_Temp.:91.40, PWM:27
16:46:20.540 -> Water_Temp.:91.40, PWM:27
16:46:21.333 -> Water_Temp.:91.40, PWM:27
16:46:22.092 -> Water_Temp.:91.40, PWM:27
16:46:22.889 -> Water_Temp.:91.40, PWM:26
16:46:23.681 -> Water_Temp.:91.40, PWM:27
16:46:24.439 -> Water_Temp.:91.40, PWM:26
16:46:25.235 -> Water_Temp.:91.40, PWM:26
16:46:25.997 -> Water_Temp.:91.40, PWM:26
16:46:26.796 -> Water_Temp.:91.51, PWM:27
16:46:27.555 -> Water_Temp.:91.51, PWM:26
16:46:28.352 -> Water_Temp.:93.09, PWM:26
16:46:29.119 -> Water_Temp.:95.11, PWM:26
16:46:29.915 -> Water_Temp.:92.75, PWM:26
16:46:30.675 -> Water_Temp.:92.30, PWM:26
16:46:31.462 -> Water_Temp.:92.19, PWM:27
16:46:32.259 -> Water_Temp.:92.07, PWM:27
16:46:33.018 -> Water_Temp.:91.96, PWM:27
16:46:33.812 -> Water_Temp.:91.85, PWM:27
16:46:35.333 -> Water_Temp.:91.85, PWM:27
16:46:36.127 -> Water_Temp.:91.74, PWM:26
16:46:36.886 -> Water_Temp.:91.74, PWM:27
16:46:37.684 -> Water_Temp.:91.74, PWM:27
16:46:39.237 -> Water_Temp.:91.74, PWM:26
16:46:39.998 -> Water_Temp.:91.62, PWM:27
16:46:40.789 -> Water_Temp.:91.62, PWM:26
16:46:41.547 -> Water_Temp.:91.62, PWM:26
16:46:42.341 -> Water_Temp.:91.62, PWM:26
16:46:43.137 -> Water_Temp.:91.62, PWM:26
16:46:43.896 -> Water_Temp.:91.62, PWM:26
16:46:44.687 -> Water_Temp.:91.62, PWM:26
16:46:45.478 -> Water_Temp.:91.62, PWM:26
16:46:46.241 -> Water_Temp.:91.62, PWM:26
16:46:47.035 -> Water_Temp.:91.62, PWM:26
16:46:47.795 -> Water_Temp.:91.74, PWM:26
16:46:48.591 -> Water_Temp.:91.74, PWM:27
16:46:49.353 -> Water_Temp.:91.74, PWM:26
16:46:50.150 -> Water_Temp.:91.74, PWM:26
16:46:50.910 -> Water_Temp.:91.74, PWM:26
16:46:51.704 -> Water_Temp.:91.74, PWM:26
16:46:53.258 -> Water_Temp.:91.74, PWM:26
16:46:54.050 -> Water_Temp.:91.74, PWM:27
16:46:54.813 -> Water_Temp.:91.74, PWM:27
16:46:55.609 -> Water_Temp.:91.74, PWM:26
16:46:56.366 -> Water_Temp.:91.74, PWM:26
16:46:57.158 -> Water_Temp.:91.74, PWM:27
16:46:57.950 -> Water_Temp.:91.74, PWM:26
16:46:58.711 -> Water_Temp.:91.74, PWM:26
16:46:59.507 -> Water_Temp.:91.74, PWM:26
16:47:00.266 -> Water_Temp.:91.74, PWM:26
16:47:01.058 -> Water_Temp.:91.74, PWM:27
16:47:03.375 -> Water_Temp.:91.74, PWM:27
16:47:04.127 -> Water_Temp.:91.85, PWM:26
16:47:04.924 -> Water_Temp.:91.74, PWM:26
16:47:05.710 -> Water_Temp.:91.74, PWM:26
16:47:07.232 -> Water_Temp.:91.85, PWM:26
16:47:08.028 -> Water_Temp.:91.85, PWM:26
16:47:08.788 -> Water_Temp.:91.85, PWM:26
16:47:09.586 -> Water_Temp.:91.85, PWM:26
16:47:10.348 -> Water_Temp.:91.85, PWM:26
16:47:11.140 -> Water_Temp.:91.85, PWM:26
16:47:11.902 -> Water_Temp.:91.96, PWM:26
16:47:12.697 -> Water_Temp.:91.96, PWM:27
16:47:13.457 -> Water_Temp.:91.96, PWM:26
16:47:14.249 -> Water_Temp.:91.96, PWM:27
16:47:15.046 -> Water_Temp.:91.96, PWM:27
16:47:15.807 -> Water_Temp.:92.07, PWM:26
16:47:16.601 -> Water_Temp.:92.07, PWM:27
16:47:18.158 -> Water_Temp.:92.07, PWM:26
16:47:18.950 -> Water_Temp.:92.07, PWM:26
16:47:19.710 -> Water_Temp.:92.07, PWM:27
16:47:20.503 -> Water_Temp.:92.07, PWM:26
16:47:21.264 -> Water_Temp.:92.07, PWM:27
16:47:22.053 -> Water_Temp.:92.07, PWM:27
16:47:22.847 -> Water_Temp.:92.07, PWM:26
16:47:23.603 -> Water_Temp.:92.07, PWM:27
16:47:24.396 -> Water_Temp.:92.07, PWM:27
16:47:25.155 -> Water_Temp.:92.07, PWM:27
16:47:25.949 -> Water_Temp.:92.07, PWM:26
16:47:26.744 -> Water_Temp.:92.07, PWM:26
16:47:27.503 -> Water_Temp.:91.96, PWM:26
16:47:28.295 -> Water_Temp.:91.96, PWM:26
16:47:29.087 -> Water_Temp.:91.96, PWM:26
16:47:30.607 -> Water_Temp.:91.96, PWM:26
16:47:31.403 -> Water_Temp.:91.96, PWM:27
16:47:32.163 -> Water_Temp.:91.96, PWM:27
16:47:32.956 -> Water_Temp.:91.96, PWM:27
16:47:33.719 -> Water_Temp.:91.96, PWM:26
16:47:34.515 -> Water_Temp.:91.96, PWM:26
16:47:35.309 -> Water_Temp.:92.07, PWM:26
16:47:36.066 -> Water_Temp.:91.96, PWM:27
16:47:36.861 -> Water_Temp.:92.07, PWM:26
16:47:37.618 -> Water_Temp.:91.96, PWM:26
16:47:38.409 -> Water_Temp.:92.07, PWM:27
16:47:39.201 -> Water_Temp.:92.07, PWM:26
16:47:39.961 -> Water_Temp.:92.07, PWM:26
16:47:40.757 -> Water_Temp.:92.07, PWM:27
16:47:41.514 -> Water_Temp.:92.07, PWM:26
16:47:42.309 -> Water_Temp.:92.07, PWM:27
16:47:43.102 -> Water_Temp.:92.07, PWM:27
16:47:43.861 -> Water_Temp.:92.07, PWM:27
16:47:44.650 -> Water_Temp.:92.07, PWM:26
16:47:45.443 -> Water_Temp.:92.07, PWM:27
16:47:46.204 -> Water_Temp.:92.07, PWM:27
16:47:46.996 -> Water_Temp.:92.07, PWM:26
16:47:47.758 -> Water_Temp.:92.07, PWM:26
16:47:48.553 -> Water_Temp.:92.07, PWM:27
16:47:50.106 -> Water_Temp.:92.07, PWM:26
16:47:50.866 -> Water_Temp.:92.07, PWM:26
16:47:52.447 -> Water_Temp.:92.19, PWM:27
16:47:53.207 -> Water_Temp.:92.19, PWM:26
16:47:53.999 -> Water_Temp.:92.19, PWM:27
16:47:54.755 -> Water_Temp.:92.19, PWM:27
16:47:55.544 -> Water_Temp.:92.19, PWM:27
16:47:57.107 -> Water_Temp.:92.19, PWM:27
16:47:57.865 -> Water_Temp.:92.19, PWM:26
16:47:58.659 -> Water_Temp.:92.19, PWM:26
16:48:00.215 -> Water_Temp.:92.19, PWM:27
16:48:01.011 -> Water_Temp.:92.19, PWM:26
16:48:01.771 -> Water_Temp.:92.19, PWM:27
16:48:02.562 -> Water_Temp.:92.19, PWM:27
16:48:03.321 -> Water_Temp.:92.19, PWM:27
16:48:04.875 -> Water_Temp.:92.30, PWM:26
16:48:05.637 -> Water_Temp.:92.30, PWM:26
16:48:06.428 -> Water_Temp.:92.30, PWM:27
16:48:07.218 -> Water_Temp.:93.87, PWM:26
16:48:08.010 -> Water_Temp.:95.90, PWM:27
16:48:09.561 -> Water_Temp.:93.20, PWM:26
16:48:10.324 -> Water_Temp.:92.97, PWM:26
16:48:11.119 -> Water_Temp.:92.86, PWM:26
16:48:11.881 -> Water_Temp.:92.75, PWM:26
16:48:12.679 -> Water_Temp.:92.64, PWM:26
16:48:13.440 -> Water_Temp.:92.64, PWM:26
16:48:14.232 -> Water_Temp.:92.52, PWM:26
16:48:15.025 -> Water_Temp.:92.52, PWM:26
16:48:15.781 -> Water_Temp.:92.52, PWM:26
16:48:17.343 -> Water_Temp.:92.41, PWM:27
16:48:18.135 -> Water_Temp.:92.41, PWM:26
16:48:18.928 -> Water_Temp.:92.41, PWM:26
16:48:19.692 -> Water_Temp.:92.41, PWM:27
16:48:20.486 -> Water_Temp.:92.41, PWM:27
16:48:21.243 -> Water_Temp.:92.41, PWM:27
16:48:22.036 -> Water_Temp.:92.41, PWM:26
16:48:22.798 -> Water_Temp.:92.41, PWM:26
16:48:23.592 -> Water_Temp.:92.41, PWM:26
16:48:24.385 -> Water_Temp.:92.41, PWM:27
16:48:25.146 -> Water_Temp.:92.41, PWM:27
16:48:25.941 -> Water_Temp.:92.41, PWM:27
16:48:26.698 -> Water_Temp.:92.41, PWM:26
16:48:27.493 -> Water_Temp.:92.41, PWM:26
16:48:29.049 -> Water_Temp.:92.41, PWM:27
16:48:29.847 -> Water_Temp.:92.41, PWM:27
16:48:30.605 -> Water_Temp.:92.41, PWM:26
16:48:31.399 -> Water_Temp.:92.41, PWM:26
16:48:32.158 -> Water_Temp.:92.41, PWM:26
16:48:32.950 -> Water_Temp.:92.41, PWM:26
16:48:33.745 -> Water_Temp.:92.41, PWM:27
16:48:34.503 -> Water_Temp.:92.41, PWM:26
16:48:35.298 -> Water_Temp.:92.52, PWM:26
16:48:36.057 -> Water_Temp.:92.52, PWM:26
16:48:36.856 -> Water_Temp.:92.52, PWM:27
16:48:37.618 -> Water_Temp.:92.52, PWM:26
16:48:38.415 -> Water_Temp.:92.52, PWM:26
16:48:39.177 -> Water_Temp.:92.52, PWM:27
16:48:39.970 -> Water_Temp.:92.64, PWM:27
16:48:40.765 -> Water_Temp.:92.64, PWM:26
16:48:41.521 -> Water_Temp.:92.64, PWM:27
16:48:42.312 -> Water_Temp.:92.64, PWM:26
16:48:43.103 -> Water_Temp.:92.64, PWM:26
16:48:43.868 -> Water_Temp.:92.64, PWM:26
16:48:44.666 -> Water_Temp.:92.64, PWM:27
16:48:45.429 -> Water_Temp.:92.64, PWM:26
16:48:46.222 -> Water_Temp.:92.64, PWM:26
16:48:46.983 -> Water_Temp.:92.64, PWM:26
16:48:47.776 -> Water_Temp.:92.64, PWM:26
16:48:48.535 -> Water_Temp.:92.64, PWM:26
16:48:49.330 -> Water_Temp.:92.64, PWM:26
16:48:50.884 -> Water_Temp.:92.64, PWM:27
16:48:51.680 -> Water_Temp.:92.64, PWM:27
16:48:52.438 -> Water_Temp.:92.64, PWM:27
16:48:53.991 -> Water_Temp.:92.64, PWM:26
16:48:54.786 -> Water_Temp.:92.64, PWM:26
16:48:55.551 -> Water_Temp.:92.64, PWM:26
16:48:56.351 -> Water_Temp.:92.64, PWM:27
16:48:57.111 -> Water_Temp.:92.64, PWM:27
16:48:57.906 -> Water_Temp.:92.64, PWM:27
16:48:58.664 -> Water_Temp.:92.64, PWM:26
16:48:59.463 -> Water_Temp.:92.64, PWM:27
16:49:00.218 -> Water_Temp.:92.64, PWM:26
16:49:01.009 -> Water_Temp.:92.64, PWM:26
16:49:01.802 -> Water_Temp.:92.75, PWM:27
16:49:02.565 -> Water_Temp.:92.64, PWM:26
16:49:03.362 -> Water_Temp.:92.75, PWM:27
16:49:04.121 -> Water_Temp.:92.64, PWM:26
16:49:04.915 -> Water_Temp.:92.64, PWM:27
16:49:05.707 -> Water_Temp.:92.75, PWM:26
16:49:06.467 -> Water_Temp.:92.75, PWM:26
16:49:07.262 -> Water_Temp.:92.75, PWM:26
16:49:08.025 -> Water_Temp.:92.75, PWM:26
16:49:08.822 -> Water_Temp.:92.75, PWM:27
16:49:09.584 -> Water_Temp.:92.75, PWM:26
16:49:10.384 -> Water_Temp.:92.75, PWM:26
16:49:11.147 -> Water_Temp.:92.75, PWM:27
16:49:11.939 -> Water_Temp.:92.75, PWM:26
16:49:12.700 -> Water_Temp.:92.75, PWM:26
16:49:13.495 -> Water_Temp.:92.75, PWM:27
16:49:14.255 -> Water_Temp.:92.75, PWM:26
16:49:15.048 -> Water_Temp.:92.75, PWM:26
16:49:15.841 -> Water_Temp.:92.75, PWM:27
16:49:16.601 -> Water_Temp.:92.75, PWM:27
16:49:17.398 -> Water_Temp.:92.75, PWM:26
16:49:18.160 -> Water_Temp.:92.75, PWM:27
16:49:18.956 -> Water_Temp.:92.75, PWM:27
16:49:19.720 -> Water_Temp.:92.75, PWM:27
16:49:20.517 -> Water_Temp.:92.75, PWM:27
16:49:21.283 -> Water_Temp.:92.75, PWM:26
16:49:22.078 -> Water_Temp.:92.75, PWM:27
16:49:22.837 -> Water_Temp.:92.75, PWM:26
16:49:23.631 -> Water_Temp.:92.75, PWM:27
16:49:24.427 -> Water_Temp.:92.75, PWM:27
16:49:25.980 -> Water_Temp.:92.75, PWM:26
16:49:26.737 -> Water_Temp.:92.75, PWM:27
16:49:27.531 -> Water_Temp.:92.75, PWM:27
16:49:28.326 -> Water_Temp.:92.75, PWM:26
16:49:29.088 -> Water_Temp.:92.75, PWM:26
16:49:29.882 -> Water_Temp.:92.86, PWM:26
16:49:30.642 -> Water_Temp.:92.86, PWM:26
16:49:31.435 -> Water_Temp.:92.86, PWM:27
16:49:32.196 -> Water_Temp.:92.86, PWM:26
16:49:32.988 -> Water_Temp.:92.86, PWM:26
16:49:33.785 -> Water_Temp.:92.86, PWM:27
16:49:35.344 -> Water_Temp.:92.86, PWM:27
16:49:36.100 -> Water_Temp.:92.86, PWM:27
16:49:36.888 -> Water_Temp.:92.97, PWM:27
16:49:37.678 -> Water_Temp.:92.86, PWM:26
16:49:38.440 -> Water_Temp.:92.97, PWM:26
16:49:39.230 -> Water_Temp.:92.86, PWM:27
16:49:40.023 -> Water_Temp.:92.86, PWM:27
16:49:40.782 -> Water_Temp.:92.86, PWM:26
16:49:41.574 -> Water_Temp.:92.97, PWM:26
16:49:42.337 -> Water_Temp.:92.97, PWM:27
16:49:43.131 -> Water_Temp.:92.97, PWM:26
16:49:43.922 -> Water_Temp.:92.86, PWM:26
16:49:44.678 -> Water_Temp.:92.86, PWM:26
16:49:45.469 -> Water_Temp.:92.86, PWM:26
16:49:46.263 -> Water_Temp.:92.86, PWM:27
16:49:47.024 -> Water_Temp.:92.86, PWM:27
16:49:47.818 -> Water_Temp.:92.86, PWM:26
16:49:48.577 -> Water_Temp.:92.86, PWM:27
16:49:49.368 -> Water_Temp.:92.86, PWM:26
16:49:50.130 -> Water_Temp.:92.86, PWM:26
16:49:50.924 -> Water_Temp.:92.86, PWM:27
16:49:51.717 -> Water_Temp.:92.86, PWM:26
16:49:52.473 -> Water_Temp.:92.86, PWM:26
16:49:53.265 -> Water_Temp.:92.86, PWM:27
16:49:54.060 -> Water_Temp.:92.86, PWM:26
16:49:55.611 -> Water_Temp.:92.86, PWM:27
16:49:56.376 -> Water_Temp.:92.86, PWM:26
16:49:57.169 -> Water_Temp.:92.86, PWM:26
16:49:57.931 -> Water_Temp.:92.86, PWM:27
16:49:58.725 -> Water_Temp.:92.86, PWM:27
16:49:59.487 -> Water_Temp.:92.86, PWM:26
16:50:00.278 -> Water_Temp.:92.86, PWM:26
16:50:01.076 -> Water_Temp.:92.86, PWM:26
16:50:01.835 -> Water_Temp.:92.86, PWM:26
16:50:02.626 -> Water_Temp.:92.97, PWM:26
16:50:03.418 -> Water_Temp.:92.97, PWM:26
16:50:04.177 -> Water_Temp.:92.97, PWM:26
16:50:04.972 -> Water_Temp.:92.97, PWM:26
16:50:05.732 -> Water_Temp.:93.09, PWM:26
16:50:06.526 -> Water_Temp.:93.09, PWM:27
16:50:07.289 -> Water_Temp.:93.09, PWM:27
16:50:08.090 -> Water_Temp.:93.09, PWM:27
16:50:08.851 -> Water_Temp.:93.09, PWM:27
16:50:09.643 -> Water_Temp.:93.09, PWM:26
16:50:10.440 -> Water_Temp.:93.09, PWM:26
16:50:11.200 -> Water_Temp.:93.09, PWM:27
16:50:11.996 -> Water_Temp.:93.09, PWM:27
16:50:12.756 -> Water_Temp.:93.09, PWM:26
16:50:13.549 -> Water_Temp.:93.09, PWM:26
16:50:14.310 -> Water_Temp.:93.09, PWM:26
16:50:15.108 -> Water_Temp.:93.09, PWM:26
16:50:15.870 -> Water_Temp.:93.09, PWM:27
16:50:16.666 -> Water_Temp.:93.09, PWM:27
16:50:17.456 -> Water_Temp.:93.09, PWM:27
16:50:19.011 -> Water_Temp.:93.09, PWM:26
16:50:19.772 -> Water_Temp.:93.09, PWM:26
16:50:20.566 -> Water_Temp.:93.09, PWM:26
16:50:21.358 -> Water_Temp.:93.09, PWM:27
16:50:22.116 -> Water_Temp.:93.09, PWM:27
16:50:22.905 -> Water_Temp.:93.09, PWM:27
16:50:23.665 -> Water_Temp.:93.09, PWM:26
16:50:24.458 -> Water_Temp.:93.09, PWM:26
16:50:25.254 -> Water_Temp.:93.09, PWM:27
16:50:26.012 -> Water_Temp.:93.09, PWM:27
16:50:26.808 -> Water_Temp.:93.09, PWM:26
16:50:27.570 -> Water_Temp.:93.09, PWM:27
16:50:28.362 -> Water_Temp.:93.09, PWM:26
16:50:29.154 -> Water_Temp.:93.09, PWM:27
16:50:29.909 -> Water_Temp.:93.09, PWM:26
16:50:31.497 -> Water_Temp.:93.09, PWM:27
16:50:32.257 -> Water_Temp.:93.20, PWM:26
16:50:33.050 -> Water_Temp.:93.20, PWM:26
16:50:33.811 -> Water_Temp.:93.20, PWM:26
16:50:34.605 -> Water_Temp.:93.20, PWM:27
16:50:35.367 -> Water_Temp.:93.20, PWM:27
16:50:36.164 -> Water_Temp.:93.20, PWM:26
16:50:36.923 -> Water_Temp.:93.20, PWM:27
16:50:37.715 -> Water_Temp.:93.20, PWM:27
16:50:38.510 -> Water_Temp.:93.20, PWM:26
16:50:39.271 -> Water_Temp.:93.20, PWM:26
16:50:40.065 -> Water_Temp.:93.20, PWM:26
16:50:41.624 -> Water_Temp.:93.20, PWM:27
16:50:42.387 -> Water_Temp.:93.20, PWM:27
16:50:43.177 -> Water_Temp.:93.20, PWM:26
16:50:43.969 -> Water_Temp.:93.20, PWM:27
16:50:44.732 -> Water_Temp.:93.20, PWM:26
16:50:45.528 -> Water_Temp.:93.20, PWM:27
16:50:46.288 -> Water_Temp.:93.20, PWM:26
16:50:47.081 -> Water_Temp.:93.20, PWM:26
16:50:47.873 -> Water_Temp.:93.20, PWM:26
16:50:48.636 -> Water_Temp.:93.20, PWM:26
16:50:49.433 -> Water_Temp.:93.20, PWM:26
16:50:50.192 -> Water_Temp.:93.20, PWM:26
16:50:50.987 -> Water_Temp.:93.20, PWM:26
16:50:51.749 -> Water_Temp.:93.20, PWM:27
16:50:52.543 -> Water_Temp.:93.20, PWM:26
16:50:53.336 -> Water_Temp.:93.20, PWM:26
16:50:54.094 -> Water_Temp.:93.20, PWM:27
16:50:54.885 -> Water_Temp.:93.20, PWM:26
16:50:55.650 -> Water_Temp.:93.20, PWM:27
16:50:57.204 -> Water_Temp.:93.31, PWM:26
16:50:57.998 -> Water_Temp.:93.31, PWM:26
16:50:58.757 -> Water_Temp.:93.31, PWM:27
16:50:59.555 -> Water_Temp.:93.31, PWM:26
16:51:00.316 -> Water_Temp.:93.31, PWM:26
16:51:01.105 -> Water_Temp.:93.43, PWM:26
16:51:01.900 -> Water_Temp.:93.43, PWM:26
16:51:02.658 -> Water_Temp.:93.43, PWM:27
16:51:03.448 -> Water_Temp.:93.43, PWM:27
16:51:04.241 -> Water_Temp.:93.43, PWM:27
16:51:05.000 -> Water_Temp.:93.43, PWM:26
16:51:05.796 -> Water_Temp.:93.43, PWM:27
16:51:06.557 -> Water_Temp.:93.43, PWM:27
16:51:07.348 -> Water_Temp.:93.43, PWM:26
16:51:08.110 -> Water_Temp.:93.31, PWM:27
16:51:08.907 -> Water_Temp.:93.31, PWM:27
16:51:09.698 -> Water_Temp.:93.31, PWM:26
16:51:10.460 -> Water_Temp.:93.31, PWM:26
16:51:11.258 -> Water_Temp.:93.31, PWM:26
16:51:12.776 -> Water_Temp.:93.31, PWM:27
16:51:13.572 -> Water_Temp.:93.31, PWM:26
16:51:14.333 -> Water_Temp.:93.31, PWM:26
16:51:15.886 -> Water_Temp.:93.31, PWM:26
16:51:16.647 -> Water_Temp.:93.31, PWM:26
16:51:18.196 -> Water_Temp.:93.31, PWM:27
16:51:18.989 -> Water_Temp.:93.31, PWM:26
16:51:19.752 -> Water_Temp.:93.31, PWM:26
16:51:20.549 -> Water_Temp.:93.31, PWM:26
16:51:21.309 -> Water_Temp.:93.43, PWM:26
16:51:22.107 -> Water_Temp.:93.43, PWM:26
16:51:23.656 -> Water_Temp.:93.43, PWM:27
16:51:24.449 -> Water_Temp.:93.43, PWM:27
16:51:25.208 -> Water_Temp.:93.43, PWM:26
16:51:25.996 -> Water_Temp.:93.31, PWM:27
16:51:26.788 -> Water_Temp.:93.31, PWM:27
16:51:27.546 -> Water_Temp.:93.31, PWM:26
16:51:29.101 -> Water_Temp.:93.31, PWM:27
16:51:29.897 -> Water_Temp.:93.31, PWM:26
16:51:30.656 -> Water_Temp.:93.31, PWM:26
16:51:31.449 -> Water_Temp.:93.31, PWM:26
16:51:32.241 -> Water_Temp.:93.31, PWM:26
16:51:33.003 -> Water_Temp.:93.31, PWM:26
16:51:33.801 -> Water_Temp.:93.31, PWM:26
16:51:34.561 -> Water_Temp.:93.31, PWM:26
16:51:35.359 -> Water_Temp.:93.31, PWM:26
16:51:36.118 -> Water_Temp.:93.31, PWM:26
16:51:36.914 -> Water_Temp.:93.31, PWM:26
16:51:37.708 -> Water_Temp.:93.31, PWM:27
16:51:38.466 -> Water_Temp.:93.31, PWM:27
16:51:39.263 -> Water_Temp.:93.31, PWM:27
16:51:40.023 -> Water_Temp.:93.31, PWM:27
16:51:40.815 -> Water_Temp.:93.31, PWM:26
16:51:41.576 -> Water_Temp.:93.31, PWM:27
16:51:42.366 -> Water_Temp.:93.31, PWM:27
16:51:43.160 -> Water_Temp.:93.31, PWM:27
16:51:43.921 -> Water_Temp.:93.31, PWM:26
16:51:44.715 -> Water_Temp.:93.31, PWM:27
16:51:45.479 -> Water_Temp.:93.31, PWM:27
16:51:46.273 -> Water_Temp.:93.43, PWM:27
16:51:47.798 -> Water_Temp.:93.31, PWM:26
16:51:48.594 -> Water_Temp.:93.31, PWM:26
16:51:49.358 -> Water_Temp.:93.31, PWM:26
16:51:50.156 -> Water_Temp.:93.43, PWM:27
16:51:50.920 -> Water_Temp.:93.43, PWM:27
16:51:51.713 -> Water_Temp.:93.43, PWM:26
16:51:52.510 -> Water_Temp.:93.31, PWM:26
16:51:53.270 -> Water_Temp.:93.31, PWM:26
16:51:54.061 -> Water_Temp.:93.43, PWM:26
16:51:55.611 -> Water_Temp.:93.43, PWM:26
16:51:56.408 -> Water_Temp.:93.43, PWM:26
16:51:57.165 -> Water_Temp.:93.43, PWM:26
16:51:57.963 -> Water_Temp.:93.31, PWM:26
16:51:58.723 -> Water_Temp.:93.43, PWM:26
16:51:59.517 -> Water_Temp.:93.43, PWM:27
16:52:00.307 -> Water_Temp.:93.43, PWM:26
16:52:01.068 -> Water_Temp.:93.43, PWM:26
16:52:01.859 -> Water_Temp.:93.43, PWM:26
16:52:02.622 -> Water_Temp.:93.43, PWM:26
16:52:03.413 -> Water_Temp.:93.43, PWM:26
16:52:04.176 -> Water_Temp.:93.31, PWM:26
16:52:04.972 -> Water_Temp.:93.31, PWM:26
16:52:05.767 -> Water_Temp.:93.31, PWM:26
16:52:06.530 -> Water_Temp.:93.43, PWM:26
16:52:08.087 -> Water_Temp.:93.43, PWM:27
16:52:09.637 -> Water_Temp.:93.31, PWM:27
16:52:10.397 -> Water_Temp.:93.31, PWM:27
16:52:11.190 -> Water_Temp.:93.31, PWM:26
16:52:11.983 -> Water_Temp.:93.43, PWM:26
16:52:12.743 -> Water_Temp.:93.43, PWM:27
16:52:13.536 -> Water_Temp.:93.43, PWM:26
16:52:14.323 -> Water_Temp.:93.43, PWM:27
16:52:15.082 -> Water_Temp.:93.31, PWM:26
16:52:15.872 -> Water_Temp.:93.43, PWM:26
16:52:16.634 -> Water_Temp.:93.43, PWM:27
16:52:17.426 -> Water_Temp.:93.43, PWM:27
16:52:18.222 -> Water_Temp.:93.43, PWM:26
16:52:18.978 -> Water_Temp.:93.43, PWM:27
16:52:19.776 -> Water_Temp.:93.43, PWM:27
16:52:20.566 -> Water_Temp.:93.43, PWM:26
16:52:21.325 -> Water_Temp.:93.43, PWM:26
16:52:22.118 -> Water_Temp.:93.43, PWM:27
16:52:22.880 -> Water_Temp.:93.43, PWM:27
16:52:23.677 -> Water_Temp.:93.43, PWM:27
16:52:24.439 -> Water_Temp.:93.43, PWM:26
16:52:25.231 -> Water_Temp.:93.43, PWM:26
16:52:26.779 -> Water_Temp.:93.43, PWM:26
16:52:27.539 -> Water_Temp.:93.43, PWM:26
16:52:29.128 -> Water_Temp.:93.43, PWM:27
16:52:29.884 -> Water_Temp.:93.43, PWM:27
16:52:30.677 -> Water_Temp.:93.43, PWM:26
16:52:31.441 -> Water_Temp.:93.43, PWM:27
16:52:32.235 -> Water_Temp.:93.43, PWM:27
16:52:33.031 -> Water_Temp.:93.43, PWM:26
16:52:33.795 -> Water_Temp.:93.43, PWM:26
16:52:34.588 -> Water_Temp.:93.43, PWM:26
16:52:35.350 -> Water_Temp.:93.43, PWM:27
16:52:36.143 -> Water_Temp.:93.43, PWM:26
16:52:37.695 -> Water_Temp.:93.54, PWM:27
16:52:38.488 -> Water_Temp.:93.54, PWM:27
16:52:39.249 -> Water_Temp.:93.54, PWM:26
16:52:40.040 -> Water_Temp.:93.54, PWM:27
16:52:40.803 -> Water_Temp.:95.11, PWM:26
16:52:41.595 -> Water_Temp.:97.14, PWM:26
16:52:42.386 -> Water_Temp.:94.77, PWM:26
16:52:43.145 -> Water_Temp.:94.32, PWM:26
16:52:43.938 -> Water_Temp.:94.21, PWM:27
16:52:44.702 -> Water_Temp.:93.99, PWM:27
16:52:45.495 -> Water_Temp.:93.99, PWM:27
16:52:46.257 -> Water_Temp.:93.87, PWM:26
16:52:47.050 -> Water_Temp.:93.87, PWM:26
16:52:47.846 -> Water_Temp.:93.76, PWM:26
16:52:48.605 -> Water_Temp.:93.76, PWM:26
16:52:49.396 -> Water_Temp.:93.76, PWM:27
16:52:50.155 -> Water_Temp.:93.76, PWM:26
16:52:50.949 -> Water_Temp.:93.65, PWM:26
16:52:51.744 -> Water_Temp.:93.65, PWM:27
16:52:52.503 -> Water_Temp.:93.65, PWM:26
16:52:53.300 -> Water_Temp.:93.65, PWM:26
16:52:54.821 -> Water_Temp.:93.65, PWM:26
16:52:55.615 -> Water_Temp.:93.65, PWM:27
16:52:56.376 -> Water_Temp.:93.65, PWM:27
16:52:57.169 -> Water_Temp.:93.65, PWM:26
16:52:57.932 -> Water_Temp.:93.65, PWM:27
16:52:58.730 -> Water_Temp.:93.65, PWM:26
16:52:59.523 -> Water_Temp.:93.65, PWM:27
16:53:00.285 -> Water_Temp.:93.65, PWM:27
16:53:01.083 -> Water_Temp.:93.65, PWM:26
16:53:01.843 -> Water_Temp.:93.65, PWM:26
16:53:02.634 -> Water_Temp.:93.65, PWM:27
16:53:03.392 -> Water_Temp.:93.65, PWM:27
16:53:04.186 -> Water_Temp.:93.65, PWM:26
16:53:04.982 -> Water_Temp.:93.65, PWM:27
16:53:05.746 -> Water_Temp.:93.65, PWM:26
16:53:06.542 -> Water_Temp.:93.65, PWM:26
16:53:07.298 -> Water_Temp.:93.65, PWM:26
16:53:08.091 -> Water_Temp.:93.65, PWM:26
16:53:08.853 -> Water_Temp.:93.65, PWM:27
16:53:09.647 -> Water_Temp.:93.65, PWM:26
16:53:10.445 -> Water_Temp.:93.65, PWM:27
16:53:11.202 -> Water_Temp.:93.65, PWM:26
16:53:11.999 -> Water_Temp.:93.65, PWM:27
16:53:12.758 -> Water_Temp.:93.65, PWM:27
16:53:13.548 -> Water_Temp.:93.65, PWM:26
16:53:14.342 -> Water_Temp.:93.65, PWM:27
16:53:15.102 -> Water_Temp.:93.65, PWM:27
16:53:15.892 -> Water_Temp.:93.65, PWM:27
16:53:16.653 -> Water_Temp.:93.65, PWM:26
16:53:17.445 -> Water_Temp.:93.65, PWM:26
16:53:18.241 -> Water_Temp.:93.65, PWM:27
16:53:18.999 -> Water_Temp.:93.65, PWM:26
16:53:19.792 -> Water_Temp.:93.65, PWM:26
16:53:20.586 -> Water_Temp.:93.65, PWM:27
16:53:22.138 -> Water_Temp.:93.65, PWM:26
16:53:23.695 -> Water_Temp.:93.65, PWM:27
16:53:24.454 -> Water_Temp.:93.65, PWM:26
16:53:25.250 -> Water_Temp.:93.65, PWM:27
16:53:26.012 -> Water_Temp.:93.65, PWM:26
16:53:26.809 -> Water_Temp.:93.65, PWM:26
16:53:27.603 -> Water_Temp.:93.65, PWM:26
16:53:28.364 -> Water_Temp.:93.65, PWM:26
16:53:29.156 -> Water_Temp.:93.65, PWM:26
16:53:29.916 -> Water_Temp.:93.65, PWM:27
16:53:30.715 -> Water_Temp.:93.65, PWM:26
16:53:32.239 -> Water_Temp.:93.65, PWM:27
16:53:33.032 -> Water_Temp.:93.65, PWM:27
16:53:33.790 -> Water_Temp.:93.65, PWM:26
16:53:34.587 -> Water_Temp.:93.65, PWM:27
16:53:35.348 -> Water_Temp.:93.76, PWM:26
16:53:37.665 -> Water_Temp.:93.65, PWM:27
16:53:38.461 -> Water_Temp.:93.76, PWM:27
16:53:39.218 -> Water_Temp.:93.76, PWM:27
16:53:40.012 -> Water_Temp.:93.76, PWM:27
16:53:40.775 -> Water_Temp.:93.76, PWM:27
16:53:41.561 -> Water_Temp.:93.76, PWM:27
16:53:42.358 -> Water_Temp.:93.76, PWM:26
16:53:43.117 -> Water_Temp.:93.76, PWM:27
16:53:43.914 -> Water_Temp.:93.65, PWM:27
16:53:44.673 -> Water_Temp.:93.76, PWM:27
16:53:45.472 -> Water_Temp.:93.76, PWM:27
16:53:47.032 -> Water_Temp.:93.76, PWM:26
16:53:48.582 -> Water_Temp.:93.76, PWM:27
16:53:49.382 -> Water_Temp.:93.76, PWM:26
16:53:50.143 -> Water_Temp.:93.76, PWM:26
16:53:50.938 -> Water_Temp.:93.76, PWM:26
16:53:51.702 -> Water_Temp.:93.76, PWM:26
16:53:52.498 -> Water_Temp.:93.76, PWM:26
16:53:53.252 -> Water_Temp.:93.76, PWM:26
16:53:54.041 -> Water_Temp.:93.76, PWM:26
16:53:54.836 -> Water_Temp.:93.76, PWM:27
16:53:55.598 -> Water_Temp.:93.76, PWM:27
16:53:56.392 -> Water_Temp.:93.76, PWM:27
16:53:57.150 -> Water_Temp.:93.76, PWM:27
16:53:57.945 -> Water_Temp.:93.76, PWM:27
16:53:58.738 -> Water_Temp.:93.76, PWM:26
16:53:59.501 -> Water_Temp.:93.76, PWM:27
16:54:00.294 -> Water_Temp.:93.76, PWM:27
16:54:01.054 -> Water_Temp.:93.87, PWM:27
16:54:01.845 -> Water_Temp.:93.87, PWM:26
16:54:02.635 -> Water_Temp.:93.87, PWM:26
16:54:03.396 -> Water_Temp.:93.87, PWM:26
16:54:04.187 -> Water_Temp.:93.87, PWM:27
16:54:04.948 -> Water_Temp.:93.87, PWM:26
16:54:06.503 -> Water_Temp.:93.87, PWM:26
16:54:07.288 -> Water_Temp.:93.87, PWM:26
16:54:08.080 -> Water_Temp.:93.87, PWM:27
16:54:08.836 -> Water_Temp.:93.87, PWM:26
16:54:09.625 -> Water_Temp.:93.87, PWM:27
16:54:10.416 -> Water_Temp.:93.87, PWM:26
16:54:11.176 -> Water_Temp.:93.87, PWM:27
16:54:11.967 -> Water_Temp.:93.87, PWM:26
16:54:12.762 -> Water_Temp.:93.87, PWM:26
16:54:13.526 -> Water_Temp.:93.87, PWM:26
16:54:14.319 -> Water_Temp.:93.87, PWM:26
16:54:15.076 -> Water_Temp.:93.87, PWM:26
16:54:15.871 -> Water_Temp.:93.87, PWM:27
16:54:16.635 -> Water_Temp.:93.87, PWM:27
16:54:17.428 -> Water_Temp.:93.87, PWM:26
16:54:18.982 -> Water_Temp.:93.87, PWM:26
16:54:19.773 -> Water_Temp.:93.87, PWM:26
16:54:21.324 -> Water_Temp.:93.87, PWM:26
16:54:22.878 -> Water_Temp.:93.99, PWM:26
16:54:23.637 -> Water_Temp.:93.99, PWM:27
16:54:25.198 -> Water_Temp.:93.99, PWM:27
16:54:25.993 -> Water_Temp.:93.99, PWM:26
16:54:26.756 -> Water_Temp.:93.99, PWM:26
16:54:27.551 -> Water_Temp.:93.99, PWM:26
16:54:28.342 -> Water_Temp.:93.99, PWM:27
16:54:29.100 -> Water_Temp.:93.99, PWM:27
16:54:30.650 -> Water_Temp.:93.99, PWM:26
16:54:31.445 -> Water_Temp.:93.99, PWM:27
16:54:32.236 -> Water_Temp.:93.99, PWM:26
16:54:32.993 -> Water_Temp.:93.99, PWM:26
16:54:33.788 -> Water_Temp.:93.99, PWM:27
16:54:34.552 -> Water_Temp.:93.99, PWM:27
16:54:35.350 -> Water_Temp.:93.99, PWM:27
16:54:36.141 -> Water_Temp.:93.99, PWM:26
16:54:37.664 -> Water_Temp.:93.99, PWM:27
16:54:38.460 -> Water_Temp.:93.99, PWM:26
16:54:39.221 -> Water_Temp.:93.99, PWM:26
16:54:40.016 -> Water_Temp.:93.99, PWM:26
16:54:41.567 -> Water_Temp.:93.99, PWM:26
16:54:42.323 -> Water_Temp.:93.99, PWM:26
16:54:43.117 -> Water_Temp.:93.99, PWM:26
16:54:43.914 -> Water_Temp.:93.99, PWM:27
16:54:44.673 -> Water_Temp.:93.99, PWM:26
16:54:45.463 -> Water_Temp.:93.99, PWM:26
16:54:46.221 -> Water_Temp.:93.99, PWM:26
16:54:47.013 -> Water_Temp.:93.99, PWM:26
16:54:47.807 -> Water_Temp.:93.99, PWM:27
16:54:48.569 -> Water_Temp.:93.99, PWM:26
16:54:49.362 -> Water_Temp.:93.99, PWM:26
16:54:50.150 -> Water_Temp.:93.99, PWM:26
16:54:50.909 -> Water_Temp.:93.99, PWM:26
16:54:51.705 -> Water_Temp.:93.99, PWM:27
16:54:52.495 -> Water_Temp.:93.99, PWM:27
16:54:53.256 -> Water_Temp.:93.99, PWM:26
16:54:54.051 -> Water_Temp.:93.99, PWM:26
16:54:54.817 -> Water_Temp.:93.99, PWM:26
16:54:55.612 -> Water_Temp.:93.99, PWM:26
16:54:56.378 -> Water_Temp.:93.99, PWM:27
16:54:57.143 -> Water_Temp.:93.99, PWM:27
16:54:57.937 -> Water_Temp.:95.68, PWM:26
16:54:58.702 -> Water_Temp.:97.70, PWM:27
16:54:59.493 -> Water_Temp.:95.34, PWM:27
16:55:00.285 -> Water_Temp.:94.89, PWM:26
16:55:01.040 -> Water_Temp.:94.66, PWM:27
16:55:01.835 -> Water_Temp.:94.55, PWM:26
16:55:02.625 -> Water_Temp.:94.44, PWM:26
16:55:03.383 -> Water_Temp.:94.44, PWM:27
16:55:04.935 -> Water_Temp.:94.32, PWM:26
16:55:05.733 -> Water_Temp.:94.32, PWM:26
16:55:06.494 -> Water_Temp.:94.21, PWM:26
16:55:07.286 -> Water_Temp.:94.21, PWM:27
16:55:08.078 -> Water_Temp.:94.21, PWM:27
16:55:08.841 -> Water_Temp.:94.21, PWM:26
16:55:09.632 -> Water_Temp.:94.21, PWM:27
16:55:10.392 -> Water_Temp.:94.21, PWM:26
16:55:11.946 -> Water_Temp.:94.21, PWM:27
16:55:12.705 -> Water_Temp.:94.21, PWM:26
16:55:13.494 -> Water_Temp.:94.21, PWM:26
16:55:14.290 -> Water_Temp.:94.21, PWM:26
16:55:15.052 -> Water_Temp.:94.21, PWM:27
16:55:15.849 -> Water_Temp.:94.10, PWM:27
16:55:17.362 -> Water_Temp.:94.21, PWM:27
16:55:18.158 -> Water_Temp.:94.10, PWM:26
16:55:22.026 -> Water_Temp.:94.10, PWM:26
16:55:23.578 -> Water_Temp.:94.10, PWM:26
16:55:24.369 -> Water_Temp.:94.10, PWM:27
16:55:25.132 -> Water_Temp.:94.10, PWM:26
16:55:25.923 -> Water_Temp.:94.10, PWM:27
16:55:26.720 -> Water_Temp.:94.10, PWM:26
16:55:27.475 -> Water_Temp.:94.10, PWM:27
16:55:28.268 -> Water_Temp.:94.10, PWM:27
16:55:29.031 -> Water_Temp.:94.10, PWM:27
16:55:29.825 -> Water_Temp.:94.21, PWM:27
16:55:30.620 -> Water_Temp.:94.10, PWM:26
16:55:31.384 -> Water_Temp.:94.10, PWM:26
16:55:32.177 -> Water_Temp.:94.21, PWM:27
16:55:32.934 -> Water_Temp.:94.21, PWM:27
16:55:34.523 -> Water_Temp.:94.21, PWM:27
16:55:35.280 -> Water_Temp.:94.21, PWM:27
16:55:36.071 -> Water_Temp.:94.21, PWM:27
16:55:36.832 -> Water_Temp.:94.21, PWM:26
16:55:37.627 -> Water_Temp.:94.21, PWM:26
16:55:38.422 -> Water_Temp.:94.21, PWM:26
16:55:39.178 -> Water_Temp.:94.21, PWM:26
16:55:39.970 -> Water_Temp.:94.21, PWM:26
16:55:40.730 -> Water_Temp.:94.21, PWM:26
16:55:41.523 -> Water_Temp.:94.21, PWM:26
16:55:42.319 -> Water_Temp.:94.21, PWM:26
16:55:43.073 -> Water_Temp.:94.21, PWM:26
16:55:43.864 -> Water_Temp.:94.21, PWM:27
16:55:44.657 -> Water_Temp.:94.21, PWM:27
16:55:45.419 -> Water_Temp.:94.21, PWM:27
16:55:46.216 -> Water_Temp.:94.21, PWM:27
16:55:46.974 -> Water_Temp.:94.21, PWM:26
16:55:47.770 -> Water_Temp.:94.21, PWM:26
16:55:48.532 -> Water_Temp.:94.21, PWM:26
16:55:50.089 -> Water_Temp.:94.21, PWM:27
16:55:50.885 -> Water_Temp.:94.21, PWM:26
16:55:51.648 -> Water_Temp.:94.21, PWM:26
16:55:52.442 -> Water_Temp.:94.21, PWM:26
16:55:53.235 -> Water_Temp.:94.21, PWM:26
16:55:53.990 -> Water_Temp.:94.21, PWM:26
16:55:54.782 -> Water_Temp.:94.21, PWM:26
16:55:55.579 -> Water_Temp.:94.21, PWM:26
16:55:56.338 -> Water_Temp.:94.21, PWM:26
16:55:57.131 -> Water_Temp.:94.21, PWM:26
16:55:57.892 -> Water_Temp.:94.21, PWM:26
16:55:58.687 -> Water_Temp.:95.79, PWM:27
16:55:59.445 -> Water_Temp.:97.81, PWM:27
16:56:00.242 -> Water_Temp.:95.45, PWM:26
16:56:01.038 -> Water_Temp.:95.11, PWM:26
16:56:01.800 -> Water_Temp.:94.89, PWM:26
16:56:02.593 -> Water_Temp.:94.77, PWM:26
16:56:03.353 -> Water_Temp.:94.66, PWM:27
16:56:04.150 -> Water_Temp.:94.55, PWM:27
16:56:04.912 -> Water_Temp.:94.55, PWM:26
16:56:05.705 -> Water_Temp.:94.55, PWM:27
16:56:06.498 -> Water_Temp.:94.44, PWM:26
16:56:07.257 -> Water_Temp.:94.44, PWM:26
16:56:08.051 -> Water_Temp.:94.44, PWM:26
16:56:09.598 -> Water_Temp.:94.44, PWM:26
16:56:11.155 -> Water_Temp.:94.32, PWM:26
16:56:11.913 -> Water_Temp.:94.32, PWM:26
16:56:12.704 -> Water_Temp.:94.32, PWM:27
16:56:13.465 -> Water_Temp.:94.32, PWM:26
16:56:15.013 -> Water_Temp.:94.32, PWM:27
16:56:16.606 -> Water_Temp.:94.32, PWM:26
16:56:17.363 -> Water_Temp.:94.32, PWM:26
16:56:18.158 -> Water_Temp.:94.32, PWM:26
16:56:18.920 -> Water_Temp.:94.32, PWM:26
16:56:20.503 -> Water_Temp.:94.32, PWM:27
16:56:21.262 -> Water_Temp.:94.32, PWM:26
16:56:22.056 -> Water_Temp.:94.32, PWM:26
16:56:22.813 -> Water_Temp.:94.32, PWM:27
16:56:23.606 -> Water_Temp.:94.32, PWM:26
16:56:24.396 -> Water_Temp.:94.32, PWM:26
16:56:25.156 -> Water_Temp.:94.32, PWM:27
16:56:25.948 -> Water_Temp.:94.32, PWM:26
16:56:26.711 -> Water_Temp.:94.32, PWM:27
16:56:27.508 -> Water_Temp.:94.32, PWM:27
16:56:28.303 -> Water_Temp.:94.32, PWM:26
16:56:29.066 -> Water_Temp.:94.32, PWM:27
16:56:29.860 -> Water_Temp.:94.32, PWM:26
16:56:30.620 -> Water_Temp.:94.32, PWM:27
16:56:31.409 -> Water_Temp.:94.32, PWM:27
16:56:32.201 -> Water_Temp.:94.32, PWM:27
16:56:32.962 -> Water_Temp.:94.32, PWM:26
16:56:33.755 -> Water_Temp.:94.32, PWM:26
16:56:34.514 -> Water_Temp.:94.32, PWM:26
16:56:35.309 -> Water_Temp.:94.32, PWM:26
16:56:36.069 -> Water_Temp.:94.32, PWM:27
16:56:36.863 -> Water_Temp.:94.32, PWM:27
16:56:37.656 -> Water_Temp.:94.32, PWM:27
16:56:38.415 -> Water_Temp.:94.32, PWM:26
16:56:39.211 -> Water_Temp.:94.32, PWM:27
16:56:39.971 -> Water_Temp.:94.32, PWM:26
16:56:40.765 -> Water_Temp.:94.32, PWM:26
16:56:41.558 -> Water_Temp.:94.32, PWM:26
16:56:42.316 -> Water_Temp.:94.32, PWM:27
16:56:43.111 -> Water_Temp.:94.32, PWM:27
16:56:43.873 -> Water_Temp.:94.32, PWM:26
16:56:44.664 -> Water_Temp.:94.32, PWM:27
16:56:45.457 -> Water_Temp.:95.90, PWM:26
16:56:46.221 -> Water_Temp.:97.92, PWM:27
16:56:47.018 -> Water_Temp.:95.56, PWM:27
16:56:47.777 -> Water_Temp.:95.22, PWM:27
16:56:48.571 -> Water_Temp.:95.00, PWM:26
16:56:49.328 -> Water_Temp.:94.89, PWM:27
16:56:50.912 -> Water_Temp.:94.66, PWM:26
16:56:51.670 -> Water_Temp.:94.66, PWM:27
16:56:52.459 -> Water_Temp.:94.55, PWM:27
16:56:53.255 -> Water_Temp.:94.55, PWM:26
16:56:54.011 -> Water_Temp.:94.55, PWM:26
16:56:54.808 -> Water_Temp.:94.55, PWM:27
16:56:55.598 -> Water_Temp.:96.12, PWM:26
16:56:56.360 -> Water_Temp.:98.15, PWM:26
16:56:57.151 -> Water_Temp.:95.79, PWM:26
16:56:57.912 -> Water_Temp.:95.34, PWM:27
16:56:58.702 -> Water_Temp.:95.11, PWM:26
16:56:59.499 -> Water_Temp.:95.00, PWM:26
16:57:00.258 -> Water_Temp.:94.89, PWM:26
16:57:01.056 -> Water_Temp.:94.77, PWM:26
16:57:01.815 -> Water_Temp.:94.77, PWM:27
16:57:02.607 -> Water_Temp.:94.66, PWM:26
16:57:03.368 -> Water_Temp.:94.66, PWM:26
16:57:04.164 -> Water_Temp.:94.66, PWM:27
16:57:05.689 -> Water_Temp.:94.55, PWM:26
16:57:06.482 -> Water_Temp.:94.55, PWM:26
16:57:07.239 -> Water_Temp.:94.55, PWM:27
16:57:08.033 -> Water_Temp.:94.55, PWM:26
16:57:08.822 -> Water_Temp.:94.55, PWM:26
16:57:11.136 -> Water_Temp.:94.55, PWM:27
16:57:11.895 -> Water_Temp.:94.55, PWM:27
16:57:12.686 -> Water_Temp.:94.55, PWM:26
16:57:13.483 -> Water_Temp.:94.55, PWM:26
16:57:14.242 -> Water_Temp.:94.55, PWM:27
16:57:15.037 -> Water_Temp.:94.55, PWM:27
16:57:15.799 -> Water_Temp.:94.55, PWM:26
16:57:16.591 -> Water_Temp.:94.55, PWM:26
16:57:17.351 -> Water_Temp.:94.55, PWM:26
16:57:18.143 -> Water_Temp.:94.55, PWM:26
16:57:18.936 -> Water_Temp.:94.55, PWM:27
16:57:19.723 -> Water_Temp.:94.55, PWM:26
16:57:20.478 -> Water_Temp.:94.55, PWM:27
16:57:21.277 -> Water_Temp.:94.55, PWM:27
16:57:22.033 -> Water_Temp.:94.55, PWM:27
16:57:22.827 -> Water_Temp.:94.44, PWM:26
16:57:23.619 -> Water_Temp.:94.44, PWM:26
16:57:24.376 -> Water_Temp.:94.55, PWM:27
16:57:25.169 -> Water_Temp.:94.55, PWM:27
16:57:25.959 -> Water_Temp.:94.44, PWM:27
16:57:26.719 -> Water_Temp.:94.55, PWM:27
16:57:27.514 -> Water_Temp.:94.44, PWM:27
16:57:28.274 -> Water_Temp.:94.44, PWM:26
16:57:29.068 -> Water_Temp.:94.44, PWM:27
16:57:29.860 -> Water_Temp.:94.55, PWM:26
16:57:31.374 -> Water_Temp.:94.44, PWM:26
16:57:32.930 -> Water_Temp.:94.44, PWM:26
16:57:33.723 -> Water_Temp.:94.44, PWM:26
16:57:34.486 -> Water_Temp.:94.44, PWM:27
16:57:35.279 -> Water_Temp.:94.44, PWM:26
16:57:36.074 -> Water_Temp.:94.44, PWM:27
16:57:36.834 -> Water_Temp.:94.44, PWM:27
16:57:37.628 -> Water_Temp.:94.44, PWM:27
16:57:39.143 -> Water_Temp.:94.44, PWM:26
16:57:39.936 -> Water_Temp.:94.44, PWM:26
16:57:40.726 -> Water_Temp.:94.44, PWM:27
16:57:41.486 -> Water_Temp.:94.44, PWM:26
16:57:42.282 -> Water_Temp.:94.44, PWM:26
16:57:43.043 -> Water_Temp.:94.44, PWM:26
16:57:43.836 -> Water_Temp.:94.44, PWM:27
16:57:44.596 -> Water_Temp.:94.44, PWM:26
16:57:45.393 -> Water_Temp.:94.44, PWM:26
16:57:46.186 -> Water_Temp.:94.44, PWM:27
16:57:46.945 -> Water_Temp.:94.44, PWM:26
16:57:47.738 -> Water_Temp.:94.44, PWM:27
16:57:48.499 -> Water_Temp.:94.44, PWM:27
16:57:49.292 -> Water_Temp.:94.44, PWM:26
16:57:50.086 -> Water_Temp.:94.44, PWM:27
16:57:50.849 -> Water_Temp.:94.44, PWM:27
16:57:51.642 -> Water_Temp.:94.44, PWM:26
16:57:52.396 -> Water_Temp.:94.44, PWM:26
16:57:53.189 -> Water_Temp.:94.44, PWM:27
16:57:53.984 -> Water_Temp.:94.44, PWM:26
16:57:54.748 -> Water_Temp.:94.44, PWM:27
16:57:55.537 -> Water_Temp.:94.44, PWM:27
16:57:56.299 -> Water_Temp.:94.44, PWM:27
16:57:57.096 -> Water_Temp.:94.44, PWM:26
16:57:57.856 -> Water_Temp.:94.44, PWM:26
16:57:58.651 -> Water_Temp.:94.44, PWM:26
16:57:59.415 -> Water_Temp.:94.44, PWM:26
16:58:00.207 -> Water_Temp.:94.44, PWM:26
16:58:01.005 -> Water_Temp.:94.44, PWM:27
16:58:01.765 -> Water_Temp.:94.55, PWM:26
16:58:02.559 -> Water_Temp.:94.55, PWM:27
16:58:03.319 -> Water_Temp.:94.44, PWM:27
16:58:04.113 -> Water_Temp.:94.44, PWM:27
16:58:04.903 -> Water_Temp.:94.55, PWM:27
16:58:05.667 -> Water_Temp.:94.44, PWM:27
16:58:06.465 -> Water_Temp.:94.55, PWM:26
16:58:07.220 -> Water_Temp.:94.55, PWM:27
16:58:08.018 -> Water_Temp.:94.44, PWM:26
16:58:08.775 -> Water_Temp.:94.44, PWM:27
16:58:09.572 -> Water_Temp.:94.55, PWM:27
16:58:11.126 -> Water_Temp.:94.55, PWM:26
16:58:11.923 -> Water_Temp.:94.55, PWM:26
16:58:12.687 -> Water_Temp.:94.55, PWM:26
16:58:13.484 -> Water_Temp.:94.55, PWM:27
16:58:14.248 -> Water_Temp.:94.55, PWM:26
16:58:15.043 -> Water_Temp.:94.55, PWM:26
16:58:15.804 -> Water_Temp.:94.55, PWM:26
16:58:16.598 -> Water_Temp.:94.55, PWM:26
16:58:17.361 -> Water_Temp.:96.12, PWM:27
16:58:18.156 -> Water_Temp.:98.15, PWM:26
16:58:18.918 -> Water_Temp.:95.79, PWM:27
16:58:19.715 -> Water_Temp.:95.34, PWM:26
16:58:20.473 -> Water_Temp.:95.22, PWM:27
16:58:21.272 -> Water_Temp.:95.00, PWM:26
16:58:22.032 -> Water_Temp.:95.00, PWM:27
16:58:22.829 -> Water_Temp.:94.89, PWM:26
16:58:23.594 -> Water_Temp.:94.89, PWM:27
16:58:24.389 -> Water_Temp.:94.77, PWM:27
16:58:25.153 -> Water_Temp.:94.77, PWM:26
16:58:25.954 -> Water_Temp.:94.77, PWM:27
16:58:26.718 -> Water_Temp.:94.66, PWM:27
16:58:27.508 -> Water_Temp.:94.66, PWM:26
16:58:28.270 -> Water_Temp.:94.66, PWM:26
16:58:29.063 -> Water_Temp.:96.35, PWM:26
16:58:29.858 -> Water_Temp.:98.26, PWM:26
16:58:30.616 -> Water_Temp.:95.90, PWM:26
16:58:31.408 -> Water_Temp.:95.56, PWM:26
16:58:32.201 -> Water_Temp.:95.34, PWM:26
16:58:32.961 -> Water_Temp.:95.11, PWM:26
16:58:33.753 -> Water_Temp.:95.00, PWM:26
16:58:34.513 -> Water_Temp.:95.00, PWM:27
16:58:35.310 -> Water_Temp.:94.89, PWM:27
16:58:36.103 -> Water_Temp.:94.89, PWM:27
16:58:36.861 -> Water_Temp.:94.77, PWM:27
16:58:37.655 -> Water_Temp.:94.77, PWM:27
16:58:38.418 -> Water_Temp.:94.77, PWM:27
16:58:39.205 -> Water_Temp.:94.77, PWM:26
16:58:39.995 -> Water_Temp.:94.66, PWM:27
16:58:40.753 -> Water_Temp.:94.66, PWM:27
16:58:41.548 -> Water_Temp.:94.66, PWM:27
16:58:42.340 -> Water_Temp.:94.66, PWM:26
16:58:43.892 -> Water_Temp.:94.66, PWM:26
16:58:44.654 -> Water_Temp.:94.66, PWM:27
16:58:45.444 -> Water_Temp.:94.66, PWM:26
16:58:46.234 -> Water_Temp.:94.66, PWM:27
16:58:46.993 -> Water_Temp.:94.66, PWM:26
16:58:47.783 -> Water_Temp.:94.66, PWM:26
16:58:48.577 -> Water_Temp.:94.66, PWM:26
16:58:49.337 -> Water_Temp.:94.66, PWM:27
16:58:50.129 -> Water_Temp.:94.66, PWM:26
16:58:50.892 -> Water_Temp.:94.66, PWM:27
16:58:51.685 -> Water_Temp.:94.66, PWM:26
16:58:52.475 -> Water_Temp.:94.66, PWM:27
16:58:53.237 -> Water_Temp.:94.66, PWM:26
16:58:54.036 -> Water_Temp.:94.66, PWM:27
16:58:54.798 -> Water_Temp.:94.66, PWM:27
16:58:55.584 -> Water_Temp.:94.66, PWM:27
16:58:56.379 -> Water_Temp.:94.66, PWM:27
16:58:57.135 -> Water_Temp.:94.66, PWM:27
16:58:57.930 -> Water_Temp.:94.66, PWM:26
16:58:58.689 -> Water_Temp.:94.66, PWM:26
16:58:59.482 -> Water_Temp.:94.66, PWM:26
16:59:00.275 -> Water_Temp.:94.66, PWM:26
16:59:01.034 -> Water_Temp.:94.66, PWM:27
16:59:01.827 -> Water_Temp.:94.66, PWM:26
16:59:02.621 -> Water_Temp.:94.55, PWM:26
16:59:03.383 -> Water_Temp.:94.66, PWM:26
16:59:04.178 -> Water_Temp.:94.55, PWM:27
16:59:04.936 -> Water_Temp.:94.66, PWM:26
16:59:05.728 -> Water_Temp.:94.66, PWM:26
16:59:06.490 -> Water_Temp.:94.66, PWM:26
16:59:07.285 -> Water_Temp.:94.66, PWM:26
16:59:08.077 -> Water_Temp.:94.55, PWM:26
16:59:08.834 -> Water_Temp.:94.66, PWM:26
16:59:09.631 -> Water_Temp.:94.66, PWM:26
16:59:10.416 -> Water_Temp.:94.66, PWM:26
16:59:11.177 -> Water_Temp.:94.66, PWM:26
16:59:11.972 -> Water_Temp.:94.66, PWM:26
16:59:12.732 -> Water_Temp.:94.66, PWM:27
16:59:13.524 -> Water_Temp.:94.66, PWM:26
16:59:14.316 -> Water_Temp.:94.66, PWM:27
16:59:15.072 -> Water_Temp.:94.66, PWM:27
16:59:15.868 -> Water_Temp.:94.66, PWM:26
16:59:16.661 -> Water_Temp.:94.66, PWM:26
16:59:18.178 -> Water_Temp.:94.66, PWM:27
16:59:18.971 -> Water_Temp.:94.66, PWM:26
16:59:19.730 -> Water_Temp.:94.66, PWM:27
16:59:20.527 -> Water_Temp.:94.66, PWM:27
16:59:21.289 -> Water_Temp.:94.66, PWM:26
16:59:22.080 -> Water_Temp.:94.66, PWM:27
16:59:22.838 -> Water_Temp.:94.66, PWM:27
16:59:23.633 -> Water_Temp.:94.66, PWM:27
16:59:24.426 -> Water_Temp.:94.66, PWM:26
16:59:25.191 -> Water_Temp.:94.66, PWM:26
16:59:25.986 -> Water_Temp.:94.66, PWM:27
16:59:26.746 -> Water_Temp.:94.66, PWM:27
16:59:27.538 -> Water_Temp.:94.66, PWM:26
16:59:28.331 -> Water_Temp.:94.66, PWM:27
16:59:29.090 -> Water_Temp.:94.66, PWM:26
16:59:29.884 -> Water_Temp.:94.66, PWM:26
16:59:30.643 -> Water_Temp.:94.66, PWM:26
16:59:31.437 -> Water_Temp.:94.66, PWM:26
16:59:32.229 -> Water_Temp.:94.66, PWM:27
16:59:32.986 -> Water_Temp.:94.66, PWM:26
16:59:33.777 -> Water_Temp.:94.66, PWM:27
16:59:35.328 -> Water_Temp.:94.66, PWM:26
16:59:36.123 -> Water_Temp.:94.66, PWM:27
16:59:37.672 -> Water_Temp.:94.66, PWM:27
16:59:38.468 -> Water_Temp.:94.66, PWM:26
16:59:39.230 -> Water_Temp.:94.66, PWM:27
16:59:40.021 -> Water_Temp.:94.66, PWM:26
16:59:40.811 -> Water_Temp.:94.66, PWM:26
16:59:41.571 -> Water_Temp.:94.66, PWM:26
16:59:42.365 -> Water_Temp.:94.66, PWM:26
16:59:43.122 -> Water_Temp.:94.66, PWM:26
16:59:43.918 -> Water_Temp.:94.66, PWM:27
16:59:44.710 -> Water_Temp.:94.66, PWM:26
16:59:45.472 -> Water_Temp.:94.66, PWM:26
16:59:46.236 -> Water_Temp.:94.66, PWM:26
16:59:47.032 -> Water_Temp.:94.66, PWM:26
16:59:47.829 -> Water_Temp.:94.66, PWM:27
16:59:48.592 -> Water_Temp.:94.66, PWM:27
16:59:49.383 -> Water_Temp.:94.66, PWM:26
16:59:50.141 -> Water_Temp.:94.66, PWM:26
16:59:50.935 -> Water_Temp.:94.66, PWM:26
16:59:51.730 -> Water_Temp.:94.66, PWM:26
16:59:52.487 -> Water_Temp.:94.66, PWM:26
16:59:53.281 -> Water_Temp.:94.66, PWM:27
16:59:54.043 -> Water_Temp.:94.77, PWM:26
16:59:54.835 -> Water_Temp.:94.77, PWM:27
16:59:55.596 -> Water_Temp.:94.77, PWM:27
16:59:57.151 -> Water_Temp.:94.77, PWM:26
16:59:57.945 -> Water_Temp.:94.77, PWM:26
16:59:58.703 -> Water_Temp.:94.77, PWM:26
16:59:59.491 -> Water_Temp.:94.77, PWM:27
17:00:00.286 -> Water_Temp.:94.77, PWM:26
17:00:01.044 -> Water_Temp.:94.77, PWM:26
17:00:01.841 -> Water_Temp.:94.77, PWM:27
17:00:02.605 -> Water_Temp.:94.77, PWM:26
17:00:03.393 -> Water_Temp.:94.77, PWM:27
17:00:04.186 -> Water_Temp.:94.77, PWM:27
17:00:04.947 -> Water_Temp.:94.77, PWM:26
17:00:05.733 -> Water_Temp.:94.77, PWM:26
17:00:06.522 -> Water_Temp.:94.77, PWM:26
17:00:07.281 -> Water_Temp.:94.77, PWM:26
17:00:08.073 -> Water_Temp.:94.77, PWM:27
17:00:08.865 -> Water_Temp.:94.77, PWM:26
17:00:09.629 -> Water_Temp.:94.77, PWM:26
17:00:10.421 -> Water_Temp.:94.77, PWM:27
17:00:11.181 -> Water_Temp.:94.77, PWM:26
17:00:11.977 -> Water_Temp.:94.77, PWM:27
17:00:13.499 -> Water_Temp.:94.77, PWM:26
17:00:14.288 -> Water_Temp.:94.77, PWM:26
17:00:15.845 -> Water_Temp.:94.77, PWM:27
17:00:16.642 -> Water_Temp.:94.77, PWM:27
17:00:17.403 -> Water_Temp.:94.77, PWM:27
17:00:18.194 -> Water_Temp.:94.77, PWM:26
17:00:18.985 -> Water_Temp.:94.77, PWM:26
17:00:19.742 -> Water_Temp.:94.77, PWM:27
17:00:21.285 -> Water_Temp.:94.77, PWM:27
17:00:22.076 -> Water_Temp.:94.77, PWM:27
17:00:22.839 -> Water_Temp.:94.77, PWM:26
17:00:23.633 -> Water_Temp.:94.77, PWM:26
17:00:24.392 -> Water_Temp.:94.77, PWM:26
17:00:25.187 -> Water_Temp.:94.77, PWM:26
17:00:25.948 -> Water_Temp.:94.77, PWM:27
17:00:26.739 -> Water_Temp.:94.77, PWM:26
17:00:28.290 -> Water_Temp.:94.77, PWM:26
17:00:29.080 -> Water_Temp.:94.77, PWM:26
17:00:29.869 -> Water_Temp.:94.77, PWM:27
17:00:30.630 -> Water_Temp.:94.77, PWM:26
17:00:31.422 -> Water_Temp.:94.77, PWM:26
17:00:32.212 -> Water_Temp.:94.77, PWM:26
17:00:32.972 -> Water_Temp.:94.77, PWM:26
17:00:33.765 -> Water_Temp.:94.89, PWM:26
17:00:34.559 -> Water_Temp.:94.89, PWM:26
17:00:35.309 -> Water_Temp.:94.89, PWM:27
17:00:36.100 -> Water_Temp.:94.89, PWM:26
17:00:36.893 -> Water_Temp.:94.89, PWM:26
17:00:37.649 -> Water_Temp.:94.89, PWM:27
17:00:38.443 -> Water_Temp.:94.89, PWM:27
17:00:39.234 -> Water_Temp.:94.89, PWM:26
17:00:39.991 -> Water_Temp.:94.89, PWM:26
17:00:40.786 -> Water_Temp.:94.89, PWM:26
17:00:41.579 -> Water_Temp.:94.89, PWM:27
17:00:42.336 -> Water_Temp.:94.89, PWM:26
17:00:43.130 -> Water_Temp.:94.89, PWM:26
17:00:43.889 -> Water_Temp.:94.89, PWM:26
17:00:44.683 -> Water_Temp.:94.89, PWM:26
17:00:45.473 -> Water_Temp.:94.89, PWM:27
17:00:46.231 -> Water_Temp.:94.89, PWM:27
17:00:47.025 -> Water_Temp.:94.89, PWM:27
17:00:47.817 -> Water_Temp.:94.89, PWM:27
17:00:49.333 -> Water_Temp.:94.89, PWM:26
17:00:50.125 -> Water_Temp.:94.89, PWM:26
17:00:50.921 -> Water_Temp.:94.89, PWM:26
17:00:51.676 -> Water_Temp.:94.89, PWM:26
17:00:52.464 -> Water_Temp.:94.89, PWM:26
17:00:54.021 -> Water_Temp.:94.89, PWM:26
17:00:54.817 -> Water_Temp.:94.89, PWM:26
17:00:55.580 -> Water_Temp.:94.89, PWM:27
17:00:56.375 -> Water_Temp.:94.89, PWM:27
17:00:57.137 -> Water_Temp.:94.89, PWM:27
17:00:57.932 -> Water_Temp.:94.89, PWM:27
17:00:58.691 -> Water_Temp.:94.89, PWM:26
17:00:59.484 -> Water_Temp.:95.00, PWM:26
17:01:01.036 -> Water_Temp.:95.00, PWM:26
17:01:01.793 -> Water_Temp.:94.89, PWM:26
17:01:02.583 -> Water_Temp.:94.89, PWM:26
17:01:03.375 -> Water_Temp.:95.00, PWM:27
17:01:04.892 -> Water_Temp.:94.89, PWM:26
17:01:05.683 -> Water_Temp.:94.89, PWM:26
17:01:06.478 -> Water_Temp.:94.89, PWM:27
17:01:07.239 -> Water_Temp.:94.89, PWM:27
17:01:08.034 -> Water_Temp.:94.89, PWM:27
17:01:08.791 -> Water_Temp.:94.89, PWM:27
17:01:09.586 -> Water_Temp.:94.89, PWM:26
17:01:10.380 -> Water_Temp.:94.89, PWM:27
17:01:11.142 -> Water_Temp.:94.89, PWM:26
17:01:11.931 -> Water_Temp.:94.89, PWM:26
17:01:12.692 -> Water_Temp.:94.89, PWM:26
17:01:13.483 -> Water_Temp.:94.89, PWM:26
17:01:14.275 -> Water_Temp.:94.89, PWM:26
17:01:15.034 -> Water_Temp.:94.89, PWM:26
17:01:15.830 -> Water_Temp.:94.89, PWM:26
17:01:16.591 -> Water_Temp.:94.89, PWM:27
17:01:17.385 -> Water_Temp.:94.89, PWM:27
17:01:18.147 -> Water_Temp.:94.89, PWM:26
17:01:18.943 -> Water_Temp.:94.89, PWM:27
17:01:19.738 -> Water_Temp.:94.89, PWM:27
17:01:20.498 -> Water_Temp.:94.89, PWM:26
17:01:21.288 -> Water_Temp.:94.89, PWM:27
17:01:22.047 -> Water_Temp.:94.89, PWM:26
17:01:22.841 -> Water_Temp.:94.89, PWM:26
17:01:23.635 -> Water_Temp.:94.89, PWM:27
17:01:24.394 -> Water_Temp.:94.89, PWM:26
17:01:25.186 -> Water_Temp.:94.89, PWM:26
17:01:25.978 -> Water_Temp.:94.89, PWM:26
17:01:26.740 -> Water_Temp.:94.89, PWM:26
17:01:27.535 -> Water_Temp.:94.89, PWM:26
17:01:28.294 -> Water_Temp.:95.00, PWM:26
17:01:29.085 -> Water_Temp.:94.89, PWM:26
17:01:29.878 -> Water_Temp.:95.00, PWM:26
17:01:30.637 -> Water_Temp.:94.89, PWM:27
17:01:31.427 -> Water_Temp.:94.89, PWM:27
17:01:32.188 -> Water_Temp.:94.89, PWM:27
17:01:32.985 -> Water_Temp.:95.00, PWM:26
17:01:33.776 -> Water_Temp.:94.89, PWM:26
17:01:34.537 -> Water_Temp.:94.89, PWM:26
17:01:36.086 -> Water_Temp.:94.89, PWM:26
17:01:36.878 -> Water_Temp.:95.00, PWM:26
17:01:37.638 -> Water_Temp.:94.89, PWM:26
17:01:38.423 -> Water_Temp.:95.00, PWM:27
17:01:39.214 -> Water_Temp.:95.00, PWM:26
17:01:39.971 -> Water_Temp.:95.00, PWM:26
17:01:40.763 -> Water_Temp.:95.00, PWM:27
17:01:41.555 -> Water_Temp.:95.00, PWM:26
17:01:42.312 -> Water_Temp.:95.00, PWM:26
17:01:43.105 -> Water_Temp.:95.00, PWM:26
17:01:43.898 -> Water_Temp.:95.00, PWM:27
17:01:44.658 -> Water_Temp.:95.00, PWM:27
17:01:45.451 -> Water_Temp.:95.00, PWM:27
17:01:46.209 -> Water_Temp.:95.00, PWM:26
17:01:47.002 -> Water_Temp.:95.00, PWM:26
17:01:47.796 -> Water_Temp.:95.00, PWM:26
17:01:48.552 -> Water_Temp.:95.00, PWM:27
17:01:49.345 -> Water_Temp.:95.00, PWM:27
17:01:50.138 -> Water_Temp.:95.00, PWM:27
17:01:50.894 -> Water_Temp.:95.00, PWM:26
17:01:51.689 -> Water_Temp.:95.00, PWM:26
17:01:52.445 -> Water_Temp.:95.00, PWM:26
17:01:53.239 -> Water_Temp.:95.00, PWM:26
17:01:54.032 -> Water_Temp.:95.00, PWM:26
17:01:54.791 -> Water_Temp.:95.00, PWM:26
17:01:55.583 -> Water_Temp.:95.00, PWM:26
17:01:56.376 -> Water_Temp.:95.00, PWM:26
17:01:57.139 -> Water_Temp.:95.00, PWM:26
17:01:57.936 -> Water_Temp.:95.00, PWM:27
17:01:58.697 -> Water_Temp.:95.00, PWM:27
17:01:59.487 -> Water_Temp.:95.00, PWM:26
17:02:00.249 -> Water_Temp.:95.00, PWM:26
17:02:01.044 -> Water_Temp.:95.00, PWM:26
17:02:01.804 -> Water_Temp.:95.00, PWM:26
17:02:02.601 -> Water_Temp.:95.00, PWM:26
17:02:03.393 -> Water_Temp.:95.00, PWM:27
17:02:04.153 -> Water_Temp.:95.11, PWM:26
17:02:04.950 -> Water_Temp.:95.00, PWM:26
17:02:05.711 -> Water_Temp.:95.00, PWM:26
17:02:06.505 -> Water_Temp.:95.00, PWM:26
17:02:07.267 -> Water_Temp.:95.00, PWM:26
17:02:08.059 -> Water_Temp.:96.69, PWM:27
17:02:08.856 -> Water_Temp.:98.71, PWM:27
17:02:09.617 -> Water_Temp.:96.35, PWM:26
17:02:10.412 -> Water_Temp.:95.90, PWM:26
17:02:11.167 -> Water_Temp.:95.68, PWM:26
17:02:11.960 -> Water_Temp.:95.56, PWM:26
17:02:12.755 -> Water_Temp.:95.45, PWM:27
17:02:13.516 -> Water_Temp.:95.45, PWM:26
17:02:14.312 -> Water_Temp.:95.34, PWM:26
17:02:15.075 -> Water_Temp.:95.34, PWM:27
17:02:15.873 -> Water_Temp.:95.34, PWM:30
17:02:16.630 -> Water_Temp.:95.34, PWM:30
17:02:17.421 -> Water_Temp.:95.22, PWM:30
17:02:18.215 -> Water_Temp.:95.22, PWM:30
17:02:18.975 -> Water_Temp.:95.22, PWM:30
17:02:19.762 -> Water_Temp.:95.22, PWM:30
17:02:20.520 -> Water_Temp.:95.22, PWM:30
17:02:21.306 -> Water_Temp.:95.22, PWM:30
17:02:22.094 -> Water_Temp.:95.22, PWM:30
17:02:22.887 -> Water_Temp.:95.22, PWM:30
17:02:23.647 -> Water_Temp.:95.11, PWM:30
17:02:24.441 -> Water_Temp.:95.11, PWM:30
17:02:25.231 -> Water_Temp.:95.11, PWM:30
17:02:25.990 -> Water_Temp.:95.11, PWM:30
17:02:26.784 -> Water_Temp.:95.11, PWM:30
17:02:27.544 -> Water_Temp.:95.11, PWM:30
17:02:28.338 -> Water_Temp.:95.11, PWM:30
17:02:29.131 -> Water_Temp.:95.11, PWM:30
17:02:29.887 -> Water_Temp.:95.11, PWM:30
17:02:30.682 -> Water_Temp.:95.11, PWM:30
17:02:31.472 -> Water_Temp.:95.11, PWM:30
17:02:32.227 -> Water_Temp.:95.11, PWM:30
17:02:33.020 -> Water_Temp.:95.11, PWM:30
17:02:33.810 -> Water_Temp.:95.11, PWM:30
17:02:34.569 -> Water_Temp.:95.11, PWM:30
17:02:35.363 -> Water_Temp.:95.11, PWM:30
17:02:36.123 -> Water_Temp.:95.11, PWM:30
17:02:36.919 -> Water_Temp.:95.11, PWM:30
17:02:37.709 -> Water_Temp.:95.11, PWM:30
17:02:38.468 -> Water_Temp.:95.11, PWM:30
17:02:39.266 -> Water_Temp.:95.11, PWM:30
17:02:40.028 -> Water_Temp.:95.11, PWM:30
17:02:40.823 -> Water_Temp.:95.11, PWM:30
17:02:41.584 -> Water_Temp.:95.11, PWM:30
17:02:42.381 -> Water_Temp.:95.11, PWM:30
17:02:43.174 -> Water_Temp.:95.11, PWM:30
17:02:43.930 -> Water_Temp.:95.11, PWM:30
17:02:44.723 -> Water_Temp.:95.11, PWM:30
17:02:45.487 -> Water_Temp.:95.22, PWM:30
17:02:46.283 -> Water_Temp.:95.11, PWM:30
17:02:47.042 -> Water_Temp.:95.11, PWM:30
17:02:47.829 -> Water_Temp.:95.11, PWM:30
17:02:49.384 -> Water_Temp.:95.22, PWM:30
17:02:50.175 -> Water_Temp.:95.22, PWM:30
17:02:50.971 -> Water_Temp.:95.22, PWM:30
17:02:51.727 -> Water_Temp.:95.22, PWM:30
17:02:52.519 -> Water_Temp.:95.22, PWM:30
17:02:53.278 -> Water_Temp.:95.22, PWM:30
17:02:54.072 -> Water_Temp.:95.22, PWM:30
17:02:54.869 -> Water_Temp.:95.22, PWM:30
17:02:55.630 -> Water_Temp.:95.22, PWM:30
17:02:56.424 -> Water_Temp.:95.22, PWM:30
17:02:57.936 -> Water_Temp.:95.22, PWM:30
17:02:58.727 -> Water_Temp.:95.22, PWM:30
17:02:59.520 -> Water_Temp.:95.22, PWM:30
17:03:00.279 -> Water_Temp.:95.22, PWM:30
17:03:01.073 -> Water_Temp.:95.22, PWM:30
17:03:01.867 -> Water_Temp.:95.22, PWM:30
17:03:02.623 -> Water_Temp.:95.22, PWM:30
17:03:03.418 -> Water_Temp.:95.22, PWM:30
17:03:04.211 -> Water_Temp.:95.22, PWM:30
17:03:04.972 -> Water_Temp.:95.22, PWM:30
17:03:05.766 -> Water_Temp.:95.22, PWM:30
17:03:06.528 -> Water_Temp.:95.22, PWM:30
17:03:07.318 -> Water_Temp.:95.22, PWM:30
17:03:08.081 -> Water_Temp.:95.22, PWM:30
17:03:08.875 -> Water_Temp.:95.22, PWM:30
17:03:09.638 -> Water_Temp.:95.22, PWM:30
17:03:10.429 -> Water_Temp.:95.22, PWM:30
17:03:11.223 -> Water_Temp.:95.22, PWM:30
17:03:12.776 -> Water_Temp.:95.22, PWM:30
17:03:13.536 -> Water_Temp.:95.22, PWM:30
17:03:14.332 -> Water_Temp.:95.22, PWM:30
17:03:15.095 -> Water_Temp.:95.22, PWM:30
17:03:15.889 -> Water_Temp.:95.22, PWM:30
17:03:16.685 -> Water_Temp.:95.22, PWM:30
17:03:17.447 -> Water_Temp.:95.22, PWM:30
17:03:18.245 -> Water_Temp.:95.22, PWM:30
17:03:18.999 -> Water_Temp.:95.22, PWM:30
17:03:19.794 -> Water_Temp.:95.34, PWM:30
17:03:20.557 -> Water_Temp.:95.22, PWM:30
17:03:21.354 -> Water_Temp.:95.34, PWM:30
17:03:22.143 -> Water_Temp.:95.34, PWM:30
17:03:22.899 -> Water_Temp.:95.34, PWM:30
17:03:23.695 -> Water_Temp.:95.34, PWM:30
17:03:24.456 -> Water_Temp.:95.34, PWM:30
17:03:25.252 -> Water_Temp.:95.34, PWM:30
17:03:26.015 -> Water_Temp.:95.34, PWM:30
17:03:26.808 -> Water_Temp.:95.34, PWM:30
17:03:27.599 -> Water_Temp.:95.45, PWM:30
17:03:28.359 -> Water_Temp.:95.45, PWM:30
17:03:29.149 -> Water_Temp.:95.45, PWM:30
17:03:29.944 -> Water_Temp.:95.45, PWM:30
17:03:30.703 -> Water_Temp.:95.45, PWM:30
17:03:31.495 -> Water_Temp.:95.45, PWM:30
17:03:32.287 -> Water_Temp.:95.45, PWM:30
17:03:33.046 -> Water_Temp.:95.45, PWM:30
17:03:33.838 -> Water_Temp.:95.45, PWM:30
17:03:34.601 -> Water_Temp.:95.45, PWM:30
17:03:35.397 -> Water_Temp.:95.45, PWM:30
17:03:36.157 -> Water_Temp.:95.45, PWM:30
17:03:36.950 -> Water_Temp.:95.45, PWM:30
17:03:37.741 -> Water_Temp.:95.45, PWM:30
17:03:38.496 -> Water_Temp.:95.45, PWM:30
17:03:39.291 -> Water_Temp.:95.45, PWM:30
17:03:40.844 -> Water_Temp.:95.34, PWM:30
17:03:41.601 -> Water_Temp.:95.45, PWM:30
17:03:42.392 -> Water_Temp.:95.45, PWM:30
17:03:43.183 -> Water_Temp.:95.34, PWM:30
17:03:43.945 -> Water_Temp.:95.34, PWM:30
17:03:44.742 -> Water_Temp.:95.34, PWM:30
17:03:45.530 -> Water_Temp.:95.45, PWM:30
17:03:46.291 -> Water_Temp.:95.45, PWM:30
17:03:47.085 -> Water_Temp.:95.34, PWM:30
17:03:48.600 -> Water_Temp.:95.34, PWM:30
17:03:50.156 -> Water_Temp.:95.34, PWM:30
17:03:50.914 -> Water_Temp.:95.34, PWM:30
17:03:51.707 -> Water_Temp.:95.34, PWM:30
17:03:52.499 -> Water_Temp.:95.34, PWM:30
17:03:53.261 -> Water_Temp.:95.34, PWM:30
17:03:54.053 -> Water_Temp.:95.45, PWM:30
17:03:54.813 -> Water_Temp.:95.45, PWM:30
17:03:55.611 -> Water_Temp.:95.34, PWM:30
17:03:56.405 -> Water_Temp.:95.34, PWM:30
17:03:57.167 -> Water_Temp.:95.45, PWM:30
17:03:57.966 -> Water_Temp.:95.45, PWM:30
17:03:58.730 -> Water_Temp.:95.45, PWM:30
17:03:59.493 -> Water_Temp.:95.34, PWM:30
17:04:00.292 -> Water_Temp.:95.45, PWM:30
17:04:01.054 -> Water_Temp.:95.45, PWM:30
17:04:01.848 -> Water_Temp.:95.45, PWM:30
17:04:02.641 -> Water_Temp.:95.45, PWM:30
17:04:04.201 -> Water_Temp.:95.34, PWM:30
17:04:04.963 -> Water_Temp.:95.34, PWM:30
17:04:05.759 -> Water_Temp.:95.34, PWM:30
17:04:06.518 -> Water_Temp.:95.34, PWM:30
17:04:07.314 -> Water_Temp.:95.34, PWM:30
17:04:08.076 -> Water_Temp.:95.34, PWM:30
17:04:08.865 -> Water_Temp.:95.34, PWM:30
17:04:09.657 -> Water_Temp.:95.34, PWM:30
17:04:10.416 -> Water_Temp.:95.34, PWM:30
17:04:11.210 -> Water_Temp.:95.34, PWM:30
17:04:11.972 -> Water_Temp.:95.34, PWM:30
17:04:12.763 -> Water_Temp.:95.34, PWM:30
17:04:13.560 -> Water_Temp.:95.34, PWM:30
17:04:14.320 -> Water_Temp.:95.34, PWM:30
17:04:15.111 -> Water_Temp.:95.34, PWM:30
17:04:16.659 -> Water_Temp.:95.45, PWM:30
17:04:17.453 -> Water_Temp.:95.45, PWM:30
17:04:18.217 -> Water_Temp.:95.45, PWM:30
17:04:19.773 -> Water_Temp.:95.45, PWM:30
17:04:20.569 -> Water_Temp.:95.45, PWM:30
17:04:21.331 -> Water_Temp.:95.45, PWM:30
17:04:22.129 -> Water_Temp.:95.45, PWM:30
17:04:22.887 -> Water_Temp.:95.45, PWM:30
17:04:23.686 -> Water_Temp.:95.45, PWM:30
17:04:24.447 -> Water_Temp.:95.45, PWM:30
17:04:25.240 -> Water_Temp.:95.45, PWM:30
17:04:26.035 -> Water_Temp.:95.45, PWM:30
17:04:26.798 -> Water_Temp.:95.45, PWM:30
17:04:27.597 -> Water_Temp.:95.45, PWM:30
17:04:28.356 -> Water_Temp.:95.45, PWM:30
17:04:29.155 -> Water_Temp.:95.45, PWM:30
17:04:29.915 -> Water_Temp.:95.45, PWM:30
17:04:30.709 -> Water_Temp.:95.45, PWM:30
17:04:31.472 -> Water_Temp.:95.45, PWM:30
17:04:32.264 -> Water_Temp.:95.45, PWM:30
17:04:33.060 -> Water_Temp.:95.45, PWM:30
17:04:33.817 -> Water_Temp.:95.45, PWM:30
17:04:34.608 -> Water_Temp.:95.45, PWM:30
17:04:35.371 -> Water_Temp.:95.45, PWM:30
17:04:36.160 -> Water_Temp.:95.45, PWM:30
17:04:36.949 -> Water_Temp.:95.45, PWM:30
17:04:37.708 -> Water_Temp.:97.02, PWM:30
17:04:39.264 -> Water_Temp.:96.80, PWM:30
17:04:40.021 -> Water_Temp.:96.35, PWM:30
17:04:40.810 -> Water_Temp.:96.12, PWM:30
17:04:41.602 -> Water_Temp.:96.01, PWM:30
17:04:42.363 -> Water_Temp.:95.90, PWM:30
17:04:43.923 -> Water_Temp.:95.79, PWM:30
17:04:44.683 -> Water_Temp.:95.79, PWM:30
17:04:45.481 -> Water_Temp.:95.79, PWM:30
17:04:46.272 -> Water_Temp.:95.68, PWM:30
17:04:47.033 -> Water_Temp.:95.68, PWM:30
17:04:48.590 -> Water_Temp.:95.68, PWM:30
17:04:49.384 -> Water_Temp.:95.68, PWM:30
17:04:50.944 -> Water_Temp.:95.68, PWM:30
17:04:52.502 -> Water_Temp.:95.68, PWM:30
17:04:53.261 -> Water_Temp.:95.56, PWM:30
17:04:54.055 -> Water_Temp.:95.56, PWM:30
17:04:54.849 -> Water_Temp.:95.56, PWM:30
17:04:55.608 -> Water_Temp.:95.56, PWM:30
17:04:56.403 -> Water_Temp.:95.56, PWM:30
17:04:57.163 -> Water_Temp.:95.56, PWM:30
17:04:57.957 -> Water_Temp.:95.56, PWM:30
17:04:58.716 -> Water_Temp.:95.56, PWM:30
17:04:59.510 -> Water_Temp.:95.56, PWM:30
17:05:00.303 -> Water_Temp.:95.56, PWM:30
17:05:01.066 -> Water_Temp.:95.56, PWM:30
17:05:01.858 -> Water_Temp.:95.56, PWM:30
17:05:02.616 -> Water_Temp.:95.56, PWM:30
17:05:03.409 -> Water_Temp.:95.56, PWM:30
17:05:04.198 -> Water_Temp.:95.56, PWM:30
17:05:04.962 -> Water_Temp.:95.56, PWM:30
17:05:06.546 -> Water_Temp.:95.56, PWM:30
17:05:07.304 -> Water_Temp.:95.56, PWM:30
17:05:08.098 -> Water_Temp.:95.56, PWM:30
17:05:08.859 -> Water_Temp.:95.56, PWM:30
17:05:09.657 -> Water_Temp.:95.56, PWM:30
17:05:10.415 -> Water_Temp.:95.56, PWM:30
17:05:11.210 -> Water_Temp.:95.56, PWM:30
17:05:12.007 -> Water_Temp.:95.56, PWM:30
17:05:12.768 -> Water_Temp.:95.56, PWM:30
17:05:13.564 -> Water_Temp.:95.56, PWM:30
17:05:14.323 -> Water_Temp.:95.56, PWM:30
17:05:15.116 -> Water_Temp.:95.56, PWM:30
17:05:15.878 -> Water_Temp.:95.56, PWM:30
17:05:16.672 -> Water_Temp.:95.56, PWM:30
17:05:18.224 -> Water_Temp.:95.56, PWM:30
17:05:19.021 -> Water_Temp.:95.56, PWM:30
17:05:19.782 -> Water_Temp.:95.56, PWM:30
17:05:20.580 -> Water_Temp.:95.56, PWM:30
17:05:22.103 -> Water_Temp.:95.56, PWM:30
17:05:22.899 -> Water_Temp.:95.56, PWM:30
17:05:23.656 -> Water_Temp.:95.56, PWM:30
17:05:24.451 -> Water_Temp.:95.56, PWM:30
17:05:25.210 -> Water_Temp.:95.56, PWM:30
17:05:26.005 -> Water_Temp.:95.56, PWM:30
17:05:27.525 -> Water_Temp.:95.56, PWM:30
17:05:28.321 -> Water_Temp.:95.56, PWM:30
17:05:29.116 -> Water_Temp.:95.56, PWM:30
17:05:29.876 -> Water_Temp.:95.56, PWM:30
17:05:30.668 -> Water_Temp.:95.56, PWM:30
17:05:31.425 -> Water_Temp.:95.56, PWM:30
17:05:32.219 -> Water_Temp.:95.56, PWM:30
17:05:33.767 -> Water_Temp.:95.56, PWM:30
17:05:34.562 -> Water_Temp.:95.56, PWM:30
17:05:35.327 -> Water_Temp.:95.56, PWM:30
17:05:36.118 -> Water_Temp.:95.56, PWM:30
17:05:36.882 -> Water_Temp.:95.56, PWM:30
17:05:37.673 -> Water_Temp.:95.56, PWM:30
17:05:38.466 -> Water_Temp.:95.56, PWM:30
17:05:39.224 -> Water_Temp.:95.68, PWM:30
17:05:40.017 -> Water_Temp.:95.68, PWM:30
17:05:41.563 -> Water_Temp.:95.68, PWM:30
17:05:42.358 -> Water_Temp.:95.56, PWM:30
17:05:43.118 -> Water_Temp.:95.68, PWM:30
17:05:43.912 -> Water_Temp.:95.68, PWM:30
17:05:44.705 -> Water_Temp.:95.56, PWM:30
17:05:45.463 -> Water_Temp.:95.68, PWM:30
17:05:46.259 -> Water_Temp.:95.68, PWM:30
17:05:47.023 -> Water_Temp.:95.68, PWM:30
17:05:47.813 -> Water_Temp.:95.68, PWM:30
17:05:48.578 -> Water_Temp.:95.68, PWM:30
17:05:49.377 -> Water_Temp.:95.68, PWM:30
17:05:50.172 -> Water_Temp.:95.68, PWM:30
17:05:50.933 -> Water_Temp.:95.68, PWM:30
17:05:51.697 -> Water_Temp.:95.68, PWM:30
17:05:52.492 -> Water_Temp.:95.68, PWM:30
17:05:53.286 -> Water_Temp.:95.68, PWM:30
17:05:54.047 -> Water_Temp.:95.68, PWM:30
17:05:54.838 -> Water_Temp.:95.68, PWM:30
17:05:55.630 -> Water_Temp.:95.68, PWM:30
17:05:56.390 -> Water_Temp.:95.68, PWM:30
17:05:57.945 -> Water_Temp.:95.68, PWM:30
17:05:58.736 -> Water_Temp.:95.68, PWM:30
17:06:00.281 -> Water_Temp.:95.68, PWM:30
17:06:01.074 -> Water_Temp.:95.68, PWM:30
17:06:01.866 -> Water_Temp.:95.68, PWM:30
17:06:03.415 -> Water_Temp.:95.68, PWM:30
17:06:04.207 -> Water_Temp.:95.68, PWM:30
17:06:04.969 -> Water_Temp.:95.68, PWM:30
17:06:05.762 -> Water_Temp.:95.68, PWM:30
17:06:07.287 -> Water_Temp.:95.68, PWM:30
17:06:08.079 -> Water_Temp.:95.68, PWM:30
17:06:08.836 -> Water_Temp.:95.68, PWM:30
17:06:09.633 -> Water_Temp.:95.68, PWM:30
17:06:10.396 -> Water_Temp.:95.68, PWM:30
17:06:11.192 -> Water_Temp.:95.68, PWM:30
17:06:11.957 -> Water_Temp.:95.68, PWM:30
17:06:12.756 -> Water_Temp.:95.68, PWM:30
17:06:14.270 -> Water_Temp.:95.68, PWM:30
17:06:15.069 -> Water_Temp.:95.56, PWM:30
17:06:15.831 -> Water_Temp.:95.68, PWM:30
17:06:16.624 -> Water_Temp.:95.68, PWM:30
17:06:17.385 -> Water_Temp.:95.68, PWM:30
17:06:18.180 -> Water_Temp.:95.68, PWM:30
17:06:18.975 -> Water_Temp.:95.68, PWM:30
17:06:20.494 -> Water_Temp.:95.68, PWM:30
17:06:21.289 -> Water_Temp.:95.68, PWM:30
17:06:22.049 -> Water_Temp.:95.68, PWM:30
17:06:22.841 -> Water_Temp.:95.68, PWM:30
17:06:23.602 -> Water_Temp.:95.68, PWM:30
17:06:24.395 -> Water_Temp.:95.68, PWM:30
17:06:25.192 -> Water_Temp.:95.68, PWM:30
17:06:25.949 -> Water_Temp.:95.68, PWM:30
17:06:26.741 -> Water_Temp.:95.68, PWM:30
17:06:27.534 -> Water_Temp.:95.68, PWM:30
17:06:28.291 -> Water_Temp.:95.68, PWM:30
17:06:29.082 -> Water_Temp.:95.68, PWM:30
17:06:29.872 -> Water_Temp.:95.68, PWM:30
17:06:30.630 -> Water_Temp.:95.68, PWM:30
17:06:31.425 -> Water_Temp.:95.68, PWM:30
17:06:32.942 -> Water_Temp.:95.68, PWM:30
17:06:33.735 -> Water_Temp.:95.68, PWM:30
17:06:34.527 -> Water_Temp.:95.68, PWM:30
17:06:35.284 -> Water_Temp.:95.68, PWM:30
17:06:36.079 -> Water_Temp.:95.68, PWM:30
17:06:36.840 -> Water_Temp.:95.68, PWM:30
17:06:38.391 -> Water_Temp.:95.68, PWM:30
17:06:39.946 -> Water_Temp.:95.68, PWM:30
17:06:40.740 -> Water_Temp.:95.68, PWM:30
17:06:41.499 -> Water_Temp.:95.68, PWM:30
17:06:42.294 -> Water_Temp.:95.68, PWM:30
17:06:43.056 -> Water_Temp.:95.68, PWM:30
17:06:43.851 -> Water_Temp.:95.68, PWM:30
17:06:44.643 -> Water_Temp.:95.68, PWM:30
17:06:45.404 -> Water_Temp.:95.68, PWM:30
17:06:46.198 -> Water_Temp.:95.68, PWM:30
17:06:46.955 -> Water_Temp.:95.68, PWM:30
17:06:47.751 -> Water_Temp.:95.68, PWM:30
17:06:48.540 -> Water_Temp.:95.68, PWM:30
17:06:49.298 -> Water_Temp.:95.68, PWM:30
17:06:50.094 -> Water_Temp.:95.68, PWM:30
17:06:50.885 -> Water_Temp.:95.68, PWM:30
17:06:51.643 -> Water_Temp.:95.68, PWM:30
17:06:52.435 -> Water_Temp.:95.68, PWM:30
17:06:53.194 -> Water_Temp.:95.68, PWM:30
17:06:53.988 -> Water_Temp.:95.68, PWM:30
17:06:54.783 -> Water_Temp.:95.68, PWM:30
17:06:55.543 -> Water_Temp.:95.68, PWM:30
17:06:56.334 -> Water_Temp.:95.68, PWM:30
17:06:57.857 -> Water_Temp.:95.68, PWM:30
17:06:58.650 -> Water_Temp.:95.68, PWM:30
17:06:59.414 -> Water_Temp.:95.68, PWM:30
17:07:00.206 -> Water_Temp.:95.68, PWM:30
17:07:01.004 -> Water_Temp.:95.68, PWM:30
17:07:01.762 -> Water_Temp.:95.68, PWM:30
17:07:02.554 -> Water_Temp.:95.68, PWM:30
17:07:03.347 -> Water_Temp.:95.68, PWM:30
17:07:04.106 -> Water_Temp.:95.68, PWM:30
17:07:04.899 -> Water_Temp.:95.68, PWM:30
17:07:05.690 -> Water_Temp.:95.68, PWM:30
17:07:07.204 -> Water_Temp.:95.68, PWM:30
17:07:08.001 -> Water_Temp.:95.79, PWM:30
17:07:08.762 -> Water_Temp.:95.79, PWM:30
17:07:09.557 -> Water_Temp.:95.79, PWM:30
17:07:11.111 -> Water_Temp.:95.79, PWM:30
17:07:11.872 -> Water_Temp.:95.79, PWM:30
17:07:12.667 -> Water_Temp.:95.79, PWM:30
17:07:13.426 -> Water_Temp.:95.79, PWM:30
17:07:14.219 -> Water_Temp.:95.79, PWM:30
17:07:15.019 -> Water_Temp.:95.79, PWM:30
17:07:15.778 -> Water_Temp.:95.79, PWM:30
17:07:16.570 -> Water_Temp.:95.79, PWM:30
17:07:17.325 -> Water_Temp.:95.79, PWM:30
17:07:18.117 -> Water_Temp.:95.79, PWM:30
17:07:18.912 -> Water_Temp.:95.79, PWM:30
17:07:19.673 -> Water_Temp.:95.79, PWM:30
17:07:21.256 -> Water_Temp.:95.79, PWM:30
17:07:22.013 -> Water_Temp.:95.79, PWM:30
17:07:22.811 -> Water_Temp.:95.79, PWM:30
17:07:23.571 -> Water_Temp.:95.79, PWM:30
17:07:24.365 -> Water_Temp.:95.90, PWM:30
17:07:25.127 -> Water_Temp.:95.79, PWM:30
17:07:25.922 -> Water_Temp.:95.90, PWM:30
17:07:26.718 -> Water_Temp.:95.90, PWM:30
17:07:27.481 -> Water_Temp.:95.90, PWM:30
17:07:28.274 -> Water_Temp.:95.90, PWM:30
17:07:29.032 -> Water_Temp.:95.90, PWM:30
17:07:29.825 -> Water_Temp.:95.79, PWM:30
17:07:30.615 -> Water_Temp.:95.79, PWM:30
17:07:31.377 -> Water_Temp.:95.79, PWM:30
17:07:32.171 -> Water_Temp.:95.79, PWM:30
17:07:33.695 -> Water_Temp.:95.79, PWM:30
17:07:34.487 -> Water_Temp.:95.79, PWM:30
17:07:35.249 -> Water_Temp.:95.79, PWM:30
17:07:36.044 -> Water_Temp.:95.79, PWM:30
17:07:36.806 -> Water_Temp.:95.79, PWM:30
17:07:37.594 -> Water_Temp.:95.79, PWM:30
17:07:38.388 -> Water_Temp.:95.90, PWM:30
17:07:39.907 -> Water_Temp.:95.90, PWM:30
17:07:40.698 -> Water_Temp.:95.90, PWM:30
17:07:42.252 -> Water_Temp.:95.90, PWM:30
17:07:43.011 -> Water_Temp.:95.90, PWM:30
17:07:43.808 -> Water_Temp.:95.90, PWM:30
17:07:44.601 -> Water_Temp.:95.90, PWM:30
17:07:45.358 -> Water_Temp.:95.90, PWM:30
17:07:46.151 -> Water_Temp.:95.90, PWM:30
17:07:46.914 -> Water_Temp.:95.90, PWM:30
17:07:47.708 -> Water_Temp.:95.90, PWM:30
17:07:48.500 -> Water_Temp.:95.90, PWM:30
17:07:49.264 -> Water_Temp.:95.90, PWM:30
17:07:50.055 -> Water_Temp.:95.90, PWM:30
17:07:50.813 -> Water_Temp.:95.90, PWM:30
17:07:51.609 -> Water_Temp.:95.90, PWM:30
17:07:52.371 -> Water_Temp.:95.90, PWM:30
17:07:53.164 -> Water_Temp.:95.90, PWM:30
17:07:53.962 -> Water_Temp.:95.90, PWM:30
17:07:54.725 -> Water_Temp.:95.90, PWM:30
17:07:55.520 -> Water_Temp.:95.90, PWM:30
17:07:56.281 -> Water_Temp.:95.90, PWM:30
17:07:57.077 -> Water_Temp.:95.90, PWM:30
17:07:57.833 -> Water_Temp.:97.47, PWM:30
17:07:58.627 -> Water_Temp.:99.50, PWM:30
17:07:59.418 -> Water_Temp.:97.14, PWM:30
17:08:00.183 -> Water_Temp.:96.69, PWM:30
17:08:00.973 -> Water_Temp.:96.46, PWM:30
17:08:02.494 -> Water_Temp.:96.24, PWM:30
17:08:03.290 -> Water_Temp.:96.24, PWM:30
17:08:04.046 -> Water_Temp.:96.12, PWM:30
17:08:04.837 -> Water_Temp.:96.12, PWM:30
17:08:05.630 -> Water_Temp.:96.12, PWM:30
17:08:06.387 -> Water_Temp.:96.01, PWM:30
17:08:07.178 -> Water_Temp.:96.01, PWM:30
17:08:07.971 -> Water_Temp.:96.01, PWM:30
17:08:08.730 -> Water_Temp.:96.01, PWM:30
17:08:09.525 -> Water_Temp.:96.01, PWM:30
17:08:10.283 -> Water_Temp.:96.01, PWM:30
17:08:11.078 -> Water_Temp.:96.01, PWM:30
17:08:11.871 -> Water_Temp.:96.01, PWM:30
17:08:12.630 -> Water_Temp.:95.90, PWM:30
17:08:13.423 -> Water_Temp.:96.01, PWM:30
17:08:14.215 -> Water_Temp.:95.90, PWM:30
17:08:14.974 -> Water_Temp.:95.90, PWM:30
17:08:15.767 -> Water_Temp.:95.90, PWM:30
17:08:16.529 -> Water_Temp.:95.90, PWM:30
17:08:17.321 -> Water_Temp.:95.90, PWM:30
17:08:18.082 -> Water_Temp.:95.90, PWM:30
17:08:18.876 -> Water_Temp.:96.01, PWM:30
17:08:19.667 -> Water_Temp.:95.90, PWM:30
17:08:20.431 -> Water_Temp.:95.90, PWM:30
17:08:21.224 -> Water_Temp.:96.01, PWM:30
17:08:21.983 -> Water_Temp.:95.90, PWM:30
17:08:22.776 -> Water_Temp.:96.01, PWM:30
17:08:23.573 -> Water_Temp.:96.01, PWM:30
17:08:24.330 -> Water_Temp.:96.01, PWM:30
17:08:25.914 -> Water_Temp.:96.01, PWM:30
17:08:26.674 -> Water_Temp.:96.01, PWM:30
17:08:27.468 -> Water_Temp.:96.01, PWM:30
17:08:28.227 -> Water_Temp.:96.01, PWM:30
17:08:29.019 -> Water_Temp.:96.01, PWM:30
17:08:29.781 -> Water_Temp.:96.01, PWM:30
17:08:30.572 -> Water_Temp.:96.01, PWM:30
17:08:31.367 -> Water_Temp.:96.01, PWM:30
17:08:32.127 -> Water_Temp.:96.01, PWM:30
17:08:32.917 -> Water_Temp.:97.59, PWM:30
17:08:33.710 -> Water_Temp.:99.61, PWM:30
17:08:34.468 -> Water_Temp.:97.25, PWM:30
17:08:35.264 -> Water_Temp.:96.80, PWM:30
17:08:36.023 -> Water_Temp.:96.69, PWM:30
17:08:36.817 -> Water_Temp.:96.57, PWM:30
17:08:37.612 -> Water_Temp.:96.46, PWM:30
17:08:38.373 -> Water_Temp.:96.46, PWM:30
17:08:39.164 -> Water_Temp.:96.35, PWM:30
17:08:39.924 -> Water_Temp.:96.24, PWM:30
17:08:40.716 -> Water_Temp.:96.24, PWM:30
17:08:41.510 -> Water_Temp.:96.24, PWM:30
17:08:42.266 -> Water_Temp.:96.12, PWM:30
17:08:43.057 -> Water_Temp.:96.12, PWM:30
17:08:43.851 -> Water_Temp.:96.12, PWM:30
17:08:45.370 -> Water_Temp.:96.12, PWM:30
17:08:46.164 -> Water_Temp.:96.12, PWM:30
17:08:46.959 -> Water_Temp.:96.12, PWM:30
17:08:47.717 -> Water_Temp.:96.01, PWM:30
17:08:48.512 -> Water_Temp.:96.01, PWM:30
17:08:49.275 -> Water_Temp.:96.01, PWM:30
17:08:50.062 -> Water_Temp.:96.01, PWM:30
17:08:50.855 -> Water_Temp.:96.01, PWM:30
17:08:51.619 -> Water_Temp.:96.01, PWM:30
17:08:52.408 -> Water_Temp.:96.01, PWM:30
17:08:53.199 -> Water_Temp.:96.01, PWM:30
17:08:53.959 -> Water_Temp.:96.01, PWM:30
17:08:54.752 -> Water_Temp.:96.01, PWM:30
17:08:55.507 -> Water_Temp.:96.01, PWM:30
17:08:56.302 -> Water_Temp.:96.01, PWM:30
17:08:57.097 -> Water_Temp.:96.01, PWM:30
17:08:57.856 -> Water_Temp.:96.01, PWM:30
17:08:58.648 -> Water_Temp.:96.01, PWM:30
17:08:59.414 -> Water_Temp.:96.01, PWM:30
17:09:00.207 -> Water_Temp.:96.01, PWM:30
17:09:00.995 -> Water_Temp.:96.01, PWM:30
17:09:01.750 -> Water_Temp.:96.01, PWM:30
17:09:02.539 -> Water_Temp.:96.01, PWM:30
17:09:03.332 -> Water_Temp.:96.01, PWM:30
17:09:04.093 -> Water_Temp.:96.01, PWM:30
17:09:04.886 -> Water_Temp.:96.01, PWM:30
17:09:05.649 -> Water_Temp.:96.01, PWM:30
17:09:06.439 -> Water_Temp.:96.01, PWM:30
17:09:07.234 -> Water_Temp.:96.01, PWM:30
17:09:07.996 -> Water_Temp.:96.01, PWM:30
17:09:08.792 -> Water_Temp.:96.01, PWM:30
17:09:09.547 -> Water_Temp.:96.01, PWM:30
17:09:10.339 -> Water_Temp.:96.01, PWM:30
17:09:11.134 -> Water_Temp.:96.01, PWM:30
17:09:11.892 -> Water_Temp.:96.01, PWM:30
17:09:12.684 -> Water_Temp.:96.01, PWM:30
17:09:13.446 -> Water_Temp.:96.01, PWM:30
17:09:14.239 -> Water_Temp.:96.01, PWM:30
17:09:15.032 -> Water_Temp.:96.01, PWM:30
17:09:15.791 -> Water_Temp.:96.01, PWM:30
17:09:16.584 -> Water_Temp.:96.01, PWM:30
17:09:18.105 -> Water_Temp.:96.01, PWM:30
17:09:18.897 -> Water_Temp.:96.01, PWM:30
17:09:19.658 -> Water_Temp.:96.01, PWM:30
17:09:20.453 -> Water_Temp.:96.01, PWM:30
17:09:22.007 -> Water_Temp.:96.01, PWM:30
17:09:22.767 -> Water_Temp.:96.01, PWM:30
17:09:23.564 -> Water_Temp.:96.01, PWM:30
17:09:24.325 -> Water_Temp.:96.01, PWM:30
17:09:25.121 -> Water_Temp.:96.01, PWM:30
17:09:25.883 -> Water_Temp.:96.01, PWM:30
17:09:26.678 -> Water_Temp.:96.01, PWM:30
17:09:27.436 -> Water_Temp.:96.01, PWM:30
17:09:28.231 -> Water_Temp.:96.01, PWM:30
17:09:29.022 -> Water_Temp.:96.01, PWM:30
17:09:29.782 -> Water_Temp.:96.01, PWM:30
17:09:31.336 -> Water_Temp.:96.01, PWM:30
17:09:32.132 -> Water_Temp.:96.01, PWM:30
17:09:32.925 -> Water_Temp.:96.01, PWM:30
17:09:33.682 -> Water_Temp.:96.01, PWM:30
17:09:34.474 -> Water_Temp.:96.12, PWM:30
17:09:35.235 -> Water_Temp.:96.01, PWM:30
17:09:36.027 -> Water_Temp.:96.12, PWM:30
17:09:37.580 -> Water_Temp.:96.12, PWM:30
17:09:38.373 -> Water_Temp.:96.12, PWM:30
17:09:39.166 -> Water_Temp.:96.12, PWM:30
17:09:39.925 -> Water_Temp.:96.12, PWM:30
17:09:40.715 -> Water_Temp.:96.12, PWM:29
17:09:41.476 -> Water_Temp.:96.12, PWM:30
17:09:43.062 -> Water_Temp.:96.12, PWM:30
17:09:43.816 -> Water_Temp.:96.12, PWM:30
17:09:45.403 -> Water_Temp.:96.12, PWM:30
17:09:46.160 -> Water_Temp.:96.12, PWM:32
17:09:46.954 -> Water_Temp.:96.12, PWM:32
17:09:47.715 -> Water_Temp.:96.12, PWM:32
17:09:48.509 -> Water_Temp.:96.12, PWM:32
17:09:49.303 -> Water_Temp.:96.12, PWM:32
17:09:50.062 -> Water_Temp.:96.12, PWM:32
17:09:50.856 -> Water_Temp.:96.12, PWM:32
17:09:51.614 -> Water_Temp.:96.12, PWM:32
17:09:52.407 -> Water_Temp.:96.12, PWM:32
17:09:53.203 -> Water_Temp.:96.12, PWM:32
17:09:54.754 -> Water_Temp.:96.12, PWM:32
17:09:55.509 -> Water_Temp.:96.12, PWM:32
17:09:56.304 -> Water_Temp.:96.12, PWM:32
17:09:57.096 -> Water_Temp.:96.12, PWM:32
17:09:57.885 -> Water_Temp.:96.12, PWM:32
17:09:58.641 -> Water_Temp.:96.12, PWM:32
17:10:00.196 -> Water_Temp.:96.12, PWM:32
17:10:01.752 -> Water_Temp.:96.12, PWM:32
17:10:02.542 -> Water_Temp.:96.12, PWM:32
17:10:03.337 -> Water_Temp.:96.12, PWM:32
17:10:04.097 -> Water_Temp.:96.12, PWM:32
17:10:04.889 -> Water_Temp.:96.12, PWM:32
17:10:05.651 -> Water_Temp.:96.12, PWM:32
17:10:06.446 -> Water_Temp.:96.12, PWM:32
17:10:07.238 -> Water_Temp.:96.12, PWM:32
17:10:07.997 -> Water_Temp.:96.12, PWM:32
17:10:08.792 -> Water_Temp.:96.12, PWM:32
17:10:09.578 -> Water_Temp.:96.12, PWM:32
17:10:10.338 -> Water_Temp.:96.12, PWM:32
17:10:11.130 -> Water_Temp.:96.12, PWM:32
17:10:11.889 -> Water_Temp.:96.12, PWM:32
17:10:12.685 -> Water_Temp.:96.12, PWM:32
17:10:13.474 -> Water_Temp.:96.12, PWM:32
17:10:14.233 -> Water_Temp.:96.12, PWM:32
17:10:15.024 -> Water_Temp.:96.12, PWM:32
17:10:15.820 -> Water_Temp.:96.12, PWM:32
17:10:17.371 -> Water_Temp.:96.12, PWM:32
17:10:18.130 -> Water_Temp.:96.12, PWM:32
17:10:18.924 -> Water_Temp.:96.12, PWM:32
17:10:19.712 -> Water_Temp.:96.12, PWM:32
17:10:20.470 -> Water_Temp.:96.12, PWM:32
17:10:21.263 -> Water_Temp.:96.12, PWM:32
17:10:22.057 -> Water_Temp.:96.12, PWM:32
17:10:22.817 -> Water_Temp.:96.24, PWM:32
17:10:23.612 -> Water_Temp.:96.24, PWM:32
17:10:24.370 -> Water_Temp.:96.12, PWM:32
17:10:25.161 -> Water_Temp.:96.12, PWM:32
17:10:25.925 -> Water_Temp.:96.24, PWM:32
17:10:26.722 -> Water_Temp.:96.24, PWM:32
17:10:27.512 -> Water_Temp.:96.24, PWM:32
17:10:28.272 -> Water_Temp.:96.24, PWM:32
17:10:29.823 -> Water_Temp.:96.24, PWM:32
17:10:30.615 -> Water_Temp.:96.24, PWM:32
17:10:31.408 -> Water_Temp.:96.24, PWM:32
17:10:32.200 -> Water_Temp.:96.24, PWM:32
17:10:32.959 -> Water_Temp.:96.24, PWM:32
17:10:33.755 -> Water_Temp.:96.24, PWM:32
17:10:34.513 -> Water_Temp.:96.24, PWM:32
17:10:35.307 -> Water_Temp.:96.24, PWM:32
17:10:36.093 -> Water_Temp.:96.24, PWM:32
17:10:36.849 -> Water_Temp.:96.24, PWM:32
17:10:37.640 -> Water_Temp.:96.24, PWM:32
17:10:38.434 -> Water_Temp.:96.24, PWM:32
17:10:39.187 -> Water_Temp.:96.24, PWM:32
17:10:39.982 -> Water_Temp.:96.24, PWM:32
17:10:40.745 -> Water_Temp.:96.24, PWM:32
17:10:41.545 -> Water_Temp.:96.24, PWM:32
17:10:42.310 -> Water_Temp.:96.24, PWM:32
17:10:43.100 -> Water_Temp.:96.24, PWM:32
17:10:43.892 -> Water_Temp.:96.24, PWM:32
17:10:44.652 -> Water_Temp.:96.24, PWM:32
17:10:45.444 -> Water_Temp.:96.24, PWM:32
17:10:46.205 -> Water_Temp.:96.24, PWM:32
17:10:47.754 -> Water_Temp.:96.24, PWM:32
17:10:48.549 -> Water_Temp.:96.24, PWM:32
17:10:49.310 -> Water_Temp.:96.24, PWM:32
17:10:50.104 -> Water_Temp.:96.24, PWM:32
17:10:50.865 -> Water_Temp.:96.24, PWM:32
17:10:52.426 -> Water_Temp.:96.24, PWM:32
17:10:53.218 -> Water_Temp.:96.24, PWM:32
17:10:53.977 -> Water_Temp.:96.24, PWM:32
17:10:54.768 -> Water_Temp.:96.24, PWM:32
17:10:55.560 -> Water_Temp.:96.24, PWM:32
17:10:56.322 -> Water_Temp.:96.24, PWM:32
17:10:57.116 -> Water_Temp.:96.24, PWM:32
17:10:57.907 -> Water_Temp.:96.24, PWM:32
17:10:58.666 -> Water_Temp.:96.24, PWM:32
17:10:59.458 -> Water_Temp.:96.24, PWM:32
17:11:00.215 -> Water_Temp.:96.24, PWM:32
17:11:01.007 -> Water_Temp.:96.24, PWM:32
17:11:01.802 -> Water_Temp.:96.24, PWM:32
17:11:02.566 -> Water_Temp.:96.24, PWM:32
17:11:03.363 -> Water_Temp.:96.24, PWM:32
17:11:04.123 -> Water_Temp.:96.24, PWM:32
17:11:04.914 -> Water_Temp.:96.24, PWM:32
17:11:05.700 -> Water_Temp.:96.24, PWM:32
17:11:07.217 -> Water_Temp.:96.35, PWM:32
17:11:08.010 -> Water_Temp.:96.35, PWM:32
17:11:08.772 -> Water_Temp.:96.35, PWM:32
17:11:09.562 -> Water_Temp.:96.35, PWM:32
17:11:10.353 -> Water_Temp.:96.35, PWM:32
17:11:11.113 -> Water_Temp.:96.35, PWM:32
17:11:11.904 -> Water_Temp.:96.35, PWM:32
17:11:12.697 -> Water_Temp.:96.35, PWM:32
17:11:13.455 -> Water_Temp.:96.35, PWM:32
17:11:14.247 -> Water_Temp.:96.35, PWM:32
17:11:15.041 -> Water_Temp.:96.35, PWM:32
17:11:15.801 -> Water_Temp.:96.35, PWM:32
17:11:16.594 -> Water_Temp.:96.35, PWM:32
17:11:17.352 -> Water_Temp.:96.35, PWM:32
17:11:18.145 -> Water_Temp.:96.35, PWM:32
17:11:18.940 -> Water_Temp.:96.35, PWM:32
17:11:19.702 -> Water_Temp.:96.35, PWM:32
17:11:20.495 -> Water_Temp.:96.35, PWM:32
17:11:22.803 -> Water_Temp.:96.35, PWM:32
17:11:23.565 -> Water_Temp.:96.35, PWM:32
17:11:24.353 -> Water_Temp.:96.35, PWM:32
17:11:25.143 -> Water_Temp.:96.35, PWM:32
17:11:25.903 -> Water_Temp.:96.35, PWM:32
17:11:26.698 -> Water_Temp.:96.35, PWM:32
17:11:27.489 -> Water_Temp.:96.35, PWM:32
17:11:28.251 -> Water_Temp.:96.35, PWM:32
17:11:29.044 -> Water_Temp.:96.35, PWM:32
17:11:29.805 -> Water_Temp.:96.35, PWM:32
17:11:30.598 -> Water_Temp.:96.35, PWM:32
17:11:32.151 -> Water_Temp.:96.35, PWM:32
17:11:32.910 -> Water_Temp.:96.35, PWM:32
17:11:33.700 -> Water_Temp.:96.35, PWM:32
17:11:34.492 -> Water_Temp.:96.35, PWM:33
17:11:35.252 -> Water_Temp.:96.35, PWM:32
17:11:36.048 -> Water_Temp.:96.35, PWM:32
17:11:36.807 -> Water_Temp.:96.35, PWM:32
17:11:37.603 -> Water_Temp.:96.35, PWM:32
17:11:38.398 -> Water_Temp.:96.35, PWM:32
17:11:39.163 -> Water_Temp.:96.35, PWM:32
17:11:39.922 -> Water_Temp.:96.35, PWM:32
17:11:40.713 -> Water_Temp.:96.35, PWM:32
17:11:41.510 -> Water_Temp.:96.35, PWM:32
17:11:42.273 -> Water_Temp.:96.35, PWM:32
17:11:43.066 -> Water_Temp.:96.35, PWM:32
17:11:43.826 -> Water_Temp.:96.35, PWM:32
17:11:44.623 -> Water_Temp.:96.35, PWM:32
17:11:45.411 -> Water_Temp.:96.35, PWM:32
17:11:46.167 -> Water_Temp.:96.35, PWM:32
17:11:46.964 -> Water_Temp.:96.35, PWM:32
17:11:47.754 -> Water_Temp.:96.35, PWM:32
17:11:48.510 -> Water_Temp.:96.35, PWM:32
17:11:49.301 -> Water_Temp.:96.35, PWM:32
17:11:50.848 -> Water_Temp.:96.35, PWM:32
17:11:51.638 -> Water_Temp.:96.35, PWM:32
17:11:52.433 -> Water_Temp.:96.35, PWM:32
17:11:53.191 -> Water_Temp.:96.35, PWM:32
17:11:53.988 -> Water_Temp.:96.35, PWM:32
17:11:54.752 -> Water_Temp.:96.35, PWM:32
17:11:55.543 -> Water_Temp.:96.35, PWM:32
17:11:56.305 -> Water_Temp.:96.35, PWM:32
17:11:57.096 -> Water_Temp.:96.35, PWM:32
17:11:57.889 -> Water_Temp.:96.46, PWM:32
17:11:58.650 -> Water_Temp.:96.35, PWM:32
17:11:59.437 -> Water_Temp.:96.35, PWM:32
17:12:00.230 -> Water_Temp.:96.46, PWM:32
17:12:00.990 -> Water_Temp.:96.35, PWM:32
17:12:01.787 -> Water_Temp.:96.46, PWM:32
17:12:02.546 -> Water_Temp.:96.35, PWM:32
17:12:03.342 -> Water_Temp.:96.46, PWM:32
17:12:04.101 -> Water_Temp.:96.35, PWM:32
17:12:04.897 -> Water_Temp.:96.46, PWM:32
17:12:05.691 -> Water_Temp.:96.35, PWM:32
17:12:06.449 -> Water_Temp.:96.46, PWM:32
17:12:07.241 -> Water_Temp.:96.46, PWM:32
17:12:08.000 -> Water_Temp.:96.46, PWM:32
17:12:08.793 -> Water_Temp.:96.46, PWM:32
17:12:09.584 -> Water_Temp.:96.46, PWM:32
17:12:10.346 -> Water_Temp.:96.46, PWM:32
17:12:11.900 -> Water_Temp.:96.46, PWM:32
17:12:12.661 -> Water_Temp.:96.46, PWM:32
17:12:13.454 -> Water_Temp.:96.46, PWM:32
17:12:14.250 -> Water_Temp.:96.46, PWM:32
17:12:15.010 -> Water_Temp.:96.46, PWM:32
17:12:15.802 -> Water_Temp.:96.46, PWM:32
17:12:16.561 -> Water_Temp.:96.46, PWM:32
17:12:17.355 -> Water_Temp.:96.46, PWM:32
17:12:18.152 -> Water_Temp.:96.46, PWM:32
17:12:18.912 -> Water_Temp.:96.46, PWM:32
17:12:19.707 -> Water_Temp.:96.46, PWM:32
17:12:20.466 -> Water_Temp.:96.46, PWM:32
17:12:21.257 -> Water_Temp.:96.46, PWM:32
17:12:22.049 -> Water_Temp.:96.46, PWM:32
17:12:22.809 -> Water_Temp.:96.46, PWM:32
17:12:23.603 -> Water_Temp.:96.46, PWM:32
17:12:24.369 -> Water_Temp.:96.46, PWM:32
17:12:25.166 -> Water_Temp.:96.46, PWM:32
17:12:25.925 -> Water_Temp.:96.46, PWM:32
17:12:26.719 -> Water_Temp.:96.46, PWM:32
17:12:27.483 -> Water_Temp.:96.46, PWM:32
17:12:28.279 -> Water_Temp.:96.46, PWM:32
17:12:29.038 -> Water_Temp.:96.46, PWM:32
17:12:29.832 -> Water_Temp.:96.46, PWM:32
17:12:30.628 -> Water_Temp.:96.46, PWM:32
17:12:31.389 -> Water_Temp.:96.57, PWM:32
17:12:32.185 -> Water_Temp.:96.57, PWM:32
17:12:32.943 -> Water_Temp.:96.57, PWM:32
17:12:33.738 -> Water_Temp.:96.57, PWM:32
17:12:34.498 -> Water_Temp.:96.57, PWM:32
17:12:36.060 -> Water_Temp.:96.57, PWM:32
17:12:36.822 -> Water_Temp.:96.57, PWM:32
17:12:37.616 -> Water_Temp.:96.57, PWM:32
17:12:38.376 -> Water_Temp.:96.57, PWM:32
17:12:39.170 -> Water_Temp.:96.57, PWM:32
17:12:40.724 -> Water_Temp.:96.57, PWM:32
17:12:41.519 -> Water_Temp.:96.57, PWM:32
17:12:42.280 -> Water_Temp.:96.46, PWM:32
17:12:43.066 -> Water_Temp.:96.46, PWM:32
17:12:43.861 -> Water_Temp.:96.57, PWM:32
17:12:45.376 -> Water_Temp.:96.46, PWM:32
17:12:46.170 -> Water_Temp.:96.46, PWM:32
17:12:46.926 -> Water_Temp.:96.57, PWM:32
17:12:47.718 -> Water_Temp.:96.46, PWM:32
17:12:48.515 -> Water_Temp.:96.46, PWM:32
17:12:49.273 -> Water_Temp.:96.46, PWM:32
17:12:50.065 -> Water_Temp.:96.57, PWM:32
17:12:50.858 -> Water_Temp.:96.46, PWM:32
17:12:51.617 -> Water_Temp.:96.46, PWM:32
17:12:52.410 -> Water_Temp.:96.57, PWM:32
17:12:53.170 -> Water_Temp.:96.57, PWM:32
17:12:53.961 -> Water_Temp.:96.46, PWM:32
17:12:54.723 -> Water_Temp.:96.57, PWM:32
17:12:55.520 -> Water_Temp.:96.57, PWM:32
17:12:56.318 -> Water_Temp.:98.15, PWM:32
17:12:57.077 -> Water_Temp.:100.17, PWM:32
17:12:57.868 -> Water_Temp.:97.81, PWM:32
17:12:58.630 -> Water_Temp.:97.36, PWM:32
17:12:59.417 -> Water_Temp.:97.25, PWM:32
17:13:00.206 -> Water_Temp.:97.02, PWM:32
17:13:00.962 -> Water_Temp.:96.91, PWM:32
17:13:01.756 -> Water_Temp.:96.91, PWM:32
17:13:02.550 -> Water_Temp.:96.80, PWM:32
17:13:03.310 -> Water_Temp.:96.80, PWM:32
17:13:04.106 -> Water_Temp.:96.80, PWM:32
17:13:04.866 -> Water_Temp.:96.69, PWM:32
17:13:05.657 -> Water_Temp.:96.69, PWM:32
17:13:06.421 -> Water_Temp.:96.69, PWM:32
17:13:08.005 -> Water_Temp.:96.69, PWM:32
17:13:08.764 -> Water_Temp.:96.69, PWM:32
17:13:09.564 -> Water_Temp.:96.69, PWM:32
17:13:10.350 -> Water_Temp.:96.69, PWM:32
17:13:11.107 -> Water_Temp.:96.69, PWM:32
17:13:11.902 -> Water_Temp.:96.69, PWM:32
17:13:12.662 -> Water_Temp.:96.69, PWM:32
17:13:13.457 -> Water_Temp.:96.69, PWM:32
17:13:14.248 -> Water_Temp.:96.57, PWM:32
17:13:15.010 -> Water_Temp.:96.57, PWM:32
17:13:15.803 -> Water_Temp.:96.57, PWM:32
17:13:16.592 -> Water_Temp.:96.57, PWM:32
17:13:17.352 -> Water_Temp.:96.57, PWM:32
17:13:18.145 -> Water_Temp.:96.57, PWM:32
17:13:18.903 -> Water_Temp.:96.57, PWM:32
17:13:19.699 -> Water_Temp.:96.57, PWM:32
17:13:20.493 -> Water_Temp.:96.57, PWM:32
17:13:21.249 -> Water_Temp.:96.57, PWM:32
17:13:22.044 -> Water_Temp.:96.57, PWM:32
17:13:22.799 -> Water_Temp.:96.57, PWM:32
17:13:23.593 -> Water_Temp.:96.57, PWM:32
17:13:24.391 -> Water_Temp.:96.57, PWM:32
17:13:25.153 -> Water_Temp.:96.57, PWM:32
17:13:25.942 -> Water_Temp.:96.69, PWM:32
17:13:26.703 -> Water_Temp.:96.69, PWM:32
17:13:27.497 -> Water_Temp.:96.69, PWM:32
17:13:29.048 -> Water_Temp.:96.69, PWM:32
17:13:29.812 -> Water_Temp.:96.69, PWM:32
17:13:30.605 -> Water_Temp.:96.69, PWM:32
17:13:31.396 -> Water_Temp.:96.69, PWM:32
17:13:32.153 -> Water_Temp.:96.69, PWM:32
17:13:32.949 -> Water_Temp.:96.69, PWM:32
17:13:33.710 -> Water_Temp.:96.69, PWM:32
17:13:34.499 -> Water_Temp.:96.69, PWM:32
17:13:35.294 -> Water_Temp.:96.69, PWM:32
17:13:36.844 -> Water_Temp.:96.69, PWM:35
17:13:37.604 -> Water_Temp.:96.69, PWM:35
17:13:38.401 -> Water_Temp.:96.69, PWM:35
17:13:39.189 -> Water_Temp.:96.69, PWM:35
17:13:39.946 -> Water_Temp.:96.69, PWM:35
17:13:41.531 -> Water_Temp.:96.69, PWM:35
17:13:42.292 -> Water_Temp.:96.69, PWM:35
17:13:43.845 -> Water_Temp.:96.69, PWM:35
17:13:44.636 -> Water_Temp.:96.69, PWM:35
17:13:45.397 -> Water_Temp.:96.69, PWM:35
17:13:46.189 -> Water_Temp.:96.69, PWM:35
17:13:46.946 -> Water_Temp.:96.69, PWM:35
17:13:48.526 -> Water_Temp.:96.69, PWM:35
17:13:49.284 -> Water_Temp.:96.69, PWM:35
17:13:50.080 -> Water_Temp.:96.69, PWM:35
17:13:50.841 -> Water_Temp.:96.69, PWM:35
17:13:51.637 -> Water_Temp.:96.69, PWM:35
17:13:52.429 -> Water_Temp.:96.69, PWM:35
17:13:53.187 -> Water_Temp.:96.69, PWM:35
17:13:53.984 -> Water_Temp.:96.69, PWM:35
17:13:54.743 -> Water_Temp.:96.69, PWM:35
17:13:55.537 -> Water_Temp.:96.69, PWM:35
17:13:56.301 -> Water_Temp.:96.80, PWM:35
17:13:57.094 -> Water_Temp.:96.80, PWM:35
17:13:57.889 -> Water_Temp.:96.80, PWM:35
17:13:58.644 -> Water_Temp.:96.80, PWM:35
17:13:59.437 -> Water_Temp.:96.80, PWM:35
17:14:00.229 -> Water_Temp.:96.80, PWM:35
17:14:00.992 -> Water_Temp.:96.80, PWM:35
17:14:01.787 -> Water_Temp.:96.80, PWM:35
17:14:02.548 -> Water_Temp.:96.80, PWM:35
17:14:03.343 -> Water_Temp.:96.80, PWM:35
17:14:04.133 -> Water_Temp.:96.80, PWM:35
17:14:04.895 -> Water_Temp.:96.80, PWM:35
17:14:05.692 -> Water_Temp.:96.80, PWM:35
17:14:06.452 -> Water_Temp.:96.80, PWM:35
17:14:07.245 -> Water_Temp.:96.80, PWM:35
17:14:08.007 -> Water_Temp.:96.80, PWM:35
17:14:08.801 -> Water_Temp.:96.69, PWM:35
17:14:09.595 -> Water_Temp.:96.69, PWM:35
17:14:10.355 -> Water_Temp.:96.69, PWM:35
17:14:11.120 -> Water_Temp.:96.80, PWM:35
17:14:11.911 -> Water_Temp.:96.69, PWM:35
17:14:14.255 -> Water_Temp.:96.69, PWM:35
17:14:15.049 -> Water_Temp.:96.69, PWM:35
17:14:15.808 -> Water_Temp.:96.69, PWM:35
17:14:16.602 -> Water_Temp.:96.69, PWM:35
17:14:17.359 -> Water_Temp.:96.69, PWM:35
17:14:18.154 -> Water_Temp.:96.69, PWM:35
17:14:19.708 -> Water_Temp.:96.69, PWM:35
17:14:20.502 -> Water_Temp.:96.69, PWM:35
17:14:21.260 -> Water_Temp.:96.69, PWM:35
17:14:22.053 -> Water_Temp.:96.69, PWM:35
17:14:22.848 -> Water_Temp.:96.69, PWM:35
17:14:23.608 -> Water_Temp.:96.69, PWM:35
17:14:24.404 -> Water_Temp.:96.69, PWM:35
17:14:25.162 -> Water_Temp.:96.69, PWM:35
17:14:25.953 -> Water_Temp.:96.69, PWM:35
17:14:27.509 -> Water_Temp.:96.69, PWM:35
17:14:28.266 -> Water_Temp.:96.69, PWM:35
17:14:29.064 -> Water_Temp.:96.69, PWM:35
17:14:29.822 -> Water_Temp.:96.69, PWM:35
17:14:30.612 -> Water_Temp.:96.69, PWM:35
17:14:31.404 -> Water_Temp.:96.69, PWM:35
17:14:32.163 -> Water_Temp.:96.69, PWM:35
17:14:32.952 -> Water_Temp.:96.69, PWM:35
17:14:33.747 -> Water_Temp.:96.69, PWM:35
17:14:34.509 -> Water_Temp.:96.69, PWM:35
17:14:35.300 -> Water_Temp.:96.69, PWM:35
17:14:36.059 -> Water_Temp.:96.69, PWM:35
17:14:36.853 -> Water_Temp.:96.80, PWM:35
17:14:37.642 -> Water_Temp.:96.69, PWM:35
17:14:38.401 -> Water_Temp.:96.80, PWM:35
17:14:39.197 -> Water_Temp.:96.69, PWM:35
17:14:39.954 -> Water_Temp.:96.69, PWM:35
17:14:40.753 -> Water_Temp.:96.69, PWM:35
17:14:41.546 -> Water_Temp.:96.69, PWM:35
17:14:42.305 -> Water_Temp.:96.80, PWM:35
17:14:43.859 -> Water_Temp.:96.80, PWM:35
17:14:44.619 -> Water_Temp.:96.69, PWM:35
17:14:45.414 -> Water_Temp.:96.80, PWM:35
17:14:46.172 -> Water_Temp.:96.80, PWM:35
17:14:46.963 -> Water_Temp.:96.80, PWM:35
17:14:47.757 -> Water_Temp.:96.80, PWM:35
17:14:48.517 -> Water_Temp.:96.80, PWM:35
17:14:49.312 -> Water_Temp.:96.80, PWM:35
17:14:50.070 -> Water_Temp.:96.80, PWM:35
17:14:50.865 -> Water_Temp.:96.80, PWM:35
17:14:51.625 -> Water_Temp.:96.80, PWM:35
17:14:52.417 -> Water_Temp.:96.80, PWM:35
17:14:53.209 -> Water_Temp.:96.80, PWM:35
17:14:53.968 -> Water_Temp.:96.80, PWM:35
17:14:54.759 -> Water_Temp.:96.80, PWM:35
17:14:55.552 -> Water_Temp.:96.80, PWM:35
17:14:56.311 -> Water_Temp.:96.80, PWM:35
17:14:57.104 -> Water_Temp.:96.80, PWM:35
17:14:57.898 -> Water_Temp.:96.80, PWM:35
17:14:58.656 -> Water_Temp.:96.80, PWM:35
17:14:59.448 -> Water_Temp.:96.80, PWM:35
17:15:00.209 -> Water_Temp.:96.80, PWM:35
17:15:01.002 -> Water_Temp.:96.80, PWM:35
17:15:01.794 -> Water_Temp.:96.80, PWM:35
17:15:02.549 -> Water_Temp.:96.80, PWM:35
17:15:03.347 -> Water_Temp.:96.80, PWM:35
17:15:04.109 -> Water_Temp.:96.80, PWM:35
17:15:04.902 -> Water_Temp.:96.80, PWM:35
17:15:05.699 -> Water_Temp.:96.80, PWM:35
17:15:06.459 -> Water_Temp.:96.80, PWM:35
17:15:07.254 -> Water_Temp.:96.80, PWM:35
17:15:08.009 -> Water_Temp.:96.80, PWM:35
17:15:08.804 -> Water_Temp.:96.80, PWM:35
17:15:09.568 -> Water_Temp.:96.80, PWM:35
17:15:11.118 -> Water_Temp.:96.80, PWM:35
17:15:11.883 -> Water_Temp.:96.91, PWM:35
17:15:12.677 -> Water_Temp.:96.91, PWM:35
17:15:13.468 -> Water_Temp.:96.91, PWM:35
17:15:14.224 -> Water_Temp.:96.91, PWM:35
17:15:15.013 -> Water_Temp.:97.02, PWM:35
17:15:16.564 -> Water_Temp.:97.02, PWM:35
17:15:17.357 -> Water_Temp.:97.02, PWM:35
17:15:18.149 -> Water_Temp.:97.02, PWM:35
17:15:18.911 -> Water_Temp.:97.02, PWM:35
17:15:19.710 -> Water_Temp.:97.02, PWM:35
17:15:22.019 -> Water_Temp.:97.02, PWM:35
17:15:22.779 -> Water_Temp.:97.02, PWM:35
17:15:23.571 -> Water_Temp.:97.02, PWM:35
17:15:24.332 -> Water_Temp.:97.02, PWM:35
17:15:25.916 -> Water_Temp.:97.02, PWM:35
17:15:26.677 -> Water_Temp.:97.02, PWM:35
17:15:27.470 -> Water_Temp.:97.02, PWM:35
17:15:28.262 -> Water_Temp.:97.02, PWM:35
17:15:29.022 -> Water_Temp.:97.02, PWM:35
17:15:29.812 -> Water_Temp.:97.02, PWM:35
17:15:30.571 -> Water_Temp.:97.02, PWM:35
17:15:31.364 -> Water_Temp.:97.02, PWM:35
17:15:32.919 -> Water_Temp.:97.02, PWM:35
17:15:33.676 -> Water_Temp.:97.02, PWM:35
17:15:34.470 -> Water_Temp.:97.02, PWM:35
17:15:35.263 -> Water_Temp.:97.02, PWM:35
17:15:36.023 -> Water_Temp.:97.02, PWM:35
17:15:37.608 -> Water_Temp.:97.02, PWM:35
17:15:38.363 -> Water_Temp.:97.02, PWM:35
17:15:39.161 -> Water_Temp.:97.02, PWM:35
17:15:39.917 -> Water_Temp.:97.02, PWM:35
17:15:40.710 -> Water_Temp.:97.02, PWM:35
17:15:41.504 -> Water_Temp.:97.02, PWM:35
17:15:42.264 -> Water_Temp.:97.02, PWM:35
17:15:43.056 -> Water_Temp.:97.02, PWM:35
17:15:44.608 -> Water_Temp.:97.02, PWM:35
17:15:45.402 -> Water_Temp.:97.02, PWM:35
17:15:46.164 -> Water_Temp.:97.02, PWM:35
17:15:46.957 -> Water_Temp.:97.02, PWM:35
17:15:47.746 -> Water_Temp.:97.02, PWM:35
17:15:48.506 -> Water_Temp.:97.02, PWM:35
17:15:49.305 -> Water_Temp.:97.02, PWM:35
17:15:50.062 -> Water_Temp.:97.02, PWM:35
17:15:50.853 -> Water_Temp.:97.02, PWM:35
17:15:51.613 -> Water_Temp.:97.02, PWM:35
17:15:52.408 -> Water_Temp.:97.02, PWM:35
17:15:53.204 -> Water_Temp.:97.02, PWM:35
17:15:53.966 -> Water_Temp.:97.02, PWM:35
17:15:54.760 -> Water_Temp.:97.02, PWM:35
17:15:55.517 -> Water_Temp.:97.02, PWM:35
17:15:56.313 -> Water_Temp.:97.02, PWM:35
17:15:57.072 -> Water_Temp.:97.02, PWM:35
17:15:57.867 -> Water_Temp.:97.02, PWM:35
17:15:58.663 -> Water_Temp.:97.02, PWM:35
17:15:59.419 -> Water_Temp.:97.02, PWM:35
17:16:00.211 -> Water_Temp.:97.02, PWM:35
17:16:01.002 -> Water_Temp.:97.02, PWM:35
17:16:01.761 -> Water_Temp.:97.02, PWM:35
17:16:02.556 -> Water_Temp.:97.02, PWM:35
17:16:03.316 -> Water_Temp.:97.02, PWM:35
17:16:04.109 -> Water_Temp.:97.02, PWM:35
17:16:04.907 -> Water_Temp.:97.02, PWM:35
17:16:05.668 -> Water_Temp.:97.02, PWM:35
17:16:06.461 -> Water_Temp.:97.02, PWM:35
17:16:07.219 -> Water_Temp.:97.02, PWM:35
17:16:08.012 -> Water_Temp.:97.02, PWM:35
17:16:08.802 -> Water_Temp.:97.02, PWM:35
17:16:09.561 -> Water_Temp.:97.02, PWM:35
17:16:10.359 -> Water_Temp.:97.02, PWM:35
17:16:11.117 -> Water_Temp.:97.02, PWM:35
17:16:11.909 -> Water_Temp.:97.02, PWM:35
17:16:12.704 -> Water_Temp.:97.02, PWM:35
17:16:13.459 -> Water_Temp.:97.02, PWM:35
17:16:14.256 -> Water_Temp.:97.02, PWM:35
17:16:15.017 -> Water_Temp.:97.02, PWM:35
17:16:15.811 -> Water_Temp.:97.02, PWM:35
17:16:16.604 -> Water_Temp.:97.02, PWM:35
17:16:17.363 -> Water_Temp.:97.02, PWM:35
17:16:18.156 -> Water_Temp.:97.02, PWM:35
17:16:18.915 -> Water_Temp.:97.02, PWM:35
17:16:19.708 -> Water_Temp.:97.02, PWM:35
17:16:22.813 -> Water_Temp.:97.14, PWM:35
17:16:23.577 -> Water_Temp.:97.14, PWM:35
17:16:24.372 -> Water_Temp.:97.14, PWM:35
17:16:25.134 -> Water_Temp.:97.14, PWM:35
17:16:25.927 -> Water_Temp.:97.14, PWM:35
17:16:26.720 -> Water_Temp.:97.14, PWM:35
17:16:28.235 -> Water_Temp.:97.14, PWM:35
17:16:29.027 -> Water_Temp.:97.25, PWM:35
17:16:29.786 -> Water_Temp.:97.25, PWM:35
17:16:30.584 -> Water_Temp.:97.25, PWM:35
17:16:31.377 -> Water_Temp.:97.25, PWM:35
17:16:32.141 -> Water_Temp.:97.25, PWM:35
17:16:32.933 -> Water_Temp.:97.25, PWM:35
17:16:33.691 -> Water_Temp.:97.25, PWM:35
17:16:34.486 -> Water_Temp.:97.25, PWM:35
17:16:35.244 -> Water_Temp.:97.25, PWM:35
17:16:36.040 -> Water_Temp.:97.25, PWM:35
17:16:36.835 -> Water_Temp.:97.25, PWM:35
17:16:37.596 -> Water_Temp.:97.25, PWM:35
17:16:38.391 -> Water_Temp.:97.25, PWM:35
17:16:39.151 -> Water_Temp.:97.25, PWM:35
17:16:39.944 -> Water_Temp.:97.25, PWM:35
17:16:40.738 -> Water_Temp.:97.25, PWM:35
17:16:41.494 -> Water_Temp.:97.25, PWM:35
17:16:42.287 -> Water_Temp.:97.25, PWM:35
17:16:43.049 -> Water_Temp.:97.25, PWM:35
17:16:43.839 -> Water_Temp.:97.25, PWM:35
17:16:44.603 -> Water_Temp.:97.25, PWM:35
17:16:45.397 -> Water_Temp.:97.25, PWM:35
17:16:46.188 -> Water_Temp.:97.25, PWM:35
17:16:46.949 -> Water_Temp.:97.25, PWM:35
17:16:47.744 -> Water_Temp.:97.25, PWM:35
17:16:48.503 -> Water_Temp.:97.25, PWM:35
17:16:49.296 -> Water_Temp.:97.25, PWM:35
17:16:50.087 -> Water_Temp.:97.25, PWM:35
17:16:50.879 -> Water_Temp.:97.25, PWM:35
17:16:51.644 -> Water_Temp.:97.25, PWM:35
17:16:52.402 -> Water_Temp.:97.25, PWM:35
17:16:53.992 -> Water_Temp.:97.25, PWM:35
17:16:54.749 -> Water_Temp.:97.25, PWM:35
17:16:55.544 -> Water_Temp.:98.82, PWM:35
17:16:56.303 -> Water_Temp.:100.85, PWM:35
17:16:57.094 -> Water_Temp.:98.49, PWM:35
17:16:57.883 -> Water_Temp.:98.15, PWM:35
17:16:58.643 -> Water_Temp.:97.92, PWM:35
17:16:59.439 -> Water_Temp.:97.81, PWM:35
17:17:00.200 -> Water_Temp.:97.70, PWM:35
17:17:00.993 -> Water_Temp.:97.59, PWM:35
17:17:01.789 -> Water_Temp.:97.59, PWM:35
17:17:02.549 -> Water_Temp.:97.47, PWM:35
17:17:03.344 -> Water_Temp.:97.47, PWM:35
17:17:04.136 -> Water_Temp.:97.47, PWM:35
17:17:04.896 -> Water_Temp.:97.36, PWM:35
17:17:05.685 -> Water_Temp.:97.36, PWM:35
17:17:06.473 -> Water_Temp.:97.36, PWM:35
17:17:08.023 -> Water_Temp.:97.36, PWM:35
17:17:08.781 -> Water_Temp.:97.36, PWM:35
17:17:09.576 -> Water_Temp.:97.36, PWM:35
17:17:10.337 -> Water_Temp.:97.36, PWM:35
17:17:11.130 -> Water_Temp.:97.36, PWM:35
17:17:11.890 -> Water_Temp.:97.25, PWM:35
17:17:12.688 -> Water_Temp.:97.25, PWM:35
17:17:13.481 -> Water_Temp.:97.25, PWM:35
17:17:14.241 -> Water_Temp.:97.36, PWM:35
17:17:15.035 -> Water_Temp.:97.36, PWM:35
17:17:15.797 -> Water_Temp.:97.36, PWM:35
17:17:16.592 -> Water_Temp.:97.36, PWM:35
17:17:17.353 -> Water_Temp.:97.36, PWM:35
17:17:18.144 -> Water_Temp.:97.36, PWM:35
17:17:18.938 -> Water_Temp.:97.36, PWM:35
17:17:19.698 -> Water_Temp.:97.47, PWM:35
17:17:20.492 -> Water_Temp.:97.47, PWM:35
17:17:21.250 -> Water_Temp.:97.47, PWM:35
17:17:22.045 -> Water_Temp.:97.47, PWM:35
17:17:23.599 -> Water_Temp.:97.47, PWM:35
17:17:24.359 -> Water_Temp.:97.47, PWM:35
17:17:25.152 -> Water_Temp.:97.47, PWM:35
17:17:25.909 -> Water_Temp.:97.47, PWM:35
17:17:26.703 -> Water_Temp.:97.36, PWM:35
17:17:27.491 -> Water_Temp.:97.36, PWM:35
17:17:28.252 -> Water_Temp.:97.36, PWM:35
17:17:29.046 -> Water_Temp.:97.36, PWM:35
17:17:29.842 -> Water_Temp.:97.36, PWM:35
17:17:30.598 -> Water_Temp.:97.36, PWM:35
17:17:31.392 -> Water_Temp.:97.36, PWM:35
17:17:32.148 -> Water_Temp.:97.36, PWM:35
17:17:32.943 -> Water_Temp.:97.36, PWM:35
17:17:33.734 -> Water_Temp.:97.36, PWM:35
17:17:34.496 -> Water_Temp.:97.36, PWM:35
17:17:35.294 -> Water_Temp.:97.36, PWM:35
17:17:36.053 -> Water_Temp.:97.36, PWM:35
17:17:38.403 -> Water_Temp.:97.36, PWM:35
17:17:39.200 -> Water_Temp.:97.36, PWM:35
17:17:39.957 -> Water_Temp.:97.36, PWM:35
17:17:40.745 -> Water_Temp.:97.36, PWM:35
17:17:41.506 -> Water_Temp.:97.36, PWM:35
17:17:42.299 -> Water_Temp.:97.36, PWM:35
17:17:43.086 -> Water_Temp.:97.36, PWM:35
17:17:44.639 -> Water_Temp.:97.36, PWM:35
17:17:45.400 -> Water_Temp.:97.36, PWM:35
17:17:46.196 -> Water_Temp.:97.47, PWM:35
17:17:47.753 -> Water_Temp.:97.36, PWM:35
17:17:48.516 -> Water_Temp.:97.36, PWM:35
17:17:49.309 -> Water_Temp.:97.36, PWM:35
17:17:51.616 -> Water_Temp.:97.47, PWM:35
17:17:52.412 -> Water_Temp.:97.36, PWM:35
17:17:53.171 -> Water_Temp.:97.47, PWM:35
17:17:53.968 -> Water_Temp.:97.36, PWM:35
17:17:54.724 -> Water_Temp.:97.47, PWM:35
17:17:55.520 -> Water_Temp.:97.47, PWM:35
17:17:56.315 -> Water_Temp.:97.47, PWM:35
17:17:57.074 -> Water_Temp.:97.36, PWM:35
17:17:57.866 -> Water_Temp.:97.47, PWM:35
17:17:58.624 -> Water_Temp.:97.36, PWM:35
17:17:59.418 -> Water_Temp.:97.36, PWM:35
17:18:00.210 -> Water_Temp.:97.36, PWM:35
17:18:00.970 -> Water_Temp.:97.47, PWM:35
17:18:01.760 -> Water_Temp.:97.36, PWM:35
17:18:02.550 -> Water_Temp.:97.36, PWM:35
17:18:03.312 -> Water_Temp.:97.47, PWM:35
17:18:04.104 -> Water_Temp.:97.36, PWM:35
17:18:04.863 -> Water_Temp.:97.36, PWM:35
17:18:05.656 -> Water_Temp.:97.36, PWM:35
17:18:06.452 -> Water_Temp.:97.36, PWM:35
17:18:07.213 -> Water_Temp.:97.47, PWM:35
17:18:08.008 -> Water_Temp.:97.47, PWM:35
17:18:08.772 -> Water_Temp.:97.47, PWM:35
17:18:09.570 -> Water_Temp.:97.47, PWM:35
17:18:10.330 -> Water_Temp.:97.47, PWM:35
17:18:11.125 -> Water_Temp.:97.47, PWM:35
17:18:11.885 -> Water_Temp.:97.47, PWM:35
17:18:12.676 -> Water_Temp.:97.47, PWM:35
17:18:13.466 -> Water_Temp.:97.47, PWM:35
17:18:14.223 -> Water_Temp.:97.47, PWM:35
17:18:15.017 -> Water_Temp.:97.47, PWM:35
17:18:15.810 -> Water_Temp.:97.47, PWM:35
17:18:16.571 -> Water_Temp.:97.47, PWM:35
17:18:17.365 -> Water_Temp.:97.47, PWM:35
17:18:18.125 -> Water_Temp.:97.47, PWM:35
17:18:18.919 -> Water_Temp.:97.47, PWM:35
17:18:19.683 -> Water_Temp.:97.47, PWM:35
17:18:20.479 -> Water_Temp.:97.47, PWM:35
17:18:21.243 -> Water_Temp.:97.47, PWM:35
17:18:22.036 -> Water_Temp.:97.47, PWM:35
17:18:22.831 -> Water_Temp.:97.47, PWM:35
17:18:23.595 -> Water_Temp.:97.47, PWM:35
17:18:24.391 -> Water_Temp.:97.47, PWM:35
17:18:25.151 -> Water_Temp.:97.47, PWM:35
17:18:25.943 -> Water_Temp.:97.47, PWM:35
17:18:26.703 -> Water_Temp.:97.47, PWM:35
17:18:27.496 -> Water_Temp.:97.47, PWM:35
17:18:28.287 -> Water_Temp.:97.47, PWM:35
17:18:29.049 -> Water_Temp.:97.47, PWM:35
17:18:29.838 -> Water_Temp.:97.47, PWM:35
17:18:30.632 -> Water_Temp.:97.47, PWM:35
17:18:31.388 -> Water_Temp.:97.47, PWM:35
17:18:32.176 -> Water_Temp.:97.47, PWM:35
17:18:32.964 -> Water_Temp.:97.47, PWM:35
17:18:33.724 -> Water_Temp.:97.47, PWM:35
17:18:34.517 -> Water_Temp.:97.59, PWM:35
17:18:35.312 -> Water_Temp.:97.47, PWM:35
17:18:36.072 -> Water_Temp.:97.47, PWM:35
17:18:36.863 -> Water_Temp.:97.59, PWM:35
17:18:37.656 -> Water_Temp.:97.47, PWM:35
17:18:38.418 -> Water_Temp.:97.59, PWM:35
17:18:39.180 -> Water_Temp.:97.59, PWM:35
17:18:39.976 -> Water_Temp.:97.59, PWM:35
17:18:40.739 -> Water_Temp.:97.47, PWM:35
17:18:41.530 -> Water_Temp.:97.47, PWM:35
17:18:43.845 -> Water_Temp.:97.59, PWM:35
17:18:44.643 -> Water_Temp.:97.59, PWM:35
17:18:45.403 -> Water_Temp.:97.59, PWM:35
17:18:46.989 -> Water_Temp.:101.19, PWM:35
17:18:47.746 -> Water_Temp.:98.82, PWM:35
17:18:48.539 -> Water_Temp.:98.49, PWM:35
17:18:49.301 -> Water_Temp.:98.26, PWM:35
17:18:50.094 -> Water_Temp.:98.15, PWM:35
17:18:50.882 -> Water_Temp.:98.04, PWM:35
17:18:51.639 -> Water_Temp.:97.92, PWM:35
17:18:52.436 -> Water_Temp.:97.92, PWM:35
17:18:53.985 -> Water_Temp.:97.81, PWM:35
17:18:55.540 -> Water_Temp.:97.70, PWM:35
17:18:56.301 -> Water_Temp.:97.70, PWM:35
17:18:57.093 -> Water_Temp.:97.70, PWM:35
17:18:57.851 -> Water_Temp.:97.70, PWM:35
17:18:58.649 -> Water_Temp.:97.59, PWM:35
17:18:59.440 -> Water_Temp.:97.59, PWM:35
17:19:00.198 -> Water_Temp.:97.59, PWM:35
17:19:00.987 -> Water_Temp.:97.59, PWM:35
17:19:01.782 -> Water_Temp.:97.59, PWM:35
17:19:02.542 -> Water_Temp.:97.59, PWM:35
17:19:03.333 -> Water_Temp.:97.59, PWM:35
17:19:04.093 -> Water_Temp.:97.59, PWM:35
17:19:04.881 -> Water_Temp.:97.59, PWM:35
17:19:05.673 -> Water_Temp.:97.59, PWM:35
17:19:06.433 -> Water_Temp.:97.59, PWM:35
17:19:07.228 -> Water_Temp.:97.59, PWM:35
17:19:08.018 -> Water_Temp.:97.59, PWM:35
17:19:08.778 -> Water_Temp.:97.59, PWM:35
17:19:10.338 -> Water_Temp.:97.59, PWM:35
17:19:11.131 -> Water_Temp.:97.59, PWM:35
17:19:12.647 -> Water_Temp.:97.59, PWM:35
17:19:13.447 -> Water_Temp.:97.59, PWM:35
17:19:14.201 -> Water_Temp.:97.59, PWM:35
17:19:14.997 -> Water_Temp.:97.59, PWM:35
17:19:15.793 -> Water_Temp.:97.47, PWM:35
17:19:16.553 -> Water_Temp.:97.47, PWM:35
17:19:17.347 -> Water_Temp.:97.47, PWM:35
17:19:18.867 -> Water_Temp.:97.47, PWM:35
17:19:19.665 -> Water_Temp.:97.47, PWM:35
17:19:20.429 -> Water_Temp.:97.47, PWM:35
17:19:21.224 -> Water_Temp.:97.47, PWM:35
17:19:21.985 -> Water_Temp.:97.47, PWM:35
17:19:22.779 -> Water_Temp.:97.59, PWM:35
17:19:23.538 -> Water_Temp.:97.59, PWM:35
17:19:24.327 -> Water_Temp.:97.59, PWM:35
17:19:25.124 -> Water_Temp.:97.59, PWM:35
17:19:25.880 -> Water_Temp.:97.59, PWM:35
17:19:26.673 -> Water_Temp.:97.59, PWM:35
17:19:27.464 -> Water_Temp.:97.59, PWM:35
17:19:28.224 -> Water_Temp.:97.59, PWM:35
17:19:29.016 -> Water_Temp.:97.59, PWM:35
17:19:29.805 -> Water_Temp.:97.59, PWM:35
17:19:30.563 -> Water_Temp.:97.59, PWM:35
17:19:31.354 -> Water_Temp.:97.59, PWM:35
17:19:32.146 -> Water_Temp.:97.59, PWM:35
17:19:33.666 -> Water_Temp.:97.59, PWM:35
17:19:34.455 -> Water_Temp.:97.59, PWM:35
17:19:35.976 -> Water_Temp.:97.59, PWM:35
17:19:36.766 -> Water_Temp.:97.59, PWM:35
17:19:37.563 -> Water_Temp.:97.59, PWM:35
17:19:38.319 -> Water_Temp.:97.70, PWM:35
17:19:39.875 -> Water_Temp.:97.70, PWM:35
17:19:40.667 -> Water_Temp.:97.70, PWM:35
17:19:41.427 -> Water_Temp.:97.70, PWM:35
17:19:42.217 -> Water_Temp.:97.70, PWM:35
17:19:43.010 -> Water_Temp.:97.70, PWM:35
17:19:43.770 -> Water_Temp.:97.70, PWM:35
17:19:44.565 -> Water_Temp.:97.70, PWM:35
17:19:45.324 -> Water_Temp.:97.70, PWM:35
17:19:46.880 -> Water_Temp.:97.70, PWM:35
17:19:47.675 -> Water_Temp.:97.70, PWM:35
17:19:48.471 -> Water_Temp.:97.70, PWM:35
17:19:49.229 -> Water_Temp.:97.70, PWM:35
17:19:50.021 -> Water_Temp.:97.70, PWM:35
17:19:50.778 -> Water_Temp.:97.70, PWM:35
17:19:51.572 -> Water_Temp.:97.70, PWM:35
17:19:52.365 -> Water_Temp.:97.81, PWM:35
17:19:53.127 -> Water_Temp.:97.70, PWM:35
17:19:53.925 -> Water_Temp.:97.70, PWM:35
17:19:54.687 -> Water_Temp.:97.70, PWM:35
17:19:56.240 -> Water_Temp.:97.70, PWM:35
17:19:57.032 -> Water_Temp.:97.70, PWM:35
17:19:57.792 -> Water_Temp.:97.70, PWM:35
17:19:58.585 -> Water_Temp.:97.70, PWM:35
17:19:59.340 -> Water_Temp.:97.70, PWM:35
17:20:00.135 -> Water_Temp.:97.70, PWM:35
17:20:00.928 -> Water_Temp.:97.70, PWM:35
17:20:01.685 -> Water_Temp.:97.70, PWM:35
17:20:02.481 -> Water_Temp.:97.70, PWM:35
17:20:04.037 -> Water_Temp.:97.70, PWM:35
17:20:04.829 -> Water_Temp.:97.70, PWM:35
17:20:05.585 -> Water_Temp.:97.70, PWM:35
17:20:06.377 -> Water_Temp.:97.70, PWM:35
17:20:07.169 -> Water_Temp.:97.70, PWM:35
17:20:07.928 -> Water_Temp.:97.70, PWM:35
17:20:08.721 -> Water_Temp.:97.70, PWM:35
17:20:09.485 -> Water_Temp.:97.70, PWM:35
17:20:10.277 -> Water_Temp.:97.70, PWM:35
17:20:11.825 -> Water_Temp.:97.70, PWM:35
17:20:12.619 -> Water_Temp.:97.70, PWM:35
17:20:13.378 -> Water_Temp.:97.70, PWM:35
17:20:14.172 -> Water_Temp.:97.70, PWM:35
17:20:14.965 -> Water_Temp.:97.70, PWM:35
17:20:15.726 -> Water_Temp.:97.70, PWM:35
17:20:16.522 -> Water_Temp.:97.70, PWM:35
17:20:17.283 -> Water_Temp.:97.70, PWM:35
17:20:18.074 -> Water_Temp.:97.70, PWM:35
17:20:18.869 -> Water_Temp.:97.70, PWM:35
17:20:19.625 -> Water_Temp.:97.70, PWM:35
17:20:20.417 -> Water_Temp.:97.70, PWM:35
17:20:21.178 -> Water_Temp.:97.70, PWM:35
17:20:21.973 -> Water_Temp.:97.70, PWM:35
17:20:22.765 -> Water_Temp.:97.70, PWM:35
17:20:24.319 -> Water_Temp.:97.70, PWM:35
17:20:25.080 -> Water_Temp.:97.70, PWM:35
17:20:25.877 -> Water_Temp.:97.70, PWM:35
17:20:26.639 -> Water_Temp.:97.70, PWM:35
17:20:27.435 -> Water_Temp.:97.70, PWM:35
17:20:28.197 -> Water_Temp.:97.70, PWM:35
17:20:28.991 -> Water_Temp.:97.70, PWM:35
17:20:29.756 -> Water_Temp.:97.70, PWM:35
17:20:30.545 -> Water_Temp.:97.81, PWM:35
17:20:31.336 -> Water_Temp.:97.81, PWM:35
17:20:32.095 -> Water_Temp.:97.81, PWM:35
17:20:32.887 -> Water_Temp.:97.81, PWM:35
17:20:33.680 -> Water_Temp.:97.81, PWM:35
17:20:34.442 -> Water_Temp.:97.81, PWM:35
17:20:35.238 -> Water_Temp.:97.81, PWM:35
17:20:35.999 -> Water_Temp.:97.81, PWM:35
17:20:36.793 -> Water_Temp.:97.81, PWM:35
17:20:37.586 -> Water_Temp.:97.81, PWM:35
17:20:38.342 -> Water_Temp.:97.81, PWM:35
17:20:39.138 -> Water_Temp.:97.81, PWM:35
17:20:39.898 -> Water_Temp.:97.81, PWM:35
17:20:40.696 -> Water_Temp.:97.81, PWM:35
17:20:42.211 -> Water_Temp.:97.81, PWM:35
17:20:43.004 -> Water_Temp.:97.81, PWM:35
17:20:43.797 -> Water_Temp.:97.81, PWM:35
17:20:44.556 -> Water_Temp.:97.81, PWM:35
17:20:45.353 -> Water_Temp.:97.81, PWM:35
17:20:46.113 -> Water_Temp.:97.81, PWM:35
17:20:46.906 -> Water_Temp.:97.81, PWM:35
17:20:47.701 -> Water_Temp.:97.81, PWM:35
17:20:48.460 -> Water_Temp.:97.81, PWM:35
17:20:49.250 -> Water_Temp.:97.81, PWM:35
17:20:50.009 -> Water_Temp.:97.81, PWM:35
17:20:50.802 -> Water_Temp.:97.81, PWM:35
17:20:51.596 -> Water_Temp.:97.92, PWM:35
17:20:52.357 -> Water_Temp.:97.92, PWM:35
17:20:53.149 -> Water_Temp.:97.92, PWM:35
17:20:53.908 -> Water_Temp.:97.92, PWM:35
17:20:54.698 -> Water_Temp.:97.92, PWM:35
17:20:55.490 -> Water_Temp.:97.92, PWM:35
17:20:56.253 -> Water_Temp.:97.92, PWM:35
17:20:57.049 -> Water_Temp.:97.92, PWM:35
17:20:57.811 -> Water_Temp.:97.92, PWM:35
17:20:58.606 -> Water_Temp.:97.92, PWM:35
17:20:59.367 -> Water_Temp.:97.92, PWM:35
17:21:00.162 -> Water_Temp.:97.92, PWM:35
17:21:00.959 -> Water_Temp.:97.92, PWM:35
17:21:01.714 -> Water_Temp.:97.92, PWM:35
17:21:02.507 -> Water_Temp.:97.92, PWM:35
17:21:03.301 -> Water_Temp.:97.92, PWM:35
17:21:04.057 -> Water_Temp.:97.92, PWM:35
17:21:04.852 -> Water_Temp.:97.92, PWM:35
17:21:05.615 -> Water_Temp.:97.92, PWM:35
17:21:06.409 -> Water_Temp.:97.92, PWM:35
17:21:07.168 -> Water_Temp.:98.04, PWM:35
17:21:07.958 -> Water_Temp.:97.92, PWM:35
17:21:09.511 -> Water_Temp.:98.04, PWM:35
17:21:10.306 -> Water_Temp.:97.92, PWM:35
17:21:11.066 -> Water_Temp.:98.04, PWM:35
17:21:11.856 -> Water_Temp.:98.04, PWM:35
17:21:13.403 -> Water_Temp.:97.92, PWM:35
17:21:14.198 -> Water_Temp.:97.92, PWM:35
17:21:14.956 -> Water_Temp.:97.92, PWM:35
17:21:15.746 -> Water_Temp.:97.92, PWM:35
17:21:16.505 -> Water_Temp.:97.92, PWM:35
17:21:17.299 -> Water_Temp.:97.92, PWM:35
17:21:18.853 -> Water_Temp.:97.92, PWM:35
17:21:19.610 -> Water_Temp.:97.92, PWM:35
17:21:20.407 -> Water_Temp.:97.92, PWM:35
17:21:21.169 -> Water_Temp.:97.92, PWM:35
17:21:21.964 -> Water_Temp.:97.92, PWM:35
17:21:22.724 -> Water_Temp.:97.92, PWM:35
17:21:23.515 -> Water_Temp.:97.92, PWM:35
17:21:24.313 -> Water_Temp.:97.92, PWM:35
17:21:25.834 -> Water_Temp.:97.92, PWM:35
17:21:26.594 -> Water_Temp.:97.92, PWM:35
17:21:27.387 -> Water_Temp.:97.92, PWM:35
17:21:28.181 -> Water_Temp.:97.92, PWM:35
17:21:28.939 -> Water_Temp.:97.92, PWM:35
17:21:29.731 -> Water_Temp.:97.92, PWM:35
17:21:30.525 -> Water_Temp.:97.92, PWM:35
17:21:32.041 -> Water_Temp.:97.92, PWM:35
17:21:32.835 -> Water_Temp.:97.92, PWM:35
17:21:33.624 -> Water_Temp.:97.92, PWM:35
17:21:34.384 -> Water_Temp.:97.92, PWM:35
17:21:35.178 -> Water_Temp.:98.04, PWM:35
17:21:35.972 -> Water_Temp.:97.92, PWM:35
17:21:36.731 -> Water_Temp.:97.92, PWM:35
17:21:37.526 -> Water_Temp.:97.92, PWM:35
17:21:38.280 -> Water_Temp.:97.92, PWM:35
17:21:39.073 -> Water_Temp.:98.04, PWM:35
17:21:40.632 -> Water_Temp.:97.92, PWM:35
17:21:41.424 -> Water_Temp.:98.04, PWM:35
17:21:42.182 -> Water_Temp.:97.92, PWM:35
17:21:42.975 -> Water_Temp.:97.92, PWM:35
17:21:43.769 -> Water_Temp.:98.04, PWM:35
17:21:44.533 -> Water_Temp.:98.04, PWM:35
17:21:45.332 -> Water_Temp.:98.04, PWM:35
17:21:46.096 -> Water_Temp.:98.04, PWM:35
17:21:46.855 -> Water_Temp.:98.04, PWM:35
17:21:47.645 -> Water_Temp.:98.04, PWM:35
17:21:48.441 -> Water_Temp.:98.04, PWM:35
17:21:49.204 -> Water_Temp.:97.92, PWM:35
17:21:49.998 -> Water_Temp.:98.04, PWM:35
17:21:50.759 -> Water_Temp.:97.92, PWM:35
17:21:51.555 -> Water_Temp.:97.92, PWM:35
17:21:53.112 -> Water_Temp.:97.92, PWM:35
17:21:53.909 -> Water_Temp.:97.92, PWM:35
17:21:55.463 -> Water_Temp.:98.04, PWM:35
17:21:56.225 -> Water_Temp.:97.92, PWM:35
17:21:57.023 -> Water_Temp.:97.92, PWM:35
17:21:57.781 -> Water_Temp.:97.92, PWM:35
17:21:58.576 -> Water_Temp.:98.04, PWM:35
17:21:59.339 -> Water_Temp.:98.04, PWM:35
17:22:00.892 -> Water_Temp.:98.04, PWM:35
17:22:01.687 -> Water_Temp.:98.04, PWM:35
17:22:02.450 -> Water_Temp.:98.04, PWM:35
17:22:03.242 -> Water_Temp.:97.92, PWM:35
17:22:04.000 -> Water_Temp.:97.92, PWM:35
17:22:04.797 -> Water_Temp.:98.04, PWM:35
17:22:06.352 -> Water_Temp.:97.92, PWM:35
17:22:07.145 -> Water_Temp.:97.92, PWM:35
17:22:07.905 -> Water_Temp.:97.92, PWM:35
17:22:08.698 -> Water_Temp.:98.04, PWM:35
17:22:09.464 -> Water_Temp.:98.04, PWM:35
17:22:10.256 -> Water_Temp.:97.92, PWM:35
17:22:11.018 -> Water_Temp.:97.92, PWM:35
17:22:12.569 -> Water_Temp.:97.92, PWM:35
17:22:14.124 -> Water_Temp.:97.92, PWM:35
17:22:14.919 -> Water_Temp.:97.92, PWM:35
17:22:15.680 -> Water_Temp.:97.92, PWM:35
17:22:16.477 -> Water_Temp.:97.92, PWM:35
17:22:17.237 -> Water_Temp.:97.92, PWM:35
17:22:18.036 -> Water_Temp.:97.92, PWM:35
17:22:18.797 -> Water_Temp.:97.92, PWM:35
17:22:19.593 -> Water_Temp.:97.92, PWM:35
17:22:20.357 -> Water_Temp.:97.92, PWM:35
17:22:21.151 -> Water_Temp.:97.92, PWM:35
17:22:21.916 -> Water_Temp.:97.92, PWM:35
17:22:22.710 -> Water_Temp.:97.92, PWM:35
17:22:23.473 -> Water_Temp.:97.92, PWM:35
17:22:24.269 -> Water_Temp.:97.92, PWM:35
17:22:25.030 -> Water_Temp.:97.92, PWM:35
17:22:25.828 -> Water_Temp.:97.92, PWM:35
17:22:26.622 -> Water_Temp.:97.92, PWM:35
17:22:27.380 -> Water_Temp.:97.92, PWM:35
17:22:28.172 -> Water_Temp.:97.92, PWM:35
17:22:28.932 -> Water_Temp.:97.92, PWM:35
17:22:29.729 -> Water_Temp.:97.92, PWM:35
17:22:30.490 -> Water_Temp.:98.04, PWM:35
17:22:31.288 -> Water_Temp.:98.04, PWM:35
17:22:32.081 -> Water_Temp.:98.04, PWM:35
17:22:32.837 -> Water_Temp.:98.04, PWM:35
17:22:33.633 -> Water_Temp.:98.04, PWM:35
17:22:34.395 -> Water_Temp.:98.04, PWM:35
17:22:35.191 -> Water_Temp.:98.04, PWM:35
17:22:35.947 -> Water_Temp.:98.04, PWM:35
17:22:37.504 -> Water_Temp.:98.04, PWM:35
17:22:38.300 -> Water_Temp.:98.04, PWM:35
17:22:39.063 -> Water_Temp.:98.04, PWM:35
17:22:39.861 -> Water_Temp.:98.04, PWM:35
17:22:40.654 -> Water_Temp.:98.04, PWM:35
17:22:41.415 -> Water_Temp.:98.04, PWM:35
17:22:42.209 -> Water_Temp.:98.04, PWM:35
17:22:42.972 -> Water_Temp.:98.04, PWM:35
17:22:43.767 -> Water_Temp.:98.04, PWM:35
17:22:44.557 -> Water_Temp.:98.04, PWM:35
17:22:45.320 -> Water_Temp.:98.04, PWM:35
17:22:46.116 -> Water_Temp.:98.04, PWM:35
17:22:46.876 -> Water_Temp.:98.04, PWM:35
17:22:47.669 -> Water_Temp.:98.15, PWM:35
17:22:48.431 -> Water_Temp.:98.04, PWM:35
17:22:49.226 -> Water_Temp.:98.15, PWM:35
17:22:49.987 -> Water_Temp.:98.15, PWM:35
17:22:50.781 -> Water_Temp.:98.15, PWM:35
17:22:51.573 -> Water_Temp.:98.15, PWM:35
17:22:52.333 -> Water_Temp.:98.15, PWM:35
17:22:53.126 -> Water_Temp.:98.15, PWM:35
17:22:53.886 -> Water_Temp.:98.15, PWM:35
17:22:54.679 -> Water_Temp.:98.15, PWM:35
17:22:55.444 -> Water_Temp.:98.15, PWM:35
17:22:56.239 -> Water_Temp.:98.15, PWM:35
17:22:57.034 -> Water_Temp.:98.15, PWM:35
17:22:57.797 -> Water_Temp.:98.15, PWM:35
17:22:59.323 -> Water_Temp.:98.15, PWM:35
17:23:00.122 -> Water_Temp.:98.15, PWM:35
17:23:00.881 -> Water_Temp.:98.15, PWM:35
17:23:01.678 -> Water_Temp.:98.15, PWM:35
17:23:02.439 -> Water_Temp.:98.15, PWM:35
17:23:03.231 -> Water_Temp.:98.15, PWM:35
17:23:03.995 -> Water_Temp.:98.15, PWM:35
17:23:04.793 -> Water_Temp.:98.15, PWM:35
17:23:05.587 -> Water_Temp.:98.15, PWM:35
17:23:06.350 -> Water_Temp.:98.15, PWM:35
17:23:07.143 -> Water_Temp.:98.15, PWM:35
17:23:07.903 -> Water_Temp.:98.15, PWM:35
17:23:08.696 -> Water_Temp.:98.15, PWM:35
17:23:09.459 -> Water_Temp.:98.15, PWM:35
17:23:10.251 -> Water_Temp.:98.15, PWM:35
17:23:11.015 -> Water_Temp.:98.15, PWM:35
17:23:11.811 -> Water_Temp.:98.15, PWM:35
17:23:12.606 -> Water_Temp.:98.26, PWM:35
17:23:13.364 -> Water_Temp.:98.26, PWM:35
17:23:14.153 -> Water_Temp.:98.26, PWM:35
17:23:14.916 -> Water_Temp.:98.26, PWM:35
17:23:15.709 -> Water_Temp.:98.26, PWM:35
17:23:17.233 -> Water_Temp.:98.26, PWM:35
17:23:18.032 -> Water_Temp.:98.26, PWM:35
17:23:18.796 -> Water_Temp.:98.26, PWM:35
17:23:19.591 -> Water_Temp.:98.26, PWM:35
17:23:20.352 -> Water_Temp.:98.26, PWM:35
17:23:21.146 -> Water_Temp.:98.37, PWM:35
17:23:21.942 -> Water_Temp.:99.95, PWM:35
17:23:22.704 -> Water_Temp.:101.97, PWM:35
17:23:23.501 -> Water_Temp.:99.61, PWM:35
17:23:24.266 -> Water_Temp.:99.16, PWM:35
17:23:25.029 -> Water_Temp.:98.94, PWM:35
17:23:25.824 -> Water_Temp.:98.82, PWM:35
17:23:26.618 -> Water_Temp.:98.71, PWM:35
17:23:28.172 -> Water_Temp.:98.60, PWM:35
17:23:28.933 -> Water_Temp.:98.60, PWM:35
17:23:29.725 -> Water_Temp.:98.60, PWM:35
17:23:30.520 -> Water_Temp.:98.49, PWM:35
17:23:32.075 -> Water_Temp.:98.49, PWM:35
17:23:33.632 -> Water_Temp.:98.37, PWM:35
17:23:34.391 -> Water_Temp.:98.37, PWM:35
17:23:35.190 -> Water_Temp.:98.37, PWM:35
17:23:36.743 -> Water_Temp.:98.37, PWM:35
17:23:37.505 -> Water_Temp.:98.37, PWM:35
17:23:38.301 -> Water_Temp.:98.37, PWM:35
17:23:39.096 -> Water_Temp.:98.37, PWM:35
17:23:39.855 -> Water_Temp.:98.37, PWM:35
17:23:40.649 -> Water_Temp.:98.37, PWM:35
17:23:41.412 -> Water_Temp.:98.37, PWM:35
17:23:42.209 -> Water_Temp.:98.37, PWM:35
17:23:42.971 -> Water_Temp.:98.37, PWM:35
17:23:43.765 -> Water_Temp.:98.37, PWM:35
17:23:44.529 -> Water_Temp.:98.37, PWM:35
17:23:45.326 -> Water_Temp.:98.37, PWM:35
17:23:46.881 -> Water_Temp.:98.37, PWM:35
17:23:47.646 -> Water_Temp.:98.37, PWM:35
17:23:48.442 -> Water_Temp.:98.37, PWM:35
17:23:50.004 -> Water_Temp.:98.37, PWM:35
17:23:50.763 -> Water_Temp.:98.37, PWM:35
17:23:51.557 -> Water_Temp.:98.37, PWM:35
17:23:52.320 -> Water_Temp.:98.37, PWM:35
17:23:53.111 -> Water_Temp.:98.37, PWM:35
17:23:53.909 -> Water_Temp.:98.37, PWM:35
17:23:54.673 -> Water_Temp.:98.37, PWM:35
17:23:55.472 -> Water_Temp.:98.37, PWM:35
17:23:56.233 -> Water_Temp.:98.37, PWM:35
17:23:57.026 -> Water_Temp.:98.37, PWM:35
17:23:57.786 -> Water_Temp.:98.49, PWM:35
17:23:58.579 -> Water_Temp.:98.49, PWM:35
17:23:59.341 -> Water_Temp.:98.49, PWM:35
17:24:00.135 -> Water_Temp.:98.49, PWM:35
17:24:01.692 -> Water_Temp.:98.37, PWM:35
17:24:02.455 -> Water_Temp.:98.37, PWM:35
17:24:03.245 -> Water_Temp.:98.37, PWM:35
17:24:04.001 -> Water_Temp.:98.37, PWM:35
17:24:04.798 -> Water_Temp.:98.37, PWM:35
17:24:05.591 -> Water_Temp.:98.37, PWM:35
17:24:06.347 -> Water_Temp.:98.37, PWM:35
17:24:07.142 -> Water_Temp.:98.37, PWM:35
17:24:07.902 -> Water_Temp.:98.37, PWM:35
17:24:08.694 -> Water_Temp.:98.37, PWM:35
17:24:09.491 -> Water_Temp.:98.37, PWM:35
17:24:10.248 -> Water_Temp.:98.37, PWM:35
17:24:11.039 -> Water_Temp.:98.37, PWM:35
17:24:11.831 -> Water_Temp.:98.37, PWM:35
17:24:12.592 -> Water_Temp.:98.37, PWM:35
17:24:13.386 -> Water_Temp.:98.37, PWM:35
17:24:14.143 -> Water_Temp.:98.37, PWM:35
17:24:14.938 -> Water_Temp.:98.37, PWM:35
17:24:15.733 -> Water_Temp.:98.37, PWM:35
17:24:16.491 -> Water_Temp.:98.37, PWM:35
17:24:17.286 -> Water_Temp.:98.37, PWM:35
17:24:18.045 -> Water_Temp.:98.37, PWM:35
17:24:18.841 -> Water_Temp.:98.37, PWM:35
17:24:19.606 -> Water_Temp.:98.37, PWM:35
17:24:20.400 -> Water_Temp.:98.37, PWM:35
17:24:21.160 -> Water_Temp.:98.37, PWM:35
17:24:22.714 -> Water_Temp.:98.37, PWM:35
17:24:23.474 -> Water_Temp.:98.37, PWM:35
17:24:24.268 -> Water_Temp.:98.37, PWM:35
17:24:25.062 -> Water_Temp.:98.37, PWM:35
17:24:25.816 -> Water_Temp.:98.37, PWM:35
17:24:26.611 -> Water_Temp.:98.37, PWM:35
17:24:27.405 -> Water_Temp.:98.37, PWM:35
17:24:28.163 -> Water_Temp.:98.37, PWM:35
17:24:28.956 -> Water_Temp.:98.37, PWM:35
17:24:29.718 -> Water_Temp.:98.37, PWM:35
17:24:30.513 -> Water_Temp.:98.37, PWM:35
17:24:31.273 -> Water_Temp.:98.37, PWM:35
17:24:32.068 -> Water_Temp.:98.37, PWM:35
17:24:32.863 -> Water_Temp.:98.37, PWM:35
17:24:33.625 -> Water_Temp.:98.37, PWM:35
17:24:34.415 -> Water_Temp.:98.37, PWM:35
17:24:35.174 -> Water_Temp.:98.37, PWM:35
17:24:35.969 -> Water_Temp.:98.37, PWM:35
17:24:36.764 -> Water_Temp.:98.37, PWM:35
17:24:37.520 -> Water_Temp.:98.37, PWM:35
17:24:38.317 -> Water_Temp.:98.37, PWM:35
17:24:39.074 -> Water_Temp.:98.49, PWM:35
17:24:39.872 -> Water_Temp.:98.37, PWM:35
17:24:40.632 -> Water_Temp.:98.37, PWM:35
17:24:41.422 -> Water_Temp.:98.37, PWM:35
17:24:42.216 -> Water_Temp.:100.06, PWM:35
17:24:43.006 -> Water_Temp.:102.09, PWM:35
17:24:43.763 -> Water_Temp.:99.72, PWM:35
17:24:44.559 -> Water_Temp.:99.27, PWM:35
17:24:45.319 -> Water_Temp.:99.05, PWM:35
17:24:46.107 -> Water_Temp.:98.94, PWM:35
17:24:46.899 -> Water_Temp.:98.82, PWM:35
17:24:47.661 -> Water_Temp.:98.82, PWM:35
17:24:48.456 -> Water_Temp.:98.71, PWM:35
17:24:49.220 -> Water_Temp.:98.71, PWM:35
17:24:50.016 -> Water_Temp.:98.60, PWM:35
17:24:50.772 -> Water_Temp.:98.60, PWM:35
17:24:51.562 -> Water_Temp.:98.60, PWM:35
17:24:52.352 -> Water_Temp.:98.60, PWM:35
17:24:53.115 -> Water_Temp.:98.49, PWM:35
17:24:53.910 -> Water_Temp.:98.49, PWM:35
17:24:54.699 -> Water_Temp.:98.49, PWM:35
17:24:55.463 -> Water_Temp.:98.49, PWM:35
17:24:56.261 -> Water_Temp.:98.49, PWM:35
17:24:57.016 -> Water_Temp.:98.49, PWM:35
17:24:57.805 -> Water_Temp.:98.49, PWM:35
17:24:58.602 -> Water_Temp.:98.49, PWM:35
17:24:59.364 -> Water_Temp.:98.37, PWM:35
17:25:00.156 -> Water_Temp.:98.37, PWM:35
17:25:00.916 -> Water_Temp.:98.37, PWM:35
17:25:01.708 -> Water_Temp.:98.37, PWM:35
17:25:02.471 -> Water_Temp.:98.37, PWM:35
17:25:04.060 -> Water_Temp.:98.49, PWM:35
17:25:04.820 -> Water_Temp.:98.49, PWM:35
17:25:05.614 -> Water_Temp.:98.49, PWM:35
17:25:06.371 -> Water_Temp.:98.49, PWM:35
17:25:07.164 -> Water_Temp.:98.49, PWM:35
17:25:07.956 -> Water_Temp.:98.49, PWM:35
17:25:08.710 -> Water_Temp.:98.49, PWM:35
17:25:09.504 -> Water_Temp.:98.49, PWM:35
17:25:11.057 -> Water_Temp.:98.49, PWM:35
17:25:11.851 -> Water_Temp.:98.49, PWM:35
17:25:12.609 -> Water_Temp.:98.49, PWM:35
17:25:13.402 -> Water_Temp.:98.49, PWM:35
17:25:14.193 -> Water_Temp.:98.49, PWM:35
17:25:14.949 -> Water_Temp.:98.49, PWM:35
17:25:15.743 -> Water_Temp.:98.49, PWM:35
17:25:16.537 -> Water_Temp.:98.49, PWM:35
17:25:17.295 -> Water_Temp.:98.49, PWM:35
17:25:18.090 -> Water_Temp.:98.49, PWM:35
17:25:18.852 -> Water_Temp.:98.49, PWM:35
17:25:19.648 -> Water_Temp.:98.49, PWM:35
17:25:20.442 -> Water_Temp.:98.49, PWM:35
17:25:21.201 -> Water_Temp.:98.49, PWM:35
17:25:21.998 -> Water_Temp.:98.49, PWM:35
17:25:22.757 -> Water_Temp.:98.49, PWM:35
17:25:23.548 -> Water_Temp.:98.49, PWM:35
17:25:24.340 -> Water_Temp.:98.49, PWM:35
17:25:25.102 -> Water_Temp.:98.49, PWM:35
17:25:26.655 -> Water_Temp.:98.49, PWM:35
17:25:27.448 -> Water_Temp.:98.49, PWM:35
17:25:28.972 -> Water_Temp.:98.49, PWM:35
17:25:29.768 -> Water_Temp.:98.49, PWM:35
17:25:31.321 -> Water_Temp.:98.49, PWM:35
17:25:32.078 -> Water_Temp.:98.49, PWM:35
17:25:32.874 -> Water_Temp.:98.49, PWM:35
17:25:33.635 -> Water_Temp.:98.49, PWM:35
17:25:34.430 -> Water_Temp.:98.49, PWM:35
17:25:35.190 -> Water_Temp.:98.49, PWM:35
17:25:35.987 -> Water_Temp.:98.49, PWM:35
17:25:36.747 -> Water_Temp.:98.49, PWM:35
17:25:37.539 -> Water_Temp.:98.49, PWM:35
17:25:39.094 -> Water_Temp.:98.49, PWM:35
17:25:39.889 -> Water_Temp.:98.49, PWM:35
17:25:40.652 -> Water_Temp.:98.49, PWM:35
17:25:41.448 -> Water_Temp.:98.49, PWM:35
17:25:42.210 -> Water_Temp.:98.49, PWM:35
17:25:43.005 -> Water_Temp.:98.49, PWM:35
17:25:43.768 -> Water_Temp.:98.49, PWM:35
17:25:44.559 -> Water_Temp.:98.49, PWM:35
17:25:45.355 -> Water_Temp.:98.49, PWM:35
17:25:46.115 -> Water_Temp.:98.49, PWM:35
17:25:46.907 -> Water_Temp.:98.49, PWM:35
17:25:47.699 -> Water_Temp.:98.49, PWM:35
17:25:48.458 -> Water_Temp.:98.49, PWM:35
17:25:49.255 -> Water_Temp.:98.49, PWM:35
17:25:50.012 -> Water_Temp.:98.49, PWM:35
17:25:50.803 -> Water_Temp.:98.49, PWM:35
17:25:51.566 -> Water_Temp.:98.49, PWM:35
17:25:52.360 -> Water_Temp.:98.49, PWM:35
17:25:53.155 -> Water_Temp.:98.49, PWM:35
17:25:53.917 -> Water_Temp.:98.49, PWM:35
17:25:54.709 -> Water_Temp.:98.49, PWM:35
17:25:55.474 -> Water_Temp.:98.49, PWM:35
17:25:56.269 -> Water_Temp.:98.49, PWM:35
17:25:57.029 -> Water_Temp.:98.49, PWM:35
17:25:58.580 -> Water_Temp.:98.49, PWM:35
17:25:59.377 -> Water_Temp.:98.49, PWM:35
17:26:00.133 -> Water_Temp.:98.49, PWM:35
17:26:00.926 -> Water_Temp.:98.49, PWM:35
17:26:01.720 -> Water_Temp.:98.49, PWM:35
17:26:02.477 -> Water_Temp.:98.49, PWM:35
17:26:03.272 -> Water_Temp.:98.49, PWM:35
17:26:04.034 -> Water_Temp.:98.49, PWM:35
17:26:04.831 -> Water_Temp.:98.49, PWM:35
17:26:06.380 -> Water_Temp.:98.49, PWM:35
17:26:07.171 -> Water_Temp.:98.49, PWM:35
17:26:07.925 -> Water_Temp.:98.49, PWM:35
17:26:08.714 -> Water_Temp.:98.49, PWM:35
17:26:09.511 -> Water_Temp.:98.49, PWM:35
17:26:11.058 -> Water_Temp.:98.49, PWM:35
17:26:11.851 -> Water_Temp.:98.49, PWM:35
17:26:12.606 -> Water_Temp.:98.49, PWM:35
17:26:13.403 -> Water_Temp.:98.49, PWM:35
17:26:14.194 -> Water_Temp.:98.37, PWM:35
17:26:14.956 -> Water_Temp.:98.37, PWM:35
17:26:15.749 -> Water_Temp.:98.49, PWM:35
17:26:16.506 -> Water_Temp.:98.37, PWM:35
17:26:17.299 -> Water_Temp.:98.37, PWM:35
17:26:18.091 -> Water_Temp.:98.37, PWM:35
17:26:18.850 -> Water_Temp.:98.37, PWM:35
17:26:19.642 -> Water_Temp.:98.37, PWM:35
17:26:20.405 -> Water_Temp.:98.37, PWM:35
17:26:21.197 -> Water_Temp.:98.37, PWM:35
17:26:21.995 -> Water_Temp.:98.37, PWM:35
17:26:22.758 -> Water_Temp.:98.37, PWM:35
17:26:23.553 -> Water_Temp.:98.37, PWM:35
17:26:24.314 -> Water_Temp.:98.37, PWM:35
17:26:25.838 -> Water_Temp.:98.37, PWM:35
17:26:26.633 -> Water_Temp.:98.37, PWM:35
17:26:27.428 -> Water_Temp.:98.37, PWM:35
17:26:28.188 -> Water_Temp.:98.37, PWM:35
17:26:28.977 -> Water_Temp.:98.49, PWM:35
17:26:29.768 -> Water_Temp.:98.49, PWM:35
17:26:30.529 -> Water_Temp.:98.49, PWM:35
17:26:31.323 -> Water_Temp.:98.49, PWM:35
17:26:32.085 -> Water_Temp.:98.49, PWM:35
17:26:32.881 -> Water_Temp.:98.49, PWM:35
17:26:33.642 -> Water_Temp.:98.49, PWM:35
17:26:34.431 -> Water_Temp.:98.49, PWM:35
17:26:35.223 -> Water_Temp.:98.49, PWM:35
17:26:35.982 -> Water_Temp.:98.49, PWM:35
17:26:37.531 -> Water_Temp.:98.49, PWM:35
17:26:38.324 -> Water_Temp.:98.49, PWM:35
17:26:39.084 -> Water_Temp.:98.49, PWM:35
17:26:39.877 -> Water_Temp.:98.49, PWM:35
17:26:40.637 -> Water_Temp.:98.49, PWM:35
17:26:42.220 -> Water_Temp.:98.49, PWM:35
17:26:42.976 -> Water_Temp.:98.49, PWM:35
17:26:43.771 -> Water_Temp.:98.49, PWM:35
17:26:44.566 -> Water_Temp.:98.49, PWM:35
17:26:45.324 -> Water_Temp.:98.49, PWM:35
17:26:46.115 -> Water_Temp.:98.49, PWM:35
17:26:46.880 -> Water_Temp.:98.49, PWM:35
17:26:47.676 -> Water_Temp.:98.49, PWM:35
17:26:48.439 -> Water_Temp.:98.49, PWM:35
17:26:49.230 -> Water_Temp.:98.49, PWM:35
17:26:50.025 -> Water_Temp.:98.49, PWM:35
17:26:50.787 -> Water_Temp.:98.49, PWM:35
17:26:51.578 -> Water_Temp.:98.49, PWM:35
17:26:52.337 -> Water_Temp.:98.49, PWM:35
17:26:53.130 -> Water_Temp.:98.49, PWM:35
17:26:53.923 -> Water_Temp.:98.49, PWM:35
17:26:54.683 -> Water_Temp.:98.49, PWM:35
17:26:55.476 -> Water_Temp.:98.49, PWM:35
17:26:56.266 -> Water_Temp.:98.49, PWM:35
17:26:57.025 -> Water_Temp.:98.49, PWM:35
17:26:57.819 -> Water_Temp.:98.49, PWM:35
17:26:58.576 -> Water_Temp.:98.49, PWM:35
17:26:59.370 -> Water_Temp.:98.49, PWM:35
17:27:00.160 -> Water_Temp.:98.49, PWM:35
17:27:00.920 -> Water_Temp.:98.49, PWM:35
17:27:01.716 -> Water_Temp.:98.49, PWM:35
17:27:02.474 -> Water_Temp.:98.49, PWM:35
17:27:03.269 -> Water_Temp.:98.49, PWM:35
17:27:04.064 -> Water_Temp.:98.49, PWM:35
17:27:04.826 -> Water_Temp.:98.49, PWM:35
17:27:05.620 -> Water_Temp.:98.49, PWM:35
17:27:06.378 -> Water_Temp.:98.49, PWM:35
17:27:07.176 -> Water_Temp.:98.49, PWM:35
17:27:07.938 -> Water_Temp.:98.49, PWM:35
17:27:08.732 -> Water_Temp.:98.49, PWM:35
17:27:09.525 -> Water_Temp.:98.49, PWM:35
17:27:10.277 -> Water_Temp.:98.49, PWM:35
17:27:11.072 -> Water_Temp.:98.49, PWM:35
17:27:11.832 -> Water_Temp.:98.49, PWM:35
17:27:12.626 -> Water_Temp.:98.49, PWM:35
17:27:13.421 -> Water_Temp.:98.49, PWM:35
17:27:14.181 -> Water_Temp.:98.49, PWM:35
17:27:14.974 -> Water_Temp.:98.49, PWM:35
17:27:15.733 -> Water_Temp.:98.49, PWM:35
17:27:16.523 -> Water_Temp.:98.49, PWM:35
17:27:18.077 -> Water_Temp.:98.49, PWM:35
17:27:18.837 -> Water_Temp.:98.49, PWM:35
17:27:19.633 -> Water_Temp.:98.49, PWM:35
17:27:20.394 -> Water_Temp.:98.49, PWM:35
17:27:21.185 -> Water_Temp.:98.49, PWM:35
17:27:21.979 -> Water_Temp.:98.49, PWM:35
17:27:22.737 -> Water_Temp.:98.49, PWM:35
17:27:23.531 -> Water_Temp.:98.49, PWM:35
17:27:24.327 -> Water_Temp.:98.49, PWM:35
17:27:25.086 -> Water_Temp.:98.49, PWM:35
17:27:25.881 -> Water_Temp.:98.49, PWM:35
17:27:26.640 -> Water_Temp.:98.49, PWM:35
17:27:27.439 -> Water_Temp.:98.49, PWM:35
17:27:28.200 -> Water_Temp.:98.49, PWM:35
17:27:28.996 -> Water_Temp.:98.49, PWM:35
17:27:29.785 -> Water_Temp.:98.49, PWM:35
17:27:30.547 -> Water_Temp.:98.49, PWM:35
17:27:31.341 -> Water_Temp.:98.49, PWM:35
17:27:32.094 -> Water_Temp.:98.49, PWM:35
17:27:32.888 -> Water_Temp.:98.49, PWM:35
17:27:33.684 -> Water_Temp.:98.49, PWM:35
17:27:34.443 -> Water_Temp.:98.49, PWM:35
17:27:35.236 -> Water_Temp.:98.60, PWM:35
17:27:35.997 -> Water_Temp.:98.60, PWM:35
17:27:36.788 -> Water_Temp.:98.60, PWM:35
17:27:37.579 -> Water_Temp.:98.60, PWM:35
17:27:38.339 -> Water_Temp.:100.17, PWM:35
17:27:39.134 -> Water_Temp.:102.20, PWM:35
17:27:39.894 -> Water_Temp.:99.84, PWM:35
17:27:40.682 -> Water_Temp.:99.39, PWM:35
17:27:41.471 -> Water_Temp.:99.27, PWM:35
17:27:42.265 -> Water_Temp.:99.05, PWM:35
17:27:43.028 -> Water_Temp.:99.05, PWM:35
17:27:43.814 -> Water_Temp.:98.94, PWM:35
17:27:44.604 -> Water_Temp.:98.94, PWM:35
17:27:45.359 -> Water_Temp.:98.82, PWM:35
17:27:46.156 -> Water_Temp.:98.82, PWM:35
17:27:46.915 -> Water_Temp.:98.82, PWM:35
17:27:47.709 -> Water_Temp.:98.71, PWM:35
17:27:48.472 -> Water_Temp.:98.71, PWM:35
17:27:49.265 -> Water_Temp.:98.71, PWM:35
17:27:50.060 -> Water_Temp.:98.71, PWM:35
17:27:50.823 -> Water_Temp.:98.71, PWM:35
17:27:51.616 -> Water_Temp.:98.71, PWM:35
17:27:53.138 -> Water_Temp.:98.60, PWM:35
17:27:53.930 -> Water_Temp.:98.60, PWM:35
17:27:56.249 -> Water_Temp.:98.60, PWM:35
17:27:57.042 -> Water_Temp.:98.60, PWM:35
17:27:58.596 -> Water_Temp.:98.60, PWM:39
17:27:59.361 -> Water_Temp.:98.60, PWM:39
17:28:00.154 -> Water_Temp.:98.60, PWM:39
17:28:00.950 -> Water_Temp.:98.60, PWM:39
17:28:01.710 -> Water_Temp.:98.60, PWM:39
17:28:02.503 -> Water_Temp.:98.60, PWM:39
17:28:03.260 -> Water_Temp.:98.60, PWM:37
17:28:04.049 -> Water_Temp.:98.60, PWM:37
17:28:05.609 -> Water_Temp.:98.60, PWM:37
17:28:06.397 -> Water_Temp.:98.60, PWM:37
17:28:07.156 -> Water_Temp.:98.60, PWM:37
17:28:07.955 -> Water_Temp.:98.60, PWM:37
17:28:08.716 -> Water_Temp.:98.60, PWM:37
17:28:09.509 -> Water_Temp.:98.60, PWM:37
17:28:11.059 -> Water_Temp.:98.60, PWM:37
17:28:11.818 -> Water_Temp.:98.60, PWM:37
17:28:12.611 -> Water_Temp.:98.60, PWM:37
17:28:13.404 -> Water_Temp.:98.60, PWM:37
17:28:14.166 -> Water_Temp.:98.60, PWM:37
17:28:14.956 -> Water_Temp.:98.60, PWM:37
17:28:15.747 -> Water_Temp.:98.60, PWM:37
17:28:17.268 -> Water_Temp.:98.60, PWM:37
17:28:18.063 -> Water_Temp.:98.60, PWM:37
17:28:18.817 -> Water_Temp.:98.60, PWM:37
17:28:19.616 -> Water_Temp.:98.60, PWM:37
17:28:20.377 -> Water_Temp.:98.60, PWM:37
17:28:21.166 -> Water_Temp.:98.60, PWM:37
17:28:21.960 -> Water_Temp.:98.71, PWM:37
17:28:22.716 -> Water_Temp.:98.71, PWM:37
17:28:23.510 -> Water_Temp.:98.71, PWM:37
17:28:24.303 -> Water_Temp.:98.71, PWM:37
17:28:25.062 -> Water_Temp.:98.71, PWM:37
17:28:25.857 -> Water_Temp.:98.71, PWM:37
17:28:26.618 -> Water_Temp.:98.71, PWM:37
17:28:27.417 -> Water_Temp.:98.71, PWM:37
17:28:28.177 -> Water_Temp.:98.60, PWM:37
17:28:28.971 -> Water_Temp.:98.60, PWM:37
17:28:29.766 -> Water_Temp.:98.60, PWM:37
17:28:30.526 -> Water_Temp.:98.60, PWM:37
17:28:31.313 -> Water_Temp.:98.60, PWM:37
17:28:32.075 -> Water_Temp.:98.60, PWM:37
17:28:32.873 -> Water_Temp.:98.60, PWM:37
17:28:33.666 -> Water_Temp.:98.71, PWM:37
17:28:34.422 -> Water_Temp.:98.71, PWM:37
17:28:35.209 -> Water_Temp.:98.60, PWM:37
17:28:36.003 -> Water_Temp.:98.71, PWM:37
17:28:36.759 -> Water_Temp.:98.71, PWM:37
17:28:37.550 -> Water_Temp.:98.71, PWM:37
17:28:38.343 -> Water_Temp.:98.71, PWM:37
17:28:39.106 -> Water_Temp.:98.71, PWM:37
17:28:39.891 -> Water_Temp.:98.71, PWM:37
17:28:41.445 -> Water_Temp.:98.71, PWM:37
17:28:42.207 -> Water_Temp.:98.71, PWM:37
17:28:43.003 -> Water_Temp.:98.82, PWM:37
17:28:43.797 -> Water_Temp.:98.82, PWM:37
17:28:44.558 -> Water_Temp.:98.82, PWM:37
17:28:45.354 -> Water_Temp.:98.82, PWM:37
17:28:46.110 -> Water_Temp.:98.82, PWM:37
17:28:46.903 -> Water_Temp.:98.82, PWM:38
17:28:47.663 -> Water_Temp.:98.94, PWM:37
17:28:48.459 -> Water_Temp.:98.94, PWM:37
17:28:49.246 -> Water_Temp.:98.94, PWM:37
17:28:50.006 -> Water_Temp.:98.94, PWM:37
17:28:50.797 -> Water_Temp.:98.94, PWM:37
17:28:51.588 -> Water_Temp.:98.94, PWM:37
17:28:52.352 -> Water_Temp.:98.94, PWM:37
17:28:53.147 -> Water_Temp.:98.94, PWM:37
17:28:53.906 -> Water_Temp.:98.94, PWM:37
17:28:54.698 -> Water_Temp.:98.94, PWM:37
17:28:55.493 -> Water_Temp.:98.94, PWM:37
17:28:56.252 -> Water_Temp.:98.94, PWM:37
17:28:57.044 -> Water_Temp.:98.94, PWM:37
17:28:57.803 -> Water_Temp.:98.94, PWM:37
17:28:58.596 -> Water_Temp.:98.82, PWM:37
17:28:59.388 -> Water_Temp.:98.82, PWM:37
17:29:00.177 -> Water_Temp.:98.82, PWM:37
17:29:00.938 -> Water_Temp.:98.82, PWM:37
17:29:01.733 -> Water_Temp.:98.82, PWM:37
17:29:02.493 -> Water_Temp.:98.82, PWM:37
17:29:03.283 -> Water_Temp.:98.82, PWM:37
17:29:04.074 -> Water_Temp.:98.71, PWM:37
17:29:05.592 -> Water_Temp.:98.82, PWM:37
17:29:06.388 -> Water_Temp.:98.82, PWM:37
17:29:07.179 -> Water_Temp.:98.82, PWM:37
17:29:07.937 -> Water_Temp.:98.82, PWM:37
17:29:09.485 -> Water_Temp.:98.82, PWM:37
17:29:10.276 -> Water_Temp.:98.82, PWM:37
17:29:11.037 -> Water_Temp.:98.82, PWM:37
17:29:11.822 -> Water_Temp.:98.82, PWM:37
17:29:12.618 -> Water_Temp.:98.82, PWM:37
17:29:13.378 -> Water_Temp.:98.82, PWM:37
17:29:14.170 -> Water_Temp.:100.40, PWM:37
17:29:14.963 -> Water_Temp.:102.42, PWM:37
17:29:15.721 -> Water_Temp.:100.06, PWM:37
17:29:16.510 -> Water_Temp.:99.61, PWM:37
17:29:17.300 -> Water_Temp.:99.39, PWM:37
17:29:18.058 -> Water_Temp.:99.27, PWM:37
17:29:18.850 -> Water_Temp.:99.16, PWM:37
17:29:19.640 -> Water_Temp.:99.05, PWM:37
17:29:20.400 -> Water_Temp.:99.05, PWM:37
17:29:21.193 -> Water_Temp.:98.94, PWM:37
17:29:21.985 -> Water_Temp.:98.94, PWM:37
17:29:22.747 -> Water_Temp.:98.94, PWM:37
17:29:23.538 -> Water_Temp.:98.94, PWM:37
17:29:24.299 -> Water_Temp.:98.82, PWM:37
17:29:25.095 -> Water_Temp.:98.82, PWM:37
17:29:25.855 -> Water_Temp.:98.82, PWM:37
17:29:26.652 -> Water_Temp.:98.82, PWM:37
17:29:28.207 -> Water_Temp.:98.82, PWM:37
17:29:28.992 -> Water_Temp.:98.82, PWM:37
17:29:30.541 -> Water_Temp.:98.82, PWM:37
17:29:31.302 -> Water_Temp.:98.82, PWM:37
17:29:32.093 -> Water_Temp.:98.71, PWM:37
17:29:32.886 -> Water_Temp.:98.82, PWM:37
17:29:33.648 -> Water_Temp.:98.71, PWM:37
17:29:34.437 -> Water_Temp.:98.71, PWM:37
17:29:35.197 -> Water_Temp.:98.71, PWM:37
17:29:35.994 -> Water_Temp.:98.71, PWM:37
17:29:36.787 -> Water_Temp.:98.71, PWM:37
17:29:38.342 -> Water_Temp.:98.71, PWM:37
17:29:39.100 -> Water_Temp.:98.71, PWM:37
17:29:39.897 -> Water_Temp.:98.71, PWM:37
17:29:40.686 -> Water_Temp.:98.71, PWM:37
17:29:41.449 -> Water_Temp.:98.71, PWM:37
17:29:42.241 -> Water_Temp.:98.71, PWM:37
17:29:43.000 -> Water_Temp.:98.71, PWM:37
17:29:43.788 -> Water_Temp.:98.71, PWM:37
17:29:44.577 -> Water_Temp.:98.71, PWM:37
17:29:45.368 -> Water_Temp.:98.71, PWM:37
17:29:46.887 -> Water_Temp.:98.71, PWM:37
17:29:47.679 -> Water_Temp.:98.71, PWM:37
17:29:48.437 -> Water_Temp.:98.82, PWM:37
17:29:49.234 -> Water_Temp.:98.82, PWM:37
17:29:49.996 -> Water_Temp.:98.82, PWM:37
17:29:50.788 -> Water_Temp.:98.82, PWM:37
17:29:51.577 -> Water_Temp.:98.82, PWM:37
17:29:52.339 -> Water_Temp.:98.82, PWM:37
17:29:53.131 -> Water_Temp.:98.82, PWM:37
17:29:53.892 -> Water_Temp.:98.82, PWM:37
17:29:54.680 -> Water_Temp.:98.82, PWM:37
17:29:55.473 -> Water_Temp.:98.82, PWM:37
17:29:56.233 -> Water_Temp.:98.82, PWM:37
17:29:57.023 -> Water_Temp.:98.82, PWM:37
17:29:57.816 -> Water_Temp.:98.82, PWM:37
17:29:58.573 -> Water_Temp.:98.82, PWM:37
17:29:59.364 -> Water_Temp.:98.82, PWM:37
17:30:00.161 -> Water_Temp.:98.82, PWM:37
17:30:00.918 -> Water_Temp.:98.82, PWM:37
17:30:01.709 -> Water_Temp.:98.82, PWM:37
17:30:02.501 -> Water_Temp.:98.82, PWM:37
17:30:03.261 -> Water_Temp.:98.82, PWM:37
17:30:04.053 -> Water_Temp.:98.82, PWM:37
17:30:04.811 -> Water_Temp.:98.82, PWM:37
17:30:05.608 -> Water_Temp.:98.94, PWM:37
17:30:06.368 -> Water_Temp.:98.82, PWM:37
17:30:07.161 -> Water_Temp.:98.82, PWM:37
17:30:07.955 -> Water_Temp.:98.82, PWM:37
17:30:08.716 -> Water_Temp.:98.82, PWM:37
17:30:09.508 -> Water_Temp.:98.82, PWM:37
17:30:10.303 -> Water_Temp.:98.82, PWM:37
17:30:11.062 -> Water_Temp.:98.82, PWM:37
17:30:11.859 -> Water_Temp.:98.82, PWM:37
17:30:12.618 -> Water_Temp.:98.82, PWM:37
17:30:13.413 -> Water_Temp.:98.82, PWM:37
17:30:14.172 -> Water_Temp.:98.82, PWM:37
17:30:14.967 -> Water_Temp.:98.82, PWM:37
17:30:15.764 -> Water_Temp.:98.82, PWM:37
17:30:16.519 -> Water_Temp.:98.82, PWM:37
17:30:17.312 -> Water_Temp.:98.82, PWM:37
17:30:18.070 -> Water_Temp.:98.82, PWM:37
17:30:18.863 -> Water_Temp.:98.82, PWM:37
17:30:19.655 -> Water_Temp.:98.82, PWM:37
17:30:20.411 -> Water_Temp.:98.82, PWM:37
17:30:21.206 -> Water_Temp.:98.82, PWM:37
17:30:21.997 -> Water_Temp.:98.82, PWM:37
17:30:22.760 -> Water_Temp.:98.82, PWM:37
17:30:23.552 -> Water_Temp.:98.82, PWM:37
17:30:24.308 -> Water_Temp.:98.82, PWM:37
17:30:25.864 -> Water_Temp.:98.82, PWM:37
17:30:27.422 -> Water_Temp.:98.82, PWM:37
17:30:28.180 -> Water_Temp.:98.82, PWM:37
17:30:28.974 -> Water_Temp.:98.82, PWM:37
17:30:29.738 -> Water_Temp.:98.82, PWM:37
17:30:30.531 -> Water_Temp.:98.82, PWM:37
17:30:31.325 -> Water_Temp.:98.82, PWM:37
17:30:32.085 -> Water_Temp.:98.82, PWM:37
17:30:32.877 -> Water_Temp.:98.82, PWM:37
17:30:33.639 -> Water_Temp.:98.82, PWM:37
17:30:34.432 -> Water_Temp.:98.82, PWM:37
17:30:35.223 -> Water_Temp.:98.82, PWM:37
17:30:35.977 -> Water_Temp.:98.82, PWM:37
17:30:36.773 -> Water_Temp.:98.82, PWM:37
17:30:37.567 -> Water_Temp.:98.82, PWM:37
17:30:38.327 -> Water_Temp.:98.82, PWM:37
17:30:39.880 -> Water_Temp.:98.82, PWM:37
17:30:40.672 -> Water_Temp.:98.82, PWM:37
17:30:41.430 -> Water_Temp.:98.82, PWM:37
17:30:42.224 -> Water_Temp.:98.82, PWM:37
17:30:43.012 -> Water_Temp.:98.82, PWM:37
17:30:43.768 -> Water_Temp.:98.82, PWM:37
17:30:44.563 -> Water_Temp.:98.82, PWM:37
17:30:45.323 -> Water_Temp.:98.82, PWM:37
17:30:46.114 -> Water_Temp.:98.82, PWM:37
17:30:46.906 -> Water_Temp.:98.82, PWM:37
17:30:47.666 -> Water_Temp.:98.82, PWM:37
17:30:48.460 -> Water_Temp.:98.82, PWM:37
17:30:49.223 -> Water_Temp.:98.82, PWM:37
17:30:50.018 -> Water_Temp.:98.82, PWM:37
17:30:50.813 -> Water_Temp.:98.82, PWM:37
17:30:51.569 -> Water_Temp.:98.82, PWM:37
17:30:52.364 -> Water_Temp.:98.82, PWM:37
17:30:53.121 -> Water_Temp.:98.82, PWM:37
17:30:53.910 -> Water_Temp.:98.82, PWM:37
17:30:55.460 -> Water_Temp.:98.82, PWM:37
17:30:56.257 -> Water_Temp.:98.82, PWM:37
17:30:57.051 -> Water_Temp.:98.82, PWM:37
17:30:57.802 -> Water_Temp.:98.94, PWM:37
17:30:58.596 -> Water_Temp.:98.94, PWM:37
17:30:59.391 -> Water_Temp.:98.94, PWM:37
17:31:00.149 -> Water_Temp.:98.94, PWM:37
17:31:00.942 -> Water_Temp.:98.94, PWM:37
17:31:01.700 -> Water_Temp.:98.94, PWM:37
17:31:02.498 -> Water_Temp.:98.94, PWM:37
17:31:03.287 -> Water_Temp.:98.94, PWM:37
17:31:04.046 -> Water_Temp.:98.94, PWM:37
17:31:04.842 -> Water_Temp.:98.94, PWM:37
17:31:05.602 -> Water_Temp.:98.94, PWM:37
17:31:06.399 -> Water_Temp.:98.94, PWM:37
17:31:07.162 -> Water_Temp.:98.94, PWM:37
17:31:07.953 -> Water_Temp.:98.94, PWM:37
17:31:08.743 -> Water_Temp.:98.94, PWM:37
17:31:09.500 -> Water_Temp.:98.94, PWM:37
17:31:10.296 -> Water_Temp.:98.94, PWM:37
17:31:11.091 -> Water_Temp.:98.94, PWM:37
17:31:11.851 -> Water_Temp.:98.94, PWM:37
17:31:12.646 -> Water_Temp.:98.94, PWM:37
17:31:13.406 -> Water_Temp.:98.94, PWM:37
17:31:14.203 -> Water_Temp.:98.94, PWM:37
17:31:14.963 -> Water_Temp.:98.94, PWM:37
17:31:15.752 -> Water_Temp.:98.94, PWM:37
17:31:16.545 -> Water_Temp.:98.94, PWM:37
17:31:17.306 -> Water_Temp.:98.94, PWM:37
17:31:18.097 -> Water_Temp.:98.94, PWM:37
17:31:18.861 -> Water_Temp.:98.94, PWM:37
17:31:19.661 -> Water_Temp.:98.94, PWM:37
17:31:20.451 -> Water_Temp.:98.94, PWM:37
17:31:21.214 -> Water_Temp.:98.94, PWM:37
17:31:22.005 -> Water_Temp.:98.94, PWM:37
17:31:22.766 -> Water_Temp.:98.94, PWM:37
17:31:23.556 -> Water_Temp.:98.94, PWM:37
17:31:25.105 -> Water_Temp.:98.94, PWM:37
17:31:25.864 -> Water_Temp.:98.94, PWM:37
17:31:26.656 -> Water_Temp.:98.94, PWM:37
17:31:27.414 -> Water_Temp.:98.94, PWM:37
17:31:28.206 -> Water_Temp.:98.94, PWM:37
17:31:29.005 -> Water_Temp.:98.94, PWM:37
17:31:29.770 -> Water_Temp.:98.94, PWM:37
17:31:30.562 -> Water_Temp.:98.94, PWM:37
17:31:31.323 -> Water_Temp.:98.94, PWM:37
17:31:32.116 -> Water_Temp.:98.94, PWM:37
17:31:32.904 -> Water_Temp.:98.94, PWM:37
17:31:33.658 -> Water_Temp.:98.94, PWM:37
17:31:34.454 -> Water_Temp.:98.94, PWM:37
17:31:35.214 -> Water_Temp.:98.94, PWM:37
17:31:36.008 -> Water_Temp.:98.94, PWM:37
17:31:36.800 -> Water_Temp.:98.94, PWM:37
17:31:38.315 -> Water_Temp.:99.05, PWM:37
17:31:39.110 -> Water_Temp.:99.05, PWM:37
17:31:39.870 -> Water_Temp.:99.05, PWM:37
17:31:40.666 -> Water_Temp.:99.05, PWM:37
17:31:41.457 -> Water_Temp.:99.05, PWM:37
17:31:42.220 -> Water_Temp.:99.05, PWM:37
17:31:43.013 -> Water_Temp.:99.05, PWM:37
17:31:43.771 -> Water_Temp.:99.05, PWM:37
17:31:44.567 -> Water_Temp.:99.05, PWM:37
17:31:45.326 -> Water_Temp.:99.05, PWM:37
17:31:46.121 -> Water_Temp.:99.05, PWM:37
17:31:46.917 -> Water_Temp.:99.05, PWM:37
17:31:47.673 -> Water_Temp.:99.05, PWM:37
17:31:49.228 -> Water_Temp.:99.05, PWM:37
17:31:50.021 -> Water_Temp.:99.05, PWM:37
17:31:50.812 -> Water_Temp.:99.05, PWM:37
17:31:51.576 -> Water_Temp.:99.05, PWM:37
17:31:53.127 -> Water_Temp.:99.05, PWM:37
17:31:53.921 -> Water_Temp.:99.05, PWM:37
17:31:55.481 -> Water_Temp.:99.05, PWM:37
17:31:56.245 -> Water_Temp.:99.05, PWM:37
17:31:57.042 -> Water_Temp.:99.05, PWM:37
17:31:57.835 -> Water_Temp.:99.05, PWM:37
17:31:58.595 -> Water_Temp.:98.94, PWM:37
17:31:59.391 -> Water_Temp.:99.05, PWM:37
17:32:00.150 -> Water_Temp.:98.94, PWM:37
17:32:00.942 -> Water_Temp.:98.94, PWM:37
17:32:01.734 -> Water_Temp.:98.94, PWM:37
17:32:02.496 -> Water_Temp.:98.94, PWM:37
17:32:03.290 -> Water_Temp.:98.94, PWM:37
17:32:04.838 -> Water_Temp.:98.94, PWM:37
17:32:05.602 -> Water_Temp.:98.94, PWM:37
17:32:06.389 -> Water_Temp.:98.94, PWM:37
17:32:07.147 -> Water_Temp.:98.94, PWM:37
17:32:07.943 -> Water_Temp.:98.94, PWM:37
17:32:08.738 -> Water_Temp.:98.94, PWM:37
17:32:09.498 -> Water_Temp.:98.94, PWM:37
17:32:10.292 -> Water_Temp.:98.94, PWM:37
17:32:11.053 -> Water_Temp.:98.94, PWM:37
17:32:11.845 -> Water_Temp.:98.94, PWM:37
17:32:12.604 -> Water_Temp.:98.94, PWM:37
17:32:13.398 -> Water_Temp.:98.94, PWM:37
17:32:14.193 -> Water_Temp.:98.94, PWM:37
17:32:14.957 -> Water_Temp.:98.94, PWM:37
17:32:15.749 -> Water_Temp.:98.94, PWM:37
17:32:17.274 -> Water_Temp.:98.94, PWM:37
17:32:18.070 -> Water_Temp.:98.94, PWM:37
17:32:18.830 -> Water_Temp.:98.94, PWM:37
17:32:20.351 -> Water_Temp.:98.94, PWM:37
17:32:21.141 -> Water_Temp.:98.94, PWM:37
17:32:21.934 -> Water_Temp.:98.94, PWM:37
17:32:22.692 -> Water_Temp.:98.94, PWM:37
17:32:23.485 -> Water_Temp.:98.94, PWM:37
17:32:24.277 -> Water_Temp.:98.94, PWM:37
17:32:25.040 -> Water_Temp.:99.05, PWM:37
17:32:25.825 -> Water_Temp.:99.05, PWM:37
17:32:26.620 -> Water_Temp.:99.05, PWM:37
17:32:27.379 -> Water_Temp.:99.05, PWM:37
17:32:28.174 -> Water_Temp.:99.05, PWM:37
17:32:28.929 -> Water_Temp.:99.05, PWM:37
17:32:29.726 -> Water_Temp.:99.05, PWM:37
17:32:30.515 -> Water_Temp.:99.05, PWM:37
17:32:31.277 -> Water_Temp.:99.05, PWM:37
17:32:32.067 -> Water_Temp.:99.05, PWM:37
17:32:32.862 -> Water_Temp.:99.05, PWM:37
17:32:33.623 -> Water_Temp.:99.05, PWM:37
17:32:34.421 -> Water_Temp.:99.05, PWM:37
17:32:35.184 -> Water_Temp.:99.16, PWM:37
17:32:35.978 -> Water_Temp.:99.16, PWM:37
17:32:36.742 -> Water_Temp.:99.05, PWM:37
17:32:37.536 -> Water_Temp.:99.16, PWM:37
17:32:38.294 -> Water_Temp.:99.16, PWM:37
17:32:39.087 -> Water_Temp.:99.16, PWM:37
17:32:39.880 -> Water_Temp.:99.16, PWM:37
17:32:40.638 -> Water_Temp.:99.16, PWM:37
17:32:41.431 -> Water_Temp.:99.16, PWM:37
17:32:42.188 -> Water_Temp.:99.16, PWM:37
17:32:42.982 -> Water_Temp.:99.16, PWM:37
17:32:43.773 -> Water_Temp.:99.16, PWM:37
17:32:44.563 -> Water_Temp.:99.16, PWM:37
17:32:45.319 -> Water_Temp.:99.16, PWM:37
17:32:46.110 -> Water_Temp.:99.16, PWM:37
17:32:46.869 -> Water_Temp.:99.16, PWM:37
17:32:48.454 -> Water_Temp.:99.16, PWM:37
17:32:50.767 -> Water_Temp.:99.16, PWM:37
17:32:51.560 -> Water_Temp.:99.05, PWM:37
17:32:52.355 -> Water_Temp.:99.05, PWM:37
17:32:53.117 -> Water_Temp.:99.05, PWM:37
17:32:53.913 -> Water_Temp.:99.16, PWM:37
17:32:54.667 -> Water_Temp.:99.05, PWM:37
17:32:55.458 -> Water_Temp.:99.16, PWM:37
17:32:56.251 -> Water_Temp.:99.16, PWM:37
17:32:57.009 -> Water_Temp.:99.05, PWM:37
17:32:57.803 -> Water_Temp.:99.16, PWM:37
17:32:59.352 -> Water_Temp.:99.16, PWM:41
17:33:00.112 -> Water_Temp.:99.16, PWM:43
17:33:00.905 -> Water_Temp.:99.16, PWM:44
17:33:01.696 -> Water_Temp.:100.74, PWM:42
17:33:02.458 -> Water_Temp.:102.76, PWM:43
17:33:03.252 -> Water_Temp.:100.40, PWM:43
17:33:04.011 -> Water_Temp.:99.95, PWM:42
17:33:04.803 -> Water_Temp.:99.84, PWM:42
17:33:05.598 -> Water_Temp.:99.72, PWM:41
17:33:06.358 -> Water_Temp.:99.61, PWM:41
17:33:07.913 -> Water_Temp.:99.50, PWM:41
17:33:08.706 -> Water_Temp.:99.39, PWM:41
17:33:09.463 -> Water_Temp.:99.39, PWM:41
17:33:10.257 -> Water_Temp.:99.39, PWM:41
17:33:11.017 -> Water_Temp.:99.39, PWM:41
17:33:11.809 -> Water_Temp.:99.39, PWM:41
17:33:12.606 -> Water_Temp.:99.27, PWM:41
17:33:13.363 -> Water_Temp.:99.27, PWM:41
17:33:14.153 -> Water_Temp.:99.27, PWM:41
17:33:15.702 -> Water_Temp.:99.27, PWM:41
17:33:16.493 -> Water_Temp.:99.27, PWM:41
17:33:17.284 -> Water_Temp.:99.27, PWM:41
17:33:18.045 -> Water_Temp.:99.27, PWM:41
17:33:18.841 -> Water_Temp.:99.27, PWM:41
17:33:19.603 -> Water_Temp.:99.27, PWM:41
17:33:20.393 -> Water_Temp.:99.27, PWM:41
17:33:21.946 -> Water_Temp.:99.27, PWM:41
17:33:22.706 -> Water_Temp.:99.27, PWM:41
17:33:23.498 -> Water_Temp.:99.27, PWM:41
17:33:24.288 -> Water_Temp.:99.27, PWM:41
17:33:25.051 -> Water_Temp.:99.27, PWM:41
17:33:25.846 -> Water_Temp.:99.27, PWM:41
17:33:26.603 -> Water_Temp.:99.27, PWM:41
17:33:27.398 -> Water_Temp.:99.27, PWM:41
17:33:28.154 -> Water_Temp.:99.27, PWM:41
17:33:28.944 -> Water_Temp.:99.27, PWM:41
17:33:29.736 -> Water_Temp.:99.27, PWM:41
17:33:30.531 -> Water_Temp.:99.27, PWM:41
17:33:31.287 -> Water_Temp.:99.27, PWM:41
17:33:32.076 -> Water_Temp.:99.27, PWM:41
17:33:32.836 -> Water_Temp.:99.27, PWM:41
17:33:33.633 -> Water_Temp.:99.27, PWM:41
17:33:34.429 -> Water_Temp.:99.27, PWM:41
17:33:35.191 -> Water_Temp.:99.27, PWM:41
17:33:36.744 -> Water_Temp.:99.27, PWM:41
17:33:37.537 -> Water_Temp.:100.85, PWM:41
17:33:38.299 -> Water_Temp.:102.87, PWM:41
17:33:39.085 -> Water_Temp.:100.51, PWM:41
17:33:39.878 -> Water_Temp.:100.17, PWM:41
17:33:40.641 -> Water_Temp.:99.95, PWM:41
17:33:41.434 -> Water_Temp.:99.84, PWM:41
17:33:42.197 -> Water_Temp.:99.72, PWM:41
17:33:42.990 -> Water_Temp.:99.61, PWM:41
17:33:43.783 -> Water_Temp.:99.61, PWM:41
17:33:44.541 -> Water_Temp.:99.50, PWM:41
17:33:45.337 -> Water_Temp.:99.50, PWM:41
17:33:46.127 -> Water_Temp.:99.50, PWM:41
17:33:46.885 -> Water_Temp.:99.50, PWM:41
17:33:47.677 -> Water_Temp.:99.39, PWM:41
17:33:48.466 -> Water_Temp.:99.39, PWM:41
17:33:49.223 -> Water_Temp.:99.39, PWM:41
17:33:50.013 -> Water_Temp.:99.39, PWM:41
17:33:50.772 -> Water_Temp.:99.39, PWM:41
17:33:51.569 -> Water_Temp.:99.39, PWM:41
17:33:52.363 -> Water_Temp.:99.39, PWM:41
17:33:53.123 -> Water_Temp.:99.27, PWM:41
17:33:53.915 -> Water_Temp.:99.27, PWM:41
17:33:54.706 -> Water_Temp.:99.27, PWM:41
17:33:55.469 -> Water_Temp.:99.27, PWM:41
17:33:56.263 -> Water_Temp.:99.27, PWM:41
17:33:57.019 -> Water_Temp.:99.39, PWM:41
17:33:57.812 -> Water_Temp.:99.39, PWM:41
17:33:58.604 -> Water_Temp.:99.39, PWM:41
17:33:59.367 -> Water_Temp.:99.39, PWM:41
17:34:00.157 -> Water_Temp.:99.39, PWM:41
17:34:00.917 -> Water_Temp.:99.39, PWM:41
17:34:01.707 -> Water_Temp.:99.39, PWM:41
17:34:02.498 -> Water_Temp.:99.39, PWM:41
17:34:03.258 -> Water_Temp.:99.39, PWM:41
17:34:04.051 -> Water_Temp.:99.39, PWM:41
17:34:04.814 -> Water_Temp.:99.39, PWM:41
17:34:05.606 -> Water_Temp.:99.39, PWM:41
17:34:06.397 -> Water_Temp.:99.39, PWM:41
17:34:07.157 -> Water_Temp.:99.39, PWM:41
17:34:07.946 -> Water_Temp.:99.39, PWM:41
17:34:08.737 -> Water_Temp.:99.39, PWM:41
17:34:09.494 -> Water_Temp.:99.39, PWM:41
17:34:10.289 -> Water_Temp.:99.39, PWM:41
17:34:11.082 -> Water_Temp.:99.39, PWM:41
17:34:11.841 -> Water_Temp.:99.39, PWM:41
17:34:12.631 -> Water_Temp.:99.39, PWM:41
17:34:13.422 -> Water_Temp.:99.39, PWM:41
17:34:14.181 -> Water_Temp.:99.39, PWM:41
17:34:14.976 -> Water_Temp.:99.39, PWM:41
17:34:15.737 -> Water_Temp.:99.39, PWM:41
17:34:16.536 -> Water_Temp.:99.39, PWM:41
17:34:17.296 -> Water_Temp.:99.39, PWM:41
17:34:18.087 -> Water_Temp.:99.39, PWM:41
17:34:18.882 -> Water_Temp.:99.39, PWM:41
17:34:19.640 -> Water_Temp.:99.39, PWM:41
17:34:20.438 -> Water_Temp.:99.39, PWM:41
17:34:21.200 -> Water_Temp.:99.39, PWM:41
17:34:21.993 -> Water_Temp.:99.39, PWM:41
17:34:22.754 -> Water_Temp.:99.39, PWM:41
17:34:23.547 -> Water_Temp.:99.39, PWM:41
17:34:24.337 -> Water_Temp.:99.39, PWM:41
17:34:25.092 -> Water_Temp.:99.39, PWM:41
17:34:25.886 -> Water_Temp.:99.39, PWM:41
17:34:26.651 -> Water_Temp.:99.39, PWM:41
17:34:27.444 -> Water_Temp.:99.39, PWM:41
17:34:28.241 -> Water_Temp.:99.50, PWM:41
17:34:28.999 -> Water_Temp.:99.50, PWM:41
17:34:29.795 -> Water_Temp.:99.50, PWM:41
17:34:30.583 -> Water_Temp.:99.50, PWM:41
17:34:31.342 -> Water_Temp.:99.50, PWM:41
17:34:32.137 -> Water_Temp.:99.50, PWM:41
17:34:32.894 -> Water_Temp.:99.50, PWM:41
17:34:33.688 -> Water_Temp.:99.50, PWM:41
17:34:34.477 -> Water_Temp.:99.50, PWM:41
17:34:35.234 -> Water_Temp.:99.50, PWM:41
17:34:36.023 -> Water_Temp.:99.50, PWM:41
17:34:36.815 -> Water_Temp.:99.61, PWM:41
17:34:37.576 -> Water_Temp.:99.61, PWM:41
17:34:38.374 -> Water_Temp.:99.61, PWM:41
17:34:39.134 -> Water_Temp.:99.61, PWM:41
17:34:39.930 -> Water_Temp.:99.61, PWM:41
17:34:40.723 -> Water_Temp.:99.61, PWM:41
17:34:41.484 -> Water_Temp.:99.61, PWM:41
17:34:42.247 -> Water_Temp.:99.61, PWM:41
17:34:43.044 -> Water_Temp.:99.61, PWM:41
17:34:43.841 -> Water_Temp.:99.61, PWM:41
17:34:44.602 -> Water_Temp.:99.61, PWM:41
17:34:45.396 -> Water_Temp.:99.61, PWM:41
17:34:46.154 -> Water_Temp.:99.61, PWM:41
17:34:46.945 -> Water_Temp.:99.61, PWM:41
17:34:47.740 -> Water_Temp.:99.61, PWM:41
17:34:48.498 -> Water_Temp.:99.61, PWM:41
17:34:49.288 -> Water_Temp.:99.61, PWM:41
17:34:50.078 -> Water_Temp.:99.61, PWM:41
17:34:50.838 -> Water_Temp.:99.61, PWM:41
17:34:51.632 -> Water_Temp.:99.61, PWM:41
17:34:52.390 -> Water_Temp.:99.61, PWM:41
17:34:53.186 -> Water_Temp.:99.61, PWM:41
17:34:53.977 -> Water_Temp.:99.61, PWM:41
17:34:54.732 -> Water_Temp.:99.61, PWM:41
17:34:55.524 -> Water_Temp.:99.61, PWM:41
17:34:56.318 -> Water_Temp.:99.72, PWM:41
17:34:57.867 -> Water_Temp.:99.72, PWM:41
17:34:58.627 -> Water_Temp.:99.72, PWM:41
17:34:59.425 -> Water_Temp.:99.72, PWM:41
17:35:00.185 -> Water_Temp.:99.72, PWM:41
17:35:00.975 -> Water_Temp.:99.72, PWM:41
17:35:01.736 -> Water_Temp.:99.72, PWM:41
17:35:02.526 -> Water_Temp.:99.72, PWM:41
17:35:03.315 -> Water_Temp.:99.72, PWM:41
17:35:04.073 -> Water_Temp.:99.72, PWM:41
17:35:04.868 -> Water_Temp.:99.72, PWM:41
17:35:05.662 -> Water_Temp.:99.72, PWM:41
17:35:06.421 -> Water_Temp.:99.72, PWM:41
17:35:07.214 -> Water_Temp.:99.72, PWM:41
17:35:07.973 -> Water_Temp.:99.72, PWM:41
17:35:08.762 -> Water_Temp.:99.72, PWM:41
17:35:09.557 -> Water_Temp.:99.72, PWM:41
17:35:10.319 -> Water_Temp.:99.72, PWM:41
17:35:11.109 -> Water_Temp.:99.72, PWM:41
17:35:11.905 -> Water_Temp.:99.72, PWM:41
17:35:12.667 -> Water_Temp.:99.72, PWM:41
17:35:13.458 -> Water_Temp.:99.61, PWM:41
17:35:14.218 -> Water_Temp.:99.61, PWM:41
17:35:15.798 -> Water_Temp.:99.61, PWM:41
17:35:16.556 -> Water_Temp.:99.61, PWM:41
17:35:17.349 -> Water_Temp.:99.61, PWM:41
17:35:18.902 -> Water_Temp.:99.61, PWM:41
17:35:19.696 -> Water_Temp.:99.61, PWM:41
17:35:20.452 -> Water_Temp.:99.61, PWM:41
17:35:21.247 -> Water_Temp.:99.61, PWM:41
17:35:22.042 -> Water_Temp.:99.61, PWM:41
17:35:22.800 -> Water_Temp.:99.61, PWM:41
17:35:23.593 -> Water_Temp.:99.50, PWM:41
17:35:24.352 -> Water_Temp.:99.50, PWM:41
17:35:25.146 -> Water_Temp.:99.50, PWM:41
17:35:25.943 -> Water_Temp.:99.50, PWM:41
17:35:26.699 -> Water_Temp.:99.50, PWM:41
17:35:27.490 -> Water_Temp.:99.50, PWM:41
17:35:28.252 -> Water_Temp.:99.50, PWM:41
17:35:29.046 -> Water_Temp.:99.50, PWM:41
17:35:29.842 -> Water_Temp.:99.50, PWM:41
17:35:30.597 -> Water_Temp.:99.50, PWM:41
17:35:31.390 -> Water_Temp.:99.50, PWM:41
17:35:32.183 -> Water_Temp.:99.50, PWM:41
17:35:32.942 -> Water_Temp.:99.50, PWM:41
17:35:33.737 -> Water_Temp.:99.50, PWM:41
17:35:34.494 -> Water_Temp.:99.50, PWM:41
17:35:35.288 -> Water_Temp.:99.50, PWM:41
17:35:36.081 -> Water_Temp.:99.50, PWM:41
17:35:36.839 -> Water_Temp.:99.50, PWM:41
17:35:37.628 -> Water_Temp.:99.50, PWM:41
17:35:38.389 -> Water_Temp.:99.50, PWM:41
17:35:39.181 -> Water_Temp.:99.50, PWM:41
17:35:39.976 -> Water_Temp.:99.50, PWM:41
17:35:40.734 -> Water_Temp.:99.50, PWM:41
17:35:42.320 -> Water_Temp.:99.50, PWM:41
17:35:43.078 -> Water_Temp.:99.50, PWM:41
17:35:43.872 -> Water_Temp.:99.50, PWM:41
17:35:44.634 -> Water_Temp.:99.50, PWM:41
17:35:45.428 -> Water_Temp.:99.50, PWM:41
17:35:46.217 -> Water_Temp.:99.50, PWM:41
17:35:46.979 -> Water_Temp.:99.50, PWM:41
17:35:47.768 -> Water_Temp.:99.50, PWM:41
17:35:48.530 -> Water_Temp.:99.50, PWM:41
17:35:49.319 -> Water_Temp.:99.61, PWM:41
17:35:50.115 -> Water_Temp.:99.61, PWM:41
17:35:50.874 -> Water_Temp.:99.50, PWM:41
17:35:51.665 -> Water_Temp.:99.50, PWM:41
17:35:52.456 -> Water_Temp.:99.61, PWM:41
17:35:53.214 -> Water_Temp.:99.50, PWM:41
17:35:54.006 -> Water_Temp.:99.50, PWM:41
17:35:54.768 -> Water_Temp.:99.61, PWM:41
17:35:55.559 -> Water_Temp.:99.50, PWM:41
17:35:56.356 -> Water_Temp.:99.50, PWM:41
17:35:57.116 -> Water_Temp.:99.50, PWM:41
17:35:57.911 -> Water_Temp.:99.61, PWM:41
17:35:58.670 -> Water_Temp.:99.61, PWM:41
17:36:00.986 -> Water_Temp.:99.61, PWM:41
17:36:01.743 -> Water_Temp.:99.61, PWM:41
17:36:02.536 -> Water_Temp.:99.61, PWM:41
17:36:03.325 -> Water_Temp.:99.61, PWM:41
17:36:04.082 -> Water_Temp.:99.61, PWM:41
17:36:04.876 -> Water_Temp.:99.61, PWM:41
17:36:05.638 -> Water_Temp.:99.61, PWM:41
17:36:06.438 -> Water_Temp.:99.61, PWM:41
17:36:07.230 -> Water_Temp.:99.61, PWM:41
17:36:07.990 -> Water_Temp.:99.61, PWM:41
17:36:08.785 -> Water_Temp.:99.61, PWM:41
17:36:09.550 -> Water_Temp.:99.61, PWM:41
17:36:10.343 -> Water_Temp.:99.61, PWM:41
17:36:11.103 -> Water_Temp.:99.61, PWM:41
17:36:11.895 -> Water_Temp.:99.61, PWM:41
17:36:12.686 -> Water_Temp.:99.61, PWM:41
17:36:13.448 -> Water_Temp.:99.61, PWM:41
17:36:14.243 -> Water_Temp.:99.61, PWM:41
17:36:15.002 -> Water_Temp.:99.61, PWM:41
17:36:15.798 -> Water_Temp.:99.61, PWM:41
17:36:16.589 -> Water_Temp.:99.61, PWM:41
17:36:17.346 -> Water_Temp.:99.61, PWM:41
17:36:18.143 -> Water_Temp.:99.61, PWM:41
17:36:18.902 -> Water_Temp.:99.61, PWM:41
17:36:19.702 -> Water_Temp.:99.61, PWM:41
17:36:20.466 -> Water_Temp.:99.61, PWM:41
17:36:21.256 -> Water_Temp.:99.61, PWM:41
17:36:22.018 -> Water_Temp.:99.61, PWM:41
17:36:22.810 -> Water_Temp.:99.61, PWM:41
17:36:23.604 -> Water_Temp.:99.61, PWM:41
17:36:24.366 -> Water_Temp.:99.61, PWM:41
17:36:25.161 -> Water_Temp.:99.61, PWM:41
17:36:25.924 -> Water_Temp.:99.61, PWM:41
17:36:26.718 -> Water_Temp.:99.61, PWM:41
17:36:28.237 -> Water_Temp.:99.61, PWM:41
17:36:29.029 -> Water_Temp.:99.61, PWM:41
17:36:29.822 -> Water_Temp.:99.61, PWM:41
17:36:30.583 -> Water_Temp.:99.61, PWM:41
17:36:31.376 -> Water_Temp.:99.61, PWM:41
17:36:32.927 -> Water_Temp.:99.61, PWM:41
17:36:33.686 -> Water_Temp.:99.61, PWM:41
17:36:35.241 -> Water_Temp.:99.61, PWM:41
17:36:36.031 -> Water_Temp.:99.61, PWM:41
17:36:36.823 -> Water_Temp.:99.61, PWM:41
17:36:37.579 -> Water_Temp.:99.61, PWM:41
17:36:38.367 -> Water_Temp.:99.61, PWM:41
17:36:39.160 -> Water_Temp.:99.61, PWM:41
17:36:39.945 -> Water_Temp.:99.61, PWM:41
17:36:40.705 -> Water_Temp.:99.61, PWM:41
17:36:41.495 -> Water_Temp.:99.61, PWM:41
17:36:42.283 -> Water_Temp.:99.61, PWM:41
17:36:43.044 -> Water_Temp.:99.61, PWM:41
17:36:43.838 -> Water_Temp.:99.61, PWM:41
17:36:44.600 -> Water_Temp.:99.72, PWM:41
17:36:45.396 -> Water_Temp.:99.72, PWM:41
17:36:46.153 -> Water_Temp.:99.72, PWM:41
17:36:46.948 -> Water_Temp.:99.72, PWM:41
17:36:47.742 -> Water_Temp.:99.72, PWM:41
17:36:48.503 -> Water_Temp.:99.72, PWM:41
17:36:50.055 -> Water_Temp.:99.72, PWM:41
17:36:50.850 -> Water_Temp.:99.72, PWM:41
17:36:51.642 -> Water_Temp.:99.72, PWM:41
17:36:52.397 -> Water_Temp.:99.72, PWM:41
17:36:53.190 -> Water_Temp.:99.72, PWM:41
17:36:53.952 -> Water_Temp.:99.72, PWM:41
17:36:54.748 -> Water_Temp.:99.72, PWM:41
17:36:55.539 -> Water_Temp.:99.72, PWM:41
17:36:56.299 -> Water_Temp.:99.72, PWM:41
17:36:57.087 -> Water_Temp.:99.72, PWM:41
17:36:57.879 -> Water_Temp.:99.72, PWM:41
17:36:58.640 -> Water_Temp.:99.72, PWM:41
17:36:59.434 -> Water_Temp.:99.72, PWM:41
17:37:00.225 -> Water_Temp.:99.72, PWM:41
17:37:00.985 -> Water_Temp.:99.72, PWM:41
17:37:01.778 -> Water_Temp.:99.72, PWM:41
17:37:02.532 -> Water_Temp.:99.72, PWM:41
17:37:03.325 -> Water_Temp.:99.72, PWM:41
17:37:04.872 -> Water_Temp.:99.72, PWM:41
17:37:06.414 -> Water_Temp.:99.72, PWM:41
17:37:07.172 -> Water_Temp.:99.72, PWM:41
17:37:07.963 -> Water_Temp.:99.72, PWM:41
17:37:08.757 -> Water_Temp.:99.72, PWM:41
17:37:10.272 -> Water_Temp.:99.72, PWM:41
17:37:11.065 -> Water_Temp.:99.72, PWM:41
17:37:11.826 -> Water_Temp.:99.72, PWM:41
17:37:12.621 -> Water_Temp.:99.72, PWM:41
17:37:13.384 -> Water_Temp.:99.72, PWM:41
17:37:14.177 -> Water_Temp.:99.72, PWM:41
17:37:14.970 -> Water_Temp.:99.72, PWM:41
17:37:15.727 -> Water_Temp.:99.72, PWM:41
17:37:16.523 -> Water_Temp.:99.72, PWM:41
17:37:17.282 -> Water_Temp.:99.72, PWM:41
17:37:18.075 -> Water_Temp.:99.72, PWM:41
17:37:18.869 -> Water_Temp.:99.72, PWM:41
17:37:19.628 -> Water_Temp.:99.72, PWM:41
17:37:20.423 -> Water_Temp.:99.72, PWM:41
17:37:21.182 -> Water_Temp.:99.84, PWM:41
17:37:21.977 -> Water_Temp.:99.84, PWM:41
17:37:22.771 -> Water_Temp.:99.84, PWM:41
17:37:23.535 -> Water_Temp.:99.84, PWM:41
17:37:24.328 -> Water_Temp.:99.84, PWM:41
17:37:25.083 -> Water_Temp.:99.84, PWM:41
17:37:25.879 -> Water_Temp.:99.84, PWM:41
17:37:27.427 -> Water_Temp.:99.84, PWM:41
17:37:28.185 -> Water_Temp.:99.84, PWM:41
17:37:28.981 -> Water_Temp.:99.84, PWM:41
17:37:29.775 -> Water_Temp.:99.84, PWM:41
17:37:30.534 -> Water_Temp.:99.84, PWM:41
17:37:31.326 -> Water_Temp.:99.72, PWM:41
17:37:32.087 -> Water_Temp.:99.84, PWM:41
17:37:32.882 -> Water_Temp.:99.84, PWM:41
17:37:33.675 -> Water_Temp.:99.84, PWM:41
17:37:34.437 -> Water_Temp.:99.72, PWM:41
17:37:35.230 -> Water_Temp.:99.84, PWM:41
17:37:35.985 -> Water_Temp.:99.72, PWM:41
17:37:36.777 -> Water_Temp.:99.72, PWM:41
17:37:37.568 -> Water_Temp.:99.72, PWM:41
17:37:38.324 -> Water_Temp.:99.72, PWM:41
17:37:39.119 -> Water_Temp.:99.72, PWM:41
17:37:39.913 -> Water_Temp.:99.84, PWM:41
17:37:40.671 -> Water_Temp.:99.84, PWM:41
17:37:41.468 -> Water_Temp.:99.84, PWM:41
17:37:42.228 -> Water_Temp.:99.72, PWM:41
17:37:43.023 -> Water_Temp.:99.72, PWM:41
17:37:43.783 -> Water_Temp.:99.84, PWM:41
17:37:45.370 -> Water_Temp.:99.84, PWM:41
17:37:46.890 -> Water_Temp.:99.84, PWM:41
17:37:47.684 -> Water_Temp.:99.84, PWM:41
17:37:48.475 -> Water_Temp.:99.84, PWM:41
17:37:49.232 -> Water_Temp.:99.84, PWM:41
17:37:50.024 -> Water_Temp.:99.84, PWM:41
17:37:50.782 -> Water_Temp.:99.84, PWM:41
17:37:51.575 -> Water_Temp.:99.84, PWM:41
17:37:52.371 -> Water_Temp.:99.84, PWM:41
17:37:53.128 -> Water_Temp.:99.84, PWM:41
17:37:53.917 -> Water_Temp.:99.84, PWM:41
17:37:54.708 -> Water_Temp.:99.84, PWM:41
17:37:55.467 -> Water_Temp.:99.84, PWM:41
17:37:56.261 -> Water_Temp.:99.84, PWM:41
17:37:57.055 -> Water_Temp.:99.84, PWM:41
17:37:58.573 -> Water_Temp.:99.84, PWM:41
17:37:59.370 -> Water_Temp.:99.84, PWM:41
17:38:00.130 -> Water_Temp.:99.84, PWM:41
17:38:00.929 -> Water_Temp.:99.84, PWM:41
17:38:01.688 -> Water_Temp.:99.84, PWM:41
17:38:02.480 -> Water_Temp.:99.84, PWM:41
17:38:03.235 -> Water_Temp.:99.95, PWM:41
17:38:04.781 -> Water_Temp.:99.84, PWM:41
17:38:05.578 -> Water_Temp.:99.95, PWM:41
17:38:06.371 -> Water_Temp.:99.95, PWM:41
17:38:07.128 -> Water_Temp.:99.95, PWM:41
17:38:07.921 -> Water_Temp.:99.95, PWM:41
17:38:08.680 -> Water_Temp.:99.95, PWM:41
17:38:09.471 -> Water_Temp.:99.95, PWM:41
17:38:10.265 -> Water_Temp.:99.95, PWM:41
17:38:11.022 -> Water_Temp.:99.95, PWM:41
17:38:11.819 -> Water_Temp.:99.95, PWM:41
17:38:12.578 -> Water_Temp.:99.95, PWM:41
17:38:13.373 -> Water_Temp.:99.95, PWM:41
17:38:14.168 -> Water_Temp.:99.95, PWM:41
17:38:14.925 -> Water_Temp.:99.95, PWM:41
17:38:15.718 -> Water_Temp.:99.95, PWM:41
17:38:17.234 -> Water_Temp.:99.95, PWM:41
17:38:18.026 -> Water_Temp.:99.95, PWM:41
17:38:18.824 -> Water_Temp.:99.95, PWM:41
17:38:19.583 -> Water_Temp.:99.95, PWM:41
17:38:20.378 -> Water_Temp.:99.95, PWM:41
17:38:21.135 -> Water_Temp.:99.95, PWM:41
17:38:21.924 -> Water_Temp.:99.95, PWM:41
17:38:22.713 -> Water_Temp.:99.95, PWM:41
17:38:23.474 -> Water_Temp.:99.95, PWM:41
17:38:24.267 -> Water_Temp.:99.95, PWM:41
17:38:25.030 -> Water_Temp.:99.95, PWM:41
17:38:25.828 -> Water_Temp.:99.95, PWM:41
17:38:27.375 -> Water_Temp.:99.95, PWM:41
17:38:28.931 -> Water_Temp.:99.95, PWM:41
17:38:29.730 -> Water_Temp.:99.95, PWM:41
17:38:30.489 -> Water_Temp.:99.95, PWM:41
17:38:31.282 -> Water_Temp.:99.95, PWM:41
17:38:32.076 -> Water_Temp.:99.95, PWM:41
17:38:32.836 -> Water_Temp.:99.95, PWM:41
17:38:33.628 -> Water_Temp.:99.95, PWM:41
17:38:34.392 -> Water_Temp.:99.95, PWM:41
17:38:35.184 -> Water_Temp.:99.95, PWM:41
17:38:35.975 -> Water_Temp.:100.06, PWM:41
17:38:36.733 -> Water_Temp.:99.95, PWM:41
17:38:37.525 -> Water_Temp.:99.95, PWM:41
17:38:38.317 -> Water_Temp.:100.06, PWM:41
17:38:39.071 -> Water_Temp.:100.06, PWM:41
17:38:39.864 -> Water_Temp.:100.06, PWM:41
17:38:40.658 -> Water_Temp.:100.06, PWM:41
17:38:41.419 -> Water_Temp.:100.06, PWM:41
17:38:42.210 -> Water_Temp.:100.06, PWM:41
17:38:43.765 -> Water_Temp.:100.06, PWM:41
17:38:44.558 -> Water_Temp.:100.06, PWM:41
17:38:45.318 -> Water_Temp.:100.06, PWM:41
17:38:46.117 -> Water_Temp.:100.06, PWM:41
17:38:46.877 -> Water_Temp.:100.06, PWM:41
17:38:48.427 -> Water_Temp.:100.06, PWM:41
17:38:49.219 -> Water_Temp.:100.06, PWM:41
17:38:50.014 -> Water_Temp.:100.06, PWM:41
17:38:50.775 -> Water_Temp.:100.06, PWM:41
17:38:51.567 -> Water_Temp.:100.06, PWM:41
17:38:52.330 -> Water_Temp.:100.06, PWM:41
17:38:53.123 -> Water_Temp.:100.06, PWM:41
17:38:53.883 -> Water_Temp.:100.06, PWM:41
17:38:54.681 -> Water_Temp.:100.06, PWM:41
17:38:55.466 -> Water_Temp.:100.06, PWM:41
17:38:56.259 -> Water_Temp.:100.06, PWM:41
17:38:57.015 -> Water_Temp.:100.06, PWM:41
17:38:57.809 -> Water_Temp.:100.06, PWM:41
17:38:58.568 -> Water_Temp.:100.06, PWM:41
17:38:59.360 -> Water_Temp.:100.06, PWM:41
17:39:00.154 -> Water_Temp.:100.06, PWM:41
17:39:00.913 -> Water_Temp.:100.06, PWM:41
17:39:01.711 -> Water_Temp.:100.06, PWM:41
17:39:02.476 -> Water_Temp.:99.95, PWM:41
17:39:03.273 -> Water_Temp.:100.06, PWM:41
17:39:04.032 -> Water_Temp.:100.06, PWM:41
17:39:04.829 -> Water_Temp.:100.06, PWM:41
17:39:05.592 -> Water_Temp.:100.06, PWM:41
17:39:06.391 -> Water_Temp.:100.06, PWM:41
17:39:07.179 -> Water_Temp.:100.06, PWM:41
17:39:07.935 -> Water_Temp.:100.06, PWM:41
17:39:08.729 -> Water_Temp.:100.06, PWM:41
17:39:09.491 -> Water_Temp.:100.06, PWM:41
17:39:11.076 -> Water_Temp.:100.06, PWM:41
17:39:12.630 -> Water_Temp.:100.06, PWM:41
17:39:13.386 -> Water_Temp.:100.06, PWM:41
17:39:14.177 -> Water_Temp.:100.06, PWM:41
17:39:14.968 -> Water_Temp.:100.06, PWM:41
17:39:15.726 -> Water_Temp.:100.06, PWM:41
17:39:16.520 -> Water_Temp.:100.06, PWM:41
17:39:17.310 -> Water_Temp.:100.06, PWM:41
17:39:18.069 -> Water_Temp.:100.06, PWM:41
17:39:18.862 -> Water_Temp.:100.06, PWM:41
17:39:19.624 -> Water_Temp.:100.06, PWM:41
17:39:20.422 -> Water_Temp.:100.06, PWM:41
17:39:21.212 -> Water_Temp.:100.06, PWM:41
17:39:21.969 -> Water_Temp.:100.06, PWM:41
17:39:22.762 -> Water_Temp.:100.06, PWM:41
17:39:23.521 -> Water_Temp.:100.06, PWM:41
17:39:24.314 -> Water_Temp.:100.06, PWM:41
17:39:25.106 -> Water_Temp.:100.06, PWM:41
17:39:25.864 -> Water_Temp.:100.06, PWM:41
17:39:26.660 -> Water_Temp.:100.06, PWM:41
17:39:28.211 -> Water_Temp.:100.17, PWM:41
17:39:28.964 -> Water_Temp.:100.17, PWM:41
17:39:29.760 -> Water_Temp.:100.17, PWM:41
17:39:30.554 -> Water_Temp.:100.17, PWM:41
17:39:31.316 -> Water_Temp.:100.17, PWM:41
17:39:32.111 -> Water_Temp.:100.17, PWM:41
17:39:32.870 -> Water_Temp.:100.17, PWM:41
17:39:33.663 -> Water_Temp.:100.17, PWM:41
17:39:34.454 -> Water_Temp.:100.17, PWM:41
17:39:35.214 -> Water_Temp.:100.17, PWM:41
17:39:36.012 -> Water_Temp.:100.17, PWM:41
17:39:36.771 -> Water_Temp.:100.17, PWM:41
17:39:37.559 -> Water_Temp.:100.17, PWM:41
17:39:38.351 -> Water_Temp.:100.17, PWM:41
17:39:39.108 -> Water_Temp.:100.17, PWM:41
17:39:39.898 -> Water_Temp.:100.17, PWM:41
17:39:41.423 -> Water_Temp.:100.17, PWM:41
17:39:42.219 -> Water_Temp.:100.17, PWM:41
17:39:42.976 -> Water_Temp.:100.17, PWM:41
17:39:44.534 -> Water_Temp.:100.17, PWM:41
17:39:45.294 -> Water_Temp.:100.17, PWM:41
17:39:46.088 -> Water_Temp.:100.17, PWM:41
17:39:46.878 -> Water_Temp.:100.17, PWM:41
17:39:47.634 -> Water_Temp.:100.17, PWM:41
17:39:48.433 -> Water_Temp.:100.17, PWM:41
17:39:49.190 -> Water_Temp.:100.17, PWM:41
17:39:49.983 -> Water_Temp.:100.17, PWM:41
17:39:50.775 -> Water_Temp.:100.17, PWM:41
17:39:51.533 -> Water_Temp.:100.17, PWM:41
17:39:52.328 -> Water_Temp.:100.17, PWM:41
17:39:53.117 -> Water_Temp.:100.17, PWM:41
17:39:53.874 -> Water_Temp.:100.17, PWM:41
17:39:55.427 -> Water_Temp.:100.17, PWM:41
17:39:56.225 -> Water_Temp.:100.17, PWM:41
17:39:56.981 -> Water_Temp.:100.17, PWM:41
17:39:57.775 -> Water_Temp.:100.17, PWM:41
17:39:58.569 -> Water_Temp.:100.17, PWM:41
17:39:59.326 -> Water_Temp.:100.17, PWM:41
17:40:00.120 -> Water_Temp.:100.17, PWM:41
17:40:00.878 -> Water_Temp.:100.17, PWM:41
17:40:01.670 -> Water_Temp.:100.17, PWM:41
17:40:02.464 -> Water_Temp.:100.17, PWM:41
17:40:03.222 -> Water_Temp.:100.17, PWM:41
17:40:04.016 -> Water_Temp.:100.29, PWM:41
17:40:04.812 -> Water_Temp.:100.29, PWM:41
17:40:05.569 -> Water_Temp.:100.29, PWM:41
17:40:06.367 -> Water_Temp.:100.29, PWM:41
17:40:07.129 -> Water_Temp.:100.29, PWM:41
17:40:07.922 -> Water_Temp.:100.29, PWM:41
17:40:08.687 -> Water_Temp.:100.29, PWM:41
17:40:09.481 -> Water_Temp.:100.29, PWM:41
17:40:10.241 -> Water_Temp.:100.29, PWM:41
17:40:11.033 -> Water_Temp.:100.29, PWM:41
17:40:11.825 -> Water_Temp.:100.29, PWM:41
17:40:12.587 -> Water_Temp.:100.29, PWM:41
17:40:13.383 -> Water_Temp.:100.29, PWM:41
17:40:14.143 -> Water_Temp.:100.29, PWM:41
17:40:14.935 -> Water_Temp.:100.29, PWM:41
17:40:15.730 -> Water_Temp.:100.17, PWM:41
17:40:16.486 -> Water_Temp.:100.17, PWM:41
17:40:17.274 -> Water_Temp.:100.17, PWM:41
17:40:18.068 -> Water_Temp.:100.17, PWM:41
17:40:18.822 -> Water_Temp.:100.17, PWM:41
17:40:19.612 -> Water_Temp.:100.17, PWM:41
17:40:20.405 -> Water_Temp.:100.17, PWM:41
17:40:21.161 -> Water_Temp.:100.17, PWM:41
17:40:21.949 -> Water_Temp.:100.17, PWM:41
17:40:22.747 -> Water_Temp.:100.17, PWM:41
17:40:23.508 -> Water_Temp.:100.29, PWM:41
17:40:25.063 -> Water_Temp.:100.29, PWM:41
17:40:25.859 -> Water_Temp.:100.17, PWM:41
17:40:26.619 -> Water_Temp.:100.29, PWM:41
17:40:27.412 -> Water_Temp.:100.29, PWM:41
17:40:28.207 -> Water_Temp.:100.29, PWM:41
17:40:28.971 -> Water_Temp.:100.29, PWM:41
17:40:29.757 -> Water_Temp.:100.29, PWM:41
17:40:31.308 -> Water_Temp.:100.29, PWM:41
17:40:32.861 -> Water_Temp.:100.29, PWM:41
17:40:33.621 -> Water_Temp.:100.29, PWM:41
17:40:34.413 -> Water_Temp.:100.29, PWM:41
17:40:35.208 -> Water_Temp.:100.29, PWM:41
17:40:35.972 -> Water_Temp.:100.29, PWM:41
17:40:36.767 -> Water_Temp.:100.29, PWM:41
17:40:38.288 -> Water_Temp.:100.29, PWM:41
17:40:39.085 -> Water_Temp.:100.29, PWM:41
17:40:39.843 -> Water_Temp.:100.29, PWM:41
17:40:40.631 -> Water_Temp.:100.29, PWM:41
17:40:41.427 -> Water_Temp.:100.29, PWM:41
17:40:42.188 -> Water_Temp.:100.29, PWM:41
17:40:42.982 -> Water_Temp.:100.29, PWM:41
17:40:43.737 -> Water_Temp.:100.29, PWM:41
17:40:44.532 -> Water_Temp.:100.29, PWM:41
17:40:45.327 -> Water_Temp.:100.29, PWM:41
17:40:46.089 -> Water_Temp.:100.29, PWM:41
17:40:46.884 -> Water_Temp.:100.29, PWM:41
17:40:47.640 -> Water_Temp.:100.29, PWM:41
17:40:48.431 -> Water_Temp.:100.29, PWM:41
17:40:49.223 -> Water_Temp.:100.29, PWM:41
17:40:49.979 -> Water_Temp.:100.29, PWM:41
17:40:50.774 -> Water_Temp.:100.29, PWM:41
17:40:51.568 -> Water_Temp.:100.29, PWM:41
17:40:52.329 -> Water_Temp.:100.29, PWM:41
17:40:53.124 -> Water_Temp.:100.29, PWM:41
17:40:53.883 -> Water_Temp.:101.86, PWM:41
17:40:54.672 -> Water_Temp.:103.89, PWM:41
17:40:55.470 -> Water_Temp.:101.52, PWM:41
17:40:56.227 -> Water_Temp.:101.19, PWM:41
17:40:57.017 -> Water_Temp.:100.96, PWM:41
17:40:57.778 -> Water_Temp.:100.85, PWM:41
17:40:58.571 -> Water_Temp.:100.74, PWM:41
17:40:59.362 -> Water_Temp.:100.62, PWM:41
17:41:00.125 -> Water_Temp.:100.62, PWM:41
17:41:00.919 -> Water_Temp.:100.51, PWM:41
17:41:01.685 -> Water_Temp.:100.51, PWM:41
17:41:02.481 -> Water_Temp.:100.51, PWM:41
17:41:03.241 -> Water_Temp.:100.51, PWM:41
17:41:04.034 -> Water_Temp.:100.51, PWM:41
17:41:04.830 -> Water_Temp.:100.51, PWM:41
17:41:05.594 -> Water_Temp.:100.40, PWM:41
17:41:06.388 -> Water_Temp.:100.40, PWM:41
17:41:07.149 -> Water_Temp.:100.40, PWM:41
17:41:07.935 -> Water_Temp.:100.40, PWM:41
17:41:08.730 -> Water_Temp.:100.40, PWM:41
17:41:09.486 -> Water_Temp.:100.40, PWM:41
17:41:10.277 -> Water_Temp.:100.40, PWM:41
17:41:11.073 -> Water_Temp.:100.40, PWM:41
17:41:11.828 -> Water_Temp.:100.40, PWM:41
17:41:12.620 -> Water_Temp.:100.40, PWM:41
17:41:13.379 -> Water_Temp.:100.40, PWM:41
17:41:14.174 -> Water_Temp.:100.40, PWM:41
17:41:14.971 -> Water_Temp.:100.40, PWM:41
17:41:15.727 -> Water_Temp.:100.40, PWM:41
17:41:16.522 -> Water_Temp.:100.40, PWM:41
17:41:17.282 -> Water_Temp.:100.40, PWM:41
17:41:18.073 -> Water_Temp.:100.40, PWM:41
17:41:18.865 -> Water_Temp.:100.40, PWM:41
17:41:19.627 -> Water_Temp.:100.40, PWM:41
17:41:20.417 -> Water_Temp.:102.09, PWM:41
17:41:21.207 -> Water_Temp.:104.11, PWM:41
17:41:21.972 -> Water_Temp.:101.75, PWM:41
17:41:22.760 -> Water_Temp.:101.30, PWM:41
17:41:23.551 -> Water_Temp.:101.07, PWM:41
17:41:24.315 -> Water_Temp.:100.96, PWM:41
17:41:27.392 -> Water_Temp.:100.74, PWM:41
17:41:28.153 -> Water_Temp.:100.62, PWM:41
17:41:28.944 -> Water_Temp.:100.62, PWM:41
17:41:29.734 -> Water_Temp.:100.62, PWM:41
17:41:30.520 -> Water_Temp.:100.62, PWM:41
17:41:31.279 -> Water_Temp.:100.62, PWM:41
17:41:32.070 -> Water_Temp.:100.51, PWM:41
17:41:32.858 -> Water_Temp.:100.51, PWM:41
17:41:33.619 -> Water_Temp.:100.51, PWM:41
17:41:34.414 -> Water_Temp.:100.51, PWM:41
17:41:35.171 -> Water_Temp.:100.51, PWM:41
17:41:36.730 -> Water_Temp.:100.51, PWM:41
17:41:37.519 -> Water_Temp.:100.51, PWM:41
17:41:38.317 -> Water_Temp.:100.51, PWM:41
17:41:39.078 -> Water_Temp.:100.40, PWM:41
17:41:39.876 -> Water_Temp.:100.40, PWM:41
17:41:40.639 -> Water_Temp.:100.40, PWM:41
17:41:41.433 -> Water_Temp.:100.40, PWM:41
17:41:42.194 -> Water_Temp.:100.40, PWM:41
17:41:42.985 -> Water_Temp.:100.40, PWM:41
17:41:43.777 -> Water_Temp.:100.40, PWM:41
17:41:44.534 -> Water_Temp.:100.40, PWM:41
17:41:45.326 -> Water_Temp.:100.40, PWM:41
17:41:46.120 -> Water_Temp.:100.40, PWM:41
17:41:46.880 -> Water_Temp.:100.40, PWM:41
17:41:48.433 -> Water_Temp.:100.40, PWM:41
17:41:49.192 -> Water_Temp.:100.40, PWM:41
17:41:49.988 -> Water_Temp.:100.40, PWM:41
17:41:50.747 -> Water_Temp.:100.40, PWM:41
17:41:51.540 -> Water_Temp.:100.40, PWM:41
17:41:52.334 -> Water_Temp.:100.40, PWM:41
17:41:53.092 -> Water_Temp.:100.40, PWM:41
17:41:53.883 -> Water_Temp.:100.40, PWM:41
17:41:54.676 -> Water_Temp.:100.40, PWM:41
17:41:55.433 -> Water_Temp.:100.40, PWM:41
17:41:56.228 -> Water_Temp.:100.40, PWM:41
17:41:56.986 -> Water_Temp.:100.40, PWM:41
17:41:57.782 -> Water_Temp.:100.40, PWM:41
17:41:58.574 -> Water_Temp.:100.40, PWM:41
17:41:59.336 -> Water_Temp.:100.40, PWM:41
17:42:00.126 -> Water_Temp.:100.40, PWM:41
17:42:00.885 -> Water_Temp.:100.40, PWM:41
17:42:01.678 -> Water_Temp.:100.40, PWM:41
17:42:03.231 -> Water_Temp.:100.40, PWM:41
17:42:04.023 -> Water_Temp.:100.40, PWM:41
17:42:04.815 -> Water_Temp.:100.40, PWM:41
17:42:05.574 -> Water_Temp.:100.40, PWM:41
17:42:06.366 -> Water_Temp.:100.40, PWM:41
17:42:07.124 -> Water_Temp.:100.40, PWM:41
17:42:07.919 -> Water_Temp.:100.40, PWM:41
17:42:08.714 -> Water_Temp.:100.40, PWM:41
17:42:09.474 -> Water_Temp.:100.40, PWM:41
17:42:10.265 -> Water_Temp.:100.40, PWM:41
17:42:11.058 -> Water_Temp.:100.40, PWM:41
17:42:11.813 -> Water_Temp.:100.40, PWM:41
17:42:12.608 -> Water_Temp.:100.40, PWM:41
17:42:13.365 -> Water_Temp.:100.40, PWM:41
17:42:14.159 -> Water_Temp.:100.40, PWM:41
17:42:14.949 -> Water_Temp.:100.40, PWM:41
17:42:15.737 -> Water_Temp.:100.40, PWM:41
17:42:16.498 -> Water_Temp.:100.40, PWM:41
17:42:17.292 -> Water_Temp.:100.40, PWM:41
17:42:18.056 -> Water_Temp.:100.40, PWM:41
17:42:18.847 -> Water_Temp.:100.51, PWM:41
17:42:19.609 -> Water_Temp.:100.51, PWM:41
17:42:20.397 -> Water_Temp.:100.51, PWM:41
17:42:21.187 -> Water_Temp.:100.51, PWM:41
17:42:21.949 -> Water_Temp.:100.51, PWM:41
17:42:22.743 -> Water_Temp.:100.51, PWM:41
17:42:23.539 -> Water_Temp.:100.51, PWM:41
17:42:25.058 -> Water_Temp.:100.51, PWM:41
17:42:25.848 -> Water_Temp.:100.51, PWM:41
17:42:26.637 -> Water_Temp.:100.51, PWM:41
17:42:27.396 -> Water_Temp.:100.51, PWM:41
17:42:28.192 -> Water_Temp.:100.51, PWM:41
17:42:28.978 -> Water_Temp.:100.51, PWM:41
17:42:30.494 -> Water_Temp.:100.51, PWM:41
17:42:31.285 -> Water_Temp.:100.51, PWM:41
17:42:32.077 -> Water_Temp.:100.51, PWM:41
17:42:32.836 -> Water_Temp.:100.51, PWM:41
17:42:33.632 -> Water_Temp.:100.51, PWM:41
17:42:34.426 -> Water_Temp.:100.51, PWM:41
17:42:35.186 -> Water_Temp.:100.51, PWM:41
17:42:36.731 -> Water_Temp.:100.51, PWM:41
17:42:37.525 -> Water_Temp.:100.51, PWM:41
17:42:38.318 -> Water_Temp.:100.51, PWM:41
17:42:39.076 -> Water_Temp.:100.40, PWM:41
17:42:39.871 -> Water_Temp.:100.40, PWM:41
17:42:40.660 -> Water_Temp.:100.40, PWM:41
17:42:41.420 -> Water_Temp.:100.40, PWM:41
17:42:42.208 -> Water_Temp.:100.51, PWM:41
17:42:43.002 -> Water_Temp.:100.51, PWM:41
17:42:43.761 -> Water_Temp.:100.51, PWM:41
17:42:44.552 -> Water_Temp.:100.51, PWM:41
17:42:45.340 -> Water_Temp.:100.51, PWM:41
17:42:46.100 -> Water_Temp.:100.51, PWM:41
17:42:46.897 -> Water_Temp.:100.51, PWM:41
17:42:47.658 -> Water_Temp.:100.51, PWM:41
17:42:48.448 -> Water_Temp.:100.51, PWM:41
17:42:49.244 -> Water_Temp.:100.51, PWM:41
17:42:50.002 -> Water_Temp.:100.51, PWM:41
17:42:52.343 -> Water_Temp.:100.40, PWM:41
17:42:53.133 -> Water_Temp.:100.51, PWM:41
17:42:54.681 -> Water_Temp.:100.40, PWM:41
17:42:55.471 -> Water_Temp.:100.40, PWM:41
17:42:56.228 -> Water_Temp.:100.40, PWM:41
17:42:57.019 -> Water_Temp.:100.40, PWM:41
17:42:57.813 -> Water_Temp.:100.40, PWM:41
17:42:59.334 -> Water_Temp.:100.40, PWM:41
17:43:00.131 -> Water_Temp.:100.40, PWM:41
17:43:00.891 -> Water_Temp.:100.40, PWM:41
17:43:01.687 -> Water_Temp.:100.40, PWM:41
17:43:02.480 -> Water_Temp.:100.40, PWM:41
17:43:03.235 -> Water_Temp.:100.40, PWM:41
17:43:04.029 -> Water_Temp.:100.40, PWM:41
17:43:04.819 -> Water_Temp.:100.40, PWM:41
17:43:05.583 -> Water_Temp.:100.40, PWM:41
17:43:06.380 -> Water_Temp.:100.40, PWM:41
17:43:07.140 -> Water_Temp.:100.40, PWM:41
17:43:07.935 -> Water_Temp.:100.51, PWM:41
17:43:08.694 -> Water_Temp.:100.51, PWM:41
17:43:09.490 -> Water_Temp.:102.09, PWM:41
17:43:10.249 -> Water_Temp.:104.11, PWM:41
17:43:11.043 -> Water_Temp.:101.75, PWM:41
17:43:11.838 -> Water_Temp.:101.41, PWM:41
17:43:12.596 -> Water_Temp.:101.19, PWM:41
17:43:13.388 -> Water_Temp.:101.07, PWM:41
17:43:14.181 -> Water_Temp.:100.96, PWM:41
17:43:14.941 -> Water_Temp.:100.85, PWM:41
17:43:15.733 -> Water_Temp.:100.85, PWM:41
17:43:16.492 -> Water_Temp.:100.74, PWM:41
17:43:17.283 -> Water_Temp.:100.74, PWM:41
17:43:18.078 -> Water_Temp.:100.74, PWM:41
17:43:19.636 -> Water_Temp.:100.62, PWM:41
17:43:20.393 -> Water_Temp.:100.62, PWM:41
17:43:21.189 -> Water_Temp.:100.62, PWM:41
17:43:21.949 -> Water_Temp.:100.62, PWM:41
17:43:22.739 -> Water_Temp.:100.62, PWM:41
17:43:23.534 -> Water_Temp.:100.51, PWM:41
17:43:24.293 -> Water_Temp.:100.51, PWM:41
17:43:25.090 -> Water_Temp.:100.51, PWM:41
17:43:25.848 -> Water_Temp.:100.51, PWM:41
17:43:26.644 -> Water_Temp.:100.51, PWM:41
17:43:27.434 -> Water_Temp.:100.51, PWM:41
17:43:28.195 -> Water_Temp.:100.51, PWM:41
17:43:28.989 -> Water_Temp.:100.51, PWM:46
17:43:29.751 -> Water_Temp.:100.51, PWM:46
17:43:31.330 -> Water_Temp.:100.51, PWM:46
17:43:32.091 -> Water_Temp.:100.51, PWM:46
17:43:32.885 -> Water_Temp.:100.51, PWM:46
17:43:33.647 -> Water_Temp.:100.51, PWM:46
17:43:34.439 -> Water_Temp.:100.51, PWM:46
17:43:35.236 -> Water_Temp.:100.51, PWM:46
17:43:35.997 -> Water_Temp.:100.51, PWM:46
17:43:36.792 -> Water_Temp.:100.51, PWM:46
17:43:37.549 -> Water_Temp.:100.51, PWM:46
17:43:38.347 -> Water_Temp.:100.62, PWM:46
17:43:39.867 -> Water_Temp.:100.62, PWM:46
17:43:40.656 -> Water_Temp.:100.62, PWM:46
17:43:41.450 -> Water_Temp.:100.62, PWM:46
17:43:42.207 -> Water_Temp.:100.62, PWM:46
17:43:42.997 -> Water_Temp.:100.62, PWM:46
17:43:43.791 -> Water_Temp.:100.62, PWM:46
17:43:44.549 -> Water_Temp.:100.62, PWM:46
17:43:45.342 -> Water_Temp.:100.62, PWM:46
17:43:46.131 -> Water_Temp.:100.62, PWM:46
17:43:46.889 -> Water_Temp.:102.20, PWM:46
17:43:47.683 -> Water_Temp.:104.22, PWM:46
17:43:48.441 -> Water_Temp.:101.86, PWM:46
17:43:49.237 -> Water_Temp.:101.41, PWM:46
17:43:50.033 -> Water_Temp.:101.30, PWM:46
17:43:50.787 -> Water_Temp.:101.19, PWM:46
17:43:51.578 -> Water_Temp.:101.07, PWM:46
17:43:52.368 -> Water_Temp.:100.96, PWM:46
17:43:53.127 -> Water_Temp.:100.85, PWM:46
17:43:53.919 -> Water_Temp.:100.85, PWM:46
17:43:55.469 -> Water_Temp.:100.74, PWM:46
17:43:56.231 -> Water_Temp.:100.74, PWM:46
17:43:57.019 -> Water_Temp.:100.74, PWM:46
17:43:57.779 -> Water_Temp.:100.74, PWM:46
17:43:58.574 -> Water_Temp.:100.74, PWM:46
17:43:59.365 -> Water_Temp.:100.62, PWM:46
17:44:00.127 -> Water_Temp.:100.62, PWM:46
17:44:00.921 -> Water_Temp.:100.62, PWM:46
17:44:01.708 -> Water_Temp.:100.62, PWM:46
17:44:02.467 -> Water_Temp.:100.62, PWM:46
17:44:03.263 -> Water_Temp.:100.62, PWM:46
17:44:04.051 -> Water_Temp.:100.62, PWM:46
17:44:04.813 -> Water_Temp.:100.62, PWM:46
17:44:05.609 -> Water_Temp.:100.62, PWM:46
17:44:06.367 -> Water_Temp.:100.62, PWM:46
17:44:07.162 -> Water_Temp.:100.62, PWM:46
17:44:07.922 -> Water_Temp.:100.51, PWM:46
17:44:08.713 -> Water_Temp.:100.51, PWM:46
17:44:09.505 -> Water_Temp.:100.51, PWM:46
17:44:10.264 -> Water_Temp.:100.51, PWM:46
17:44:11.059 -> Water_Temp.:100.51, PWM:46
17:44:11.819 -> Water_Temp.:100.51, PWM:46
17:44:12.614 -> Water_Temp.:100.62, PWM:46
17:44:13.408 -> Water_Temp.:100.51, PWM:46
17:44:14.924 -> Water_Temp.:100.62, PWM:46
17:44:15.713 -> Water_Temp.:100.62, PWM:46
17:44:16.472 -> Water_Temp.:100.62, PWM:46
17:44:17.264 -> Water_Temp.:100.62, PWM:46
17:44:18.057 -> Water_Temp.:100.62, PWM:46
17:44:18.818 -> Water_Temp.:100.62, PWM:46
17:44:19.612 -> Water_Temp.:100.62, PWM:46
17:44:20.402 -> Water_Temp.:100.62, PWM:46
17:44:21.159 -> Water_Temp.:100.62, PWM:46
17:44:21.954 -> Water_Temp.:100.62, PWM:46
17:44:22.713 -> Water_Temp.:100.62, PWM:46
17:44:23.508 -> Water_Temp.:100.62, PWM:46
17:44:24.302 -> Water_Temp.:100.62, PWM:46
17:44:25.823 -> Water_Temp.:100.62, PWM:46
17:44:28.173 -> Water_Temp.:100.62, PWM:46
17:44:28.931 -> Water_Temp.:100.62, PWM:46
17:44:29.723 -> Water_Temp.:100.62, PWM:46
17:44:30.513 -> Water_Temp.:100.62, PWM:46
17:44:31.273 -> Water_Temp.:100.62, PWM:46
17:44:32.068 -> Water_Temp.:100.62, PWM:46
17:44:32.855 -> Water_Temp.:100.62, PWM:46
17:44:33.616 -> Water_Temp.:100.62, PWM:46
17:44:34.409 -> Water_Temp.:100.62, PWM:46
17:44:35.171 -> Water_Temp.:100.51, PWM:46
17:44:35.963 -> Water_Temp.:100.51, PWM:46
17:44:36.756 -> Water_Temp.:100.51, PWM:46
17:44:37.513 -> Water_Temp.:100.62, PWM:46
17:44:38.305 -> Water_Temp.:100.51, PWM:46
17:44:39.099 -> Water_Temp.:100.51, PWM:46
17:44:39.856 -> Water_Temp.:100.62, PWM:46
17:44:40.646 -> Water_Temp.:100.51, PWM:46
17:44:41.436 -> Water_Temp.:100.62, PWM:46
17:44:42.197 -> Water_Temp.:100.51, PWM:46
17:44:42.991 -> Water_Temp.:100.51, PWM:46
17:44:43.750 -> Water_Temp.:100.62, PWM:46
17:44:45.301 -> Water_Temp.:100.62, PWM:46
17:44:46.093 -> Water_Temp.:100.62, PWM:46
17:44:46.848 -> Water_Temp.:100.62, PWM:46
17:44:47.639 -> Water_Temp.:100.62, PWM:46
17:44:48.431 -> Water_Temp.:100.62, PWM:46
17:44:49.188 -> Water_Temp.:100.62, PWM:46
17:44:49.980 -> Water_Temp.:100.62, PWM:46
17:44:50.773 -> Water_Temp.:100.62, PWM:46
17:44:51.532 -> Water_Temp.:100.62, PWM:46
17:44:52.323 -> Water_Temp.:100.62, PWM:46
17:44:53.117 -> Water_Temp.:100.62, PWM:46
17:44:53.873 -> Water_Temp.:100.62, PWM:46
17:44:54.667 -> Water_Temp.:100.62, PWM:46
17:44:55.426 -> Water_Temp.:100.62, PWM:46
17:44:56.982 -> Water_Temp.:100.62, PWM:46
17:44:57.775 -> Water_Temp.:100.62, PWM:46
17:44:58.535 -> Water_Temp.:100.62, PWM:46
17:44:59.327 -> Water_Temp.:100.62, PWM:46
17:45:00.085 -> Water_Temp.:100.62, PWM:46
17:45:00.883 -> Water_Temp.:100.62, PWM:46
17:45:01.647 -> Water_Temp.:100.62, PWM:46
17:45:02.439 -> Water_Temp.:100.62, PWM:46
17:45:03.233 -> Water_Temp.:100.62, PWM:46
17:45:03.990 -> Water_Temp.:100.62, PWM:46
17:45:04.783 -> Water_Temp.:100.62, PWM:46
17:45:05.572 -> Water_Temp.:100.62, PWM:46
17:45:06.329 -> Water_Temp.:100.74, PWM:46
17:45:07.125 -> Water_Temp.:100.74, PWM:46
17:45:07.919 -> Water_Temp.:100.74, PWM:46
17:45:08.681 -> Water_Temp.:100.74, PWM:46
17:45:09.443 -> Water_Temp.:100.74, PWM:46
17:45:10.237 -> Water_Temp.:100.74, PWM:46
17:45:11.032 -> Water_Temp.:100.74, PWM:46
17:45:11.791 -> Water_Temp.:100.74, PWM:46
17:45:12.582 -> Water_Temp.:100.74, PWM:46
17:45:13.377 -> Water_Temp.:100.74, PWM:46
17:45:14.139 -> Water_Temp.:100.74, PWM:46
17:45:14.903 -> Water_Temp.:100.74, PWM:46
17:45:15.698 -> Water_Temp.:100.74, PWM:46
17:45:16.488 -> Water_Temp.:100.74, PWM:46
17:45:17.250 -> Water_Temp.:100.74, PWM:46
17:45:18.047 -> Water_Temp.:100.74, PWM:46
17:45:18.838 -> Water_Temp.:100.74, PWM:46
17:45:19.598 -> Water_Temp.:100.74, PWM:46
17:45:20.392 -> Water_Temp.:100.74, PWM:46
17:45:21.152 -> Water_Temp.:100.74, PWM:46
17:45:21.950 -> Water_Temp.:100.74, PWM:46
17:45:22.711 -> Water_Temp.:100.74, PWM:46
17:45:23.502 -> Water_Temp.:100.74, PWM:46
17:45:24.297 -> Water_Temp.:100.74, PWM:46
17:45:25.055 -> Water_Temp.:100.74, PWM:46
17:45:25.853 -> Water_Temp.:100.74, PWM:46
17:45:26.613 -> Water_Temp.:100.74, PWM:46
17:45:27.406 -> Water_Temp.:100.74, PWM:46
17:45:28.196 -> Water_Temp.:100.74, PWM:46
17:45:29.722 -> Water_Temp.:100.74, PWM:46
17:45:30.479 -> Water_Temp.:100.74, PWM:46
17:45:31.267 -> Water_Temp.:100.74, PWM:46
17:45:32.056 -> Water_Temp.:100.74, PWM:46
17:45:32.847 -> Water_Temp.:100.74, PWM:46
17:45:33.606 -> Water_Temp.:100.74, PWM:46
17:45:34.396 -> Water_Temp.:100.74, PWM:46
17:45:35.160 -> Water_Temp.:100.74, PWM:46
17:45:36.717 -> Water_Temp.:100.74, PWM:46
17:45:37.511 -> Water_Temp.:100.74, PWM:46
17:45:38.267 -> Water_Temp.:100.74, PWM:46
17:45:39.060 -> Water_Temp.:100.74, PWM:46
17:45:40.579 -> Water_Temp.:100.74, PWM:46
17:45:41.373 -> Water_Temp.:100.74, PWM:46
17:45:42.166 -> Water_Temp.:100.74, PWM:46
17:45:42.929 -> Water_Temp.:100.74, PWM:46
17:45:43.724 -> Water_Temp.:100.74, PWM:46
17:45:44.479 -> Water_Temp.:100.74, PWM:46
17:45:46.033 -> Water_Temp.:100.74, PWM:46
17:45:46.828 -> Water_Temp.:100.74, PWM:46
17:45:47.585 -> Water_Temp.:100.74, PWM:46
17:45:48.380 -> Water_Temp.:100.74, PWM:46
17:45:49.142 -> Water_Temp.:100.74, PWM:46
17:45:49.939 -> Water_Temp.:100.74, PWM:46
17:45:50.731 -> Water_Temp.:100.74, PWM:46
17:45:51.488 -> Water_Temp.:100.74, PWM:46
17:45:52.277 -> Water_Temp.:100.74, PWM:46
17:45:53.070 -> Water_Temp.:100.74, PWM:46
17:45:53.825 -> Water_Temp.:100.74, PWM:46
17:45:54.615 -> Water_Temp.:100.74, PWM:46
17:45:55.409 -> Water_Temp.:100.74, PWM:46
17:45:56.168 -> Water_Temp.:100.85, PWM:46
17:45:56.963 -> Water_Temp.:100.85, PWM:46
17:45:57.754 -> Water_Temp.:100.85, PWM:46
17:45:58.515 -> Water_Temp.:100.85, PWM:46
17:45:59.301 -> Water_Temp.:100.85, PWM:46
17:46:00.090 -> Water_Temp.:100.85, PWM:46
17:46:00.851 -> Water_Temp.:100.85, PWM:46
17:46:01.642 -> Water_Temp.:100.85, PWM:46
17:46:02.433 -> Water_Temp.:100.85, PWM:46
17:46:03.194 -> Water_Temp.:100.85, PWM:46
17:46:03.989 -> Water_Temp.:100.85, PWM:46
17:46:04.748 -> Water_Temp.:100.85, PWM:46
17:46:05.543 -> Water_Temp.:100.85, PWM:46
17:46:06.300 -> Water_Temp.:100.85, PWM:46
17:46:07.094 -> Water_Temp.:100.85, PWM:46
17:46:07.889 -> Water_Temp.:100.85, PWM:46
17:46:08.649 -> Water_Temp.:100.85, PWM:46
17:46:09.445 -> Water_Temp.:100.85, PWM:46
17:46:10.205 -> Water_Temp.:100.85, PWM:46
17:46:11.001 -> Water_Temp.:100.85, PWM:46
17:46:11.792 -> Water_Temp.:100.85, PWM:46
17:46:12.550 -> Water_Temp.:100.85, PWM:46
17:46:13.343 -> Water_Temp.:100.85, PWM:46
17:46:14.132 -> Water_Temp.:100.85, PWM:46
17:46:14.893 -> Water_Temp.:100.85, PWM:46
17:46:15.685 -> Water_Temp.:100.85, PWM:46
17:46:16.446 -> Water_Temp.:100.85, PWM:46
17:46:17.238 -> Water_Temp.:100.85, PWM:46
17:46:18.031 -> Water_Temp.:100.85, PWM:46
17:46:18.789 -> Water_Temp.:100.96, PWM:46
17:46:20.369 -> Water_Temp.:100.96, PWM:46
17:46:21.128 -> Water_Temp.:100.96, PWM:46
17:46:21.917 -> Water_Temp.:100.96, PWM:46
17:46:22.711 -> Water_Temp.:100.96, PWM:46
17:46:23.472 -> Water_Temp.:100.96, PWM:46
17:46:25.029 -> Water_Temp.:100.96, PWM:46
17:46:25.790 -> Water_Temp.:100.96, PWM:46
17:46:26.587 -> Water_Temp.:100.96, PWM:46
17:46:27.344 -> Water_Temp.:100.96, PWM:46
17:46:28.133 -> Water_Temp.:100.96, PWM:46
17:46:28.927 -> Water_Temp.:100.96, PWM:46
17:46:29.713 -> Water_Temp.:100.96, PWM:46
17:46:30.473 -> Water_Temp.:100.96, PWM:46
17:46:32.026 -> Water_Temp.:100.96, PWM:46
17:46:32.819 -> Water_Temp.:100.96, PWM:46
17:46:33.580 -> Water_Temp.:100.96, PWM:46
17:46:34.369 -> Water_Temp.:100.96, PWM:46
17:46:35.127 -> Water_Temp.:100.85, PWM:46
17:46:35.917 -> Water_Temp.:100.96, PWM:46
17:46:36.714 -> Water_Temp.:100.96, PWM:46
17:46:37.470 -> Water_Temp.:100.96, PWM:46
17:46:38.265 -> Water_Temp.:100.96, PWM:46
17:46:39.055 -> Water_Temp.:100.96, PWM:46
17:46:39.814 -> Water_Temp.:100.96, PWM:46
17:46:40.606 -> Water_Temp.:100.96, PWM:46
17:46:41.398 -> Water_Temp.:100.96, PWM:46
17:46:42.161 -> Water_Temp.:100.96, PWM:46
17:46:42.951 -> Water_Temp.:100.96, PWM:46
17:46:44.475 -> Water_Temp.:100.96, PWM:46
17:46:45.269 -> Water_Temp.:100.96, PWM:46
17:46:46.033 -> Water_Temp.:100.96, PWM:46
17:46:46.828 -> Water_Temp.:100.96, PWM:46
17:46:47.583 -> Water_Temp.:100.96, PWM:46
17:46:48.379 -> Water_Temp.:100.96, PWM:46
17:46:49.167 -> Water_Temp.:100.96, PWM:46
17:46:49.921 -> Water_Temp.:100.96, PWM:46
17:46:50.711 -> Water_Temp.:100.96, PWM:46
17:46:51.503 -> Water_Temp.:100.96, PWM:46
17:46:52.291 -> Water_Temp.:100.85, PWM:46
17:46:53.809 -> Water_Temp.:100.85, PWM:46
17:46:54.598 -> Water_Temp.:100.85, PWM:46
17:46:55.391 -> Water_Temp.:100.85, PWM:46
17:46:56.150 -> Water_Temp.:100.85, PWM:46
17:46:56.943 -> Water_Temp.:100.85, PWM:46
17:46:57.699 -> Water_Temp.:100.85, PWM:46
17:46:58.491 -> Water_Temp.:100.85, PWM:46
17:47:00.042 -> Water_Temp.:100.85, PWM:46
17:47:00.803 -> Water_Temp.:100.96, PWM:46
17:47:01.599 -> Water_Temp.:100.85, PWM:46
17:47:02.360 -> Water_Temp.:100.96, PWM:46
17:47:03.157 -> Water_Temp.:100.96, PWM:46
17:47:03.950 -> Water_Temp.:100.96, PWM:46
17:47:04.709 -> Water_Temp.:100.96, PWM:46
17:47:05.503 -> Water_Temp.:100.96, PWM:46
17:47:06.291 -> Water_Temp.:100.96, PWM:46
17:47:07.051 -> Water_Temp.:100.96, PWM:46
17:47:07.851 -> Water_Temp.:100.96, PWM:46
17:47:08.614 -> Water_Temp.:100.96, PWM:46
17:47:09.412 -> Water_Temp.:100.96, PWM:46
17:47:10.179 -> Water_Temp.:100.96, PWM:46
17:47:10.942 -> Water_Temp.:101.07, PWM:46
17:47:11.738 -> Water_Temp.:101.07, PWM:46
17:47:12.503 -> Water_Temp.:101.07, PWM:46
17:47:13.298 -> Water_Temp.:101.07, PWM:46
17:47:14.056 -> Water_Temp.:101.07, PWM:46
17:47:14.852 -> Water_Temp.:101.07, PWM:46
17:47:15.648 -> Water_Temp.:101.07, PWM:46
17:47:16.410 -> Water_Temp.:101.07, PWM:46
17:47:17.209 -> Water_Temp.:101.07, PWM:46
17:47:17.971 -> Water_Temp.:101.07, PWM:46
17:47:18.765 -> Water_Temp.:101.07, PWM:46
17:47:19.530 -> Water_Temp.:101.07, PWM:46
17:47:20.324 -> Water_Temp.:101.07, PWM:46
17:47:21.084 -> Water_Temp.:102.65, PWM:46
17:47:21.879 -> Water_Temp.:104.67, PWM:46
17:47:22.639 -> Water_Temp.:102.31, PWM:46
17:47:23.436 -> Water_Temp.:101.86, PWM:46
17:47:24.198 -> Water_Temp.:101.75, PWM:46
17:47:24.993 -> Water_Temp.:101.64, PWM:46
17:47:25.789 -> Water_Temp.:101.52, PWM:46
17:47:26.551 -> Water_Temp.:101.41, PWM:46
17:47:27.345 -> Water_Temp.:101.41, PWM:46
17:47:28.108 -> Water_Temp.:101.30, PWM:46
17:47:28.905 -> Water_Temp.:101.30, PWM:46
17:47:29.672 -> Water_Temp.:101.30, PWM:46
17:47:30.468 -> Water_Temp.:101.19, PWM:46
17:47:31.230 -> Water_Temp.:101.19, PWM:46
17:47:32.028 -> Water_Temp.:101.19, PWM:46
17:47:32.787 -> Water_Temp.:101.19, PWM:46
17:47:33.583 -> Water_Temp.:101.19, PWM:46
17:47:34.346 -> Water_Temp.:102.76, PWM:46
17:47:35.143 -> Water_Temp.:104.90, PWM:46
17:47:35.905 -> Water_Temp.:102.54, PWM:46
17:47:36.700 -> Water_Temp.:102.09, PWM:46
17:47:37.460 -> Water_Temp.:101.86, PWM:46
17:47:38.253 -> Water_Temp.:101.75, PWM:46
17:47:39.015 -> Water_Temp.:101.64, PWM:46
17:47:39.812 -> Water_Temp.:101.64, PWM:46
17:47:40.610 -> Water_Temp.:101.52, PWM:46
17:47:41.373 -> Water_Temp.:101.52, PWM:46
17:47:42.167 -> Water_Temp.:101.52, PWM:46
17:47:43.725 -> Water_Temp.:101.41, PWM:46
17:47:44.488 -> Water_Temp.:101.41, PWM:46
17:47:45.287 -> Water_Temp.:101.41, PWM:46
17:47:46.045 -> Water_Temp.:101.41, PWM:46
17:47:46.842 -> Water_Temp.:101.41, PWM:46
17:47:47.602 -> Water_Temp.:101.41, PWM:46
17:47:48.396 -> Water_Temp.:101.30, PWM:46
17:47:49.163 -> Water_Temp.:101.30, PWM:46
17:47:49.962 -> Water_Temp.:101.30, PWM:46
17:47:50.725 -> Water_Temp.:101.30, PWM:46
17:47:51.522 -> Water_Temp.:101.30, PWM:46
17:47:52.284 -> Water_Temp.:101.30, PWM:46
17:47:53.083 -> Water_Temp.:101.30, PWM:46
17:47:53.843 -> Water_Temp.:101.30, PWM:46
17:47:54.639 -> Water_Temp.:101.30, PWM:46
17:47:55.399 -> Water_Temp.:101.30, PWM:46
17:47:56.194 -> Water_Temp.:101.30, PWM:46
17:47:56.956 -> Water_Temp.:101.30, PWM:46
17:47:57.750 -> Water_Temp.:101.30, PWM:46
17:47:58.542 -> Water_Temp.:101.30, PWM:46
17:47:59.301 -> Water_Temp.:101.30, PWM:46
17:48:00.099 -> Water_Temp.:101.30, PWM:46
17:48:00.893 -> Water_Temp.:101.30, PWM:46
17:48:01.651 -> Water_Temp.:101.30, PWM:46
17:48:02.447 -> Water_Temp.:101.19, PWM:46
17:48:03.206 -> Water_Temp.:101.19, PWM:46
17:48:04.000 -> Water_Temp.:101.19, PWM:46
17:48:04.790 -> Water_Temp.:101.30, PWM:46
17:48:05.556 -> Water_Temp.:101.30, PWM:46
17:48:06.318 -> Water_Temp.:101.30, PWM:46
17:48:07.116 -> Water_Temp.:101.30, PWM:46
17:48:07.908 -> Water_Temp.:101.30, PWM:46
17:48:08.670 -> Water_Temp.:101.30, PWM:46
17:48:09.467 -> Water_Temp.:101.30, PWM:46
17:48:10.229 -> Water_Temp.:101.30, PWM:46
17:48:11.787 -> Water_Temp.:101.30, PWM:46
17:48:12.582 -> Water_Temp.:101.30, PWM:46
17:48:13.344 -> Water_Temp.:101.30, PWM:46
17:48:14.141 -> Water_Temp.:101.30, PWM:46
17:48:14.905 -> Water_Temp.:101.30, PWM:46
17:48:15.706 -> Water_Temp.:101.30, PWM:46
17:48:16.469 -> Water_Temp.:101.30, PWM:46
17:48:17.263 -> Water_Temp.:101.30, PWM:46
17:48:18.021 -> Water_Temp.:101.30, PWM:46
17:48:18.814 -> Water_Temp.:101.30, PWM:46
17:48:19.577 -> Water_Temp.:101.30, PWM:46
17:48:20.374 -> Water_Temp.:101.30, PWM:46
17:48:21.164 -> Water_Temp.:101.30, PWM:46
17:48:21.923 -> Water_Temp.:101.30, PWM:46
17:48:23.476 -> Water_Temp.:101.30, PWM:46
17:48:24.264 -> Water_Temp.:101.30, PWM:46
17:48:25.060 -> Water_Temp.:101.30, PWM:46
17:48:25.823 -> Water_Temp.:101.30, PWM:46
17:48:27.380 -> Water_Temp.:101.30, PWM:46
17:48:28.143 -> Water_Temp.:101.30, PWM:46
17:48:28.939 -> Water_Temp.:101.30, PWM:46
17:48:29.702 -> Water_Temp.:101.30, PWM:46
17:48:30.496 -> Water_Temp.:101.30, PWM:46
17:48:31.290 -> Water_Temp.:101.30, PWM:46
17:48:32.052 -> Water_Temp.:101.30, PWM:46
17:48:32.845 -> Water_Temp.:101.30, PWM:46
17:48:33.605 -> Water_Temp.:101.30, PWM:46
17:48:35.159 -> Water_Temp.:101.30, PWM:46
17:48:35.953 -> Water_Temp.:101.30, PWM:46
17:48:36.745 -> Water_Temp.:101.30, PWM:46
17:48:37.506 -> Water_Temp.:101.30, PWM:46
17:48:38.299 -> Water_Temp.:101.30, PWM:46
17:48:39.093 -> Water_Temp.:101.30, PWM:46
17:48:39.853 -> Water_Temp.:101.30, PWM:46
17:48:40.645 -> Water_Temp.:101.30, PWM:46
17:48:41.402 -> Water_Temp.:101.30, PWM:46
17:48:42.196 -> Water_Temp.:101.30, PWM:46
17:48:42.989 -> Water_Temp.:101.30, PWM:46
17:48:43.750 -> Water_Temp.:101.30, PWM:46
17:48:44.545 -> Water_Temp.:101.30, PWM:46
17:48:45.307 -> Water_Temp.:101.30, PWM:46
17:48:46.103 -> Water_Temp.:101.30, PWM:46
17:48:46.864 -> Water_Temp.:101.30, PWM:46
17:48:47.660 -> Water_Temp.:101.30, PWM:46
17:48:48.451 -> Water_Temp.:101.30, PWM:46
17:48:49.217 -> Water_Temp.:101.30, PWM:46
17:48:50.010 -> Water_Temp.:101.30, PWM:46
17:48:50.768 -> Water_Temp.:101.30, PWM:46
17:48:51.563 -> Water_Temp.:101.30, PWM:46
17:48:52.319 -> Water_Temp.:101.41, PWM:46
17:48:53.111 -> Water_Temp.:101.41, PWM:46
17:48:53.903 -> Water_Temp.:101.41, PWM:46
17:48:54.664 -> Water_Temp.:101.41, PWM:46
17:48:55.459 -> Water_Temp.:101.41, PWM:46
17:48:56.218 -> Water_Temp.:101.41, PWM:46
17:48:57.017 -> Water_Temp.:101.41, PWM:46
17:48:57.805 -> Water_Temp.:101.41, PWM:46
17:48:58.567 -> Water_Temp.:101.41, PWM:46
17:48:59.357 -> Water_Temp.:101.41, PWM:46
17:49:00.116 -> Water_Temp.:101.41, PWM:46
17:49:00.910 -> Water_Temp.:101.41, PWM:46
17:49:01.701 -> Water_Temp.:101.41, PWM:46
17:49:02.490 -> Water_Temp.:101.41, PWM:46
17:49:03.249 -> Water_Temp.:101.41, PWM:46
17:49:04.044 -> Water_Temp.:101.41, PWM:46
17:49:04.802 -> Water_Temp.:101.41, PWM:46
17:49:05.593 -> Water_Temp.:101.41, PWM:46
17:49:06.382 -> Water_Temp.:101.41, PWM:46
17:49:07.139 -> Water_Temp.:101.41, PWM:46
17:49:07.937 -> Water_Temp.:101.41, PWM:46
17:49:08.727 -> Water_Temp.:101.41, PWM:46
17:49:09.490 -> Water_Temp.:101.41, PWM:46
17:49:10.289 -> Water_Temp.:101.41, PWM:46
17:49:11.052 -> Water_Temp.:101.41, PWM:46
17:49:11.849 -> Water_Temp.:101.41, PWM:46
17:49:12.603 -> Water_Temp.:101.30, PWM:46
17:49:13.398 -> Water_Temp.:101.41, PWM:46
17:49:14.189 -> Water_Temp.:101.41, PWM:46
17:49:14.945 -> Water_Temp.:101.41, PWM:46
17:49:15.742 -> Water_Temp.:101.30, PWM:46
17:49:16.504 -> Water_Temp.:101.41, PWM:46
17:49:17.295 -> Water_Temp.:101.41, PWM:46
17:49:18.091 -> Water_Temp.:101.41, PWM:46
17:49:18.851 -> Water_Temp.:101.41, PWM:46
17:49:20.401 -> Water_Temp.:101.41, PWM:46
17:49:21.161 -> Water_Temp.:101.41, PWM:46
17:49:21.959 -> Water_Temp.:101.41, PWM:46
17:49:22.717 -> Water_Temp.:101.41, PWM:46
17:49:23.511 -> Water_Temp.:101.41, PWM:46
17:49:24.302 -> Water_Temp.:101.52, PWM:46
17:49:25.059 -> Water_Temp.:101.52, PWM:46
17:49:25.855 -> Water_Temp.:101.52, PWM:46
17:49:26.649 -> Water_Temp.:101.52, PWM:46
17:49:27.409 -> Water_Temp.:101.52, PWM:46
17:49:28.208 -> Water_Temp.:101.52, PWM:46
17:49:28.969 -> Water_Temp.:101.52, PWM:46
17:49:29.763 -> Water_Temp.:101.52, PWM:46
17:49:30.521 -> Water_Temp.:101.52, PWM:46
17:49:31.310 -> Water_Temp.:101.52, PWM:46
17:49:32.104 -> Water_Temp.:101.52, PWM:46
17:49:32.858 -> Water_Temp.:101.52, PWM:46
17:49:33.655 -> Water_Temp.:101.41, PWM:46
17:49:34.446 -> Water_Temp.:101.41, PWM:46
17:49:35.204 -> Water_Temp.:101.41, PWM:46
17:49:36.001 -> Water_Temp.:101.41, PWM:46
17:49:36.758 -> Water_Temp.:101.41, PWM:46
17:49:37.553 -> Water_Temp.:101.41, PWM:46
17:49:38.348 -> Water_Temp.:101.41, PWM:46
17:49:39.109 -> Water_Temp.:101.41, PWM:46
17:49:39.903 -> Water_Temp.:101.52, PWM:46
17:49:40.661 -> Water_Temp.:101.52, PWM:46
17:49:41.452 -> Water_Temp.:101.52, PWM:46
17:49:42.243 -> Water_Temp.:101.52, PWM:46
17:49:43.004 -> Water_Temp.:101.52, PWM:46
17:49:43.799 -> Water_Temp.:101.52, PWM:46
17:49:44.591 -> Water_Temp.:101.52, PWM:46
17:49:45.352 -> Water_Temp.:101.52, PWM:46
17:49:46.145 -> Water_Temp.:101.52, PWM:46
17:49:46.907 -> Water_Temp.:101.52, PWM:46
17:49:47.697 -> Water_Temp.:101.52, PWM:46
17:49:48.492 -> Water_Temp.:101.52, PWM:46
17:49:49.251 -> Water_Temp.:101.52, PWM:46
17:49:50.047 -> Water_Temp.:101.52, PWM:46
17:49:50.810 -> Water_Temp.:101.52, PWM:46
17:49:51.600 -> Water_Temp.:101.52, PWM:46
17:49:52.360 -> Water_Temp.:101.52, PWM:46
17:49:53.153 -> Water_Temp.:101.52, PWM:46
17:49:53.944 -> Water_Temp.:101.52, PWM:46
17:49:54.705 -> Water_Temp.:101.41, PWM:46
17:49:56.266 -> Water_Temp.:101.52, PWM:46
17:49:57.029 -> Water_Temp.:101.52, PWM:46
17:49:57.827 -> Water_Temp.:101.52, PWM:46
17:49:58.589 -> Water_Temp.:101.52, PWM:46
17:49:59.382 -> Water_Temp.:101.52, PWM:46
17:50:00.172 -> Water_Temp.:101.52, PWM:46
17:50:00.932 -> Water_Temp.:101.52, PWM:46
17:50:01.720 -> Water_Temp.:101.52, PWM:46
17:50:02.512 -> Water_Temp.:101.52, PWM:46
17:50:03.271 -> Water_Temp.:101.52, PWM:46
17:50:04.827 -> Water_Temp.:101.64, PWM:46
17:50:05.625 -> Water_Temp.:101.64, PWM:46
17:50:06.384 -> Water_Temp.:101.64, PWM:46
17:50:07.178 -> Water_Temp.:101.64, PWM:46
17:50:07.940 -> Water_Temp.:101.64, PWM:46
17:50:08.734 -> Water_Temp.:101.64, PWM:46
17:50:09.526 -> Water_Temp.:101.64, PWM:46
17:50:10.288 -> Water_Temp.:101.64, PWM:46
17:50:11.079 -> Water_Temp.:101.64, PWM:46
17:50:11.838 -> Water_Temp.:101.64, PWM:46
17:50:12.635 -> Water_Temp.:101.64, PWM:46
17:50:13.428 -> Water_Temp.:101.64, PWM:46
17:50:14.188 -> Water_Temp.:101.64, PWM:46
17:50:14.983 -> Water_Temp.:101.64, PWM:46
17:50:15.742 -> Water_Temp.:101.64, PWM:46
17:50:16.535 -> Water_Temp.:101.64, PWM:46
17:50:17.325 -> Water_Temp.:101.64, PWM:46
17:50:18.082 -> Water_Temp.:101.64, PWM:46
17:50:18.875 -> Water_Temp.:101.64, PWM:46
17:50:19.666 -> Water_Temp.:101.64, PWM:46
17:50:20.428 -> Water_Temp.:101.64, PWM:46
17:50:21.223 -> Water_Temp.:101.64, PWM:46
17:50:22.779 -> Water_Temp.:101.64, PWM:46
17:50:24.298 -> Water_Temp.:101.64, PWM:46
17:50:25.094 -> Water_Temp.:101.64, PWM:46
17:50:25.849 -> Water_Temp.:101.75, PWM:46
17:50:26.646 -> Water_Temp.:101.75, PWM:46
17:50:27.442 -> Water_Temp.:101.75, PWM:46
17:50:28.201 -> Water_Temp.:101.75, PWM:46
17:50:28.997 -> Water_Temp.:101.75, PWM:46
17:50:30.514 -> Water_Temp.:101.75, PWM:46
17:50:31.312 -> Water_Temp.:101.75, PWM:46
17:50:32.071 -> Water_Temp.:101.75, PWM:46
17:50:32.862 -> Water_Temp.:101.75, PWM:46
17:50:33.657 -> Water_Temp.:101.75, PWM:46
17:50:34.417 -> Water_Temp.:101.75, PWM:46
17:50:35.210 -> Water_Temp.:101.75, PWM:46
17:50:35.973 -> Water_Temp.:101.75, PWM:46
17:50:36.771 -> Water_Temp.:101.75, PWM:46
17:50:38.291 -> Water_Temp.:101.64, PWM:46
17:50:39.091 -> Water_Temp.:101.64, PWM:46
17:50:39.856 -> Water_Temp.:101.64, PWM:46
17:50:40.650 -> Water_Temp.:101.64, PWM:46
17:50:41.408 -> Water_Temp.:101.64, PWM:46
17:50:42.199 -> Water_Temp.:101.64, PWM:46
17:50:42.994 -> Water_Temp.:101.64, PWM:46
17:50:43.753 -> Water_Temp.:101.64, PWM:46
17:50:44.545 -> Water_Temp.:101.64, PWM:46
17:50:45.335 -> Water_Temp.:101.64, PWM:46
17:50:46.093 -> Water_Temp.:101.64, PWM:46
17:50:46.887 -> Water_Temp.:101.64, PWM:46
17:50:47.645 -> Water_Temp.:101.75, PWM:46
17:50:48.442 -> Water_Temp.:101.64, PWM:46
17:50:49.204 -> Water_Temp.:101.75, PWM:46
17:50:49.995 -> Water_Temp.:101.75, PWM:46
17:50:50.790 -> Water_Temp.:101.75, PWM:46
17:50:51.547 -> Water_Temp.:101.75, PWM:46
17:50:52.342 -> Water_Temp.:101.75, PWM:46
17:50:53.135 -> Water_Temp.:101.75, PWM:46
17:50:53.893 -> Water_Temp.:101.75, PWM:46
17:50:54.687 -> Water_Temp.:101.75, PWM:46
17:50:55.479 -> Water_Temp.:101.75, PWM:46
17:50:56.239 -> Water_Temp.:101.75, PWM:46
17:50:57.031 -> Water_Temp.:101.75, PWM:46
17:50:57.788 -> Water_Temp.:101.75, PWM:46
17:50:58.584 -> Water_Temp.:101.75, PWM:46
17:51:00.137 -> Water_Temp.:101.75, PWM:46
17:51:00.897 -> Water_Temp.:101.75, PWM:46
17:51:01.696 -> Water_Temp.:101.75, PWM:46
17:51:02.454 -> Water_Temp.:101.75, PWM:46
17:51:03.247 -> Water_Temp.:101.75, PWM:46
17:51:04.043 -> Water_Temp.:101.75, PWM:46
17:51:04.799 -> Water_Temp.:101.75, PWM:46
17:51:05.586 -> Water_Temp.:101.75, PWM:46
17:51:06.381 -> Water_Temp.:101.75, PWM:46
17:51:07.139 -> Water_Temp.:101.75, PWM:46
17:51:07.930 -> Water_Temp.:101.75, PWM:46
17:51:08.721 -> Water_Temp.:101.75, PWM:46
17:51:09.479 -> Water_Temp.:101.75, PWM:46
17:51:10.269 -> Water_Temp.:101.75, PWM:46
17:51:11.063 -> Water_Temp.:101.75, PWM:46
17:51:11.818 -> Water_Temp.:101.75, PWM:46
17:51:12.614 -> Water_Temp.:101.75, PWM:46
17:51:13.372 -> Water_Temp.:101.75, PWM:46
17:51:14.169 -> Water_Temp.:101.75, PWM:46
17:51:14.963 -> Water_Temp.:101.75, PWM:46
17:51:15.724 -> Water_Temp.:103.32, PWM:46
17:51:16.514 -> Water_Temp.:105.35, PWM:46
17:51:18.061 -> Water_Temp.:102.65, PWM:46
17:51:18.853 -> Water_Temp.:102.42, PWM:46
17:51:19.613 -> Water_Temp.:102.31, PWM:46
17:51:20.408 -> Water_Temp.:102.31, PWM:46
17:51:21.167 -> Water_Temp.:102.20, PWM:46
17:51:21.961 -> Water_Temp.:102.20, PWM:46
17:51:22.757 -> Water_Temp.:102.20, PWM:46
17:51:23.518 -> Water_Temp.:102.20, PWM:46
17:51:24.314 -> Water_Temp.:102.20, PWM:46
17:51:25.073 -> Water_Temp.:102.20, PWM:46
17:51:25.863 -> Water_Temp.:102.20, PWM:46
17:51:26.655 -> Water_Temp.:102.20, PWM:46
17:51:27.416 -> Water_Temp.:102.09, PWM:46
17:51:28.211 -> Water_Temp.:102.09, PWM:46
17:51:28.993 -> Water_Temp.:102.09, PWM:46
17:51:29.754 -> Water_Temp.:102.09, PWM:46
17:51:30.546 -> Water_Temp.:102.09, PWM:46
17:51:31.338 -> Water_Temp.:101.97, PWM:46
17:51:32.100 -> Water_Temp.:101.97, PWM:46
17:51:32.892 -> Water_Temp.:101.97, PWM:46
17:51:33.657 -> Water_Temp.:101.97, PWM:46
17:51:34.452 -> Water_Temp.:101.97, PWM:46
17:51:35.212 -> Water_Temp.:101.97, PWM:46
17:51:36.006 -> Water_Temp.:101.97, PWM:46
17:51:36.800 -> Water_Temp.:101.97, PWM:46
17:51:37.559 -> Water_Temp.:101.97, PWM:46
17:51:38.355 -> Water_Temp.:101.97, PWM:46
17:51:39.118 -> Water_Temp.:101.97, PWM:46
17:51:39.914 -> Water_Temp.:101.97, PWM:46
17:51:40.670 -> Water_Temp.:101.97, PWM:46
17:51:41.466 -> Water_Temp.:101.97, PWM:46
17:51:42.260 -> Water_Temp.:101.97, PWM:46
17:51:43.017 -> Water_Temp.:101.97, PWM:46
17:51:43.808 -> Water_Temp.:101.86, PWM:46
17:51:44.567 -> Water_Temp.:101.86, PWM:46
17:51:45.363 -> Water_Temp.:101.86, PWM:46
17:51:46.153 -> Water_Temp.:101.86, PWM:46
17:51:46.916 -> Water_Temp.:101.86, PWM:46
17:51:47.708 -> Water_Temp.:101.86, PWM:46
17:51:48.465 -> Water_Temp.:101.86, PWM:46
17:51:49.256 -> Water_Temp.:101.86, PWM:46
17:51:50.048 -> Water_Temp.:101.86, PWM:46
17:51:50.808 -> Water_Temp.:101.86, PWM:46
17:51:51.603 -> Water_Temp.:101.86, PWM:46
17:51:52.396 -> Water_Temp.:101.86, PWM:46
17:51:53.155 -> Water_Temp.:101.86, PWM:46
17:51:53.950 -> Water_Temp.:101.86, PWM:46
17:51:54.714 -> Water_Temp.:101.86, PWM:46
17:51:55.508 -> Water_Temp.:101.86, PWM:46
17:51:56.301 -> Water_Temp.:101.86, PWM:46
17:51:57.062 -> Water_Temp.:101.86, PWM:46
17:51:57.850 -> Water_Temp.:101.86, PWM:46
17:51:58.611 -> Water_Temp.:101.86, PWM:46
17:52:00.157 -> Water_Temp.:101.86, PWM:46
17:52:00.947 -> Water_Temp.:101.86, PWM:46
17:52:01.739 -> Water_Temp.:101.86, PWM:46
17:52:02.498 -> Water_Temp.:101.86, PWM:46
17:52:03.291 -> Water_Temp.:101.86, PWM:46
17:52:04.081 -> Water_Temp.:101.86, PWM:46
17:52:04.838 -> Water_Temp.:101.86, PWM:46
17:52:05.628 -> Water_Temp.:101.86, PWM:46
17:52:06.422 -> Water_Temp.:101.86, PWM:46
17:52:07.179 -> Water_Temp.:101.86, PWM:46
17:52:07.968 -> Water_Temp.:101.86, PWM:46
17:52:08.764 -> Water_Temp.:101.86, PWM:46
17:52:09.526 -> Water_Temp.:101.86, PWM:46
17:52:10.324 -> Water_Temp.:101.86, PWM:46
17:52:11.085 -> Water_Temp.:101.86, PWM:46
17:52:11.881 -> Water_Temp.:101.86, PWM:46
17:52:13.440 -> Water_Temp.:101.86, PWM:46
17:52:14.197 -> Water_Temp.:101.86, PWM:46
17:52:14.995 -> Water_Temp.:101.86, PWM:46
17:52:15.757 -> Water_Temp.:101.86, PWM:46
17:52:16.555 -> Water_Temp.:101.86, PWM:46
17:52:17.313 -> Water_Temp.:101.86, PWM:46
17:52:18.107 -> Water_Temp.:101.86, PWM:46
17:52:18.906 -> Water_Temp.:101.86, PWM:46
17:52:19.671 -> Water_Temp.:101.86, PWM:46
17:52:20.436 -> Water_Temp.:101.86, PWM:46
17:52:21.232 -> Water_Temp.:101.86, PWM:46
17:52:22.027 -> Water_Temp.:101.86, PWM:46
17:52:22.785 -> Water_Temp.:101.86, PWM:46
17:52:23.580 -> Water_Temp.:101.86, PWM:46
17:52:25.101 -> Water_Temp.:101.97, PWM:46
17:52:25.896 -> Water_Temp.:101.97, PWM:46
17:52:26.653 -> Water_Temp.:101.97, PWM:46
17:52:27.441 -> Water_Temp.:101.97, PWM:46
17:52:28.206 -> Water_Temp.:101.97, PWM:46
17:52:28.997 -> Water_Temp.:102.09, PWM:46
17:52:29.789 -> Water_Temp.:101.97, PWM:46
17:52:30.549 -> Water_Temp.:102.09, PWM:46
17:52:31.339 -> Water_Temp.:102.09, PWM:46
17:52:32.133 -> Water_Temp.:102.09, PWM:46
17:52:32.893 -> Water_Temp.:102.09, PWM:46
17:52:33.689 -> Water_Temp.:101.97, PWM:46
17:52:34.478 -> Water_Temp.:101.97, PWM:46
17:52:35.239 -> Water_Temp.:101.97, PWM:46
17:52:36.035 -> Water_Temp.:101.97, PWM:46
17:52:36.794 -> Water_Temp.:101.97, PWM:46
17:52:37.590 -> Water_Temp.:101.97, PWM:46
17:52:38.353 -> Water_Temp.:101.97, PWM:46
17:52:39.148 -> Water_Temp.:101.97, PWM:46
17:52:40.665 -> Water_Temp.:101.97, PWM:46
17:52:41.460 -> Water_Temp.:101.97, PWM:46
17:52:42.250 -> Water_Temp.:101.97, PWM:46
17:52:43.011 -> Water_Temp.:101.97, PWM:46
17:52:43.807 -> Water_Temp.:101.97, PWM:46
17:52:44.564 -> Water_Temp.:101.97, PWM:46
17:52:45.359 -> Water_Temp.:101.97, PWM:46
17:52:46.123 -> Water_Temp.:101.97, PWM:46
17:52:46.918 -> Water_Temp.:101.97, PWM:46
17:52:47.712 -> Water_Temp.:101.97, PWM:46
17:52:49.267 -> Water_Temp.:101.97, PWM:46
17:52:50.027 -> Water_Temp.:101.97, PWM:46
17:52:50.821 -> Water_Temp.:101.97, PWM:46
17:52:51.615 -> Water_Temp.:101.97, PWM:46
17:52:52.373 -> Water_Temp.:101.97, PWM:46
17:52:53.170 -> Water_Temp.:101.97, PWM:46
17:52:53.928 -> Water_Temp.:101.97, PWM:46
17:52:54.722 -> Water_Temp.:101.97, PWM:46
17:52:55.485 -> Water_Temp.:101.97, PWM:46
17:52:56.275 -> Water_Temp.:101.97, PWM:46
17:52:57.067 -> Water_Temp.:101.97, PWM:46
17:52:58.616 -> Water_Temp.:101.97, PWM:46
17:52:59.380 -> Water_Temp.:101.97, PWM:46
17:53:00.170 -> Water_Temp.:101.97, PWM:46
17:53:00.927 -> Water_Temp.:101.97, PWM:46
17:53:01.722 -> Water_Temp.:101.97, PWM:46
17:53:02.517 -> Water_Temp.:101.97, PWM:46
17:53:03.279 -> Water_Temp.:101.97, PWM:46
17:53:04.071 -> Water_Temp.:101.97, PWM:46
17:53:04.825 -> Water_Temp.:101.97, PWM:46
17:53:05.622 -> Water_Temp.:101.97, PWM:46
17:53:06.383 -> Water_Temp.:101.97, PWM:46
17:53:07.179 -> Water_Temp.:101.97, PWM:46
17:53:07.977 -> Water_Temp.:101.97, PWM:46
17:53:08.735 -> Water_Temp.:101.97, PWM:46
17:53:09.531 -> Water_Temp.:101.97, PWM:46
17:53:10.289 -> Water_Temp.:101.97, PWM:46
17:53:11.080 -> Water_Temp.:101.97, PWM:46
17:53:11.876 -> Water_Temp.:101.97, PWM:46
17:53:12.634 -> Water_Temp.:101.97, PWM:46
17:53:13.431 -> Water_Temp.:101.97, PWM:46
17:53:14.193 -> Water_Temp.:101.97, PWM:46
17:53:14.987 -> Water_Temp.:101.97, PWM:46
17:53:15.779 -> Water_Temp.:101.97, PWM:46
17:53:16.540 -> Water_Temp.:101.97, PWM:50
17:53:17.335 -> Water_Temp.:102.09, PWM:51
17:53:18.858 -> Water_Temp.:102.09, PWM:51
17:53:19.623 -> Water_Temp.:101.97, PWM:50
17:53:20.420 -> Water_Temp.:101.97, PWM:49
17:53:21.183 -> Water_Temp.:102.09, PWM:49
17:53:21.977 -> Water_Temp.:102.09, PWM:49
17:53:22.772 -> Water_Temp.:101.97, PWM:49
17:53:23.536 -> Water_Temp.:101.97, PWM:49
17:53:24.324 -> Water_Temp.:102.09, PWM:49
17:53:25.118 -> Water_Temp.:102.09, PWM:49
17:53:25.882 -> Water_Temp.:101.97, PWM:49
17:53:26.643 -> Water_Temp.:102.09, PWM:49
17:53:28.233 -> Water_Temp.:102.09, PWM:49
17:53:28.994 -> Water_Temp.:101.97, PWM:49
17:53:29.792 -> Water_Temp.:101.97, PWM:49
17:53:30.553 -> Water_Temp.:102.09, PWM:49
17:53:31.349 -> Water_Temp.:102.09, PWM:49
17:53:32.112 -> Water_Temp.:101.97, PWM:49
17:53:32.904 -> Water_Temp.:101.97, PWM:49
17:53:33.665 -> Water_Temp.:101.97, PWM:49
17:53:34.457 -> Water_Temp.:101.97, PWM:49
17:53:35.251 -> Water_Temp.:101.97, PWM:49
17:53:36.007 -> Water_Temp.:101.97, PWM:49
17:53:36.799 -> Water_Temp.:101.97, PWM:49
17:53:37.588 -> Water_Temp.:101.97, PWM:49
17:53:38.349 -> Water_Temp.:101.97, PWM:49
17:53:39.140 -> Water_Temp.:101.97, PWM:49
17:53:39.932 -> Water_Temp.:101.97, PWM:49
17:53:40.694 -> Water_Temp.:101.97, PWM:49
17:53:41.490 -> Water_Temp.:101.97, PWM:49
17:53:42.249 -> Water_Temp.:102.09, PWM:49
17:53:43.041 -> Water_Temp.:102.09, PWM:49
17:53:43.835 -> Water_Temp.:102.09, PWM:49
17:53:44.597 -> Water_Temp.:102.09, PWM:49
17:53:45.393 -> Water_Temp.:102.09, PWM:49
17:53:46.150 -> Water_Temp.:102.09, PWM:49
17:53:46.941 -> Water_Temp.:102.09, PWM:49
17:53:47.704 -> Water_Temp.:102.09, PWM:49
17:53:48.496 -> Water_Temp.:102.09, PWM:49
17:53:49.293 -> Water_Temp.:102.09, PWM:49
17:53:50.049 -> Water_Temp.:102.09, PWM:49
17:53:50.843 -> Water_Temp.:102.09, PWM:49
17:53:51.604 -> Water_Temp.:102.09, PWM:49
17:53:52.394 -> Water_Temp.:102.09, PWM:49
17:53:53.194 -> Water_Temp.:102.09, PWM:49
17:53:53.957 -> Water_Temp.:102.09, PWM:49
17:53:55.508 -> Water_Temp.:101.97, PWM:49
17:53:56.303 -> Water_Temp.:101.97, PWM:49
17:53:57.061 -> Water_Temp.:101.97, PWM:49
17:53:57.854 -> Water_Temp.:101.97, PWM:49
17:53:58.646 -> Water_Temp.:101.97, PWM:49
17:53:59.410 -> Water_Temp.:101.97, PWM:49
17:54:00.200 -> Water_Temp.:101.97, PWM:49
17:54:00.994 -> Water_Temp.:102.09, PWM:49
17:54:01.753 -> Water_Temp.:101.97, PWM:49
17:54:02.548 -> Water_Temp.:101.97, PWM:49
17:54:03.309 -> Water_Temp.:101.97, PWM:49
17:54:04.108 -> Water_Temp.:101.97, PWM:49
17:54:04.868 -> Water_Temp.:101.97, PWM:49
17:54:05.664 -> Water_Temp.:101.97, PWM:49
17:54:06.423 -> Water_Temp.:101.97, PWM:49
17:54:07.217 -> Water_Temp.:101.97, PWM:49
17:54:08.007 -> Water_Temp.:101.97, PWM:49
17:54:08.770 -> Water_Temp.:101.97, PWM:49
17:54:10.329 -> Water_Temp.:101.97, PWM:49
17:54:11.126 -> Water_Temp.:101.97, PWM:49
17:54:12.679 -> Water_Temp.:102.09, PWM:49
17:54:13.440 -> Water_Temp.:102.09, PWM:49
17:54:14.231 -> Water_Temp.:102.09, PWM:49
17:54:15.027 -> Water_Temp.:102.09, PWM:49
17:54:15.790 -> Water_Temp.:102.09, PWM:49
17:54:16.589 -> Water_Temp.:102.09, PWM:49
17:54:18.144 -> Water_Temp.:102.09, PWM:49
17:54:18.903 -> Water_Temp.:102.09, PWM:49
17:54:20.455 -> Water_Temp.:102.09, PWM:49
17:54:21.250 -> Water_Temp.:102.09, PWM:49
17:54:22.010 -> Water_Temp.:102.09, PWM:49
17:54:22.799 -> Water_Temp.:102.09, PWM:49
17:54:23.593 -> Water_Temp.:102.09, PWM:49
17:54:24.355 -> Water_Temp.:102.09, PWM:49
17:54:25.145 -> Water_Temp.:102.09, PWM:49
17:54:25.935 -> Water_Temp.:102.09, PWM:49
17:54:26.698 -> Water_Temp.:102.09, PWM:49
17:54:27.461 -> Water_Temp.:102.20, PWM:49
17:54:28.254 -> Water_Temp.:102.20, PWM:49
17:54:29.045 -> Water_Temp.:102.20, PWM:49
17:54:29.806 -> Water_Temp.:102.20, PWM:49
17:54:30.596 -> Water_Temp.:102.20, PWM:49
17:54:31.387 -> Water_Temp.:102.20, PWM:49
17:54:32.149 -> Water_Temp.:102.20, PWM:49
17:54:32.946 -> Water_Temp.:102.20, PWM:49
17:54:34.497 -> Water_Temp.:102.20, PWM:49
17:54:35.262 -> Water_Temp.:102.20, PWM:49
17:54:36.057 -> Water_Temp.:102.20, PWM:49
17:54:36.849 -> Water_Temp.:102.20, PWM:49
17:54:37.609 -> Water_Temp.:102.20, PWM:49
17:54:38.400 -> Water_Temp.:102.20, PWM:49
17:54:39.188 -> Water_Temp.:102.20, PWM:49
17:54:39.948 -> Water_Temp.:102.20, PWM:49
17:54:40.742 -> Water_Temp.:102.20, PWM:49
17:54:41.500 -> Water_Temp.:102.20, PWM:49
17:54:42.295 -> Water_Temp.:102.20, PWM:49
17:54:43.092 -> Water_Temp.:102.20, PWM:49
17:54:43.851 -> Water_Temp.:102.20, PWM:49
17:54:44.643 -> Water_Temp.:102.20, PWM:49
17:54:45.404 -> Water_Temp.:103.77, PWM:49
17:54:46.198 -> Water_Temp.:105.80, PWM:49
17:54:46.989 -> Water_Temp.:103.44, PWM:49
17:54:47.746 -> Water_Temp.:103.10, PWM:49
17:54:48.540 -> Water_Temp.:102.87, PWM:49
17:54:49.336 -> Water_Temp.:102.76, PWM:49
17:54:50.092 -> Water_Temp.:102.65, PWM:49
17:54:50.886 -> Water_Temp.:102.65, PWM:49
17:54:51.646 -> Water_Temp.:102.65, PWM:49
17:54:52.438 -> Water_Temp.:102.54, PWM:49
17:54:53.989 -> Water_Temp.:102.54, PWM:49
17:54:54.746 -> Water_Temp.:102.54, PWM:49
17:54:55.540 -> Water_Temp.:102.54, PWM:49
17:54:56.331 -> Water_Temp.:102.54, PWM:49
17:54:57.090 -> Water_Temp.:102.42, PWM:49
17:54:57.881 -> Water_Temp.:102.42, PWM:49
17:54:58.643 -> Water_Temp.:102.42, PWM:49
17:54:59.434 -> Water_Temp.:102.42, PWM:49
17:55:00.980 -> Water_Temp.:102.42, PWM:49
17:55:02.530 -> Water_Temp.:102.31, PWM:49
17:55:03.328 -> Water_Temp.:102.31, PWM:49
17:55:04.086 -> Water_Temp.:102.31, PWM:49
17:55:04.880 -> Water_Temp.:102.31, PWM:49
17:55:05.642 -> Water_Temp.:102.31, PWM:49
17:55:06.434 -> Water_Temp.:102.31, PWM:49
17:55:07.225 -> Water_Temp.:102.31, PWM:49
17:55:07.986 -> Water_Temp.:102.31, PWM:49
17:55:08.782 -> Water_Temp.:102.31, PWM:49
17:55:09.544 -> Water_Temp.:102.31, PWM:49
17:55:10.343 -> Water_Temp.:102.31, PWM:49
17:55:11.106 -> Water_Temp.:102.31, PWM:49
17:55:11.902 -> Water_Temp.:102.31, PWM:49
17:55:14.215 -> Water_Temp.:102.31, PWM:49
17:55:15.768 -> Water_Temp.:102.31, PWM:49
17:55:16.562 -> Water_Temp.:102.31, PWM:49
17:55:17.325 -> Water_Temp.:102.31, PWM:49
17:55:18.118 -> Water_Temp.:102.20, PWM:49
17:55:18.878 -> Water_Temp.:102.31, PWM:49
17:55:20.466 -> Water_Temp.:102.31, PWM:49
17:55:21.222 -> Water_Temp.:102.31, PWM:49
17:55:22.779 -> Water_Temp.:102.31, PWM:49
17:55:23.571 -> Water_Temp.:102.31, PWM:49
17:55:24.330 -> Water_Temp.:102.31, PWM:49
17:55:25.127 -> Water_Temp.:102.31, PWM:49
17:55:25.889 -> Water_Temp.:102.31, PWM:49
17:55:26.686 -> Water_Temp.:102.31, PWM:49
17:55:27.444 -> Water_Temp.:102.31, PWM:49
17:55:29.002 -> Water_Temp.:102.31, PWM:49
17:55:29.799 -> Water_Temp.:102.31, PWM:49
17:55:30.591 -> Water_Temp.:102.31, PWM:49
17:55:31.353 -> Water_Temp.:102.31, PWM:49
17:55:32.146 -> Water_Temp.:102.31, PWM:49
17:55:32.904 -> Water_Temp.:102.31, PWM:49
17:55:33.702 -> Water_Temp.:102.31, PWM:49
17:55:34.466 -> Water_Temp.:102.31, PWM:49
17:55:35.259 -> Water_Temp.:102.31, PWM:49
17:55:36.020 -> Water_Temp.:102.31, PWM:49
17:55:36.808 -> Water_Temp.:102.31, PWM:49
17:55:37.603 -> Water_Temp.:102.31, PWM:49
17:55:38.365 -> Water_Temp.:102.31, PWM:49
17:55:39.159 -> Water_Temp.:102.31, PWM:49
17:55:39.918 -> Water_Temp.:102.31, PWM:49
17:55:40.715 -> Water_Temp.:102.31, PWM:49
17:55:41.507 -> Water_Temp.:102.31, PWM:49
17:55:42.265 -> Water_Temp.:102.31, PWM:49
17:55:43.060 -> Water_Temp.:102.31, PWM:49
17:55:43.817 -> Water_Temp.:102.31, PWM:49
17:55:44.612 -> Water_Temp.:102.31, PWM:49
17:55:45.410 -> Water_Temp.:102.31, PWM:49
17:55:46.169 -> Water_Temp.:102.31, PWM:49
17:55:47.726 -> Water_Temp.:102.31, PWM:49
17:55:48.487 -> Water_Temp.:102.31, PWM:49
17:55:49.275 -> Water_Temp.:102.31, PWM:49
17:55:50.070 -> Water_Temp.:102.31, PWM:49
17:55:50.831 -> Water_Temp.:102.31, PWM:49
17:55:51.625 -> Water_Temp.:102.31, PWM:49
17:55:52.385 -> Water_Temp.:102.31, PWM:49
17:55:53.175 -> Water_Temp.:102.31, PWM:49
17:55:53.969 -> Water_Temp.:102.31, PWM:49
17:55:54.727 -> Water_Temp.:102.31, PWM:49
17:55:55.516 -> Water_Temp.:102.31, PWM:49
17:55:56.310 -> Water_Temp.:102.31, PWM:49
17:55:57.069 -> Water_Temp.:102.31, PWM:49
17:55:58.620 -> Water_Temp.:102.31, PWM:49
17:55:59.380 -> Water_Temp.:102.31, PWM:49
17:56:00.175 -> Water_Temp.:102.31, PWM:49
17:56:00.937 -> Water_Temp.:102.31, PWM:49
17:56:01.730 -> Water_Temp.:102.31, PWM:49
17:56:03.280 -> Water_Temp.:102.31, PWM:49
17:56:04.073 -> Water_Temp.:102.31, PWM:49
17:56:04.832 -> Water_Temp.:102.31, PWM:49
17:56:05.629 -> Water_Temp.:102.31, PWM:49
17:56:06.425 -> Water_Temp.:102.31, PWM:49
17:56:07.184 -> Water_Temp.:102.31, PWM:49
17:56:07.978 -> Water_Temp.:102.42, PWM:49
17:56:08.740 -> Water_Temp.:102.31, PWM:49
17:56:09.536 -> Water_Temp.:102.42, PWM:49
17:56:10.297 -> Water_Temp.:102.42, PWM:49
17:56:11.090 -> Water_Temp.:102.42, PWM:49
17:56:11.881 -> Water_Temp.:102.42, PWM:49
17:56:12.643 -> Water_Temp.:102.42, PWM:49
17:56:14.194 -> Water_Temp.:102.42, PWM:49
17:56:14.955 -> Water_Temp.:102.42, PWM:49
17:56:15.753 -> Water_Temp.:102.42, PWM:49
17:56:16.515 -> Water_Temp.:102.42, PWM:49
17:56:17.310 -> Water_Temp.:102.42, PWM:49
17:56:18.074 -> Water_Temp.:102.42, PWM:49
17:56:18.867 -> Water_Temp.:102.42, PWM:49
17:56:19.626 -> Water_Temp.:102.42, PWM:49
17:56:20.422 -> Water_Temp.:102.42, PWM:49
17:56:21.214 -> Water_Temp.:102.42, PWM:49
17:56:21.977 -> Water_Temp.:102.42, PWM:49
17:56:23.535 -> Water_Temp.:102.42, PWM:49
17:56:24.332 -> Water_Temp.:102.42, PWM:49
17:56:25.095 -> Water_Temp.:102.42, PWM:49
17:56:25.896 -> Water_Temp.:102.42, PWM:49
17:56:26.655 -> Water_Temp.:102.42, PWM:49
17:56:27.444 -> Water_Temp.:102.42, PWM:49
17:56:28.238 -> Water_Temp.:102.54, PWM:49
17:56:28.999 -> Water_Temp.:102.54, PWM:49
17:56:29.795 -> Water_Temp.:102.54, PWM:49
17:56:31.312 -> Water_Temp.:102.54, PWM:49
17:56:32.106 -> Water_Temp.:102.54, PWM:49
17:56:32.868 -> Water_Temp.:102.54, PWM:49
17:56:33.661 -> Water_Temp.:102.54, PWM:49
17:56:34.459 -> Water_Temp.:102.54, PWM:49
17:56:35.224 -> Water_Temp.:102.54, PWM:49
17:56:36.020 -> Water_Temp.:102.54, PWM:49
17:56:36.778 -> Water_Temp.:102.54, PWM:49
17:56:37.571 -> Water_Temp.:102.54, PWM:49
17:56:38.332 -> Water_Temp.:102.54, PWM:49
17:56:39.130 -> Water_Temp.:102.54, PWM:49
17:56:39.896 -> Water_Temp.:102.54, PWM:49
17:56:40.689 -> Water_Temp.:102.42, PWM:49
17:56:41.448 -> Water_Temp.:102.42, PWM:49
17:56:42.243 -> Water_Temp.:102.42, PWM:49
17:56:43.039 -> Water_Temp.:102.42, PWM:49
17:56:43.802 -> Water_Temp.:102.42, PWM:49
17:56:44.598 -> Water_Temp.:102.42, PWM:49
17:56:45.358 -> Water_Temp.:102.42, PWM:49
17:56:46.153 -> Water_Temp.:102.42, PWM:49
17:56:46.913 -> Water_Temp.:102.42, PWM:49
17:56:47.709 -> Water_Temp.:102.42, PWM:49
17:56:48.500 -> Water_Temp.:102.42, PWM:49
17:56:49.261 -> Water_Temp.:102.42, PWM:49
17:56:50.052 -> Water_Temp.:102.42, PWM:49
17:56:50.812 -> Water_Temp.:102.42, PWM:49
17:56:52.371 -> Water_Temp.:102.42, PWM:49
17:56:53.163 -> Water_Temp.:102.42, PWM:49
17:56:53.924 -> Water_Temp.:102.42, PWM:49
17:56:54.716 -> Water_Temp.:102.42, PWM:49
17:56:55.482 -> Water_Temp.:102.42, PWM:49
17:56:56.280 -> Water_Temp.:102.42, PWM:49
17:56:57.043 -> Water_Temp.:102.42, PWM:49
17:56:57.839 -> Water_Temp.:102.42, PWM:49
17:56:58.599 -> Water_Temp.:102.42, PWM:49
17:56:59.397 -> Water_Temp.:102.42, PWM:49
17:57:00.151 -> Water_Temp.:102.42, PWM:49
17:57:00.946 -> Water_Temp.:102.42, PWM:49
17:57:01.737 -> Water_Temp.:102.42, PWM:49
17:57:02.497 -> Water_Temp.:102.42, PWM:49
17:57:03.293 -> Water_Temp.:102.42, PWM:49
17:57:04.056 -> Water_Temp.:102.42, PWM:49
17:57:04.853 -> Water_Temp.:102.42, PWM:49
17:57:05.616 -> Water_Temp.:102.42, PWM:49
17:57:06.412 -> Water_Temp.:102.42, PWM:49
17:57:07.173 -> Water_Temp.:102.42, PWM:49
17:57:07.967 -> Water_Temp.:102.42, PWM:49
17:57:08.759 -> Water_Temp.:102.42, PWM:49
17:57:10.316 -> Water_Temp.:102.42, PWM:49
17:57:11.078 -> Water_Temp.:102.42, PWM:49
17:57:11.874 -> Water_Temp.:102.42, PWM:49
17:57:13.429 -> Water_Temp.:102.42, PWM:49
17:57:14.187 -> Water_Temp.:102.42, PWM:49
17:57:14.980 -> Water_Temp.:102.42, PWM:49
17:57:15.774 -> Water_Temp.:102.42, PWM:49
17:57:16.530 -> Water_Temp.:102.42, PWM:49
17:57:17.321 -> Water_Temp.:102.42, PWM:49
17:57:18.112 -> Water_Temp.:102.42, PWM:49
17:57:18.876 -> Water_Temp.:102.42, PWM:49
17:57:19.669 -> Water_Temp.:102.42, PWM:49
17:57:20.428 -> Water_Temp.:102.42, PWM:49
17:57:21.221 -> Water_Temp.:102.42, PWM:49
17:57:22.014 -> Water_Temp.:102.42, PWM:49
17:57:22.777 -> Water_Temp.:102.42, PWM:49
17:57:23.574 -> Water_Temp.:102.42, PWM:49
17:57:24.334 -> Water_Temp.:102.42, PWM:49
17:57:25.894 -> Water_Temp.:102.42, PWM:49
17:57:26.691 -> Water_Temp.:102.42, PWM:49
17:57:27.453 -> Water_Temp.:102.42, PWM:49
17:57:28.248 -> Water_Temp.:102.42, PWM:49
17:57:29.010 -> Water_Temp.:102.42, PWM:49
17:57:29.805 -> Water_Temp.:102.54, PWM:49
17:57:30.564 -> Water_Temp.:102.42, PWM:49
17:57:31.354 -> Water_Temp.:102.42, PWM:49
17:57:32.149 -> Water_Temp.:102.54, PWM:49
17:57:32.911 -> Water_Temp.:102.54, PWM:49
17:57:33.707 -> Water_Temp.:102.54, PWM:49
17:57:34.506 -> Water_Temp.:102.54, PWM:49
17:57:35.265 -> Water_Temp.:102.54, PWM:49
17:57:36.822 -> Water_Temp.:102.54, PWM:49
17:57:37.580 -> Water_Temp.:102.54, PWM:49
17:57:38.366 -> Water_Temp.:102.54, PWM:49
17:57:39.164 -> Water_Temp.:102.54, PWM:49
17:57:39.927 -> Water_Temp.:102.54, PWM:49
17:57:40.724 -> Water_Temp.:102.54, PWM:49
17:57:41.486 -> Water_Temp.:102.54, PWM:49
17:57:42.282 -> Water_Temp.:102.54, PWM:49
17:57:43.042 -> Water_Temp.:102.54, PWM:49
17:57:43.834 -> Water_Temp.:102.54, PWM:49
17:57:44.627 -> Water_Temp.:102.54, PWM:49
17:57:45.391 -> Water_Temp.:102.54, PWM:49
17:57:46.185 -> Water_Temp.:102.54, PWM:49
17:57:47.710 -> Water_Temp.:102.54, PWM:49
17:57:48.504 -> Water_Temp.:102.54, PWM:49
17:57:49.267 -> Water_Temp.:102.54, PWM:49
17:57:50.819 -> Water_Temp.:102.54, PWM:49
17:57:51.615 -> Water_Temp.:102.54, PWM:49
17:57:52.377 -> Water_Temp.:102.54, PWM:49
17:57:53.172 -> Water_Temp.:102.54, PWM:49
17:57:53.931 -> Water_Temp.:102.54, PWM:49
17:57:54.724 -> Water_Temp.:102.54, PWM:49
17:57:55.524 -> Water_Temp.:102.54, PWM:49
17:57:56.283 -> Water_Temp.:102.65, PWM:49
17:57:57.079 -> Water_Temp.:102.65, PWM:49
17:57:57.838 -> Water_Temp.:102.65, PWM:49
17:57:58.628 -> Water_Temp.:102.65, PWM:49
17:57:59.393 -> Water_Temp.:102.65, PWM:49
17:58:00.184 -> Water_Temp.:102.65, PWM:49
17:58:00.982 -> Water_Temp.:102.65, PWM:49
17:58:01.742 -> Water_Temp.:102.65, PWM:49
17:58:02.533 -> Water_Temp.:102.65, PWM:49
17:58:04.086 -> Water_Temp.:102.65, PWM:49
17:58:04.876 -> Water_Temp.:102.65, PWM:49
17:58:05.639 -> Water_Temp.:102.65, PWM:49
17:58:06.432 -> Water_Temp.:102.65, PWM:49
17:58:07.196 -> Water_Temp.:102.65, PWM:49
17:58:07.989 -> Water_Temp.:102.65, PWM:49
17:58:08.782 -> Water_Temp.:102.65, PWM:49
17:58:09.545 -> Water_Temp.:102.65, PWM:49
17:58:10.339 -> Water_Temp.:102.65, PWM:49
17:58:11.099 -> Water_Temp.:102.76, PWM:49
17:58:11.895 -> Water_Temp.:102.65, PWM:49
17:58:12.658 -> Water_Temp.:102.76, PWM:49
17:58:13.451 -> Water_Temp.:102.76, PWM:49
17:58:14.241 -> Water_Temp.:102.76, PWM:49
17:58:14.998 -> Water_Temp.:102.76, PWM:49
17:58:15.794 -> Water_Temp.:102.76, PWM:49
17:58:16.553 -> Water_Temp.:102.76, PWM:49
17:58:17.342 -> Water_Temp.:102.76, PWM:49
17:58:18.139 -> Water_Temp.:102.76, PWM:49
17:58:18.900 -> Water_Temp.:102.76, PWM:49
17:58:19.697 -> Water_Temp.:102.76, PWM:49
17:58:20.456 -> Water_Temp.:102.76, PWM:49
17:58:21.249 -> Water_Temp.:102.65, PWM:49
17:58:22.012 -> Water_Temp.:102.76, PWM:49
17:58:23.599 -> Water_Temp.:102.65, PWM:49
17:58:24.363 -> Water_Temp.:102.65, PWM:49
17:58:25.162 -> Water_Temp.:102.65, PWM:49
17:58:26.678 -> Water_Temp.:102.65, PWM:49
17:58:27.470 -> Water_Temp.:102.65, PWM:49
17:58:28.262 -> Water_Temp.:102.65, PWM:49
17:58:29.022 -> Water_Temp.:102.65, PWM:49
17:58:29.821 -> Water_Temp.:102.76, PWM:49
17:58:30.582 -> Water_Temp.:102.76, PWM:49
17:58:31.376 -> Water_Temp.:102.76, PWM:49
17:58:32.138 -> Water_Temp.:102.76, PWM:49
17:58:32.934 -> Water_Temp.:102.76, PWM:49
17:58:33.697 -> Water_Temp.:102.76, PWM:49
17:58:35.257 -> Water_Temp.:102.76, PWM:49
17:58:36.055 -> Water_Temp.:102.76, PWM:49
17:58:36.814 -> Water_Temp.:102.76, PWM:49
17:58:38.403 -> Water_Temp.:102.76, PWM:49
17:58:39.164 -> Water_Temp.:102.76, PWM:49
17:58:39.954 -> Water_Temp.:102.76, PWM:49
17:58:40.717 -> Water_Temp.:102.76, PWM:49
17:58:41.512 -> Water_Temp.:102.76, PWM:49
17:58:42.277 -> Water_Temp.:102.76, PWM:49
17:58:43.074 -> Water_Temp.:102.76, PWM:49
17:58:43.834 -> Water_Temp.:102.76, PWM:49
17:58:44.628 -> Water_Temp.:102.76, PWM:49
17:58:45.388 -> Water_Temp.:102.87, PWM:49
17:58:46.185 -> Water_Temp.:102.87, PWM:49
17:58:46.977 -> Water_Temp.:102.87, PWM:49
17:58:47.741 -> Water_Temp.:102.87, PWM:49
17:58:48.530 -> Water_Temp.:102.87, PWM:49
17:58:49.323 -> Water_Temp.:102.87, PWM:49
17:58:50.078 -> Water_Temp.:102.87, PWM:49
17:58:51.626 -> Water_Temp.:102.76, PWM:49
17:58:52.386 -> Water_Temp.:102.76, PWM:49
17:58:53.182 -> Water_Temp.:102.76, PWM:49
17:58:53.974 -> Water_Temp.:102.76, PWM:49
17:58:54.732 -> Water_Temp.:102.76, PWM:49
17:58:55.525 -> Water_Temp.:102.76, PWM:49
17:58:56.281 -> Water_Temp.:102.76, PWM:49
17:58:57.075 -> Water_Temp.:102.76, PWM:49
17:58:57.869 -> Water_Temp.:102.76, PWM:49
17:58:58.630 -> Water_Temp.:102.76, PWM:49
17:58:59.424 -> Water_Temp.:102.76, PWM:49
17:59:00.215 -> Water_Temp.:102.76, PWM:49
17:59:00.973 -> Water_Temp.:102.76, PWM:49
17:59:01.765 -> Water_Temp.:102.76, PWM:49
17:59:02.525 -> Water_Temp.:102.76, PWM:49
17:59:03.317 -> Water_Temp.:102.76, PWM:49
17:59:04.111 -> Water_Temp.:102.76, PWM:49
17:59:04.867 -> Water_Temp.:102.76, PWM:49
17:59:05.662 -> Water_Temp.:102.76, PWM:49
17:59:06.423 -> Water_Temp.:102.76, PWM:49
17:59:07.220 -> Water_Temp.:102.76, PWM:49
17:59:08.012 -> Water_Temp.:102.76, PWM:49
17:59:08.771 -> Water_Temp.:102.76, PWM:49
17:59:09.571 -> Water_Temp.:102.76, PWM:49
17:59:10.327 -> Water_Temp.:102.76, PWM:49
17:59:11.122 -> Water_Temp.:102.76, PWM:49
17:59:11.882 -> Water_Temp.:102.76, PWM:49
17:59:12.674 -> Water_Temp.:102.76, PWM:49
17:59:13.467 -> Water_Temp.:102.76, PWM:49
17:59:14.228 -> Water_Temp.:102.76, PWM:49
17:59:15.021 -> Water_Temp.:102.76, PWM:49
17:59:15.813 -> Water_Temp.:102.76, PWM:49
17:59:16.569 -> Water_Temp.:102.76, PWM:49
17:59:17.363 -> Water_Temp.:102.76, PWM:49
17:59:18.123 -> Water_Temp.:102.76, PWM:49
17:59:18.917 -> Water_Temp.:102.76, PWM:48
17:59:19.681 -> Water_Temp.:102.76, PWM:49
17:59:20.474 -> Water_Temp.:102.76, PWM:49
17:59:21.266 -> Water_Temp.:102.76, PWM:49
17:59:22.027 -> Water_Temp.:102.76, PWM:49
17:59:22.822 -> Water_Temp.:102.76, PWM:48
17:59:23.586 -> Water_Temp.:102.76, PWM:49
17:59:24.382 -> Water_Temp.:102.76, PWM:49
17:59:25.171 -> Water_Temp.:102.87, PWM:49
17:59:25.933 -> Water_Temp.:102.87, PWM:49
17:59:27.481 -> Water_Temp.:102.87, PWM:49
17:59:28.278 -> Water_Temp.:102.87, PWM:49
17:59:29.041 -> Water_Temp.:102.87, PWM:49
17:59:29.834 -> Water_Temp.:102.87, PWM:49
17:59:30.590 -> Water_Temp.:102.87, PWM:49
17:59:31.383 -> Water_Temp.:102.87, PWM:49
17:59:32.144 -> Water_Temp.:102.87, PWM:49
17:59:34.455 -> Water_Temp.:102.99, PWM:49
17:59:35.248 -> Water_Temp.:102.99, PWM:49
17:59:36.041 -> Water_Temp.:102.99, PWM:49
17:59:36.804 -> Water_Temp.:102.99, PWM:49
17:59:39.147 -> Water_Temp.:102.99, PWM:49
17:59:39.910 -> Water_Temp.:102.87, PWM:49
17:59:40.706 -> Water_Temp.:102.87, PWM:49
17:59:41.467 -> Water_Temp.:104.45, PWM:49
17:59:42.258 -> Water_Temp.:106.47, PWM:49
17:59:43.051 -> Water_Temp.:104.11, PWM:49
17:59:43.812 -> Water_Temp.:103.77, PWM:49
17:59:44.605 -> Water_Temp.:103.55, PWM:49
17:59:45.402 -> Water_Temp.:103.44, PWM:49
17:59:46.159 -> Water_Temp.:103.32, PWM:49
17:59:46.953 -> Water_Temp.:103.32, PWM:49
17:59:47.712 -> Water_Temp.:103.21, PWM:49
17:59:48.507 -> Water_Temp.:103.21, PWM:49
17:59:49.298 -> Water_Temp.:103.10, PWM:49
17:59:50.056 -> Water_Temp.:103.10, PWM:49
17:59:51.605 -> Water_Temp.:103.10, PWM:49
17:59:52.398 -> Water_Temp.:102.99, PWM:49
17:59:53.159 -> Water_Temp.:102.99, PWM:49
17:59:53.948 -> Water_Temp.:102.99, PWM:49
17:59:54.712 -> Water_Temp.:102.99, PWM:49
17:59:55.499 -> Water_Temp.:102.99, PWM:49
17:59:56.293 -> Water_Temp.:102.99, PWM:49
17:59:57.057 -> Water_Temp.:102.99, PWM:49
17:59:57.847 -> Water_Temp.:102.99, PWM:49
17:59:58.638 -> Water_Temp.:102.99, PWM:49
17:59:59.398 -> Water_Temp.:102.99, PWM:49
18:00:00.951 -> Water_Temp.:102.99, PWM:49
18:00:01.707 -> Water_Temp.:102.99, PWM:49
18:00:02.504 -> Water_Temp.:102.99, PWM:49
18:00:03.265 -> Water_Temp.:102.87, PWM:49
18:00:04.061 -> Water_Temp.:102.87, PWM:49
18:00:04.854 -> Water_Temp.:102.87, PWM:49
18:00:05.616 -> Water_Temp.:102.87, PWM:49
18:00:06.397 -> Water_Temp.:102.87, PWM:49
18:00:07.186 -> Water_Temp.:102.87, PWM:49
18:00:07.944 -> Water_Temp.:102.87, PWM:49
18:00:08.737 -> Water_Temp.:102.87, PWM:49
18:00:09.503 -> Water_Temp.:102.87, PWM:49
18:00:10.300 -> Water_Temp.:102.87, PWM:49
18:00:11.091 -> Water_Temp.:102.87, PWM:49
18:00:11.854 -> Water_Temp.:102.87, PWM:49
18:00:12.651 -> Water_Temp.:102.87, PWM:49
18:00:13.407 -> Water_Temp.:102.87, PWM:49
18:00:14.198 -> Water_Temp.:102.87, PWM:49
18:00:14.991 -> Water_Temp.:102.87, PWM:49
18:00:16.544 -> Water_Temp.:102.87, PWM:49
18:00:17.334 -> Water_Temp.:102.87, PWM:49
18:00:18.097 -> Water_Temp.:102.87, PWM:49
18:00:18.862 -> Water_Temp.:102.87, PWM:49
18:00:19.659 -> Water_Temp.:102.87, PWM:49
18:00:20.423 -> Water_Temp.:102.87, PWM:49
18:00:21.217 -> Water_Temp.:102.87, PWM:49
18:00:22.013 -> Water_Temp.:102.87, PWM:49
18:00:22.772 -> Water_Temp.:102.87, PWM:49
18:00:23.569 -> Water_Temp.:102.87, PWM:49
18:00:24.323 -> Water_Temp.:102.87, PWM:49
18:00:25.118 -> Water_Temp.:102.87, PWM:49
18:00:25.914 -> Water_Temp.:102.87, PWM:49
18:00:26.673 -> Water_Temp.:102.87, PWM:49
18:00:27.465 -> Water_Temp.:102.99, PWM:49
18:00:28.228 -> Water_Temp.:102.99, PWM:49
18:00:29.020 -> Water_Temp.:102.99, PWM:49
18:00:29.810 -> Water_Temp.:102.99, PWM:49
18:00:30.566 -> Water_Temp.:102.99, PWM:49
18:00:31.352 -> Water_Temp.:102.87, PWM:49
18:00:32.142 -> Water_Temp.:102.87, PWM:49
18:00:32.933 -> Water_Temp.:102.87, PWM:49
18:00:34.448 -> Water_Temp.:102.87, PWM:49
18:00:35.244 -> Water_Temp.:102.87, PWM:49
18:00:36.037 -> Water_Temp.:102.87, PWM:49
18:00:36.796 -> Water_Temp.:102.87, PWM:49
18:00:37.587 -> Water_Temp.:102.87, PWM:49
18:00:38.349 -> Water_Temp.:102.87, PWM:49
18:00:39.139 -> Water_Temp.:102.87, PWM:49
18:00:41.453 -> Water_Temp.:102.87, PWM:49
18:00:43.043 -> Water_Temp.:102.87, PWM:49
18:00:43.804 -> Water_Temp.:102.87, PWM:49
18:00:44.602 -> Water_Temp.:102.87, PWM:49
18:00:45.366 -> Water_Temp.:102.87, PWM:49
18:00:46.132 -> Water_Temp.:102.87, PWM:49
18:00:46.927 -> Water_Temp.:102.87, PWM:49
18:00:47.718 -> Water_Temp.:102.87, PWM:49
18:00:49.242 -> Water_Temp.:102.87, PWM:49
18:00:50.037 -> Water_Temp.:102.87, PWM:49
18:00:50.797 -> Water_Temp.:102.87, PWM:49
18:00:51.593 -> Water_Temp.:102.87, PWM:49
18:00:52.367 -> Water_Temp.:102.87, PWM:49
18:00:53.160 -> Water_Temp.:102.87, PWM:49
18:00:54.712 -> Water_Temp.:102.87, PWM:49
18:00:55.507 -> Water_Temp.:102.99, PWM:49
18:00:56.267 -> Water_Temp.:102.87, PWM:49
18:00:57.062 -> Water_Temp.:102.87, PWM:49
18:00:57.821 -> Water_Temp.:102.99, PWM:49
18:00:58.618 -> Water_Temp.:102.99, PWM:49
18:01:00.139 -> Water_Temp.:102.99, PWM:49
18:01:00.934 -> Water_Temp.:102.87, PWM:49
18:01:01.700 -> Water_Temp.:102.99, PWM:49
18:01:02.491 -> Water_Temp.:102.99, PWM:49
18:01:03.250 -> Water_Temp.:102.87, PWM:49
18:01:04.045 -> Water_Temp.:102.99, PWM:49
18:01:04.843 -> Water_Temp.:102.99, PWM:49
18:01:05.604 -> Water_Temp.:102.99, PWM:49
18:01:06.395 -> Water_Temp.:102.87, PWM:49
18:01:07.153 -> Water_Temp.:102.87, PWM:49
18:01:07.946 -> Water_Temp.:102.99, PWM:49
18:01:08.742 -> Water_Temp.:102.87, PWM:49
18:01:09.505 -> Water_Temp.:102.99, PWM:49
18:01:11.056 -> Water_Temp.:106.59, PWM:49
18:01:11.849 -> Water_Temp.:104.22, PWM:49
18:01:12.612 -> Water_Temp.:103.77, PWM:49
18:01:13.411 -> Water_Temp.:103.66, PWM:49
18:01:14.204 -> Water_Temp.:103.55, PWM:49
18:01:14.967 -> Water_Temp.:103.44, PWM:49
18:01:15.762 -> Water_Temp.:103.32, PWM:49
18:01:16.523 -> Water_Temp.:103.32, PWM:49
18:01:17.316 -> Water_Temp.:103.21, PWM:49
18:01:18.076 -> Water_Temp.:103.21, PWM:49
18:01:18.873 -> Water_Temp.:103.21, PWM:49
18:01:19.632 -> Water_Temp.:103.10, PWM:49
18:01:20.422 -> Water_Temp.:103.10, PWM:49
18:01:21.217 -> Water_Temp.:103.10, PWM:49
18:01:21.976 -> Water_Temp.:103.10, PWM:49
18:01:22.771 -> Water_Temp.:103.10, PWM:49
18:01:23.529 -> Water_Temp.:103.10, PWM:49
18:01:24.322 -> Water_Temp.:103.10, PWM:49
18:01:25.121 -> Water_Temp.:103.10, PWM:49
18:01:25.882 -> Water_Temp.:103.10, PWM:49
18:01:26.678 -> Water_Temp.:103.10, PWM:48
18:01:27.431 -> Water_Temp.:103.10, PWM:49
18:01:28.227 -> Water_Temp.:103.10, PWM:49
18:01:29.020 -> Water_Temp.:103.10, PWM:49
18:01:29.783 -> Water_Temp.:103.10, PWM:49
18:01:30.574 -> Water_Temp.:103.10, PWM:49
18:01:32.131 -> Water_Temp.:103.10, PWM:49
18:01:32.890 -> Water_Temp.:103.10, PWM:49
18:01:33.682 -> Water_Temp.:103.10, PWM:49
18:01:34.472 -> Water_Temp.:103.10, PWM:49
18:01:35.232 -> Water_Temp.:103.10, PWM:49
18:01:36.028 -> Water_Temp.:103.10, PWM:49
18:01:36.818 -> Water_Temp.:103.10, PWM:49
18:01:39.124 -> Water_Temp.:103.10, PWM:49
18:01:39.881 -> Water_Temp.:103.10, PWM:49
18:01:40.673 -> Water_Temp.:103.10, PWM:49
18:01:41.463 -> Water_Temp.:103.10, PWM:49
18:01:42.223 -> Water_Temp.:103.10, PWM:49
18:01:43.014 -> Water_Temp.:102.99, PWM:49
18:01:43.776 -> Water_Temp.:102.99, PWM:49
18:01:44.572 -> Water_Temp.:102.99, PWM:49
18:01:45.365 -> Water_Temp.:102.99, PWM:49
18:01:46.123 -> Water_Temp.:102.99, PWM:48
18:01:47.677 -> Water_Temp.:102.99, PWM:49
18:01:48.439 -> Water_Temp.:102.99, PWM:49
18:01:49.229 -> Water_Temp.:102.99, PWM:49
18:01:50.023 -> Water_Temp.:102.99, PWM:49
18:01:50.782 -> Water_Temp.:102.99, PWM:49
18:01:51.579 -> Water_Temp.:102.99, PWM:49
18:01:52.373 -> Water_Temp.:102.87, PWM:49
18:01:53.133 -> Water_Temp.:102.87, PWM:49
18:01:53.894 -> Water_Temp.:102.87, PWM:49
18:01:54.692 -> Water_Temp.:102.87, PWM:49
18:01:55.487 -> Water_Temp.:102.99, PWM:49
18:01:57.043 -> Water_Temp.:102.99, PWM:49
18:01:57.802 -> Water_Temp.:102.87, PWM:49
18:01:58.598 -> Water_Temp.:102.99, PWM:49
18:01:59.357 -> Water_Temp.:102.99, PWM:49
18:02:00.148 -> Water_Temp.:102.99, PWM:49
18:02:00.941 -> Water_Temp.:102.99, PWM:49
18:02:01.703 -> Water_Temp.:102.99, PWM:49
18:02:02.493 -> Water_Temp.:102.99, PWM:49
18:02:03.288 -> Water_Temp.:102.99, PWM:49
18:02:04.048 -> Water_Temp.:102.99, PWM:49
18:02:04.840 -> Water_Temp.:102.99, PWM:49
18:02:05.601 -> Water_Temp.:102.99, PWM:49
18:02:06.395 -> Water_Temp.:102.99, PWM:49
18:02:07.185 -> Water_Temp.:102.99, PWM:49
18:02:07.947 -> Water_Temp.:102.99, PWM:49
18:02:09.500 -> Water_Temp.:102.99, PWM:49
18:02:10.293 -> Water_Temp.:102.99, PWM:49
18:02:11.084 -> Water_Temp.:102.99, PWM:49
18:02:11.846 -> Water_Temp.:102.99, PWM:49
18:02:12.639 -> Water_Temp.:102.99, PWM:49
18:02:13.400 -> Water_Temp.:102.99, PWM:49
18:02:14.191 -> Water_Temp.:102.99, PWM:49
18:02:14.984 -> Water_Temp.:102.99, PWM:49
18:02:15.746 -> Water_Temp.:102.99, PWM:49
18:02:16.535 -> Water_Temp.:102.99, PWM:49
18:02:18.052 -> Water_Temp.:102.99, PWM:49
18:02:18.845 -> Water_Temp.:102.99, PWM:49
18:02:19.639 -> Water_Temp.:102.99, PWM:49
18:02:20.398 -> Water_Temp.:102.99, PWM:49
18:02:21.191 -> Water_Temp.:102.99, PWM:49
18:02:21.953 -> Water_Temp.:102.99, PWM:49
18:02:22.742 -> Water_Temp.:102.99, PWM:49
18:02:25.094 -> Water_Temp.:103.10, PWM:49
18:02:25.857 -> Water_Temp.:103.10, PWM:49
18:02:26.655 -> Water_Temp.:103.10, PWM:49
18:02:27.418 -> Water_Temp.:103.10, PWM:49
18:02:28.210 -> Water_Temp.:103.10, PWM:49
18:02:29.001 -> Water_Temp.:103.10, PWM:49
18:02:29.765 -> Water_Temp.:103.10, PWM:49
18:02:30.526 -> Water_Temp.:103.10, PWM:49
18:02:31.325 -> Water_Temp.:103.10, PWM:49
18:02:32.083 -> Water_Temp.:103.21, PWM:49
18:02:32.875 -> Water_Temp.:104.79, PWM:49
18:02:33.667 -> Water_Temp.:106.81, PWM:49
18:02:34.455 -> Water_Temp.:104.45, PWM:49
18:02:35.221 -> Water_Temp.:104.00, PWM:49
18:02:36.017 -> Water_Temp.:103.77, PWM:49
18:02:36.777 -> Water_Temp.:103.66, PWM:49
18:02:37.572 -> Water_Temp.:103.55, PWM:49
18:02:38.329 -> Water_Temp.:103.44, PWM:49
18:02:39.124 -> Water_Temp.:103.44, PWM:49
18:02:40.672 -> Water_Temp.:103.32, PWM:49
18:02:41.466 -> Water_Temp.:103.32, PWM:49
18:02:42.225 -> Water_Temp.:103.21, PWM:49
18:02:43.016 -> Water_Temp.:103.21, PWM:49
18:02:43.807 -> Water_Temp.:103.21, PWM:49
18:02:44.567 -> Water_Temp.:103.21, PWM:49
18:02:45.366 -> Water_Temp.:103.21, PWM:49
18:02:46.917 -> Water_Temp.:103.21, PWM:49
18:02:47.713 -> Water_Temp.:103.10, PWM:49
18:02:48.472 -> Water_Temp.:103.10, PWM:49
18:02:49.997 -> Water_Temp.:103.10, PWM:49
18:02:50.793 -> Water_Temp.:103.10, PWM:49
18:02:51.584 -> Water_Temp.:103.10, PWM:49
18:02:52.343 -> Water_Temp.:103.10, PWM:49
18:02:53.134 -> Water_Temp.:103.10, PWM:49
18:02:53.932 -> Water_Temp.:102.99, PWM:49
18:02:54.687 -> Water_Temp.:102.99, PWM:49
18:02:55.485 -> Water_Temp.:102.99, PWM:49
18:02:56.245 -> Water_Temp.:103.10, PWM:49
18:02:57.038 -> Water_Temp.:102.99, PWM:49
18:02:57.801 -> Water_Temp.:102.99, PWM:49
18:02:58.596 -> Water_Temp.:103.10, PWM:49
18:02:59.387 -> Water_Temp.:103.10, PWM:49
18:03:00.147 -> Water_Temp.:103.10, PWM:49
18:03:00.938 -> Water_Temp.:103.10, PWM:49
18:03:01.697 -> Water_Temp.:103.10, PWM:49
18:03:02.496 -> Water_Temp.:103.10, PWM:49
18:03:03.257 -> Water_Temp.:103.10, PWM:49
18:03:04.052 -> Water_Temp.:103.10, PWM:49
18:03:04.843 -> Water_Temp.:103.10, PWM:49
18:03:05.604 -> Water_Temp.:103.10, PWM:49
18:03:06.392 -> Water_Temp.:103.10, PWM:49
18:03:07.156 -> Water_Temp.:103.10, PWM:49
18:03:07.955 -> Water_Temp.:103.10, PWM:49
18:03:08.717 -> Water_Temp.:103.10, PWM:49
18:03:09.511 -> Water_Temp.:103.10, PWM:49
18:03:10.305 -> Water_Temp.:103.10, PWM:49
18:03:11.068 -> Water_Temp.:103.10, PWM:49
18:03:11.862 -> Water_Temp.:103.10, PWM:49
18:03:12.624 -> Water_Temp.:103.10, PWM:49
18:03:13.416 -> Water_Temp.:103.10, PWM:49
18:03:14.178 -> Water_Temp.:103.10, PWM:49
18:03:14.971 -> Water_Temp.:103.10, PWM:49
18:03:15.759 -> Water_Temp.:102.99, PWM:49
18:03:16.517 -> Water_Temp.:102.99, PWM:49
18:03:17.312 -> Water_Temp.:102.99, PWM:49
18:03:18.107 -> Water_Temp.:102.99, PWM:49
18:03:19.629 -> Water_Temp.:102.99, PWM:49
18:03:20.389 -> Water_Temp.:102.99, PWM:49
18:03:21.182 -> Water_Temp.:102.99, PWM:49
18:03:21.976 -> Water_Temp.:102.99, PWM:49
18:03:22.736 -> Water_Temp.:102.99, PWM:49
18:03:23.530 -> Water_Temp.:102.99, PWM:49
18:03:24.293 -> Water_Temp.:102.99, PWM:49
18:03:25.089 -> Water_Temp.:103.10, PWM:49
18:03:25.851 -> Water_Temp.:103.10, PWM:49
18:03:26.647 -> Water_Temp.:103.10, PWM:49
18:03:27.435 -> Water_Temp.:103.10, PWM:49
18:03:28.195 -> Water_Temp.:103.10, PWM:49
18:03:28.991 -> Water_Temp.:103.10, PWM:49
18:03:29.751 -> Water_Temp.:103.10, PWM:49
18:03:30.541 -> Water_Temp.:103.10, PWM:49
18:03:31.334 -> Water_Temp.:103.10, PWM:49
18:03:32.096 -> Water_Temp.:103.10, PWM:49
18:03:32.891 -> Water_Temp.:103.10, PWM:49
18:03:33.649 -> Water_Temp.:103.10, PWM:49
18:03:34.442 -> Water_Temp.:103.10, PWM:49
18:03:35.231 -> Water_Temp.:103.21, PWM:49
18:03:35.991 -> Water_Temp.:103.21, PWM:49
18:03:36.782 -> Water_Temp.:103.21, PWM:49
18:03:37.575 -> Water_Temp.:103.21, PWM:49
18:03:38.337 -> Water_Temp.:103.21, PWM:49
18:03:39.128 -> Water_Temp.:103.21, PWM:49
18:03:39.919 -> Water_Temp.:104.79, PWM:49
18:03:40.676 -> Water_Temp.:106.81, PWM:49
18:03:41.467 -> Water_Temp.:104.45, PWM:49
18:03:42.228 -> Water_Temp.:104.00, PWM:49
18:03:43.020 -> Water_Temp.:103.77, PWM:49
18:03:43.812 -> Water_Temp.:103.77, PWM:49
18:03:44.604 -> Water_Temp.:103.66, PWM:49
18:03:45.361 -> Water_Temp.:103.55, PWM:49
18:03:46.914 -> Water_Temp.:103.44, PWM:49
18:03:47.708 -> Water_Temp.:103.44, PWM:49
18:03:48.465 -> Water_Temp.:103.32, PWM:49
18:03:50.012 -> Water_Temp.:103.32, PWM:49
18:03:51.570 -> Water_Temp.:103.21, PWM:49
18:03:53.123 -> Water_Temp.:103.21, PWM:49
18:03:53.917 -> Water_Temp.:103.21, PWM:49
18:03:54.711 -> Water_Temp.:103.21, PWM:49
18:03:55.475 -> Water_Temp.:103.10, PWM:53
18:03:56.272 -> Water_Temp.:103.10, PWM:52
18:03:57.034 -> Water_Temp.:103.10, PWM:52
18:03:57.829 -> Water_Temp.:103.10, PWM:52
18:03:58.591 -> Water_Temp.:103.10, PWM:52
18:03:59.385 -> Water_Temp.:103.21, PWM:52
18:04:00.900 -> Water_Temp.:103.21, PWM:52
18:04:01.692 -> Water_Temp.:103.21, PWM:52
18:04:02.482 -> Water_Temp.:103.21, PWM:52
18:04:03.239 -> Water_Temp.:103.21, PWM:52
18:04:04.030 -> Water_Temp.:103.32, PWM:52
18:04:04.819 -> Water_Temp.:103.32, PWM:52
18:04:05.579 -> Water_Temp.:103.32, PWM:52
18:04:06.371 -> Water_Temp.:103.32, PWM:52
18:04:07.160 -> Water_Temp.:103.32, PWM:52
18:04:07.918 -> Water_Temp.:103.32, PWM:52
18:04:10.262 -> Water_Temp.:103.32, PWM:52
18:04:11.057 -> Water_Temp.:103.21, PWM:52
18:04:11.813 -> Water_Temp.:103.21, PWM:52
18:04:12.609 -> Water_Temp.:103.21, PWM:52
18:04:13.402 -> Water_Temp.:103.32, PWM:52
18:04:14.162 -> Water_Temp.:103.32, PWM:52
18:04:14.955 -> Water_Temp.:103.32, PWM:52
18:04:15.719 -> Water_Temp.:103.32, PWM:52
18:04:16.508 -> Water_Temp.:103.32, PWM:52
18:04:17.301 -> Water_Temp.:103.32, PWM:52
18:04:18.059 -> Water_Temp.:103.32, PWM:52
18:04:18.853 -> Water_Temp.:103.32, PWM:52
18:04:19.613 -> Water_Temp.:103.32, PWM:53
18:04:20.405 -> Water_Temp.:103.21, PWM:52
18:04:21.198 -> Water_Temp.:103.21, PWM:52
18:04:21.960 -> Water_Temp.:103.21, PWM:52
18:04:22.752 -> Water_Temp.:103.21, PWM:52
18:04:23.514 -> Water_Temp.:103.21, PWM:52
18:04:25.070 -> Water_Temp.:103.21, PWM:52
18:04:25.866 -> Water_Temp.:103.21, PWM:52
18:04:27.422 -> Water_Temp.:103.21, PWM:52
18:04:28.184 -> Water_Temp.:103.21, PWM:52
18:04:28.980 -> Water_Temp.:103.21, PWM:52
18:04:30.531 -> Water_Temp.:103.21, PWM:52
18:04:31.324 -> Water_Temp.:103.21, PWM:52
18:04:32.080 -> Water_Temp.:103.21, PWM:52
18:04:32.872 -> Water_Temp.:103.21, PWM:53
18:04:33.664 -> Water_Temp.:103.21, PWM:52
18:04:35.218 -> Water_Temp.:103.21, PWM:52
18:04:35.978 -> Water_Temp.:103.21, PWM:52
18:04:37.526 -> Water_Temp.:103.21, PWM:52
18:04:38.321 -> Water_Temp.:103.21, PWM:52
18:04:39.079 -> Water_Temp.:103.21, PWM:52
18:04:39.867 -> Water_Temp.:103.21, PWM:52
18:04:40.659 -> Water_Temp.:103.21, PWM:53
18:04:41.417 -> Water_Temp.:103.21, PWM:52
18:04:42.972 -> Water_Temp.:103.21, PWM:53
18:04:43.766 -> Water_Temp.:103.21, PWM:52
18:04:44.529 -> Water_Temp.:103.21, PWM:52
18:04:45.325 -> Water_Temp.:103.21, PWM:53
18:04:46.881 -> Water_Temp.:103.21, PWM:52
18:04:47.641 -> Water_Temp.:103.21, PWM:52
18:04:48.439 -> Water_Temp.:103.32, PWM:52
18:04:49.231 -> Water_Temp.:103.32, PWM:52
18:04:49.992 -> Water_Temp.:103.32, PWM:52
18:04:50.786 -> Water_Temp.:103.32, PWM:52
18:04:51.541 -> Water_Temp.:103.32, PWM:52
18:04:53.126 -> Water_Temp.:103.32, PWM:53
18:04:53.883 -> Water_Temp.:103.32, PWM:52
18:04:54.673 -> Water_Temp.:103.21, PWM:53
18:04:55.436 -> Water_Temp.:103.21, PWM:52
18:04:56.233 -> Water_Temp.:103.21, PWM:52
18:04:57.028 -> Water_Temp.:103.21, PWM:52
18:04:57.787 -> Water_Temp.:103.21, PWM:52
18:04:58.577 -> Water_Temp.:103.21, PWM:52
18:04:59.336 -> Water_Temp.:103.21, PWM:52
18:05:00.133 -> Water_Temp.:103.21, PWM:52
18:05:00.926 -> Water_Temp.:103.21, PWM:52
18:05:01.679 -> Water_Temp.:103.21, PWM:52
18:05:02.473 -> Water_Temp.:103.21, PWM:52
18:05:03.265 -> Water_Temp.:103.21, PWM:52
18:05:04.813 -> Water_Temp.:103.21, PWM:52
18:05:05.577 -> Water_Temp.:103.21, PWM:52
18:05:06.373 -> Water_Temp.:103.21, PWM:52
18:05:07.166 -> Water_Temp.:103.21, PWM:52
18:05:08.715 -> Water_Temp.:103.10, PWM:52
18:05:09.483 -> Water_Temp.:103.21, PWM:52
18:05:10.276 -> Water_Temp.:103.21, PWM:52
18:05:11.036 -> Water_Temp.:103.21, PWM:52
18:05:11.822 -> Water_Temp.:103.21, PWM:53
18:05:12.582 -> Water_Temp.:103.21, PWM:53
18:05:13.374 -> Water_Temp.:103.21, PWM:52
18:05:14.169 -> Water_Temp.:103.21, PWM:52
18:05:14.926 -> Water_Temp.:103.21, PWM:52
18:05:15.718 -> Water_Temp.:103.21, PWM:52
18:05:16.511 -> Water_Temp.:103.21, PWM:52
18:05:17.269 -> Water_Temp.:103.32, PWM:52
18:05:18.063 -> Water_Temp.:103.32, PWM:52
18:05:18.854 -> Water_Temp.:103.32, PWM:53
18:05:19.613 -> Water_Temp.:103.44, PWM:52
18:05:20.409 -> Water_Temp.:103.44, PWM:52
18:05:21.162 -> Water_Temp.:105.01, PWM:52
18:05:21.956 -> Water_Temp.:107.04, PWM:52
18:05:22.751 -> Water_Temp.:104.56, PWM:52
18:05:23.511 -> Water_Temp.:104.56, PWM:52
18:05:24.304 -> Water_Temp.:104.00, PWM:52
18:05:25.064 -> Water_Temp.:103.89, PWM:52
18:05:25.862 -> Water_Temp.:103.77, PWM:52
18:05:26.621 -> Water_Temp.:103.66, PWM:52
18:05:27.417 -> Water_Temp.:103.55, PWM:52
18:05:28.213 -> Water_Temp.:103.55, PWM:52
18:05:28.980 -> Water_Temp.:103.55, PWM:52
18:05:29.774 -> Water_Temp.:103.44, PWM:52
18:05:31.328 -> Water_Temp.:103.44, PWM:52
18:05:32.086 -> Water_Temp.:103.44, PWM:52
18:05:32.883 -> Water_Temp.:103.32, PWM:52
18:05:33.645 -> Water_Temp.:103.32, PWM:53
18:05:34.438 -> Water_Temp.:103.32, PWM:52
18:05:35.232 -> Water_Temp.:103.32, PWM:52
18:05:35.996 -> Water_Temp.:103.32, PWM:53
18:05:36.791 -> Water_Temp.:103.32, PWM:52
18:05:37.553 -> Water_Temp.:103.32, PWM:52
18:05:38.349 -> Water_Temp.:103.32, PWM:53
18:05:39.110 -> Water_Temp.:103.32, PWM:52
18:05:39.907 -> Water_Temp.:103.32, PWM:52
18:05:40.669 -> Water_Temp.:103.32, PWM:52
18:05:41.466 -> Water_Temp.:103.32, PWM:52
18:05:42.227 -> Water_Temp.:103.32, PWM:52
18:05:43.020 -> Water_Temp.:103.32, PWM:52
18:05:44.579 -> Water_Temp.:103.21, PWM:53
18:05:45.344 -> Water_Temp.:103.21, PWM:52
18:05:46.142 -> Water_Temp.:103.21, PWM:52
18:05:46.898 -> Water_Temp.:103.21, PWM:53
18:05:47.687 -> Water_Temp.:103.21, PWM:52
18:05:49.210 -> Water_Temp.:103.21, PWM:52
18:05:50.006 -> Water_Temp.:103.21, PWM:52
18:05:50.799 -> Water_Temp.:103.21, PWM:52
18:05:51.558 -> Water_Temp.:103.21, PWM:52
18:05:52.353 -> Water_Temp.:103.21, PWM:52
18:05:53.114 -> Water_Temp.:103.21, PWM:53
18:05:53.906 -> Water_Temp.:103.21, PWM:52
18:05:54.699 -> Water_Temp.:103.21, PWM:52
18:05:55.464 -> Water_Temp.:103.21, PWM:52
18:05:56.259 -> Water_Temp.:103.21, PWM:52
18:05:57.021 -> Water_Temp.:103.21, PWM:52
18:05:57.817 -> Water_Temp.:103.21, PWM:52
18:05:58.573 -> Water_Temp.:103.21, PWM:52
18:05:59.368 -> Water_Temp.:103.21, PWM:52
18:06:00.162 -> Water_Temp.:103.21, PWM:52
18:06:01.685 -> Water_Temp.:103.21, PWM:52
18:06:02.481 -> Water_Temp.:103.21, PWM:52
18:06:03.244 -> Water_Temp.:103.21, PWM:53
18:06:04.040 -> Water_Temp.:103.21, PWM:52
18:06:04.802 -> Water_Temp.:103.21, PWM:52
18:06:05.601 -> Water_Temp.:103.21, PWM:52
18:06:06.362 -> Water_Temp.:103.21, PWM:52
18:06:07.911 -> Water_Temp.:103.21, PWM:52
18:06:08.673 -> Water_Temp.:103.32, PWM:53
18:06:09.468 -> Water_Temp.:103.32, PWM:52
18:06:10.230 -> Water_Temp.:103.32, PWM:52
18:06:11.027 -> Water_Temp.:103.32, PWM:52
18:06:11.820 -> Water_Temp.:103.32, PWM:52
18:06:12.579 -> Water_Temp.:103.32, PWM:52
18:06:13.372 -> Water_Temp.:103.21, PWM:52
18:06:14.135 -> Water_Temp.:103.21, PWM:52
18:06:14.925 -> Water_Temp.:103.21, PWM:52
18:06:15.690 -> Water_Temp.:103.21, PWM:53
18:06:16.483 -> Water_Temp.:103.21, PWM:52
18:06:17.277 -> Water_Temp.:103.21, PWM:53
18:06:18.040 -> Water_Temp.:103.21, PWM:52
18:06:18.836 -> Water_Temp.:103.21, PWM:52
18:06:20.360 -> Water_Temp.:103.21, PWM:52
18:06:21.151 -> Water_Temp.:103.21, PWM:52
18:06:21.909 -> Water_Temp.:103.21, PWM:52
18:06:22.705 -> Water_Temp.:103.32, PWM:52
18:06:23.498 -> Water_Temp.:103.32, PWM:52
18:06:24.256 -> Water_Temp.:103.44, PWM:52
18:06:25.047 -> Water_Temp.:103.44, PWM:53
18:06:25.811 -> Water_Temp.:103.44, PWM:52
18:06:26.603 -> Water_Temp.:103.44, PWM:52
18:06:27.400 -> Water_Temp.:103.44, PWM:52
18:06:28.155 -> Water_Temp.:103.44, PWM:52
18:06:28.946 -> Water_Temp.:103.44, PWM:52
18:06:29.740 -> Water_Temp.:103.44, PWM:52
18:06:30.500 -> Water_Temp.:103.32, PWM:52
18:06:31.292 -> Water_Temp.:103.32, PWM:52
18:06:32.053 -> Water_Temp.:103.32, PWM:52
18:06:32.849 -> Water_Temp.:103.32, PWM:52
18:06:33.641 -> Water_Temp.:103.32, PWM:52
18:06:34.406 -> Water_Temp.:103.32, PWM:52
18:06:35.198 -> Water_Temp.:103.32, PWM:52
18:06:35.954 -> Water_Temp.:103.32, PWM:52
18:06:36.752 -> Water_Temp.:103.32, PWM:52
18:06:37.539 -> Water_Temp.:103.32, PWM:52
18:06:38.299 -> Water_Temp.:103.32, PWM:52
18:06:39.089 -> Water_Temp.:103.32, PWM:52
18:06:39.850 -> Water_Temp.:103.32, PWM:52
18:06:40.642 -> Water_Temp.:103.32, PWM:52
18:06:41.437 -> Water_Temp.:103.32, PWM:53
18:06:42.197 -> Water_Temp.:103.32, PWM:52
18:06:42.995 -> Water_Temp.:103.32, PWM:52
18:06:44.513 -> Water_Temp.:103.44, PWM:52
18:06:45.311 -> Water_Temp.:103.44, PWM:52
18:06:46.073 -> Water_Temp.:103.44, PWM:52
18:06:46.869 -> Water_Temp.:103.44, PWM:52
18:06:47.660 -> Water_Temp.:103.44, PWM:52
18:06:48.419 -> Water_Temp.:103.44, PWM:52
18:06:49.213 -> Water_Temp.:103.44, PWM:52
18:06:50.762 -> Water_Temp.:103.44, PWM:52
18:06:51.519 -> Water_Temp.:103.44, PWM:52
18:06:52.310 -> Water_Temp.:103.44, PWM:52
18:06:53.102 -> Water_Temp.:103.44, PWM:52
18:06:54.656 -> Water_Temp.:103.44, PWM:52
18:06:55.414 -> Water_Temp.:103.44, PWM:52
18:06:56.208 -> Water_Temp.:103.44, PWM:52
18:06:57.000 -> Water_Temp.:103.44, PWM:52
18:06:57.762 -> Water_Temp.:103.44, PWM:52
18:06:58.556 -> Water_Temp.:103.44, PWM:52
18:06:59.318 -> Water_Temp.:103.44, PWM:52
18:07:00.109 -> Water_Temp.:103.44, PWM:52
18:07:01.661 -> Water_Temp.:103.32, PWM:52
18:07:02.453 -> Water_Temp.:103.32, PWM:52
18:07:03.211 -> Water_Temp.:103.32, PWM:52
18:07:04.005 -> Water_Temp.:103.44, PWM:52
18:07:04.799 -> Water_Temp.:103.44, PWM:52
18:07:06.347 -> Water_Temp.:103.44, PWM:52
18:07:07.107 -> Water_Temp.:103.44, PWM:52
18:07:07.900 -> Water_Temp.:103.44, PWM:52
18:07:08.693 -> Water_Temp.:103.44, PWM:52
18:07:09.454 -> Water_Temp.:103.44, PWM:52
18:07:11.033 -> Water_Temp.:103.44, PWM:52
18:07:11.794 -> Water_Temp.:103.44, PWM:53
18:07:12.587 -> Water_Temp.:103.44, PWM:52
18:07:13.379 -> Water_Temp.:103.44, PWM:52
18:07:14.143 -> Water_Temp.:103.44, PWM:52
18:07:14.907 -> Water_Temp.:103.44, PWM:52
18:07:15.700 -> Water_Temp.:103.32, PWM:52
18:07:16.494 -> Water_Temp.:103.32, PWM:52
18:07:17.249 -> Water_Temp.:103.32, PWM:52
18:07:18.045 -> Water_Temp.:103.32, PWM:52
18:07:18.842 -> Water_Temp.:103.32, PWM:52
18:07:19.600 -> Water_Temp.:103.32, PWM:52
18:07:20.395 -> Water_Temp.:103.32, PWM:52
18:07:21.154 -> Water_Temp.:103.32, PWM:52
18:07:21.945 -> Water_Temp.:103.32, PWM:53
18:07:22.740 -> Water_Temp.:103.21, PWM:52
18:07:23.496 -> Water_Temp.:103.21, PWM:52
18:07:25.048 -> Water_Temp.:103.32, PWM:52
18:07:25.840 -> Water_Temp.:103.21, PWM:52
18:07:26.637 -> Water_Temp.:103.32, PWM:53
18:07:27.397 -> Water_Temp.:103.32, PWM:52
18:07:28.187 -> Water_Temp.:103.21, PWM:52
18:07:28.946 -> Water_Temp.:103.21, PWM:52
18:07:29.741 -> Water_Temp.:103.32, PWM:52
18:07:30.540 -> Water_Temp.:103.32, PWM:52
18:07:31.303 -> Water_Temp.:103.32, PWM:52
18:07:32.098 -> Water_Temp.:103.32, PWM:52
18:07:32.855 -> Water_Temp.:103.32, PWM:52
18:07:33.650 -> Water_Temp.:103.44, PWM:52
18:07:34.410 -> Water_Temp.:103.44, PWM:52
18:07:35.206 -> Water_Temp.:103.44, PWM:52
18:07:35.998 -> Water_Temp.:103.44, PWM:52
18:07:36.758 -> Water_Temp.:103.32, PWM:52
18:07:37.547 -> Water_Temp.:103.32, PWM:52
18:07:38.305 -> Water_Temp.:103.32, PWM:52
18:07:39.096 -> Water_Temp.:103.32, PWM:52
18:07:39.889 -> Water_Temp.:103.32, PWM:52
18:07:40.680 -> Water_Temp.:103.32, PWM:52
18:07:41.437 -> Water_Temp.:103.32, PWM:52
18:07:42.230 -> Water_Temp.:103.32, PWM:52
18:07:42.989 -> Water_Temp.:103.32, PWM:52
18:07:43.785 -> Water_Temp.:103.32, PWM:52
18:07:44.579 -> Water_Temp.:103.32, PWM:53
18:07:45.336 -> Water_Temp.:103.32, PWM:52
18:07:46.126 -> Water_Temp.:103.32, PWM:52
18:07:46.918 -> Water_Temp.:103.32, PWM:53
18:07:48.433 -> Water_Temp.:103.21, PWM:52
18:07:49.229 -> Water_Temp.:103.21, PWM:52
18:07:49.987 -> Water_Temp.:103.32, PWM:52
18:07:50.778 -> Water_Temp.:103.21, PWM:52
18:07:51.575 -> Water_Temp.:103.21, PWM:52
18:07:52.333 -> Water_Temp.:103.21, PWM:52
18:07:53.125 -> Water_Temp.:103.21, PWM:52
18:07:53.917 -> Water_Temp.:103.21, PWM:52
18:07:54.678 -> Water_Temp.:103.21, PWM:53
18:07:55.476 -> Water_Temp.:103.21, PWM:52
18:07:56.237 -> Water_Temp.:103.21, PWM:52
18:07:57.031 -> Water_Temp.:103.21, PWM:52
18:07:57.789 -> Water_Temp.:103.21, PWM:52
18:07:58.586 -> Water_Temp.:103.21, PWM:52
18:07:59.345 -> Water_Temp.:103.21, PWM:52
18:08:00.907 -> Water_Temp.:103.21, PWM:52
18:08:01.700 -> Water_Temp.:103.21, PWM:53
18:08:02.493 -> Water_Temp.:103.21, PWM:52
18:08:03.252 -> Water_Temp.:103.21, PWM:52
18:08:04.046 -> Water_Temp.:103.21, PWM:52
18:08:04.805 -> Water_Temp.:103.21, PWM:52
18:08:05.600 -> Water_Temp.:103.21, PWM:52
18:08:06.395 -> Water_Temp.:103.21, PWM:52
18:08:07.158 -> Water_Temp.:103.32, PWM:52
18:08:07.953 -> Water_Temp.:103.32, PWM:52
18:08:08.711 -> Water_Temp.:103.32, PWM:52
18:08:09.504 -> Water_Temp.:103.32, PWM:52
18:08:10.296 -> Water_Temp.:103.32, PWM:52
18:08:11.058 -> Water_Temp.:103.32, PWM:52
18:08:11.856 -> Water_Temp.:103.32, PWM:52
18:08:12.614 -> Water_Temp.:103.32, PWM:52
18:08:13.408 -> Water_Temp.:103.32, PWM:52
18:08:14.164 -> Water_Temp.:103.32, PWM:52
18:08:14.962 -> Water_Temp.:103.32, PWM:52
18:08:15.757 -> Water_Temp.:103.32, PWM:52
18:08:16.519 -> Water_Temp.:103.32, PWM:52
18:08:17.314 -> Water_Temp.:103.32, PWM:52
18:08:18.073 -> Water_Temp.:103.32, PWM:52
18:08:18.866 -> Water_Temp.:103.32, PWM:52
18:08:19.628 -> Water_Temp.:103.32, PWM:52
18:08:20.423 -> Water_Temp.:103.32, PWM:53
18:08:21.212 -> Water_Temp.:103.32, PWM:52
18:08:21.969 -> Water_Temp.:103.32, PWM:52
18:08:22.762 -> Water_Temp.:103.32, PWM:52
18:08:23.555 -> Water_Temp.:103.32, PWM:52
18:08:24.308 -> Water_Temp.:103.32, PWM:53
18:08:25.104 -> Water_Temp.:103.32, PWM:52
18:08:25.898 -> Water_Temp.:103.32, PWM:52
18:08:26.657 -> Water_Temp.:103.32, PWM:52
18:08:27.449 -> Water_Temp.:103.32, PWM:52
18:08:28.211 -> Water_Temp.:103.32, PWM:52
18:08:29.003 -> Water_Temp.:103.32, PWM:52
18:08:30.554 -> Water_Temp.:103.32, PWM:52
18:08:31.350 -> Water_Temp.:103.44, PWM:53
18:08:32.108 -> Water_Temp.:103.44, PWM:52
18:08:32.899 -> Water_Temp.:103.44, PWM:52
18:08:33.691 -> Water_Temp.:103.44, PWM:53
18:08:34.451 -> Water_Temp.:103.44, PWM:52
18:08:35.245 -> Water_Temp.:103.44, PWM:52
18:08:36.006 -> Water_Temp.:103.44, PWM:52
18:08:36.801 -> Water_Temp.:103.44, PWM:52
18:08:37.593 -> Water_Temp.:103.44, PWM:52
18:08:38.357 -> Water_Temp.:103.44, PWM:52
18:08:39.150 -> Water_Temp.:103.44, PWM:52
18:08:39.906 -> Water_Temp.:103.44, PWM:52
18:08:40.700 -> Water_Temp.:103.44, PWM:53
18:08:41.491 -> Water_Temp.:103.44, PWM:52
18:08:42.250 -> Water_Temp.:103.55, PWM:52
18:08:43.042 -> Water_Temp.:103.55, PWM:54
18:08:43.805 -> Water_Temp.:103.55, PWM:54
18:08:44.601 -> Water_Temp.:103.55, PWM:55
18:08:45.365 -> Water_Temp.:103.55, PWM:54
18:08:46.161 -> Water_Temp.:103.55, PWM:54
18:08:46.923 -> Water_Temp.:103.55, PWM:54
18:08:47.716 -> Water_Temp.:103.55, PWM:54
18:08:48.513 -> Water_Temp.:103.44, PWM:55
18:08:49.272 -> Water_Temp.:103.44, PWM:54
18:08:50.058 -> Water_Temp.:103.44, PWM:54
18:08:50.850 -> Water_Temp.:103.44, PWM:54
18:08:51.611 -> Water_Temp.:103.44, PWM:55
18:08:52.405 -> Water_Temp.:103.44, PWM:54
18:08:53.166 -> Water_Temp.:103.44, PWM:54
18:08:53.960 -> Water_Temp.:103.44, PWM:54
18:08:54.757 -> Water_Temp.:103.44, PWM:55
18:08:55.516 -> Water_Temp.:103.44, PWM:55
18:08:56.313 -> Water_Temp.:103.44, PWM:54
18:08:57.073 -> Water_Temp.:103.32, PWM:55
18:08:57.870 -> Water_Temp.:103.32, PWM:55
18:08:58.629 -> Water_Temp.:103.32, PWM:55
18:08:59.418 -> Water_Temp.:103.32, PWM:54
18:09:00.212 -> Water_Temp.:103.32, PWM:55
18:09:00.971 -> Water_Temp.:103.32, PWM:55
18:09:01.764 -> Water_Temp.:103.32, PWM:54
18:09:02.556 -> Water_Temp.:103.32, PWM:55
18:09:03.315 -> Water_Temp.:103.32, PWM:54
18:09:04.110 -> Water_Temp.:103.32, PWM:54
18:09:04.866 -> Water_Temp.:103.32, PWM:55
18:09:05.662 -> Water_Temp.:103.32, PWM:55
18:09:07.183 -> Water_Temp.:103.32, PWM:55
18:09:07.972 -> Water_Temp.:103.32, PWM:54
18:09:08.768 -> Water_Temp.:103.32, PWM:55
18:09:09.530 -> Water_Temp.:103.32, PWM:55
18:09:11.115 -> Water_Temp.:103.32, PWM:54
18:09:11.880 -> Water_Temp.:103.32, PWM:54
18:09:12.643 -> Water_Temp.:103.32, PWM:54
18:09:14.989 -> Water_Temp.:103.32, PWM:55
18:09:15.744 -> Water_Temp.:103.32, PWM:55
18:09:16.536 -> Water_Temp.:103.32, PWM:54
18:09:17.331 -> Water_Temp.:103.32, PWM:54
18:09:18.091 -> Water_Temp.:103.32, PWM:55
18:09:18.889 -> Water_Temp.:103.32, PWM:54
18:09:19.651 -> Water_Temp.:103.32, PWM:54
18:09:20.446 -> Water_Temp.:103.44, PWM:54
18:09:21.206 -> Water_Temp.:103.44, PWM:55
18:09:22.001 -> Water_Temp.:103.44, PWM:54
18:09:22.792 -> Water_Temp.:103.44, PWM:54
18:09:23.550 -> Water_Temp.:103.44, PWM:55
18:09:24.342 -> Water_Temp.:103.44, PWM:54
18:09:25.136 -> Water_Temp.:103.44, PWM:54
18:09:25.899 -> Water_Temp.:103.44, PWM:54
18:09:26.695 -> Water_Temp.:103.44, PWM:55
18:09:27.453 -> Water_Temp.:103.44, PWM:54
18:09:28.247 -> Water_Temp.:103.44, PWM:55
18:09:29.010 -> Water_Temp.:103.44, PWM:54
18:09:29.805 -> Water_Temp.:103.44, PWM:54
18:09:30.563 -> Water_Temp.:103.44, PWM:54
18:09:31.361 -> Water_Temp.:103.32, PWM:55
18:09:32.157 -> Water_Temp.:103.32, PWM:54
18:09:32.916 -> Water_Temp.:103.32, PWM:54
18:09:33.710 -> Water_Temp.:103.32, PWM:54
18:09:34.466 -> Water_Temp.:103.32, PWM:54
18:09:35.265 -> Water_Temp.:103.32, PWM:55
18:09:36.786 -> Water_Temp.:103.32, PWM:55
18:09:37.576 -> Water_Temp.:103.32, PWM:54
18:09:38.369 -> Water_Temp.:103.32, PWM:55
18:09:39.132 -> Water_Temp.:103.32, PWM:54
18:09:39.927 -> Water_Temp.:103.32, PWM:54
18:09:40.718 -> Water_Temp.:103.44, PWM:54
18:09:41.477 -> Water_Temp.:103.44, PWM:55
18:09:42.271 -> Water_Temp.:103.44, PWM:55
18:09:44.592 -> Water_Temp.:103.44, PWM:55
18:09:45.353 -> Water_Temp.:103.44, PWM:54
18:09:46.147 -> Water_Temp.:103.44, PWM:55
18:09:46.908 -> Water_Temp.:103.44, PWM:54
18:09:47.706 -> Water_Temp.:103.44, PWM:54
18:09:48.498 -> Water_Temp.:103.44, PWM:55
18:09:49.256 -> Water_Temp.:103.44, PWM:54
18:09:50.050 -> Water_Temp.:103.44, PWM:55
18:09:50.811 -> Water_Temp.:103.44, PWM:54
18:09:51.604 -> Water_Temp.:103.44, PWM:55
18:09:52.394 -> Water_Temp.:103.44, PWM:54
18:09:53.153 -> Water_Temp.:103.44, PWM:54
18:09:53.946 -> Water_Temp.:103.44, PWM:55
18:09:54.707 -> Water_Temp.:103.44, PWM:55
18:09:55.501 -> Water_Temp.:103.44, PWM:54
18:09:56.297 -> Water_Temp.:103.44, PWM:54
18:09:57.056 -> Water_Temp.:103.44, PWM:55
18:09:57.849 -> Water_Temp.:103.44, PWM:54
18:09:58.610 -> Water_Temp.:103.44, PWM:54
18:09:59.400 -> Water_Temp.:103.44, PWM:54
18:10:00.189 -> Water_Temp.:103.55, PWM:54
18:10:00.948 -> Water_Temp.:103.44, PWM:54
18:10:01.743 -> Water_Temp.:103.55, PWM:54
18:10:02.538 -> Water_Temp.:103.55, PWM:55
18:10:03.298 -> Water_Temp.:103.55, PWM:54
18:10:04.091 -> Water_Temp.:103.55, PWM:54
18:10:04.848 -> Water_Temp.:103.55, PWM:54
18:10:05.640 -> Water_Temp.:103.55, PWM:54
18:10:06.436 -> Water_Temp.:103.55, PWM:54
18:10:07.192 -> Water_Temp.:103.55, PWM:55
18:10:07.984 -> Water_Temp.:103.55, PWM:54
18:10:08.780 -> Water_Temp.:103.55, PWM:55
18:10:09.542 -> Water_Temp.:103.55, PWM:55
18:10:10.338 -> Water_Temp.:103.44, PWM:54
18:10:11.100 -> Water_Temp.:103.44, PWM:54
18:10:11.893 -> Water_Temp.:103.44, PWM:55
18:10:12.649 -> Water_Temp.:103.44, PWM:55
18:10:13.440 -> Water_Temp.:103.44, PWM:54
18:10:14.233 -> Water_Temp.:103.44, PWM:55
18:10:14.990 -> Water_Temp.:103.44, PWM:54
18:10:15.786 -> Water_Temp.:103.44, PWM:54
18:10:16.545 -> Water_Temp.:103.44, PWM:55
18:10:17.340 -> Water_Temp.:103.44, PWM:54
18:10:18.135 -> Water_Temp.:103.44, PWM:54
18:10:18.895 -> Water_Temp.:103.44, PWM:55
18:10:19.690 -> Water_Temp.:103.44, PWM:54
18:10:21.242 -> Water_Temp.:103.44, PWM:54
18:10:22.001 -> Water_Temp.:103.44, PWM:55
18:10:22.796 -> Water_Temp.:103.44, PWM:54
18:10:23.551 -> Water_Temp.:103.44, PWM:55
18:10:24.342 -> Water_Temp.:103.44, PWM:55
18:10:25.132 -> Water_Temp.:103.44, PWM:54
18:10:25.920 -> Water_Temp.:103.44, PWM:54
18:10:26.681 -> Water_Temp.:103.44, PWM:55
18:10:28.234 -> Water_Temp.:103.44, PWM:54
18:10:29.034 -> Water_Temp.:103.55, PWM:54
18:10:29.794 -> Water_Temp.:103.44, PWM:54
18:10:30.582 -> Water_Temp.:103.55, PWM:54
18:10:32.925 -> Water_Temp.:103.55, PWM:54
18:10:33.719 -> Water_Temp.:103.55, PWM:55
18:10:34.477 -> Water_Temp.:103.55, PWM:54
18:10:35.268 -> Water_Temp.:103.55, PWM:54
18:10:36.060 -> Water_Temp.:103.55, PWM:55
18:10:36.820 -> Water_Temp.:103.55, PWM:54
18:10:37.614 -> Water_Temp.:103.55, PWM:54
18:10:38.380 -> Water_Temp.:103.55, PWM:54
18:10:39.166 -> Water_Temp.:103.55, PWM:55
18:10:39.957 -> Water_Temp.:103.55, PWM:54
18:10:40.717 -> Water_Temp.:103.55, PWM:55
18:10:41.509 -> Water_Temp.:103.55, PWM:54
18:10:42.267 -> Water_Temp.:103.55, PWM:55
18:10:43.059 -> Water_Temp.:103.55, PWM:55
18:10:43.856 -> Water_Temp.:103.55, PWM:55
18:10:44.613 -> Water_Temp.:103.55, PWM:54
18:10:45.411 -> Water_Temp.:103.55, PWM:55
18:10:46.167 -> Water_Temp.:103.55, PWM:55
18:10:48.480 -> Water_Temp.:103.55, PWM:54
18:10:49.275 -> Water_Temp.:103.55, PWM:54
18:10:50.065 -> Water_Temp.:103.55, PWM:55
18:10:50.822 -> Water_Temp.:103.55, PWM:55
18:10:51.618 -> Water_Temp.:103.55, PWM:54
18:10:52.412 -> Water_Temp.:103.55, PWM:54
18:10:53.172 -> Water_Temp.:103.55, PWM:55
18:10:53.967 -> Water_Temp.:103.55, PWM:54
18:10:54.725 -> Water_Temp.:103.55, PWM:55
18:10:55.520 -> Water_Temp.:103.66, PWM:54
18:10:56.279 -> Water_Temp.:103.66, PWM:55
18:10:57.827 -> Water_Temp.:103.66, PWM:54
18:10:58.622 -> Water_Temp.:103.55, PWM:54
18:10:59.408 -> Water_Temp.:103.55, PWM:55
18:11:00.167 -> Water_Temp.:103.55, PWM:55
18:11:00.962 -> Water_Temp.:103.55, PWM:55
18:11:01.718 -> Water_Temp.:103.55, PWM:54
18:11:02.513 -> Water_Temp.:103.55, PWM:54
18:11:03.304 -> Water_Temp.:103.55, PWM:54
18:11:04.068 -> Water_Temp.:103.55, PWM:55
18:11:04.864 -> Water_Temp.:103.55, PWM:55
18:11:05.627 -> Water_Temp.:103.55, PWM:54
18:11:06.420 -> Water_Temp.:103.55, PWM:55
18:11:07.184 -> Water_Temp.:103.44, PWM:54
18:11:07.977 -> Water_Temp.:103.44, PWM:54
18:11:08.739 -> Water_Temp.:103.44, PWM:55
18:11:09.535 -> Water_Temp.:103.44, PWM:54
18:11:10.296 -> Water_Temp.:103.44, PWM:54
18:11:11.088 -> Water_Temp.:103.44, PWM:54
18:11:11.885 -> Water_Temp.:103.44, PWM:54
18:11:12.648 -> Water_Temp.:103.44, PWM:54
18:11:13.442 -> Water_Temp.:103.44, PWM:54
18:11:14.200 -> Water_Temp.:103.44, PWM:54
18:11:14.990 -> Water_Temp.:103.44, PWM:54
18:11:15.754 -> Water_Temp.:103.44, PWM:54
18:11:16.548 -> Water_Temp.:103.44, PWM:54
18:11:17.341 -> Water_Temp.:103.44, PWM:55
18:11:18.101 -> Water_Temp.:103.44, PWM:55
18:11:18.893 -> Water_Temp.:103.44, PWM:55
18:11:19.685 -> Water_Temp.:103.55, PWM:55
18:11:20.445 -> Water_Temp.:103.44, PWM:54
18:11:21.241 -> Water_Temp.:103.55, PWM:54
18:11:22.000 -> Water_Temp.:103.55, PWM:55
18:11:22.796 -> Water_Temp.:103.55, PWM:55
18:11:23.589 -> Water_Temp.:103.55, PWM:54
18:11:24.349 -> Water_Temp.:103.55, PWM:54
18:11:25.142 -> Water_Temp.:103.55, PWM:54
18:11:25.898 -> Water_Temp.:103.55, PWM:54
18:11:26.698 -> Water_Temp.:103.55, PWM:54
18:11:27.457 -> Water_Temp.:103.55, PWM:55
18:11:28.254 -> Water_Temp.:103.55, PWM:55
18:11:29.018 -> Water_Temp.:103.55, PWM:54
18:11:29.812 -> Water_Temp.:103.55, PWM:54
18:11:30.605 -> Water_Temp.:103.55, PWM:55
18:11:31.365 -> Water_Temp.:103.55, PWM:54
18:11:32.163 -> Water_Temp.:103.55, PWM:54
18:11:32.917 -> Water_Temp.:103.55, PWM:54
18:11:33.710 -> Water_Temp.:103.55, PWM:54
18:11:34.501 -> Water_Temp.:103.55, PWM:54
18:11:35.260 -> Water_Temp.:103.55, PWM:54
18:11:36.052 -> Water_Temp.:103.55, PWM:54
18:11:36.847 -> Water_Temp.:103.55, PWM:54
18:11:37.606 -> Water_Temp.:103.55, PWM:54
18:11:38.400 -> Water_Temp.:103.55, PWM:54
18:11:39.160 -> Water_Temp.:103.55, PWM:55
18:11:39.951 -> Water_Temp.:103.55, PWM:55
18:11:40.745 -> Water_Temp.:103.55, PWM:55
18:11:41.503 -> Water_Temp.:103.55, PWM:55
18:11:42.295 -> Water_Temp.:103.66, PWM:55
18:11:43.090 -> Water_Temp.:103.55, PWM:55
18:11:43.853 -> Water_Temp.:103.55, PWM:54
18:11:44.644 -> Water_Temp.:103.55, PWM:54
18:11:45.404 -> Water_Temp.:103.55, PWM:55
18:11:46.195 -> Water_Temp.:103.55, PWM:54
18:11:46.956 -> Water_Temp.:103.55, PWM:54
18:11:47.752 -> Water_Temp.:103.55, PWM:55
18:11:48.549 -> Water_Temp.:103.55, PWM:54
18:11:49.305 -> Water_Temp.:103.55, PWM:54
18:11:50.096 -> Water_Temp.:103.55, PWM:54
18:11:50.889 -> Water_Temp.:103.55, PWM:54
18:11:52.440 -> Water_Temp.:103.55, PWM:55
18:11:53.197 -> Water_Temp.:103.55, PWM:54
18:11:53.986 -> Water_Temp.:103.55, PWM:54
18:11:54.775 -> Water_Temp.:103.55, PWM:55
18:11:55.536 -> Water_Temp.:103.55, PWM:54
18:11:56.331 -> Water_Temp.:103.55, PWM:54
18:11:57.120 -> Water_Temp.:103.55, PWM:55
18:11:57.880 -> Water_Temp.:103.55, PWM:54
18:11:59.435 -> Water_Temp.:103.55, PWM:54
18:12:00.226 -> Water_Temp.:103.55, PWM:54
18:12:01.023 -> Water_Temp.:103.55, PWM:55
18:12:01.785 -> Water_Temp.:103.55, PWM:54
18:12:02.578 -> Water_Temp.:103.55, PWM:54
18:12:04.125 -> Water_Temp.:103.55, PWM:55
18:12:04.885 -> Water_Temp.:103.55, PWM:55
18:12:05.676 -> Water_Temp.:103.55, PWM:54
18:12:06.437 -> Water_Temp.:103.55, PWM:54
18:12:07.232 -> Water_Temp.:103.55, PWM:54
18:12:08.026 -> Water_Temp.:103.55, PWM:54
18:12:08.788 -> Water_Temp.:103.55, PWM:54
18:12:09.575 -> Water_Temp.:103.55, PWM:54
18:12:10.365 -> Water_Temp.:103.55, PWM:55
18:12:11.125 -> Water_Temp.:103.66, PWM:55
18:12:11.921 -> Water_Temp.:103.66, PWM:54
18:12:12.680 -> Water_Temp.:103.66, PWM:55
18:12:13.472 -> Water_Temp.:103.66, PWM:54
18:12:14.267 -> Water_Temp.:103.66, PWM:55
18:12:15.023 -> Water_Temp.:103.66, PWM:55
18:12:15.813 -> Water_Temp.:103.66, PWM:54
18:12:16.608 -> Water_Temp.:103.66, PWM:54
18:12:17.367 -> Water_Temp.:103.66, PWM:55
18:12:18.162 -> Water_Temp.:103.66, PWM:55
18:12:18.921 -> Water_Temp.:103.66, PWM:54
18:12:19.715 -> Water_Temp.:103.66, PWM:55
18:12:20.508 -> Water_Temp.:103.66, PWM:54
18:12:21.268 -> Water_Temp.:103.66, PWM:55
18:12:22.059 -> Water_Temp.:103.77, PWM:54
18:12:23.577 -> Water_Temp.:103.66, PWM:54
18:12:24.375 -> Water_Temp.:103.77, PWM:55
18:12:25.168 -> Water_Temp.:103.66, PWM:54
18:12:25.931 -> Water_Temp.:103.66, PWM:55
18:12:26.723 -> Water_Temp.:103.66, PWM:55
18:12:27.481 -> Water_Temp.:103.66, PWM:55
18:12:28.274 -> Water_Temp.:103.66, PWM:54
18:12:29.067 -> Water_Temp.:103.66, PWM:55
18:12:29.827 -> Water_Temp.:103.66, PWM:54
18:12:30.622 -> Water_Temp.:103.66, PWM:55
18:12:31.382 -> Water_Temp.:103.66, PWM:55
18:12:32.178 -> Water_Temp.:103.66, PWM:55
18:12:32.968 -> Water_Temp.:103.66, PWM:54
18:12:33.727 -> Water_Temp.:103.66, PWM:54
18:12:34.522 -> Water_Temp.:103.66, PWM:54
18:12:35.282 -> Water_Temp.:103.66, PWM:54
18:12:36.069 -> Water_Temp.:103.66, PWM:54
18:12:36.865 -> Water_Temp.:103.66, PWM:54
18:12:37.626 -> Water_Temp.:103.66, PWM:55
18:12:38.419 -> Water_Temp.:103.66, PWM:54
18:12:39.209 -> Water_Temp.:103.66, PWM:55
18:12:39.967 -> Water_Temp.:103.66, PWM:54
18:12:41.546 -> Water_Temp.:103.66, PWM:55
18:12:42.305 -> Water_Temp.:103.66, PWM:54
18:12:43.100 -> Water_Temp.:103.66, PWM:54
18:12:43.855 -> Water_Temp.:103.66, PWM:55
18:12:45.439 -> Water_Temp.:103.77, PWM:54
18:12:46.990 -> Water_Temp.:103.77, PWM:54
18:12:47.783 -> Water_Temp.:103.77, PWM:54
18:12:48.544 -> Water_Temp.:103.77, PWM:55
18:12:49.341 -> Water_Temp.:103.77, PWM:54
18:12:50.102 -> Water_Temp.:103.77, PWM:55
18:12:50.895 -> Water_Temp.:103.77, PWM:54
18:12:51.655 -> Water_Temp.:103.77, PWM:54
18:12:52.452 -> Water_Temp.:103.77, PWM:54
18:12:53.243 -> Water_Temp.:103.77, PWM:54
18:12:54.003 -> Water_Temp.:103.77, PWM:55
18:12:54.798 -> Water_Temp.:103.77, PWM:54
18:12:55.561 -> Water_Temp.:103.77, PWM:54
18:12:56.356 -> Water_Temp.:103.77, PWM:55
18:12:57.122 -> Water_Temp.:103.77, PWM:55
18:12:57.918 -> Water_Temp.:103.77, PWM:54
18:12:58.678 -> Water_Temp.:103.77, PWM:54
18:12:59.475 -> Water_Temp.:103.77, PWM:55
18:13:00.235 -> Water_Temp.:103.77, PWM:54
18:13:01.028 -> Water_Temp.:103.77, PWM:55
18:13:01.822 -> Water_Temp.:103.77, PWM:54
18:13:02.582 -> Water_Temp.:103.89, PWM:54
18:13:03.374 -> Water_Temp.:103.89, PWM:54
18:13:04.132 -> Water_Temp.:103.89, PWM:54
18:13:04.924 -> Water_Temp.:103.89, PWM:55
18:13:05.717 -> Water_Temp.:103.89, PWM:54
18:13:06.472 -> Water_Temp.:103.89, PWM:55
18:13:07.268 -> Water_Temp.:103.89, PWM:55
18:13:08.059 -> Water_Temp.:103.89, PWM:54
18:13:08.817 -> Water_Temp.:103.89, PWM:54
18:13:09.612 -> Water_Temp.:103.89, PWM:54
18:13:10.374 -> Water_Temp.:103.89, PWM:55
18:13:11.166 -> Water_Temp.:103.89, PWM:54
18:13:11.960 -> Water_Temp.:103.89, PWM:55
18:13:12.717 -> Water_Temp.:103.89, PWM:55
18:13:13.509 -> Water_Temp.:103.89, PWM:55
18:13:14.302 -> Water_Temp.:103.89, PWM:54
18:13:15.062 -> Water_Temp.:103.89, PWM:55
18:13:15.862 -> Water_Temp.:103.77, PWM:55
18:13:16.621 -> Water_Temp.:103.77, PWM:54
18:13:17.414 -> Water_Temp.:103.77, PWM:54
18:13:18.173 -> Water_Temp.:103.77, PWM:54
18:13:18.966 -> Water_Temp.:103.77, PWM:54
18:13:19.761 -> Water_Temp.:103.77, PWM:54
18:13:20.521 -> Water_Temp.:103.77, PWM:55
18:13:21.314 -> Water_Temp.:103.77, PWM:54
18:13:22.073 -> Water_Temp.:103.77, PWM:55
18:13:22.865 -> Water_Temp.:103.77, PWM:55
18:13:23.652 -> Water_Temp.:103.77, PWM:55
18:13:24.412 -> Water_Temp.:103.77, PWM:54
18:13:25.209 -> Water_Temp.:103.77, PWM:54
18:13:26.005 -> Water_Temp.:103.77, PWM:55
18:13:27.526 -> Water_Temp.:103.77, PWM:55
18:13:28.322 -> Water_Temp.:103.77, PWM:55
18:13:29.078 -> Water_Temp.:103.77, PWM:54
18:13:29.874 -> Water_Temp.:103.77, PWM:55
18:13:30.665 -> Water_Temp.:103.77, PWM:55
18:13:31.421 -> Water_Temp.:103.77, PWM:54
18:13:32.213 -> Water_Temp.:103.77, PWM:55
18:13:33.005 -> Water_Temp.:103.77, PWM:54
18:13:33.762 -> Water_Temp.:103.77, PWM:55
18:13:34.555 -> Water_Temp.:103.77, PWM:55
18:13:35.314 -> Water_Temp.:103.77, PWM:55
18:13:36.112 -> Water_Temp.:103.89, PWM:55
18:13:36.903 -> Water_Temp.:103.89, PWM:55
18:13:37.660 -> Water_Temp.:103.89, PWM:54
18:13:38.453 -> Water_Temp.:103.89, PWM:54
18:13:39.245 -> Water_Temp.:103.89, PWM:54
18:13:40.004 -> Water_Temp.:103.89, PWM:54
18:13:40.799 -> Water_Temp.:103.89, PWM:54
18:13:41.560 -> Water_Temp.:103.89, PWM:54
18:13:42.353 -> Water_Temp.:103.89, PWM:54
18:13:43.147 -> Water_Temp.:103.89, PWM:55
18:13:43.905 -> Water_Temp.:104.00, PWM:54
18:13:44.699 -> Water_Temp.:103.89, PWM:55
18:13:45.462 -> Water_Temp.:103.89, PWM:54
18:13:47.012 -> Water_Temp.:103.89, PWM:55
18:13:47.807 -> Water_Temp.:103.89, PWM:55
18:13:48.564 -> Water_Temp.:103.89, PWM:54
18:13:49.355 -> Water_Temp.:103.89, PWM:54
18:13:50.146 -> Water_Temp.:103.89, PWM:54
18:13:50.907 -> Water_Temp.:103.89, PWM:55
18:13:51.703 -> Water_Temp.:103.89, PWM:54
18:13:52.463 -> Water_Temp.:103.89, PWM:55
18:13:53.256 -> Water_Temp.:103.89, PWM:55
18:13:54.049 -> Water_Temp.:103.89, PWM:54
18:13:54.807 -> Water_Temp.:103.89, PWM:54
18:13:55.594 -> Water_Temp.:103.89, PWM:55
18:13:56.358 -> Water_Temp.:103.89, PWM:55
18:13:57.152 -> Water_Temp.:104.00, PWM:55
18:13:57.944 -> Water_Temp.:103.89, PWM:54
18:13:58.701 -> Water_Temp.:104.00, PWM:54
18:13:59.497 -> Water_Temp.:104.00, PWM:55
18:14:00.291 -> Water_Temp.:104.00, PWM:55
18:14:01.050 -> Water_Temp.:104.00, PWM:54
18:14:01.843 -> Water_Temp.:104.00, PWM:54
18:14:02.605 -> Water_Temp.:104.00, PWM:55
18:14:03.400 -> Water_Temp.:104.00, PWM:54
18:14:04.159 -> Water_Temp.:104.00, PWM:54
18:14:04.948 -> Water_Temp.:104.00, PWM:55
18:14:06.497 -> Water_Temp.:104.00, PWM:54
18:14:07.289 -> Water_Temp.:104.00, PWM:55
18:14:08.047 -> Water_Temp.:103.89, PWM:54
18:14:08.840 -> Water_Temp.:104.00, PWM:54
18:14:09.601 -> Water_Temp.:104.00, PWM:54
18:14:10.393 -> Water_Temp.:104.00, PWM:55
18:14:11.186 -> Water_Temp.:104.00, PWM:55
18:14:11.948 -> Water_Temp.:104.00, PWM:55
18:14:12.737 -> Water_Temp.:104.00, PWM:54
18:14:13.496 -> Water_Temp.:104.00, PWM:54
18:14:14.291 -> Water_Temp.:104.00, PWM:54
18:14:15.085 -> Water_Temp.:103.89, PWM:54
18:14:15.848 -> Water_Temp.:103.89, PWM:54
18:14:16.640 -> Water_Temp.:103.89, PWM:54
18:14:17.399 -> Water_Temp.:103.89, PWM:54
18:14:18.191 -> Water_Temp.:103.89, PWM:55
18:14:18.985 -> Water_Temp.:103.89, PWM:54
18:14:20.504 -> Water_Temp.:103.89, PWM:54
18:14:21.294 -> Water_Temp.:103.89, PWM:55
18:14:22.055 -> Water_Temp.:103.89, PWM:55
18:14:22.853 -> Water_Temp.:103.89, PWM:54
18:14:23.617 -> Water_Temp.:103.89, PWM:55
18:14:24.410 -> Water_Temp.:103.89, PWM:54
18:14:25.201 -> Water_Temp.:103.89, PWM:54
18:14:25.960 -> Water_Temp.:103.89, PWM:54
18:14:26.756 -> Water_Temp.:103.89, PWM:54
18:14:27.521 -> Water_Temp.:103.89, PWM:54
18:14:28.319 -> Water_Temp.:103.89, PWM:55
18:14:29.074 -> Water_Temp.:103.89, PWM:54
18:14:29.868 -> Water_Temp.:103.89, PWM:54
18:14:30.663 -> Water_Temp.:103.89, PWM:54
18:14:31.427 -> Water_Temp.:103.89, PWM:55
18:14:32.189 -> Water_Temp.:103.89, PWM:54
18:14:32.982 -> Water_Temp.:103.89, PWM:55
18:14:33.773 -> Water_Temp.:103.89, PWM:54
18:14:34.562 -> Water_Temp.:103.89, PWM:55
18:14:35.322 -> Water_Temp.:103.89, PWM:54
18:14:36.110 -> Water_Temp.:103.89, PWM:54
18:14:36.869 -> Water_Temp.:103.89, PWM:55
18:14:37.666 -> Water_Temp.:103.89, PWM:54
18:14:38.462 -> Water_Temp.:103.89, PWM:55
18:14:39.221 -> Water_Temp.:103.89, PWM:54
18:14:40.007 -> Water_Temp.:103.89, PWM:54
18:14:40.768 -> Water_Temp.:103.89, PWM:55
18:14:41.561 -> Water_Temp.:103.89, PWM:54
18:14:43.115 -> Water_Temp.:103.89, PWM:54
18:14:43.908 -> Water_Temp.:103.89, PWM:54
18:14:44.700 -> Water_Temp.:103.89, PWM:54
18:14:45.460 -> Water_Temp.:103.89, PWM:54
18:14:46.253 -> Water_Temp.:103.89, PWM:54
18:14:47.013 -> Water_Temp.:103.89, PWM:55
18:14:47.807 -> Water_Temp.:103.89, PWM:55
18:14:48.599 -> Water_Temp.:103.89, PWM:55
18:14:49.365 -> Water_Temp.:103.89, PWM:55
18:14:50.155 -> Water_Temp.:103.89, PWM:54
18:14:50.915 -> Water_Temp.:103.89, PWM:54
18:14:51.708 -> Water_Temp.:103.89, PWM:55
18:14:52.470 -> Water_Temp.:103.89, PWM:54
18:14:53.265 -> Water_Temp.:103.89, PWM:55
18:14:54.817 -> Water_Temp.:103.89, PWM:54
18:14:55.583 -> Water_Temp.:103.89, PWM:55
18:14:56.379 -> Water_Temp.:103.89, PWM:54
18:14:57.138 -> Water_Temp.:103.89, PWM:55
18:14:57.928 -> Water_Temp.:104.00, PWM:54
18:14:58.690 -> Water_Temp.:103.89, PWM:55
18:14:59.477 -> Water_Temp.:104.00, PWM:54
18:15:00.274 -> Water_Temp.:104.00, PWM:54
18:15:01.036 -> Water_Temp.:104.00, PWM:54
18:15:01.828 -> Water_Temp.:104.00, PWM:54
18:15:03.377 -> Water_Temp.:104.00, PWM:54
18:15:04.139 -> Water_Temp.:104.00, PWM:55
18:15:04.929 -> Water_Temp.:104.00, PWM:55
18:15:05.688 -> Water_Temp.:104.00, PWM:55
18:15:07.241 -> Water_Temp.:104.00, PWM:55
18:15:08.000 -> Water_Temp.:104.00, PWM:54
18:15:08.796 -> Water_Temp.:104.00, PWM:55
18:15:09.558 -> Water_Temp.:104.00, PWM:54
18:15:10.352 -> Water_Temp.:104.00, PWM:54
18:15:11.146 -> Water_Temp.:104.00, PWM:55
18:15:11.905 -> Water_Temp.:104.00, PWM:54
18:15:12.699 -> Water_Temp.:104.00, PWM:55
18:15:13.461 -> Water_Temp.:104.00, PWM:55
18:15:14.255 -> Water_Temp.:104.00, PWM:54
18:15:15.018 -> Water_Temp.:104.00, PWM:54
18:15:15.808 -> Water_Temp.:104.00, PWM:55
18:15:16.599 -> Water_Temp.:104.00, PWM:55
18:15:17.359 -> Water_Temp.:104.00, PWM:55
18:15:18.156 -> Water_Temp.:104.00, PWM:54
18:15:18.916 -> Water_Temp.:104.00, PWM:54
18:15:20.502 -> Water_Temp.:104.00, PWM:55
18:15:21.262 -> Water_Temp.:104.00, PWM:54
18:15:22.057 -> Water_Temp.:104.00, PWM:55
18:15:22.819 -> Water_Temp.:104.00, PWM:54
18:15:23.613 -> Water_Temp.:104.00, PWM:54
18:15:24.400 -> Water_Temp.:104.00, PWM:54
18:15:25.162 -> Water_Temp.:104.00, PWM:54
18:15:25.950 -> Water_Temp.:104.00, PWM:54
18:15:26.749 -> Water_Temp.:104.00, PWM:54
18:15:27.509 -> Water_Temp.:103.89, PWM:54
18:15:28.298 -> Water_Temp.:104.00, PWM:54
18:15:29.058 -> Water_Temp.:103.89, PWM:54
18:15:29.856 -> Water_Temp.:103.89, PWM:54
18:15:30.618 -> Water_Temp.:103.89, PWM:54
18:15:31.410 -> Water_Temp.:104.00, PWM:54
18:15:32.205 -> Water_Temp.:103.89, PWM:55
18:15:32.962 -> Water_Temp.:104.00, PWM:55
18:15:33.753 -> Water_Temp.:104.00, PWM:54
18:15:35.308 -> Water_Temp.:104.00, PWM:54
18:15:36.099 -> Water_Temp.:104.00, PWM:54
18:15:36.860 -> Water_Temp.:104.00, PWM:55
18:15:37.656 -> Water_Temp.:104.00, PWM:55
18:15:38.418 -> Water_Temp.:104.00, PWM:54
18:15:39.207 -> Water_Temp.:104.00, PWM:54
18:15:40.004 -> Water_Temp.:104.00, PWM:54
18:15:40.761 -> Water_Temp.:104.00, PWM:54
18:15:41.555 -> Water_Temp.:104.00, PWM:54
18:15:42.346 -> Water_Temp.:104.00, PWM:54
18:15:43.107 -> Water_Temp.:104.00, PWM:55
18:15:43.904 -> Water_Temp.:104.00, PWM:55
18:15:44.659 -> Water_Temp.:104.00, PWM:54
18:15:45.454 -> Water_Temp.:104.00, PWM:54
18:15:46.216 -> Water_Temp.:104.00, PWM:55
18:15:47.010 -> Water_Temp.:104.00, PWM:55
18:15:47.802 -> Water_Temp.:104.00, PWM:54
18:15:48.558 -> Water_Temp.:104.00, PWM:54
18:15:49.354 -> Water_Temp.:104.00, PWM:54
18:15:50.147 -> Water_Temp.:104.00, PWM:55
18:15:50.899 -> Water_Temp.:104.00, PWM:54
18:15:51.689 -> Water_Temp.:104.00, PWM:55
18:15:52.452 -> Water_Temp.:104.00, PWM:54
18:15:53.242 -> Water_Temp.:104.11, PWM:54
18:15:54.033 -> Water_Temp.:104.11, PWM:54
18:15:54.828 -> Water_Temp.:104.11, PWM:55
18:15:55.593 -> Water_Temp.:104.11, PWM:54
18:15:56.386 -> Water_Temp.:104.22, PWM:54
18:15:57.148 -> Water_Temp.:104.22, PWM:54
18:15:57.939 -> Water_Temp.:104.22, PWM:54
18:15:58.699 -> Water_Temp.:104.22, PWM:55
18:15:59.491 -> Water_Temp.:104.22, PWM:54
18:16:00.284 -> Water_Temp.:104.22, PWM:55
18:16:01.041 -> Water_Temp.:104.22, PWM:54
18:16:01.836 -> Water_Temp.:104.22, PWM:55
18:16:02.595 -> Water_Temp.:104.22, PWM:55
18:16:03.386 -> Water_Temp.:104.22, PWM:54
18:16:04.180 -> Water_Temp.:104.22, PWM:55
18:16:04.940 -> Water_Temp.:104.22, PWM:54
18:16:06.492 -> Water_Temp.:104.22, PWM:54
18:16:07.281 -> Water_Temp.:104.22, PWM:54
18:16:08.043 -> Water_Temp.:104.22, PWM:54
18:16:08.838 -> Water_Temp.:104.22, PWM:55
18:16:09.594 -> Water_Temp.:104.11, PWM:55
18:16:11.182 -> Water_Temp.:104.22, PWM:54
18:16:11.942 -> Water_Temp.:104.11, PWM:54
18:16:12.738 -> Water_Temp.:104.11, PWM:54
18:16:14.253 -> Water_Temp.:104.11, PWM:55
18:16:15.044 -> Water_Temp.:104.11, PWM:54
18:16:15.836 -> Water_Temp.:104.11, PWM:55
18:16:16.596 -> Water_Temp.:104.11, PWM:54
18:16:17.385 -> Water_Temp.:104.11, PWM:54
18:16:18.142 -> Water_Temp.:104.11, PWM:54
18:16:18.932 -> Water_Temp.:104.11, PWM:55
18:16:19.729 -> Water_Temp.:104.11, PWM:55
18:16:20.491 -> Water_Temp.:104.11, PWM:54
18:16:21.283 -> Water_Temp.:104.11, PWM:54
18:16:22.075 -> Water_Temp.:104.11, PWM:54
18:16:22.834 -> Water_Temp.:104.11, PWM:54
18:16:23.625 -> Water_Temp.:104.11, PWM:55
18:16:24.384 -> Water_Temp.:104.11, PWM:54
18:16:25.173 -> Water_Temp.:104.11, PWM:54
18:16:25.963 -> Water_Temp.:104.11, PWM:54
18:16:26.725 -> Water_Temp.:104.11, PWM:54
18:16:27.514 -> Water_Temp.:104.11, PWM:55
18:16:28.311 -> Water_Temp.:104.11, PWM:55
18:16:29.071 -> Water_Temp.:104.11, PWM:54
18:16:30.625 -> Water_Temp.:104.11, PWM:54
18:16:31.419 -> Water_Temp.:104.11, PWM:55
18:16:32.966 -> Water_Temp.:104.11, PWM:55
18:16:33.763 -> Water_Temp.:104.11, PWM:55
18:16:34.520 -> Water_Temp.:104.11, PWM:54
18:16:35.317 -> Water_Temp.:104.22, PWM:54
18:16:36.115 -> Water_Temp.:104.11, PWM:54
18:16:36.874 -> Water_Temp.:104.22, PWM:55
18:16:37.664 -> Water_Temp.:104.22, PWM:54
18:16:38.422 -> Water_Temp.:104.22, PWM:54
18:16:39.217 -> Water_Temp.:104.22, PWM:54
18:16:40.010 -> Water_Temp.:104.22, PWM:55
18:16:40.771 -> Water_Temp.:104.22, PWM:54
18:16:42.325 -> Water_Temp.:104.22, PWM:55
18:16:43.117 -> Water_Temp.:104.22, PWM:54
18:16:43.874 -> Water_Temp.:104.22, PWM:54
18:16:44.669 -> Water_Temp.:104.22, PWM:54
18:16:45.423 -> Water_Temp.:104.22, PWM:54
18:16:46.215 -> Water_Temp.:104.22, PWM:54
18:16:47.009 -> Water_Temp.:104.22, PWM:55
18:16:47.768 -> Water_Temp.:104.22, PWM:54
18:16:48.562 -> Water_Temp.:104.22, PWM:54
18:16:49.348 -> Water_Temp.:104.22, PWM:54
18:16:50.109 -> Water_Temp.:104.22, PWM:54
18:16:51.697 -> Water_Temp.:104.22, PWM:55
18:16:52.455 -> Water_Temp.:104.22, PWM:54
18:16:53.249 -> Water_Temp.:104.22, PWM:55
18:16:54.009 -> Water_Temp.:104.22, PWM:55
18:16:54.808 -> Water_Temp.:104.22, PWM:55
18:16:55.571 -> Water_Temp.:104.22, PWM:54
18:16:56.363 -> Water_Temp.:104.22, PWM:55
18:16:57.155 -> Water_Temp.:104.22, PWM:55
18:16:57.914 -> Water_Temp.:104.22, PWM:54
18:16:58.712 -> Water_Temp.:104.22, PWM:54
18:16:59.465 -> Water_Temp.:104.22, PWM:55
18:17:00.254 -> Water_Temp.:104.22, PWM:55
18:17:01.048 -> Water_Temp.:104.22, PWM:54
18:17:01.805 -> Water_Temp.:104.22, PWM:54
18:17:02.602 -> Water_Temp.:104.22, PWM:54
18:17:03.395 -> Water_Temp.:104.22, PWM:54
18:17:04.152 -> Water_Temp.:104.22, PWM:55
18:17:04.946 -> Water_Temp.:104.22, PWM:55
18:17:05.702 -> Water_Temp.:104.22, PWM:54
18:17:06.495 -> Water_Temp.:104.22, PWM:54
18:17:07.289 -> Water_Temp.:104.22, PWM:54
18:17:08.048 -> Water_Temp.:104.22, PWM:55
18:17:08.842 -> Water_Temp.:104.22, PWM:55
18:17:09.638 -> Water_Temp.:104.22, PWM:54
18:17:10.399 -> Water_Temp.:104.22, PWM:54
18:17:11.194 -> Water_Temp.:104.22, PWM:55
18:17:11.950 -> Water_Temp.:104.22, PWM:54
18:17:12.747 -> Water_Temp.:104.22, PWM:54
18:17:13.507 -> Water_Temp.:104.22, PWM:54
18:17:14.303 -> Water_Temp.:104.22, PWM:54
18:17:15.094 -> Water_Temp.:104.22, PWM:54
18:17:15.852 -> Water_Temp.:104.22, PWM:55
18:17:16.642 -> Water_Temp.:104.22, PWM:54
18:17:17.404 -> Water_Temp.:104.22, PWM:55
18:17:18.196 -> Water_Temp.:104.22, PWM:54
18:17:19.748 -> Water_Temp.:104.22, PWM:59
18:17:20.507 -> Water_Temp.:104.22, PWM:58
18:17:21.299 -> Water_Temp.:104.22, PWM:59
18:17:22.061 -> Water_Temp.:104.22, PWM:58
18:17:22.856 -> Water_Temp.:104.22, PWM:58
18:17:23.620 -> Water_Temp.:104.22, PWM:58
18:17:25.166 -> Water_Temp.:104.22, PWM:59
18:17:25.960 -> Water_Temp.:104.22, PWM:58
18:17:27.510 -> Water_Temp.:104.22, PWM:58
18:17:28.305 -> Water_Temp.:104.22, PWM:58
18:17:29.094 -> Water_Temp.:104.22, PWM:58
18:17:29.855 -> Water_Temp.:104.22, PWM:58
18:17:30.648 -> Water_Temp.:104.22, PWM:58
18:17:31.405 -> Water_Temp.:104.22, PWM:58
18:17:32.200 -> Water_Temp.:104.22, PWM:58
18:17:32.994 -> Water_Temp.:104.22, PWM:58
18:17:33.754 -> Water_Temp.:104.22, PWM:58
18:17:34.548 -> Water_Temp.:104.22, PWM:58
18:17:35.339 -> Water_Temp.:104.22, PWM:58
18:17:36.892 -> Water_Temp.:104.22, PWM:58
18:17:37.650 -> Water_Temp.:104.22, PWM:58
18:17:38.444 -> Water_Temp.:104.22, PWM:58
18:17:39.230 -> Water_Temp.:104.22, PWM:58
18:17:39.991 -> Water_Temp.:104.22, PWM:58
18:17:40.789 -> Water_Temp.:104.22, PWM:58
18:17:41.548 -> Water_Temp.:104.22, PWM:58
18:17:43.131 -> Water_Temp.:104.22, PWM:58
18:17:43.890 -> Water_Temp.:104.34, PWM:58
18:17:44.684 -> Water_Temp.:105.91, PWM:58
18:17:45.448 -> Water_Temp.:107.94, PWM:58
18:17:46.245 -> Water_Temp.:105.57, PWM:58
18:17:47.003 -> Water_Temp.:105.12, PWM:58
18:17:47.798 -> Water_Temp.:104.90, PWM:58
18:17:48.591 -> Water_Temp.:104.79, PWM:58
18:17:49.354 -> Water_Temp.:104.67, PWM:58
18:17:50.150 -> Water_Temp.:104.67, PWM:58
18:17:50.913 -> Water_Temp.:104.56, PWM:58
18:17:52.471 -> Water_Temp.:104.56, PWM:58
18:17:53.231 -> Water_Temp.:104.56, PWM:58
18:17:54.025 -> Water_Temp.:104.45, PWM:58
18:17:54.784 -> Water_Temp.:104.45, PWM:58
18:17:55.584 -> Water_Temp.:104.45, PWM:58
18:17:56.349 -> Water_Temp.:104.45, PWM:58
18:17:57.141 -> Water_Temp.:104.45, PWM:58
18:17:57.934 -> Water_Temp.:104.45, PWM:58
18:17:58.693 -> Water_Temp.:104.45, PWM:58
18:17:59.493 -> Water_Temp.:104.45, PWM:58
18:18:01.010 -> Water_Temp.:104.34, PWM:58
18:18:01.808 -> Water_Temp.:104.34, PWM:58
18:18:02.567 -> Water_Temp.:104.34, PWM:58
18:18:03.361 -> Water_Temp.:104.34, PWM:58
18:18:04.155 -> Water_Temp.:104.34, PWM:58
18:18:04.911 -> Water_Temp.:104.34, PWM:58
18:18:05.707 -> Water_Temp.:104.34, PWM:58
18:18:06.466 -> Water_Temp.:104.34, PWM:58
18:18:07.261 -> Water_Temp.:104.34, PWM:58
18:18:08.056 -> Water_Temp.:104.34, PWM:58
18:18:08.817 -> Water_Temp.:104.34, PWM:58
18:18:09.613 -> Water_Temp.:104.22, PWM:58
18:18:10.374 -> Water_Temp.:104.22, PWM:58
18:18:11.168 -> Water_Temp.:104.22, PWM:58
18:18:11.933 -> Water_Temp.:104.22, PWM:58
18:18:12.728 -> Water_Temp.:104.22, PWM:58
18:18:13.485 -> Water_Temp.:104.34, PWM:58
18:18:14.281 -> Water_Temp.:104.34, PWM:58
18:18:15.044 -> Water_Temp.:104.34, PWM:58
18:18:15.839 -> Water_Temp.:104.34, PWM:58
18:18:16.636 -> Water_Temp.:104.34, PWM:58
18:18:17.396 -> Water_Temp.:104.34, PWM:58
18:18:18.950 -> Water_Temp.:104.34, PWM:58
18:18:19.750 -> Water_Temp.:104.34, PWM:58
18:18:20.514 -> Water_Temp.:104.34, PWM:58
18:18:21.311 -> Water_Temp.:104.34, PWM:58
18:18:22.075 -> Water_Temp.:104.34, PWM:58
18:18:22.869 -> Water_Temp.:104.34, PWM:58
18:18:24.380 -> Water_Temp.:104.34, PWM:58
18:18:25.176 -> Water_Temp.:104.34, PWM:58
18:18:25.935 -> Water_Temp.:104.34, PWM:58
18:18:26.733 -> Water_Temp.:104.34, PWM:58
18:18:27.498 -> Water_Temp.:104.34, PWM:58
18:18:28.288 -> Water_Temp.:104.34, PWM:58
18:18:29.085 -> Water_Temp.:104.34, PWM:58
18:18:29.846 -> Water_Temp.:104.34, PWM:58
18:18:30.640 -> Water_Temp.:104.45, PWM:58
18:18:31.396 -> Water_Temp.:104.45, PWM:58
18:18:32.191 -> Water_Temp.:104.45, PWM:58
18:18:32.986 -> Water_Temp.:104.45, PWM:58
18:18:33.746 -> Water_Temp.:104.45, PWM:58
18:18:34.541 -> Water_Temp.:104.45, PWM:58
18:18:36.064 -> Water_Temp.:104.45, PWM:58
18:18:36.861 -> Water_Temp.:104.45, PWM:58
18:18:37.624 -> Water_Temp.:104.45, PWM:58
18:18:39.177 -> Water_Temp.:104.45, PWM:58
18:18:39.972 -> Water_Temp.:104.45, PWM:58
18:18:40.728 -> Water_Temp.:104.45, PWM:58
18:18:41.525 -> Water_Temp.:104.45, PWM:58
18:18:42.321 -> Water_Temp.:104.45, PWM:58
18:18:43.082 -> Water_Temp.:104.45, PWM:58
18:18:43.873 -> Water_Temp.:104.45, PWM:58
18:18:44.632 -> Water_Temp.:104.45, PWM:58
18:18:46.191 -> Water_Temp.:104.34, PWM:58
18:18:46.954 -> Water_Temp.:104.34, PWM:58
18:18:47.747 -> Water_Temp.:104.34, PWM:58
18:18:48.505 -> Water_Temp.:104.45, PWM:58
18:18:49.297 -> Water_Temp.:104.34, PWM:58
18:18:50.091 -> Water_Temp.:104.34, PWM:58
18:18:50.852 -> Water_Temp.:104.34, PWM:58
18:18:51.639 -> Water_Temp.:104.45, PWM:58
18:18:53.981 -> Water_Temp.:104.34, PWM:58
18:18:54.743 -> Water_Temp.:104.34, PWM:58
18:18:55.537 -> Water_Temp.:104.34, PWM:58
18:18:56.298 -> Water_Temp.:104.34, PWM:58
18:18:57.088 -> Water_Temp.:104.34, PWM:58
18:18:57.884 -> Water_Temp.:104.34, PWM:58
18:18:59.404 -> Water_Temp.:104.34, PWM:58
18:19:00.195 -> Water_Temp.:104.34, PWM:58
18:19:01.749 -> Water_Temp.:104.34, PWM:58
18:19:02.541 -> Water_Temp.:104.34, PWM:58
18:19:03.300 -> Water_Temp.:104.34, PWM:58
18:19:04.094 -> Water_Temp.:104.34, PWM:58
18:19:04.856 -> Water_Temp.:104.34, PWM:58
18:19:05.652 -> Water_Temp.:104.34, PWM:58
18:19:07.201 -> Water_Temp.:104.45, PWM:58
18:19:07.961 -> Water_Temp.:104.45, PWM:58
18:19:08.753 -> Water_Temp.:104.45, PWM:58
18:19:09.546 -> Water_Temp.:104.45, PWM:58
18:19:10.303 -> Water_Temp.:104.45, PWM:58
18:19:11.100 -> Water_Temp.:104.45, PWM:58
18:19:11.860 -> Water_Temp.:104.45, PWM:58
18:19:12.655 -> Water_Temp.:104.45, PWM:58
18:19:13.448 -> Water_Temp.:104.45, PWM:58
18:19:14.963 -> Water_Temp.:104.45, PWM:58
18:19:15.759 -> Water_Temp.:104.45, PWM:58
18:19:16.519 -> Water_Temp.:104.45, PWM:58
18:19:17.310 -> Water_Temp.:104.45, PWM:58
18:19:18.072 -> Water_Temp.:104.45, PWM:58
18:19:18.863 -> Water_Temp.:104.56, PWM:58
18:19:19.657 -> Water_Temp.:106.14, PWM:58
18:19:20.415 -> Water_Temp.:108.16, PWM:58
18:19:21.204 -> Water_Temp.:105.80, PWM:58
18:19:21.994 -> Water_Temp.:105.35, PWM:58
18:19:22.754 -> Water_Temp.:105.12, PWM:58
18:19:23.547 -> Water_Temp.:105.01, PWM:58
18:19:24.337 -> Water_Temp.:104.90, PWM:58
18:19:25.092 -> Water_Temp.:104.90, PWM:58
18:19:25.886 -> Water_Temp.:104.79, PWM:58
18:19:26.650 -> Water_Temp.:104.79, PWM:58
18:19:27.434 -> Water_Temp.:104.79, PWM:58
18:19:28.228 -> Water_Temp.:104.79, PWM:58
18:19:29.022 -> Water_Temp.:104.79, PWM:58
18:19:29.781 -> Water_Temp.:104.67, PWM:58
18:19:30.574 -> Water_Temp.:104.67, PWM:59
18:19:31.337 -> Water_Temp.:104.67, PWM:58
18:19:32.133 -> Water_Temp.:104.67, PWM:58
18:19:32.894 -> Water_Temp.:104.67, PWM:58
18:19:33.687 -> Water_Temp.:104.56, PWM:58
18:19:34.482 -> Water_Temp.:104.56, PWM:58
18:19:35.240 -> Water_Temp.:104.56, PWM:58
18:19:36.039 -> Water_Temp.:104.56, PWM:58
18:19:36.794 -> Water_Temp.:104.56, PWM:58
18:19:37.588 -> Water_Temp.:104.56, PWM:58
18:19:38.385 -> Water_Temp.:104.56, PWM:58
18:19:39.146 -> Water_Temp.:104.56, PWM:58
18:19:39.938 -> Water_Temp.:104.56, PWM:58
18:19:40.700 -> Water_Temp.:104.56, PWM:58
18:19:41.497 -> Water_Temp.:104.56, PWM:58
18:19:42.260 -> Water_Temp.:104.56, PWM:58
18:19:43.055 -> Water_Temp.:104.56, PWM:58
18:19:43.816 -> Water_Temp.:104.56, PWM:58
18:19:44.612 -> Water_Temp.:104.56, PWM:58
18:19:45.402 -> Water_Temp.:104.56, PWM:58
18:19:46.156 -> Water_Temp.:104.56, PWM:58
18:19:46.947 -> Water_Temp.:104.56, PWM:58
18:19:47.739 -> Water_Temp.:104.67, PWM:58
18:19:48.502 -> Water_Temp.:104.67, PWM:58
18:19:49.295 -> Water_Temp.:104.56, PWM:58
18:19:50.050 -> Water_Temp.:104.67, PWM:58
18:19:50.843 -> Water_Temp.:104.56, PWM:58
18:19:52.398 -> Water_Temp.:104.56, PWM:58
18:19:53.195 -> Water_Temp.:104.56, PWM:58
18:19:53.954 -> Water_Temp.:104.56, PWM:58
18:19:54.746 -> Water_Temp.:104.56, PWM:58
18:19:55.538 -> Water_Temp.:104.56, PWM:58
18:19:57.092 -> Water_Temp.:104.56, PWM:58
18:19:57.852 -> Water_Temp.:104.56, PWM:58
18:19:58.646 -> Water_Temp.:104.56, PWM:58
18:19:59.409 -> Water_Temp.:104.56, PWM:58
18:20:00.962 -> Water_Temp.:104.56, PWM:58
18:20:01.758 -> Water_Temp.:104.56, PWM:58
18:20:02.517 -> Water_Temp.:104.67, PWM:58
18:20:03.309 -> Water_Temp.:104.67, PWM:58
18:20:04.071 -> Water_Temp.:104.67, PWM:58
18:20:04.865 -> Water_Temp.:104.67, PWM:58
18:20:05.659 -> Water_Temp.:104.67, PWM:58
18:20:06.415 -> Water_Temp.:104.67, PWM:58
18:20:07.212 -> Water_Temp.:104.67, PWM:58
18:20:07.972 -> Water_Temp.:104.67, PWM:58
18:20:08.768 -> Water_Temp.:104.67, PWM:58
18:20:09.531 -> Water_Temp.:104.67, PWM:58
18:20:10.326 -> Water_Temp.:104.67, PWM:58
18:20:11.087 -> Water_Temp.:104.67, PWM:58
18:20:11.880 -> Water_Temp.:104.67, PWM:58
18:20:12.678 -> Water_Temp.:104.67, PWM:58
18:20:13.441 -> Water_Temp.:104.67, PWM:58
18:20:14.237 -> Water_Temp.:104.56, PWM:58
18:20:14.993 -> Water_Temp.:104.56, PWM:58
18:20:15.792 -> Water_Temp.:104.56, PWM:58
18:20:16.551 -> Water_Temp.:104.56, PWM:58
18:20:17.344 -> Water_Temp.:104.67, PWM:58
18:20:18.137 -> Water_Temp.:104.67, PWM:58
18:20:18.900 -> Water_Temp.:104.67, PWM:58
18:20:19.697 -> Water_Temp.:104.67, PWM:58
18:20:20.453 -> Water_Temp.:104.67, PWM:58
18:20:21.244 -> Water_Temp.:104.67, PWM:58
18:20:23.550 -> Water_Temp.:104.67, PWM:58
18:20:24.310 -> Water_Temp.:104.67, PWM:58
18:20:25.101 -> Water_Temp.:104.67, PWM:58
18:20:25.894 -> Water_Temp.:104.67, PWM:58
18:20:26.654 -> Water_Temp.:104.67, PWM:58
18:20:27.449 -> Water_Temp.:104.67, PWM:58
18:20:28.213 -> Water_Temp.:104.67, PWM:58
18:20:29.001 -> Water_Temp.:104.67, PWM:58
18:20:29.791 -> Water_Temp.:104.67, PWM:58
18:20:30.551 -> Water_Temp.:104.67, PWM:58
18:20:31.349 -> Water_Temp.:104.67, PWM:58
18:20:32.110 -> Water_Temp.:104.56, PWM:59
18:20:32.900 -> Water_Temp.:104.67, PWM:58
18:20:33.692 -> Water_Temp.:104.56, PWM:58
18:20:34.452 -> Water_Temp.:104.56, PWM:58
18:20:35.242 -> Water_Temp.:104.67, PWM:58
18:20:36.031 -> Water_Temp.:104.56, PWM:58
18:20:36.792 -> Water_Temp.:104.56, PWM:58
18:20:37.586 -> Water_Temp.:104.56, PWM:58
18:20:38.345 -> Water_Temp.:104.56, PWM:58
18:20:39.137 -> Water_Temp.:104.56, PWM:58
18:20:39.926 -> Water_Temp.:104.56, PWM:58
18:20:40.681 -> Water_Temp.:104.56, PWM:58
18:20:41.472 -> Water_Temp.:104.56, PWM:58
18:20:42.263 -> Water_Temp.:104.56, PWM:58
18:20:43.054 -> Water_Temp.:104.56, PWM:58
18:20:44.572 -> Water_Temp.:104.56, PWM:58
18:20:45.363 -> Water_Temp.:104.56, PWM:58
18:20:46.119 -> Water_Temp.:104.67, PWM:58
18:20:46.913 -> Water_Temp.:104.67, PWM:58
18:20:47.708 -> Water_Temp.:104.67, PWM:58
18:20:48.472 -> Water_Temp.:104.67, PWM:58
18:20:50.025 -> Water_Temp.:104.67, PWM:58
18:20:50.817 -> Water_Temp.:104.67, PWM:58
18:20:51.608 -> Water_Temp.:104.67, PWM:58
18:20:53.159 -> Water_Temp.:104.67, PWM:58
18:20:53.951 -> Water_Temp.:104.67, PWM:58
18:20:54.712 -> Water_Temp.:104.79, PWM:58
18:20:55.502 -> Water_Temp.:104.79, PWM:58
18:20:56.264 -> Water_Temp.:104.79, PWM:58
18:20:57.058 -> Water_Temp.:104.79, PWM:58
18:20:57.820 -> Water_Temp.:104.79, PWM:59
18:20:58.615 -> Water_Temp.:104.79, PWM:58
18:20:59.410 -> Water_Temp.:104.79, PWM:58
18:21:00.164 -> Water_Temp.:104.79, PWM:58
18:21:00.955 -> Water_Temp.:104.79, PWM:59
18:21:01.717 -> Water_Temp.:104.79, PWM:58
18:21:02.512 -> Water_Temp.:104.79, PWM:58
18:21:03.305 -> Water_Temp.:104.79, PWM:58
18:21:04.066 -> Water_Temp.:104.79, PWM:58
18:21:04.862 -> Water_Temp.:104.79, PWM:58
18:21:05.623 -> Water_Temp.:104.79, PWM:58
18:21:06.417 -> Water_Temp.:104.67, PWM:58
18:21:07.969 -> Water_Temp.:104.67, PWM:58
18:21:08.762 -> Water_Temp.:104.67, PWM:58
18:21:09.523 -> Water_Temp.:104.79, PWM:58
18:21:10.320 -> Water_Temp.:104.79, PWM:58
18:21:11.873 -> Water_Temp.:104.79, PWM:58
18:21:12.667 -> Water_Temp.:104.79, PWM:58
18:21:14.222 -> Water_Temp.:104.79, PWM:58
18:21:14.983 -> Water_Temp.:104.79, PWM:58
18:21:15.774 -> Water_Temp.:104.79, PWM:58
18:21:16.563 -> Water_Temp.:104.79, PWM:58
18:21:17.325 -> Water_Temp.:104.90, PWM:58
18:21:18.114 -> Water_Temp.:104.90, PWM:58
18:21:18.875 -> Water_Temp.:104.90, PWM:58
18:21:19.668 -> Water_Temp.:104.90, PWM:58
18:21:21.216 -> Water_Temp.:104.90, PWM:58
18:21:21.973 -> Water_Temp.:104.90, PWM:58
18:21:22.765 -> Water_Temp.:104.90, PWM:58
18:21:23.530 -> Water_Temp.:104.90, PWM:58
18:21:24.325 -> Water_Temp.:104.79, PWM:58
18:21:25.087 -> Water_Temp.:104.79, PWM:58
18:21:25.882 -> Water_Temp.:104.79, PWM:58
18:21:26.672 -> Water_Temp.:104.90, PWM:58
18:21:27.459 -> Water_Temp.:104.79, PWM:58
18:21:28.218 -> Water_Temp.:104.90, PWM:58
18:21:29.009 -> Water_Temp.:104.79, PWM:58
18:21:29.800 -> Water_Temp.:104.79, PWM:58
18:21:30.560 -> Water_Temp.:104.90, PWM:58
18:21:31.358 -> Water_Temp.:104.90, PWM:58
18:21:32.912 -> Water_Temp.:104.90, PWM:58
18:21:33.670 -> Water_Temp.:104.79, PWM:58
18:21:34.465 -> Water_Temp.:104.79, PWM:58
18:21:35.229 -> Water_Temp.:104.79, PWM:58
18:21:36.017 -> Water_Temp.:104.79, PWM:58
18:21:36.816 -> Water_Temp.:104.79, PWM:58
18:21:37.569 -> Water_Temp.:104.79, PWM:58
18:21:38.364 -> Water_Temp.:104.79, PWM:58
18:21:39.157 -> Water_Temp.:104.79, PWM:58
18:21:39.919 -> Water_Temp.:104.79, PWM:58
18:21:40.708 -> Water_Temp.:104.79, PWM:58
18:21:41.470 -> Water_Temp.:104.79, PWM:58
18:21:42.264 -> Water_Temp.:104.79, PWM:58
18:21:43.057 -> Water_Temp.:104.79, PWM:58
18:21:43.820 -> Water_Temp.:104.79, PWM:58
18:21:44.615 -> Water_Temp.:104.79, PWM:58
18:21:45.376 -> Water_Temp.:104.79, PWM:58
18:21:46.171 -> Water_Temp.:104.90, PWM:58
18:21:46.935 -> Water_Temp.:104.90, PWM:58
18:21:47.729 -> Water_Temp.:104.90, PWM:58
18:21:49.255 -> Water_Temp.:104.79, PWM:58
18:21:50.817 -> Water_Temp.:104.79, PWM:58
18:21:51.611 -> Water_Temp.:104.79, PWM:58
18:21:52.398 -> Water_Temp.:104.90, PWM:58
18:21:53.915 -> Water_Temp.:104.79, PWM:58
18:21:55.463 -> Water_Temp.:104.90, PWM:58
18:21:56.256 -> Water_Temp.:104.90, PWM:58
18:21:57.054 -> Water_Temp.:104.90, PWM:58
18:21:57.812 -> Water_Temp.:104.90, PWM:58
18:21:58.608 -> Water_Temp.:104.90, PWM:58
18:21:59.365 -> Water_Temp.:104.90, PWM:58
18:22:00.161 -> Water_Temp.:104.90, PWM:58
18:22:00.953 -> Water_Temp.:104.90, PWM:58
18:22:01.714 -> Water_Temp.:104.90, PWM:58
18:22:02.512 -> Water_Temp.:104.90, PWM:58
18:22:03.276 -> Water_Temp.:104.90, PWM:58
18:22:04.072 -> Water_Temp.:104.79, PWM:58
18:22:04.834 -> Water_Temp.:104.90, PWM:58
18:22:05.629 -> Water_Temp.:104.90, PWM:58
18:22:06.387 -> Water_Temp.:104.90, PWM:58
18:22:07.184 -> Water_Temp.:104.90, PWM:58
18:22:07.944 -> Water_Temp.:104.90, PWM:58
18:22:08.741 -> Water_Temp.:104.90, PWM:58
18:22:09.536 -> Water_Temp.:104.90, PWM:58
18:22:10.288 -> Water_Temp.:104.90, PWM:58
18:22:11.085 -> Water_Temp.:104.90, PWM:58
18:22:11.844 -> Water_Temp.:104.90, PWM:58
18:22:12.641 -> Water_Temp.:104.90, PWM:58
18:22:13.432 -> Water_Temp.:104.90, PWM:58
18:22:14.188 -> Water_Temp.:104.90, PWM:58
18:22:14.983 -> Water_Temp.:104.90, PWM:58
18:22:15.746 -> Water_Temp.:104.90, PWM:58
18:22:16.538 -> Water_Temp.:104.90, PWM:58
18:22:17.332 -> Water_Temp.:104.90, PWM:58
18:22:18.089 -> Water_Temp.:105.01, PWM:58
18:22:18.888 -> Water_Temp.:105.01, PWM:58
18:22:19.650 -> Water_Temp.:105.01, PWM:58
18:22:20.444 -> Water_Temp.:105.01, PWM:58
18:22:21.206 -> Water_Temp.:105.01, PWM:58
18:22:22.003 -> Water_Temp.:105.01, PWM:58
18:22:22.766 -> Water_Temp.:105.01, PWM:58
18:22:24.330 -> Water_Temp.:105.01, PWM:58
18:22:25.093 -> Water_Temp.:104.90, PWM:58
18:22:25.886 -> Water_Temp.:104.90, PWM:58
18:22:26.680 -> Water_Temp.:104.90, PWM:58
18:22:27.440 -> Water_Temp.:104.90, PWM:58
18:22:28.229 -> Water_Temp.:104.90, PWM:58
18:22:28.992 -> Water_Temp.:104.90, PWM:30
18:22:29.791 -> Water_Temp.:104.90, PWM:29
18:22:30.579 -> Water_Temp.:104.90, PWM:0
18:22:31.341 -> Water_Temp.:104.90, PWM:0
18:22:32.137 -> Water_Temp.:104.90, PWM:0
18:22:32.896 -> Water_Temp.:104.90, PWM:0
18:22:33.690 -> Water_Temp.:104.90, PWM:0
18:22:34.452 -> Water_Temp.:104.90, PWM:0
18:22:35.247 -> Water_Temp.:104.90, PWM:0
18:22:36.009 -> Water_Temp.:104.90, PWM:0
18:22:36.807 -> Water_Temp.:104.90, PWM:0
18:22:37.567 -> Water_Temp.:104.90, PWM:0
18:22:38.359 -> Water_Temp.:105.01, PWM:0
18:22:39.152 -> Water_Temp.:105.01, PWM:0
18:22:39.908 -> Water_Temp.:105.01, PWM:0
18:22:40.705 -> Water_Temp.:105.01, PWM:0
18:22:41.498 -> Water_Temp.:105.01, PWM:0
18:22:42.254 -> Water_Temp.:105.01, PWM:0
18:22:43.048 -> Water_Temp.:105.01, PWM:0
18:22:43.809 -> Water_Temp.:105.01, PWM:0
18:22:44.604 -> Water_Temp.:105.01, PWM:0
18:22:45.364 -> Water_Temp.:105.01, PWM:0
18:22:46.160 -> Water_Temp.:105.01, PWM:0
18:22:46.956 -> Water_Temp.:105.01, PWM:0
18:22:47.719 -> Water_Temp.:105.01, PWM:0
18:22:48.511 -> Water_Temp.:105.01, PWM:0
18:22:49.275 -> Water_Temp.:105.01, PWM:0
18:22:50.074 -> Water_Temp.:105.01, PWM:0
18:22:50.836 -> Water_Temp.:104.90, PWM:0
18:22:51.633 -> Water_Temp.:104.90, PWM:0
18:22:52.393 -> Water_Temp.:104.90, PWM:0
18:22:53.189 -> Water_Temp.:104.90, PWM:0
18:22:53.947 -> Water_Temp.:104.90, PWM:0
18:22:54.742 -> Water_Temp.:104.90, PWM:0
18:22:55.539 -> Water_Temp.:104.90, PWM:0
18:22:56.300 -> Water_Temp.:104.90, PWM:0
18:22:57.091 -> Water_Temp.:104.90, PWM:0
18:22:57.853 -> Water_Temp.:104.90, PWM:0
18:22:58.647 -> Water_Temp.:104.90, PWM:0
18:22:59.412 -> Water_Temp.:104.79, PWM:0
18:23:00.205 -> Water_Temp.:104.79, PWM:0
18:23:00.999 -> Water_Temp.:104.79, PWM:0
18:23:01.761 -> Water_Temp.:104.79, PWM:0
18:23:02.522 -> Water_Temp.:104.79, PWM:0
18:23:03.319 -> Water_Temp.:104.79, PWM:0
18:23:04.115 -> Water_Temp.:104.79, PWM:0
18:23:04.873 -> Water_Temp.:104.79, PWM:0
18:23:05.672 -> Water_Temp.:104.79, PWM:0
18:23:06.435 -> Water_Temp.:104.79, PWM:0
18:23:07.226 -> Water_Temp.:104.79, PWM:0
18:23:07.985 -> Water_Temp.:104.79, PWM:0
18:23:08.779 -> Water_Temp.:104.79, PWM:0
18:23:09.573 -> Water_Temp.:104.67, PWM:0
18:23:10.333 -> Water_Temp.:104.67, PWM:0
18:23:11.124 -> Water_Temp.:104.67, PWM:0
18:23:11.918 -> Water_Temp.:104.67, PWM:0
18:23:12.679 -> Water_Temp.:104.67, PWM:0
18:23:13.444 -> Water_Temp.:104.67, PWM:0
18:23:14.238 -> Water_Temp.:104.67, PWM:0
18:23:15.032 -> Water_Temp.:104.67, PWM:0
18:23:15.794 -> Water_Temp.:104.67, PWM:0
18:23:16.590 -> Water_Temp.:104.67, PWM:0
18:23:17.348 -> Water_Temp.:104.67, PWM:0
18:23:18.142 -> Water_Temp.:104.56, PWM:0
18:23:18.904 -> Water_Temp.:104.56, PWM:0
18:23:19.704 -> Water_Temp.:104.56, PWM:0
18:23:20.463 -> Water_Temp.:104.56, PWM:0
18:23:21.261 -> Water_Temp.:104.56, PWM:0
18:23:22.056 -> Water_Temp.:104.56, PWM:0
18:23:22.818 -> Water_Temp.:104.56, PWM:0
18:23:23.612 -> Water_Temp.:104.56, PWM:0
18:23:24.371 -> Water_Temp.:104.56, PWM:0
18:23:25.169 -> Water_Temp.:104.56, PWM:0
18:23:25.932 -> Water_Temp.:104.45, PWM:0
18:23:26.725 -> Water_Temp.:104.45, PWM:0
18:23:27.487 -> Water_Temp.:104.45, PWM:0
18:23:28.283 -> Water_Temp.:104.45, PWM:0
18:23:29.043 -> Water_Temp.:104.45, PWM:0
18:23:29.840 -> Water_Temp.:104.45, PWM:0
18:23:30.633 -> Water_Temp.:104.45, PWM:0
18:23:31.387 -> Water_Temp.:104.45, PWM:0
18:23:32.182 -> Water_Temp.:104.45, PWM:0
18:23:32.975 -> Water_Temp.:104.34, PWM:0
18:23:33.734 -> Water_Temp.:104.34, PWM:0
18:23:34.530 -> Water_Temp.:104.34, PWM:0
18:23:35.290 -> Water_Temp.:104.34, PWM:0
18:23:36.087 -> Water_Temp.:104.34, PWM:0
18:23:36.848 -> Water_Temp.:104.34, PWM:0
18:23:37.646 -> Water_Temp.:104.34, PWM:0
18:23:38.408 -> Water_Temp.:104.34, PWM:0
18:23:39.206 -> Water_Temp.:104.34, PWM:0
18:23:39.965 -> Water_Temp.:104.34, PWM:0
18:23:40.759 -> Water_Temp.:104.22, PWM:0
18:23:41.553 -> Water_Temp.:104.22, PWM:0
18:23:42.310 -> Water_Temp.:104.22, PWM:0
18:23:43.106 -> Water_Temp.:104.22, PWM:0
18:23:43.866 -> Water_Temp.:104.22, PWM:0
18:23:44.664 -> Water_Temp.:104.22, PWM:0
18:23:45.428 -> Water_Temp.:104.22, PWM:0
18:23:46.220 -> Water_Temp.:104.22, PWM:0
18:23:47.008 -> Water_Temp.:104.11, PWM:0
18:23:47.770 -> Water_Temp.:104.11, PWM:0
18:23:48.565 -> Water_Temp.:104.11, PWM:0
18:23:49.329 -> Water_Temp.:104.11, PWM:0
18:23:50.124 -> Water_Temp.:104.11, PWM:0
18:23:50.889 -> Water_Temp.:104.11, PWM:0
18:23:51.686 -> Water_Temp.:104.11, PWM:0
18:23:52.449 -> Water_Temp.:104.11, PWM:0
18:23:53.244 -> Water_Temp.:104.00, PWM:0
18:23:54.002 -> Water_Temp.:104.11, PWM:0
18:23:54.797 -> Water_Temp.:104.00, PWM:0
18:23:55.561 -> Water_Temp.:104.00, PWM:0
18:23:56.357 -> Water_Temp.:104.00, PWM:0
18:23:57.147 -> Water_Temp.:104.00, PWM:0
18:23:57.909 -> Water_Temp.:104.00, PWM:0
18:23:58.706 -> Water_Temp.:104.00, PWM:0
18:23:59.470 -> Water_Temp.:104.00, PWM:0
18:24:00.270 -> Water_Temp.:103.89, PWM:0
18:24:01.031 -> Water_Temp.:103.89, PWM:0
18:24:01.830 -> Water_Temp.:103.89, PWM:0
18:24:02.591 -> Water_Temp.:103.89, PWM:0
18:24:03.389 -> Water_Temp.:103.89, PWM:0
18:24:04.155 -> Water_Temp.:103.89, PWM:0
18:24:04.953 -> Water_Temp.:103.89, PWM:0
18:24:05.715 -> Water_Temp.:103.89, PWM:0
18:24:06.477 -> Water_Temp.:103.89, PWM:0
18:24:07.272 -> Water_Temp.:103.77, PWM:0
18:24:08.067 -> Water_Temp.:103.77, PWM:0
18:24:08.826 -> Water_Temp.:103.77, PWM:0
18:24:09.623 -> Water_Temp.:103.77, PWM:0
18:24:10.385 -> Water_Temp.:103.77, PWM:0
18:24:11.181 -> Water_Temp.:103.77, PWM:0
18:24:11.940 -> Water_Temp.:103.77, PWM:0
18:24:12.730 -> Water_Temp.:103.66, PWM:0
18:24:13.526 -> Water_Temp.:103.66, PWM:0
18:24:14.285 -> Water_Temp.:103.66, PWM:0
18:24:15.077 -> Water_Temp.:103.66, PWM:0
18:24:15.840 -> Water_Temp.:103.66, PWM:0
18:24:16.640 -> Water_Temp.:103.66, PWM:0
18:24:17.401 -> Water_Temp.:103.66, PWM:0
18:24:18.194 -> Water_Temp.:103.66, PWM:0
18:24:18.957 -> Water_Temp.:103.66, PWM:0
18:24:19.754 -> Water_Temp.:103.66, PWM:0
18:24:20.540 -> Water_Temp.:103.66, PWM:0
18:24:21.299 -> Water_Temp.:103.55, PWM:0
18:24:22.089 -> Water_Temp.:103.55, PWM:0
18:24:22.883 -> Water_Temp.:103.55, PWM:0
18:24:23.642 -> Water_Temp.:103.55, PWM:0
18:24:24.438 -> Water_Temp.:103.55, PWM:0
18:24:25.203 -> Water_Temp.:103.55, PWM:0
18:24:26.001 -> Water_Temp.:103.55, PWM:0
18:24:26.761 -> Water_Temp.:103.44, PWM:0
18:24:27.555 -> Water_Temp.:103.44, PWM:0
18:24:28.347 -> Water_Temp.:103.44, PWM:0
18:24:29.109 -> Water_Temp.:103.44, PWM:0
18:24:29.905 -> Water_Temp.:103.44, PWM:0
18:24:30.667 -> Water_Temp.:103.44, PWM:0
18:24:31.465 -> Water_Temp.:103.44, PWM:0
18:24:32.225 -> Water_Temp.:103.44, PWM:0
18:24:33.019 -> Water_Temp.:103.32, PWM:0
18:24:33.780 -> Water_Temp.:103.32, PWM:0
18:24:34.573 -> Water_Temp.:103.32, PWM:0
18:24:35.369 -> Water_Temp.:103.32, PWM:0
18:24:36.130 -> Water_Temp.:103.32, PWM:0
18:24:36.919 -> Water_Temp.:103.32, PWM:0
18:24:37.682 -> Water_Temp.:103.21, PWM:0
18:24:38.479 -> Water_Temp.:103.21, PWM:0
18:24:39.243 -> Water_Temp.:103.21, PWM:0
18:24:40.037 -> Water_Temp.:103.21, PWM:0
18:24:40.796 -> Water_Temp.:103.21, PWM:0
18:24:41.594 -> Water_Temp.:103.21, PWM:0
18:24:42.357 -> Water_Temp.:103.21, PWM:0
18:24:43.152 -> Water_Temp.:103.21, PWM:0
18:24:43.949 -> Water_Temp.:103.21, PWM:0
18:24:44.714 -> Water_Temp.:103.10, PWM:0
18:24:45.478 -> Water_Temp.:103.10, PWM:0
18:24:46.275 -> Water_Temp.:103.10, PWM:0
18:24:47.070 -> Water_Temp.:103.10, PWM:0
18:24:47.833 -> Water_Temp.:103.10, PWM:0
18:24:48.626 -> Water_Temp.:103.10, PWM:0
18:24:49.391 -> Water_Temp.:103.10, PWM:0
18:24:50.184 -> Water_Temp.:103.10, PWM:0
18:24:50.945 -> Water_Temp.:102.99, PWM:0
18:24:51.736 -> Water_Temp.:103.10, PWM:0
18:24:52.497 -> Water_Temp.:102.99, PWM:0
18:24:53.293 -> Water_Temp.:102.99, PWM:0
18:24:54.056 -> Water_Temp.:102.99, PWM:0
18:24:54.847 -> Water_Temp.:102.99, PWM:0
18:24:55.648 -> Water_Temp.:102.99, PWM:0
18:24:56.412 -> Water_Temp.:102.99, PWM:0
18:24:57.205 -> Water_Temp.:102.99, PWM:0
18:24:57.963 -> Water_Temp.:102.99, PWM:0
18:24:58.755 -> Water_Temp.:102.99, PWM:0
18:24:59.519 -> Water_Temp.:102.99, PWM:0
18:25:00.311 -> Water_Temp.:102.99, PWM:0
18:25:01.103 -> Water_Temp.:102.99, PWM:0
18:25:01.866 -> Water_Temp.:102.99, PWM:0
18:25:02.659 -> Water_Temp.:102.87, PWM:0
18:25:03.419 -> Water_Temp.:102.87, PWM:0
18:25:04.214 -> Water_Temp.:102.87, PWM:0
18:25:04.975 -> Water_Temp.:102.87, PWM:0
18:25:05.776 -> Water_Temp.:102.87, PWM:0
18:25:06.536 -> Water_Temp.:102.87, PWM:0
18:25:07.334 -> Water_Temp.:102.87, PWM:0
18:25:08.098 -> Water_Temp.:102.87, PWM:0
18:25:08.893 -> Water_Temp.:102.87, PWM:0
18:25:09.688 -> Water_Temp.:102.87, PWM:0
18:25:10.450 -> Water_Temp.:102.87, PWM:0
18:25:11.243 -> Water_Temp.:102.87, PWM:0
18:25:12.003 -> Water_Temp.:102.87, PWM:0
18:25:12.796 -> Water_Temp.:102.87, PWM:0
18:25:13.554 -> Water_Temp.:102.76, PWM:0
18:25:14.346 -> Water_Temp.:102.76, PWM:0
18:25:15.139 -> Water_Temp.:102.76, PWM:0
18:25:15.901 -> Water_Temp.:102.76, PWM:0
18:25:16.695 -> Water_Temp.:102.76, PWM:0
18:25:17.455 -> Water_Temp.:102.76, PWM:0
18:25:18.248 -> Water_Temp.:102.76, PWM:0
18:25:19.047 -> Water_Temp.:102.76, PWM:0
18:25:19.809 -> Water_Temp.:102.76, PWM:0
18:25:20.605 -> Water_Temp.:102.76, PWM:0
18:25:21.361 -> Water_Temp.:102.76, PWM:0
18:25:22.158 -> Water_Temp.:102.76, PWM:0
18:25:22.918 -> Water_Temp.:102.76, PWM:0
18:25:23.714 -> Water_Temp.:102.76, PWM:0
18:25:24.475 -> Water_Temp.:102.76, PWM:0
18:25:25.271 -> Water_Temp.:102.76, PWM:0
18:25:26.036 -> Water_Temp.:102.76, PWM:0
18:25:26.831 -> Water_Temp.:102.76, PWM:0
18:25:27.621 -> Water_Temp.:102.65, PWM:0
18:25:28.384 -> Water_Temp.:102.65, PWM:0
18:25:29.182 -> Water_Temp.:102.65, PWM:0
18:25:29.941 -> Water_Temp.:102.65, PWM:0
18:25:30.735 -> Water_Temp.:102.65, PWM:0
18:25:31.496 -> Water_Temp.:102.65, PWM:0
18:25:32.291 -> Water_Temp.:102.65, PWM:0
18:25:33.087 -> Water_Temp.:102.65, PWM:0
18:25:33.841 -> Water_Temp.:102.65, PWM:0
18:25:34.635 -> Water_Temp.:102.65, PWM:0
18:25:35.399 -> Water_Temp.:102.65, PWM:0
18:25:36.196 -> Water_Temp.:102.65, PWM:0
18:25:36.956 -> Water_Temp.:102.65, PWM:0
18:25:37.751 -> Water_Temp.:102.65, PWM:0
18:25:38.514 -> Water_Temp.:102.54, PWM:0
18:25:39.313 -> Water_Temp.:102.65, PWM:0
18:25:40.103 -> Water_Temp.:102.54, PWM:0
18:25:40.858 -> Water_Temp.:102.54, PWM:0
18:25:41.653 -> Water_Temp.:102.54, PWM:0
18:25:42.412 -> Water_Temp.:102.54, PWM:0
18:25:43.208 -> Water_Temp.:102.54, PWM:0
18:25:43.998 -> Water_Temp.:102.54, PWM:0
18:25:44.757 -> Water_Temp.:102.54, PWM:0
18:25:45.551 -> Water_Temp.:102.54, PWM:0
18:25:46.312 -> Water_Temp.:102.54, PWM:0
18:25:47.108 -> Water_Temp.:102.54, PWM:0
18:25:47.905 -> Water_Temp.:102.42, PWM:0
18:25:48.666 -> Water_Temp.:102.42, PWM:0
18:25:49.465 -> Water_Temp.:102.42, PWM:0
18:25:50.229 -> Water_Temp.:102.42, PWM:0
18:25:50.990 -> Water_Temp.:102.42, PWM:0
18:25:51.787 -> Water_Temp.:102.42, PWM:0
18:25:52.580 -> Water_Temp.:102.42, PWM:0
18:25:53.344 -> Water_Temp.:102.31, PWM:0
18:25:54.144 -> Water_Temp.:102.31, PWM:0
18:25:54.902 -> Water_Temp.:102.31, PWM:0
18:25:55.701 -> Water_Temp.:102.31, PWM:0
18:25:56.462 -> Water_Temp.:102.31, PWM:0
18:25:57.254 -> Water_Temp.:102.31, PWM:0
18:25:58.013 -> Water_Temp.:102.31, PWM:0
18:25:58.809 -> Water_Temp.:102.31, PWM:0
18:25:59.603 -> Water_Temp.:102.20, PWM:0
18:26:00.360 -> Water_Temp.:102.20, PWM:0
18:26:01.154 -> Water_Temp.:102.20, PWM:0
18:26:01.915 -> Water_Temp.:102.20, PWM:0
18:26:02.714 -> Water_Temp.:102.20, PWM:0
18:26:03.473 -> Water_Temp.:102.20, PWM:0
18:26:04.270 -> Water_Temp.:102.20, PWM:0
18:26:05.030 -> Water_Temp.:102.20, PWM:0
18:26:05.827 -> Water_Temp.:102.20, PWM:0
18:26:06.619 -> Water_Temp.:102.20, PWM:0
18:26:07.380 -> Water_Temp.:102.20, PWM:0
18:26:08.171 -> Water_Temp.:102.20, PWM:0
18:26:08.930 -> Water_Temp.:102.20, PWM:0
18:26:09.731 -> Water_Temp.:102.20, PWM:0
18:26:10.523 -> Water_Temp.:102.20, PWM:0
18:26:11.281 -> Water_Temp.:102.20, PWM:0
18:26:12.079 -> Water_Temp.:102.20, PWM:0
18:26:12.837 -> Water_Temp.:102.20, PWM:0
18:26:13.629 -> Water_Temp.:102.20, PWM:0
18:26:14.390 -> Water_Temp.:102.20, PWM:0
18:26:15.184 -> Water_Temp.:102.20, PWM:0
18:26:15.977 -> Water_Temp.:102.20, PWM:0
18:26:16.740 -> Water_Temp.:102.20, PWM:0
18:26:17.535 -> Water_Temp.:102.20, PWM:0
18:26:18.291 -> Water_Temp.:102.20, PWM:0
18:26:19.084 -> Water_Temp.:102.20, PWM:0
18:26:19.847 -> Water_Temp.:102.20, PWM:0
18:26:20.644 -> Water_Temp.:102.20, PWM:0
18:26:21.435 -> Water_Temp.:102.20, PWM:0
18:26:22.193 -> Water_Temp.:102.20, PWM:0
18:26:22.988 -> Water_Temp.:102.20, PWM:0
18:26:23.748 -> Water_Temp.:102.20, PWM:0
18:26:24.543 -> Water_Temp.:102.09, PWM:0
18:26:25.340 -> Water_Temp.:102.09, PWM:0
18:26:26.097 -> Water_Temp.:102.09, PWM:0
18:26:26.897 -> Water_Temp.:102.09, PWM:0
18:26:27.658 -> Water_Temp.:102.09, PWM:0
18:26:28.449 -> Water_Temp.:102.09, PWM:0
18:26:29.210 -> Water_Temp.:102.09, PWM:0
18:26:30.008 -> Water_Temp.:102.09, PWM:0
18:26:30.768 -> Water_Temp.:102.09, PWM:0
18:26:31.560 -> Water_Temp.:101.97, PWM:0
18:26:32.354 -> Water_Temp.:102.09, PWM:0
18:26:33.116 -> Water_Temp.:101.97, PWM:0
18:26:33.909 -> Water_Temp.:101.97, PWM:0
18:26:34.669 -> Water_Temp.:101.97, PWM:0
18:26:35.466 -> Water_Temp.:101.97, PWM:0
18:26:36.259 -> Water_Temp.:101.97, PWM:0
18:26:37.017 -> Water_Temp.:101.97, PWM:0
18:26:37.810 -> Water_Temp.:101.97, PWM:0
18:26:38.570 -> Water_Temp.:101.97, PWM:0
18:26:39.370 -> Water_Temp.:101.97, PWM:0
18:26:40.132 -> Water_Temp.:101.97, PWM:0
18:26:40.923 -> Water_Temp.:101.97, PWM:0
18:26:41.719 -> Water_Temp.:101.86, PWM:0
18:26:42.476 -> Water_Temp.:101.86, PWM:0
18:26:43.270 -> Water_Temp.:101.86, PWM:0
18:26:44.026 -> Water_Temp.:101.86, PWM:0
18:26:44.820 -> Water_Temp.:101.86, PWM:0
18:26:45.620 -> Water_Temp.:101.86, PWM:0
18:26:46.380 -> Water_Temp.:101.86, PWM:0
18:26:47.168 -> Water_Temp.:101.86, PWM:0
18:26:47.931 -> Water_Temp.:101.86, PWM:0
18:26:48.722 -> Water_Temp.:101.86, PWM:0
18:26:49.518 -> Water_Temp.:101.86, PWM:0
18:26:50.277 -> Water_Temp.:101.86, PWM:0
18:26:51.071 -> Water_Temp.:101.86, PWM:0
18:26:51.831 -> Water_Temp.:101.86, PWM:0
18:26:52.627 -> Water_Temp.:101.86, PWM:0
18:26:53.384 -> Water_Temp.:101.86, PWM:0
18:26:54.182 -> Water_Temp.:101.86, PWM:0
18:26:54.978 -> Water_Temp.:101.86, PWM:0
18:26:55.745 -> Water_Temp.:101.86, PWM:0
18:26:56.506 -> Water_Temp.:101.86, PWM:0
18:26:57.302 -> Water_Temp.:101.86, PWM:0
18:26:58.099 -> Water_Temp.:101.86, PWM:0
18:26:58.854 -> Water_Temp.:101.86, PWM:0
18:26:59.651 -> Water_Temp.:101.86, PWM:0
18:27:00.413 -> Water_Temp.:101.86, PWM:0
18:27:01.204 -> Water_Temp.:101.86, PWM:0
18:27:01.996 -> Water_Temp.:101.86, PWM:0
18:27:02.757 -> Water_Temp.:101.86, PWM:0
18:27:03.550 -> Water_Temp.:101.86, PWM:0
18:27:04.304 -> Water_Temp.:101.75, PWM:0
18:27:05.099 -> Water_Temp.:101.75, PWM:0
18:27:05.898 -> Water_Temp.:101.75, PWM:0
18:27:06.657 -> Water_Temp.:101.75, PWM:0
18:27:07.452 -> Water_Temp.:101.75, PWM:0
18:27:08.213 -> Water_Temp.:101.75, PWM:0
18:27:09.008 -> Water_Temp.:101.75, PWM:0
18:27:09.767 -> Water_Temp.:101.75, PWM:0
18:27:10.561 -> Water_Temp.:101.75, PWM:0
18:27:11.354 -> Water_Temp.:101.75, PWM:0
18:27:12.116 -> Water_Temp.:101.75, PWM:0
18:27:12.909 -> Water_Temp.:101.75, PWM:0
18:27:13.699 -> Water_Temp.:101.75, PWM:0
18:27:14.460 -> Water_Temp.:101.75, PWM:0
18:27:15.223 -> Water_Temp.:101.75, PWM:0
18:27:16.019 -> Water_Temp.:101.75, PWM:0
18:27:16.813 -> Water_Temp.:101.75, PWM:0
18:27:17.566 -> Water_Temp.:101.75, PWM:0
18:27:18.361 -> Water_Temp.:101.75, PWM:0
18:27:19.154 -> Water_Temp.:101.64, PWM:0
18:27:19.917 -> Water_Temp.:101.64, PWM:0
18:27:20.704 -> Water_Temp.:101.64, PWM:0
18:27:21.464 -> Water_Temp.:101.64, PWM:0
18:27:22.253 -> Water_Temp.:101.64, PWM:0
18:27:23.052 -> Water_Temp.:101.64, PWM:0
18:27:23.808 -> Water_Temp.:101.64, PWM:0
18:27:24.601 -> Water_Temp.:101.64, PWM:0
18:27:25.396 -> Water_Temp.:101.64, PWM:0
18:27:26.157 -> Water_Temp.:101.52, PWM:0
18:27:26.954 -> Water_Temp.:101.52, PWM:0
18:27:27.713 -> Water_Temp.:101.52, PWM:0
18:27:28.506 -> Water_Temp.:101.52, PWM:0
18:27:29.266 -> Water_Temp.:101.52, PWM:0
18:27:30.059 -> Water_Temp.:101.52, PWM:0
18:27:30.852 -> Water_Temp.:101.52, PWM:0
18:27:31.614 -> Water_Temp.:101.52, PWM:0
18:27:32.409 -> Water_Temp.:101.41, PWM:0
18:27:33.196 -> Water_Temp.:101.41, PWM:0
18:27:33.953 -> Water_Temp.:101.41, PWM:0
18:27:34.745 -> Water_Temp.:101.41, PWM:0
18:27:35.508 -> Water_Temp.:101.41, PWM:0
18:27:36.303 -> Water_Temp.:101.41, PWM:0
18:27:37.064 -> Water_Temp.:101.41, PWM:0
18:27:37.858 -> Water_Temp.:101.41, PWM:0
18:27:38.655 -> Water_Temp.:101.41, PWM:0
18:27:39.417 -> Water_Temp.:101.41, PWM:0
18:27:40.211 -> Water_Temp.:101.30, PWM:0
18:27:40.972 -> Water_Temp.:101.30, PWM:0
18:27:41.764 -> Water_Temp.:101.30, PWM:0
18:27:42.554 -> Water_Temp.:101.30, PWM:0
18:27:43.314 -> Water_Temp.:101.30, PWM:0
18:27:44.112 -> Water_Temp.:101.30, PWM:0
18:27:44.873 -> Water_Temp.:101.30, PWM:0
18:27:45.670 -> Water_Temp.:101.30, PWM:0
18:27:46.432 -> Water_Temp.:101.30, PWM:0
18:27:47.227 -> Water_Temp.:101.30, PWM:0
18:27:47.987 -> Water_Temp.:101.30, PWM:0
18:27:48.780 -> Water_Temp.:101.30, PWM:0
18:27:49.541 -> Water_Temp.:101.30, PWM:0
18:27:50.339 -> Water_Temp.:101.30, PWM:0
18:27:51.133 -> Water_Temp.:101.30, PWM:0
18:27:51.897 -> Water_Temp.:101.30, PWM:0
18:27:52.690 -> Water_Temp.:101.30, PWM:0
18:27:53.447 -> Water_Temp.:101.30, PWM:0
18:27:54.241 -> Water_Temp.:101.30, PWM:0
18:27:55.002 -> Water_Temp.:101.30, PWM:0
18:27:55.800 -> Water_Temp.:101.30, PWM:0
18:27:56.594 -> Water_Temp.:101.30, PWM:0
18:27:57.357 -> Water_Temp.:101.30, PWM:0
18:27:58.151 -> Water_Temp.:101.30, PWM:0
18:27:58.913 -> Water_Temp.:101.30, PWM:0
18:27:59.711 -> Water_Temp.:101.30, PWM:0
18:28:00.466 -> Water_Temp.:101.19, PWM:0
18:28:01.260 -> Water_Temp.:101.19, PWM:0
18:28:02.053 -> Water_Temp.:101.19, PWM:0
18:28:02.813 -> Water_Temp.:101.19, PWM:0
18:28:03.608 -> Water_Temp.:101.19, PWM:0
18:28:04.370 -> Water_Temp.:101.19, PWM:0
18:28:05.169 -> Water_Temp.:101.19, PWM:0
18:28:05.934 -> Water_Temp.:101.19, PWM:0
18:28:06.723 -> Water_Temp.:101.19, PWM:0
18:28:07.514 -> Water_Temp.:101.19, PWM:0
18:28:08.273 -> Water_Temp.:101.19, PWM:0
18:28:09.069 -> Water_Temp.:101.19, PWM:0
18:28:09.834 -> Water_Temp.:101.19, PWM:0
18:28:10.627 -> Water_Temp.:101.19, PWM:0
18:28:11.391 -> Water_Temp.:101.19, PWM:0
18:28:12.187 -> Water_Temp.:101.19, PWM:0
18:28:12.947 -> Water_Temp.:101.19, PWM:0
18:28:13.737 -> Water_Temp.:101.07, PWM:0
18:28:14.530 -> Water_Temp.:101.19, PWM:0
18:28:15.295 -> Water_Temp.:101.07, PWM:0
18:28:16.089 -> Water_Temp.:101.07, PWM:0
18:28:16.848 -> Water_Temp.:101.07, PWM:0
18:28:17.639 -> Water_Temp.:101.07, PWM:0
18:28:18.431 -> Water_Temp.:101.07, PWM:0
18:28:19.191 -> Water_Temp.:101.07, PWM:0
18:28:19.985 -> Water_Temp.:101.07, PWM:0
18:28:20.750 -> Water_Temp.:101.07, PWM:0
18:28:21.544 -> Water_Temp.:101.07, PWM:0
18:28:22.304 -> Water_Temp.:101.07, PWM:0
18:28:23.099 -> Water_Temp.:101.07, PWM:0
18:28:23.860 -> Water_Temp.:101.07, PWM:0
18:28:24.656 -> Water_Temp.:100.96, PWM:0
18:28:25.452 -> Water_Temp.:101.07, PWM:0
18:28:26.214 -> Water_Temp.:101.07, PWM:0
18:28:27.008 -> Water_Temp.:100.96, PWM:0
18:28:27.769 -> Water_Temp.:100.96, PWM:0
18:28:28.558 -> Water_Temp.:100.96, PWM:0
18:28:29.320 -> Water_Temp.:100.96, PWM:0
18:28:30.110 -> Water_Temp.:100.96, PWM:0
18:28:30.902 -> Water_Temp.:100.96, PWM:0
18:28:31.665 -> Water_Temp.:100.96, PWM:0
18:28:32.459 -> Water_Temp.:100.96, PWM:0
18:28:33.251 -> Water_Temp.:100.96, PWM:0
18:28:34.010 -> Water_Temp.:100.96, PWM:0
18:28:34.807 -> Water_Temp.:100.96, PWM:0
18:28:35.571 -> Water_Temp.:100.96, PWM:0
18:28:36.369 -> Water_Temp.:100.96, PWM:0
18:28:37.127 -> Water_Temp.:100.96, PWM:0
18:28:37.923 -> Water_Temp.:100.96, PWM:0
18:28:38.712 -> Water_Temp.:100.96, PWM:0
18:28:39.473 -> Water_Temp.:100.85, PWM:0
18:28:40.269 -> Water_Temp.:100.85, PWM:0
18:28:41.031 -> Water_Temp.:100.85, PWM:0
18:28:41.828 -> Water_Temp.:100.85, PWM:0
18:28:42.593 -> Water_Temp.:100.85, PWM:0
18:28:43.388 -> Water_Temp.:100.85, PWM:0
18:28:44.149 -> Water_Temp.:100.85, PWM:0
18:28:44.944 -> Water_Temp.:100.85, PWM:0
18:28:45.709 -> Water_Temp.:100.85, PWM:0
18:28:46.503 -> Water_Temp.:100.74, PWM:0
18:28:47.267 -> Water_Temp.:100.74, PWM:0
18:28:48.058 -> Water_Temp.:100.74, PWM:0
18:28:48.849 -> Water_Temp.:100.74, PWM:0
18:28:49.611 -> Water_Temp.:100.74, PWM:0
18:28:50.407 -> Water_Temp.:100.74, PWM:0
18:28:51.170 -> Water_Temp.:100.74, PWM:0
18:28:51.960 -> Water_Temp.:100.74, PWM:0
18:28:52.723 -> Water_Temp.:100.74, PWM:0
18:28:53.517 -> Water_Temp.:100.74, PWM:0
18:28:54.313 -> Water_Temp.:100.74, PWM:0
18:28:55.075 -> Water_Temp.:100.62, PWM:0
18:28:55.868 -> Water_Temp.:100.62, PWM:0
18:28:56.634 -> Water_Temp.:100.62, PWM:0
18:28:57.426 -> Water_Temp.:100.62, PWM:0
18:28:58.181 -> Water_Temp.:100.62, PWM:0
18:28:58.973 -> Water_Temp.:100.62, PWM:0
18:28:59.771 -> Water_Temp.:100.62, PWM:0
18:29:00.529 -> Water_Temp.:100.62, PWM:0
18:29:01.324 -> Water_Temp.:100.62, PWM:0
18:29:02.085 -> Water_Temp.:100.62, PWM:0
18:29:02.885 -> Water_Temp.:100.62, PWM:0
18:29:03.646 -> Water_Temp.:100.62, PWM:0
18:29:04.441 -> Water_Temp.:100.62, PWM:0
18:29:05.205 -> Water_Temp.:100.62, PWM:0
18:29:05.996 -> Water_Temp.:100.62, PWM:0
18:29:06.759 -> Water_Temp.:100.62, PWM:0
18:29:07.556 -> Water_Temp.:100.62, PWM:0
18:29:08.347 -> Water_Temp.:100.62, PWM:0
18:29:09.105 -> Water_Temp.:100.62, PWM:0
18:29:09.901 -> Water_Temp.:100.62, PWM:0
18:29:10.656 -> Water_Temp.:100.62, PWM:0
18:29:11.452 -> Water_Temp.:100.62, PWM:0
18:29:12.215 -> Water_Temp.:100.62, PWM:0
18:29:13.012 -> Water_Temp.:100.62, PWM:0
18:29:13.808 -> Water_Temp.:100.51, PWM:0
18:29:14.570 -> Water_Temp.:100.51, PWM:0
18:29:15.367 -> Water_Temp.:100.51, PWM:0
18:29:16.127 -> Water_Temp.:100.51, PWM:0
18:29:16.921 -> Water_Temp.:100.51, PWM:0
18:29:17.681 -> Water_Temp.:100.51, PWM:0
18:29:18.475 -> Water_Temp.:100.51, PWM:0
18:29:19.237 -> Water_Temp.:100.51, PWM:0
18:29:20.033 -> Water_Temp.:100.51, PWM:0
18:29:20.827 -> Water_Temp.:100.51, PWM:0
18:29:21.587 -> Water_Temp.:100.51, PWM:0
18:29:22.379 -> Water_Temp.:100.51, PWM:0
18:29:23.141 -> Water_Temp.:100.51, PWM:0
18:29:23.932 -> Water_Temp.:100.51, PWM:0
18:29:24.695 -> Water_Temp.:100.51, PWM:0
18:29:25.494 -> Water_Temp.:100.40, PWM:0
18:29:26.256 -> Water_Temp.:100.40, PWM:0
18:29:27.055 -> Water_Temp.:100.40, PWM:0
18:29:27.848 -> Water_Temp.:100.40, PWM:0
18:29:28.608 -> Water_Temp.:100.40, PWM:0
18:29:29.402 -> Water_Temp.:100.40, PWM:0
18:29:30.164 -> Water_Temp.:100.40, PWM:0
18:29:30.961 -> Water_Temp.:100.40, PWM:0
18:29:31.716 -> Water_Temp.:100.40, PWM:0
18:29:32.510 -> Water_Temp.:100.40, PWM:0
18:29:33.305 -> Water_Temp.:100.40, PWM:0
18:29:34.065 -> Water_Temp.:100.40, PWM:0
18:29:34.857 -> Water_Temp.:100.40, PWM:0
18:29:35.621 -> Water_Temp.:100.40, PWM:0
18:29:36.414 -> Water_Temp.:100.40, PWM:0
18:29:37.208 -> Water_Temp.:100.29, PWM:0
18:29:37.967 -> Water_Temp.:100.29, PWM:0
18:29:38.761 -> Water_Temp.:100.29, PWM:0
18:29:39.525 -> Water_Temp.:100.29, PWM:0
18:29:40.320 -> Water_Temp.:100.29, PWM:0
18:29:41.082 -> Water_Temp.:100.29, PWM:0
18:29:41.876 -> Water_Temp.:100.29, PWM:0
18:29:42.638 -> Water_Temp.:100.29, PWM:0
18:29:43.432 -> Water_Temp.:100.29, PWM:0
18:29:44.228 -> Water_Temp.:100.29, PWM:0
18:29:44.990 -> Water_Temp.:100.29, PWM:0
18:29:45.755 -> Water_Temp.:100.29, PWM:0
18:29:46.548 -> Water_Temp.:100.29, PWM:0
18:29:47.343 -> Water_Temp.:100.29, PWM:0
18:29:48.101 -> Water_Temp.:100.29, PWM:0
18:29:48.895 -> Water_Temp.:100.29, PWM:0
18:29:49.657 -> Water_Temp.:100.29, PWM:0
18:29:50.455 -> Water_Temp.:100.29, PWM:0
18:29:51.214 -> Water_Temp.:100.17, PWM:0
18:29:52.009 -> Water_Temp.:100.17, PWM:0
18:29:52.803 -> Water_Temp.:100.17, PWM:0
18:29:53.560 -> Water_Temp.:100.17, PWM:0
18:29:54.355 -> Water_Temp.:100.17, PWM:0
18:29:55.119 -> Water_Temp.:100.17, PWM:0
18:29:55.919 -> Water_Temp.:100.17, PWM:0
18:29:56.680 -> Water_Temp.:100.17, PWM:0
18:29:57.476 -> Water_Temp.:100.17, PWM:0
18:29:58.266 -> Water_Temp.:100.17, PWM:0
18:29:59.028 -> Water_Temp.:100.17, PWM:0
18:29:59.792 -> Water_Temp.:100.17, PWM:0
18:30:00.584 -> Water_Temp.:100.17, PWM:0
18:30:01.380 -> Water_Temp.:100.17, PWM:0
18:30:02.140 -> Water_Temp.:100.17, PWM:0
18:30:02.930 -> Water_Temp.:100.17, PWM:0
18:30:03.721 -> Water_Temp.:100.17, PWM:0
18:30:04.483 -> Water_Temp.:100.17, PWM:0
18:30:05.283 -> Water_Temp.:100.17, PWM:0
18:30:06.044 -> Water_Temp.:100.17, PWM:0
18:30:06.841 -> Water_Temp.:100.17, PWM:0
18:30:07.599 -> Water_Temp.:100.06, PWM:0
18:30:08.394 -> Water_Temp.:100.06, PWM:0
18:30:09.156 -> Water_Temp.:100.06, PWM:0
18:30:09.955 -> Water_Temp.:100.06, PWM:0
18:30:10.720 -> Water_Temp.:100.06, PWM:0
18:30:11.511 -> Water_Temp.:100.06, PWM:0
18:30:12.272 -> Water_Temp.:100.06, PWM:0
18:30:13.063 -> Water_Temp.:100.06, PWM:0
18:30:13.853 -> Water_Temp.:100.06, PWM:0
18:30:14.611 -> Water_Temp.:100.06, PWM:0
18:30:15.409 -> Water_Temp.:100.06, PWM:0
18:30:16.203 -> Water_Temp.:100.06, PWM:0
18:30:16.967 -> Water_Temp.:100.06, PWM:0
18:30:17.759 -> Water_Temp.:100.06, PWM:0
18:30:18.520 -> Water_Temp.:100.06, PWM:0
18:30:19.318 -> Water_Temp.:100.06, PWM:0
18:30:20.079 -> Water_Temp.:100.06, PWM:0
18:30:20.873 -> Water_Temp.:99.95, PWM:0
18:30:21.635 -> Water_Temp.:99.95, PWM:0
18:30:22.426 -> Water_Temp.:99.95, PWM:0
18:30:23.221 -> Water_Temp.:99.95, PWM:0
18:30:23.981 -> Water_Temp.:99.95, PWM:0
18:30:24.770 -> Water_Temp.:99.95, PWM:0
18:30:25.536 -> Water_Temp.:99.95, PWM:0
18:30:26.334 -> Water_Temp.:99.95, PWM:0
18:30:27.092 -> Water_Temp.:99.95, PWM:0
18:30:27.887 -> Water_Temp.:99.95, PWM:0
18:30:28.680 -> Water_Temp.:99.95, PWM:0
18:30:29.441 -> Water_Temp.:99.95, PWM:0
18:30:30.237 -> Water_Temp.:99.95, PWM:0
18:30:30.995 -> Water_Temp.:99.95, PWM:0
18:30:31.792 -> Water_Temp.:99.95, PWM:0
18:30:32.553 -> Water_Temp.:99.95, PWM:0
18:30:33.352 -> Water_Temp.:99.95, PWM:0
18:30:34.110 -> Water_Temp.:99.95, PWM:0
18:30:34.902 -> Water_Temp.:99.95, PWM:0
18:30:35.699 -> Water_Temp.:99.95, PWM:0
18:30:36.455 -> Water_Temp.:99.95, PWM:0
18:30:37.255 -> Water_Temp.:99.95, PWM:0
18:30:38.016 -> Water_Temp.:99.95, PWM:0
18:30:38.810 -> Water_Temp.:99.84, PWM:0
18:30:39.574 -> Water_Temp.:99.84, PWM:0
18:30:40.368 -> Water_Temp.:99.84, PWM:0
18:30:41.160 -> Water_Temp.:99.84, PWM:0
18:30:41.920 -> Water_Temp.:99.84, PWM:0
18:30:42.716 -> Water_Temp.:99.84, PWM:0
18:30:43.474 -> Water_Temp.:99.84, PWM:0
18:30:44.270 -> Water_Temp.:99.84, PWM:0
18:30:45.031 -> Water_Temp.:99.84, PWM:0
18:30:45.832 -> Water_Temp.:99.84, PWM:0
18:30:46.593 -> Water_Temp.:99.84, PWM:0
18:30:47.392 -> Water_Temp.:99.84, PWM:0
18:30:48.152 -> Water_Temp.:99.84, PWM:0
18:30:48.947 -> Water_Temp.:99.84, PWM:0
18:30:49.715 -> Water_Temp.:99.72, PWM:0
18:30:50.512 -> Water_Temp.:99.72, PWM:0
18:30:51.271 -> Water_Temp.:99.72, PWM:0
18:30:52.066 -> Water_Temp.:99.72, PWM:0
18:30:52.860 -> Water_Temp.:99.72, PWM:0
18:30:53.623 -> Water_Temp.:99.72, PWM:0
18:30:54.415 -> Water_Temp.:99.72, PWM:0
18:30:55.174 -> Water_Temp.:99.72, PWM:0
18:30:55.969 -> Water_Temp.:99.72, PWM:0
18:30:56.731 -> Water_Temp.:99.72, PWM:0
18:30:57.526 -> Water_Temp.:99.72, PWM:0
18:30:58.321 -> Water_Temp.:99.72, PWM:0
18:30:59.082 -> Water_Temp.:99.72, PWM:0
18:30:59.877 -> Water_Temp.:99.72, PWM:0
18:31:00.640 -> Water_Temp.:99.72, PWM:0
18:31:01.435 -> Water_Temp.:99.61, PWM:0
18:31:02.193 -> Water_Temp.:99.61, PWM:0
18:31:02.987 -> Water_Temp.:99.61, PWM:0
18:31:03.777 -> Water_Temp.:99.61, PWM:0
18:31:04.539 -> Water_Temp.:99.61, PWM:0
18:31:05.303 -> Water_Temp.:99.61, PWM:0
18:31:06.102 -> Water_Temp.:99.61, PWM:0
18:31:06.894 -> Water_Temp.:99.61, PWM:0
18:31:07.647 -> Water_Temp.:99.61, PWM:0
18:31:08.444 -> Water_Temp.:99.61, PWM:0
18:31:09.238 -> Water_Temp.:99.61, PWM:0
18:31:10.001 -> Water_Temp.:99.61, PWM:0
18:31:10.795 -> Water_Temp.:99.61, PWM:0
18:31:11.550 -> Water_Temp.:99.61, PWM:0
18:31:12.345 -> Water_Temp.:99.50, PWM:0
18:31:13.105 -> Water_Temp.:99.61, PWM:0
18:31:13.903 -> Water_Temp.:99.50, PWM:0
18:31:14.666 -> Water_Temp.:99.61, PWM:0
18:31:15.461 -> Water_Temp.:99.50, PWM:0
18:31:16.253 -> Water_Temp.:99.50, PWM:0
18:31:17.016 -> Water_Temp.:99.50, PWM:0
18:31:17.811 -> Water_Temp.:99.50, PWM:0
18:31:18.571 -> Water_Temp.:99.50, PWM:0
18:31:19.369 -> Water_Temp.:99.50, PWM:0
18:31:20.131 -> Water_Temp.:99.50, PWM:0
18:31:20.930 -> Water_Temp.:99.50, PWM:0
18:31:21.687 -> Water_Temp.:99.50, PWM:0
18:31:22.482 -> Water_Temp.:99.50, PWM:0
18:31:23.275 -> Water_Temp.:99.50, PWM:0
18:31:24.029 -> Water_Temp.:99.50, PWM:0
18:31:24.825 -> Water_Temp.:99.50, PWM:0
18:31:25.584 -> Water_Temp.:99.50, PWM:0
18:31:26.377 -> Water_Temp.:99.50, PWM:0
18:31:27.167 -> Water_Temp.:99.50, PWM:0
18:31:27.929 -> Water_Temp.:99.50, PWM:0
18:31:28.725 -> Water_Temp.:99.50, PWM:0
18:31:29.490 -> Water_Temp.:99.50, PWM:0
18:31:30.284 -> Water_Temp.:99.50, PWM:0
18:31:31.046 -> Water_Temp.:99.50, PWM:0
18:31:31.843 -> Water_Temp.:99.50, PWM:0
18:31:32.606 -> Water_Temp.:99.50, PWM:0
18:31:33.401 -> Water_Temp.:99.50, PWM:0
18:31:34.192 -> Water_Temp.:99.39, PWM:0
18:31:34.951 -> Water_Temp.:99.39, PWM:0
18:31:35.747 -> Water_Temp.:99.39, PWM:0
18:31:36.501 -> Water_Temp.:99.39, PWM:0
18:31:37.291 -> Water_Temp.:99.39, PWM:0
18:31:38.084 -> Water_Temp.:99.39, PWM:0
18:31:38.848 -> Water_Temp.:99.39, PWM:0
18:31:39.646 -> Water_Temp.:99.39, PWM:0
18:31:40.401 -> Water_Temp.:99.39, PWM:0
18:31:41.196 -> Water_Temp.:99.39, PWM:0
18:31:41.991 -> Water_Temp.:99.39, PWM:0
18:31:42.751 -> Water_Temp.:99.39, PWM:0
18:31:43.543 -> Water_Temp.:99.39, PWM:0
18:31:44.300 -> Water_Temp.:99.39, PWM:0
18:31:45.097 -> Water_Temp.:99.39, PWM:0
18:31:45.892 -> Water_Temp.:99.39, PWM:0
18:31:46.654 -> Water_Temp.:99.39, PWM:0
18:31:47.448 -> Water_Temp.:99.39, PWM:0
18:31:48.207 -> Water_Temp.:99.27, PWM:0
18:31:49.002 -> Water_Temp.:99.27, PWM:0
18:31:49.766 -> Water_Temp.:99.27, PWM:0
18:31:50.560 -> Water_Temp.:99.27, PWM:0
18:31:51.347 -> Water_Temp.:99.27, PWM:0
18:31:52.102 -> Water_Temp.:99.27, PWM:0
18:31:52.893 -> Water_Temp.:99.27, PWM:0
18:31:53.684 -> Water_Temp.:99.27, PWM:0
18:31:54.443 -> Water_Temp.:99.27, PWM:0
18:31:55.240 -> Water_Temp.:99.27, PWM:0
18:31:56.004 -> Water_Temp.:99.27, PWM:0
18:31:56.797 -> Water_Temp.:99.27, PWM:0
18:31:57.589 -> Water_Temp.:99.27, PWM:0
18:31:58.347 -> Water_Temp.:99.27, PWM:0
18:31:59.142 -> Water_Temp.:99.16, PWM:0
18:31:59.908 -> Water_Temp.:99.16, PWM:0
18:32:00.699 -> Water_Temp.:99.16, PWM:0
18:32:01.462 -> Water_Temp.:99.16, PWM:0
18:32:02.258 -> Water_Temp.:99.16, PWM:0
18:32:03.019 -> Water_Temp.:99.16, PWM:0
18:32:03.813 -> Water_Temp.:99.16, PWM:0
18:32:04.604 -> Water_Temp.:99.16, PWM:0
18:32:05.365 -> Water_Temp.:99.16, PWM:0
18:32:06.161 -> Water_Temp.:99.16, PWM:0
18:32:06.923 -> Water_Temp.:99.16, PWM:0
18:32:07.721 -> Water_Temp.:99.16, PWM:0
18:32:08.481 -> Water_Temp.:99.16, PWM:0
18:32:09.274 -> Water_Temp.:99.16, PWM:0
18:32:10.038 -> Water_Temp.:99.16, PWM:0
18:32:10.833 -> Water_Temp.:99.16, PWM:0
18:32:11.627 -> Water_Temp.:99.16, PWM:0
18:32:12.387 -> Water_Temp.:99.16, PWM:0
18:32:13.186 -> Water_Temp.:99.16, PWM:0
18:32:13.945 -> Water_Temp.:99.05, PWM:0
18:32:14.740 -> Water_Temp.:99.16, PWM:0
18:32:15.506 -> Water_Temp.:99.05, PWM:0
18:32:16.302 -> Water_Temp.:99.05, PWM:0
18:32:17.065 -> Water_Temp.:99.05, PWM:0
18:32:17.864 -> Water_Temp.:99.05, PWM:0
18:32:18.622 -> Water_Temp.:99.05, PWM:0
18:32:19.421 -> Water_Temp.:99.05, PWM:0
18:32:20.181 -> Water_Temp.:99.05, PWM:0
18:32:20.978 -> Water_Temp.:99.05, PWM:0
18:32:21.740 -> Water_Temp.:99.05, PWM:0
18:32:22.538 -> Water_Temp.:99.05, PWM:0
18:32:23.298 -> Water_Temp.:99.05, PWM:0
18:32:24.094 -> Water_Temp.:99.05, PWM:0
18:32:24.855 -> Water_Temp.:99.05, PWM:0
18:32:25.651 -> Water_Temp.:99.05, PWM:0
18:32:26.414 -> Water_Temp.:99.05, PWM:0
18:32:27.212 -> Water_Temp.:99.05, PWM:0
18:32:28.005 -> Water_Temp.:99.05, PWM:0
18:32:28.766 -> Water_Temp.:98.94, PWM:0
18:32:29.566 -> Water_Temp.:98.94, PWM:0
18:32:30.324 -> Water_Temp.:98.94, PWM:0
18:32:31.116 -> Water_Temp.:98.94, PWM:0
18:32:31.872 -> Water_Temp.:98.94, PWM:0
18:32:32.669 -> Water_Temp.:98.94, PWM:0
18:32:33.465 -> Water_Temp.:98.94, PWM:0
18:32:34.221 -> Water_Temp.:98.94, PWM:0
18:32:35.018 -> Water_Temp.:98.94, PWM:0
18:32:35.784 -> Water_Temp.:98.94, PWM:0
18:32:36.584 -> Water_Temp.:98.94, PWM:0
18:32:37.342 -> Water_Temp.:98.94, PWM:0
18:32:38.134 -> Water_Temp.:98.94, PWM:0
18:32:38.897 -> Water_Temp.:98.82, PWM:0
18:32:39.698 -> Water_Temp.:98.82, PWM:0
18:32:40.485 -> Water_Temp.:98.82, PWM:0
18:32:41.241 -> Water_Temp.:98.82, PWM:0
18:32:42.034 -> Water_Temp.:98.82, PWM:0
18:32:42.792 -> Water_Temp.:98.82, PWM:0
18:32:43.587 -> Water_Temp.:98.82, PWM:0
18:32:44.383 -> Water_Temp.:98.82, PWM:0
18:32:45.144 -> Water_Temp.:98.82, PWM:0
18:32:45.942 -> Water_Temp.:98.82, PWM:0
18:32:46.706 -> Water_Temp.:98.82, PWM:0
18:32:47.503 -> Water_Temp.:98.82, PWM:0
18:32:48.263 -> Water_Temp.:98.82, PWM:0
18:32:49.060 -> Water_Temp.:98.82, PWM:0
18:32:49.824 -> Water_Temp.:98.82, PWM:0
18:32:50.618 -> Water_Temp.:98.82, PWM:0
18:32:51.380 -> Water_Temp.:98.82, PWM:0
18:32:52.175 -> Water_Temp.:98.82, PWM:0
18:32:52.934 -> Water_Temp.:98.82, PWM:0
18:32:53.728 -> Water_Temp.:98.82, PWM:0
18:32:54.492 -> Water_Temp.:98.82, PWM:0
18:32:55.284 -> Water_Temp.:98.82, PWM:0
18:32:56.051 -> Water_Temp.:98.71, PWM:0
18:32:56.846 -> Water_Temp.:98.71, PWM:0
18:32:57.642 -> Water_Temp.:98.71, PWM:0
18:32:58.404 -> Water_Temp.:98.71, PWM:0
18:32:59.202 -> Water_Temp.:98.71, PWM:0
18:32:59.968 -> Water_Temp.:98.71, PWM:0
18:33:00.764 -> Water_Temp.:98.71, PWM:0
18:33:01.522 -> Water_Temp.:98.71, PWM:0
18:33:02.313 -> Water_Temp.:98.71, PWM:0
18:33:03.074 -> Water_Temp.:98.71, PWM:0
18:33:03.867 -> Water_Temp.:98.71, PWM:0
18:33:04.629 -> Water_Temp.:98.71, PWM:0
18:33:05.426 -> Water_Temp.:98.71, PWM:0
18:33:06.221 -> Water_Temp.:98.60, PWM:0
18:33:06.980 -> Water_Temp.:98.60, PWM:0
18:33:07.775 -> Water_Temp.:98.60, PWM:0
18:33:08.534 -> Water_Temp.:98.60, PWM:0
18:33:09.333 -> Water_Temp.:98.60, PWM:0
18:33:10.098 -> Water_Temp.:98.60, PWM:0
18:33:10.895 -> Water_Temp.:98.60, PWM:0
18:33:11.655 -> Water_Temp.:98.60, PWM:0
18:33:12.449 -> Water_Temp.:98.60, PWM:0
18:33:13.238 -> Water_Temp.:98.60, PWM:0
18:33:13.996 -> Water_Temp.:98.60, PWM:0
18:33:14.787 -> Water_Temp.:98.60, PWM:0
18:33:15.552 -> Water_Temp.:98.49, PWM:0
18:33:16.347 -> Water_Temp.:98.49, PWM:0
18:33:17.109 -> Water_Temp.:98.49, PWM:0
18:33:17.905 -> Water_Temp.:98.49, PWM:0
18:33:18.667 -> Water_Temp.:98.49, PWM:0
18:33:19.468 -> Water_Temp.:98.49, PWM:0
18:33:20.258 -> Water_Temp.:98.37, PWM:0
18:33:21.018 -> Water_Temp.:98.37, PWM:0
18:33:21.809 -> Water_Temp.:98.37, PWM:0
18:33:22.597 -> Water_Temp.:98.37, PWM:0
18:33:23.360 -> Water_Temp.:98.37, PWM:0
18:33:24.150 -> Water_Temp.:98.37, PWM:0
18:33:24.913 -> Water_Temp.:98.37, PWM:0
18:33:25.715 -> Water_Temp.:98.37, PWM:0
18:33:26.473 -> Water_Temp.:98.37, PWM:0
18:33:27.271 -> Water_Temp.:98.37, PWM:0
18:33:28.030 -> Water_Temp.:98.37, PWM:0
18:33:28.828 -> Water_Temp.:98.26, PWM:0
18:33:29.585 -> Water_Temp.:98.26, PWM:0
18:33:30.380 -> Water_Temp.:98.26, PWM:0
18:33:31.168 -> Water_Temp.:98.26, PWM:0
18:33:31.929 -> Water_Temp.:98.26, PWM:0
18:33:32.719 -> Water_Temp.:98.26, PWM:0
18:33:33.509 -> Water_Temp.:98.26, PWM:0
18:33:34.267 -> Water_Temp.:98.26, PWM:0
18:33:35.058 -> Water_Temp.:98.37, PWM:0
18:33:35.854 -> Water_Temp.:98.37, PWM:0
18:33:36.613 -> Water_Temp.:98.37, PWM:0
18:33:37.403 -> Water_Temp.:98.37, PWM:0
18:33:38.163 -> Water_Temp.:98.37, PWM:0
18:33:38.958 -> Water_Temp.:98.37, PWM:0
18:33:39.751 -> Water_Temp.:98.37, PWM:0
18:33:40.515 -> Water_Temp.:98.37, PWM:0
18:33:41.310 -> Water_Temp.:98.37, PWM:0
18:33:42.070 -> Water_Temp.:98.37, PWM:0
18:33:42.866 -> Water_Temp.:98.37, PWM:0
18:33:43.653 -> Water_Temp.:98.37, PWM:0
18:33:44.415 -> Water_Temp.:98.37, PWM:0
18:33:45.209 -> Water_Temp.:98.37, PWM:0
18:33:45.973 -> Water_Temp.:98.37, PWM:0
18:33:46.762 -> Water_Temp.:98.37, PWM:0
18:33:47.523 -> Water_Temp.:98.37, PWM:0
18:33:48.320 -> Water_Temp.:98.37, PWM:0
18:33:49.082 -> Water_Temp.:98.37, PWM:0
18:33:49.881 -> Water_Temp.:98.37, PWM:0
18:33:50.669 -> Water_Temp.:98.37, PWM:0
18:33:51.430 -> Water_Temp.:98.37, PWM:0
18:33:52.222 -> Water_Temp.:98.37, PWM:0
18:33:52.981 -> Water_Temp.:98.26, PWM:0
18:33:53.775 -> Water_Temp.:98.26, PWM:0
18:33:54.566 -> Water_Temp.:98.26, PWM:0
18:33:55.326 -> Water_Temp.:98.26, PWM:0
18:33:56.124 -> Water_Temp.:98.26, PWM:0
18:33:56.886 -> Water_Temp.:98.26, PWM:0
18:33:57.681 -> Water_Temp.:98.26, PWM:0
18:33:58.473 -> Water_Temp.:98.26, PWM:0
18:33:59.232 -> Water_Temp.:98.26, PWM:0
18:34:00.029 -> Water_Temp.:98.26, PWM:0
18:34:00.789 -> Water_Temp.:98.26, PWM:0
18:34:01.582 -> Water_Temp.:98.26, PWM:0
18:34:02.342 -> Water_Temp.:98.26, PWM:0
18:34:03.138 -> Water_Temp.:98.26, PWM:0
18:34:03.901 -> Water_Temp.:98.26, PWM:0
18:34:04.699 -> Water_Temp.:98.26, PWM:0
18:34:05.494 -> Water_Temp.:98.26, PWM:0
18:34:06.254 -> Water_Temp.:98.26, PWM:0
18:34:07.046 -> Water_Temp.:98.26, PWM:0
18:34:07.802 -> Water_Temp.:98.15, PWM:0
18:34:08.601 -> Water_Temp.:98.15, PWM:0
18:34:09.364 -> Water_Temp.:98.15, PWM:0
18:34:10.160 -> Water_Temp.:98.15, PWM:0
18:34:10.954 -> Water_Temp.:98.15, PWM:0
18:34:11.715 -> Water_Temp.:98.15, PWM:0
18:34:12.504 -> Water_Temp.:98.15, PWM:0
18:34:13.265 -> Water_Temp.:98.15, PWM:0
18:34:14.056 -> Water_Temp.:98.15, PWM:0
18:34:14.852 -> Water_Temp.:98.15, PWM:0
18:34:15.612 -> Water_Temp.:98.15, PWM:0
18:34:16.406 -> Water_Temp.:98.15, PWM:0
18:34:17.167 -> Water_Temp.:98.15, PWM:0
18:34:17.963 -> Water_Temp.:98.15, PWM:0
18:34:18.751 -> Water_Temp.:98.04, PWM:0
18:34:19.510 -> Water_Temp.:98.04, PWM:0
18:34:20.302 -> Water_Temp.:98.04, PWM:0
18:34:21.062 -> Water_Temp.:98.04, PWM:0
18:34:21.855 -> Water_Temp.:98.04, PWM:0
18:34:22.647 -> Water_Temp.:98.04, PWM:0
18:34:23.402 -> Water_Temp.:98.04, PWM:0
18:34:24.200 -> Water_Temp.:98.04, PWM:0
18:34:24.958 -> Water_Temp.:98.04, PWM:0
18:34:25.751 -> Water_Temp.:98.04, PWM:0
18:34:26.551 -> Water_Temp.:98.04, PWM:0
18:34:27.310 -> Water_Temp.:98.04, PWM:0
18:34:28.108 -> Water_Temp.:98.04, PWM:0
18:34:28.865 -> Water_Temp.:98.04, PWM:0
18:34:29.661 -> Water_Temp.:98.04, PWM:0
18:34:30.423 -> Water_Temp.:98.04, PWM:0
18:34:31.218 -> Water_Temp.:98.04, PWM:0
18:34:31.977 -> Water_Temp.:98.04, PWM:0
18:34:32.771 -> Water_Temp.:98.04, PWM:0
18:34:33.564 -> Water_Temp.:97.92, PWM:0
18:34:34.328 -> Water_Temp.:97.92, PWM:0
18:34:35.124 -> Water_Temp.:97.92, PWM:0
18:34:35.885 -> Water_Temp.:97.92, PWM:0
18:34:36.678 -> Water_Temp.:97.92, PWM:0
18:34:37.440 -> Water_Temp.:97.92, PWM:0
18:34:38.232 -> Water_Temp.:97.92, PWM:0
18:34:39.024 -> Water_Temp.:97.92, PWM:0
18:34:39.789 -> Water_Temp.:97.92, PWM:0
18:34:40.586 -> Water_Temp.:97.92, PWM:0
18:34:41.346 -> Water_Temp.:97.92, PWM:0
18:34:42.141 -> Water_Temp.:97.92, PWM:0
18:34:42.898 -> Water_Temp.:97.92, PWM:0
18:34:43.691 -> Water_Temp.:97.92, PWM:0
18:34:44.488 -> Water_Temp.:97.92, PWM:0
18:34:45.249 -> Water_Temp.:97.92, PWM:0
18:34:46.013 -> Water_Temp.:97.92, PWM:0
18:34:46.806 -> Water_Temp.:97.92, PWM:0
18:34:47.595 -> Water_Temp.:97.92, PWM:0
18:34:48.384 -> Water_Temp.:97.81, PWM:0
18:34:49.148 -> Water_Temp.:97.81, PWM:0
18:34:49.941 -> Water_Temp.:97.81, PWM:0
18:34:50.699 -> Water_Temp.:97.81, PWM:0
18:34:51.490 -> Water_Temp.:97.81, PWM:0
18:34:52.281 -> Water_Temp.:97.81, PWM:0
18:34:53.041 -> Water_Temp.:97.81, PWM:0
18:34:53.835 -> Water_Temp.:97.81, PWM:0
18:34:54.596 -> Water_Temp.:97.81, PWM:0
18:34:55.396 -> Water_Temp.:97.81, PWM:0
18:34:56.155 -> Water_Temp.:97.81, PWM:0
18:34:56.953 -> Water_Temp.:97.81, PWM:0
18:34:57.747 -> Water_Temp.:97.81, PWM:0
18:34:58.504 -> Water_Temp.:97.81, PWM:0
18:34:59.298 -> Water_Temp.:97.81, PWM:0
18:35:00.059 -> Water_Temp.:97.81, PWM:0
18:35:00.858 -> Water_Temp.:97.81, PWM:0
18:35:01.619 -> Water_Temp.:97.81, PWM:0
18:35:02.412 -> Water_Temp.:97.81, PWM:0
18:35:03.174 -> Water_Temp.:97.81, PWM:0
18:35:03.972 -> Water_Temp.:97.70, PWM:0
18:35:04.735 -> Water_Temp.:97.70, PWM:0
18:35:05.534 -> Water_Temp.:97.70, PWM:0
18:35:06.296 -> Water_Temp.:97.70, PWM:0
18:35:07.093 -> Water_Temp.:97.70, PWM:0
18:35:07.854 -> Water_Temp.:97.70, PWM:0
18:35:08.649 -> Water_Temp.:97.70, PWM:0
18:35:09.412 -> Water_Temp.:97.70, PWM:0
18:35:10.207 -> Water_Temp.:97.70, PWM:0
18:35:11.006 -> Water_Temp.:97.70, PWM:0
18:35:11.768 -> Water_Temp.:97.70, PWM:0
18:35:12.561 -> Water_Temp.:97.70, PWM:0
18:35:13.324 -> Water_Temp.:97.70, PWM:0
18:35:14.121 -> Water_Temp.:97.70, PWM:0
18:35:14.882 -> Water_Temp.:97.70, PWM:0
18:35:15.677 -> Water_Temp.:97.70, PWM:0
18:35:16.439 -> Water_Temp.:97.70, PWM:0
18:35:17.233 -> Water_Temp.:97.70, PWM:0
18:35:17.995 -> Water_Temp.:97.70, PWM:0
18:35:18.789 -> Water_Temp.:97.70, PWM:0
18:35:19.551 -> Water_Temp.:97.70, PWM:0
18:35:20.340 -> Water_Temp.:97.70, PWM:0
18:35:21.135 -> Water_Temp.:97.70, PWM:0
18:35:21.894 -> Water_Temp.:97.59, PWM:0
18:35:22.691 -> Water_Temp.:97.59, PWM:0
18:35:23.456 -> Water_Temp.:97.59, PWM:0
18:35:24.253 -> Water_Temp.:97.59, PWM:0
18:35:25.013 -> Water_Temp.:97.59, PWM:0
18:35:25.810 -> Water_Temp.:97.59, PWM:0
18:35:26.572 -> Water_Temp.:97.59, PWM:0
18:35:27.363 -> Water_Temp.:97.59, PWM:0
18:35:28.161 -> Water_Temp.:97.59, PWM:0
18:35:28.924 -> Water_Temp.:97.59, PWM:0
18:35:29.720 -> Water_Temp.:97.59, PWM:0
18:35:30.483 -> Water_Temp.:97.59, PWM:0
18:35:31.277 -> Water_Temp.:97.59, PWM:0
18:35:32.033 -> Water_Temp.:97.59, PWM:0
18:35:32.825 -> Water_Temp.:97.47, PWM:0
18:35:33.618 -> Water_Temp.:97.47, PWM:0
18:35:34.381 -> Water_Temp.:97.47, PWM:0
18:35:35.178 -> Water_Temp.:97.47, PWM:0
18:35:35.941 -> Water_Temp.:97.47, PWM:0
18:35:36.737 -> Water_Temp.:97.47, PWM:0
18:35:37.495 -> Water_Temp.:97.47, PWM:0
18:35:38.288 -> Water_Temp.:97.47, PWM:0
18:35:39.049 -> Water_Temp.:97.47, PWM:0
18:35:39.849 -> Water_Temp.:97.47, PWM:0
18:35:40.607 -> Water_Temp.:97.47, PWM:0
18:35:41.400 -> Water_Temp.:97.47, PWM:0
18:35:42.194 -> Water_Temp.:97.47, PWM:0
18:35:42.956 -> Water_Temp.:97.47, PWM:0
18:35:43.750 -> Water_Temp.:97.47, PWM:0
18:35:44.507 -> Water_Temp.:97.47, PWM:0
18:35:45.303 -> Water_Temp.:97.36, PWM:0
18:35:46.096 -> Water_Temp.:97.36, PWM:0
18:35:46.856 -> Water_Temp.:97.36, PWM:0
18:35:47.650 -> Water_Temp.:97.36, PWM:0
18:35:48.409 -> Water_Temp.:97.36, PWM:0
18:35:49.203 -> Water_Temp.:97.36, PWM:0
18:35:49.968 -> Water_Temp.:97.36, PWM:0
18:35:50.760 -> Water_Temp.:97.36, PWM:0
18:35:51.557 -> Water_Temp.:97.36, PWM:0
18:35:52.316 -> Water_Temp.:97.36, PWM:0
18:35:53.113 -> Water_Temp.:97.25, PWM:0
18:35:53.877 -> Water_Temp.:97.25, PWM:0
18:35:54.677 -> Water_Temp.:97.25, PWM:0
18:35:55.437 -> Water_Temp.:97.25, PWM:0
18:35:56.233 -> Water_Temp.:97.25, PWM:0
18:35:56.998 -> Water_Temp.:97.25, PWM:0
18:35:57.789 -> Water_Temp.:97.25, PWM:0
18:35:58.548 -> Water_Temp.:97.25, PWM:0
18:35:59.342 -> Water_Temp.:97.25, PWM:0
18:36:00.106 -> Water_Temp.:97.25, PWM:0
18:36:00.901 -> Water_Temp.:97.25, PWM:0
18:36:01.669 -> Water_Temp.:97.25, PWM:0
18:36:02.461 -> Water_Temp.:97.25, PWM:0
18:36:03.226 -> Water_Temp.:97.25, PWM:0
18:36:04.015 -> Water_Temp.:97.25, PWM:0
18:36:04.809 -> Water_Temp.:97.14, PWM:0
18:36:05.574 -> Water_Temp.:97.25, PWM:0
18:36:06.371 -> Water_Temp.:97.14, PWM:0
18:36:07.131 -> Water_Temp.:97.14, PWM:0
18:36:07.926 -> Water_Temp.:97.14, PWM:0
18:36:08.688 -> Water_Temp.:97.25, PWM:0
18:36:09.484 -> Water_Temp.:97.14, PWM:0
18:36:10.243 -> Water_Temp.:97.14, PWM:0
18:36:11.041 -> Water_Temp.:97.14, PWM:0
18:36:11.803 -> Water_Temp.:97.14, PWM:0
18:36:12.601 -> Water_Temp.:97.14, PWM:0
18:36:13.366 -> Water_Temp.:97.14, PWM:0
18:36:14.160 -> Water_Temp.:97.14, PWM:0
18:36:14.954 -> Water_Temp.:97.14, PWM:0
18:36:15.717 -> Water_Temp.:97.14, PWM:0
18:36:16.505 -> Water_Temp.:97.14, PWM:0
18:36:17.295 -> Water_Temp.:97.14, PWM:0
18:36:18.054 -> Water_Temp.:97.14, PWM:0
18:36:18.851 -> Water_Temp.:97.14, PWM:0
18:36:19.615 -> Water_Temp.:97.14, PWM:0
18:36:20.379 -> Water_Temp.:97.14, PWM:0
18:36:21.174 -> Water_Temp.:97.14, PWM:0
18:36:21.968 -> Water_Temp.:97.14, PWM:0
18:36:22.726 -> Water_Temp.:97.14, PWM:0
18:36:23.515 -> Water_Temp.:97.14, PWM:0
18:36:24.310 -> Water_Temp.:97.14, PWM:0
18:36:25.073 -> Water_Temp.:97.14, PWM:0
18:36:25.839 -> Water_Temp.:97.14, PWM:0
18:36:26.635 -> Water_Temp.:97.14, PWM:0
18:36:27.426 -> Water_Temp.:97.14, PWM:0
18:36:28.185 -> Water_Temp.:97.14, PWM:0
18:36:28.981 -> Water_Temp.:97.02, PWM:0
18:36:29.746 -> Water_Temp.:97.14, PWM:0
18:36:30.540 -> Water_Temp.:97.02, PWM:0
18:36:31.300 -> Water_Temp.:97.14, PWM:0
18:36:32.094 -> Water_Temp.:97.02, PWM:0
18:36:32.890 -> Water_Temp.:97.02, PWM:0
18:36:33.649 -> Water_Temp.:97.02, PWM:0
18:36:34.443 -> Water_Temp.:97.14, PWM:0
18:36:35.205 -> Water_Temp.:97.02, PWM:0
18:36:36.001 -> Water_Temp.:97.02, PWM:0
18:36:36.763 -> Water_Temp.:97.02, PWM:0
18:36:37.557 -> Water_Temp.:97.02, PWM:0
18:36:38.322 -> Water_Temp.:97.02, PWM:0
18:36:39.117 -> Water_Temp.:97.02, PWM:0
18:36:39.910 -> Water_Temp.:97.02, PWM:0
18:36:40.672 -> Water_Temp.:97.02, PWM:0
18:36:41.466 -> Water_Temp.:97.02, PWM:0
18:36:42.225 -> Water_Temp.:97.02, PWM:0
18:36:43.018 -> Water_Temp.:97.02, PWM:0
18:36:43.811 -> Water_Temp.:97.02, PWM:0
18:36:44.571 -> Water_Temp.:97.02, PWM:0
18:36:45.371 -> Water_Temp.:97.02, PWM:0
18:36:46.127 -> Water_Temp.:97.02, PWM:0
18:36:46.920 -> Water_Temp.:97.02, PWM:0
18:36:47.681 -> Water_Temp.:97.02, PWM:0
18:36:48.467 -> Water_Temp.:97.02, PWM:0
18:36:49.260 -> Water_Temp.:97.02, PWM:0
18:36:50.024 -> Water_Temp.:97.02, PWM:0
18:36:50.819 -> Water_Temp.:97.02, PWM:0
18:36:51.610 -> Water_Temp.:97.02, PWM:0
18:36:52.368 -> Water_Temp.:97.02, PWM:0
18:36:53.165 -> Water_Temp.:97.02, PWM:0
18:36:53.930 -> Water_Temp.:97.02, PWM:0
18:36:54.727 -> Water_Temp.:97.02, PWM:0
18:36:55.493 -> Water_Temp.:97.02, PWM:0
18:36:56.288 -> Water_Temp.:97.02, PWM:0
18:36:57.050 -> Water_Temp.:97.02, PWM:0
18:36:57.847 -> Water_Temp.:97.02, PWM:0
18:36:58.607 -> Water_Temp.:97.02, PWM:0
18:36:59.408 -> Water_Temp.:97.02, PWM:0
18:37:00.171 -> Water_Temp.:97.02, PWM:0
18:37:00.968 -> Water_Temp.:96.91, PWM:0
18:37:01.731 -> Water_Temp.:96.91, PWM:0
18:37:02.526 -> Water_Temp.:96.91, PWM:0
18:37:03.286 -> Water_Temp.:96.91, PWM:0
18:37:04.080 -> Water_Temp.:96.91, PWM:0
18:37:04.837 -> Water_Temp.:96.91, PWM:0
18:37:05.638 -> Water_Temp.:96.91, PWM:0
18:37:06.397 -> Water_Temp.:96.91, PWM:0
18:37:07.193 -> Water_Temp.:96.91, PWM:0
18:37:07.984 -> Water_Temp.:96.91, PWM:0
18:37:08.744 -> Water_Temp.:96.91, PWM:0
18:37:09.540 -> Water_Temp.:96.91, PWM:0
18:37:10.294 -> Water_Temp.:96.91, PWM:0
18:37:11.088 -> Water_Temp.:96.91, PWM:0
18:37:11.883 -> Water_Temp.:96.91, PWM:0
18:37:12.637 -> Water_Temp.:96.91, PWM:0
18:37:13.430 -> Water_Temp.:96.91, PWM:0
18:37:14.223 -> Water_Temp.:96.91, PWM:0
18:37:14.984 -> Water_Temp.:96.91, PWM:0
18:37:15.780 -> Water_Temp.:96.91, PWM:0
18:37:16.539 -> Water_Temp.:96.91, PWM:0
18:37:17.332 -> Water_Temp.:96.91, PWM:0
18:37:18.127 -> Water_Temp.:96.91, PWM:0
18:37:18.889 -> Water_Temp.:96.91, PWM:0
18:37:19.685 -> Water_Temp.:96.91, PWM:0
18:37:20.446 -> Water_Temp.:96.91, PWM:0
18:37:21.241 -> Water_Temp.:96.91, PWM:0
18:37:22.002 -> Water_Temp.:96.80, PWM:0
18:37:22.795 -> Water_Temp.:96.91, PWM:0
18:37:23.553 -> Water_Temp.:96.80, PWM:0
18:37:24.348 -> Water_Temp.:96.80, PWM:0
18:37:25.142 -> Water_Temp.:96.80, PWM:0
18:37:25.905 -> Water_Temp.:96.80, PWM:0
18:37:26.702 -> Water_Temp.:96.80, PWM:0
18:37:27.463 -> Water_Temp.:96.80, PWM:0
18:37:28.256 -> Water_Temp.:96.80, PWM:0
18:37:29.013 -> Water_Temp.:96.80, PWM:0
18:37:29.807 -> Water_Temp.:96.80, PWM:0
18:37:30.597 -> Water_Temp.:96.80, PWM:0
18:37:31.355 -> Water_Temp.:96.80, PWM:0
18:37:32.152 -> Water_Temp.:96.80, PWM:0
18:37:32.944 -> Water_Temp.:96.80, PWM:0
18:37:33.702 -> Water_Temp.:96.80, PWM:0
18:37:34.494 -> Water_Temp.:96.80, PWM:0
18:37:35.255 -> Water_Temp.:96.80, PWM:0
18:37:36.051 -> Water_Temp.:96.80, PWM:0
18:37:36.813 -> Water_Temp.:96.80, PWM:0
18:37:37.610 -> Water_Temp.:96.80, PWM:0
18:37:38.403 -> Water_Temp.:96.80, PWM:0
18:37:39.163 -> Water_Temp.:96.80, PWM:0
18:37:39.962 -> Water_Temp.:96.80, PWM:0
18:37:40.722 -> Water_Temp.:96.80, PWM:0
18:37:41.521 -> Water_Temp.:96.80, PWM:0
18:37:42.282 -> Water_Temp.:96.80, PWM:0
18:37:43.074 -> Water_Temp.:96.80, PWM:0
18:37:43.836 -> Water_Temp.:96.80, PWM:0
18:37:44.631 -> Water_Temp.:96.80, PWM:0
18:37:45.423 -> Water_Temp.:96.80, PWM:0
18:37:46.180 -> Water_Temp.:96.80, PWM:0
18:37:46.972 -> Water_Temp.:96.69, PWM:0
18:37:47.731 -> Water_Temp.:96.69, PWM:0
18:37:48.522 -> Water_Temp.:96.80, PWM:0
18:37:49.317 -> Water_Temp.:96.69, PWM:0
18:37:50.070 -> Water_Temp.:96.80, PWM:0
18:37:50.868 -> Water_Temp.:96.69, PWM:0
18:37:51.661 -> Water_Temp.:96.69, PWM:0
18:37:52.419 -> Water_Temp.:96.69, PWM:0
18:37:53.215 -> Water_Temp.:96.69, PWM:0
18:37:53.973 -> Water_Temp.:96.69, PWM:0
18:37:54.769 -> Water_Temp.:96.69, PWM:0
18:37:55.561 -> Water_Temp.:96.69, PWM:0
18:37:56.319 -> Water_Temp.:96.69, PWM:0
18:37:57.112 -> Water_Temp.:96.69, PWM:0
18:37:57.871 -> Water_Temp.:96.69, PWM:0
18:37:58.669 -> Water_Temp.:96.69, PWM:6
18:37:59.433 -> Water_Temp.:96.69, PWM:6
18:38:00.227 -> Water_Temp.:96.69, PWM:7
18:38:00.988 -> Water_Temp.:96.69, PWM:28
18:38:01.787 -> Water_Temp.:96.69, PWM:58
18:38:02.583 -> Water_Temp.:96.69, PWM:103
18:38:04.099 -> Water_Temp.:96.69, PWM:109
18:38:04.894 -> Water_Temp.:96.69, PWM:109
18:38:05.656 -> Water_Temp.:96.69, PWM:109
18:38:06.447 -> Water_Temp.:96.69, PWM:109
18:38:07.237 -> Water_Temp.:96.69, PWM:109
18:38:07.998 -> Water_Temp.:96.69, PWM:109
18:38:08.788 -> Water_Temp.:96.69, PWM:109
18:38:09.576 -> Water_Temp.:96.69, PWM:109
18:38:10.338 -> Water_Temp.:96.69, PWM:109
18:38:11.132 -> Water_Temp.:96.69, PWM:109
18:38:11.890 -> Water_Temp.:96.69, PWM:109
18:38:12.683 -> Water_Temp.:96.69, PWM:109
18:38:13.479 -> Water_Temp.:96.69, PWM:109
18:38:14.238 -> Water_Temp.:96.57, PWM:109
18:38:15.030 -> Water_Temp.:96.57, PWM:109
18:38:15.788 -> Water_Temp.:96.57, PWM:109
18:38:16.579 -> Water_Temp.:96.69, PWM:109
18:38:17.375 -> Water_Temp.:96.57, PWM:109
18:38:18.133 -> Water_Temp.:96.57, PWM:109
18:38:18.925 -> Water_Temp.:96.57, PWM:109
18:38:19.720 -> Water_Temp.:96.57, PWM:109
18:38:20.480 -> Water_Temp.:96.57, PWM:109
18:38:22.030 -> Water_Temp.:96.57, PWM:109
18:38:22.790 -> Water_Temp.:96.57, PWM:163
18:38:23.583 -> Water_Temp.:96.57, PWM:254
18:38:24.376 -> Water_Temp.:96.57, PWM:254
18:38:25.134 -> Water_Temp.:98.15, PWM:254
18:38:25.931 -> Water_Temp.:100.17, PWM:254
18:38:26.724 -> Water_Temp.:97.81, PWM:254
18:38:27.484 -> Water_Temp.:97.47, PWM:254
18:38:28.278 -> Water_Temp.:97.25, PWM:254
18:38:29.037 -> Water_Temp.:97.14, PWM:254
18:38:29.831 -> Water_Temp.:97.02, PWM:254
18:38:30.624 -> Water_Temp.:96.91, PWM:254
18:38:31.385 -> Water_Temp.:96.91, PWM:254
18:38:32.177 -> Water_Temp.:96.80, PWM:254
18:38:32.936 -> Water_Temp.:96.80, PWM:254
18:38:33.729 -> Water_Temp.:96.80, PWM:254
18:38:34.521 -> Water_Temp.:96.80, PWM:254
18:38:35.291 -> Water_Temp.:96.80, PWM:254
18:38:36.052 -> Water_Temp.:96.80, PWM:254
18:38:36.838 -> Water_Temp.:96.80, PWM:254
18:38:37.638 -> Water_Temp.:96.80, PWM:254
18:38:38.395 -> Water_Temp.:96.91, PWM:254
18:38:39.192 -> Water_Temp.:96.91, PWM:254
18:38:39.952 -> Water_Temp.:96.91, PWM:254
18:38:40.744 -> Water_Temp.:96.91, PWM:254
18:38:41.533 -> Water_Temp.:97.02, PWM:254
18:38:42.291 -> Water_Temp.:97.02, PWM:254
18:38:43.080 -> Water_Temp.:97.02, PWM:254
18:38:43.878 -> Water_Temp.:97.02, PWM:254
18:38:44.640 -> Water_Temp.:97.02, PWM:254
18:38:45.434 -> Water_Temp.:97.02, PWM:254
18:38:46.226 -> Water_Temp.:97.14, PWM:254
18:38:46.985 -> Water_Temp.:97.14, PWM:254
18:38:47.776 -> Water_Temp.:97.14, PWM:254
18:38:48.535 -> Water_Temp.:97.14, PWM:254
18:38:49.324 -> Water_Temp.:97.14, PWM:254
18:38:50.119 -> Water_Temp.:97.14, PWM:254
18:38:51.638 -> Water_Temp.:97.25, PWM:254
18:38:52.430 -> Water_Temp.:97.25, PWM:254
18:38:53.226 -> Water_Temp.:97.25, PWM:254
18:38:53.985 -> Water_Temp.:97.36, PWM:254
18:38:54.783 -> Water_Temp.:97.36, PWM:254
18:38:55.545 -> Water_Temp.:97.47, PWM:254
18:38:56.338 -> Water_Temp.:97.47, PWM:254
18:38:57.095 -> Water_Temp.:97.47, PWM:254
18:38:57.889 -> Water_Temp.:97.59, PWM:254
18:38:58.686 -> Water_Temp.:97.59, PWM:254
18:38:59.447 -> Water_Temp.:97.59, PWM:254
18:39:00.236 -> Water_Temp.:97.59, PWM:254
18:39:00.997 -> Water_Temp.:97.70, PWM:254
18:39:01.789 -> Water_Temp.:97.70, PWM:254
18:39:02.585 -> Water_Temp.:97.70, PWM:254
18:39:03.345 -> Water_Temp.:97.70, PWM:254
18:39:04.136 -> Water_Temp.:97.70, PWM:254
18:39:04.895 -> Water_Temp.:97.81, PWM:254
18:39:05.694 -> Water_Temp.:97.92, PWM:254
18:39:06.452 -> Water_Temp.:97.92, PWM:254
18:39:07.246 -> Water_Temp.:98.04, PWM:254
18:39:08.038 -> Water_Temp.:98.15, PWM:254
18:39:08.799 -> Water_Temp.:98.15, PWM:254
18:39:09.591 -> Water_Temp.:98.26, PWM:254
18:39:10.350 -> Water_Temp.:98.26, PWM:254
18:39:11.146 -> Water_Temp.:98.37, PWM:254
18:39:11.940 -> Water_Temp.:98.49, PWM:254
18:39:12.699 -> Water_Temp.:98.60, PWM:254
18:39:13.494 -> Water_Temp.:98.71, PWM:254
18:39:15.016 -> Water_Temp.:98.82, PWM:254
18:39:15.817 -> Water_Temp.:98.82, PWM:254
18:39:16.577 -> Water_Temp.:98.82, PWM:254
18:39:17.372 -> Water_Temp.:98.82, PWM:254
18:39:18.131 -> Water_Temp.:98.94, PWM:254
18:39:18.926 -> Water_Temp.:98.94, PWM:254
18:39:19.722 -> Water_Temp.:98.94, PWM:254
18:39:20.481 -> Water_Temp.:98.94, PWM:254
18:39:21.272 -> Water_Temp.:98.94, PWM:254
18:39:22.032 -> Water_Temp.:99.05, PWM:254
18:39:22.824 -> Water_Temp.:99.05, PWM:254
18:39:23.616 -> Water_Temp.:99.05, PWM:254
18:39:24.376 -> Water_Temp.:99.05, PWM:254
18:39:25.174 -> Water_Temp.:99.16, PWM:254
18:39:25.935 -> Water_Temp.:99.16, PWM:254
18:39:26.735 -> Water_Temp.:99.16, PWM:254
18:39:27.490 -> Water_Temp.:99.27, PWM:254
18:39:28.284 -> Water_Temp.:99.27, PWM:254
18:39:29.077 -> Water_Temp.:99.27, PWM:254
18:39:29.835 -> Water_Temp.:99.39, PWM:254
18:39:30.632 -> Water_Temp.:99.39, PWM:254
18:39:31.393 -> Water_Temp.:99.39, PWM:254
18:39:32.183 -> Water_Temp.:99.50, PWM:254
18:39:32.974 -> Water_Temp.:99.61, PWM:254
18:39:33.734 -> Water_Temp.:99.72, PWM:254
18:39:35.286 -> Water_Temp.:99.84, PWM:254
18:39:36.078 -> Water_Temp.:99.95, PWM:254
18:39:36.867 -> Water_Temp.:99.95, PWM:254
18:39:37.628 -> Water_Temp.:100.06, PWM:254
18:39:38.421 -> Water_Temp.:100.17, PWM:254
18:39:39.215 -> Water_Temp.:100.17, PWM:254
18:39:39.979 -> Water_Temp.:100.29, PWM:254
18:39:41.561 -> Water_Temp.:100.40, PWM:254
18:39:42.319 -> Water_Temp.:100.51, PWM:254
18:39:43.111 -> Water_Temp.:100.51, PWM:254
18:39:43.875 -> Water_Temp.:100.62, PWM:254
18:39:44.669 -> Water_Temp.:100.62, PWM:254
18:39:45.427 -> Water_Temp.:100.62, PWM:254
18:39:46.216 -> Water_Temp.:100.62, PWM:254
18:39:47.771 -> Water_Temp.:100.74, PWM:254
18:39:48.566 -> Water_Temp.:100.74, PWM:254
18:39:49.329 -> Water_Temp.:100.85, PWM:254
18:39:50.118 -> Water_Temp.:100.85, PWM:254
18:39:50.912 -> Water_Temp.:100.96, PWM:254
18:39:51.672 -> Water_Temp.:100.96, PWM:254
18:39:52.469 -> Water_Temp.:101.07, PWM:254
18:39:53.225 -> Water_Temp.:101.07, PWM:254
18:39:54.024 -> Water_Temp.:101.07, PWM:254
18:39:54.817 -> Water_Temp.:101.19, PWM:254
18:39:55.579 -> Water_Temp.:101.19, PWM:194
18:39:56.367 -> Water_Temp.:101.30, PWM:194
18:39:57.124 -> Water_Temp.:101.30, PWM:194
18:39:58.707 -> Water_Temp.:101.41, PWM:194
18:39:59.466 -> Water_Temp.:101.41, PWM:150
18:40:00.263 -> Water_Temp.:101.52, PWM:119
18:40:01.025 -> Water_Temp.:101.64, PWM:109
18:40:01.821 -> Water_Temp.:101.64, PWM:109
18:40:02.616 -> Water_Temp.:101.75, PWM:97
18:40:03.372 -> Water_Temp.:101.75, PWM:96
18:40:04.161 -> Water_Temp.:101.75, PWM:96
18:40:04.953 -> Water_Temp.:101.86, PWM:96
18:40:05.716 -> Water_Temp.:101.86, PWM:96
18:40:06.508 -> Water_Temp.:101.86, PWM:96
18:40:07.299 -> Water_Temp.:101.86, PWM:96
18:40:08.054 -> Water_Temp.:101.97, PWM:96
18:40:08.853 -> Water_Temp.:101.97, PWM:96
18:40:09.613 -> Water_Temp.:101.97, PWM:96
18:40:10.403 -> Water_Temp.:101.97, PWM:129
18:40:11.169 -> Water_Temp.:101.97, PWM:129
18:40:12.719 -> Water_Temp.:102.20, PWM:129
18:40:14.270 -> Water_Temp.:102.31, PWM:129
18:40:15.061 -> Water_Temp.:102.31, PWM:129
18:40:15.854 -> Water_Temp.:102.31, PWM:129
18:40:16.611 -> Water_Temp.:102.42, PWM:129
18:40:17.403 -> Water_Temp.:102.42, PWM:109
18:40:18.166 -> Water_Temp.:102.42, PWM:101
18:40:19.723 -> Water_Temp.:102.42, PWM:101
18:40:20.480 -> Water_Temp.:102.42, PWM:94
18:40:21.269 -> Water_Temp.:102.54, PWM:94
18:40:22.817 -> Water_Temp.:102.54, PWM:94
18:40:23.609 -> Water_Temp.:102.65, PWM:84
18:40:24.403 -> Water_Temp.:102.65, PWM:74
18:40:25.166 -> Water_Temp.:102.65, PWM:67
18:40:25.964 -> Water_Temp.:102.76, PWM:54
18:40:26.724 -> Water_Temp.:102.76, PWM:54
18:40:27.517 -> Water_Temp.:102.76, PWM:54
18:40:28.308 -> Water_Temp.:102.76, PWM:54
18:40:29.066 -> Water_Temp.:102.76, PWM:52
18:40:29.861 -> Water_Temp.:102.76, PWM:51
18:40:30.621 -> Water_Temp.:102.76, PWM:51
18:40:31.409 -> Water_Temp.:102.76, PWM:51
18:40:32.202 -> Water_Temp.:102.76, PWM:51
18:40:32.963 -> Water_Temp.:102.76, PWM:51
18:40:33.755 -> Water_Temp.:102.87, PWM:51
18:40:34.518 -> Water_Temp.:102.87, PWM:51
18:40:35.312 -> Water_Temp.:102.87, PWM:51
18:40:36.075 -> Water_Temp.:102.99, PWM:51
18:40:36.870 -> Water_Temp.:102.99, PWM:51
18:40:37.666 -> Water_Temp.:102.99, PWM:51
18:40:39.193 -> Water_Temp.:103.10, PWM:51
18:40:39.992 -> Water_Temp.:103.10, PWM:51
18:40:40.756 -> Water_Temp.:103.10, PWM:51
18:40:41.550 -> Water_Temp.:103.21, PWM:51
18:40:42.309 -> Water_Temp.:103.21, PWM:51
18:40:43.101 -> Water_Temp.:103.21, PWM:51
18:40:43.860 -> Water_Temp.:103.32, PWM:51
18:40:44.654 -> Water_Temp.:103.32, PWM:51
18:40:45.448 -> Water_Temp.:103.44, PWM:51
18:40:46.203 -> Water_Temp.:103.44, PWM:51
18:40:46.997 -> Water_Temp.:103.44, PWM:51
18:40:47.759 -> Water_Temp.:103.55, PWM:51
18:40:48.555 -> Water_Temp.:103.55, PWM:0
18:40:49.349 -> Water_Temp.:103.55, PWM:0
18:40:50.104 -> Water_Temp.:103.55, PWM:0
18:40:50.900 -> Water_Temp.:103.55, PWM:0
18:40:51.658 -> Water_Temp.:103.55, PWM:0
18:40:52.453 -> Water_Temp.:103.55, PWM:0
18:40:53.248 -> Water_Temp.:103.55, PWM:0
18:40:54.010 -> Water_Temp.:103.55, PWM:0
18:40:54.801 -> Water_Temp.:103.66, PWM:0
18:40:55.564 -> Water_Temp.:103.66, PWM:0
18:40:56.360 -> Water_Temp.:103.66, PWM:0
18:40:57.116 -> Water_Temp.:103.66, PWM:0
18:40:57.911 -> Water_Temp.:103.66, PWM:0
18:40:58.709 -> Water_Temp.:103.66, PWM:0
18:40:59.470 -> Water_Temp.:103.66, PWM:0
18:41:00.266 -> Water_Temp.:103.77, PWM:0
18:41:01.023 -> Water_Temp.:103.77, PWM:0
18:41:01.819 -> Water_Temp.:103.77, PWM:0
18:41:02.579 -> Water_Temp.:103.77, PWM:0
18:41:03.372 -> Water_Temp.:103.77, PWM:0
18:41:04.165 -> Water_Temp.:103.89, PWM:0
18:41:04.922 -> Water_Temp.:103.89, PWM:0
18:41:05.716 -> Water_Temp.:103.89, PWM:0
18:41:06.481 -> Water_Temp.:103.89, PWM:0
18:41:07.273 -> Water_Temp.:103.89, PWM:0
18:41:08.068 -> Water_Temp.:103.89, PWM:0
18:41:08.826 -> Water_Temp.:103.89, PWM:0
18:41:09.619 -> Water_Temp.:103.89, PWM:0
18:41:10.379 -> Water_Temp.:103.89, PWM:0
18:41:11.173 -> Water_Temp.:103.89, PWM:0
18:41:11.967 -> Water_Temp.:103.89, PWM:0
18:41:12.726 -> Water_Temp.:103.89, PWM:0
18:41:13.524 -> Water_Temp.:103.89, PWM:0
18:41:14.283 -> Water_Temp.:103.89, PWM:0
18:41:15.074 -> Water_Temp.:103.89, PWM:0
18:41:15.869 -> Water_Temp.:103.89, PWM:0
18:41:16.631 -> Water_Temp.:103.89, PWM:0
18:41:17.427 -> Water_Temp.:103.89, PWM:0
18:41:18.185 -> Water_Temp.:103.89, PWM:0
18:41:18.982 -> Water_Temp.:103.89, PWM:0
18:41:19.742 -> Water_Temp.:103.89, PWM:0
18:41:20.534 -> Water_Temp.:103.89, PWM:0
18:41:21.327 -> Water_Temp.:104.00, PWM:0
18:41:22.085 -> Water_Temp.:104.00, PWM:0
18:41:22.878 -> Water_Temp.:104.00, PWM:0
18:41:23.639 -> Water_Temp.:104.00, PWM:0
18:41:24.432 -> Water_Temp.:104.00, PWM:0
18:41:25.229 -> Water_Temp.:104.00, PWM:0
18:41:25.983 -> Water_Temp.:104.00, PWM:0
18:41:26.777 -> Water_Temp.:104.00, PWM:0
18:41:27.535 -> Water_Temp.:104.00, PWM:0
18:41:28.326 -> Water_Temp.:104.00, PWM:0
18:41:29.120 -> Water_Temp.:104.00, PWM:0
18:41:29.879 -> Water_Temp.:104.00, PWM:0
18:41:30.666 -> Water_Temp.:104.00, PWM:0
18:41:31.459 -> Water_Temp.:104.00, PWM:0
18:41:32.217 -> Water_Temp.:104.00, PWM:0
18:41:33.011 -> Water_Temp.:104.00, PWM:0
18:41:33.772 -> Water_Temp.:104.00, PWM:0
18:41:34.569 -> Water_Temp.:104.00, PWM:0
18:41:35.366 -> Water_Temp.:104.00, PWM:0
18:41:36.126 -> Water_Temp.:104.00, PWM:0
18:41:36.921 -> Water_Temp.:104.00, PWM:0
18:41:37.683 -> Water_Temp.:104.00, PWM:0
18:41:38.479 -> Water_Temp.:104.00, PWM:0
18:41:39.242 -> Water_Temp.:104.00, PWM:0
18:41:40.034 -> Water_Temp.:104.00, PWM:0
18:41:40.798 -> Water_Temp.:104.00, PWM:0
18:41:41.591 -> Water_Temp.:104.00, PWM:0
18:41:42.353 -> Water_Temp.:104.00, PWM:0
18:41:43.144 -> Water_Temp.:104.00, PWM:0
18:41:43.937 -> Water_Temp.:104.00, PWM:0
18:41:44.694 -> Water_Temp.:104.00, PWM:0
18:41:45.485 -> Water_Temp.:104.00, PWM:0
18:41:46.277 -> Water_Temp.:104.00, PWM:0
18:41:47.035 -> Water_Temp.:104.00, PWM:0
18:41:47.824 -> Water_Temp.:104.00, PWM:0
18:41:48.619 -> Water_Temp.:104.00, PWM:0
18:41:49.375 -> Water_Temp.:104.00, PWM:0
18:41:50.172 -> Water_Temp.:104.00, PWM:0
18:41:50.967 -> Water_Temp.:104.00, PWM:0
18:41:51.728 -> Water_Temp.:104.00, PWM:0
18:41:52.523 -> Water_Temp.:104.00, PWM:0
18:41:53.285 -> Water_Temp.:104.00, PWM:0
18:41:54.078 -> Water_Temp.:104.00, PWM:0
18:41:54.836 -> Water_Temp.:104.00, PWM:0
18:41:55.634 -> Water_Temp.:104.00, PWM:0
18:41:56.426 -> Water_Temp.:104.00, PWM:0
18:41:57.184 -> Water_Temp.:104.00, PWM:0
18:41:57.974 -> Water_Temp.:104.00, PWM:0
18:41:58.732 -> Water_Temp.:104.00, PWM:0
18:41:59.530 -> Water_Temp.:104.00, PWM:0
18:42:00.321 -> Water_Temp.:104.00, PWM:0
18:42:01.072 -> Water_Temp.:104.00, PWM:0
18:42:01.863 -> Water_Temp.:104.00, PWM:0
18:42:02.655 -> Water_Temp.:104.00, PWM:0
18:42:03.443 -> Water_Temp.:104.00, PWM:0
18:42:04.206 -> Water_Temp.:104.00, PWM:0
18:42:05.000 -> Water_Temp.:104.00, PWM:0
18:42:05.763 -> Water_Temp.:104.00, PWM:3
18:42:06.553 -> Water_Temp.:104.00, PWM:5
18:42:08.102 -> Water_Temp.:104.00, PWM:5
18:42:08.865 -> Water_Temp.:103.89, PWM:17
18:42:09.654 -> Water_Temp.:104.00, PWM:19
18:42:10.413 -> Water_Temp.:104.00, PWM:21
18:42:11.206 -> Water_Temp.:104.00, PWM:25
18:42:12.003 -> Water_Temp.:103.89, PWM:29
18:42:12.760 -> Water_Temp.:104.00, PWM:32
18:42:13.557 -> Water_Temp.:103.89, PWM:35
18:42:14.318 -> Water_Temp.:104.00, PWM:36
18:42:15.112 -> Water_Temp.:103.89, PWM:39
18:42:15.904 -> Water_Temp.:103.89, PWM:41
18:42:17.451 -> Water_Temp.:103.89, PWM:45
18:42:18.211 -> Water_Temp.:104.00, PWM:49
18:42:19.007 -> Water_Temp.:103.89, PWM:50
18:42:19.774 -> Water_Temp.:103.89, PWM:50
18:42:20.569 -> Water_Temp.:103.89, PWM:51
18:42:22.127 -> Water_Temp.:103.89, PWM:51
18:42:22.916 -> Water_Temp.:103.89, PWM:51
18:42:23.674 -> Water_Temp.:103.89, PWM:51
18:42:24.467 -> Water_Temp.:103.89, PWM:51
18:42:25.257 -> Water_Temp.:103.89, PWM:53
18:42:26.021 -> Water_Temp.:103.89, PWM:53
18:42:26.817 -> Water_Temp.:103.89, PWM:55
18:42:27.576 -> Water_Temp.:103.89, PWM:56
18:42:28.374 -> Water_Temp.:103.89, PWM:56
18:42:29.134 -> Water_Temp.:103.89, PWM:56
18:42:29.930 -> Water_Temp.:103.89, PWM:55
18:42:30.724 -> Water_Temp.:103.89, PWM:55
18:42:31.482 -> Water_Temp.:103.89, PWM:56
18:42:32.273 -> Water_Temp.:103.89, PWM:56
18:42:33.817 -> Water_Temp.:104.00, PWM:55
18:42:34.610 -> Water_Temp.:104.00, PWM:56
18:42:36.163 -> Water_Temp.:104.00, PWM:56
18:42:36.922 -> Water_Temp.:104.00, PWM:56
18:42:37.718 -> Water_Temp.:104.00, PWM:56
18:42:38.476 -> Water_Temp.:104.00, PWM:56
18:42:39.271 -> Water_Temp.:104.00, PWM:56
18:42:40.032 -> Water_Temp.:104.00, PWM:55
18:42:40.829 -> Water_Temp.:104.00, PWM:56
18:42:41.622 -> Water_Temp.:104.00, PWM:56
18:42:42.383 -> Water_Temp.:104.00, PWM:56
18:42:43.176 -> Water_Temp.:104.00, PWM:56
18:42:43.940 -> Water_Temp.:104.00, PWM:56
18:42:44.731 -> Water_Temp.:104.00, PWM:55
18:42:45.490 -> Water_Temp.:104.11, PWM:56
18:42:46.282 -> Water_Temp.:104.11, PWM:55
18:42:47.077 -> Water_Temp.:104.11, PWM:56
18:42:47.832 -> Water_Temp.:104.11, PWM:56
18:42:48.625 -> Water_Temp.:104.11, PWM:55
18:42:49.419 -> Water_Temp.:104.11, PWM:56
18:42:50.177 -> Water_Temp.:104.11, PWM:56
18:42:50.970 -> Water_Temp.:104.11, PWM:56
18:42:51.763 -> Water_Temp.:104.11, PWM:56
18:42:52.523 -> Water_Temp.:104.11, PWM:55
18:42:53.316 -> Water_Temp.:104.11, PWM:56
18:42:54.077 -> Water_Temp.:104.11, PWM:55
18:42:54.870 -> Water_Temp.:104.22, PWM:56
18:42:56.420 -> Water_Temp.:104.22, PWM:55
18:42:57.217 -> Water_Temp.:104.22, PWM:56
18:42:57.972 -> Water_Temp.:104.22, PWM:56
18:42:58.770 -> Water_Temp.:104.22, PWM:56
18:42:59.530 -> Water_Temp.:104.22, PWM:56
18:43:01.117 -> Water_Temp.:104.11, PWM:56
18:43:01.879 -> Water_Temp.:104.11, PWM:55
18:43:02.678 -> Water_Temp.:104.11, PWM:56
18:43:03.437 -> Water_Temp.:104.11, PWM:56
18:43:04.230 -> Water_Temp.:104.11, PWM:56
18:43:04.989 -> Water_Temp.:105.69, PWM:56
18:43:05.787 -> Water_Temp.:107.82, PWM:55
18:43:06.549 -> Water_Temp.:105.46, PWM:56
18:43:07.342 -> Water_Temp.:105.01, PWM:55
18:43:08.140 -> Water_Temp.:104.79, PWM:55
18:43:08.901 -> Water_Temp.:104.67, PWM:55
18:43:09.693 -> Water_Temp.:104.56, PWM:56
18:43:10.457 -> Water_Temp.:104.56, PWM:56
18:43:11.253 -> Water_Temp.:104.45, PWM:55
18:43:12.011 -> Water_Temp.:104.45, PWM:56
18:43:12.805 -> Water_Temp.:104.34, PWM:55
18:43:13.598 -> Water_Temp.:104.34, PWM:55
18:43:14.360 -> Water_Temp.:104.34, PWM:56
18:43:15.155 -> Water_Temp.:104.34, PWM:56
18:43:15.911 -> Water_Temp.:104.34, PWM:55
18:43:16.706 -> Water_Temp.:104.34, PWM:56
18:43:17.499 -> Water_Temp.:104.34, PWM:55
18:43:18.258 -> Water_Temp.:104.34, PWM:56
18:43:19.053 -> Water_Temp.:104.34, PWM:56
18:43:19.811 -> Water_Temp.:104.34, PWM:56
18:43:20.601 -> Water_Temp.:104.34, PWM:56
18:43:21.393 -> Water_Temp.:104.34, PWM:55
18:43:22.152 -> Water_Temp.:104.34, PWM:56
18:43:22.942 -> Water_Temp.:104.45, PWM:55
18:43:23.734 -> Water_Temp.:104.45, PWM:56
18:43:24.493 -> Water_Temp.:104.45, PWM:56
18:43:25.290 -> Water_Temp.:104.45, PWM:55
18:43:26.055 -> Water_Temp.:104.45, PWM:55
18:43:27.603 -> Water_Temp.:104.45, PWM:55
18:43:28.396 -> Water_Temp.:104.45, PWM:56
18:43:29.152 -> Water_Temp.:104.45, PWM:56
18:43:29.945 -> Water_Temp.:104.45, PWM:56
18:43:30.741 -> Water_Temp.:104.45, PWM:56
18:43:31.500 -> Water_Temp.:104.45, PWM:56
18:43:32.291 -> Water_Temp.:104.45, PWM:55
18:43:33.051 -> Water_Temp.:104.45, PWM:56
18:43:33.846 -> Water_Temp.:104.45, PWM:56
18:43:34.643 -> Water_Temp.:104.45, PWM:56
18:43:35.403 -> Water_Temp.:104.45, PWM:55
18:43:36.199 -> Water_Temp.:104.45, PWM:56
18:43:36.954 -> Water_Temp.:104.45, PWM:56
18:43:37.745 -> Water_Temp.:104.45, PWM:55
18:43:38.540 -> Water_Temp.:104.45, PWM:56
18:43:39.296 -> Water_Temp.:104.45, PWM:55
18:43:40.088 -> Water_Temp.:104.45, PWM:56
18:43:40.879 -> Water_Temp.:104.45, PWM:55
18:43:41.637 -> Water_Temp.:104.45, PWM:55
18:43:42.430 -> Water_Temp.:104.45, PWM:56
18:43:43.219 -> Water_Temp.:104.45, PWM:56
18:43:44.773 -> Water_Temp.:104.56, PWM:56
18:43:45.532 -> Water_Temp.:104.56, PWM:55
18:43:46.325 -> Water_Temp.:104.56, PWM:55
18:43:47.119 -> Water_Temp.:104.56, PWM:56
18:43:47.876 -> Water_Temp.:104.56, PWM:55
18:43:48.671 -> Water_Temp.:104.67, PWM:56
18:43:49.432 -> Water_Temp.:104.67, PWM:55
18:43:50.228 -> Water_Temp.:104.67, PWM:56
18:43:51.021 -> Water_Temp.:104.67, PWM:55
18:43:52.537 -> Water_Temp.:104.67, PWM:55
18:43:53.334 -> Water_Temp.:104.67, PWM:55
18:43:54.091 -> Water_Temp.:104.67, PWM:55
18:43:54.882 -> Water_Temp.:104.67, PWM:56
18:43:55.677 -> Water_Temp.:104.67, PWM:55
18:43:56.439 -> Water_Temp.:104.67, PWM:55
18:43:57.231 -> Water_Temp.:104.67, PWM:56
18:43:57.993 -> Water_Temp.:104.67, PWM:56
18:43:58.789 -> Water_Temp.:104.67, PWM:56
18:44:00.334 -> Water_Temp.:104.79, PWM:55
18:44:01.132 -> Water_Temp.:104.67, PWM:56
18:44:01.889 -> Water_Temp.:104.79, PWM:56
18:44:02.685 -> Water_Temp.:104.67, PWM:55
18:44:03.477 -> Water_Temp.:104.67, PWM:55
18:44:05.000 -> Water_Temp.:104.79, PWM:55
18:44:05.797 -> Water_Temp.:104.67, PWM:55
18:44:06.555 -> Water_Temp.:104.67, PWM:56
18:44:07.351 -> Water_Temp.:104.67, PWM:55
18:44:08.146 -> Water_Temp.:104.67, PWM:55
18:44:08.909 -> Water_Temp.:104.67, PWM:56
18:44:09.670 -> Water_Temp.:104.67, PWM:56
18:44:10.463 -> Water_Temp.:104.67, PWM:55
18:44:11.251 -> Water_Temp.:104.67, PWM:56
18:44:12.042 -> Water_Temp.:104.67, PWM:55
18:44:13.561 -> Water_Temp.:104.67, PWM:55
18:44:14.354 -> Water_Temp.:104.67, PWM:55
18:44:15.111 -> Water_Temp.:104.67, PWM:55
18:44:15.904 -> Water_Temp.:104.67, PWM:55
18:44:16.699 -> Water_Temp.:104.67, PWM:56
18:44:17.456 -> Water_Temp.:104.79, PWM:55
18:44:18.251 -> Water_Temp.:104.79, PWM:55
18:44:19.010 -> Water_Temp.:104.79, PWM:56
18:44:19.806 -> Water_Temp.:104.79, PWM:56
18:44:20.601 -> Water_Temp.:104.79, PWM:55
18:44:21.358 -> Water_Temp.:104.79, PWM:56
18:44:22.151 -> Water_Temp.:104.79, PWM:56
18:44:23.698 -> Water_Temp.:104.79, PWM:55
18:44:24.461 -> Water_Temp.:104.79, PWM:56
18:44:25.253 -> Water_Temp.:104.79, PWM:55
18:44:26.013 -> Water_Temp.:104.79, PWM:56
18:44:26.810 -> Water_Temp.:104.79, PWM:56
18:44:27.573 -> Water_Temp.:104.90, PWM:56
18:44:28.365 -> Water_Temp.:104.90, PWM:56
18:44:29.126 -> Water_Temp.:104.90, PWM:55
18:44:29.917 -> Water_Temp.:104.90, PWM:56
18:44:31.471 -> Water_Temp.:104.90, PWM:55
18:44:32.264 -> Water_Temp.:104.90, PWM:56
18:44:33.057 -> Water_Temp.:104.90, PWM:56
18:44:33.814 -> Water_Temp.:104.90, PWM:56
18:44:34.606 -> Water_Temp.:104.90, PWM:55
18:44:35.367 -> Water_Temp.:104.90, PWM:56
18:44:36.164 -> Water_Temp.:105.01, PWM:55
18:44:36.923 -> Water_Temp.:105.01, PWM:56
18:44:37.718 -> Water_Temp.:105.01, PWM:55
18:44:38.511 -> Water_Temp.:105.01, PWM:56
18:44:39.273 -> Water_Temp.:105.01, PWM:55
18:44:40.071 -> Water_Temp.:105.01, PWM:56
18:44:40.830 -> Water_Temp.:105.01, PWM:56
18:44:41.626 -> Water_Temp.:105.01, PWM:56
18:44:42.389 -> Water_Temp.:105.01, PWM:55
18:44:43.181 -> Water_Temp.:105.01, PWM:55
18:44:44.728 -> Water_Temp.:105.01, PWM:56
18:44:45.520 -> Water_Temp.:105.01, PWM:55
18:44:46.314 -> Water_Temp.:105.01, PWM:56
18:44:47.071 -> Water_Temp.:105.01, PWM:56
18:44:47.866 -> Water_Temp.:105.01, PWM:55
18:44:48.627 -> Water_Temp.:105.01, PWM:56
18:44:49.418 -> Water_Temp.:105.01, PWM:56
18:44:50.180 -> Water_Temp.:105.01, PWM:56
18:44:50.967 -> Water_Temp.:105.01, PWM:56
18:44:51.759 -> Water_Temp.:105.01, PWM:56
18:44:52.551 -> Water_Temp.:105.01, PWM:55
18:44:53.307 -> Water_Temp.:105.01, PWM:56
18:44:54.099 -> Water_Temp.:105.01, PWM:55
18:44:54.888 -> Water_Temp.:105.01, PWM:55
18:44:55.645 -> Water_Temp.:105.01, PWM:56
18:44:56.439 -> Water_Temp.:105.01, PWM:56
18:44:57.230 -> Water_Temp.:105.01, PWM:56
18:44:57.991 -> Water_Temp.:105.12, PWM:55
18:44:58.786 -> Water_Temp.:105.01, PWM:56
18:44:59.545 -> Water_Temp.:105.12, PWM:56
18:45:00.337 -> Water_Temp.:105.12, PWM:55
18:45:01.133 -> Water_Temp.:105.12, PWM:56
18:45:01.894 -> Water_Temp.:105.12, PWM:55
18:45:02.684 -> Water_Temp.:105.12, PWM:56
18:45:03.443 -> Water_Temp.:105.12, PWM:56
18:45:05.796 -> Water_Temp.:105.12, PWM:56
18:45:06.552 -> Water_Temp.:105.12, PWM:56
18:45:07.347 -> Water_Temp.:105.12, PWM:55
18:45:08.134 -> Water_Temp.:105.12, PWM:56
18:45:09.687 -> Water_Temp.:105.12, PWM:55
18:45:10.449 -> Water_Temp.:105.24, PWM:55
18:45:11.239 -> Water_Temp.:105.12, PWM:55
18:45:12.031 -> Water_Temp.:105.12, PWM:56
18:45:12.788 -> Water_Temp.:105.24, PWM:56
18:45:13.580 -> Water_Temp.:105.24, PWM:56
18:45:14.372 -> Water_Temp.:105.24, PWM:56
18:45:15.134 -> Water_Temp.:105.24, PWM:56
18:45:15.924 -> Water_Temp.:105.24, PWM:56
18:45:17.468 -> Water_Temp.:105.24, PWM:56
18:45:18.227 -> Water_Temp.:105.24, PWM:55
18:45:19.023 -> Water_Temp.:105.24, PWM:56
18:45:19.815 -> Water_Temp.:105.24, PWM:55
18:45:20.569 -> Water_Temp.:105.24, PWM:55
18:45:21.359 -> Water_Temp.:105.24, PWM:56
18:45:22.154 -> Water_Temp.:105.24, PWM:55
18:45:22.918 -> Water_Temp.:105.24, PWM:55
18:45:23.711 -> Water_Temp.:105.24, PWM:55
18:45:24.469 -> Water_Temp.:105.35, PWM:56
18:45:25.262 -> Water_Temp.:105.35, PWM:56
18:45:26.058 -> Water_Temp.:105.35, PWM:55
18:45:26.820 -> Water_Temp.:105.35, PWM:56
18:45:27.614 -> Water_Temp.:105.35, PWM:55
18:45:28.377 -> Water_Temp.:105.35, PWM:56
18:45:29.170 -> Water_Temp.:105.35, PWM:55
18:45:29.931 -> Water_Temp.:105.35, PWM:56
18:45:30.722 -> Water_Temp.:105.46, PWM:56
18:45:31.515 -> Water_Temp.:105.46, PWM:56
18:45:32.275 -> Water_Temp.:105.46, PWM:55
18:45:33.072 -> Water_Temp.:107.04, PWM:56
18:45:33.834 -> Water_Temp.:109.06, PWM:56
18:45:34.627 -> Water_Temp.:106.70, PWM:56
18:45:35.418 -> Water_Temp.:106.36, PWM:55
18:45:36.180 -> Water_Temp.:106.14, PWM:56
18:45:36.973 -> Water_Temp.:106.02, PWM:56
18:45:37.734 -> Water_Temp.:105.91, PWM:56
18:45:38.528 -> Water_Temp.:105.91, PWM:55
18:45:39.319 -> Water_Temp.:105.80, PWM:56
18:45:40.078 -> Water_Temp.:105.80, PWM:56
18:45:40.869 -> Water_Temp.:105.69, PWM:56
18:45:41.660 -> Water_Temp.:105.69, PWM:55
18:45:42.424 -> Water_Temp.:105.69, PWM:55
18:45:43.217 -> Water_Temp.:105.69, PWM:56
18:45:43.977 -> Water_Temp.:105.69, PWM:56
18:45:44.774 -> Water_Temp.:105.69, PWM:55
18:45:46.284 -> Water_Temp.:105.57, PWM:56
18:45:47.075 -> Water_Temp.:105.57, PWM:56
18:45:47.867 -> Water_Temp.:105.57, PWM:56
18:45:48.628 -> Water_Temp.:105.57, PWM:55
18:45:49.426 -> Water_Temp.:105.57, PWM:55
18:45:50.978 -> Water_Temp.:105.57, PWM:56
18:45:51.735 -> Water_Temp.:105.57, PWM:56
18:45:52.530 -> Water_Temp.:105.69, PWM:55
18:45:53.289 -> Water_Temp.:105.57, PWM:55
18:45:54.075 -> Water_Temp.:105.57, PWM:55
18:45:55.661 -> Water_Temp.:105.69, PWM:55
18:45:57.971 -> Water_Temp.:105.69, PWM:56
18:45:58.766 -> Water_Temp.:105.69, PWM:55
18:45:59.529 -> Water_Temp.:105.69, PWM:55
18:46:00.324 -> Water_Temp.:105.69, PWM:56
18:46:01.084 -> Water_Temp.:105.69, PWM:56
18:46:02.635 -> Water_Temp.:105.69, PWM:56
18:46:03.429 -> Water_Temp.:105.69, PWM:56
18:46:04.189 -> Water_Temp.:105.69, PWM:56
18:46:06.505 -> Water_Temp.:105.69, PWM:55
18:46:07.300 -> Water_Temp.:105.69, PWM:56
18:46:08.058 -> Water_Temp.:105.69, PWM:56
18:46:08.851 -> Water_Temp.:105.69, PWM:56
18:46:09.639 -> Water_Temp.:105.69, PWM:56
18:46:11.197 -> Water_Temp.:105.69, PWM:55
18:46:11.953 -> Water_Temp.:105.69, PWM:56
18:46:12.753 -> Water_Temp.:105.69, PWM:55
18:46:13.511 -> Water_Temp.:105.69, PWM:55
18:46:15.097 -> Water_Temp.:105.69, PWM:56
18:46:15.855 -> Water_Temp.:105.69, PWM:0
18:46:16.648 -> Water_Temp.:105.69, PWM:0
18:46:17.411 -> Water_Temp.:105.69, PWM:0
18:46:18.207 -> Water_Temp.:105.69, PWM:0
18:46:19.001 -> Water_Temp.:105.69, PWM:0
18:46:19.758 -> Water_Temp.:105.69, PWM:0
18:46:20.552 -> Water_Temp.:105.69, PWM:0
18:46:21.309 -> Water_Temp.:105.80, PWM:0
18:46:22.099 -> Water_Temp.:105.80, PWM:0
18:46:22.889 -> Water_Temp.:105.80, PWM:0
18:46:23.682 -> Water_Temp.:105.80, PWM:0
18:46:24.441 -> Water_Temp.:105.80, PWM:0
18:46:25.235 -> Water_Temp.:105.80, PWM:0
18:46:25.992 -> Water_Temp.:105.91, PWM:0
18:46:26.787 -> Water_Temp.:105.91, PWM:0
18:46:27.582 -> Water_Temp.:105.91, PWM:0
18:46:28.338 -> Water_Temp.:105.91, PWM:0
18:46:29.131 -> Water_Temp.:105.91, PWM:0
18:46:29.893 -> Water_Temp.:105.91, PWM:0
18:46:30.691 -> Water_Temp.:105.91, PWM:0
18:46:31.448 -> Water_Temp.:105.91, PWM:0
18:46:32.244 -> Water_Temp.:105.91, PWM:0
18:46:33.035 -> Water_Temp.:105.91, PWM:0
18:46:33.795 -> Water_Temp.:105.91, PWM:0
18:46:34.585 -> Water_Temp.:105.91, PWM:0
18:46:35.380 -> Water_Temp.:105.91, PWM:0
18:46:36.144 -> Water_Temp.:105.91, PWM:0
18:46:36.939 -> Water_Temp.:105.91, PWM:0
18:46:37.696 -> Water_Temp.:105.91, PWM:0
18:46:38.495 -> Water_Temp.:105.91, PWM:0
18:46:39.259 -> Water_Temp.:105.91, PWM:0
18:46:40.052 -> Water_Temp.:105.91, PWM:0
18:46:40.810 -> Water_Temp.:105.91, PWM:0
18:46:41.605 -> Water_Temp.:105.91, PWM:0
18:46:42.398 -> Water_Temp.:105.91, PWM:0
18:46:43.159 -> Water_Temp.:105.91, PWM:0
18:46:43.954 -> Water_Temp.:106.02, PWM:0
18:46:44.711 -> Water_Temp.:106.02, PWM:0
18:46:45.508 -> Water_Temp.:106.02, PWM:0
18:46:46.301 -> Water_Temp.:106.02, PWM:0
18:46:47.064 -> Water_Temp.:106.02, PWM:0
18:46:47.857 -> Water_Temp.:106.02, PWM:0
18:46:48.618 -> Water_Temp.:106.02, PWM:0
18:46:49.411 -> Water_Temp.:106.14, PWM:0
18:46:50.172 -> Water_Temp.:106.02, PWM:0
18:46:50.962 -> Water_Temp.:106.14, PWM:0
18:46:51.756 -> Water_Temp.:106.14, PWM:0
18:46:52.516 -> Water_Temp.:106.14, PWM:0
18:46:53.313 -> Water_Temp.:106.14, PWM:0
18:46:54.073 -> Water_Temp.:106.14, PWM:0
18:46:54.871 -> Water_Temp.:106.14, PWM:0
18:46:55.633 -> Water_Temp.:106.14, PWM:0
18:46:56.432 -> Water_Temp.:106.14, PWM:0
18:46:57.191 -> Water_Temp.:106.14, PWM:0
18:46:57.977 -> Water_Temp.:106.14, PWM:0
18:46:58.772 -> Water_Temp.:106.14, PWM:0
18:46:59.534 -> Water_Temp.:106.02, PWM:0
18:47:00.327 -> Water_Temp.:106.02, PWM:0
18:47:01.086 -> Water_Temp.:106.02, PWM:0
18:47:01.879 -> Water_Temp.:106.02, PWM:0
18:47:02.672 -> Water_Temp.:106.02, PWM:0
18:47:03.430 -> Water_Temp.:106.02, PWM:0
18:47:04.224 -> Water_Temp.:106.02, PWM:0
18:47:05.015 -> Water_Temp.:106.02, PWM:0
18:47:05.772 -> Water_Temp.:106.02, PWM:0
18:47:06.570 -> Water_Temp.:106.02, PWM:0
18:47:07.330 -> Water_Temp.:106.02, PWM:0
18:47:08.124 -> Water_Temp.:106.02, PWM:0
18:47:08.915 -> Water_Temp.:106.02, PWM:0
18:47:09.669 -> Water_Temp.:106.02, PWM:0
18:47:10.463 -> Water_Temp.:105.91, PWM:0
18:47:11.258 -> Water_Temp.:105.91, PWM:0
18:47:12.018 -> Water_Temp.:105.91, PWM:0
18:47:12.812 -> Water_Temp.:105.91, PWM:0
18:47:13.570 -> Water_Temp.:105.91, PWM:0
18:47:14.363 -> Water_Temp.:105.91, PWM:0
18:47:15.159 -> Water_Temp.:105.91, PWM:0
18:47:15.918 -> Water_Temp.:105.91, PWM:0
18:47:16.713 -> Water_Temp.:105.91, PWM:0
18:47:17.470 -> Water_Temp.:105.91, PWM:0
18:47:18.267 -> Water_Temp.:105.91, PWM:0
18:47:19.059 -> Water_Temp.:105.80, PWM:0
18:47:19.815 -> Water_Temp.:105.80, PWM:0
18:47:20.606 -> Water_Temp.:105.80, PWM:0
18:47:21.367 -> Water_Temp.:105.80, PWM:0
18:47:22.162 -> Water_Temp.:105.80, PWM:0
18:47:22.956 -> Water_Temp.:105.80, PWM:0
18:47:23.718 -> Water_Temp.:105.80, PWM:0
18:47:24.513 -> Water_Temp.:105.80, PWM:0
18:47:25.271 -> Water_Temp.:105.80, PWM:0
18:47:26.068 -> Water_Temp.:105.80, PWM:0
18:47:26.828 -> Water_Temp.:105.80, PWM:0
18:47:27.616 -> Water_Temp.:105.80, PWM:0
18:47:28.407 -> Water_Temp.:105.80, PWM:0
18:47:29.171 -> Water_Temp.:105.80, PWM:0
18:47:29.965 -> Water_Temp.:105.80, PWM:0
18:47:30.728 -> Water_Temp.:105.80, PWM:0
18:47:31.523 -> Water_Temp.:105.80, PWM:0
18:47:32.286 -> Water_Temp.:105.80, PWM:0
18:47:33.080 -> Water_Temp.:105.80, PWM:0
18:47:33.876 -> Water_Temp.:105.80, PWM:0
18:47:34.635 -> Water_Temp.:105.80, PWM:0
18:47:35.428 -> Water_Temp.:105.80, PWM:0
18:47:36.186 -> Water_Temp.:105.80, PWM:0
18:47:36.979 -> Water_Temp.:105.80, PWM:0
18:47:37.774 -> Water_Temp.:105.80, PWM:0
18:47:38.533 -> Water_Temp.:105.80, PWM:0
18:47:39.325 -> Water_Temp.:105.80, PWM:0
18:47:40.114 -> Water_Temp.:105.80, PWM:0
18:47:40.871 -> Water_Temp.:105.80, PWM:0
18:47:41.668 -> Water_Temp.:105.69, PWM:0
18:47:42.427 -> Water_Temp.:105.69, PWM:0
18:47:43.217 -> Water_Temp.:105.69, PWM:0
18:47:44.010 -> Water_Temp.:105.69, PWM:0
18:47:44.768 -> Water_Temp.:105.69, PWM:0
18:47:45.563 -> Water_Temp.:105.69, PWM:0
18:47:46.355 -> Water_Temp.:105.69, PWM:0
18:47:47.113 -> Water_Temp.:105.57, PWM:0
18:47:47.908 -> Water_Temp.:105.57, PWM:0
18:47:48.667 -> Water_Temp.:105.57, PWM:0
18:47:49.461 -> Water_Temp.:105.57, PWM:0
18:47:50.225 -> Water_Temp.:105.57, PWM:0
18:47:51.024 -> Water_Temp.:105.57, PWM:0
18:47:51.787 -> Water_Temp.:105.57, PWM:0
18:47:52.580 -> Water_Temp.:105.46, PWM:0
18:47:53.372 -> Water_Temp.:105.46, PWM:0
18:47:54.133 -> Water_Temp.:105.46, PWM:0
18:47:54.925 -> Water_Temp.:105.46, PWM:0
18:47:55.687 -> Water_Temp.:105.46, PWM:0
18:47:56.481 -> Water_Temp.:105.46, PWM:0
18:47:57.243 -> Water_Temp.:105.35, PWM:0
18:47:58.033 -> Water_Temp.:105.35, PWM:0
18:47:58.829 -> Water_Temp.:105.35, PWM:0
18:47:59.591 -> Water_Temp.:105.35, PWM:0
18:48:00.382 -> Water_Temp.:105.35, PWM:0
18:48:01.142 -> Water_Temp.:105.35, PWM:0
18:48:01.936 -> Water_Temp.:105.35, PWM:0
18:48:02.733 -> Water_Temp.:105.35, PWM:0
18:48:03.492 -> Water_Temp.:105.24, PWM:0
18:48:04.287 -> Water_Temp.:105.35, PWM:0
18:48:05.047 -> Water_Temp.:105.24, PWM:0
18:48:05.843 -> Water_Temp.:105.24, PWM:0
18:48:06.604 -> Water_Temp.:105.24, PWM:0
18:48:07.399 -> Water_Temp.:105.24, PWM:0
18:48:08.196 -> Water_Temp.:105.24, PWM:0
18:48:08.959 -> Water_Temp.:105.24, PWM:0
18:48:09.754 -> Water_Temp.:105.24, PWM:0
18:48:10.514 -> Water_Temp.:105.24, PWM:0
18:48:11.311 -> Water_Temp.:105.24, PWM:0
18:48:12.074 -> Water_Temp.:105.24, PWM:0
18:48:12.868 -> Water_Temp.:105.24, PWM:0
18:48:13.629 -> Water_Temp.:105.24, PWM:0
18:48:14.426 -> Water_Temp.:105.24, PWM:0
18:48:15.190 -> Water_Temp.:105.24, PWM:0
18:48:15.986 -> Water_Temp.:105.24, PWM:0
18:48:16.750 -> Water_Temp.:105.24, PWM:0
18:48:17.542 -> Water_Temp.:105.24, PWM:0
18:48:18.331 -> Water_Temp.:105.12, PWM:0
18:48:19.090 -> Water_Temp.:105.12, PWM:0
18:48:19.886 -> Water_Temp.:105.12, PWM:0
18:48:20.645 -> Water_Temp.:105.12, PWM:0
18:48:21.436 -> Water_Temp.:105.12, PWM:0
18:48:22.231 -> Water_Temp.:105.12, PWM:0
18:48:22.994 -> Water_Temp.:105.12, PWM:0
18:48:23.785 -> Water_Temp.:105.12, PWM:0
18:48:24.550 -> Water_Temp.:105.12, PWM:0
18:48:25.346 -> Water_Temp.:105.01, PWM:0
18:48:26.104 -> Water_Temp.:105.01, PWM:0
18:48:26.898 -> Water_Temp.:105.01, PWM:0
18:48:27.692 -> Water_Temp.:105.01, PWM:0
18:48:28.454 -> Water_Temp.:105.01, PWM:0
18:48:29.251 -> Water_Temp.:104.90, PWM:0
18:48:30.010 -> Water_Temp.:104.90, PWM:0
18:48:30.805 -> Water_Temp.:104.90, PWM:0
18:48:31.567 -> Water_Temp.:104.90, PWM:0
18:48:32.366 -> Water_Temp.:104.90, PWM:0
18:48:33.127 -> Water_Temp.:104.90, PWM:0
18:48:33.919 -> Water_Temp.:104.90, PWM:0
18:48:34.684 -> Water_Temp.:104.90, PWM:0
18:48:35.483 -> Water_Temp.:104.79, PWM:0
18:48:36.247 -> Water_Temp.:104.79, PWM:0
18:48:37.041 -> Water_Temp.:104.79, PWM:0
18:48:37.805 -> Water_Temp.:104.79, PWM:0
18:48:38.597 -> Water_Temp.:104.79, PWM:0
18:48:39.392 -> Water_Temp.:104.79, PWM:0
18:48:40.149 -> Water_Temp.:104.79, PWM:0
18:48:40.944 -> Water_Temp.:104.79, PWM:0
18:48:41.704 -> Water_Temp.:104.67, PWM:0
18:48:42.498 -> Water_Temp.:104.67, PWM:0
18:48:43.260 -> Water_Temp.:104.67, PWM:0
18:48:44.056 -> Water_Temp.:104.67, PWM:0
18:48:44.848 -> Water_Temp.:104.67, PWM:0
18:48:45.609 -> Water_Temp.:104.67, PWM:0
18:48:46.401 -> Water_Temp.:104.67, PWM:0
18:48:47.162 -> Water_Temp.:104.56, PWM:0
18:48:47.956 -> Water_Temp.:104.56, PWM:0
18:48:48.746 -> Water_Temp.:104.56, PWM:0
18:48:49.507 -> Water_Temp.:104.56, PWM:0
18:48:50.304 -> Water_Temp.:104.56, PWM:0
18:48:51.067 -> Water_Temp.:104.56, PWM:0
18:48:51.862 -> Water_Temp.:104.56, PWM:0
18:48:52.629 -> Water_Temp.:104.45, PWM:0
18:48:53.422 -> Water_Temp.:104.45, PWM:0
18:48:54.183 -> Water_Temp.:104.45, PWM:0
18:48:54.974 -> Water_Temp.:104.45, PWM:0
18:48:55.769 -> Water_Temp.:104.45, PWM:0
18:48:56.526 -> Water_Temp.:104.45, PWM:0
18:48:57.318 -> Water_Temp.:104.45, PWM:0
18:48:58.080 -> Water_Temp.:104.45, PWM:0
18:48:58.875 -> Water_Temp.:104.34, PWM:0
18:48:59.673 -> Water_Temp.:104.34, PWM:0
18:49:00.436 -> Water_Temp.:104.34, PWM:0
18:49:01.199 -> Water_Temp.:104.34, PWM:0
18:49:01.994 -> Water_Temp.:104.34, PWM:0
18:49:02.790 -> Water_Temp.:104.34, PWM:0
18:49:03.552 -> Water_Temp.:104.34, PWM:0
18:49:04.345 -> Water_Temp.:104.22, PWM:0
18:49:05.105 -> Water_Temp.:104.22, PWM:0
18:49:05.902 -> Water_Temp.:104.22, PWM:0
18:49:06.661 -> Water_Temp.:104.22, PWM:0
18:49:07.456 -> Water_Temp.:104.22, PWM:0
18:49:08.250 -> Water_Temp.:104.11, PWM:0
18:49:09.011 -> Water_Temp.:104.11, PWM:0
18:49:09.810 -> Water_Temp.:104.11, PWM:0
18:49:10.578 -> Water_Temp.:104.11, PWM:0
18:49:11.346 -> Water_Temp.:104.11, PWM:0
18:49:12.146 -> Water_Temp.:104.11, PWM:0
18:49:12.908 -> Water_Temp.:104.11, PWM:0
18:49:13.703 -> Water_Temp.:104.00, PWM:0
18:49:14.459 -> Water_Temp.:104.00, PWM:0
18:49:15.252 -> Water_Temp.:104.00, PWM:0
18:49:16.049 -> Water_Temp.:104.00, PWM:0
18:49:16.809 -> Water_Temp.:104.00, PWM:0
18:49:17.605 -> Water_Temp.:103.89, PWM:0
18:49:18.369 -> Water_Temp.:103.89, PWM:0
18:49:19.167 -> Water_Temp.:103.89, PWM:0
18:49:19.928 -> Water_Temp.:103.89, PWM:0
18:49:20.720 -> Water_Temp.:103.89, PWM:0
18:49:21.478 -> Water_Temp.:103.77, PWM:0
18:49:22.277 -> Water_Temp.:103.77, PWM:0
18:49:23.070 -> Water_Temp.:103.77, PWM:0
18:49:23.831 -> Water_Temp.:103.77, PWM:0
18:49:24.625 -> Water_Temp.:103.77, PWM:0
18:49:25.385 -> Water_Temp.:103.66, PWM:0
18:49:26.179 -> Water_Temp.:103.66, PWM:0
18:49:26.938 -> Water_Temp.:103.66, PWM:0
18:49:27.736 -> Water_Temp.:103.66, PWM:0
18:49:28.498 -> Water_Temp.:103.66, PWM:0
18:49:29.293 -> Water_Temp.:103.66, PWM:0
18:49:30.091 -> Water_Temp.:103.66, PWM:0
18:49:30.851 -> Water_Temp.:103.66, PWM:0
18:49:31.645 -> Water_Temp.:103.55, PWM:0
18:49:32.403 -> Water_Temp.:103.66, PWM:0
18:49:33.200 -> Water_Temp.:103.55, PWM:0
18:49:33.957 -> Water_Temp.:103.55, PWM:0
18:49:34.752 -> Water_Temp.:103.55, PWM:0
18:49:35.547 -> Water_Temp.:103.55, PWM:0
18:49:36.310 -> Water_Temp.:103.55, PWM:0
18:49:37.103 -> Water_Temp.:103.55, PWM:0
18:49:37.863 -> Water_Temp.:103.55, PWM:0
18:49:38.653 -> Water_Temp.:103.55, PWM:0
18:49:39.450 -> Water_Temp.:103.55, PWM:0
18:49:40.216 -> Water_Temp.:103.55, PWM:0
18:49:40.977 -> Water_Temp.:103.55, PWM:0
18:49:41.775 -> Water_Temp.:103.44, PWM:0
18:49:42.563 -> Water_Temp.:103.44, PWM:0
18:49:43.319 -> Water_Temp.:103.44, PWM:0
18:49:44.112 -> Water_Temp.:103.44, PWM:0
18:49:44.910 -> Water_Temp.:103.44, PWM:0
18:49:45.671 -> Water_Temp.:103.44, PWM:0
18:49:46.463 -> Water_Temp.:103.44, PWM:0
18:49:47.227 -> Water_Temp.:103.44, PWM:0
18:49:48.022 -> Water_Temp.:103.44, PWM:0
18:49:48.781 -> Water_Temp.:103.44, PWM:0
18:49:49.575 -> Water_Temp.:103.32, PWM:0
18:49:50.337 -> Water_Temp.:103.32, PWM:0
18:49:51.126 -> Water_Temp.:103.32, PWM:0
18:49:51.924 -> Water_Temp.:103.32, PWM:0
18:49:53.441 -> Water_Temp.:103.32, PWM:8
18:49:54.232 -> Water_Temp.:103.32, PWM:24
18:49:54.993 -> Water_Temp.:103.32, PWM:31
18:49:55.791 -> Water_Temp.:103.21, PWM:31
18:49:56.556 -> Water_Temp.:103.21, PWM:30
18:49:57.347 -> Water_Temp.:103.21, PWM:32
18:49:58.143 -> Water_Temp.:103.21, PWM:32
18:49:58.905 -> Water_Temp.:103.21, PWM:36
18:49:59.700 -> Water_Temp.:103.21, PWM:40
18:50:00.461 -> Water_Temp.:103.21, PWM:45
18:50:01.260 -> Water_Temp.:103.21, PWM:45
18:50:02.020 -> Water_Temp.:103.21, PWM:45
18:50:03.571 -> Water_Temp.:103.10, PWM:45
18:50:04.360 -> Water_Temp.:103.10, PWM:45
18:50:05.158 -> Water_Temp.:103.10, PWM:45
18:50:05.921 -> Water_Temp.:103.10, PWM:45
18:50:06.717 -> Water_Temp.:103.10, PWM:45
18:50:07.478 -> Water_Temp.:102.99, PWM:45
18:50:08.271 -> Water_Temp.:102.99, PWM:45
18:50:09.032 -> Water_Temp.:102.99, PWM:45
18:50:09.831 -> Water_Temp.:102.99, PWM:45
18:50:10.589 -> Water_Temp.:102.99, PWM:45
18:50:11.384 -> Water_Temp.:102.99, PWM:45
18:50:12.181 -> Water_Temp.:102.99, PWM:45
18:50:12.941 -> Water_Temp.:102.99, PWM:45
18:50:14.502 -> Water_Temp.:102.87, PWM:45
18:50:15.295 -> Water_Temp.:102.87, PWM:45
18:50:16.055 -> Water_Temp.:102.87, PWM:45
18:50:16.848 -> Water_Temp.:102.87, PWM:45
18:50:17.607 -> Water_Temp.:102.87, PWM:45
18:50:18.399 -> Water_Temp.:102.76, PWM:45
18:50:19.199 -> Water_Temp.:102.76, PWM:45
18:50:20.754 -> Water_Temp.:102.76, PWM:45
18:50:21.510 -> Water_Temp.:102.76, PWM:45
18:50:22.305 -> Water_Temp.:102.76, PWM:45
18:50:23.857 -> Water_Temp.:102.65, PWM:45
18:50:24.651 -> Water_Temp.:102.65, PWM:45
18:50:25.411 -> Water_Temp.:102.65, PWM:45
18:50:26.205 -> Water_Temp.:102.65, PWM:45
18:50:26.970 -> Water_Temp.:102.65, PWM:45
18:50:27.765 -> Water_Temp.:102.54, PWM:45
18:50:28.555 -> Water_Temp.:102.54, PWM:45
18:50:29.317 -> Water_Temp.:102.54, PWM:45
18:50:30.111 -> Water_Temp.:102.54, PWM:45
18:50:30.870 -> Water_Temp.:102.54, PWM:45
18:50:31.666 -> Water_Temp.:102.42, PWM:45
18:50:32.457 -> Water_Temp.:102.42, PWM:45
18:50:33.219 -> Water_Temp.:102.42, PWM:45
18:50:34.013 -> Water_Temp.:102.42, PWM:45
18:50:34.769 -> Water_Temp.:102.42, PWM:45
18:50:35.564 -> Water_Temp.:102.42, PWM:45
18:50:36.323 -> Water_Temp.:102.42, PWM:45
18:50:37.116 -> Water_Temp.:102.42, PWM:45
18:50:37.912 -> Water_Temp.:102.42, PWM:45
18:50:38.674 -> Water_Temp.:102.31, PWM:45
18:50:39.470 -> Water_Temp.:102.31, PWM:45
18:50:40.232 -> Water_Temp.:102.31, PWM:45
18:50:41.031 -> Water_Temp.:102.31, PWM:45
18:50:41.793 -> Water_Temp.:102.31, PWM:45
18:50:42.587 -> Water_Temp.:102.31, PWM:45
18:50:43.345 -> Water_Temp.:102.20, PWM:45
18:50:44.142 -> Water_Temp.:102.20, PWM:53
18:50:44.937 -> Water_Temp.:102.20, PWM:52
18:50:45.700 -> Water_Temp.:102.20, PWM:53
18:50:46.498 -> Water_Temp.:102.20, PWM:53
18:50:47.259 -> Water_Temp.:102.20, PWM:56
18:50:48.813 -> Water_Temp.:102.20, PWM:56
18:50:50.368 -> Water_Temp.:102.09, PWM:56
18:50:51.167 -> Water_Temp.:102.09, PWM:53
18:50:51.931 -> Water_Temp.:102.09, PWM:53
18:50:52.726 -> Water_Temp.:102.09, PWM:53
18:50:53.511 -> Water_Temp.:102.09, PWM:53
18:50:54.271 -> Water_Temp.:102.09, PWM:53
18:50:55.069 -> Water_Temp.:102.09, PWM:55
18:50:55.833 -> Water_Temp.:102.09, PWM:57
18:50:57.384 -> Water_Temp.:101.97, PWM:59
18:50:58.177 -> Water_Temp.:101.97, PWM:57
18:50:58.941 -> Water_Temp.:101.97, PWM:57
18:50:59.735 -> Water_Temp.:101.97, PWM:56
18:51:01.296 -> Water_Temp.:101.97, PWM:56
18:51:02.058 -> Water_Temp.:101.97, PWM:56
18:51:02.851 -> Water_Temp.:101.86, PWM:56
18:51:03.610 -> Water_Temp.:101.86, PWM:55
18:51:04.406 -> Water_Temp.:101.86, PWM:56
18:51:05.199 -> Water_Temp.:101.86, PWM:56
18:51:05.961 -> Water_Temp.:101.86, PWM:55
18:51:06.758 -> Water_Temp.:101.86, PWM:56
18:51:08.275 -> Water_Temp.:101.75, PWM:55
18:51:09.072 -> Water_Temp.:101.75, PWM:55
18:51:09.837 -> Water_Temp.:101.75, PWM:55
18:51:10.632 -> Water_Temp.:101.75, PWM:56
18:51:11.392 -> Water_Temp.:101.75, PWM:56
18:51:12.189 -> Water_Temp.:101.75, PWM:56
18:51:12.949 -> Water_Temp.:101.75, PWM:55
18:51:13.744 -> Water_Temp.:101.75, PWM:56
18:51:14.540 -> Water_Temp.:101.75, PWM:56
18:51:15.301 -> Water_Temp.:101.75, PWM:55
18:51:16.095 -> Water_Temp.:101.75, PWM:56
18:51:16.854 -> Water_Temp.:101.75, PWM:56
18:51:17.649 -> Water_Temp.:101.75, PWM:56
18:51:18.445 -> Water_Temp.:101.64, PWM:55
18:51:19.203 -> Water_Temp.:101.64, PWM:55
18:51:19.991 -> Water_Temp.:101.64, PWM:55
18:51:20.777 -> Water_Temp.:101.64, PWM:55
18:51:21.537 -> Water_Temp.:101.64, PWM:55
18:51:22.327 -> Water_Temp.:101.64, PWM:55
18:51:23.121 -> Water_Temp.:101.64, PWM:55
18:51:23.883 -> Water_Temp.:101.64, PWM:55
18:51:24.675 -> Water_Temp.:101.64, PWM:55
18:51:25.438 -> Water_Temp.:101.64, PWM:55
18:51:26.231 -> Water_Temp.:101.64, PWM:55
18:51:26.992 -> Water_Temp.:101.64, PWM:55
18:51:27.787 -> Water_Temp.:101.75, PWM:55
18:51:28.551 -> Water_Temp.:101.75, PWM:55
18:51:29.348 -> Water_Temp.:101.75, PWM:56
18:51:30.141 -> Water_Temp.:101.75, PWM:55
18:51:30.903 -> Water_Temp.:101.75, PWM:55
18:51:31.696 -> Water_Temp.:101.64, PWM:55
18:51:32.454 -> Water_Temp.:101.75, PWM:55
18:51:33.245 -> Water_Temp.:101.75, PWM:55
18:51:34.041 -> Water_Temp.:101.75, PWM:55
18:51:34.802 -> Water_Temp.:101.75, PWM:55
18:51:35.600 -> Water_Temp.:101.75, PWM:55
18:51:36.357 -> Water_Temp.:101.75, PWM:55
18:51:37.152 -> Water_Temp.:101.75, PWM:55
18:51:37.911 -> Water_Temp.:101.75, PWM:53
18:51:39.501 -> Water_Temp.:101.75, PWM:53
18:51:40.261 -> Water_Temp.:101.75, PWM:52
18:51:41.816 -> Water_Temp.:101.75, PWM:51
18:51:42.609 -> Water_Temp.:101.75, PWM:51
18:51:43.369 -> Water_Temp.:101.75, PWM:51
18:51:44.167 -> Water_Temp.:101.75, PWM:51
18:51:44.960 -> Water_Temp.:101.75, PWM:51
18:51:45.722 -> Water_Temp.:101.75, PWM:51
18:51:46.518 -> Water_Temp.:101.75, PWM:51
18:51:47.279 -> Water_Temp.:101.75, PWM:51
18:51:48.073 -> Water_Temp.:101.75, PWM:51
18:51:48.834 -> Water_Temp.:101.75, PWM:51
18:51:50.391 -> Water_Temp.:101.75, PWM:51
18:51:51.151 -> Water_Temp.:101.75, PWM:51
18:51:51.942 -> Water_Temp.:101.75, PWM:51
18:51:52.733 -> Water_Temp.:101.75, PWM:51
18:51:53.495 -> Water_Temp.:101.75, PWM:51
18:51:54.284 -> Water_Temp.:101.75, PWM:51
18:51:55.047 -> Water_Temp.:101.75, PWM:51
18:51:56.601 -> Water_Temp.:101.75, PWM:51
18:51:57.360 -> Water_Temp.:101.75, PWM:51
18:51:58.156 -> Water_Temp.:101.75, PWM:51
18:51:58.914 -> Water_Temp.:101.75, PWM:51
18:51:59.709 -> Water_Temp.:101.75, PWM:51
18:52:00.503 -> Water_Temp.:101.75, PWM:51
18:52:01.263 -> Water_Temp.:101.75, PWM:51
18:52:02.818 -> Water_Temp.:101.75, PWM:51
18:52:03.580 -> Water_Temp.:101.75, PWM:51
18:52:04.374 -> Water_Temp.:101.75, PWM:51
18:52:05.168 -> Water_Temp.:101.75, PWM:51
18:52:05.932 -> Water_Temp.:101.75, PWM:51
18:52:06.729 -> Water_Temp.:101.75, PWM:51
18:52:07.486 -> Water_Temp.:101.75, PWM:51
18:52:08.283 -> Water_Temp.:101.75, PWM:51
18:52:09.045 -> Water_Temp.:101.75, PWM:51
18:52:09.842 -> Water_Temp.:101.75, PWM:51
18:52:10.603 -> Water_Temp.:101.75, PWM:51
18:52:11.398 -> Water_Temp.:101.75, PWM:51
18:52:12.194 -> Water_Temp.:101.75, PWM:51
18:52:12.954 -> Water_Temp.:101.75, PWM:51
18:52:13.748 -> Water_Temp.:101.75, PWM:51
18:52:14.510 -> Water_Temp.:101.75, PWM:51
18:52:15.308 -> Water_Temp.:101.75, PWM:51
18:52:16.831 -> Water_Temp.:101.75, PWM:51
18:52:17.627 -> Water_Temp.:101.75, PWM:51
18:52:19.181 -> Water_Temp.:101.75, PWM:51
18:52:19.942 -> Water_Temp.:101.75, PWM:51
18:52:20.739 -> Water_Temp.:101.75, PWM:51
18:52:21.502 -> Water_Temp.:101.86, PWM:51
18:52:22.296 -> Water_Temp.:101.86, PWM:51
18:52:23.057 -> Water_Temp.:101.86, PWM:55
18:52:23.850 -> Water_Temp.:101.86, PWM:55
18:52:24.645 -> Water_Temp.:101.86, PWM:55
18:52:25.404 -> Water_Temp.:101.86, PWM:61
18:52:26.201 -> Water_Temp.:101.86, PWM:61
18:52:26.962 -> Water_Temp.:101.86, PWM:61
18:52:27.754 -> Water_Temp.:101.86, PWM:60
18:52:28.515 -> Water_Temp.:101.86, PWM:60
18:52:29.312 -> Water_Temp.:101.86, PWM:60
18:52:30.107 -> Water_Temp.:101.86, PWM:60
18:52:30.865 -> Water_Temp.:101.86, PWM:59
18:52:31.659 -> Water_Temp.:101.86, PWM:59
18:52:32.419 -> Water_Temp.:101.86, PWM:59
18:52:33.218 -> Water_Temp.:101.86, PWM:59
18:52:33.978 -> Water_Temp.:101.86, PWM:59
18:52:34.773 -> Water_Temp.:101.86, PWM:59
18:52:35.537 -> Water_Temp.:101.86, PWM:59
18:52:36.331 -> Water_Temp.:101.86, PWM:59
18:52:37.127 -> Water_Temp.:101.86, PWM:59
18:52:37.884 -> Water_Temp.:101.86, PWM:59
18:52:38.675 -> Water_Temp.:101.75, PWM:59
18:52:39.435 -> Water_Temp.:101.75, PWM:59
18:52:40.232 -> Water_Temp.:101.75, PWM:59
18:52:41.023 -> Water_Temp.:101.75, PWM:59
18:52:41.783 -> Water_Temp.:101.75, PWM:59
18:52:42.582 -> Water_Temp.:101.75, PWM:59
18:52:43.341 -> Water_Temp.:101.75, PWM:59
18:52:44.895 -> Water_Temp.:101.75, PWM:59
18:52:45.692 -> Water_Temp.:101.75, PWM:59
18:52:46.487 -> Water_Temp.:101.75, PWM:59
18:52:47.248 -> Water_Temp.:101.75, PWM:59
18:52:48.045 -> Water_Temp.:101.75, PWM:59
18:52:48.803 -> Water_Temp.:101.75, PWM:59
18:52:49.603 -> Water_Temp.:101.75, PWM:56
18:52:50.364 -> Water_Temp.:101.75, PWM:56
18:52:51.156 -> Water_Temp.:101.75, PWM:56
18:52:51.918 -> Water_Temp.:101.75, PWM:56
18:52:52.712 -> Water_Temp.:101.75, PWM:56
18:52:53.474 -> Water_Temp.:101.75, PWM:56
18:52:55.027 -> Water_Temp.:101.75, PWM:56
18:52:55.825 -> Water_Temp.:101.75, PWM:56
18:52:56.589 -> Water_Temp.:101.75, PWM:56
18:52:58.142 -> Water_Temp.:101.75, PWM:56
18:52:58.936 -> Water_Temp.:101.75, PWM:56
18:52:59.699 -> Water_Temp.:101.75, PWM:56
18:53:01.249 -> Water_Temp.:101.75, PWM:56
18:53:02.011 -> Water_Temp.:101.75, PWM:56
18:53:02.809 -> Water_Temp.:101.75, PWM:56
18:53:03.569 -> Water_Temp.:101.75, PWM:56
18:53:04.360 -> Water_Temp.:101.75, PWM:56
18:53:05.155 -> Water_Temp.:101.75, PWM:56
18:53:05.912 -> Water_Temp.:101.75, PWM:56
18:53:06.709 -> Water_Temp.:101.86, PWM:56
18:53:07.471 -> Water_Temp.:101.75, PWM:56
18:53:08.267 -> Water_Temp.:101.75, PWM:56
18:53:09.060 -> Water_Temp.:101.86, PWM:56
18:53:09.824 -> Water_Temp.:101.86, PWM:56
18:53:10.587 -> Water_Temp.:101.86, PWM:56
18:53:12.169 -> Water_Temp.:101.86, PWM:56
18:53:12.931 -> Water_Temp.:101.86, PWM:56
18:53:13.729 -> Water_Temp.:101.86, PWM:56
18:53:14.487 -> Water_Temp.:101.86, PWM:56
18:53:15.286 -> Water_Temp.:101.86, PWM:56
18:53:16.050 -> Water_Temp.:101.86, PWM:56
18:53:16.844 -> Water_Temp.:101.86, PWM:56
18:53:17.607 -> Water_Temp.:101.86, PWM:56
18:53:18.403 -> Water_Temp.:101.86, PWM:56
18:53:19.199 -> Water_Temp.:101.86, PWM:56
18:53:19.955 -> Water_Temp.:101.86, PWM:56
18:53:22.301 -> Water_Temp.:101.86, PWM:56
18:53:23.064 -> Water_Temp.:101.86, PWM:56
18:53:23.863 -> Water_Temp.:101.86, PWM:56
18:53:24.623 -> Water_Temp.:101.75, PWM:56
18:53:25.416 -> Water_Temp.:101.86, PWM:56
18:53:26.176 -> Water_Temp.:101.86, PWM:56
18:53:26.971 -> Water_Temp.:101.86, PWM:56
18:53:27.734 -> Water_Temp.:101.75, PWM:56
18:53:28.528 -> Water_Temp.:101.75, PWM:56
18:53:29.326 -> Water_Temp.:101.75, PWM:56
18:53:30.088 -> Water_Temp.:101.75, PWM:56
18:53:30.881 -> Water_Temp.:101.75, PWM:56
18:53:31.641 -> Water_Temp.:101.86, PWM:56
18:53:32.439 -> Water_Temp.:101.86, PWM:56
18:53:33.200 -> Water_Temp.:101.86, PWM:56
18:53:33.995 -> Water_Temp.:101.86, PWM:56
18:53:35.551 -> Water_Temp.:101.86, PWM:57
18:53:36.311 -> Water_Temp.:101.86, PWM:60
18:53:37.108 -> Water_Temp.:101.86, PWM:60
18:53:38.628 -> Water_Temp.:101.86, PWM:59
18:53:39.427 -> Water_Temp.:101.86, PWM:59
18:53:40.189 -> Water_Temp.:101.86, PWM:60
18:53:40.981 -> Water_Temp.:101.86, PWM:60
18:53:41.777 -> Water_Temp.:101.86, PWM:59
18:53:42.539 -> Water_Temp.:101.86, PWM:60
18:53:43.327 -> Water_Temp.:101.86, PWM:60
18:53:44.084 -> Water_Temp.:101.86, PWM:60
18:53:44.878 -> Water_Temp.:101.86, PWM:60
18:53:45.676 -> Water_Temp.:101.86, PWM:60
18:53:46.434 -> Water_Temp.:101.86, PWM:60
18:53:47.231 -> Water_Temp.:101.86, PWM:60
18:53:47.992 -> Water_Temp.:101.86, PWM:59
18:53:48.785 -> Water_Temp.:101.86, PWM:59
18:53:49.544 -> Water_Temp.:101.86, PWM:60
18:53:50.336 -> Water_Temp.:101.86, PWM:60
18:53:51.130 -> Water_Temp.:101.97, PWM:59
18:53:51.885 -> Water_Temp.:101.97, PWM:60
18:53:53.439 -> Water_Temp.:101.97, PWM:59
18:53:54.199 -> Water_Temp.:101.97, PWM:59
18:53:54.992 -> Water_Temp.:101.97, PWM:60
18:53:55.788 -> Water_Temp.:101.97, PWM:59
18:53:56.550 -> Water_Temp.:101.97, PWM:59
18:53:57.347 -> Water_Temp.:101.97, PWM:59
18:53:58.106 -> Water_Temp.:101.97, PWM:59
18:53:58.899 -> Water_Temp.:101.97, PWM:60
18:53:59.663 -> Water_Temp.:101.97, PWM:59
18:54:00.461 -> Water_Temp.:101.97, PWM:59
18:54:01.252 -> Water_Temp.:101.97, PWM:60
18:54:02.010 -> Water_Temp.:101.97, PWM:59
18:54:02.806 -> Water_Temp.:101.97, PWM:59
18:54:03.568 -> Water_Temp.:101.97, PWM:60
18:54:04.363 -> Water_Temp.:101.97, PWM:60
18:54:05.123 -> Water_Temp.:101.97, PWM:60
18:54:05.920 -> Water_Temp.:101.97, PWM:60
18:54:06.681 -> Water_Temp.:101.97, PWM:59
18:54:07.476 -> Water_Temp.:101.97, PWM:59
18:54:08.239 -> Water_Temp.:101.97, PWM:60
18:54:09.037 -> Water_Temp.:101.97, PWM:59
18:54:10.592 -> Water_Temp.:101.97, PWM:59
18:54:11.355 -> Water_Temp.:101.97, PWM:59
18:54:12.148 -> Water_Temp.:101.97, PWM:59
18:54:12.904 -> Water_Temp.:101.97, PWM:60
18:54:14.495 -> Water_Temp.:101.97, PWM:60
18:54:15.257 -> Water_Temp.:101.97, PWM:60
18:54:16.054 -> Water_Temp.:101.97, PWM:59
18:54:16.816 -> Water_Temp.:101.97, PWM:59
18:54:17.612 -> Water_Temp.:101.97, PWM:59
18:54:18.371 -> Water_Temp.:101.97, PWM:59
18:54:19.163 -> Water_Temp.:101.97, PWM:59
18:54:19.925 -> Water_Temp.:101.97, PWM:59
18:54:20.721 -> Water_Temp.:101.97, PWM:59
18:54:21.513 -> Water_Temp.:101.97, PWM:60
18:54:22.273 -> Water_Temp.:101.97, PWM:59
18:54:23.070 -> Water_Temp.:101.97, PWM:60
18:54:23.830 -> Water_Temp.:101.97, PWM:59
18:54:24.628 -> Water_Temp.:101.97, PWM:59
18:54:25.390 -> Water_Temp.:101.97, PWM:59
18:54:26.185 -> Water_Temp.:101.97, PWM:60
18:54:26.948 -> Water_Temp.:101.97, PWM:59
18:54:27.746 -> Water_Temp.:102.09, PWM:59
18:54:28.504 -> Water_Temp.:101.97, PWM:60
18:54:29.302 -> Water_Temp.:102.09, PWM:60
18:54:30.829 -> Water_Temp.:102.09, PWM:59
18:54:31.622 -> Water_Temp.:102.09, PWM:59
18:54:32.422 -> Water_Temp.:102.09, PWM:60
18:54:33.182 -> Water_Temp.:102.09, PWM:59
18:54:33.977 -> Water_Temp.:102.09, PWM:59
18:54:34.736 -> Water_Temp.:102.09, PWM:60
18:54:35.534 -> Water_Temp.:102.09, PWM:59
18:54:36.296 -> Water_Temp.:102.09, PWM:59
18:54:37.086 -> Water_Temp.:102.09, PWM:59
18:54:37.848 -> Water_Temp.:102.09, PWM:60
18:54:38.639 -> Water_Temp.:102.09, PWM:60
18:54:39.433 -> Water_Temp.:102.09, PWM:60
18:54:40.194 -> Water_Temp.:102.09, PWM:59
18:54:40.982 -> Water_Temp.:102.09, PWM:59
18:54:41.777 -> Water_Temp.:102.09, PWM:59
18:54:42.536 -> Water_Temp.:102.09, PWM:60
18:54:43.331 -> Water_Temp.:102.09, PWM:59
18:54:44.092 -> Water_Temp.:102.09, PWM:59
18:54:44.888 -> Water_Temp.:102.09, PWM:59
18:54:45.649 -> Water_Temp.:102.09, PWM:59
18:54:46.446 -> Water_Temp.:102.09, PWM:60
18:54:47.235 -> Water_Temp.:102.09, PWM:59
18:54:47.993 -> Water_Temp.:102.09, PWM:60
18:54:48.784 -> Water_Temp.:102.09, PWM:59
18:54:49.546 -> Water_Temp.:102.09, PWM:59
18:54:50.341 -> Water_Temp.:102.09, PWM:60
18:54:51.134 -> Water_Temp.:102.09, PWM:60
18:54:51.896 -> Water_Temp.:102.09, PWM:59
18:54:52.694 -> Water_Temp.:102.09, PWM:60
18:54:53.454 -> Water_Temp.:102.09, PWM:60
18:54:55.011 -> Water_Temp.:102.09, PWM:59
18:54:55.777 -> Water_Temp.:102.09, PWM:59
18:54:56.569 -> Water_Temp.:102.09, PWM:60
18:54:57.330 -> Water_Temp.:102.20, PWM:59
18:54:58.121 -> Water_Temp.:102.20, PWM:60
18:54:58.879 -> Water_Temp.:102.20, PWM:60
18:54:59.674 -> Water_Temp.:102.20, PWM:60
18:55:01.229 -> Water_Temp.:102.20, PWM:60
18:55:01.991 -> Water_Temp.:102.20, PWM:60
18:55:02.786 -> Water_Temp.:102.20, PWM:60
18:55:03.548 -> Water_Temp.:102.20, PWM:59
18:55:04.344 -> Water_Temp.:102.20, PWM:59
18:55:05.106 -> Water_Temp.:102.20, PWM:60
18:55:05.902 -> Water_Temp.:102.20, PWM:60
18:55:08.220 -> Water_Temp.:102.20, PWM:59
18:55:09.015 -> Water_Temp.:102.20, PWM:59
18:55:09.778 -> Water_Temp.:102.20, PWM:60
18:55:10.571 -> Water_Temp.:102.20, PWM:60
18:55:11.336 -> Water_Temp.:102.20, PWM:60
18:55:12.101 -> Water_Temp.:102.20, PWM:60
18:55:12.893 -> Water_Temp.:102.20, PWM:59
18:55:13.682 -> Water_Temp.:102.20, PWM:59
18:55:14.443 -> Water_Temp.:102.20, PWM:59
18:55:16.000 -> Water_Temp.:102.20, PWM:60
18:55:16.797 -> Water_Temp.:102.20, PWM:59
18:55:18.355 -> Water_Temp.:102.31, PWM:60
18:55:19.121 -> Water_Temp.:102.20, PWM:59
18:55:19.882 -> Water_Temp.:102.31, PWM:59
18:55:20.675 -> Water_Temp.:102.20, PWM:59
18:55:21.467 -> Water_Temp.:102.20, PWM:60
18:55:22.227 -> Water_Temp.:102.31, PWM:59
18:55:23.022 -> Water_Temp.:102.31, PWM:60
18:55:23.815 -> Water_Temp.:102.31, PWM:60
18:55:24.574 -> Water_Temp.:102.31, PWM:60
18:55:25.375 -> Water_Temp.:102.31, PWM:59
18:55:26.137 -> Water_Temp.:102.31, PWM:60
18:55:26.934 -> Water_Temp.:102.31, PWM:60
18:55:27.695 -> Water_Temp.:102.31, PWM:60
18:55:28.491 -> Water_Temp.:102.31, PWM:60
18:55:29.250 -> Water_Temp.:102.31, PWM:60
18:55:30.044 -> Water_Temp.:102.31, PWM:60
18:55:30.833 -> Water_Temp.:102.31, PWM:60
18:55:31.595 -> Water_Temp.:102.31, PWM:60
18:55:32.387 -> Water_Temp.:102.31, PWM:59
18:55:33.152 -> Water_Temp.:102.31, PWM:60
18:55:33.946 -> Water_Temp.:102.31, PWM:60
18:55:34.706 -> Water_Temp.:102.31, PWM:60
18:55:35.504 -> Water_Temp.:102.31, PWM:59
18:55:36.298 -> Water_Temp.:102.31, PWM:60
18:55:37.060 -> Water_Temp.:102.31, PWM:59
18:55:37.854 -> Water_Temp.:102.31, PWM:60
18:55:38.610 -> Water_Temp.:102.31, PWM:60
18:55:39.405 -> Water_Temp.:102.31, PWM:60
18:55:40.165 -> Water_Temp.:102.31, PWM:60
18:55:40.958 -> Water_Temp.:102.31, PWM:59
18:55:41.756 -> Water_Temp.:102.31, PWM:59
18:55:42.516 -> Water_Temp.:102.31, PWM:60
18:55:43.310 -> Water_Temp.:102.31, PWM:60
18:55:44.069 -> Water_Temp.:102.42, PWM:60
18:55:44.863 -> Water_Temp.:102.42, PWM:60
18:55:45.626 -> Water_Temp.:102.42, PWM:60
18:55:46.420 -> Water_Temp.:102.42, PWM:60
18:55:47.214 -> Water_Temp.:102.42, PWM:59
18:55:48.766 -> Water_Temp.:102.42, PWM:60
18:55:49.531 -> Water_Temp.:102.42, PWM:60
18:55:50.323 -> Water_Temp.:102.42, PWM:60
18:55:51.084 -> Water_Temp.:102.42, PWM:59
18:55:51.876 -> Water_Temp.:102.42, PWM:60
18:55:54.225 -> Water_Temp.:102.42, PWM:60
18:55:54.989 -> Water_Temp.:102.42, PWM:59
18:55:55.786 -> Water_Temp.:102.42, PWM:59
18:55:56.548 -> Water_Temp.:102.42, PWM:59
18:55:57.343 -> Water_Temp.:102.42, PWM:59
18:55:58.103 -> Water_Temp.:102.42, PWM:60
18:55:58.894 -> Water_Temp.:102.54, PWM:59
18:55:59.660 -> Water_Temp.:102.54, PWM:59
18:56:00.454 -> Water_Temp.:102.54, PWM:59
18:56:01.247 -> Water_Temp.:102.54, PWM:60
18:56:02.009 -> Water_Temp.:102.54, PWM:60
18:56:02.804 -> Water_Temp.:102.54, PWM:59
18:56:03.565 -> Water_Temp.:102.54, PWM:59
18:56:04.361 -> Water_Temp.:102.54, PWM:60
18:56:05.122 -> Water_Temp.:102.54, PWM:60
18:56:05.918 -> Water_Temp.:102.54, PWM:59
18:56:06.712 -> Water_Temp.:102.54, PWM:60
18:56:07.468 -> Water_Temp.:102.54, PWM:59
18:56:08.256 -> Water_Temp.:102.54, PWM:59
18:56:09.048 -> Water_Temp.:102.54, PWM:60
18:56:09.807 -> Water_Temp.:102.54, PWM:60
18:56:10.603 -> Water_Temp.:102.54, PWM:60
18:56:11.362 -> Water_Temp.:102.54, PWM:59
18:56:12.159 -> Water_Temp.:102.54, PWM:60
18:56:12.952 -> Water_Temp.:102.54, PWM:60
18:56:13.711 -> Water_Temp.:102.54, PWM:60
18:56:14.503 -> Water_Temp.:102.54, PWM:60
18:56:15.293 -> Water_Temp.:102.54, PWM:59
18:56:16.818 -> Water_Temp.:102.65, PWM:60
18:56:17.580 -> Water_Temp.:102.65, PWM:60
18:56:18.376 -> Water_Temp.:102.65, PWM:60
18:56:19.138 -> Water_Temp.:102.65, PWM:59
18:56:19.938 -> Water_Temp.:102.65, PWM:60
18:56:20.704 -> Water_Temp.:102.65, PWM:60
18:56:21.497 -> Water_Temp.:102.65, PWM:59
18:56:22.258 -> Water_Temp.:102.65, PWM:60
18:56:23.845 -> Water_Temp.:102.65, PWM:60
18:56:24.600 -> Water_Temp.:102.65, PWM:60
18:56:25.394 -> Water_Temp.:102.65, PWM:59
18:56:26.159 -> Water_Temp.:102.65, PWM:60
18:56:26.954 -> Water_Temp.:102.65, PWM:60
18:56:27.747 -> Water_Temp.:102.65, PWM:59
18:56:28.505 -> Water_Temp.:102.65, PWM:59
18:56:29.303 -> Water_Temp.:102.65, PWM:60
18:56:30.066 -> Water_Temp.:102.65, PWM:59
18:56:30.861 -> Water_Temp.:102.65, PWM:59
18:56:31.623 -> Water_Temp.:102.65, PWM:59
18:56:32.418 -> Water_Temp.:102.65, PWM:60
18:56:33.178 -> Water_Temp.:102.65, PWM:60
18:56:33.972 -> Water_Temp.:102.65, PWM:60
18:56:34.764 -> Water_Temp.:102.65, PWM:59
18:56:35.520 -> Water_Temp.:102.65, PWM:60
18:56:36.315 -> Water_Temp.:102.65, PWM:60
18:56:37.078 -> Water_Temp.:102.65, PWM:59
18:56:38.625 -> Water_Temp.:102.65, PWM:60
18:56:39.419 -> Water_Temp.:102.65, PWM:59
18:56:40.181 -> Water_Temp.:102.65, PWM:60
18:56:40.973 -> Water_Temp.:102.65, PWM:60
18:56:41.736 -> Water_Temp.:102.65, PWM:60
18:56:42.528 -> Water_Temp.:102.65, PWM:60
18:56:43.325 -> Water_Temp.:102.65, PWM:59
18:56:44.084 -> Water_Temp.:102.65, PWM:60
18:56:44.877 -> Water_Temp.:102.65, PWM:59
18:56:45.639 -> Water_Temp.:102.65, PWM:59
18:56:46.431 -> Water_Temp.:102.76, PWM:59
18:56:47.227 -> Water_Temp.:102.76, PWM:60
18:56:47.984 -> Water_Temp.:102.76, PWM:59
18:56:48.780 -> Water_Temp.:102.76, PWM:60
18:56:49.542 -> Water_Temp.:102.76, PWM:60
18:56:50.338 -> Water_Temp.:102.76, PWM:60
18:56:51.099 -> Water_Temp.:102.76, PWM:59
18:56:51.893 -> Water_Temp.:102.76, PWM:60
18:56:52.685 -> Water_Temp.:102.76, PWM:60
18:56:53.445 -> Water_Temp.:102.76, PWM:60
18:56:54.239 -> Water_Temp.:102.76, PWM:60
18:56:54.996 -> Water_Temp.:102.76, PWM:59
18:56:56.554 -> Water_Temp.:102.76, PWM:60
18:56:57.350 -> Water_Temp.:102.76, PWM:60
18:56:58.104 -> Water_Temp.:102.87, PWM:60
18:56:58.898 -> Water_Temp.:102.76, PWM:60
18:56:59.663 -> Water_Temp.:102.76, PWM:60
18:57:00.459 -> Water_Temp.:102.76, PWM:60
18:57:01.254 -> Water_Temp.:102.87, PWM:60
18:57:02.016 -> Water_Temp.:102.87, PWM:59
18:57:02.812 -> Water_Temp.:102.76, PWM:60
18:57:03.573 -> Water_Temp.:102.87, PWM:59
18:57:05.124 -> Water_Temp.:102.87, PWM:60
18:57:05.885 -> Water_Temp.:102.87, PWM:59
18:57:06.679 -> Water_Temp.:102.87, PWM:59
18:57:07.474 -> Water_Temp.:102.87, PWM:60
18:57:08.232 -> Water_Temp.:102.87, PWM:60
18:57:09.024 -> Water_Temp.:102.87, PWM:59
18:57:09.820 -> Water_Temp.:102.87, PWM:60
18:57:10.580 -> Water_Temp.:102.87, PWM:60
18:57:11.376 -> Water_Temp.:102.87, PWM:59
18:57:12.135 -> Water_Temp.:102.87, PWM:60
18:57:12.930 -> Water_Temp.:102.87, PWM:59
18:57:13.691 -> Water_Temp.:102.87, PWM:59
18:57:14.486 -> Water_Temp.:102.87, PWM:60
18:57:15.248 -> Water_Temp.:102.87, PWM:60
18:57:16.047 -> Water_Temp.:102.87, PWM:59
18:57:16.808 -> Water_Temp.:102.87, PWM:60
18:57:17.602 -> Water_Temp.:102.87, PWM:59
18:57:18.390 -> Water_Temp.:102.87, PWM:60
18:57:19.154 -> Water_Temp.:102.87, PWM:60
18:57:19.952 -> Water_Temp.:102.87, PWM:60
18:57:20.709 -> Water_Temp.:102.87, PWM:60
18:57:21.502 -> Water_Temp.:102.87, PWM:60
18:57:22.295 -> Water_Temp.:102.87, PWM:60
18:57:23.056 -> Water_Temp.:102.99, PWM:60
18:57:23.850 -> Water_Temp.:102.99, PWM:59
18:57:24.608 -> Water_Temp.:102.99, PWM:60
18:57:25.404 -> Water_Temp.:102.99, PWM:60
18:57:26.197 -> Water_Temp.:102.99, PWM:60
18:57:26.961 -> Water_Temp.:102.99, PWM:60
18:57:27.753 -> Water_Temp.:102.99, PWM:59
18:57:28.514 -> Water_Temp.:102.99, PWM:59
18:57:29.308 -> Water_Temp.:102.99, PWM:59
18:57:30.066 -> Water_Temp.:102.99, PWM:60
18:57:30.864 -> Water_Temp.:102.99, PWM:59
18:57:31.659 -> Water_Temp.:102.99, PWM:60
18:57:34.700 -> Water_Temp.:102.99, PWM:60
18:57:35.495 -> Water_Temp.:102.99, PWM:60
18:57:36.258 -> Water_Temp.:103.10, PWM:59
18:57:37.049 -> Water_Temp.:103.10, PWM:60
18:57:37.846 -> Water_Temp.:103.10, PWM:59
18:57:38.607 -> Water_Temp.:103.10, PWM:60
18:57:39.404 -> Water_Temp.:103.10, PWM:60
18:57:40.165 -> Water_Temp.:103.10, PWM:60
18:57:40.959 -> Water_Temp.:103.10, PWM:60
18:57:41.716 -> Water_Temp.:103.10, PWM:59
18:57:42.509 -> Water_Temp.:103.10, PWM:59
18:57:43.304 -> Water_Temp.:103.10, PWM:60
18:57:44.063 -> Water_Temp.:103.10, PWM:60
18:57:44.855 -> Water_Temp.:103.10, PWM:60
18:57:46.380 -> Water_Temp.:103.10, PWM:59
18:57:47.176 -> Water_Temp.:103.10, PWM:59
18:57:47.936 -> Water_Temp.:103.10, PWM:60
18:57:48.726 -> Water_Temp.:103.10, PWM:60
18:57:49.521 -> Water_Temp.:103.10, PWM:59
18:57:51.042 -> Water_Temp.:103.10, PWM:60
18:57:51.836 -> Water_Temp.:103.10, PWM:60
18:57:52.597 -> Water_Temp.:103.10, PWM:59
18:57:53.385 -> Water_Temp.:103.10, PWM:60
18:57:54.934 -> Water_Temp.:103.10, PWM:60
18:57:55.729 -> Water_Temp.:103.10, PWM:60
18:57:56.490 -> Water_Temp.:103.10, PWM:59
18:57:57.286 -> Water_Temp.:103.10, PWM:60
18:57:58.044 -> Water_Temp.:103.21, PWM:60
18:57:58.840 -> Water_Temp.:103.21, PWM:60
18:57:59.602 -> Water_Temp.:103.10, PWM:59
18:58:00.396 -> Water_Temp.:103.21, PWM:60
18:58:01.192 -> Water_Temp.:103.21, PWM:59
18:58:01.954 -> Water_Temp.:103.21, PWM:60
18:58:02.747 -> Water_Temp.:103.21, PWM:60
18:58:03.508 -> Water_Temp.:103.21, PWM:60
18:58:04.302 -> Water_Temp.:103.21, PWM:60
18:58:05.063 -> Water_Temp.:103.21, PWM:59
18:58:05.861 -> Water_Temp.:103.21, PWM:60
18:58:06.655 -> Water_Temp.:103.21, PWM:59
18:58:07.412 -> Water_Temp.:103.21, PWM:60
18:58:08.208 -> Water_Temp.:103.21, PWM:60
18:58:08.972 -> Water_Temp.:103.21, PWM:59
18:58:09.769 -> Water_Temp.:103.21, PWM:60
18:58:10.531 -> Water_Temp.:103.21, PWM:60
18:58:11.327 -> Water_Temp.:103.21, PWM:60
18:58:12.090 -> Water_Temp.:103.21, PWM:59
18:58:12.887 -> Water_Temp.:103.21, PWM:59
18:58:13.649 -> Water_Temp.:103.21, PWM:59
18:58:14.444 -> Water_Temp.:103.21, PWM:60
18:58:15.234 -> Water_Temp.:103.21, PWM:60
18:58:15.993 -> Water_Temp.:103.21, PWM:60
18:58:16.785 -> Water_Temp.:103.21, PWM:59
18:58:17.542 -> Water_Temp.:103.21, PWM:59
18:58:18.341 -> Water_Temp.:103.32, PWM:60
18:58:19.864 -> Water_Temp.:103.32, PWM:60
18:58:22.215 -> Water_Temp.:103.32, PWM:60
18:58:23.013 -> Water_Temp.:103.21, PWM:59
18:58:23.773 -> Water_Temp.:103.32, PWM:60
18:58:24.566 -> Water_Temp.:103.32, PWM:60
18:58:25.325 -> Water_Temp.:103.32, PWM:60
18:58:26.118 -> Water_Temp.:103.32, PWM:59
18:58:26.879 -> Water_Temp.:103.32, PWM:60
18:58:27.669 -> Water_Temp.:103.32, PWM:59
18:58:28.463 -> Water_Temp.:103.32, PWM:60
18:58:29.224 -> Water_Temp.:103.32, PWM:60
18:58:30.016 -> Water_Temp.:103.32, PWM:60
18:58:30.808 -> Water_Temp.:103.32, PWM:60
18:58:31.569 -> Water_Temp.:103.32, PWM:60
18:58:32.364 -> Water_Temp.:103.32, PWM:60
18:58:33.154 -> Water_Temp.:103.32, PWM:60
18:58:33.913 -> Water_Temp.:103.32, PWM:60
18:58:35.473 -> Water_Temp.:103.32, PWM:59
18:58:36.266 -> Water_Temp.:103.32, PWM:60
18:58:37.017 -> Water_Temp.:103.32, PWM:59
18:58:37.810 -> Water_Temp.:103.32, PWM:60
18:58:38.600 -> Water_Temp.:103.32, PWM:60
18:58:39.362 -> Water_Temp.:103.32, PWM:59
18:58:40.160 -> Water_Temp.:103.32, PWM:60
18:58:40.920 -> Water_Temp.:103.32, PWM:59
18:58:41.713 -> Water_Temp.:103.32, PWM:60
18:58:42.507 -> Water_Temp.:103.32, PWM:59
18:58:43.265 -> Water_Temp.:103.32, PWM:60
18:58:44.062 -> Water_Temp.:103.32, PWM:59
18:58:44.821 -> Water_Temp.:103.32, PWM:60
18:58:46.377 -> Water_Temp.:103.44, PWM:60
18:58:47.172 -> Water_Temp.:103.44, PWM:60
18:58:47.967 -> Water_Temp.:103.44, PWM:60
18:58:48.730 -> Water_Temp.:103.44, PWM:60
18:58:49.530 -> Water_Temp.:103.44, PWM:59
18:58:50.285 -> Water_Temp.:103.44, PWM:60
18:58:51.077 -> Water_Temp.:103.44, PWM:60
18:58:51.839 -> Water_Temp.:103.44, PWM:60
18:58:52.629 -> Water_Temp.:103.44, PWM:60
18:58:53.425 -> Water_Temp.:103.44, PWM:60
18:58:54.188 -> Water_Temp.:103.44, PWM:59
18:58:54.982 -> Water_Temp.:103.44, PWM:59
18:58:55.743 -> Water_Temp.:103.44, PWM:60
18:58:56.541 -> Water_Temp.:103.44, PWM:60
18:58:58.093 -> Water_Temp.:103.44, PWM:60
18:58:58.885 -> Water_Temp.:103.44, PWM:60
18:59:00.408 -> Water_Temp.:103.44, PWM:59
18:59:01.203 -> Water_Temp.:103.44, PWM:60
18:59:01.967 -> Water_Temp.:103.44, PWM:60
18:59:03.525 -> Water_Temp.:103.44, PWM:60
18:59:04.285 -> Water_Temp.:103.44, PWM:60
18:59:05.077 -> Water_Temp.:103.55, PWM:60
18:59:05.839 -> Water_Temp.:103.44, PWM:59
18:59:06.630 -> Water_Temp.:103.55, PWM:60
18:59:07.424 -> Water_Temp.:103.55, PWM:60
18:59:08.183 -> Water_Temp.:103.55, PWM:60
18:59:08.981 -> Water_Temp.:103.55, PWM:59
18:59:09.746 -> Water_Temp.:103.55, PWM:59
18:59:10.541 -> Water_Temp.:105.12, PWM:60
18:59:11.298 -> Water_Temp.:107.15, PWM:60
18:59:12.093 -> Water_Temp.:104.79, PWM:60
18:59:12.888 -> Water_Temp.:104.34, PWM:60
18:59:13.650 -> Water_Temp.:104.11, PWM:59
18:59:14.444 -> Water_Temp.:104.00, PWM:60
18:59:15.205 -> Water_Temp.:103.89, PWM:60
18:59:16.005 -> Water_Temp.:103.89, PWM:60
18:59:16.765 -> Water_Temp.:103.77, PWM:60
18:59:18.322 -> Water_Temp.:103.77, PWM:59
18:59:20.660 -> Water_Temp.:103.66, PWM:59
18:59:21.457 -> Water_Temp.:103.66, PWM:60
18:59:22.215 -> Water_Temp.:103.66, PWM:60
18:59:23.009 -> Water_Temp.:103.66, PWM:60
18:59:23.800 -> Water_Temp.:103.66, PWM:60
18:59:24.560 -> Water_Temp.:103.66, PWM:60
18:59:25.358 -> Water_Temp.:103.66, PWM:60
18:59:26.120 -> Water_Temp.:103.66, PWM:60
18:59:26.917 -> Water_Temp.:103.66, PWM:60
18:59:27.680 -> Water_Temp.:103.66, PWM:60
18:59:29.238 -> Water_Temp.:103.66, PWM:60
18:59:30.034 -> Water_Temp.:103.66, PWM:59
18:59:31.590 -> Water_Temp.:103.66, PWM:60
18:59:32.384 -> Water_Temp.:103.55, PWM:60
18:59:33.146 -> Water_Temp.:103.55, PWM:60
18:59:33.938 -> Water_Temp.:103.66, PWM:60
18:59:34.698 -> Water_Temp.:103.66, PWM:59
18:59:36.258 -> Water_Temp.:103.66, PWM:60
18:59:37.019 -> Water_Temp.:103.66, PWM:60
18:59:37.817 -> Water_Temp.:103.66, PWM:60
18:59:38.579 -> Water_Temp.:103.66, PWM:60
18:59:39.377 -> Water_Temp.:103.66, PWM:59
18:59:40.137 -> Water_Temp.:103.66, PWM:60
18:59:40.933 -> Water_Temp.:103.66, PWM:60
18:59:41.700 -> Water_Temp.:103.66, PWM:60
18:59:42.495 -> Water_Temp.:103.66, PWM:60
18:59:43.285 -> Water_Temp.:103.66, PWM:60
18:59:44.808 -> Water_Temp.:103.66, PWM:60
18:59:45.604 -> Water_Temp.:103.66, PWM:60
18:59:46.361 -> Water_Temp.:103.66, PWM:60
18:59:47.155 -> Water_Temp.:103.66, PWM:59
18:59:47.951 -> Water_Temp.:103.66, PWM:60
18:59:48.710 -> Water_Temp.:103.66, PWM:60
18:59:49.506 -> Water_Temp.:103.66, PWM:60
18:59:50.266 -> Water_Temp.:103.66, PWM:60
18:59:51.060 -> Water_Temp.:103.66, PWM:59
18:59:51.851 -> Water_Temp.:103.66, PWM:60
18:59:52.615 -> Water_Temp.:103.66, PWM:60
18:59:53.411 -> Water_Temp.:103.66, PWM:60
18:59:54.172 -> Water_Temp.:103.66, PWM:60
18:59:54.965 -> Water_Temp.:103.66, PWM:60
18:59:55.726 -> Water_Temp.:103.66, PWM:60
18:59:58.046 -> Water_Temp.:103.66, PWM:60
18:59:58.804 -> Water_Temp.:103.66, PWM:60
18:59:59.601 -> Water_Temp.:103.66, PWM:60
19:00:00.395 -> Water_Temp.:103.66, PWM:59
19:00:01.151 -> Water_Temp.:103.66, PWM:60
19:00:01.944 -> Water_Temp.:103.66, PWM:60
19:00:02.738 -> Water_Temp.:103.66, PWM:60
19:00:03.499 -> Water_Temp.:103.66, PWM:59
19:00:04.294 -> Water_Temp.:103.77, PWM:60
19:00:05.051 -> Water_Temp.:103.66, PWM:60
19:00:06.604 -> Water_Temp.:103.77, PWM:59
19:00:07.398 -> Water_Temp.:103.66, PWM:60
19:00:08.160 -> Water_Temp.:103.77, PWM:60
19:00:08.957 -> Water_Temp.:103.66, PWM:59
19:00:09.754 -> Water_Temp.:103.77, PWM:60
19:00:10.512 -> Water_Temp.:103.77, PWM:59
19:00:11.307 -> Water_Temp.:103.66, PWM:60
19:00:12.069 -> Water_Temp.:103.66, PWM:60
19:00:12.864 -> Water_Temp.:103.66, PWM:60
19:00:13.620 -> Water_Temp.:103.66, PWM:60
19:00:14.411 -> Water_Temp.:103.66, PWM:60
19:00:15.206 -> Water_Temp.:103.66, PWM:60
19:00:15.969 -> Water_Temp.:103.66, PWM:60
19:00:16.763 -> Water_Temp.:103.66, PWM:59
19:00:17.556 -> Water_Temp.:103.66, PWM:60
19:00:18.314 -> Water_Temp.:103.66, PWM:60
19:00:19.104 -> Water_Temp.:103.66, PWM:60
19:00:19.866 -> Water_Temp.:103.66, PWM:60
19:00:20.660 -> Water_Temp.:103.77, PWM:60
19:00:21.452 -> Water_Temp.:103.77, PWM:60
19:00:22.211 -> Water_Temp.:103.77, PWM:60
19:00:23.006 -> Water_Temp.:103.77, PWM:60
19:00:23.761 -> Water_Temp.:103.77, PWM:59
19:00:24.556 -> Water_Temp.:103.77, PWM:60
19:00:26.111 -> Water_Temp.:103.77, PWM:60
19:00:27.635 -> Water_Temp.:103.77, PWM:59
19:00:28.396 -> Water_Temp.:103.77, PWM:60
19:00:29.951 -> Water_Temp.:103.77, PWM:60
19:00:31.469 -> Water_Temp.:103.77, PWM:60
19:00:32.258 -> Water_Temp.:103.77, PWM:60
19:00:33.053 -> Water_Temp.:103.77, PWM:60
19:00:33.812 -> Water_Temp.:103.77, PWM:60
19:00:34.609 -> Water_Temp.:103.77, PWM:60
19:00:35.402 -> Water_Temp.:103.77, PWM:60
19:00:36.162 -> Water_Temp.:103.77, PWM:60
19:00:36.961 -> Water_Temp.:103.77, PWM:60
19:00:37.721 -> Water_Temp.:103.77, PWM:60
19:00:38.517 -> Water_Temp.:103.77, PWM:59
19:00:39.277 -> Water_Temp.:103.77, PWM:59
19:00:40.070 -> Water_Temp.:103.77, PWM:60
19:00:40.833 -> Water_Temp.:103.77, PWM:60
19:00:41.631 -> Water_Temp.:103.77, PWM:60
19:00:42.395 -> Water_Temp.:103.77, PWM:60
19:00:43.190 -> Water_Temp.:103.89, PWM:59
19:00:43.949 -> Water_Temp.:103.77, PWM:60
19:00:44.744 -> Water_Temp.:103.89, PWM:60
19:00:45.542 -> Water_Temp.:103.89, PWM:59
19:00:46.303 -> Water_Temp.:103.89, PWM:60
19:00:47.096 -> Water_Temp.:103.89, PWM:60
19:00:47.856 -> Water_Temp.:103.89, PWM:60
19:00:48.651 -> Water_Temp.:103.89, PWM:60
19:00:49.413 -> Water_Temp.:103.89, PWM:60
19:00:50.207 -> Water_Temp.:103.89, PWM:60
19:00:50.969 -> Water_Temp.:103.89, PWM:60
19:00:52.519 -> Water_Temp.:103.89, PWM:59
19:00:53.311 -> Water_Temp.:103.89, PWM:60
19:00:54.070 -> Water_Temp.:103.89, PWM:59
19:00:54.865 -> Water_Temp.:103.89, PWM:60
19:00:55.629 -> Water_Temp.:103.89, PWM:59
19:00:56.422 -> Water_Temp.:103.89, PWM:60
19:00:57.218 -> Water_Temp.:103.89, PWM:60
19:00:57.979 -> Water_Temp.:103.89, PWM:60
19:00:58.772 -> Water_Temp.:103.89, PWM:59
19:00:59.531 -> Water_Temp.:103.89, PWM:59
19:01:00.330 -> Water_Temp.:103.89, PWM:60
19:01:01.123 -> Water_Temp.:103.89, PWM:59
19:01:01.886 -> Water_Temp.:103.89, PWM:60
19:01:02.679 -> Water_Temp.:103.89, PWM:59
19:01:03.434 -> Water_Temp.:103.89, PWM:59
19:01:04.229 -> Water_Temp.:103.89, PWM:60
19:01:04.989 -> Water_Temp.:103.89, PWM:60
19:01:05.785 -> Water_Temp.:103.89, PWM:60
19:01:07.336 -> Water_Temp.:103.89, PWM:60
19:01:08.128 -> Water_Temp.:103.89, PWM:59
19:01:08.886 -> Water_Temp.:104.00, PWM:59
19:01:09.680 -> Water_Temp.:103.89, PWM:60
19:01:10.438 -> Water_Temp.:104.00, PWM:60
19:01:11.234 -> Water_Temp.:104.00, PWM:60
19:01:12.031 -> Water_Temp.:104.00, PWM:59
19:01:12.793 -> Water_Temp.:104.00, PWM:60
19:01:13.586 -> Water_Temp.:104.00, PWM:59
19:01:14.344 -> Water_Temp.:104.00, PWM:60
19:01:15.138 -> Water_Temp.:104.00, PWM:60
19:01:15.901 -> Water_Temp.:104.00, PWM:60
19:01:16.698 -> Water_Temp.:104.00, PWM:59
19:01:17.461 -> Water_Temp.:104.00, PWM:60
19:01:18.258 -> Water_Temp.:104.00, PWM:60
19:01:19.025 -> Water_Temp.:104.00, PWM:59
19:01:19.825 -> Water_Temp.:104.00, PWM:60
19:01:20.583 -> Water_Temp.:104.00, PWM:59
19:01:21.376 -> Water_Temp.:104.00, PWM:60
19:01:22.138 -> Water_Temp.:104.00, PWM:60
19:01:22.930 -> Water_Temp.:104.00, PWM:60
19:01:23.722 -> Water_Temp.:104.00, PWM:60
19:01:24.481 -> Water_Temp.:104.00, PWM:60
19:01:25.273 -> Water_Temp.:104.00, PWM:60
19:01:26.827 -> Water_Temp.:104.00, PWM:60
19:01:27.622 -> Water_Temp.:104.00, PWM:60
19:01:28.384 -> Water_Temp.:104.00, PWM:60
19:01:29.182 -> Water_Temp.:104.00, PWM:60
19:01:29.946 -> Water_Temp.:104.00, PWM:60
19:01:30.739 -> Water_Temp.:104.00, PWM:59
19:01:31.496 -> Water_Temp.:104.00, PWM:60
19:01:32.290 -> Water_Temp.:104.00, PWM:60
19:01:33.816 -> Water_Temp.:104.00, PWM:59
19:01:34.611 -> Water_Temp.:104.00, PWM:59
19:01:35.370 -> Water_Temp.:104.00, PWM:60
19:01:36.169 -> Water_Temp.:104.00, PWM:60
19:01:36.929 -> Water_Temp.:104.11, PWM:60
19:01:37.720 -> Water_Temp.:104.11, PWM:60
19:01:38.514 -> Water_Temp.:104.11, PWM:60
19:01:39.275 -> Water_Temp.:104.11, PWM:60
19:01:40.067 -> Water_Temp.:104.11, PWM:60
19:01:40.861 -> Water_Temp.:104.11, PWM:60
19:01:41.622 -> Water_Temp.:104.11, PWM:60
19:01:42.414 -> Water_Temp.:104.11, PWM:60
19:01:43.175 -> Water_Temp.:104.11, PWM:60
19:01:43.970 -> Water_Temp.:104.11, PWM:60
19:01:44.730 -> Water_Temp.:104.11, PWM:60
19:01:45.526 -> Water_Temp.:104.11, PWM:60
19:01:46.291 -> Water_Temp.:104.11, PWM:59
19:01:47.085 -> Water_Temp.:104.11, PWM:60
19:01:47.880 -> Water_Temp.:104.11, PWM:60
19:01:48.639 -> Water_Temp.:104.11, PWM:60
19:01:49.434 -> Water_Temp.:104.11, PWM:60
19:01:50.221 -> Water_Temp.:104.11, PWM:60
19:01:50.983 -> Water_Temp.:104.11, PWM:59
19:01:52.533 -> Water_Temp.:104.11, PWM:60
19:01:53.325 -> Water_Temp.:104.11, PWM:59
19:01:54.084 -> Water_Temp.:104.11, PWM:60
19:01:54.878 -> Water_Temp.:104.11, PWM:60
19:01:55.639 -> Water_Temp.:104.11, PWM:59
19:01:56.434 -> Water_Temp.:104.11, PWM:60
19:01:57.227 -> Water_Temp.:104.11, PWM:60
19:01:57.984 -> Water_Temp.:104.11, PWM:59
19:01:58.774 -> Water_Temp.:104.11, PWM:60
19:01:59.570 -> Water_Temp.:104.11, PWM:60
19:02:00.326 -> Water_Temp.:104.11, PWM:60
19:02:01.117 -> Water_Temp.:104.11, PWM:60
19:02:01.877 -> Water_Temp.:104.11, PWM:60
19:02:02.664 -> Water_Temp.:104.11, PWM:60
19:02:04.222 -> Water_Temp.:104.11, PWM:60
19:02:04.983 -> Water_Temp.:104.11, PWM:60
19:02:05.775 -> Water_Temp.:104.11, PWM:60
19:02:06.568 -> Water_Temp.:104.11, PWM:60
19:02:07.328 -> Water_Temp.:104.22, PWM:60
19:02:08.123 -> Water_Temp.:104.22, PWM:60
19:02:08.883 -> Water_Temp.:104.22, PWM:60
19:02:09.680 -> Water_Temp.:104.22, PWM:60
19:02:10.442 -> Water_Temp.:104.22, PWM:60
19:02:11.235 -> Water_Temp.:104.22, PWM:60
19:02:12.033 -> Water_Temp.:104.22, PWM:60
19:02:12.791 -> Water_Temp.:104.22, PWM:60
19:02:13.587 -> Water_Temp.:104.22, PWM:60
19:02:14.349 -> Water_Temp.:104.22, PWM:60
19:02:15.143 -> Water_Temp.:104.22, PWM:60
19:02:15.906 -> Water_Temp.:104.22, PWM:60
19:02:16.701 -> Water_Temp.:104.22, PWM:60
19:02:17.460 -> Water_Temp.:104.22, PWM:60
19:02:18.256 -> Water_Temp.:104.22, PWM:60
19:02:19.015 -> Water_Temp.:104.11, PWM:60
19:02:19.813 -> Water_Temp.:104.22, PWM:60
19:02:20.579 -> Water_Temp.:104.11, PWM:60
19:02:21.374 -> Water_Temp.:104.11, PWM:60
19:02:22.171 -> Water_Temp.:104.11, PWM:60
19:02:22.933 -> Water_Temp.:104.22, PWM:60
19:02:23.728 -> Water_Temp.:104.22, PWM:60
19:02:24.489 -> Water_Temp.:104.22, PWM:60
19:02:25.285 -> Water_Temp.:104.22, PWM:60
19:02:26.047 -> Water_Temp.:104.22, PWM:60
19:02:26.845 -> Water_Temp.:104.22, PWM:60
19:02:27.603 -> Water_Temp.:104.22, PWM:60
19:02:28.396 -> Water_Temp.:104.22, PWM:60
19:02:29.158 -> Water_Temp.:104.22, PWM:60
19:02:30.714 -> Water_Temp.:104.22, PWM:60
19:02:31.474 -> Water_Temp.:104.22, PWM:60
19:02:32.267 -> Water_Temp.:104.22, PWM:60
19:02:33.061 -> Water_Temp.:104.22, PWM:60
19:02:33.821 -> Water_Temp.:104.22, PWM:60
19:02:34.612 -> Water_Temp.:104.34, PWM:60
19:02:35.375 -> Water_Temp.:104.34, PWM:60
19:02:36.165 -> Water_Temp.:104.34, PWM:60
19:02:37.720 -> Water_Temp.:104.34, PWM:60
19:02:38.512 -> Water_Temp.:104.45, PWM:60
19:02:39.275 -> Water_Temp.:104.45, PWM:60
19:02:40.070 -> Water_Temp.:104.45, PWM:60
19:02:40.862 -> Water_Temp.:104.45, PWM:60
19:02:41.622 -> Water_Temp.:104.45, PWM:60
19:02:42.418 -> Water_Temp.:104.45, PWM:60
19:02:43.175 -> Water_Temp.:104.45, PWM:60
19:02:43.971 -> Water_Temp.:104.56, PWM:60
19:02:44.735 -> Water_Temp.:104.45, PWM:60
19:02:45.531 -> Water_Temp.:104.56, PWM:60
19:02:46.291 -> Water_Temp.:104.56, PWM:60
19:02:47.086 -> Water_Temp.:106.14, PWM:60
19:02:47.879 -> Water_Temp.:108.16, PWM:60
19:02:48.636 -> Water_Temp.:105.80, PWM:60
19:02:49.435 -> Water_Temp.:105.35, PWM:60
19:02:50.198 -> Water_Temp.:105.12, PWM:60
19:02:50.992 -> Water_Temp.:105.01, PWM:60
19:02:52.512 -> Water_Temp.:104.90, PWM:60
19:02:53.302 -> Water_Temp.:104.79, PWM:60
19:02:54.094 -> Water_Temp.:104.79, PWM:60
19:02:54.855 -> Water_Temp.:104.79, PWM:60
19:02:55.651 -> Water_Temp.:104.67, PWM:60
19:02:56.412 -> Water_Temp.:104.67, PWM:60
19:02:57.207 -> Water_Temp.:104.67, PWM:60
19:02:58.000 -> Water_Temp.:104.67, PWM:60
19:02:58.759 -> Water_Temp.:104.67, PWM:60
19:03:00.312 -> Water_Temp.:104.67, PWM:60
19:03:01.102 -> Water_Temp.:104.56, PWM:60
19:03:02.656 -> Water_Temp.:104.67, PWM:60
19:03:03.446 -> Water_Temp.:104.56, PWM:60
19:03:04.238 -> Water_Temp.:104.67, PWM:60
19:03:04.998 -> Water_Temp.:104.56, PWM:60
19:03:05.799 -> Water_Temp.:104.67, PWM:60
19:03:06.554 -> Water_Temp.:104.67, PWM:60
19:03:07.345 -> Water_Temp.:104.67, PWM:60
19:03:08.107 -> Water_Temp.:106.25, PWM:60
19:03:08.904 -> Water_Temp.:108.27, PWM:60
19:03:09.668 -> Water_Temp.:105.91, PWM:60
19:03:10.462 -> Water_Temp.:105.46, PWM:60
19:03:11.253 -> Water_Temp.:105.35, PWM:60
19:03:12.017 -> Water_Temp.:105.24, PWM:60
19:03:12.813 -> Water_Temp.:105.12, PWM:60
19:03:13.570 -> Water_Temp.:105.01, PWM:60
19:03:14.363 -> Water_Temp.:105.01, PWM:60
19:03:15.156 -> Water_Temp.:105.01, PWM:60
19:03:15.920 -> Water_Temp.:104.90, PWM:60
19:03:16.713 -> Water_Temp.:104.90, PWM:60
19:03:18.229 -> Water_Temp.:104.90, PWM:60
19:03:19.025 -> Water_Temp.:104.90, PWM:60
19:03:19.817 -> Water_Temp.:104.90, PWM:60
19:03:20.575 -> Water_Temp.:104.90, PWM:60
19:03:21.373 -> Water_Temp.:104.79, PWM:60
19:03:22.131 -> Water_Temp.:104.79, PWM:60
19:03:22.922 -> Water_Temp.:104.90, PWM:60
19:03:23.717 -> Water_Temp.:104.79, PWM:60
19:03:24.476 -> Water_Temp.:104.79, PWM:60
19:03:25.271 -> Water_Temp.:104.79, PWM:60
19:03:26.030 -> Water_Temp.:104.79, PWM:60
19:03:26.822 -> Water_Temp.:104.79, PWM:60
19:03:27.617 -> Water_Temp.:104.79, PWM:60
19:03:28.378 -> Water_Temp.:104.79, PWM:60
19:03:29.174 -> Water_Temp.:104.79, PWM:60
19:03:29.938 -> Water_Temp.:104.79, PWM:60
19:03:30.736 -> Water_Temp.:104.79, PWM:60
19:03:31.492 -> Water_Temp.:104.79, PWM:60
19:03:32.288 -> Water_Temp.:104.79, PWM:60
19:03:33.081 -> Water_Temp.:104.79, PWM:60
19:03:33.841 -> Water_Temp.:104.79, PWM:60
19:03:34.634 -> Water_Temp.:104.79, PWM:60
19:03:35.395 -> Water_Temp.:104.79, PWM:60
19:03:36.193 -> Water_Temp.:104.79, PWM:60
19:03:36.952 -> Water_Temp.:104.79, PWM:60
19:03:37.742 -> Water_Temp.:104.79, PWM:60
19:03:39.291 -> Water_Temp.:104.79, PWM:60
19:03:40.050 -> Water_Temp.:104.79, PWM:60
19:03:40.841 -> Water_Temp.:104.79, PWM:60
19:03:41.636 -> Water_Temp.:104.79, PWM:60
19:03:42.398 -> Water_Temp.:104.79, PWM:60
19:03:43.949 -> Water_Temp.:104.79, PWM:60
19:03:44.708 -> Water_Temp.:104.79, PWM:60
19:03:45.506 -> Water_Temp.:104.79, PWM:60
19:03:46.302 -> Water_Temp.:104.79, PWM:60
19:03:47.061 -> Water_Temp.:104.79, PWM:60
19:03:47.854 -> Water_Temp.:104.79, PWM:60
19:03:48.614 -> Water_Temp.:104.79, PWM:60
19:03:49.408 -> Water_Temp.:104.79, PWM:60
19:03:50.202 -> Water_Temp.:104.79, PWM:60
19:03:50.960 -> Water_Temp.:104.79, PWM:60
19:03:51.752 -> Water_Temp.:104.79, PWM:60
19:03:52.514 -> Water_Temp.:104.79, PWM:60
19:03:53.311 -> Water_Temp.:104.79, PWM:60
19:03:54.103 -> Water_Temp.:104.79, PWM:60
19:03:54.861 -> Water_Temp.:104.79, PWM:60
19:03:55.657 -> Water_Temp.:104.79, PWM:60
19:03:56.418 -> Water_Temp.:104.79, PWM:60
19:03:57.215 -> Water_Temp.:104.79, PWM:60
19:03:57.976 -> Water_Temp.:104.79, PWM:60
19:03:58.770 -> Water_Temp.:104.79, PWM:60
19:03:59.535 -> Water_Temp.:104.79, PWM:60
19:04:00.328 -> Water_Temp.:104.79, PWM:60
19:04:01.092 -> Water_Temp.:104.67, PWM:60
19:04:01.883 -> Water_Temp.:104.79, PWM:60
19:04:02.675 -> Water_Temp.:104.79, PWM:60
19:04:03.435 -> Water_Temp.:104.79, PWM:60
19:04:04.227 -> Water_Temp.:104.79, PWM:60
19:04:05.021 -> Water_Temp.:104.79, PWM:60
19:04:05.785 -> Water_Temp.:104.79, PWM:60
19:04:06.581 -> Water_Temp.:104.79, PWM:60
19:04:07.339 -> Water_Temp.:104.79, PWM:60
19:04:08.131 -> Water_Temp.:104.79, PWM:60
19:04:08.890 -> Water_Temp.:104.79, PWM:60
19:04:09.685 -> Water_Temp.:104.79, PWM:60
19:04:10.479 -> Water_Temp.:104.79, PWM:60
19:04:11.240 -> Water_Temp.:104.79, PWM:60
19:04:12.036 -> Water_Temp.:104.79, PWM:60
19:04:12.794 -> Water_Temp.:104.79, PWM:60
19:04:13.589 -> Water_Temp.:104.79, PWM:60
19:04:14.350 -> Water_Temp.:104.79, PWM:60
19:04:15.142 -> Water_Temp.:104.79, PWM:60
19:04:16.705 -> Water_Temp.:104.79, PWM:60
19:04:17.464 -> Water_Temp.:104.79, PWM:60
19:04:18.257 -> Water_Temp.:104.79, PWM:60
19:04:19.020 -> Water_Temp.:104.79, PWM:60
19:04:19.820 -> Water_Temp.:104.79, PWM:60
19:04:20.581 -> Water_Temp.:104.79, PWM:60
19:04:21.374 -> Water_Temp.:104.79, PWM:60
19:04:22.137 -> Water_Temp.:104.79, PWM:60
19:04:22.931 -> Water_Temp.:104.79, PWM:60
19:04:24.475 -> Water_Temp.:104.79, PWM:60
19:04:25.274 -> Water_Temp.:104.79, PWM:60
19:04:26.035 -> Water_Temp.:104.79, PWM:60
19:04:26.834 -> Water_Temp.:104.79, PWM:60
19:04:27.586 -> Water_Temp.:104.79, PWM:60
19:04:28.377 -> Water_Temp.:104.79, PWM:60
19:04:29.171 -> Water_Temp.:104.79, PWM:60
19:04:29.934 -> Water_Temp.:104.79, PWM:60
19:04:30.725 -> Water_Temp.:104.79, PWM:60
19:04:31.487 -> Water_Temp.:104.79, PWM:60
19:04:32.276 -> Water_Temp.:104.90, PWM:60
19:04:33.069 -> Water_Temp.:104.90, PWM:60
19:04:33.826 -> Water_Temp.:104.90, PWM:60
19:04:34.619 -> Water_Temp.:104.79, PWM:60
19:04:35.384 -> Water_Temp.:104.90, PWM:60
19:04:36.179 -> Water_Temp.:104.90, PWM:60
19:04:36.971 -> Water_Temp.:104.90, PWM:60
19:04:37.734 -> Water_Temp.:104.90, PWM:60
19:04:39.285 -> Water_Temp.:104.90, PWM:60
19:04:40.078 -> Water_Temp.:104.90, PWM:60
19:04:40.873 -> Water_Temp.:104.90, PWM:60
19:04:41.634 -> Water_Temp.:104.90, PWM:60
19:04:42.424 -> Water_Temp.:104.90, PWM:60
19:04:43.186 -> Water_Temp.:104.90, PWM:60
19:04:43.978 -> Water_Temp.:104.90, PWM:60
19:04:45.539 -> Water_Temp.:104.90, PWM:60
19:04:46.333 -> Water_Temp.:104.90, PWM:60
19:04:47.095 -> Water_Temp.:104.90, PWM:60
19:04:47.890 -> Water_Temp.:104.90, PWM:60
19:04:48.654 -> Water_Temp.:104.90, PWM:60
19:04:49.448 -> Water_Temp.:104.90, PWM:60
19:04:50.207 -> Water_Temp.:104.90, PWM:60
19:04:51.003 -> Water_Temp.:104.90, PWM:60
19:04:51.794 -> Water_Temp.:104.90, PWM:60
19:04:52.552 -> Water_Temp.:104.90, PWM:60
19:04:53.344 -> Water_Temp.:105.01, PWM:60
19:04:54.106 -> Water_Temp.:105.01, PWM:60
19:04:55.666 -> Water_Temp.:105.01, PWM:60
19:04:56.428 -> Water_Temp.:105.01, PWM:60
19:04:57.223 -> Water_Temp.:105.01, PWM:60
19:04:57.983 -> Water_Temp.:105.01, PWM:60
19:04:58.776 -> Water_Temp.:105.01, PWM:60
19:04:59.538 -> Water_Temp.:105.01, PWM:60
19:05:00.329 -> Water_Temp.:104.90, PWM:60
19:05:01.118 -> Water_Temp.:104.90, PWM:60
19:05:01.881 -> Water_Temp.:104.90, PWM:60
19:05:02.677 -> Water_Temp.:104.90, PWM:60
19:05:03.468 -> Water_Temp.:104.90, PWM:60
19:05:04.225 -> Water_Temp.:104.90, PWM:60
19:05:05.019 -> Water_Temp.:104.90, PWM:60
19:05:05.780 -> Water_Temp.:104.90, PWM:60
19:05:06.574 -> Water_Temp.:104.90, PWM:60
19:05:07.367 -> Water_Temp.:104.90, PWM:60
19:05:08.124 -> Water_Temp.:106.47, PWM:60
19:05:08.918 -> Water_Temp.:108.50, PWM:60
19:05:09.681 -> Water_Temp.:106.14, PWM:60
19:05:10.475 -> Water_Temp.:105.80, PWM:60
19:05:12.033 -> Water_Temp.:105.46, PWM:60
19:05:12.797 -> Water_Temp.:105.35, PWM:60
19:05:13.590 -> Water_Temp.:105.35, PWM:60
19:05:14.348 -> Water_Temp.:105.24, PWM:60
19:05:15.144 -> Water_Temp.:105.24, PWM:60
19:05:16.694 -> Water_Temp.:105.24, PWM:60
19:05:17.490 -> Water_Temp.:105.12, PWM:60
19:05:18.252 -> Water_Temp.:105.12, PWM:60
19:05:19.045 -> Water_Temp.:105.12, PWM:60
19:05:19.805 -> Water_Temp.:105.12, PWM:60
19:05:20.602 -> Water_Temp.:105.12, PWM:60
19:05:21.363 -> Water_Temp.:105.01, PWM:60
19:05:22.159 -> Water_Temp.:105.01, PWM:60
19:05:22.958 -> Water_Temp.:105.01, PWM:60
19:05:23.723 -> Water_Temp.:105.01, PWM:60
19:05:24.481 -> Water_Temp.:105.01, PWM:60
19:05:25.280 -> Water_Temp.:105.01, PWM:60
19:05:26.072 -> Water_Temp.:105.01, PWM:60
19:05:26.835 -> Water_Temp.:105.01, PWM:60
19:05:27.627 -> Water_Temp.:105.01, PWM:60
19:05:28.385 -> Water_Temp.:105.01, PWM:60
19:05:29.179 -> Water_Temp.:105.01, PWM:60
19:05:29.940 -> Water_Temp.:105.01, PWM:60
19:05:30.735 -> Water_Temp.:105.01, PWM:60
19:05:31.530 -> Water_Temp.:105.01, PWM:60
19:05:32.289 -> Water_Temp.:105.01, PWM:60
19:05:33.084 -> Water_Temp.:105.01, PWM:60
19:05:33.844 -> Water_Temp.:105.01, PWM:60
19:05:34.636 -> Water_Temp.:105.01, PWM:60
19:05:35.437 -> Water_Temp.:105.01, PWM:60
19:05:36.196 -> Water_Temp.:105.01, PWM:60
19:05:36.990 -> Water_Temp.:105.01, PWM:60
19:05:37.754 -> Water_Temp.:105.12, PWM:60
19:05:38.545 -> Water_Temp.:105.12, PWM:60
19:05:39.304 -> Water_Temp.:105.01, PWM:60
19:05:40.100 -> Water_Temp.:105.12, PWM:60
19:05:40.894 -> Water_Temp.:105.12, PWM:60
19:05:41.655 -> Water_Temp.:105.12, PWM:60
19:05:42.451 -> Water_Temp.:105.01, PWM:60
19:05:43.209 -> Water_Temp.:105.01, PWM:60
19:05:44.003 -> Water_Temp.:105.01, PWM:60
19:05:44.797 -> Water_Temp.:105.01, PWM:60
19:05:45.560 -> Water_Temp.:105.01, PWM:60
19:05:46.350 -> Water_Temp.:105.01, PWM:60
19:05:47.110 -> Water_Temp.:105.01, PWM:60
19:05:47.906 -> Water_Temp.:105.01, PWM:60
19:05:48.666 -> Water_Temp.:105.01, PWM:60
19:05:49.463 -> Water_Temp.:105.12, PWM:60
19:05:50.229 -> Water_Temp.:105.12, PWM:60
19:05:51.026 -> Water_Temp.:105.12, PWM:60
19:05:51.788 -> Water_Temp.:105.12, PWM:60
19:05:52.581 -> Water_Temp.:105.12, PWM:60
19:05:54.140 -> Water_Temp.:105.12, PWM:60
19:05:54.933 -> Water_Temp.:105.12, PWM:60
19:05:55.694 -> Water_Temp.:105.12, PWM:60
19:05:56.489 -> Water_Temp.:105.12, PWM:60
19:05:57.248 -> Water_Temp.:105.12, PWM:60
19:05:58.041 -> Water_Temp.:105.12, PWM:60
19:05:58.801 -> Water_Temp.:105.12, PWM:60
19:05:59.600 -> Water_Temp.:105.12, PWM:60
19:06:00.360 -> Water_Temp.:105.12, PWM:60
19:06:01.154 -> Water_Temp.:105.12, PWM:60
19:06:01.947 -> Water_Temp.:105.12, PWM:60
19:06:02.708 -> Water_Temp.:105.12, PWM:60
19:06:03.509 -> Water_Temp.:105.12, PWM:60
19:06:04.270 -> Water_Temp.:105.12, PWM:60
19:06:05.064 -> Water_Temp.:105.12, PWM:60
19:06:05.828 -> Water_Temp.:105.12, PWM:60
19:06:06.620 -> Water_Temp.:105.12, PWM:60
19:06:08.143 -> Water_Temp.:105.12, PWM:60
19:06:08.938 -> Water_Temp.:105.12, PWM:60
19:06:09.732 -> Water_Temp.:105.12, PWM:60
19:06:10.494 -> Water_Temp.:105.12, PWM:60
19:06:11.286 -> Water_Temp.:105.12, PWM:60
19:06:12.048 -> Water_Temp.:105.12, PWM:60
19:06:12.842 -> Water_Temp.:105.12, PWM:60
19:06:13.600 -> Water_Temp.:105.12, PWM:60
19:06:15.156 -> Water_Temp.:105.12, PWM:60
19:06:15.953 -> Water_Temp.:105.12, PWM:60
19:06:16.707 -> Water_Temp.:105.12, PWM:60
19:06:17.499 -> Water_Temp.:105.12, PWM:60
19:06:18.261 -> Water_Temp.:105.12, PWM:60
19:06:19.052 -> Water_Temp.:105.12, PWM:60
19:06:19.845 -> Water_Temp.:105.12, PWM:60
19:06:20.603 -> Water_Temp.:105.12, PWM:60
19:06:21.401 -> Water_Temp.:105.12, PWM:60
19:06:22.163 -> Water_Temp.:105.12, PWM:60
19:06:22.956 -> Water_Temp.:105.12, PWM:60
19:06:24.508 -> Water_Temp.:105.12, PWM:60
19:06:25.271 -> Water_Temp.:105.12, PWM:60
19:06:26.071 -> Water_Temp.:105.12, PWM:60
19:06:26.834 -> Water_Temp.:105.12, PWM:60
19:06:27.627 -> Water_Temp.:105.12, PWM:60
19:06:28.388 -> Water_Temp.:105.12, PWM:60
19:06:29.184 -> Water_Temp.:105.12, PWM:60
19:06:29.942 -> Water_Temp.:105.12, PWM:60
19:06:32.258 -> Water_Temp.:105.24, PWM:60
19:06:33.016 -> Water_Temp.:105.24, PWM:60
19:06:33.809 -> Water_Temp.:105.24, PWM:60
19:06:34.605 -> Water_Temp.:105.24, PWM:60
19:06:35.366 -> Water_Temp.:105.24, PWM:60
19:06:36.166 -> Water_Temp.:105.24, PWM:60
19:06:36.927 -> Water_Temp.:105.24, PWM:60
19:06:37.720 -> Water_Temp.:105.24, PWM:60
19:06:38.479 -> Water_Temp.:105.24, PWM:60
19:06:39.276 -> Water_Temp.:105.24, PWM:60
19:06:40.038 -> Water_Temp.:105.24, PWM:60
19:06:40.832 -> Water_Temp.:105.24, PWM:60
19:06:41.629 -> Water_Temp.:105.24, PWM:60
19:06:42.392 -> Water_Temp.:105.24, PWM:60
19:06:43.183 -> Water_Temp.:105.24, PWM:60
19:06:43.944 -> Water_Temp.:105.24, PWM:60
19:06:44.740 -> Water_Temp.:105.24, PWM:60
19:06:45.500 -> Water_Temp.:105.24, PWM:60
19:06:46.295 -> Water_Temp.:105.24, PWM:60
19:06:47.089 -> Water_Temp.:105.24, PWM:60
19:06:47.845 -> Water_Temp.:105.24, PWM:60
19:06:48.640 -> Water_Temp.:105.24, PWM:60
19:06:49.403 -> Water_Temp.:105.24, PWM:60
19:06:50.958 -> Water_Temp.:105.24, PWM:60
19:06:51.720 -> Water_Temp.:105.24, PWM:60
19:06:52.516 -> Water_Temp.:105.24, PWM:60
19:06:53.278 -> Water_Temp.:105.24, PWM:60
19:06:54.076 -> Water_Temp.:105.24, PWM:60
19:06:54.836 -> Water_Temp.:105.24, PWM:60
19:06:55.629 -> Water_Temp.:105.24, PWM:60
19:06:56.390 -> Water_Temp.:105.24, PWM:60
19:06:57.184 -> Water_Temp.:105.24, PWM:60
19:06:57.976 -> Water_Temp.:105.24, PWM:60
19:06:58.737 -> Water_Temp.:105.24, PWM:60
19:06:59.536 -> Water_Temp.:105.24, PWM:60
19:07:00.296 -> Water_Temp.:105.24, PWM:60
19:07:01.090 -> Water_Temp.:105.24, PWM:60
19:07:01.884 -> Water_Temp.:105.24, PWM:60
19:07:02.645 -> Water_Temp.:105.24, PWM:60
19:07:03.440 -> Water_Temp.:105.24, PWM:60
19:07:04.200 -> Water_Temp.:105.24, PWM:60
19:07:04.993 -> Water_Temp.:105.24, PWM:60
19:07:05.756 -> Water_Temp.:105.24, PWM:60
19:07:06.550 -> Water_Temp.:105.24, PWM:60
19:07:08.101 -> Water_Temp.:105.24, PWM:60
19:07:08.895 -> Water_Temp.:105.24, PWM:60
19:07:09.657 -> Water_Temp.:105.24, PWM:60
19:07:10.453 -> Water_Temp.:105.24, PWM:60
19:07:11.214 -> Water_Temp.:105.24, PWM:60
19:07:12.011 -> Water_Temp.:105.24, PWM:60
19:07:12.774 -> Water_Temp.:105.24, PWM:60
19:07:13.568 -> Water_Temp.:105.24, PWM:60
19:07:14.329 -> Water_Temp.:105.24, PWM:60
19:07:15.123 -> Water_Temp.:105.24, PWM:60
19:07:15.917 -> Water_Temp.:105.24, PWM:60
19:07:16.674 -> Water_Temp.:105.24, PWM:60
19:07:17.470 -> Water_Temp.:105.24, PWM:60
19:07:18.232 -> Water_Temp.:105.24, PWM:60
19:07:19.025 -> Water_Temp.:105.24, PWM:60
19:07:20.584 -> Water_Temp.:105.24, PWM:60
19:07:21.377 -> Water_Temp.:105.24, PWM:60
19:07:22.931 -> Water_Temp.:105.24, PWM:60
19:07:23.687 -> Water_Temp.:105.35, PWM:60
19:07:24.481 -> Water_Temp.:105.35, PWM:60
19:07:25.277 -> Water_Temp.:105.35, PWM:61
19:07:26.040 -> Water_Temp.:105.35, PWM:61
19:07:26.833 -> Water_Temp.:105.35, PWM:61
19:07:27.591 -> Water_Temp.:105.35, PWM:61
19:07:28.388 -> Water_Temp.:105.35, PWM:61
19:07:29.150 -> Water_Temp.:105.35, PWM:61
19:07:29.945 -> Water_Temp.:105.35, PWM:61
19:07:30.738 -> Water_Temp.:105.35, PWM:61
19:07:31.503 -> Water_Temp.:105.35, PWM:61
19:07:32.295 -> Water_Temp.:105.35, PWM:61
19:07:33.055 -> Water_Temp.:105.35, PWM:61
19:07:33.854 -> Water_Temp.:105.35, PWM:61
19:07:35.408 -> Water_Temp.:105.35, PWM:61
19:07:36.171 -> Water_Temp.:105.35, PWM:61
19:07:36.966 -> Water_Temp.:105.35, PWM:61
19:07:37.725 -> Water_Temp.:105.35, PWM:61
19:07:38.523 -> Water_Temp.:105.35, PWM:61
19:07:39.319 -> Water_Temp.:105.35, PWM:61
19:07:40.079 -> Water_Temp.:105.35, PWM:61
19:07:40.871 -> Water_Temp.:105.35, PWM:61
19:07:41.630 -> Water_Temp.:105.35, PWM:61
19:07:43.194 -> Water_Temp.:105.35, PWM:61
19:07:43.990 -> Water_Temp.:105.35, PWM:61
19:07:44.751 -> Water_Temp.:105.35, PWM:61
19:07:45.548 -> Water_Temp.:105.35, PWM:61
19:07:46.308 -> Water_Temp.:105.35, PWM:61
19:07:47.100 -> Water_Temp.:105.35, PWM:61
19:07:47.895 -> Water_Temp.:105.35, PWM:61
19:07:48.652 -> Water_Temp.:105.35, PWM:61
19:07:49.451 -> Water_Temp.:105.35, PWM:61
19:07:50.209 -> Water_Temp.:105.46, PWM:61
19:07:51.001 -> Water_Temp.:105.46, PWM:61
19:07:51.789 -> Water_Temp.:105.46, PWM:61
19:07:52.548 -> Water_Temp.:105.46, PWM:61
19:07:53.343 -> Water_Temp.:105.46, PWM:61
19:07:54.105 -> Water_Temp.:105.46, PWM:61
19:07:54.900 -> Water_Temp.:105.46, PWM:61
19:07:55.665 -> Water_Temp.:105.46, PWM:61
19:07:56.458 -> Water_Temp.:105.46, PWM:61
19:07:57.256 -> Water_Temp.:105.46, PWM:61
19:07:58.017 -> Water_Temp.:105.46, PWM:61
19:07:58.806 -> Water_Temp.:105.46, PWM:61
19:07:59.566 -> Water_Temp.:105.46, PWM:61
19:08:00.359 -> Water_Temp.:105.46, PWM:61
19:08:01.151 -> Water_Temp.:105.46, PWM:61
19:08:01.909 -> Water_Temp.:105.46, PWM:61
19:08:02.701 -> Water_Temp.:105.46, PWM:61
19:08:03.465 -> Water_Temp.:105.35, PWM:61
19:08:04.262 -> Water_Temp.:105.35, PWM:61
19:08:05.024 -> Water_Temp.:105.46, PWM:61
19:08:05.821 -> Water_Temp.:105.35, PWM:61
19:08:07.375 -> Water_Temp.:105.35, PWM:61
19:08:08.137 -> Water_Temp.:105.35, PWM:61
19:08:08.929 -> Water_Temp.:105.35, PWM:61
19:08:09.691 -> Water_Temp.:105.35, PWM:61
19:08:10.487 -> Water_Temp.:105.35, PWM:61
19:08:11.248 -> Water_Temp.:105.35, PWM:61
19:08:12.044 -> Water_Temp.:105.35, PWM:61
19:08:12.806 -> Water_Temp.:105.35, PWM:61
19:08:13.600 -> Water_Temp.:105.35, PWM:61
19:08:14.395 -> Water_Temp.:105.35, PWM:61
19:08:15.157 -> Water_Temp.:105.35, PWM:61
19:08:15.948 -> Water_Temp.:105.35, PWM:61
19:08:16.709 -> Water_Temp.:105.35, PWM:61
19:08:17.500 -> Water_Temp.:107.04, PWM:61
19:08:18.292 -> Water_Temp.:109.06, PWM:61
19:08:19.056 -> Water_Temp.:106.70, PWM:61
19:08:19.850 -> Water_Temp.:106.25, PWM:61
19:08:20.608 -> Water_Temp.:106.02, PWM:61
19:08:21.402 -> Water_Temp.:105.91, PWM:61
19:08:22.198 -> Water_Temp.:105.80, PWM:61
19:08:22.956 -> Water_Temp.:105.80, PWM:61
19:08:23.750 -> Water_Temp.:105.69, PWM:61
19:08:24.511 -> Water_Temp.:105.69, PWM:61
19:08:25.306 -> Water_Temp.:105.69, PWM:61
19:08:26.099 -> Water_Temp.:105.57, PWM:61
19:08:26.859 -> Water_Temp.:105.57, PWM:61
19:08:27.654 -> Water_Temp.:105.57, PWM:61
19:08:28.413 -> Water_Temp.:105.57, PWM:61
19:08:29.209 -> Water_Temp.:105.46, PWM:61
19:08:29.967 -> Water_Temp.:105.46, PWM:61
19:08:30.762 -> Water_Temp.:107.15, PWM:61
19:08:31.556 -> Water_Temp.:109.06, PWM:61
19:08:32.311 -> Water_Temp.:106.70, PWM:61
19:08:33.107 -> Water_Temp.:106.36, PWM:61
19:08:33.869 -> Water_Temp.:106.14, PWM:61
19:08:34.659 -> Water_Temp.:106.02, PWM:61
19:08:35.455 -> Water_Temp.:105.91, PWM:61
19:08:36.220 -> Water_Temp.:105.80, PWM:61
19:08:37.017 -> Water_Temp.:105.80, PWM:61
19:08:37.776 -> Water_Temp.:105.69, PWM:61
19:08:38.568 -> Water_Temp.:105.69, PWM:61
19:08:39.331 -> Water_Temp.:105.69, PWM:61
19:08:40.129 -> Water_Temp.:105.57, PWM:61
19:08:40.886 -> Water_Temp.:105.57, PWM:61
19:08:41.682 -> Water_Temp.:105.57, PWM:61
19:08:42.476 -> Water_Temp.:105.57, PWM:61
19:08:43.236 -> Water_Temp.:105.57, PWM:61
19:08:44.033 -> Water_Temp.:105.57, PWM:61
19:08:44.789 -> Water_Temp.:105.57, PWM:61
19:08:45.587 -> Water_Temp.:105.57, PWM:61
19:08:46.345 -> Water_Temp.:105.57, PWM:61
19:08:47.138 -> Water_Temp.:105.57, PWM:61
19:08:47.935 -> Water_Temp.:105.57, PWM:61
19:08:49.492 -> Water_Temp.:105.57, PWM:61
19:08:50.250 -> Water_Temp.:105.57, PWM:61
19:08:51.044 -> Water_Temp.:105.57, PWM:61
19:08:52.596 -> Water_Temp.:105.57, PWM:61
19:08:53.356 -> Water_Temp.:105.57, PWM:61
19:08:54.154 -> Water_Temp.:105.57, PWM:61
19:08:54.913 -> Water_Temp.:105.57, PWM:61
19:08:55.707 -> Water_Temp.:105.57, PWM:61
19:08:56.500 -> Water_Temp.:105.57, PWM:61
19:08:57.261 -> Water_Temp.:105.57, PWM:61
19:08:58.052 -> Water_Temp.:105.57, PWM:61
19:08:59.575 -> Water_Temp.:105.57, PWM:61
19:09:00.366 -> Water_Temp.:105.57, PWM:61
19:09:01.161 -> Water_Temp.:105.57, PWM:61
19:09:01.921 -> Water_Temp.:105.57, PWM:61
19:09:02.716 -> Water_Temp.:105.57, PWM:61
19:09:03.477 -> Water_Temp.:105.57, PWM:61
19:09:04.269 -> Water_Temp.:105.57, PWM:61
19:09:05.064 -> Water_Temp.:105.46, PWM:61
19:09:05.828 -> Water_Temp.:105.46, PWM:61
19:09:06.621 -> Water_Temp.:105.46, PWM:61
19:09:07.384 -> Water_Temp.:105.46, PWM:61
19:09:08.182 -> Water_Temp.:105.46, PWM:61
19:09:09.738 -> Water_Temp.:105.46, PWM:61
19:09:10.499 -> Water_Temp.:105.46, PWM:61
19:09:11.293 -> Water_Temp.:105.46, PWM:61
19:09:12.050 -> Water_Temp.:105.46, PWM:61
19:09:12.847 -> Water_Temp.:105.46, PWM:61
19:09:13.640 -> Water_Temp.:105.46, PWM:61
19:09:14.404 -> Water_Temp.:105.46, PWM:61
19:09:15.198 -> Water_Temp.:105.46, PWM:61
19:09:15.965 -> Water_Temp.:105.46, PWM:61
19:09:16.757 -> Water_Temp.:105.46, PWM:61
19:09:17.516 -> Water_Temp.:105.46, PWM:61
19:09:18.308 -> Water_Temp.:105.57, PWM:61
19:09:19.104 -> Water_Temp.:105.57, PWM:61
19:09:19.867 -> Water_Temp.:105.57, PWM:61
19:09:20.663 -> Water_Temp.:105.57, PWM:61
19:09:21.428 -> Water_Temp.:105.57, PWM:61
19:09:22.978 -> Water_Temp.:105.57, PWM:61
19:09:23.738 -> Water_Temp.:105.57, PWM:61
19:09:25.293 -> Water_Temp.:105.57, PWM:61
19:09:26.087 -> Water_Temp.:105.57, PWM:61
19:09:26.849 -> Water_Temp.:105.57, PWM:61
19:09:27.646 -> Water_Temp.:105.57, PWM:61
19:09:28.440 -> Water_Temp.:105.57, PWM:61
19:09:29.203 -> Water_Temp.:105.57, PWM:61
19:09:29.999 -> Water_Temp.:105.57, PWM:61
19:09:30.756 -> Water_Temp.:105.57, PWM:61
19:09:32.313 -> Water_Temp.:105.57, PWM:61
19:09:33.108 -> Water_Temp.:105.57, PWM:61
19:09:33.866 -> Water_Temp.:105.57, PWM:61
19:09:34.662 -> Water_Temp.:105.57, PWM:61
19:09:35.421 -> Water_Temp.:105.57, PWM:61
19:09:36.221 -> Water_Temp.:105.57, PWM:61
19:09:36.983 -> Water_Temp.:105.57, PWM:61
19:09:37.780 -> Water_Temp.:105.57, PWM:61
19:09:38.542 -> Water_Temp.:105.57, PWM:61
19:09:39.338 -> Water_Temp.:105.57, PWM:61
19:09:40.097 -> Water_Temp.:105.57, PWM:61
19:09:40.892 -> Water_Temp.:105.57, PWM:61
19:09:41.684 -> Water_Temp.:105.57, PWM:61
19:09:42.446 -> Water_Temp.:105.69, PWM:61
19:09:43.243 -> Water_Temp.:105.69, PWM:61
19:09:44.001 -> Water_Temp.:105.69, PWM:61
19:09:44.793 -> Water_Temp.:105.69, PWM:61
19:09:45.557 -> Water_Temp.:105.69, PWM:61
19:09:46.355 -> Water_Temp.:105.69, PWM:61
19:09:47.118 -> Water_Temp.:105.69, PWM:61
19:09:47.912 -> Water_Temp.:105.69, PWM:61
19:09:48.676 -> Water_Temp.:105.69, PWM:61
19:09:49.478 -> Water_Temp.:105.69, PWM:61
19:09:51.037 -> Water_Temp.:105.69, PWM:61
19:09:51.800 -> Water_Temp.:105.69, PWM:61
19:09:52.591 -> Water_Temp.:105.69, PWM:61
19:09:53.383 -> Water_Temp.:105.69, PWM:61
19:09:54.143 -> Water_Temp.:105.69, PWM:61
19:09:54.938 -> Water_Temp.:105.69, PWM:61
19:09:56.465 -> Water_Temp.:105.69, PWM:61
19:09:58.019 -> Water_Temp.:105.69, PWM:61
19:09:58.816 -> Water_Temp.:105.69, PWM:61
19:09:59.578 -> Water_Temp.:105.69, PWM:61
19:10:01.127 -> Water_Temp.:105.69, PWM:61
19:10:01.890 -> Water_Temp.:105.69, PWM:61
19:10:02.684 -> Water_Temp.:105.69, PWM:61
19:10:03.476 -> Water_Temp.:105.69, PWM:61
19:10:04.234 -> Water_Temp.:105.69, PWM:61
19:10:05.026 -> Water_Temp.:105.69, PWM:61
19:10:05.793 -> Water_Temp.:105.69, PWM:61
19:10:06.587 -> Water_Temp.:105.69, PWM:61
19:10:07.350 -> Water_Temp.:105.69, PWM:61
19:10:08.146 -> Water_Temp.:105.69, PWM:61
19:10:09.703 -> Water_Temp.:105.69, PWM:61
19:10:10.463 -> Water_Temp.:105.69, PWM:61
19:10:11.256 -> Water_Temp.:105.69, PWM:61
19:10:12.017 -> Water_Temp.:105.69, PWM:61
19:10:12.816 -> Water_Temp.:105.69, PWM:61
19:10:13.574 -> Water_Temp.:105.69, PWM:61
19:10:14.370 -> Water_Temp.:105.69, PWM:61
19:10:15.130 -> Water_Temp.:105.69, PWM:61
19:10:15.926 -> Water_Temp.:105.69, PWM:61
19:10:16.717 -> Water_Temp.:105.69, PWM:61
19:10:18.241 -> Water_Temp.:105.69, PWM:61
19:10:19.033 -> Water_Temp.:105.69, PWM:61
19:10:19.828 -> Water_Temp.:105.69, PWM:61
19:10:20.588 -> Water_Temp.:105.69, PWM:61
19:10:21.384 -> Water_Temp.:105.69, PWM:61
19:10:22.145 -> Water_Temp.:105.69, PWM:61
19:10:22.938 -> Water_Temp.:105.69, PWM:61
19:10:23.695 -> Water_Temp.:105.69, PWM:61
19:10:24.493 -> Water_Temp.:105.69, PWM:61
19:10:25.257 -> Water_Temp.:105.69, PWM:61
19:10:26.052 -> Water_Temp.:105.69, PWM:61
19:10:26.815 -> Water_Temp.:105.69, PWM:61
19:10:27.611 -> Water_Temp.:105.69, PWM:61
19:10:28.407 -> Water_Temp.:105.69, PWM:61
19:10:29.166 -> Water_Temp.:105.69, PWM:61
19:10:29.962 -> Water_Temp.:105.69, PWM:61
19:10:30.723 -> Water_Temp.:105.69, PWM:61
19:10:31.518 -> Water_Temp.:105.69, PWM:61
19:10:32.278 -> Water_Temp.:105.69, PWM:61
19:10:33.073 -> Water_Temp.:105.69, PWM:61
19:10:33.866 -> Water_Temp.:105.69, PWM:61
19:10:34.626 -> Water_Temp.:105.80, PWM:61
19:10:35.421 -> Water_Temp.:105.80, PWM:61
19:10:36.944 -> Water_Temp.:105.80, PWM:61
19:10:37.738 -> Water_Temp.:105.80, PWM:61
19:10:38.537 -> Water_Temp.:105.80, PWM:61
19:10:39.299 -> Water_Temp.:105.80, PWM:61
19:10:40.099 -> Water_Temp.:105.80, PWM:61
19:10:40.860 -> Water_Temp.:105.80, PWM:61
19:10:41.654 -> Water_Temp.:105.80, PWM:61
19:10:42.413 -> Water_Temp.:105.80, PWM:61
19:10:43.210 -> Water_Temp.:105.80, PWM:61
19:10:43.971 -> Water_Temp.:105.80, PWM:61
19:10:44.766 -> Water_Temp.:105.80, PWM:61
19:10:45.526 -> Water_Temp.:105.91, PWM:61
19:10:46.317 -> Water_Temp.:105.80, PWM:61
19:10:47.109 -> Water_Temp.:105.91, PWM:61
19:10:47.868 -> Water_Temp.:105.80, PWM:61
19:10:49.457 -> Water_Temp.:105.80, PWM:61
19:10:50.218 -> Water_Temp.:105.80, PWM:61
19:10:51.012 -> Water_Temp.:105.80, PWM:61
19:10:51.772 -> Water_Temp.:105.80, PWM:61
19:10:52.565 -> Water_Temp.:105.80, PWM:61
19:10:53.326 -> Water_Temp.:105.80, PWM:61
19:10:54.120 -> Water_Temp.:105.80, PWM:61
19:10:55.681 -> Water_Temp.:105.80, PWM:61
19:10:56.444 -> Water_Temp.:105.80, PWM:61
19:10:57.240 -> Water_Temp.:105.80, PWM:61
19:10:58.033 -> Water_Temp.:105.80, PWM:61
19:10:58.793 -> Water_Temp.:105.69, PWM:61
19:10:59.589 -> Water_Temp.:105.69, PWM:61
19:11:01.102 -> Water_Temp.:105.69, PWM:61
19:11:01.896 -> Water_Temp.:105.69, PWM:61
19:11:02.662 -> Water_Temp.:105.69, PWM:61
19:11:03.457 -> Water_Temp.:105.80, PWM:61
19:11:04.219 -> Water_Temp.:105.80, PWM:61
19:11:05.015 -> Water_Temp.:105.80, PWM:61
19:11:05.779 -> Water_Temp.:105.80, PWM:61
19:11:06.571 -> Water_Temp.:105.80, PWM:61
19:11:07.366 -> Water_Temp.:105.80, PWM:61
19:11:08.123 -> Water_Temp.:105.80, PWM:61
19:11:08.921 -> Water_Temp.:105.80, PWM:61
19:11:09.682 -> Water_Temp.:105.80, PWM:61
19:11:11.270 -> Water_Temp.:105.80, PWM:61
19:11:12.032 -> Water_Temp.:105.80, PWM:61
19:11:12.826 -> Water_Temp.:105.80, PWM:61
19:11:14.343 -> Water_Temp.:105.91, PWM:61
19:11:15.140 -> Water_Temp.:105.80, PWM:61
19:11:15.902 -> Water_Temp.:105.91, PWM:61
19:11:16.697 -> Water_Temp.:105.91, PWM:61
19:11:17.458 -> Water_Temp.:105.91, PWM:61
19:11:18.250 -> Water_Temp.:105.91, PWM:61
19:11:19.045 -> Water_Temp.:105.91, PWM:61
19:11:19.805 -> Water_Temp.:105.91, PWM:61
19:11:20.599 -> Water_Temp.:105.91, PWM:61
19:11:21.359 -> Water_Temp.:105.91, PWM:61
19:11:22.154 -> Water_Temp.:105.91, PWM:61
19:11:22.917 -> Water_Temp.:105.91, PWM:61
19:11:23.713 -> Water_Temp.:105.80, PWM:61
19:11:24.503 -> Water_Temp.:105.80, PWM:61
19:11:25.266 -> Water_Temp.:105.80, PWM:61
19:11:26.060 -> Water_Temp.:105.80, PWM:61
19:11:26.820 -> Water_Temp.:105.80, PWM:61
19:11:27.614 -> Water_Temp.:105.80, PWM:61
19:11:28.374 -> Water_Temp.:105.80, PWM:61
19:11:29.169 -> Water_Temp.:105.80, PWM:61
19:11:29.934 -> Water_Temp.:105.91, PWM:61
19:11:30.724 -> Water_Temp.:105.91, PWM:61
19:11:31.515 -> Water_Temp.:105.91, PWM:61
19:11:33.038 -> Water_Temp.:105.91, PWM:61
19:11:33.830 -> Water_Temp.:105.91, PWM:61
19:11:34.624 -> Water_Temp.:105.91, PWM:61
19:11:35.387 -> Water_Temp.:105.91, PWM:61
19:11:36.186 -> Water_Temp.:105.91, PWM:61
19:11:36.945 -> Water_Temp.:105.91, PWM:61
19:11:38.500 -> Water_Temp.:105.91, PWM:61
19:11:39.262 -> Water_Temp.:105.91, PWM:61
19:11:40.060 -> Water_Temp.:105.91, PWM:61
19:11:40.856 -> Water_Temp.:105.91, PWM:61
19:11:41.617 -> Water_Temp.:105.91, PWM:61
19:11:42.413 -> Water_Temp.:105.91, PWM:61
19:11:43.930 -> Water_Temp.:105.91, PWM:61
19:11:44.726 -> Water_Temp.:105.91, PWM:61
19:11:45.487 -> Water_Temp.:105.91, PWM:61
19:11:46.279 -> Water_Temp.:105.91, PWM:61
19:11:47.039 -> Water_Temp.:105.91, PWM:61
19:11:47.833 -> Water_Temp.:105.91, PWM:61
19:11:48.624 -> Water_Temp.:105.91, PWM:61
19:11:49.388 -> Water_Temp.:105.91, PWM:61
19:11:50.183 -> Water_Temp.:105.91, PWM:61
19:11:50.937 -> Water_Temp.:105.91, PWM:61
19:11:51.730 -> Water_Temp.:105.91, PWM:61
19:11:52.524 -> Water_Temp.:105.91, PWM:61
19:11:53.284 -> Water_Temp.:105.91, PWM:61
19:11:54.078 -> Water_Temp.:105.91, PWM:61
19:11:54.841 -> Water_Temp.:105.91, PWM:61
19:11:55.636 -> Water_Temp.:105.91, PWM:61
19:11:56.396 -> Water_Temp.:106.02, PWM:61
19:11:57.191 -> Water_Temp.:106.02, PWM:61
19:11:57.982 -> Water_Temp.:106.02, PWM:61
19:11:59.529 -> Water_Temp.:109.62, PWM:61
19:12:00.292 -> Water_Temp.:107.26, PWM:61
19:12:01.086 -> Water_Temp.:106.92, PWM:61
19:12:01.877 -> Water_Temp.:106.70, PWM:61
19:12:02.632 -> Water_Temp.:106.59, PWM:61
19:12:03.424 -> Water_Temp.:106.47, PWM:61
19:12:04.220 -> Water_Temp.:106.36, PWM:61
19:12:04.983 -> Water_Temp.:106.36, PWM:61
19:12:05.777 -> Water_Temp.:106.36, PWM:61
19:12:06.539 -> Water_Temp.:106.25, PWM:61
19:12:07.330 -> Water_Temp.:106.25, PWM:61
19:12:08.120 -> Water_Temp.:106.25, PWM:61
19:12:08.879 -> Water_Temp.:106.25, PWM:61
19:12:09.673 -> Water_Temp.:106.25, PWM:61
19:12:10.464 -> Water_Temp.:106.25, PWM:61
19:12:11.986 -> Water_Temp.:106.14, PWM:61
19:12:12.781 -> Water_Temp.:106.14, PWM:61
19:12:13.541 -> Water_Temp.:106.14, PWM:61
19:12:14.335 -> Water_Temp.:106.14, PWM:61
19:12:15.094 -> Water_Temp.:106.14, PWM:61
19:12:15.891 -> Water_Temp.:106.14, PWM:61
19:12:16.683 -> Water_Temp.:106.14, PWM:61
19:12:18.197 -> Water_Temp.:106.02, PWM:61
19:12:18.989 -> Water_Temp.:106.02, PWM:61
19:12:19.756 -> Water_Temp.:106.02, PWM:61
19:12:20.552 -> Water_Temp.:106.14, PWM:61
19:12:21.313 -> Water_Temp.:106.14, PWM:61
19:12:22.109 -> Water_Temp.:106.14, PWM:61
19:12:22.873 -> Water_Temp.:106.14, PWM:61
19:12:23.666 -> Water_Temp.:106.14, PWM:61
19:12:25.988 -> Water_Temp.:106.14, PWM:61
19:12:26.751 -> Water_Temp.:106.14, PWM:61
19:12:27.546 -> Water_Temp.:106.14, PWM:61
19:12:28.307 -> Water_Temp.:106.14, PWM:61
19:12:29.100 -> Water_Temp.:106.14, PWM:61
19:12:29.897 -> Water_Temp.:106.14, PWM:61
19:12:30.658 -> Water_Temp.:106.14, PWM:61
19:12:31.455 -> Water_Temp.:106.14, PWM:61
19:12:32.221 -> Water_Temp.:106.14, PWM:61
19:12:33.017 -> Water_Temp.:106.14, PWM:61
19:12:34.532 -> Water_Temp.:106.14, PWM:61
19:12:35.329 -> Water_Temp.:106.14, PWM:61
19:12:36.123 -> Water_Temp.:106.14, PWM:61
19:12:36.878 -> Water_Temp.:106.14, PWM:61
19:12:37.674 -> Water_Temp.:106.02, PWM:61
19:12:38.434 -> Water_Temp.:106.02, PWM:61
19:12:39.227 -> Water_Temp.:106.14, PWM:61
19:12:40.022 -> Water_Temp.:106.02, PWM:61
19:12:40.781 -> Water_Temp.:106.02, PWM:61
19:12:41.573 -> Water_Temp.:106.14, PWM:61
19:12:42.364 -> Water_Temp.:106.14, PWM:61
19:12:43.125 -> Water_Temp.:106.14, PWM:61
19:12:43.914 -> Water_Temp.:106.14, PWM:61
19:12:44.674 -> Water_Temp.:106.14, PWM:61
19:12:45.469 -> Water_Temp.:106.02, PWM:61
19:12:46.231 -> Water_Temp.:106.02, PWM:61
19:12:47.022 -> Water_Temp.:106.02, PWM:61
19:12:47.818 -> Water_Temp.:106.02, PWM:61
19:12:48.576 -> Water_Temp.:106.02, PWM:61
19:12:49.371 -> Water_Temp.:106.02, PWM:61
19:12:50.163 -> Water_Temp.:106.02, PWM:61
19:12:50.922 -> Water_Temp.:106.02, PWM:61
19:12:51.720 -> Water_Temp.:106.02, PWM:61
19:12:52.478 -> Water_Temp.:106.02, PWM:61
19:12:54.794 -> Water_Temp.:106.02, PWM:61
19:12:55.591 -> Water_Temp.:106.02, PWM:61
19:12:56.350 -> Water_Temp.:106.02, PWM:61
19:12:57.145 -> Water_Temp.:106.02, PWM:61
19:12:57.936 -> Water_Temp.:106.02, PWM:61
19:12:58.696 -> Water_Temp.:106.02, PWM:61
19:12:59.491 -> Water_Temp.:106.02, PWM:61
19:13:00.250 -> Water_Temp.:106.02, PWM:61
19:13:01.044 -> Water_Temp.:106.02, PWM:61
19:13:01.836 -> Water_Temp.:106.02, PWM:61
19:13:02.596 -> Water_Temp.:106.02, PWM:61
19:13:03.385 -> Water_Temp.:106.02, PWM:61
19:13:04.936 -> Water_Temp.:106.02, PWM:61
19:13:05.732 -> Water_Temp.:106.02, PWM:61
19:13:06.515 -> Water_Temp.:106.02, PWM:61
19:13:07.277 -> Water_Temp.:106.02, PWM:61
19:13:08.073 -> Water_Temp.:106.02, PWM:61
19:13:08.829 -> Water_Temp.:106.02, PWM:61
19:13:09.621 -> Water_Temp.:106.02, PWM:61
19:13:10.412 -> Water_Temp.:106.02, PWM:61
19:13:11.932 -> Water_Temp.:106.02, PWM:61
19:13:12.726 -> Water_Temp.:106.02, PWM:61
19:13:14.279 -> Water_Temp.:106.02, PWM:61
19:13:15.070 -> Water_Temp.:106.02, PWM:61
19:13:15.831 -> Water_Temp.:106.02, PWM:61
19:13:16.625 -> Water_Temp.:106.02, PWM:61
19:13:17.420 -> Water_Temp.:106.02, PWM:61
19:13:18.182 -> Water_Temp.:106.02, PWM:61
19:13:18.971 -> Water_Temp.:106.02, PWM:61
19:13:19.727 -> Water_Temp.:106.02, PWM:61
19:13:20.521 -> Water_Temp.:105.91, PWM:61
19:13:21.316 -> Water_Temp.:106.02, PWM:61
19:13:22.076 -> Water_Temp.:105.91, PWM:61
19:13:22.866 -> Water_Temp.:105.91, PWM:61
19:13:23.626 -> Water_Temp.:105.91, PWM:61
19:13:24.419 -> Water_Temp.:105.91, PWM:61
19:13:25.215 -> Water_Temp.:105.91, PWM:61
19:13:25.978 -> Water_Temp.:105.91, PWM:61
19:13:26.775 -> Water_Temp.:105.91, PWM:61
19:13:27.537 -> Water_Temp.:105.91, PWM:61
19:13:28.329 -> Water_Temp.:105.91, PWM:61
19:13:29.089 -> Water_Temp.:105.91, PWM:61
19:13:29.889 -> Water_Temp.:105.91, PWM:61
19:13:30.679 -> Water_Temp.:105.91, PWM:61
19:13:31.442 -> Water_Temp.:105.91, PWM:61
19:13:32.234 -> Water_Temp.:105.91, PWM:61
19:13:32.995 -> Water_Temp.:105.91, PWM:61
19:13:33.791 -> Water_Temp.:105.91, PWM:61
19:13:34.551 -> Water_Temp.:105.91, PWM:61
19:13:36.134 -> Water_Temp.:106.02, PWM:61
19:13:36.890 -> Water_Temp.:106.02, PWM:61
19:13:37.683 -> Water_Temp.:105.91, PWM:61
19:13:38.474 -> Water_Temp.:105.91, PWM:61
19:13:39.233 -> Water_Temp.:105.91, PWM:61
19:13:40.031 -> Water_Temp.:105.91, PWM:61
19:13:40.790 -> Water_Temp.:105.91, PWM:61
19:13:41.583 -> Water_Temp.:105.91, PWM:61
19:13:42.375 -> Water_Temp.:105.91, PWM:61
19:13:43.132 -> Water_Temp.:105.91, PWM:61
19:13:43.924 -> Water_Temp.:105.91, PWM:61
19:13:44.717 -> Water_Temp.:105.91, PWM:61
19:13:45.472 -> Water_Temp.:105.91, PWM:61
19:13:46.267 -> Water_Temp.:105.91, PWM:61
19:13:47.058 -> Water_Temp.:105.91, PWM:61
19:13:47.816 -> Water_Temp.:105.91, PWM:61
19:13:48.610 -> Water_Temp.:105.91, PWM:61
19:13:49.371 -> Water_Temp.:105.91, PWM:61
19:13:50.164 -> Water_Temp.:105.91, PWM:61
19:13:50.924 -> Water_Temp.:105.91, PWM:61
19:13:52.506 -> Water_Temp.:106.02, PWM:61
19:13:53.297 -> Water_Temp.:106.02, PWM:61
19:13:54.056 -> Water_Temp.:106.02, PWM:61
19:13:54.850 -> Water_Temp.:106.02, PWM:61
19:13:55.607 -> Water_Temp.:106.02, PWM:61
19:13:56.399 -> Water_Temp.:106.02, PWM:61
19:13:57.192 -> Water_Temp.:106.02, PWM:61
19:13:57.949 -> Water_Temp.:106.02, PWM:61
19:13:58.745 -> Water_Temp.:106.02, PWM:61
19:13:59.507 -> Water_Temp.:106.02, PWM:61
19:14:00.300 -> Water_Temp.:106.02, PWM:61
19:14:01.090 -> Water_Temp.:106.02, PWM:61
19:14:01.850 -> Water_Temp.:106.14, PWM:61
19:14:02.647 -> Water_Temp.:106.14, PWM:61
19:14:03.410 -> Water_Temp.:106.02, PWM:61
19:14:04.202 -> Water_Temp.:106.02, PWM:61
19:14:04.994 -> Water_Temp.:106.02, PWM:61
19:14:05.756 -> Water_Temp.:106.02, PWM:61
19:14:06.550 -> Water_Temp.:106.02, PWM:61
19:14:07.305 -> Water_Temp.:106.02, PWM:61
19:14:08.098 -> Water_Temp.:106.02, PWM:61
19:14:08.892 -> Water_Temp.:106.02, PWM:61
19:14:09.647 -> Water_Temp.:106.02, PWM:61
19:14:11.202 -> Water_Temp.:106.02, PWM:61
19:14:11.990 -> Water_Temp.:106.02, PWM:61
19:14:12.787 -> Water_Temp.:106.02, PWM:61
19:14:13.578 -> Water_Temp.:106.02, PWM:61
19:14:14.339 -> Water_Temp.:106.02, PWM:61
19:14:15.134 -> Water_Temp.:106.02, PWM:61
19:14:15.897 -> Water_Temp.:106.02, PWM:61
19:14:16.691 -> Water_Temp.:106.02, PWM:61
19:14:17.451 -> Water_Temp.:106.02, PWM:61
19:14:19.004 -> Water_Temp.:106.02, PWM:61
19:14:19.759 -> Water_Temp.:106.02, PWM:61
19:14:20.551 -> Water_Temp.:105.91, PWM:61
19:14:21.344 -> Water_Temp.:106.02, PWM:61
19:14:22.104 -> Water_Temp.:105.91, PWM:61
19:14:22.896 -> Water_Temp.:106.02, PWM:61
19:14:23.690 -> Water_Temp.:106.02, PWM:61
19:14:24.450 -> Water_Temp.:106.02, PWM:61
19:14:25.245 -> Water_Temp.:106.02, PWM:61
19:14:26.004 -> Water_Temp.:106.02, PWM:61
19:14:26.796 -> Water_Temp.:106.02, PWM:61
19:14:27.557 -> Water_Temp.:106.02, PWM:61
19:14:28.350 -> Water_Temp.:106.02, PWM:61
19:14:29.143 -> Water_Temp.:106.02, PWM:61
19:14:29.904 -> Water_Temp.:106.02, PWM:61
19:14:30.694 -> Water_Temp.:106.02, PWM:61
19:14:31.483 -> Water_Temp.:106.02, PWM:61
19:14:32.246 -> Water_Temp.:106.02, PWM:61
19:14:33.038 -> Water_Temp.:106.02, PWM:61
19:14:33.798 -> Water_Temp.:106.02, PWM:61
19:14:34.594 -> Water_Temp.:106.02, PWM:61
19:14:35.356 -> Water_Temp.:106.02, PWM:61
19:14:36.906 -> Water_Temp.:106.02, PWM:61
19:14:38.458 -> Water_Temp.:106.02, PWM:61
19:14:39.251 -> Water_Temp.:106.02, PWM:61
19:14:40.016 -> Water_Temp.:106.02, PWM:61
19:14:40.814 -> Water_Temp.:106.02, PWM:61
19:14:41.570 -> Water_Temp.:106.02, PWM:61
19:14:42.364 -> Water_Temp.:106.02, PWM:61
19:14:43.159 -> Water_Temp.:106.02, PWM:61
19:14:43.921 -> Water_Temp.:106.02, PWM:61
19:14:44.711 -> Water_Temp.:106.02, PWM:61
19:14:45.474 -> Water_Temp.:106.02, PWM:61
19:14:46.272 -> Water_Temp.:106.02, PWM:61
19:14:47.033 -> Water_Temp.:106.02, PWM:61
19:14:47.833 -> Water_Temp.:106.02, PWM:61
19:14:48.590 -> Water_Temp.:106.02, PWM:61
19:14:49.384 -> Water_Temp.:106.02, PWM:61
19:14:50.180 -> Water_Temp.:106.02, PWM:61
19:14:50.935 -> Water_Temp.:106.02, PWM:61
19:14:52.489 -> Water_Temp.:106.02, PWM:61
19:14:54.044 -> Water_Temp.:106.02, PWM:61
19:14:54.804 -> Water_Temp.:106.02, PWM:61
19:14:55.601 -> Water_Temp.:106.02, PWM:61
19:14:56.364 -> Water_Temp.:106.02, PWM:61
19:14:57.156 -> Water_Temp.:106.02, PWM:61
19:14:57.948 -> Water_Temp.:106.02, PWM:61
19:14:58.703 -> Water_Temp.:106.02, PWM:61
19:14:59.500 -> Water_Temp.:106.02, PWM:61
19:15:00.294 -> Water_Temp.:106.02, PWM:61
19:15:01.058 -> Water_Temp.:106.02, PWM:61
19:15:01.854 -> Water_Temp.:106.02, PWM:61
19:15:02.613 -> Water_Temp.:106.02, PWM:61
19:15:03.406 -> Water_Temp.:106.02, PWM:61
19:15:04.165 -> Water_Temp.:106.02, PWM:61
19:15:04.959 -> Water_Temp.:106.02, PWM:61
19:15:05.723 -> Water_Temp.:106.02, PWM:61
19:15:06.518 -> Water_Temp.:106.02, PWM:61
19:15:07.310 -> Water_Temp.:106.02, PWM:61
19:15:08.065 -> Water_Temp.:106.02, PWM:61
19:15:08.861 -> Water_Temp.:106.02, PWM:61
19:15:09.625 -> Water_Temp.:106.02, PWM:61
19:15:10.419 -> Water_Temp.:106.02, PWM:61
19:15:11.212 -> Water_Temp.:106.02, PWM:61
19:15:11.968 -> Water_Temp.:106.02, PWM:61
19:15:12.761 -> Water_Temp.:106.02, PWM:61
19:15:13.553 -> Water_Temp.:106.02, PWM:61
19:15:14.312 -> Water_Temp.:106.02, PWM:61
19:15:15.105 -> Water_Temp.:106.02, PWM:61
19:15:15.865 -> Water_Temp.:106.02, PWM:61
19:15:16.658 -> Water_Temp.:106.02, PWM:61
19:15:17.449 -> Water_Temp.:106.02, PWM:61
19:15:18.211 -> Water_Temp.:106.02, PWM:61
19:15:19.003 -> Water_Temp.:106.02, PWM:61
19:15:19.794 -> Water_Temp.:106.02, PWM:61
19:15:20.555 -> Water_Temp.:106.02, PWM:61
19:15:21.348 -> Water_Temp.:106.02, PWM:61
19:15:22.898 -> Water_Temp.:106.02, PWM:61
19:15:23.694 -> Water_Temp.:106.02, PWM:61
19:15:24.451 -> Water_Temp.:106.02, PWM:61
19:15:25.242 -> Water_Temp.:106.02, PWM:61
19:15:26.000 -> Water_Temp.:106.02, PWM:61
19:15:26.794 -> Water_Temp.:106.02, PWM:61
19:15:27.590 -> Water_Temp.:106.02, PWM:61
19:15:28.353 -> Water_Temp.:106.02, PWM:61
19:15:29.904 -> Water_Temp.:106.14, PWM:61
19:15:30.696 -> Water_Temp.:106.14, PWM:61
19:15:31.487 -> Water_Temp.:106.14, PWM:61
19:15:32.246 -> Water_Temp.:106.14, PWM:61
19:15:33.039 -> Water_Temp.:106.14, PWM:61
19:15:33.798 -> Water_Temp.:106.14, PWM:61
19:15:34.595 -> Water_Temp.:106.14, PWM:61
19:15:35.359 -> Water_Temp.:106.14, PWM:61
19:15:36.915 -> Water_Temp.:106.14, PWM:61
19:15:37.677 -> Water_Temp.:106.14, PWM:61
19:15:38.468 -> Water_Temp.:106.14, PWM:61
19:15:39.262 -> Water_Temp.:106.14, PWM:61
19:15:40.021 -> Water_Temp.:106.14, PWM:61
19:15:40.813 -> Water_Temp.:106.25, PWM:61
19:15:41.608 -> Water_Temp.:106.25, PWM:61
19:15:42.369 -> Water_Temp.:106.25, PWM:61
19:15:43.162 -> Water_Temp.:106.25, PWM:61
19:15:43.918 -> Water_Temp.:106.25, PWM:61
19:15:44.714 -> Water_Temp.:106.25, PWM:61
19:15:45.508 -> Water_Temp.:106.25, PWM:61
19:15:46.267 -> Water_Temp.:106.25, PWM:61
19:15:47.060 -> Water_Temp.:106.25, PWM:61
19:15:47.819 -> Water_Temp.:106.25, PWM:61
19:15:48.610 -> Water_Temp.:106.25, PWM:61
19:15:49.407 -> Water_Temp.:106.25, PWM:61
19:15:50.169 -> Water_Temp.:106.25, PWM:61
19:15:50.956 -> Water_Temp.:106.25, PWM:61
19:15:51.745 -> Water_Temp.:106.25, PWM:61
19:15:53.297 -> Water_Temp.:109.85, PWM:61
19:15:54.058 -> Water_Temp.:107.49, PWM:61
19:15:54.853 -> Water_Temp.:107.04, PWM:61
19:15:55.648 -> Water_Temp.:106.92, PWM:61
19:15:56.413 -> Water_Temp.:106.70, PWM:61
19:15:57.205 -> Water_Temp.:106.70, PWM:61
19:15:57.966 -> Water_Temp.:106.59, PWM:61
19:15:58.760 -> Water_Temp.:106.47, PWM:61
19:15:59.520 -> Water_Temp.:106.47, PWM:61
19:16:00.309 -> Water_Temp.:106.36, PWM:61
19:16:01.100 -> Water_Temp.:106.36, PWM:61
19:16:02.649 -> Water_Temp.:106.36, PWM:61
19:16:03.409 -> Water_Temp.:106.36, PWM:61
19:16:04.203 -> Water_Temp.:106.25, PWM:61
19:16:04.961 -> Water_Temp.:106.25, PWM:61
19:16:05.764 -> Water_Temp.:106.25, PWM:61
19:16:07.313 -> Water_Temp.:106.25, PWM:61
19:16:08.107 -> Water_Temp.:106.25, PWM:61
19:16:08.868 -> Water_Temp.:106.25, PWM:61
19:16:09.662 -> Water_Temp.:106.25, PWM:61
19:16:10.425 -> Water_Temp.:106.25, PWM:61
19:16:11.217 -> Water_Temp.:106.25, PWM:61
19:16:11.979 -> Water_Temp.:106.25, PWM:61
19:16:12.771 -> Water_Temp.:106.25, PWM:61
19:16:13.564 -> Water_Temp.:106.25, PWM:61
19:16:14.327 -> Water_Temp.:106.36, PWM:61
19:16:15.122 -> Water_Temp.:106.36, PWM:61
19:16:15.885 -> Water_Temp.:106.36, PWM:61
19:16:16.676 -> Water_Temp.:106.36, PWM:61
19:16:18.224 -> Water_Temp.:106.36, PWM:61
19:16:19.020 -> Water_Temp.:106.36, PWM:61
19:16:20.576 -> Water_Temp.:106.36, PWM:61
19:16:21.336 -> Water_Temp.:106.36, PWM:61
19:16:22.128 -> Water_Temp.:106.36, PWM:61
19:16:22.921 -> Water_Temp.:106.36, PWM:61
19:16:23.685 -> Water_Temp.:106.36, PWM:61
19:16:25.239 -> Water_Temp.:106.36, PWM:61
19:16:25.998 -> Water_Temp.:106.25, PWM:61
19:16:26.793 -> Water_Temp.:106.25, PWM:61
19:16:27.554 -> Water_Temp.:106.25, PWM:61
19:16:28.349 -> Water_Temp.:106.25, PWM:61
19:16:29.142 -> Water_Temp.:106.25, PWM:61
19:16:29.902 -> Water_Temp.:106.25, PWM:61
19:16:30.691 -> Water_Temp.:106.25, PWM:61
19:16:31.487 -> Water_Temp.:106.25, PWM:61
19:16:32.243 -> Water_Temp.:106.25, PWM:61
19:16:33.039 -> Water_Temp.:106.25, PWM:61
19:16:33.797 -> Water_Temp.:106.25, PWM:61
19:16:34.596 -> Water_Temp.:106.25, PWM:61
19:16:35.360 -> Water_Temp.:106.25, PWM:61
19:16:36.913 -> Water_Temp.:106.25, PWM:61
19:16:37.704 -> Water_Temp.:106.25, PWM:61
19:16:38.457 -> Water_Temp.:106.25, PWM:61
19:16:39.252 -> Water_Temp.:106.14, PWM:61
19:16:40.044 -> Water_Temp.:106.14, PWM:61
19:16:40.803 -> Water_Temp.:106.14, PWM:61
19:16:41.601 -> Water_Temp.:106.14, PWM:61
19:16:42.361 -> Water_Temp.:106.14, PWM:61
19:16:43.158 -> Water_Temp.:106.14, PWM:61
19:16:44.678 -> Water_Temp.:106.14, PWM:61
19:16:45.474 -> Water_Temp.:106.14, PWM:61
19:16:46.232 -> Water_Temp.:106.14, PWM:61
19:16:47.022 -> Water_Temp.:106.14, PWM:61
19:16:47.819 -> Water_Temp.:106.14, PWM:61
19:16:48.582 -> Water_Temp.:106.25, PWM:61
19:16:49.377 -> Water_Temp.:106.25, PWM:61
19:16:50.135 -> Water_Temp.:106.25, PWM:61
19:16:50.931 -> Water_Temp.:106.25, PWM:61
19:16:51.692 -> Water_Temp.:106.25, PWM:61
19:16:52.487 -> Water_Temp.:106.25, PWM:61
19:16:53.283 -> Water_Temp.:106.25, PWM:61
19:16:54.037 -> Water_Temp.:106.25, PWM:61
19:16:54.830 -> Water_Temp.:106.25, PWM:61
19:16:55.591 -> Water_Temp.:106.25, PWM:61
19:16:56.388 -> Water_Temp.:106.25, PWM:61
19:16:57.181 -> Water_Temp.:106.25, PWM:61
19:16:58.737 -> Water_Temp.:106.25, PWM:61
19:16:59.494 -> Water_Temp.:106.25, PWM:61
19:17:00.288 -> Water_Temp.:106.25, PWM:61
19:17:01.083 -> Water_Temp.:106.25, PWM:61
19:17:01.844 -> Water_Temp.:106.25, PWM:61
19:17:02.636 -> Water_Temp.:106.25, PWM:61
19:17:03.394 -> Water_Temp.:106.25, PWM:61
19:17:04.183 -> Water_Temp.:106.25, PWM:61
19:17:04.976 -> Water_Temp.:106.25, PWM:61
19:17:05.739 -> Water_Temp.:106.25, PWM:61
19:17:06.536 -> Water_Temp.:106.25, PWM:61
19:17:07.296 -> Water_Temp.:106.25, PWM:61
19:17:08.091 -> Water_Temp.:106.25, PWM:61
19:17:08.850 -> Water_Temp.:106.25, PWM:61
19:17:09.647 -> Water_Temp.:106.25, PWM:61
19:17:10.440 -> Water_Temp.:106.25, PWM:61
19:17:11.196 -> Water_Temp.:106.25, PWM:61
19:17:11.991 -> Water_Temp.:106.25, PWM:61
19:17:12.751 -> Water_Temp.:106.25, PWM:61
19:17:13.547 -> Water_Temp.:106.14, PWM:61
19:17:14.340 -> Water_Temp.:106.14, PWM:61
19:17:15.100 -> Water_Temp.:106.14, PWM:61
19:17:15.895 -> Water_Temp.:106.14, PWM:61
19:17:16.653 -> Water_Temp.:106.14, PWM:61
19:17:17.448 -> Water_Temp.:106.14, PWM:61
19:17:18.211 -> Water_Temp.:106.14, PWM:61
19:17:19.006 -> Water_Temp.:106.14, PWM:61
19:17:19.798 -> Water_Temp.:106.14, PWM:61
19:17:20.553 -> Water_Temp.:106.14, PWM:61
19:17:21.345 -> Water_Temp.:106.14, PWM:61
19:17:22.140 -> Water_Temp.:106.14, PWM:61
19:17:22.899 -> Water_Temp.:106.14, PWM:61
19:17:23.693 -> Water_Temp.:106.14, PWM:61
19:17:25.211 -> Water_Temp.:106.14, PWM:61
19:17:26.004 -> Water_Temp.:106.14, PWM:61
19:17:26.767 -> Water_Temp.:106.14, PWM:61
19:17:27.556 -> Water_Temp.:106.14, PWM:61
19:17:29.139 -> Water_Temp.:106.25, PWM:61
19:17:29.893 -> Water_Temp.:106.25, PWM:61
19:17:30.683 -> Water_Temp.:106.25, PWM:61
19:17:31.475 -> Water_Temp.:106.25, PWM:61
19:17:32.233 -> Water_Temp.:106.25, PWM:61
19:17:33.028 -> Water_Temp.:106.25, PWM:61
19:17:33.788 -> Water_Temp.:106.25, PWM:61
19:17:34.580 -> Water_Temp.:106.25, PWM:61
19:17:35.376 -> Water_Temp.:106.25, PWM:61
19:17:36.135 -> Water_Temp.:106.25, PWM:61
19:17:36.927 -> Water_Temp.:106.25, PWM:61
19:17:37.689 -> Water_Temp.:106.25, PWM:61
19:17:38.483 -> Water_Temp.:106.25, PWM:61
19:17:39.273 -> Water_Temp.:106.25, PWM:61
19:17:40.033 -> Water_Temp.:106.25, PWM:61
19:17:41.618 -> Water_Temp.:106.25, PWM:61
19:17:42.380 -> Water_Temp.:106.25, PWM:61
19:17:43.174 -> Water_Temp.:106.25, PWM:61
19:17:43.937 -> Water_Temp.:106.25, PWM:61
19:17:44.729 -> Water_Temp.:106.25, PWM:61
19:17:45.487 -> Water_Temp.:106.25, PWM:61
19:17:46.283 -> Water_Temp.:106.25, PWM:61
19:17:47.076 -> Water_Temp.:106.25, PWM:61
19:17:47.841 -> Water_Temp.:106.25, PWM:61
19:17:48.636 -> Water_Temp.:106.25, PWM:61
19:17:49.393 -> Water_Temp.:106.25, PWM:61
19:17:50.187 -> Water_Temp.:106.25, PWM:61
19:17:50.978 -> Water_Temp.:106.25, PWM:61
19:17:51.737 -> Water_Temp.:106.25, PWM:61
19:17:52.529 -> Water_Temp.:106.25, PWM:61
19:17:53.286 -> Water_Temp.:106.25, PWM:61
19:17:54.079 -> Water_Temp.:107.82, PWM:61
19:17:54.876 -> Water_Temp.:109.85, PWM:61
19:17:55.634 -> Water_Temp.:107.49, PWM:61
19:17:56.430 -> Water_Temp.:107.04, PWM:61
19:17:57.187 -> Water_Temp.:106.92, PWM:61
19:17:57.978 -> Water_Temp.:106.81, PWM:61
19:17:58.774 -> Water_Temp.:106.70, PWM:61
19:17:59.535 -> Water_Temp.:106.59, PWM:61
19:18:00.328 -> Water_Temp.:106.59, PWM:61
19:18:01.119 -> Water_Temp.:106.59, PWM:61
19:18:01.880 -> Water_Temp.:106.47, PWM:61
19:18:02.676 -> Water_Temp.:106.47, PWM:61
19:18:03.437 -> Water_Temp.:106.47, PWM:61
19:18:04.225 -> Water_Temp.:106.47, PWM:61
19:18:05.778 -> Water_Temp.:106.47, PWM:61
19:18:06.533 -> Water_Temp.:106.47, PWM:61
19:18:07.326 -> Water_Temp.:106.47, PWM:61
19:18:08.122 -> Water_Temp.:106.47, PWM:61
19:18:08.886 -> Water_Temp.:106.47, PWM:61
19:18:09.677 -> Water_Temp.:106.47, PWM:61
19:18:10.440 -> Water_Temp.:106.47, PWM:61
19:18:11.234 -> Water_Temp.:106.47, PWM:61
19:18:12.026 -> Water_Temp.:106.47, PWM:61
19:18:12.790 -> Water_Temp.:108.05, PWM:61
19:18:13.582 -> Water_Temp.:110.07, PWM:61
19:18:14.342 -> Water_Temp.:107.71, PWM:61
19:18:15.133 -> Water_Temp.:107.26, PWM:61
19:18:15.895 -> Water_Temp.:107.15, PWM:61
19:18:16.688 -> Water_Temp.:106.92, PWM:61
19:18:17.477 -> Water_Temp.:106.81, PWM:61
19:18:18.241 -> Water_Temp.:106.81, PWM:61
19:18:19.036 -> Water_Temp.:106.70, PWM:61
19:18:19.794 -> Water_Temp.:106.70, PWM:61
19:18:20.588 -> Water_Temp.:106.70, PWM:61
19:18:21.383 -> Water_Temp.:106.59, PWM:61
19:18:22.139 -> Water_Temp.:106.59, PWM:61
19:18:22.931 -> Water_Temp.:106.59, PWM:61
19:18:23.691 -> Water_Temp.:106.59, PWM:61
19:18:24.485 -> Water_Temp.:106.59, PWM:61
19:18:25.278 -> Water_Temp.:106.47, PWM:61
19:18:26.041 -> Water_Temp.:106.47, PWM:61
19:18:26.836 -> Water_Temp.:106.47, PWM:61
19:18:27.595 -> Water_Temp.:106.47, PWM:61
19:18:28.391 -> Water_Temp.:106.47, PWM:61
19:18:29.183 -> Water_Temp.:106.47, PWM:61
19:18:29.943 -> Water_Temp.:106.47, PWM:61
19:18:30.737 -> Water_Temp.:106.47, PWM:61
19:18:31.497 -> Water_Temp.:106.47, PWM:61
19:18:33.049 -> Water_Temp.:106.47, PWM:61
19:18:33.842 -> Water_Temp.:106.47, PWM:61
19:18:34.602 -> Water_Temp.:106.47, PWM:61
19:18:35.402 -> Water_Temp.:106.47, PWM:61
19:18:36.155 -> Water_Temp.:106.47, PWM:61
19:18:36.949 -> Water_Temp.:106.47, PWM:61
19:18:38.495 -> Water_Temp.:106.47, PWM:61
19:18:39.285 -> Water_Temp.:106.47, PWM:61
19:18:40.078 -> Water_Temp.:106.59, PWM:61
19:18:40.838 -> Water_Temp.:106.47, PWM:61
19:18:41.629 -> Water_Temp.:106.47, PWM:61
19:18:42.389 -> Water_Temp.:106.47, PWM:61
19:18:43.182 -> Water_Temp.:106.47, PWM:61
19:18:43.978 -> Water_Temp.:106.47, PWM:61
19:18:44.733 -> Water_Temp.:106.47, PWM:61
19:18:45.528 -> Water_Temp.:106.47, PWM:61
19:18:46.295 -> Water_Temp.:106.47, PWM:61
19:18:47.085 -> Water_Temp.:106.47, PWM:61
19:18:47.880 -> Water_Temp.:106.47, PWM:61
19:18:48.639 -> Water_Temp.:106.47, PWM:61
19:18:49.429 -> Water_Temp.:106.47, PWM:61
19:18:50.218 -> Water_Temp.:106.47, PWM:61
19:18:50.981 -> Water_Temp.:106.47, PWM:61
19:18:51.775 -> Water_Temp.:106.47, PWM:61
19:18:52.528 -> Water_Temp.:106.47, PWM:61
19:18:53.324 -> Water_Temp.:106.47, PWM:61
19:18:54.119 -> Water_Temp.:106.36, PWM:61
19:18:54.876 -> Water_Temp.:106.36, PWM:61
19:18:55.672 -> Water_Temp.:106.36, PWM:61
19:18:56.432 -> Water_Temp.:106.36, PWM:61
19:18:57.227 -> Water_Temp.:106.36, PWM:61
19:18:58.018 -> Water_Temp.:106.36, PWM:61
19:18:58.774 -> Water_Temp.:106.47, PWM:61
19:18:59.569 -> Water_Temp.:106.36, PWM:61
19:19:00.328 -> Water_Temp.:106.36, PWM:61
19:19:01.122 -> Water_Temp.:106.36, PWM:61
19:19:01.913 -> Water_Temp.:106.36, PWM:61
19:19:02.671 -> Water_Temp.:106.36, PWM:61
19:19:03.463 -> Water_Temp.:106.36, PWM:61
19:19:04.263 -> Water_Temp.:106.36, PWM:61
19:19:05.025 -> Water_Temp.:106.36, PWM:61
19:19:05.818 -> Water_Temp.:106.36, PWM:61
19:19:06.580 -> Water_Temp.:106.36, PWM:61
19:19:07.373 -> Water_Temp.:106.47, PWM:61
19:19:08.131 -> Water_Temp.:106.47, PWM:61
19:19:08.924 -> Water_Temp.:106.47, PWM:61
19:19:09.715 -> Water_Temp.:106.47, PWM:61
19:19:11.269 -> Water_Temp.:106.47, PWM:61
19:19:12.031 -> Water_Temp.:106.47, PWM:61
19:19:12.822 -> Water_Temp.:106.47, PWM:61
19:19:13.614 -> Water_Temp.:106.47, PWM:61
19:19:14.370 -> Water_Temp.:106.47, PWM:61
19:19:15.160 -> Water_Temp.:106.47, PWM:61
19:19:15.953 -> Water_Temp.:106.47, PWM:61
19:19:16.711 -> Water_Temp.:106.47, PWM:61
19:19:17.503 -> Water_Temp.:106.47, PWM:61
19:19:18.298 -> Water_Temp.:106.47, PWM:61
19:19:19.055 -> Water_Temp.:106.47, PWM:61
19:19:19.850 -> Water_Temp.:106.47, PWM:61
19:19:20.610 -> Water_Temp.:106.47, PWM:61
19:19:21.405 -> Water_Temp.:106.47, PWM:61
19:19:22.199 -> Water_Temp.:106.47, PWM:61
19:19:22.957 -> Water_Temp.:106.47, PWM:61
19:19:23.750 -> Water_Temp.:106.47, PWM:61
19:19:24.508 -> Water_Temp.:106.47, PWM:61
19:19:25.296 -> Water_Temp.:106.47, PWM:61
19:19:26.090 -> Water_Temp.:106.47, PWM:61
19:19:26.853 -> Water_Temp.:106.47, PWM:61
19:19:27.649 -> Water_Temp.:106.47, PWM:61
19:19:28.410 -> Water_Temp.:106.47, PWM:61
19:19:29.203 -> Water_Temp.:106.59, PWM:61
19:19:29.998 -> Water_Temp.:106.59, PWM:61
19:19:30.759 -> Water_Temp.:106.59, PWM:61
19:19:31.553 -> Water_Temp.:106.59, PWM:61
19:19:32.309 -> Water_Temp.:106.59, PWM:61
19:19:33.103 -> Water_Temp.:106.59, PWM:61
19:19:33.894 -> Water_Temp.:106.59, PWM:61
19:19:34.652 -> Water_Temp.:106.59, PWM:61
19:19:35.446 -> Water_Temp.:106.59, PWM:61
19:19:36.238 -> Water_Temp.:106.59, PWM:61
19:19:36.995 -> Water_Temp.:106.59, PWM:61
19:19:37.790 -> Water_Temp.:106.59, PWM:61
19:19:38.549 -> Water_Temp.:106.59, PWM:61
19:19:39.341 -> Water_Temp.:106.59, PWM:61
19:19:40.897 -> Water_Temp.:106.59, PWM:61
19:19:41.692 -> Water_Temp.:106.59, PWM:61
19:19:42.452 -> Water_Temp.:106.59, PWM:61
19:19:43.240 -> Water_Temp.:106.59, PWM:61
19:19:44.033 -> Water_Temp.:106.47, PWM:61
19:19:44.790 -> Water_Temp.:106.47, PWM:0
19:19:45.588 -> Water_Temp.:106.47, PWM:0
19:19:46.349 -> Water_Temp.:106.47, PWM:0
19:19:47.143 -> Water_Temp.:106.47, PWM:0
19:19:47.936 -> Water_Temp.:106.47, PWM:0

## RAW_CAL_MLX-7
16:41:04.485 -> Average Cal sensor Temp.: 87.44 Ambient Temp.: 93.69 F
16:41:06.032 -> Average Cal sensor Temp.: 87.61 Ambient Temp.: 93.72 F
16:41:07.594 -> Average Cal sensor Temp.: 87.71 Ambient Temp.: 93.69 F
16:41:09.164 -> Average Cal sensor Temp.: 87.68 Ambient Temp.: 93.69 F
16:41:10.705 -> Average Cal sensor Temp.: 87.57 Ambient Temp.: 93.72 F
16:41:12.272 -> Average Cal sensor Temp.: 87.67 Ambient Temp.: 93.61 F
16:41:13.808 -> Average Cal sensor Temp.: 87.63 Ambient Temp.: 93.72 F
16:41:15.379 -> Average Cal sensor Temp.: 87.52 Ambient Temp.: 93.61 F
16:41:16.918 -> Average Cal sensor Temp.: 87.30 Ambient Temp.: 93.72 F
16:41:18.486 -> Average Cal sensor Temp.: 87.47 Ambient Temp.: 93.72 F
16:41:20.036 -> Average Cal sensor Temp.: 87.44 Ambient Temp.: 93.61 F
16:41:21.586 -> Average Cal sensor Temp.: 87.46 Ambient Temp.: 93.72 F
16:41:23.139 -> Average Cal sensor Temp.: 87.59 Ambient Temp.: 93.72 F
16:41:24.696 -> Average Cal sensor Temp.: 87.52 Ambient Temp.: 93.61 F
16:41:26.279 -> Average Cal sensor Temp.: 87.58 Ambient Temp.: 93.72 F
16:41:27.830 -> Average Cal sensor Temp.: 87.66 Ambient Temp.: 93.69 F
16:41:29.389 -> Average Cal sensor Temp.: 87.61 Ambient Temp.: 93.72 F
16:41:30.943 -> Average Cal sensor Temp.: 87.58 Ambient Temp.: 93.69 F
16:41:32.498 -> Average Cal sensor Temp.: 87.75 Ambient Temp.: 93.69 F
16:41:34.047 -> Average Cal sensor Temp.: 87.77 Ambient Temp.: 93.72 F
16:41:35.603 -> Average Cal sensor Temp.: 87.84 Ambient Temp.: 93.69 F
16:41:37.163 -> Average Cal sensor Temp.: 88.00 Ambient Temp.: 93.76 F
16:41:38.715 -> Average Cal sensor Temp.: 88.08 Ambient Temp.: 93.76 F
16:41:40.262 -> Average Cal sensor Temp.: 87.98 Ambient Temp.: 93.76 F
16:41:41.818 -> Average Cal sensor Temp.: 87.95 Ambient Temp.: 93.72 F
16:41:43.371 -> Average Cal sensor Temp.: 88.02 Ambient Temp.: 93.72 F
16:41:44.926 -> Average Cal sensor Temp.: 87.97 Ambient Temp.: 93.72 F
16:41:46.510 -> Average Cal sensor Temp.: 87.75 Ambient Temp.: 93.72 F
16:41:48.057 -> Average Cal sensor Temp.: 87.93 Ambient Temp.: 93.72 F
16:41:49.610 -> Average Cal sensor Temp.: 87.91 Ambient Temp.: 93.72 F
16:41:51.166 -> Average Cal sensor Temp.: 87.86 Ambient Temp.: 93.72 F
16:41:52.719 -> Average Cal sensor Temp.: 87.77 Ambient Temp.: 93.72 F
16:41:54.267 -> Average Cal sensor Temp.: 87.77 Ambient Temp.: 93.72 F
16:41:55.816 -> Average Cal sensor Temp.: 87.81 Ambient Temp.: 93.76 F
16:41:57.368 -> Average Cal sensor Temp.: 87.82 Ambient Temp.: 93.76 F
16:41:58.948 -> Average Cal sensor Temp.: 87.90 Ambient Temp.: 93.72 F
16:42:00.497 -> Average Cal sensor Temp.: 87.90 Ambient Temp.: 93.76 F
16:42:02.054 -> Average Cal sensor Temp.: 88.02 Ambient Temp.: 93.76 F
16:42:03.606 -> Average Cal sensor Temp.: 87.63 Ambient Temp.: 93.72 F
16:42:05.158 -> Average Cal sensor Temp.: 88.18 Ambient Temp.: 93.79 F
16:42:06.712 -> Average Cal sensor Temp.: 88.15 Ambient Temp.: 93.76 F
16:42:08.261 -> Average Cal sensor Temp.: 88.06 Ambient Temp.: 93.76 F
16:42:09.848 -> Average Cal sensor Temp.: 88.16 Ambient Temp.: 93.76 F
16:42:11.398 -> Average Cal sensor Temp.: 88.21 Ambient Temp.: 93.79 F
16:42:12.939 -> Average Cal sensor Temp.: 88.06 Ambient Temp.: 93.76 F
16:42:14.488 -> Average Cal sensor Temp.: 88.14 Ambient Temp.: 93.79 F
16:42:16.074 -> Average Cal sensor Temp.: 88.22 Ambient Temp.: 93.79 F
16:42:17.620 -> Average Cal sensor Temp.: 88.20 Ambient Temp.: 93.72 F
16:42:19.172 -> Average Cal sensor Temp.: 88.24 Ambient Temp.: 93.79 F
16:42:20.724 -> Average Cal sensor Temp.: 88.20 Ambient Temp.: 93.76 F
16:42:22.278 -> Average Cal sensor Temp.: 88.10 Ambient Temp.: 93.79 F
16:42:23.827 -> Average Cal sensor Temp.: 88.09 Ambient Temp.: 93.79 F
16:42:25.415 -> Average Cal sensor Temp.: 87.94 Ambient Temp.: 93.83 F
16:42:26.959 -> Average Cal sensor Temp.: 87.74 Ambient Temp.: 93.76 F
16:42:28.520 -> Average Cal sensor Temp.: 87.91 Ambient Temp.: 93.79 F
16:42:30.075 -> Average Cal sensor Temp.: 87.95 Ambient Temp.: 93.76 F
16:42:31.629 -> Average Cal sensor Temp.: 88.07 Ambient Temp.: 93.83 F
16:42:33.180 -> Average Cal sensor Temp.: 87.98 Ambient Temp.: 93.79 F
16:42:34.732 -> Average Cal sensor Temp.: 87.76 Ambient Temp.: 93.79 F
16:42:36.290 -> Average Cal sensor Temp.: 87.97 Ambient Temp.: 93.79 F
16:42:37.840 -> Average Cal sensor Temp.: 87.89 Ambient Temp.: 93.76 F
16:42:39.392 -> Average Cal sensor Temp.: 88.08 Ambient Temp.: 93.76 F
16:42:40.951 -> Average Cal sensor Temp.: 87.86 Ambient Temp.: 93.79 F
16:42:42.508 -> Average Cal sensor Temp.: 88.02 Ambient Temp.: 93.79 F
16:42:44.092 -> Average Cal sensor Temp.: 88.15 Ambient Temp.: 93.76 F
16:42:45.641 -> Average Cal sensor Temp.: 88.25 Ambient Temp.: 93.83 F
16:42:47.192 -> Average Cal sensor Temp.: 88.25 Ambient Temp.: 93.79 F
16:42:48.738 -> Average Cal sensor Temp.: 88.32 Ambient Temp.: 93.83 F
16:42:50.288 -> Average Cal sensor Temp.: 88.25 Ambient Temp.: 93.83 F
16:42:51.871 -> Average Cal sensor Temp.: 88.29 Ambient Temp.: 93.83 F
16:42:53.424 -> Average Cal sensor Temp.: 88.25 Ambient Temp.: 93.83 F
16:42:54.980 -> Average Cal sensor Temp.: 88.28 Ambient Temp.: 93.79 F
16:42:56.535 -> Average Cal sensor Temp.: 88.91 Ambient Temp.: 93.87 F
16:42:58.086 -> Average Cal sensor Temp.: 89.28 Ambient Temp.: 93.87 F
16:42:59.639 -> Average Cal sensor Temp.: 89.18 Ambient Temp.: 93.87 F
16:43:01.194 -> Average Cal sensor Temp.: 88.83 Ambient Temp.: 93.79 F
16:43:02.745 -> Average Cal sensor Temp.: 88.69 Ambient Temp.: 93.79 F
16:43:04.322 -> Average Cal sensor Temp.: 88.46 Ambient Temp.: 93.83 F
16:43:05.879 -> Average Cal sensor Temp.: 88.36 Ambient Temp.: 93.87 F
16:43:07.429 -> Average Cal sensor Temp.: 88.23 Ambient Temp.: 93.90 F
16:43:08.986 -> Average Cal sensor Temp.: 88.16 Ambient Temp.: 93.83 F
16:43:10.540 -> Average Cal sensor Temp.: 88.20 Ambient Temp.: 93.83 F
16:43:12.089 -> Average Cal sensor Temp.: 88.02 Ambient Temp.: 93.83 F
16:43:13.647 -> Average Cal sensor Temp.: 87.99 Ambient Temp.: 93.87 F
16:43:15.202 -> Average Cal sensor Temp.: 88.07 Ambient Temp.: 93.83 F
16:43:16.757 -> Average Cal sensor Temp.: 88.01 Ambient Temp.: 93.87 F
16:43:18.313 -> Average Cal sensor Temp.: 88.19 Ambient Temp.: 93.79 F
16:43:19.868 -> Average Cal sensor Temp.: 88.29 Ambient Temp.: 93.90 F
16:43:21.450 -> Average Cal sensor Temp.: 88.09 Ambient Temp.: 93.83 F
16:43:23.002 -> Average Cal sensor Temp.: 88.11 Ambient Temp.: 93.87 F
16:43:24.556 -> Average Cal sensor Temp.: 88.14 Ambient Temp.: 93.83 F
16:43:26.116 -> Average Cal sensor Temp.: 88.16 Ambient Temp.: 93.83 F
16:43:27.672 -> Average Cal sensor Temp.: 88.14 Ambient Temp.: 93.87 F
16:43:29.230 -> Average Cal sensor Temp.: 88.04 Ambient Temp.: 93.83 F
16:43:30.782 -> Average Cal sensor Temp.: 88.12 Ambient Temp.: 93.79 F
16:43:32.336 -> Average Cal sensor Temp.: 88.12 Ambient Temp.: 93.83 F
16:43:33.889 -> Average Cal sensor Temp.: 88.30 Ambient Temp.: 93.83 F
16:43:35.441 -> Average Cal sensor Temp.: 88.20 Ambient Temp.: 93.87 F
16:43:36.992 -> Average Cal sensor Temp.: 88.20 Ambient Temp.: 93.83 F
16:43:38.576 -> Average Cal sensor Temp.: 88.28 Ambient Temp.: 93.83 F
16:43:40.128 -> Average Cal sensor Temp.: 88.27 Ambient Temp.: 93.90 F
16:43:41.675 -> Average Cal sensor Temp.: 88.26 Ambient Temp.: 93.83 F
16:43:43.227 -> Average Cal sensor Temp.: 88.43 Ambient Temp.: 93.83 F
16:43:44.779 -> Average Cal sensor Temp.: 88.39 Ambient Temp.: 93.87 F
16:43:46.335 -> Average Cal sensor Temp.: 88.37 Ambient Temp.: 93.90 F
16:43:47.919 -> Average Cal sensor Temp.: 88.32 Ambient Temp.: 93.90 F
16:43:49.478 -> Average Cal sensor Temp.: 88.32 Ambient Temp.: 93.87 F
16:43:51.032 -> Average Cal sensor Temp.: 88.35 Ambient Temp.: 93.79 F
16:43:52.578 -> Average Cal sensor Temp.: 88.34 Ambient Temp.: 93.83 F
16:43:54.129 -> Average Cal sensor Temp.: 88.43 Ambient Temp.: 93.90 F
16:43:55.677 -> Average Cal sensor Temp.: 88.45 Ambient Temp.: 93.87 F
16:43:57.236 -> Average Cal sensor Temp.: 88.64 Ambient Temp.: 93.87 F
16:43:58.794 -> Average Cal sensor Temp.: 88.51 Ambient Temp.: 93.87 F
16:44:00.375 -> Average Cal sensor Temp.: 88.68 Ambient Temp.: 93.87 F
16:44:01.921 -> Average Cal sensor Temp.: 88.69 Ambient Temp.: 93.87 F
16:44:03.478 -> Average Cal sensor Temp.: 88.64 Ambient Temp.: 93.90 F
16:44:05.031 -> Average Cal sensor Temp.: 88.56 Ambient Temp.: 93.94 F
16:44:06.585 -> Average Cal sensor Temp.: 88.70 Ambient Temp.: 93.90 F
16:44:08.168 -> Average Cal sensor Temp.: 88.60 Ambient Temp.: 93.87 F
16:44:09.722 -> Average Cal sensor Temp.: 88.63 Ambient Temp.: 93.94 F
16:44:11.270 -> Average Cal sensor Temp.: 88.57 Ambient Temp.: 93.94 F
16:44:12.821 -> Average Cal sensor Temp.: 88.53 Ambient Temp.: 93.90 F
16:44:14.367 -> Average Cal sensor Temp.: 88.50 Ambient Temp.: 93.90 F
16:44:15.926 -> Average Cal sensor Temp.: 88.65 Ambient Temp.: 93.94 F
16:44:17.480 -> Average Cal sensor Temp.: 88.78 Ambient Temp.: 93.90 F
16:44:19.064 -> Average Cal sensor Temp.: 88.97 Ambient Temp.: 93.90 F
16:44:20.612 -> Average Cal sensor Temp.: 88.77 Ambient Temp.: 93.94 F
16:44:22.170 -> Average Cal sensor Temp.: 88.81 Ambient Temp.: 93.94 F
16:44:23.723 -> Average Cal sensor Temp.: 88.83 Ambient Temp.: 93.94 F
16:44:25.273 -> Average Cal sensor Temp.: 88.81 Ambient Temp.: 93.94 F
16:44:26.825 -> Average Cal sensor Temp.: 88.75 Ambient Temp.: 93.97 F
16:44:28.383 -> Average Cal sensor Temp.: 88.79 Ambient Temp.: 93.90 F
16:44:29.936 -> Average Cal sensor Temp.: 88.73 Ambient Temp.: 93.94 F
16:44:31.522 -> Average Cal sensor Temp.: 88.88 Ambient Temp.: 93.94 F
16:44:33.077 -> Average Cal sensor Temp.: 88.83 Ambient Temp.: 93.94 F
16:44:34.634 -> Average Cal sensor Temp.: 88.94 Ambient Temp.: 93.94 F
16:44:36.190 -> Average Cal sensor Temp.: 88.91 Ambient Temp.: 93.90 F
16:44:37.738 -> Average Cal sensor Temp.: 88.89 Ambient Temp.: 93.97 F
16:44:39.293 -> Average Cal sensor Temp.: 88.70 Ambient Temp.: 93.97 F
16:44:40.848 -> Average Cal sensor Temp.: 88.80 Ambient Temp.: 94.01 F
16:44:42.405 -> Average Cal sensor Temp.: 88.85 Ambient Temp.: 93.94 F
16:44:43.964 -> Average Cal sensor Temp.: 88.96 Ambient Temp.: 93.94 F
16:44:45.523 -> Average Cal sensor Temp.: 88.95 Ambient Temp.: 94.01 F
16:44:47.076 -> Average Cal sensor Temp.: 88.86 Ambient Temp.: 94.01 F
16:44:48.630 -> Average Cal sensor Temp.: 89.04 Ambient Temp.: 93.97 F
16:44:50.218 -> Average Cal sensor Temp.: 89.03 Ambient Temp.: 93.97 F
16:44:51.766 -> Average Cal sensor Temp.: 89.03 Ambient Temp.: 93.97 F
16:44:53.323 -> Average Cal sensor Temp.: 89.06 Ambient Temp.: 93.97 F
16:44:54.875 -> Average Cal sensor Temp.: 88.96 Ambient Temp.: 93.94 F
16:44:56.435 -> Average Cal sensor Temp.: 88.99 Ambient Temp.: 93.97 F
16:44:57.987 -> Average Cal sensor Temp.: 89.04 Ambient Temp.: 93.97 F
16:44:59.540 -> Average Cal sensor Temp.: 89.10 Ambient Temp.: 93.97 F
16:45:01.099 -> Average Cal sensor Temp.: 89.01 Ambient Temp.: 94.08 F
16:45:02.653 -> Average Cal sensor Temp.: 88.97 Ambient Temp.: 93.94 F
16:45:04.206 -> Average Cal sensor Temp.: 89.03 Ambient Temp.: 93.97 F
16:45:05.790 -> Average Cal sensor Temp.: 89.19 Ambient Temp.: 94.01 F
16:45:07.347 -> Average Cal sensor Temp.: 89.13 Ambient Temp.: 94.01 F
16:45:08.900 -> Average Cal sensor Temp.: 89.02 Ambient Temp.: 94.01 F
16:45:10.450 -> Average Cal sensor Temp.: 89.01 Ambient Temp.: 94.01 F
16:45:12.003 -> Average Cal sensor Temp.: 89.03 Ambient Temp.: 94.01 F
16:45:13.554 -> Average Cal sensor Temp.: 89.01 Ambient Temp.: 94.01 F
16:45:15.108 -> Average Cal sensor Temp.: 89.08 Ambient Temp.: 94.08 F
16:45:16.664 -> Average Cal sensor Temp.: 89.14 Ambient Temp.: 94.01 F
16:45:18.216 -> Average Cal sensor Temp.: 89.15 Ambient Temp.: 94.12 F
16:45:19.775 -> Average Cal sensor Temp.: 89.19 Ambient Temp.: 94.01 F
16:45:21.332 -> Average Cal sensor Temp.: 89.08 Ambient Temp.: 94.01 F
16:45:22.891 -> Average Cal sensor Temp.: 89.07 Ambient Temp.: 94.01 F
16:45:24.446 -> Average Cal sensor Temp.: 89.07 Ambient Temp.: 94.01 F
16:45:25.998 -> Average Cal sensor Temp.: 89.13 Ambient Temp.: 94.12 F
16:45:27.555 -> Average Cal sensor Temp.: 89.11 Ambient Temp.: 94.08 F
16:45:29.142 -> Average Cal sensor Temp.: 89.15 Ambient Temp.: 94.08 F
16:45:30.697 -> Average Cal sensor Temp.: 89.14 Ambient Temp.: 94.08 F
16:45:32.249 -> Average Cal sensor Temp.: 89.01 Ambient Temp.: 94.01 F
16:45:33.803 -> Average Cal sensor Temp.: 89.20 Ambient Temp.: 94.08 F
16:45:35.350 -> Average Cal sensor Temp.: 89.22 Ambient Temp.: 94.08 F
16:45:36.896 -> Average Cal sensor Temp.: 89.33 Ambient Temp.: 94.08 F
16:45:38.482 -> Average Cal sensor Temp.: 89.46 Ambient Temp.: 94.12 F
16:45:40.034 -> Average Cal sensor Temp.: 89.59 Ambient Temp.: 94.12 F
16:45:41.584 -> Average Cal sensor Temp.: 89.57 Ambient Temp.: 94.08 F
16:45:43.140 -> Average Cal sensor Temp.: 89.56 Ambient Temp.: 94.08 F
16:45:44.694 -> Average Cal sensor Temp.: 89.51 Ambient Temp.: 94.08 F
16:45:46.248 -> Average Cal sensor Temp.: 89.32 Ambient Temp.: 94.12 F
16:45:47.805 -> Average Cal sensor Temp.: 89.27 Ambient Temp.: 94.08 F
16:45:49.358 -> Average Cal sensor Temp.: 89.42 Ambient Temp.: 94.08 F
16:45:50.908 -> Average Cal sensor Temp.: 89.51 Ambient Temp.: 94.12 F
16:45:52.495 -> Average Cal sensor Temp.: 89.68 Ambient Temp.: 94.12 F
16:45:54.050 -> Average Cal sensor Temp.: 89.65 Ambient Temp.: 94.08 F
16:45:55.597 -> Average Cal sensor Temp.: 89.66 Ambient Temp.: 94.08 F
16:45:57.152 -> Average Cal sensor Temp.: 89.77 Ambient Temp.: 94.08 F
16:45:58.703 -> Average Cal sensor Temp.: 89.72 Ambient Temp.: 94.12 F
16:46:00.257 -> Average Cal sensor Temp.: 89.72 Ambient Temp.: 94.12 F
16:46:01.835 -> Average Cal sensor Temp.: 89.60 Ambient Temp.: 94.12 F
16:46:03.391 -> Average Cal sensor Temp.: 89.53 Ambient Temp.: 94.12 F
16:46:04.947 -> Average Cal sensor Temp.: 89.58 Ambient Temp.: 94.08 F
16:46:06.498 -> Average Cal sensor Temp.: 89.65 Ambient Temp.: 94.15 F
16:46:08.054 -> Average Cal sensor Temp.: 89.60 Ambient Temp.: 94.12 F
16:46:09.603 -> Average Cal sensor Temp.: 89.67 Ambient Temp.: 94.15 F
16:46:11.153 -> Average Cal sensor Temp.: 89.73 Ambient Temp.: 94.19 F
16:46:12.708 -> Average Cal sensor Temp.: 89.54 Ambient Temp.: 94.15 F
16:46:14.259 -> Average Cal sensor Temp.: 89.55 Ambient Temp.: 94.19 F
16:46:15.847 -> Average Cal sensor Temp.: 89.54 Ambient Temp.: 94.15 F
16:46:17.400 -> Average Cal sensor Temp.: 89.51 Ambient Temp.: 94.15 F
16:46:18.954 -> Average Cal sensor Temp.: 89.50 Ambient Temp.: 94.19 F
16:46:20.511 -> Average Cal sensor Temp.: 89.56 Ambient Temp.: 94.15 F
16:46:22.061 -> Average Cal sensor Temp.: 89.60 Ambient Temp.: 94.15 F
16:46:23.612 -> Average Cal sensor Temp.: 89.70 Ambient Temp.: 94.19 F
16:46:25.166 -> Average Cal sensor Temp.: 89.74 Ambient Temp.: 94.19 F
16:46:26.727 -> Average Cal sensor Temp.: 89.76 Ambient Temp.: 94.15 F
16:46:28.282 -> Average Cal sensor Temp.: 89.60 Ambient Temp.: 94.15 F
16:46:29.846 -> Average Cal sensor Temp.: 89.60 Ambient Temp.: 94.15 F
16:46:31.394 -> Average Cal sensor Temp.: 89.66 Ambient Temp.: 94.19 F
16:46:32.949 -> Average Cal sensor Temp.: 89.75 Ambient Temp.: 94.15 F
16:46:34.503 -> Average Cal sensor Temp.: 89.74 Ambient Temp.: 94.19 F
16:46:36.056 -> Average Cal sensor Temp.: 89.74 Ambient Temp.: 94.19 F
16:46:37.614 -> Average Cal sensor Temp.: 90.10 Ambient Temp.: 94.19 F
16:46:39.167 -> Average Cal sensor Temp.: 90.13 Ambient Temp.: 94.23 F
16:46:40.721 -> Average Cal sensor Temp.: 90.07 Ambient Temp.: 94.19 F
16:46:42.306 -> Average Cal sensor Temp.: 90.03 Ambient Temp.: 94.23 F
16:46:43.864 -> Average Cal sensor Temp.: 89.75 Ambient Temp.: 94.19 F
16:46:45.418 -> Average Cal sensor Temp.: 89.82 Ambient Temp.: 94.19 F
16:46:46.975 -> Average Cal sensor Temp.: 89.68 Ambient Temp.: 94.23 F
16:46:48.527 -> Average Cal sensor Temp.: 89.64 Ambient Temp.: 94.19 F
16:46:50.084 -> Average Cal sensor Temp.: 89.80 Ambient Temp.: 94.23 F
16:46:51.635 -> Average Cal sensor Temp.: 89.75 Ambient Temp.: 94.26 F
16:46:53.185 -> Average Cal sensor Temp.: 89.65 Ambient Temp.: 94.26 F
16:46:54.744 -> Average Cal sensor Temp.: 89.73 Ambient Temp.: 94.26 F
16:46:56.298 -> Average Cal sensor Temp.: 89.68 Ambient Temp.: 94.23 F
16:46:57.847 -> Average Cal sensor Temp.: 89.77 Ambient Temp.: 94.30 F
16:46:59.404 -> Average Cal sensor Temp.: 89.87 Ambient Temp.: 94.26 F
16:47:00.991 -> Average Cal sensor Temp.: 89.97 Ambient Temp.: 94.26 F
16:47:02.547 -> Average Cal sensor Temp.: 90.04 Ambient Temp.: 94.26 F
16:47:04.093 -> Average Cal sensor Temp.: 89.80 Ambient Temp.: 94.30 F
16:47:05.647 -> Average Cal sensor Temp.: 89.82 Ambient Temp.: 94.26 F
16:47:07.201 -> Average Cal sensor Temp.: 89.99 Ambient Temp.: 94.23 F
16:47:08.753 -> Average Cal sensor Temp.: 89.93 Ambient Temp.: 94.26 F
16:47:10.314 -> Average Cal sensor Temp.: 90.04 Ambient Temp.: 94.30 F
16:47:11.867 -> Average Cal sensor Temp.: 90.23 Ambient Temp.: 94.26 F
16:47:13.422 -> Average Cal sensor Temp.: 90.17 Ambient Temp.: 94.26 F
16:47:14.977 -> Average Cal sensor Temp.: 90.12 Ambient Temp.: 94.30 F
16:47:16.566 -> Average Cal sensor Temp.: 90.01 Ambient Temp.: 94.30 F
16:47:18.123 -> Average Cal sensor Temp.: 90.20 Ambient Temp.: 94.30 F
16:47:19.675 -> Average Cal sensor Temp.: 90.03 Ambient Temp.: 94.30 F
16:47:21.229 -> Average Cal sensor Temp.: 90.06 Ambient Temp.: 94.26 F
16:47:22.779 -> Average Cal sensor Temp.: 90.12 Ambient Temp.: 94.30 F
16:47:24.334 -> Average Cal sensor Temp.: 90.00 Ambient Temp.: 94.30 F
16:47:25.890 -> Average Cal sensor Temp.: 90.04 Ambient Temp.: 94.30 F
16:47:27.447 -> Average Cal sensor Temp.: 90.07 Ambient Temp.: 94.30 F
16:47:29.006 -> Average Cal sensor Temp.: 90.19 Ambient Temp.: 94.30 F
16:47:30.566 -> Average Cal sensor Temp.: 90.03 Ambient Temp.: 94.33 F
16:47:32.122 -> Average Cal sensor Temp.: 89.96 Ambient Temp.: 94.33 F
16:47:33.674 -> Average Cal sensor Temp.: 89.82 Ambient Temp.: 94.26 F
16:47:35.226 -> Average Cal sensor Temp.: 89.92 Ambient Temp.: 94.30 F
16:47:36.779 -> Average Cal sensor Temp.: 89.91 Ambient Temp.: 94.30 F
16:47:38.337 -> Average Cal sensor Temp.: 89.99 Ambient Temp.: 94.30 F
16:47:39.921 -> Average Cal sensor Temp.: 89.96 Ambient Temp.: 94.33 F
16:47:41.473 -> Average Cal sensor Temp.: 90.06 Ambient Temp.: 94.37 F
16:47:43.034 -> Average Cal sensor Temp.: 89.90 Ambient Temp.: 94.33 F
16:47:44.585 -> Average Cal sensor Temp.: 89.87 Ambient Temp.: 94.33 F
16:47:46.136 -> Average Cal sensor Temp.: 89.95 Ambient Temp.: 94.30 F
16:47:47.689 -> Average Cal sensor Temp.: 90.07 Ambient Temp.: 94.33 F
16:47:49.247 -> Average Cal sensor Temp.: 89.87 Ambient Temp.: 94.30 F
16:47:50.796 -> Average Cal sensor Temp.: 89.98 Ambient Temp.: 94.37 F
16:47:52.377 -> Average Cal sensor Temp.: 90.01 Ambient Temp.: 94.30 F
16:47:53.929 -> Average Cal sensor Temp.: 90.14 Ambient Temp.: 94.33 F
16:47:55.476 -> Average Cal sensor Temp.: 90.15 Ambient Temp.: 94.33 F
16:47:57.033 -> Average Cal sensor Temp.: 90.08 Ambient Temp.: 94.33 F
16:47:58.586 -> Average Cal sensor Temp.: 90.32 Ambient Temp.: 94.33 F
16:48:00.136 -> Average Cal sensor Temp.: 90.23 Ambient Temp.: 94.37 F
16:48:01.691 -> Average Cal sensor Temp.: 90.39 Ambient Temp.: 94.33 F
16:48:03.273 -> Average Cal sensor Temp.: 90.26 Ambient Temp.: 94.33 F
16:48:04.829 -> Average Cal sensor Temp.: 90.25 Ambient Temp.: 94.30 F
16:48:06.382 -> Average Cal sensor Temp.: 90.18 Ambient Temp.: 94.37 F
16:48:07.934 -> Average Cal sensor Temp.: 90.17 Ambient Temp.: 94.37 F
16:48:09.492 -> Average Cal sensor Temp.: 90.18 Ambient Temp.: 94.37 F
16:48:11.049 -> Average Cal sensor Temp.: 90.24 Ambient Temp.: 94.33 F
16:48:12.610 -> Average Cal sensor Temp.: 90.27 Ambient Temp.: 94.37 F
16:48:14.164 -> Average Cal sensor Temp.: 90.14 Ambient Temp.: 94.41 F
16:48:15.711 -> Average Cal sensor Temp.: 90.32 Ambient Temp.: 94.33 F
16:48:17.274 -> Average Cal sensor Temp.: 90.46 Ambient Temp.: 94.41 F
16:48:18.825 -> Average Cal sensor Temp.: 90.50 Ambient Temp.: 94.37 F
16:48:20.384 -> Average Cal sensor Temp.: 90.62 Ambient Temp.: 94.37 F
16:48:21.932 -> Average Cal sensor Temp.: 90.53 Ambient Temp.: 94.41 F
16:48:23.488 -> Average Cal sensor Temp.: 90.51 Ambient Temp.: 94.41 F
16:48:25.041 -> Average Cal sensor Temp.: 90.50 Ambient Temp.: 94.41 F
16:48:26.595 -> Average Cal sensor Temp.: 90.51 Ambient Temp.: 94.41 F
16:48:28.152 -> Average Cal sensor Temp.: 90.37 Ambient Temp.: 94.41 F
16:48:29.708 -> Average Cal sensor Temp.: 90.51 Ambient Temp.: 94.48 F
16:48:31.257 -> Average Cal sensor Temp.: 90.44 Ambient Temp.: 94.41 F
16:48:32.811 -> Average Cal sensor Temp.: 90.34 Ambient Temp.: 94.41 F
16:48:34.373 -> Average Cal sensor Temp.: 90.45 Ambient Temp.: 94.41 F
16:48:35.956 -> Average Cal sensor Temp.: 90.35 Ambient Temp.: 94.48 F
16:48:37.513 -> Average Cal sensor Temp.: 90.53 Ambient Temp.: 94.48 F
16:48:39.038 -> Average Cal sensor Temp.: 90.59 Ambient Temp.: 94.51 F
16:48:40.625 -> Average Cal sensor Temp.: 90.54 Ambient Temp.: 94.48 F
16:48:42.172 -> Average Cal sensor Temp.: 90.62 Ambient Temp.: 94.48 F
16:48:43.730 -> Average Cal sensor Temp.: 90.50 Ambient Temp.: 94.51 F
16:48:45.291 -> Average Cal sensor Temp.: 90.45 Ambient Temp.: 94.48 F
16:48:46.845 -> Average Cal sensor Temp.: 90.35 Ambient Temp.: 94.48 F
16:48:48.397 -> Average Cal sensor Temp.: 90.42 Ambient Temp.: 94.51 F
16:48:49.953 -> Average Cal sensor Temp.: 90.49 Ambient Temp.: 94.55 F
16:48:51.507 -> Average Cal sensor Temp.: 90.42 Ambient Temp.: 94.51 F
16:48:53.060 -> Average Cal sensor Temp.: 90.56 Ambient Temp.: 94.48 F
16:48:54.613 -> Average Cal sensor Temp.: 90.50 Ambient Temp.: 94.48 F
16:48:56.179 -> Average Cal sensor Temp.: 90.60 Ambient Temp.: 94.48 F
16:48:57.733 -> Average Cal sensor Temp.: 90.47 Ambient Temp.: 94.48 F
16:48:59.290 -> Average Cal sensor Temp.: 90.63 Ambient Temp.: 94.51 F
16:49:00.837 -> Average Cal sensor Temp.: 90.56 Ambient Temp.: 94.48 F
16:49:02.425 -> Average Cal sensor Temp.: 90.66 Ambient Temp.: 94.51 F
16:49:03.982 -> Average Cal sensor Temp.: 90.63 Ambient Temp.: 94.51 F
16:49:05.536 -> Average Cal sensor Temp.: 90.46 Ambient Temp.: 94.51 F
16:49:07.087 -> Average Cal sensor Temp.: 90.55 Ambient Temp.: 94.55 F
16:49:08.649 -> Average Cal sensor Temp.: 90.66 Ambient Temp.: 94.51 F
16:49:10.178 -> Average Cal sensor Temp.: 90.52 Ambient Temp.: 94.51 F
16:49:11.765 -> Average Cal sensor Temp.: 90.65 Ambient Temp.: 94.51 F
16:49:13.288 -> Average Cal sensor Temp.: 90.54 Ambient Temp.: 94.59 F
16:49:14.874 -> Average Cal sensor Temp.: 90.59 Ambient Temp.: 94.51 F
16:49:16.426 -> Average Cal sensor Temp.: 90.50 Ambient Temp.: 94.59 F
16:49:17.987 -> Average Cal sensor Temp.: 90.61 Ambient Temp.: 94.51 F
16:49:19.545 -> Average Cal sensor Temp.: 90.56 Ambient Temp.: 94.51 F
16:49:21.075 -> Average Cal sensor Temp.: 90.58 Ambient Temp.: 94.55 F
16:49:22.631 -> Average Cal sensor Temp.: 90.54 Ambient Temp.: 94.51 F
16:49:24.185 -> Average Cal sensor Temp.: 90.83 Ambient Temp.: 94.55 F
16:49:25.741 -> Average Cal sensor Temp.: 90.73 Ambient Temp.: 94.59 F
16:49:27.324 -> Average Cal sensor Temp.: 90.76 Ambient Temp.: 94.59 F
16:49:28.877 -> Average Cal sensor Temp.: 90.79 Ambient Temp.: 94.55 F
16:49:30.429 -> Average Cal sensor Temp.: 90.80 Ambient Temp.: 94.59 F
16:49:31.985 -> Average Cal sensor Temp.: 90.72 Ambient Temp.: 94.51 F
16:49:33.539 -> Average Cal sensor Temp.: 90.61 Ambient Temp.: 94.55 F
16:49:35.092 -> Average Cal sensor Temp.: 90.84 Ambient Temp.: 94.59 F
16:49:36.648 -> Average Cal sensor Temp.: 90.82 Ambient Temp.: 94.55 F
16:49:38.197 -> Average Cal sensor Temp.: 90.77 Ambient Temp.: 94.59 F
16:49:39.778 -> Average Cal sensor Temp.: 90.85 Ambient Temp.: 94.62 F
16:49:41.332 -> Average Cal sensor Temp.: 91.54 Ambient Temp.: 94.59 F
16:49:42.889 -> Average Cal sensor Temp.: 92.07 Ambient Temp.: 94.55 F
16:49:44.435 -> Average Cal sensor Temp.: 91.67 Ambient Temp.: 94.62 F
16:49:45.987 -> Average Cal sensor Temp.: 91.54 Ambient Temp.: 94.59 F
16:49:47.543 -> Average Cal sensor Temp.: 91.17 Ambient Temp.: 94.55 F
16:49:49.093 -> Average Cal sensor Temp.: 91.09 Ambient Temp.: 94.62 F
16:49:50.643 -> Average Cal sensor Temp.: 90.86 Ambient Temp.: 94.59 F
16:49:52.229 -> Average Cal sensor Temp.: 90.91 Ambient Temp.: 94.59 F
16:49:53.783 -> Average Cal sensor Temp.: 91.44 Ambient Temp.: 94.55 F
16:49:55.335 -> Average Cal sensor Temp.: 91.84 Ambient Temp.: 94.59 F
16:49:56.894 -> Average Cal sensor Temp.: 91.67 Ambient Temp.: 94.59 F
16:49:58.450 -> Average Cal sensor Temp.: 91.33 Ambient Temp.: 94.62 F
16:50:00.004 -> Average Cal sensor Temp.: 91.72 Ambient Temp.: 94.59 F
16:50:01.561 -> Average Cal sensor Temp.: 91.80 Ambient Temp.: 94.55 F
16:50:03.105 -> Average Cal sensor Temp.: 91.35 Ambient Temp.: 94.62 F
16:50:04.660 -> Average Cal sensor Temp.: 91.28 Ambient Temp.: 94.59 F
16:50:06.218 -> Average Cal sensor Temp.: 91.10 Ambient Temp.: 94.59 F
16:50:07.778 -> Average Cal sensor Temp.: 91.06 Ambient Temp.: 94.59 F
16:50:09.332 -> Average Cal sensor Temp.: 91.08 Ambient Temp.: 94.59 F
16:50:10.890 -> Average Cal sensor Temp.: 90.72 Ambient Temp.: 94.62 F
16:50:12.446 -> Average Cal sensor Temp.: 90.53 Ambient Temp.: 94.55 F
16:50:13.999 -> Average Cal sensor Temp.: 90.49 Ambient Temp.: 94.51 F
16:50:15.556 -> Average Cal sensor Temp.: 90.45 Ambient Temp.: 94.59 F
16:50:17.113 -> Average Cal sensor Temp.: 90.46 Ambient Temp.: 94.59 F
16:50:18.665 -> Average Cal sensor Temp.: 90.22 Ambient Temp.: 94.55 F
16:50:20.218 -> Average Cal sensor Temp.: 90.08 Ambient Temp.: 94.59 F
16:50:21.773 -> Average Cal sensor Temp.: 90.36 Ambient Temp.: 94.59 F
16:50:23.354 -> Average Cal sensor Temp.: 90.34 Ambient Temp.: 94.62 F
16:50:24.907 -> Average Cal sensor Temp.: 90.23 Ambient Temp.: 94.51 F
16:50:26.463 -> Average Cal sensor Temp.: 90.44 Ambient Temp.: 94.59 F
16:50:28.019 -> Average Cal sensor Temp.: 90.47 Ambient Temp.: 94.59 F
16:50:29.565 -> Average Cal sensor Temp.: 90.56 Ambient Temp.: 94.55 F
16:50:31.116 -> Average Cal sensor Temp.: 90.54 Ambient Temp.: 94.59 F
16:50:32.671 -> Average Cal sensor Temp.: 90.70 Ambient Temp.: 94.51 F
16:50:34.224 -> Average Cal sensor Temp.: 90.70 Ambient Temp.: 94.55 F
16:50:35.778 -> Average Cal sensor Temp.: 90.49 Ambient Temp.: 94.55 F
16:50:37.333 -> Average Cal sensor Temp.: 90.66 Ambient Temp.: 94.59 F
16:50:38.891 -> Average Cal sensor Temp.: 90.62 Ambient Temp.: 94.59 F
16:50:40.481 -> Average Cal sensor Temp.: 90.63 Ambient Temp.: 94.55 F
16:50:42.037 -> Average Cal sensor Temp.: 90.75 Ambient Temp.: 94.59 F
16:50:43.591 -> Average Cal sensor Temp.: 91.05 Ambient Temp.: 94.59 F
16:50:45.148 -> Average Cal sensor Temp.: 90.91 Ambient Temp.: 94.59 F
16:50:46.700 -> Average Cal sensor Temp.: 90.78 Ambient Temp.: 94.51 F
16:50:48.255 -> Average Cal sensor Temp.: 90.92 Ambient Temp.: 94.55 F
16:50:49.812 -> Average Cal sensor Temp.: 90.91 Ambient Temp.: 94.55 F
16:50:51.367 -> Average Cal sensor Temp.: 90.95 Ambient Temp.: 94.59 F
16:50:52.922 -> Average Cal sensor Temp.: 91.06 Ambient Temp.: 94.55 F
16:50:54.468 -> Average Cal sensor Temp.: 90.91 Ambient Temp.: 94.55 F
16:50:56.021 -> Average Cal sensor Temp.: 90.79 Ambient Temp.: 94.59 F
16:50:57.576 -> Average Cal sensor Temp.: 90.89 Ambient Temp.: 94.59 F
16:50:59.133 -> Average Cal sensor Temp.: 90.98 Ambient Temp.: 94.59 F
16:51:00.688 -> Average Cal sensor Temp.: 91.08 Ambient Temp.: 94.59 F
16:51:02.245 -> Average Cal sensor Temp.: 91.00 Ambient Temp.: 94.55 F
16:51:03.830 -> Average Cal sensor Temp.: 91.22 Ambient Temp.: 94.55 F
16:51:05.384 -> Average Cal sensor Temp.: 91.09 Ambient Temp.: 94.59 F
16:51:06.937 -> Average Cal sensor Temp.: 91.21 Ambient Temp.: 94.55 F
16:51:08.488 -> Average Cal sensor Temp.: 91.18 Ambient Temp.: 94.62 F
16:51:10.039 -> Average Cal sensor Temp.: 91.22 Ambient Temp.: 94.62 F
16:51:11.592 -> Average Cal sensor Temp.: 91.01 Ambient Temp.: 94.59 F
16:51:13.145 -> Average Cal sensor Temp.: 91.01 Ambient Temp.: 94.55 F
16:51:14.694 -> Average Cal sensor Temp.: 91.13 Ambient Temp.: 94.59 F
16:51:16.246 -> Average Cal sensor Temp.: 91.12 Ambient Temp.: 94.59 F
16:51:17.826 -> Average Cal sensor Temp.: 91.13 Ambient Temp.: 94.59 F
16:51:19.346 -> Average Cal sensor Temp.: 90.91 Ambient Temp.: 94.55 F
16:51:20.932 -> Average Cal sensor Temp.: 91.18 Ambient Temp.: 94.59 F
16:51:22.488 -> Average Cal sensor Temp.: 90.91 Ambient Temp.: 94.55 F
16:51:24.037 -> Average Cal sensor Temp.: 90.82 Ambient Temp.: 94.59 F
16:51:25.585 -> Average Cal sensor Temp.: 90.98 Ambient Temp.: 94.55 F
16:51:27.133 -> Average Cal sensor Temp.: 91.00 Ambient Temp.: 94.51 F
16:51:28.686 -> Average Cal sensor Temp.: 91.06 Ambient Temp.: 94.59 F
16:51:30.244 -> Average Cal sensor Temp.: 91.07 Ambient Temp.: 94.55 F
16:51:31.798 -> Average Cal sensor Temp.: 91.03 Ambient Temp.: 94.51 F
16:51:33.355 -> Average Cal sensor Temp.: 91.24 Ambient Temp.: 94.59 F
16:51:34.910 -> Average Cal sensor Temp.: 91.22 Ambient Temp.: 94.55 F
16:51:36.464 -> Average Cal sensor Temp.: 91.24 Ambient Temp.: 94.59 F
16:51:38.021 -> Average Cal sensor Temp.: 91.16 Ambient Temp.: 94.55 F
16:51:39.573 -> Average Cal sensor Temp.: 91.19 Ambient Temp.: 94.62 F
16:51:41.159 -> Average Cal sensor Temp.: 90.92 Ambient Temp.: 94.62 F
16:51:42.711 -> Average Cal sensor Temp.: 91.10 Ambient Temp.: 94.62 F
16:51:44.266 -> Average Cal sensor Temp.: 91.04 Ambient Temp.: 94.59 F
16:51:45.822 -> Average Cal sensor Temp.: 90.95 Ambient Temp.: 94.59 F
16:51:47.381 -> Average Cal sensor Temp.: 91.04 Ambient Temp.: 94.62 F
16:51:48.930 -> Average Cal sensor Temp.: 91.15 Ambient Temp.: 94.59 F
16:51:50.492 -> Average Cal sensor Temp.: 90.74 Ambient Temp.: 94.59 F
16:51:52.045 -> Average Cal sensor Temp.: 91.06 Ambient Temp.: 94.59 F
16:51:53.593 -> Average Cal sensor Temp.: 90.87 Ambient Temp.: 94.59 F
16:51:55.143 -> Average Cal sensor Temp.: 91.04 Ambient Temp.: 94.59 F
16:51:56.694 -> Average Cal sensor Temp.: 91.10 Ambient Temp.: 94.59 F
16:51:58.277 -> Average Cal sensor Temp.: 91.25 Ambient Temp.: 94.62 F
16:51:59.832 -> Average Cal sensor Temp.: 91.26 Ambient Temp.: 94.62 F
16:52:01.385 -> Average Cal sensor Temp.: 91.41 Ambient Temp.: 94.59 F
16:52:02.940 -> Average Cal sensor Temp.: 91.31 Ambient Temp.: 94.59 F
16:52:04.495 -> Average Cal sensor Temp.: 91.18 Ambient Temp.: 94.59 F
16:52:06.047 -> Average Cal sensor Temp.: 91.06 Ambient Temp.: 94.62 F
16:52:07.600 -> Average Cal sensor Temp.: 91.14 Ambient Temp.: 94.62 F
16:52:09.156 -> Average Cal sensor Temp.: 91.42 Ambient Temp.: 94.62 F
16:52:10.707 -> Average Cal sensor Temp.: 91.09 Ambient Temp.: 94.59 F
16:52:12.257 -> Average Cal sensor Temp.: 91.04 Ambient Temp.: 94.59 F
16:52:13.810 -> Average Cal sensor Temp.: 91.08 Ambient Temp.: 94.59 F
16:52:15.362 -> Average Cal sensor Temp.: 91.00 Ambient Temp.: 94.62 F
16:52:16.918 -> Average Cal sensor Temp.: 91.19 Ambient Temp.: 94.59 F
16:52:18.503 -> Average Cal sensor Temp.: 91.17 Ambient Temp.: 94.59 F
16:52:20.050 -> Average Cal sensor Temp.: 91.31 Ambient Temp.: 94.66 F
16:52:21.603 -> Average Cal sensor Temp.: 91.45 Ambient Temp.: 94.59 F
16:52:23.157 -> Average Cal sensor Temp.: 91.34 Ambient Temp.: 94.59 F
16:52:24.708 -> Average Cal sensor Temp.: 91.25 Ambient Temp.: 94.62 F
16:52:26.262 -> Average Cal sensor Temp.: 91.14 Ambient Temp.: 94.59 F
16:52:27.813 -> Average Cal sensor Temp.: 91.24 Ambient Temp.: 94.59 F
16:52:29.364 -> Average Cal sensor Temp.: 91.17 Ambient Temp.: 94.59 F
16:52:30.918 -> Average Cal sensor Temp.: 91.22 Ambient Temp.: 94.62 F
16:52:32.468 -> Average Cal sensor Temp.: 91.36 Ambient Temp.: 94.62 F
16:52:34.054 -> Average Cal sensor Temp.: 91.25 Ambient Temp.: 94.62 F
16:52:35.607 -> Average Cal sensor Temp.: 91.34 Ambient Temp.: 94.59 F
16:52:37.160 -> Average Cal sensor Temp.: 91.54 Ambient Temp.: 94.62 F
16:52:38.709 -> Average Cal sensor Temp.: 91.37 Ambient Temp.: 94.59 F
16:52:40.260 -> Average Cal sensor Temp.: 91.60 Ambient Temp.: 94.66 F
16:52:41.811 -> Average Cal sensor Temp.: 91.34 Ambient Temp.: 94.62 F
16:52:43.362 -> Average Cal sensor Temp.: 91.06 Ambient Temp.: 94.62 F
16:52:44.920 -> Average Cal sensor Temp.: 91.40 Ambient Temp.: 94.66 F
16:52:46.501 -> Average Cal sensor Temp.: 91.57 Ambient Temp.: 94.66 F
16:52:48.052 -> Average Cal sensor Temp.: 90.79 Ambient Temp.: 94.66 F
16:52:49.603 -> Average Cal sensor Temp.: 91.02 Ambient Temp.: 94.59 F
16:52:51.156 -> Average Cal sensor Temp.: 91.44 Ambient Temp.: 94.62 F
16:52:52.712 -> Average Cal sensor Temp.: 91.24 Ambient Temp.: 94.66 F
16:52:54.261 -> Average Cal sensor Temp.: 91.40 Ambient Temp.: 94.66 F
16:52:55.812 -> Average Cal sensor Temp.: 91.37 Ambient Temp.: 94.66 F
16:52:57.368 -> Average Cal sensor Temp.: 91.29 Ambient Temp.: 94.66 F
16:52:58.954 -> Average Cal sensor Temp.: 91.08 Ambient Temp.: 94.62 F
16:53:00.506 -> Average Cal sensor Temp.: 91.17 Ambient Temp.: 94.59 F
16:53:02.058 -> Average Cal sensor Temp.: 91.44 Ambient Temp.: 94.66 F
16:53:03.613 -> Average Cal sensor Temp.: 91.70 Ambient Temp.: 94.66 F
16:53:05.169 -> Average Cal sensor Temp.: 91.34 Ambient Temp.: 94.62 F
16:53:06.721 -> Average Cal sensor Temp.: 91.66 Ambient Temp.: 94.69 F
16:53:08.277 -> Average Cal sensor Temp.: 91.51 Ambient Temp.: 94.59 F
16:53:09.834 -> Average Cal sensor Temp.: 91.69 Ambient Temp.: 94.66 F
16:53:11.384 -> Average Cal sensor Temp.: 91.52 Ambient Temp.: 94.66 F
16:53:12.936 -> Average Cal sensor Temp.: 91.57 Ambient Temp.: 94.66 F
16:53:14.486 -> Average Cal sensor Temp.: 91.67 Ambient Temp.: 94.69 F
16:53:16.072 -> Average Cal sensor Temp.: 91.83 Ambient Temp.: 94.62 F
16:53:17.623 -> Average Cal sensor Temp.: 91.58 Ambient Temp.: 94.66 F
16:53:19.179 -> Average Cal sensor Temp.: 91.59 Ambient Temp.: 94.66 F
16:53:20.730 -> Average Cal sensor Temp.: 91.70 Ambient Temp.: 94.69 F
16:53:22.282 -> Average Cal sensor Temp.: 91.58 Ambient Temp.: 94.66 F
16:53:23.838 -> Average Cal sensor Temp.: 91.58 Ambient Temp.: 94.69 F
16:53:25.392 -> Average Cal sensor Temp.: 91.62 Ambient Temp.: 94.69 F
16:53:26.951 -> Average Cal sensor Temp.: 91.54 Ambient Temp.: 94.66 F
16:53:28.509 -> Average Cal sensor Temp.: 91.65 Ambient Temp.: 94.66 F
16:53:30.060 -> Average Cal sensor Temp.: 91.92 Ambient Temp.: 94.62 F
16:53:31.617 -> Average Cal sensor Temp.: 91.95 Ambient Temp.: 94.69 F
16:53:33.171 -> Average Cal sensor Temp.: 91.94 Ambient Temp.: 94.73 F
16:53:34.725 -> Average Cal sensor Temp.: 92.00 Ambient Temp.: 94.73 F
16:53:36.284 -> Average Cal sensor Temp.: 91.80 Ambient Temp.: 94.66 F
16:53:37.832 -> Average Cal sensor Temp.: 91.80 Ambient Temp.: 94.66 F
16:53:39.386 -> Average Cal sensor Temp.: 91.95 Ambient Temp.: 94.66 F
16:53:40.972 -> Average Cal sensor Temp.: 91.97 Ambient Temp.: 94.69 F
16:53:42.528 -> Average Cal sensor Temp.: 91.76 Ambient Temp.: 94.73 F
16:53:44.087 -> Average Cal sensor Temp.: 91.79 Ambient Temp.: 94.69 F
16:53:45.610 -> Average Cal sensor Temp.: 91.72 Ambient Temp.: 94.69 F
16:53:47.170 -> Average Cal sensor Temp.: 91.66 Ambient Temp.: 94.66 F
16:53:48.725 -> Average Cal sensor Temp.: 91.30 Ambient Temp.: 94.66 F
16:53:50.281 -> Average Cal sensor Temp.: 91.36 Ambient Temp.: 94.69 F
16:53:51.838 -> Average Cal sensor Temp.: 91.72 Ambient Temp.: 94.69 F
16:53:53.423 -> Average Cal sensor Temp.: 91.51 Ambient Temp.: 94.69 F
16:53:54.975 -> Average Cal sensor Temp.: 91.57 Ambient Temp.: 94.69 F
16:53:56.529 -> Average Cal sensor Temp.: 91.41 Ambient Temp.: 94.73 F
16:53:58.082 -> Average Cal sensor Temp.: 91.33 Ambient Temp.: 94.73 F
16:53:59.639 -> Average Cal sensor Temp.: 91.59 Ambient Temp.: 94.73 F
16:54:01.190 -> Average Cal sensor Temp.: 91.55 Ambient Temp.: 94.69 F
16:54:02.737 -> Average Cal sensor Temp.: 91.87 Ambient Temp.: 94.69 F
16:54:04.289 -> Average Cal sensor Temp.: 92.00 Ambient Temp.: 94.73 F
16:54:05.846 -> Average Cal sensor Temp.: 91.87 Ambient Temp.: 94.77 F
16:54:07.397 -> Average Cal sensor Temp.: 91.86 Ambient Temp.: 94.69 F
16:54:08.945 -> Average Cal sensor Temp.: 91.94 Ambient Temp.: 94.73 F
16:54:10.529 -> Average Cal sensor Temp.: 92.12 Ambient Temp.: 94.69 F
16:54:12.083 -> Average Cal sensor Temp.: 92.11 Ambient Temp.: 94.73 F
16:54:13.636 -> Average Cal sensor Temp.: 93.46 Ambient Temp.: 94.73 F
16:54:15.192 -> Average Cal sensor Temp.: 92.27 Ambient Temp.: 94.73 F
16:54:16.745 -> Average Cal sensor Temp.: 92.27 Ambient Temp.: 94.73 F
16:54:18.302 -> Average Cal sensor Temp.: 92.34 Ambient Temp.: 94.73 F
16:54:19.855 -> Average Cal sensor Temp.: 92.11 Ambient Temp.: 94.73 F
16:54:21.405 -> Average Cal sensor Temp.: 92.09 Ambient Temp.: 94.77 F
16:54:22.957 -> Average Cal sensor Temp.: 92.11 Ambient Temp.: 94.77 F
16:54:24.513 -> Average Cal sensor Temp.: 92.16 Ambient Temp.: 94.77 F
16:54:26.067 -> Average Cal sensor Temp.: 91.93 Ambient Temp.: 94.77 F
16:54:27.622 -> Average Cal sensor Temp.: 91.90 Ambient Temp.: 94.77 F
16:54:29.178 -> Average Cal sensor Temp.: 91.43 Ambient Temp.: 94.80 F
16:54:30.731 -> Average Cal sensor Temp.: 91.79 Ambient Temp.: 94.73 F
16:54:32.280 -> Average Cal sensor Temp.: 91.90 Ambient Temp.: 94.73 F
16:54:33.834 -> Average Cal sensor Temp.: 91.83 Ambient Temp.: 94.77 F
16:54:35.386 -> Average Cal sensor Temp.: 91.94 Ambient Temp.: 94.80 F
16:54:36.972 -> Average Cal sensor Temp.: 92.01 Ambient Temp.: 94.73 F
16:54:38.530 -> Average Cal sensor Temp.: 91.84 Ambient Temp.: 94.77 F
16:54:40.085 -> Average Cal sensor Temp.: 91.91 Ambient Temp.: 94.77 F
16:54:41.637 -> Average Cal sensor Temp.: 91.77 Ambient Temp.: 94.77 F
16:54:43.187 -> Average Cal sensor Temp.: 91.51 Ambient Temp.: 94.77 F
16:54:44.742 -> Average Cal sensor Temp.: 91.44 Ambient Temp.: 94.73 F
16:54:46.289 -> Average Cal sensor Temp.: 91.36 Ambient Temp.: 94.73 F
16:54:47.839 -> Average Cal sensor Temp.: 91.70 Ambient Temp.: 94.73 F
16:54:49.392 -> Average Cal sensor Temp.: 91.84 Ambient Temp.: 94.73 F
16:54:50.940 -> Average Cal sensor Temp.: 91.71 Ambient Temp.: 94.77 F
16:54:52.491 -> Average Cal sensor Temp.: 91.91 Ambient Temp.: 94.77 F
16:54:54.078 -> Average Cal sensor Temp.: 91.77 Ambient Temp.: 94.73 F
16:54:55.603 -> Average Cal sensor Temp.: 92.10 Ambient Temp.: 94.77 F
16:54:57.165 -> Average Cal sensor Temp.: 92.19 Ambient Temp.: 94.73 F
16:54:58.715 -> Average Cal sensor Temp.: 92.05 Ambient Temp.: 94.77 F
16:55:00.267 -> Average Cal sensor Temp.: 91.65 Ambient Temp.: 94.73 F
16:55:01.853 -> Average Cal sensor Temp.: 92.02 Ambient Temp.: 94.77 F
16:55:03.402 -> Average Cal sensor Temp.: 92.12 Ambient Temp.: 94.77 F
16:55:04.957 -> Average Cal sensor Temp.: 92.39 Ambient Temp.: 94.77 F
16:55:06.519 -> Average Cal sensor Temp.: 92.54 Ambient Temp.: 94.77 F
16:55:08.067 -> Average Cal sensor Temp.: 92.33 Ambient Temp.: 94.73 F
16:55:09.619 -> Average Cal sensor Temp.: 92.23 Ambient Temp.: 94.73 F
16:55:11.169 -> Average Cal sensor Temp.: 92.29 Ambient Temp.: 94.73 F
16:55:12.719 -> Average Cal sensor Temp.: 92.25 Ambient Temp.: 94.77 F
16:55:14.274 -> Average Cal sensor Temp.: 92.27 Ambient Temp.: 94.77 F
16:55:15.827 -> Average Cal sensor Temp.: 92.21 Ambient Temp.: 94.77 F
16:55:17.374 -> Average Cal sensor Temp.: 91.94 Ambient Temp.: 94.77 F
16:55:18.933 -> Average Cal sensor Temp.: 92.03 Ambient Temp.: 94.73 F
16:55:20.485 -> Average Cal sensor Temp.: 92.12 Ambient Temp.: 94.77 F
16:55:22.071 -> Average Cal sensor Temp.: 91.98 Ambient Temp.: 94.73 F
16:55:23.628 -> Average Cal sensor Temp.: 92.16 Ambient Temp.: 94.77 F
16:55:25.152 -> Average Cal sensor Temp.: 91.95 Ambient Temp.: 94.73 F
16:55:26.710 -> Average Cal sensor Temp.: 91.87 Ambient Temp.: 94.73 F
16:55:28.261 -> Average Cal sensor Temp.: 91.81 Ambient Temp.: 94.73 F
16:55:29.818 -> Average Cal sensor Temp.: 92.07 Ambient Temp.: 94.77 F
16:55:31.373 -> Average Cal sensor Temp.: 92.16 Ambient Temp.: 94.80 F
16:55:32.930 -> Average Cal sensor Temp.: 92.20 Ambient Temp.: 94.73 F
16:55:34.486 -> Average Cal sensor Temp.: 92.11 Ambient Temp.: 94.77 F
16:55:36.037 -> Average Cal sensor Temp.: 92.25 Ambient Temp.: 94.77 F
16:55:37.592 -> Average Cal sensor Temp.: 92.33 Ambient Temp.: 94.80 F
16:55:39.178 -> Average Cal sensor Temp.: 92.48 Ambient Temp.: 94.80 F
16:55:40.730 -> Average Cal sensor Temp.: 92.35 Ambient Temp.: 94.77 F
16:55:42.285 -> Average Cal sensor Temp.: 92.34 Ambient Temp.: 94.77 F
16:55:43.832 -> Average Cal sensor Temp.: 92.35 Ambient Temp.: 94.77 F
16:55:45.384 -> Average Cal sensor Temp.: 92.39 Ambient Temp.: 94.80 F
16:55:46.941 -> Average Cal sensor Temp.: 92.45 Ambient Temp.: 94.80 F
16:55:48.497 -> Average Cal sensor Temp.: 92.28 Ambient Temp.: 94.77 F
16:55:50.055 -> Average Cal sensor Temp.: 91.89 Ambient Temp.: 94.80 F
16:55:51.613 -> Average Cal sensor Temp.: 91.98 Ambient Temp.: 94.77 F
16:55:53.166 -> Average Cal sensor Temp.: 91.80 Ambient Temp.: 94.80 F
16:55:54.715 -> Average Cal sensor Temp.: 92.10 Ambient Temp.: 94.77 F
16:55:56.268 -> Average Cal sensor Temp.: 91.99 Ambient Temp.: 94.77 F
16:55:57.857 -> Average Cal sensor Temp.: 92.04 Ambient Temp.: 94.77 F
16:55:59.412 -> Average Cal sensor Temp.: 92.20 Ambient Temp.: 94.73 F
16:56:00.968 -> Average Cal sensor Temp.: 92.22 Ambient Temp.: 94.73 F
16:56:02.523 -> Average Cal sensor Temp.: 92.21 Ambient Temp.: 94.77 F
16:56:04.048 -> Average Cal sensor Temp.: 92.14 Ambient Temp.: 94.80 F
16:56:05.605 -> Average Cal sensor Temp.: 91.92 Ambient Temp.: 94.80 F
16:56:07.188 -> Average Cal sensor Temp.: 92.16 Ambient Temp.: 94.77 F
16:56:08.740 -> Average Cal sensor Temp.: 92.24 Ambient Temp.: 94.77 F
16:56:10.290 -> Average Cal sensor Temp.: 92.23 Ambient Temp.: 94.77 F
16:56:11.843 -> Average Cal sensor Temp.: 92.31 Ambient Temp.: 94.80 F
16:56:13.395 -> Average Cal sensor Temp.: 91.88 Ambient Temp.: 94.77 F
16:56:14.944 -> Average Cal sensor Temp.: 91.76 Ambient Temp.: 94.73 F
16:56:16.500 -> Average Cal sensor Temp.: 91.93 Ambient Temp.: 94.77 F
16:56:18.054 -> Average Cal sensor Temp.: 92.15 Ambient Temp.: 94.80 F
16:56:19.609 -> Average Cal sensor Temp.: 92.29 Ambient Temp.: 94.80 F
16:56:21.192 -> Average Cal sensor Temp.: 92.23 Ambient Temp.: 94.77 F
16:56:22.745 -> Average Cal sensor Temp.: 91.72 Ambient Temp.: 94.77 F
16:56:24.293 -> Average Cal sensor Temp.: 92.06 Ambient Temp.: 94.73 F
16:56:25.845 -> Average Cal sensor Temp.: 92.40 Ambient Temp.: 94.80 F
16:56:27.404 -> Average Cal sensor Temp.: 92.23 Ambient Temp.: 94.80 F
16:56:28.961 -> Average Cal sensor Temp.: 92.37 Ambient Temp.: 94.77 F
16:56:30.516 -> Average Cal sensor Temp.: 92.30 Ambient Temp.: 94.80 F
16:56:32.063 -> Average Cal sensor Temp.: 92.55 Ambient Temp.: 94.80 F
16:56:33.617 -> Average Cal sensor Temp.: 92.56 Ambient Temp.: 94.80 F
16:56:35.172 -> Average Cal sensor Temp.: 92.70 Ambient Temp.: 94.77 F
16:56:36.727 -> Average Cal sensor Temp.: 92.52 Ambient Temp.: 94.80 F
16:56:38.277 -> Average Cal sensor Temp.: 92.61 Ambient Temp.: 94.77 F
16:56:39.833 -> Average Cal sensor Temp.: 92.57 Ambient Temp.: 94.77 F
16:56:41.385 -> Average Cal sensor Temp.: 92.43 Ambient Temp.: 94.80 F
16:56:42.973 -> Average Cal sensor Temp.: 92.34 Ambient Temp.: 94.77 F
16:56:44.528 -> Average Cal sensor Temp.: 92.45 Ambient Temp.: 94.77 F
16:56:46.083 -> Average Cal sensor Temp.: 92.19 Ambient Temp.: 94.77 F
16:56:47.638 -> Average Cal sensor Temp.: 92.14 Ambient Temp.: 94.77 F
16:56:49.191 -> Average Cal sensor Temp.: 92.23 Ambient Temp.: 94.77 F
16:56:50.746 -> Average Cal sensor Temp.: 91.95 Ambient Temp.: 94.80 F
16:56:52.303 -> Average Cal sensor Temp.: 92.30 Ambient Temp.: 94.80 F
16:56:53.861 -> Average Cal sensor Temp.: 92.14 Ambient Temp.: 94.77 F
16:56:55.409 -> Average Cal sensor Temp.: 92.15 Ambient Temp.: 94.87 F
16:56:56.959 -> Average Cal sensor Temp.: 92.17 Ambient Temp.: 94.87 F
16:56:58.513 -> Average Cal sensor Temp.: 92.38 Ambient Temp.: 94.87 F
16:57:00.059 -> Average Cal sensor Temp.: 92.27 Ambient Temp.: 94.80 F
16:57:01.616 -> Average Cal sensor Temp.: 92.35 Ambient Temp.: 94.80 F
16:57:03.197 -> Average Cal sensor Temp.: 92.16 Ambient Temp.: 94.73 F
16:57:04.748 -> Average Cal sensor Temp.: 92.25 Ambient Temp.: 94.77 F
16:57:06.305 -> Average Cal sensor Temp.: 92.12 Ambient Temp.: 94.80 F
16:57:07.858 -> Average Cal sensor Temp.: 92.29 Ambient Temp.: 94.80 F
16:57:09.410 -> Average Cal sensor Temp.: 92.21 Ambient Temp.: 94.80 F
16:57:10.961 -> Average Cal sensor Temp.: 92.14 Ambient Temp.: 94.77 F
16:57:12.515 -> Average Cal sensor Temp.: 91.94 Ambient Temp.: 94.77 F
16:57:14.069 -> Average Cal sensor Temp.: 92.26 Ambient Temp.: 94.80 F
16:57:15.626 -> Average Cal sensor Temp.: 92.33 Ambient Temp.: 94.77 F
16:57:17.180 -> Average Cal sensor Temp.: 91.93 Ambient Temp.: 94.80 F
16:57:18.732 -> Average Cal sensor Temp.: 91.59 Ambient Temp.: 94.77 F
16:57:20.308 -> Average Cal sensor Temp.: 91.97 Ambient Temp.: 94.77 F
16:57:21.861 -> Average Cal sensor Temp.: 92.06 Ambient Temp.: 94.80 F
16:57:23.409 -> Average Cal sensor Temp.: 92.28 Ambient Temp.: 94.80 F
16:57:24.957 -> Average Cal sensor Temp.: 91.98 Ambient Temp.: 94.77 F
16:57:26.512 -> Average Cal sensor Temp.: 91.83 Ambient Temp.: 94.73 F
16:57:28.068 -> Average Cal sensor Temp.: 92.06 Ambient Temp.: 94.77 F
16:57:29.622 -> Average Cal sensor Temp.: 92.21 Ambient Temp.: 94.77 F
16:57:31.179 -> Average Cal sensor Temp.: 92.32 Ambient Temp.: 94.77 F
16:57:32.733 -> Average Cal sensor Temp.: 92.26 Ambient Temp.: 94.77 F
16:57:34.288 -> Average Cal sensor Temp.: 92.21 Ambient Temp.: 94.80 F
16:57:35.846 -> Average Cal sensor Temp.: 92.40 Ambient Temp.: 94.77 F
16:57:37.403 -> Average Cal sensor Temp.: 92.13 Ambient Temp.: 94.77 F
16:57:38.961 -> Average Cal sensor Temp.: 92.10 Ambient Temp.: 94.87 F
16:57:40.509 -> Average Cal sensor Temp.: 92.25 Ambient Temp.: 94.77 F
16:57:42.091 -> Average Cal sensor Temp.: 92.16 Ambient Temp.: 94.80 F
16:57:43.640 -> Average Cal sensor Temp.: 92.29 Ambient Temp.: 94.80 F
16:57:45.199 -> Average Cal sensor Temp.: 92.43 Ambient Temp.: 94.77 F
16:57:46.752 -> Average Cal sensor Temp.: 92.62 Ambient Temp.: 94.80 F
16:57:48.305 -> Average Cal sensor Temp.: 92.49 Ambient Temp.: 94.77 F
16:57:49.852 -> Average Cal sensor Temp.: 92.31 Ambient Temp.: 94.77 F
16:57:51.408 -> Average Cal sensor Temp.: 92.40 Ambient Temp.: 94.80 F
16:57:52.964 -> Average Cal sensor Temp.: 92.29 Ambient Temp.: 94.80 F
16:57:54.517 -> Average Cal sensor Temp.: 92.25 Ambient Temp.: 94.77 F
16:57:56.071 -> Average Cal sensor Temp.: 92.17 Ambient Temp.: 94.80 F
16:57:57.656 -> Average Cal sensor Temp.: 92.22 Ambient Temp.: 94.80 F
16:57:59.209 -> Average Cal sensor Temp.: 92.33 Ambient Temp.: 94.80 F
16:58:00.764 -> Average Cal sensor Temp.: 92.45 Ambient Temp.: 94.73 F
16:58:02.315 -> Average Cal sensor Temp.: 92.32 Ambient Temp.: 94.80 F
16:58:03.872 -> Average Cal sensor Temp.: 92.32 Ambient Temp.: 94.80 F
16:58:05.424 -> Average Cal sensor Temp.: 92.32 Ambient Temp.: 94.87 F
16:58:06.979 -> Average Cal sensor Temp.: 92.53 Ambient Temp.: 94.77 F
16:58:08.533 -> Average Cal sensor Temp.: 92.53 Ambient Temp.: 94.77 F
16:58:10.086 -> Average Cal sensor Temp.: 92.68 Ambient Temp.: 94.80 F
16:58:11.641 -> Average Cal sensor Temp.: 92.72 Ambient Temp.: 94.80 F
16:58:13.195 -> Average Cal sensor Temp.: 92.61 Ambient Temp.: 94.80 F
16:58:14.749 -> Average Cal sensor Temp.: 92.37 Ambient Temp.: 94.80 F
16:58:16.300 -> Average Cal sensor Temp.: 92.37 Ambient Temp.: 94.91 F
16:58:17.854 -> Average Cal sensor Temp.: 92.62 Ambient Temp.: 94.80 F
16:58:19.441 -> Average Cal sensor Temp.: 92.39 Ambient Temp.: 94.77 F
16:58:20.992 -> Average Cal sensor Temp.: 92.35 Ambient Temp.: 94.87 F
16:58:22.553 -> Average Cal sensor Temp.: 92.49 Ambient Temp.: 94.87 F
16:58:24.112 -> Average Cal sensor Temp.: 92.31 Ambient Temp.: 94.80 F
16:58:25.641 -> Average Cal sensor Temp.: 92.26 Ambient Temp.: 94.80 F
16:58:27.199 -> Average Cal sensor Temp.: 91.97 Ambient Temp.: 94.80 F
16:58:28.755 -> Average Cal sensor Temp.: 92.36 Ambient Temp.: 94.91 F
16:58:30.305 -> Average Cal sensor Temp.: 92.41 Ambient Temp.: 94.91 F
16:58:31.892 -> Average Cal sensor Temp.: 92.20 Ambient Temp.: 94.77 F
16:58:33.441 -> Average Cal sensor Temp.: 92.36 Ambient Temp.: 94.87 F
16:58:34.997 -> Average Cal sensor Temp.: 92.40 Ambient Temp.: 94.80 F
16:58:36.551 -> Average Cal sensor Temp.: 92.43 Ambient Temp.: 94.87 F
16:58:38.106 -> Average Cal sensor Temp.: 92.28 Ambient Temp.: 94.80 F
16:58:39.655 -> Average Cal sensor Temp.: 92.32 Ambient Temp.: 94.80 F
16:58:41.204 -> Average Cal sensor Temp.: 92.53 Ambient Temp.: 94.87 F
16:58:42.785 -> Average Cal sensor Temp.: 92.43 Ambient Temp.: 94.87 F
16:58:44.338 -> Average Cal sensor Temp.: 92.62 Ambient Temp.: 94.80 F
16:58:45.888 -> Average Cal sensor Temp.: 92.49 Ambient Temp.: 94.87 F
16:58:47.439 -> Average Cal sensor Temp.: 92.48 Ambient Temp.: 94.87 F
16:58:48.991 -> Average Cal sensor Temp.: 92.50 Ambient Temp.: 94.87 F
16:58:50.546 -> Average Cal sensor Temp.: 92.46 Ambient Temp.: 94.80 F
16:58:52.100 -> Average Cal sensor Temp.: 92.38 Ambient Temp.: 94.87 F
16:58:53.655 -> Average Cal sensor Temp.: 92.63 Ambient Temp.: 94.91 F
16:58:55.241 -> Average Cal sensor Temp.: 92.65 Ambient Temp.: 94.91 F
16:58:56.792 -> Average Cal sensor Temp.: 92.60 Ambient Temp.: 94.80 F
16:58:58.341 -> Average Cal sensor Temp.: 92.61 Ambient Temp.: 94.87 F
16:58:59.895 -> Average Cal sensor Temp.: 92.70 Ambient Temp.: 94.87 F
16:59:01.446 -> Average Cal sensor Temp.: 92.52 Ambient Temp.: 94.80 F
16:59:03.002 -> Average Cal sensor Temp.: 92.51 Ambient Temp.: 94.87 F
16:59:04.553 -> Average Cal sensor Temp.: 92.58 Ambient Temp.: 94.87 F
16:59:06.108 -> Average Cal sensor Temp.: 92.69 Ambient Temp.: 94.91 F
16:59:07.664 -> Average Cal sensor Temp.: 92.56 Ambient Temp.: 94.87 F
16:59:09.250 -> Average Cal sensor Temp.: 92.59 Ambient Temp.: 94.91 F
16:59:10.801 -> Average Cal sensor Temp.: 92.52 Ambient Temp.: 94.80 F
16:59:12.353 -> Average Cal sensor Temp.: 92.35 Ambient Temp.: 94.87 F
16:59:13.904 -> Average Cal sensor Temp.: 92.64 Ambient Temp.: 94.91 F
16:59:15.455 -> Average Cal sensor Temp.: 91.96 Ambient Temp.: 94.80 F
16:59:17.005 -> Average Cal sensor Temp.: 91.60 Ambient Temp.: 94.87 F
16:59:18.590 -> Average Cal sensor Temp.: 91.96 Ambient Temp.: 94.87 F
16:59:20.142 -> Average Cal sensor Temp.: 92.05 Ambient Temp.: 94.87 F
16:59:21.699 -> Average Cal sensor Temp.: 92.06 Ambient Temp.: 94.87 F
16:59:23.252 -> Average Cal sensor Temp.: 92.10 Ambient Temp.: 94.91 F
16:59:24.807 -> Average Cal sensor Temp.: 92.39 Ambient Temp.: 94.87 F
16:59:26.366 -> Average Cal sensor Temp.: 92.46 Ambient Temp.: 94.87 F
16:59:27.917 -> Average Cal sensor Temp.: 92.48 Ambient Temp.: 94.87 F
16:59:29.463 -> Average Cal sensor Temp.: 92.38 Ambient Temp.: 94.87 F
16:59:31.015 -> Average Cal sensor Temp.: 92.58 Ambient Temp.: 94.87 F
16:59:32.596 -> Average Cal sensor Temp.: 92.52 Ambient Temp.: 94.91 F
16:59:34.149 -> Average Cal sensor Temp.: 92.44 Ambient Temp.: 94.80 F
16:59:35.698 -> Average Cal sensor Temp.: 92.35 Ambient Temp.: 94.91 F
16:59:37.252 -> Average Cal sensor Temp.: 92.54 Ambient Temp.: 94.95 F
16:59:38.806 -> Average Cal sensor Temp.: 92.77 Ambient Temp.: 94.80 F
16:59:40.356 -> Average Cal sensor Temp.: 92.76 Ambient Temp.: 94.91 F
16:59:41.911 -> Average Cal sensor Temp.: 92.60 Ambient Temp.: 94.91 F
16:59:43.465 -> Average Cal sensor Temp.: 92.69 Ambient Temp.: 94.87 F
16:59:45.049 -> Average Cal sensor Temp.: 92.72 Ambient Temp.: 94.87 F
16:59:46.573 -> Average Cal sensor Temp.: 92.80 Ambient Temp.: 94.95 F
16:59:48.165 -> Average Cal sensor Temp.: 92.45 Ambient Temp.: 94.91 F
16:59:49.718 -> Average Cal sensor Temp.: 92.59 Ambient Temp.: 94.87 F
16:59:51.277 -> Average Cal sensor Temp.: 92.75 Ambient Temp.: 94.87 F
16:59:52.829 -> Average Cal sensor Temp.: 92.95 Ambient Temp.: 94.87 F
16:59:54.383 -> Average Cal sensor Temp.: 92.80 Ambient Temp.: 94.91 F
16:59:55.939 -> Average Cal sensor Temp.: 92.89 Ambient Temp.: 94.91 F
16:59:57.496 -> Average Cal sensor Temp.: 92.71 Ambient Temp.: 94.87 F
16:59:59.047 -> Average Cal sensor Temp.: 92.82 Ambient Temp.: 94.91 F
17:00:00.595 -> Average Cal sensor Temp.: 92.79 Ambient Temp.: 94.95 F
17:00:02.153 -> Average Cal sensor Temp.: 92.52 Ambient Temp.: 94.91 F
17:00:03.701 -> Average Cal sensor Temp.: 92.44 Ambient Temp.: 94.87 F
17:00:05.255 -> Average Cal sensor Temp.: 92.66 Ambient Temp.: 94.91 F
17:00:06.831 -> Average Cal sensor Temp.: 92.75 Ambient Temp.: 94.87 F
17:00:08.383 -> Average Cal sensor Temp.: 92.71 Ambient Temp.: 94.91 F
17:00:09.938 -> Average Cal sensor Temp.: 92.55 Ambient Temp.: 94.91 F
17:00:11.492 -> Average Cal sensor Temp.: 92.83 Ambient Temp.: 94.91 F
17:00:13.050 -> Average Cal sensor Temp.: 92.39 Ambient Temp.: 94.95 F
17:00:14.600 -> Average Cal sensor Temp.: 92.71 Ambient Temp.: 94.91 F
17:00:16.158 -> Average Cal sensor Temp.: 92.31 Ambient Temp.: 94.91 F
17:00:17.714 -> Average Cal sensor Temp.: 92.51 Ambient Temp.: 94.91 F
17:00:19.268 -> Average Cal sensor Temp.: 92.51 Ambient Temp.: 94.91 F
17:00:20.814 -> Average Cal sensor Temp.: 92.66 Ambient Temp.: 94.91 F
17:00:22.360 -> Average Cal sensor Temp.: 92.59 Ambient Temp.: 94.95 F
17:00:23.917 -> Average Cal sensor Temp.: 92.69 Ambient Temp.: 94.95 F
17:00:25.506 -> Average Cal sensor Temp.: 92.92 Ambient Temp.: 94.95 F
17:00:27.055 -> Average Cal sensor Temp.: 92.97 Ambient Temp.: 94.95 F
17:00:28.614 -> Average Cal sensor Temp.: 92.95 Ambient Temp.: 94.95 F
17:00:30.169 -> Average Cal sensor Temp.: 92.78 Ambient Temp.: 94.91 F
17:00:31.719 -> Average Cal sensor Temp.: 92.80 Ambient Temp.: 94.91 F
17:00:33.270 -> Average Cal sensor Temp.: 92.75 Ambient Temp.: 94.95 F
17:00:34.822 -> Average Cal sensor Temp.: 92.77 Ambient Temp.: 94.95 F
17:00:36.401 -> Average Cal sensor Temp.: 92.80 Ambient Temp.: 94.95 F
17:00:37.955 -> Average Cal sensor Temp.: 92.42 Ambient Temp.: 94.95 F
17:00:39.510 -> Average Cal sensor Temp.: 92.63 Ambient Temp.: 94.95 F
17:00:41.063 -> Average Cal sensor Temp.: 92.56 Ambient Temp.: 94.95 F
17:00:42.613 -> Average Cal sensor Temp.: 92.61 Ambient Temp.: 94.95 F
17:00:44.165 -> Average Cal sensor Temp.: 92.50 Ambient Temp.: 94.91 F
17:00:45.712 -> Average Cal sensor Temp.: 92.59 Ambient Temp.: 94.95 F
17:00:47.265 -> Average Cal sensor Temp.: 92.69 Ambient Temp.: 94.95 F
17:00:48.851 -> Average Cal sensor Temp.: 92.66 Ambient Temp.: 94.91 F
17:00:50.400 -> Average Cal sensor Temp.: 92.73 Ambient Temp.: 95.02 F
17:00:51.955 -> Average Cal sensor Temp.: 92.84 Ambient Temp.: 94.98 F
17:00:53.506 -> Average Cal sensor Temp.: 92.20 Ambient Temp.: 94.95 F
17:00:55.068 -> Average Cal sensor Temp.: 92.52 Ambient Temp.: 94.95 F
17:00:56.624 -> Average Cal sensor Temp.: 92.47 Ambient Temp.: 94.98 F
17:00:58.183 -> Average Cal sensor Temp.: 92.48 Ambient Temp.: 95.02 F
17:00:59.736 -> Average Cal sensor Temp.: 92.50 Ambient Temp.: 94.98 F
17:01:01.287 -> Average Cal sensor Temp.: 92.58 Ambient Temp.: 94.98 F
17:01:02.831 -> Average Cal sensor Temp.: 92.57 Ambient Temp.: 94.98 F
17:01:04.390 -> Average Cal sensor Temp.: 92.59 Ambient Temp.: 94.95 F
17:01:05.971 -> Average Cal sensor Temp.: 92.95 Ambient Temp.: 94.95 F
17:01:07.522 -> Average Cal sensor Temp.: 93.00 Ambient Temp.: 94.95 F
17:01:09.077 -> Average Cal sensor Temp.: 93.17 Ambient Temp.: 94.98 F
17:01:10.629 -> Average Cal sensor Temp.: 93.32 Ambient Temp.: 94.95 F
17:01:12.181 -> Average Cal sensor Temp.: 92.96 Ambient Temp.: 94.95 F
17:01:13.733 -> Average Cal sensor Temp.: 93.01 Ambient Temp.: 94.95 F
17:01:15.284 -> Average Cal sensor Temp.: 93.00 Ambient Temp.: 94.95 F
17:01:16.840 -> Average Cal sensor Temp.: 92.64 Ambient Temp.: 94.95 F
17:01:18.429 -> Average Cal sensor Temp.: 92.96 Ambient Temp.: 94.98 F
17:01:19.986 -> Average Cal sensor Temp.: 93.10 Ambient Temp.: 94.98 F
17:01:21.541 -> Average Cal sensor Temp.: 92.96 Ambient Temp.: 94.98 F
17:01:23.089 -> Average Cal sensor Temp.: 92.75 Ambient Temp.: 94.95 F
17:01:24.645 -> Average Cal sensor Temp.: 92.45 Ambient Temp.: 94.98 F
17:01:26.203 -> Average Cal sensor Temp.: 92.62 Ambient Temp.: 94.95 F
17:01:27.761 -> Average Cal sensor Temp.: 92.70 Ambient Temp.: 94.95 F
17:01:29.308 -> Average Cal sensor Temp.: 92.97 Ambient Temp.: 94.98 F
17:01:30.863 -> Average Cal sensor Temp.: 92.91 Ambient Temp.: 94.95 F
17:01:32.416 -> Average Cal sensor Temp.: 92.90 Ambient Temp.: 94.95 F
17:01:33.974 -> Average Cal sensor Temp.: 92.48 Ambient Temp.: 94.98 F
17:01:35.533 -> Average Cal sensor Temp.: 92.79 Ambient Temp.: 94.98 F
17:01:37.086 -> Average Cal sensor Temp.: 92.51 Ambient Temp.: 94.95 F
17:01:38.664 -> Average Cal sensor Temp.: 92.42 Ambient Temp.: 95.02 F
17:01:40.211 -> Average Cal sensor Temp.: 92.77 Ambient Temp.: 94.95 F
17:01:41.763 -> Average Cal sensor Temp.: 92.93 Ambient Temp.: 94.98 F
17:01:43.312 -> Average Cal sensor Temp.: 92.88 Ambient Temp.: 95.02 F
17:01:44.865 -> Average Cal sensor Temp.: 93.00 Ambient Temp.: 95.02 F
17:01:46.449 -> Average Cal sensor Temp.: 93.00 Ambient Temp.: 94.95 F
17:01:48.003 -> Average Cal sensor Temp.: 92.98 Ambient Temp.: 95.02 F
17:01:49.552 -> Average Cal sensor Temp.: 92.45 Ambient Temp.: 94.98 F
17:01:51.103 -> Average Cal sensor Temp.: 92.81 Ambient Temp.: 94.95 F
17:01:52.650 -> Average Cal sensor Temp.: 92.94 Ambient Temp.: 94.98 F
17:01:54.205 -> Average Cal sensor Temp.: 92.99 Ambient Temp.: 95.02 F
17:01:55.789 -> Average Cal sensor Temp.: 93.06 Ambient Temp.: 94.98 F
17:01:57.348 -> Average Cal sensor Temp.: 92.88 Ambient Temp.: 94.95 F
17:01:58.903 -> Average Cal sensor Temp.: 93.01 Ambient Temp.: 94.98 F
17:02:00.458 -> Average Cal sensor Temp.: 92.97 Ambient Temp.: 94.98 F
17:02:02.012 -> Average Cal sensor Temp.: 93.01 Ambient Temp.: 95.02 F
17:02:03.568 -> Average Cal sensor Temp.: 92.93 Ambient Temp.: 94.98 F
17:02:05.125 -> Average Cal sensor Temp.: 93.12 Ambient Temp.: 95.02 F
17:02:06.679 -> Average Cal sensor Temp.: 93.17 Ambient Temp.: 94.95 F
17:02:08.234 -> Average Cal sensor Temp.: 93.34 Ambient Temp.: 94.98 F
17:02:09.793 -> Average Cal sensor Temp.: 93.21 Ambient Temp.: 94.98 F
17:02:11.339 -> Average Cal sensor Temp.: 93.44 Ambient Temp.: 95.02 F
17:02:12.894 -> Average Cal sensor Temp.: 93.22 Ambient Temp.: 95.02 F
17:02:14.446 -> Average Cal sensor Temp.: 93.22 Ambient Temp.: 95.02 F
17:02:16.002 -> Average Cal sensor Temp.: 93.46 Ambient Temp.: 95.02 F
17:02:17.556 -> Average Cal sensor Temp.: 93.49 Ambient Temp.: 95.02 F
17:02:19.114 -> Average Cal sensor Temp.: 93.31 Ambient Temp.: 95.02 F
17:02:20.692 -> Average Cal sensor Temp.: 93.08 Ambient Temp.: 95.02 F
17:02:22.232 -> Average Cal sensor Temp.: 93.32 Ambient Temp.: 95.05 F
17:02:23.785 -> Average Cal sensor Temp.: 93.50 Ambient Temp.: 95.02 F
17:02:25.336 -> Average Cal sensor Temp.: 93.08 Ambient Temp.: 95.05 F
17:02:26.888 -> Average Cal sensor Temp.: 92.78 Ambient Temp.: 95.02 F
17:02:28.442 -> Average Cal sensor Temp.: 93.00 Ambient Temp.: 94.98 F
17:02:30.027 -> Average Cal sensor Temp.: 92.95 Ambient Temp.: 95.02 F
17:02:31.575 -> Average Cal sensor Temp.: 93.15 Ambient Temp.: 95.02 F
17:02:33.123 -> Average Cal sensor Temp.: 93.25 Ambient Temp.: 95.02 F
17:02:34.673 -> Average Cal sensor Temp.: 93.20 Ambient Temp.: 95.02 F
17:02:36.227 -> Average Cal sensor Temp.: 93.20 Ambient Temp.: 95.05 F
17:02:37.778 -> Average Cal sensor Temp.: 92.92 Ambient Temp.: 95.02 F
17:02:39.336 -> Average Cal sensor Temp.: 93.07 Ambient Temp.: 94.98 F
17:02:40.892 -> Average Cal sensor Temp.: 93.11 Ambient Temp.: 95.02 F
17:02:42.448 -> Average Cal sensor Temp.: 92.90 Ambient Temp.: 95.02 F
17:02:43.999 -> Average Cal sensor Temp.: 92.93 Ambient Temp.: 95.09 F
17:02:45.556 -> Average Cal sensor Temp.: 92.81 Ambient Temp.: 95.02 F
17:02:47.112 -> Average Cal sensor Temp.: 92.87 Ambient Temp.: 95.02 F
17:02:48.668 -> Average Cal sensor Temp.: 92.75 Ambient Temp.: 95.05 F
17:02:50.226 -> Average Cal sensor Temp.: 92.95 Ambient Temp.: 95.05 F
17:02:51.810 -> Average Cal sensor Temp.: 92.77 Ambient Temp.: 95.02 F
17:02:53.334 -> Average Cal sensor Temp.: 93.09 Ambient Temp.: 95.05 F
17:02:54.893 -> Average Cal sensor Temp.: 93.11 Ambient Temp.: 95.05 F
17:02:56.449 -> Average Cal sensor Temp.: 93.22 Ambient Temp.: 95.05 F
17:02:57.998 -> Average Cal sensor Temp.: 93.33 Ambient Temp.: 95.05 F
17:02:59.583 -> Average Cal sensor Temp.: 93.15 Ambient Temp.: 95.02 F
17:03:01.134 -> Average Cal sensor Temp.: 93.46 Ambient Temp.: 95.02 F
17:03:02.685 -> Average Cal sensor Temp.: 93.23 Ambient Temp.: 95.05 F
17:03:04.238 -> Average Cal sensor Temp.: 93.18 Ambient Temp.: 95.05 F
17:03:05.784 -> Average Cal sensor Temp.: 93.10 Ambient Temp.: 95.02 F
17:03:07.337 -> Average Cal sensor Temp.: 93.19 Ambient Temp.: 95.02 F
17:03:08.894 -> Average Cal sensor Temp.: 93.16 Ambient Temp.: 95.02 F
17:03:10.446 -> Average Cal sensor Temp.: 93.24 Ambient Temp.: 95.02 F
17:03:12.000 -> Average Cal sensor Temp.: 93.47 Ambient Temp.: 94.98 F
17:03:13.556 -> Average Cal sensor Temp.: 93.39 Ambient Temp.: 95.09 F
17:03:15.112 -> Average Cal sensor Temp.: 93.41 Ambient Temp.: 95.05 F
17:03:16.668 -> Average Cal sensor Temp.: 93.29 Ambient Temp.: 95.05 F
17:03:18.252 -> Average Cal sensor Temp.: 93.20 Ambient Temp.: 94.98 F
17:03:19.803 -> Average Cal sensor Temp.: 93.04 Ambient Temp.: 95.02 F
17:03:21.358 -> Average Cal sensor Temp.: 92.93 Ambient Temp.: 95.05 F
17:03:22.917 -> Average Cal sensor Temp.: 92.97 Ambient Temp.: 95.02 F
17:03:24.467 -> Average Cal sensor Temp.: 92.93 Ambient Temp.: 95.05 F
17:03:26.023 -> Average Cal sensor Temp.: 92.86 Ambient Temp.: 95.02 F
17:03:27.568 -> Average Cal sensor Temp.: 92.63 Ambient Temp.: 95.02 F
17:03:29.121 -> Average Cal sensor Temp.: 92.63 Ambient Temp.: 95.05 F
17:03:30.673 -> Average Cal sensor Temp.: 92.65 Ambient Temp.: 95.05 F
17:03:32.232 -> Average Cal sensor Temp.: 92.90 Ambient Temp.: 95.02 F
17:03:33.784 -> Average Cal sensor Temp.: 92.98 Ambient Temp.: 95.05 F
17:03:35.341 -> Average Cal sensor Temp.: 93.25 Ambient Temp.: 95.05 F
17:03:36.926 -> Average Cal sensor Temp.: 93.37 Ambient Temp.: 95.02 F
17:03:38.483 -> Average Cal sensor Temp.: 93.10 Ambient Temp.: 95.09 F
17:03:40.034 -> Average Cal sensor Temp.: 93.10 Ambient Temp.: 95.02 F
17:03:41.589 -> Average Cal sensor Temp.: 93.13 Ambient Temp.: 95.09 F
17:03:43.146 -> Average Cal sensor Temp.: 92.67 Ambient Temp.: 95.05 F
17:03:44.704 -> Average Cal sensor Temp.: 92.87 Ambient Temp.: 95.02 F
17:03:46.256 -> Average Cal sensor Temp.: 93.01 Ambient Temp.: 95.02 F
17:03:47.810 -> Average Cal sensor Temp.: 92.85 Ambient Temp.: 95.05 F
17:03:49.357 -> Average Cal sensor Temp.: 92.83 Ambient Temp.: 95.02 F
17:03:50.914 -> Average Cal sensor Temp.: 92.96 Ambient Temp.: 95.05 F
17:03:52.467 -> Average Cal sensor Temp.: 93.01 Ambient Temp.: 95.05 F
17:03:54.019 -> Average Cal sensor Temp.: 93.02 Ambient Temp.: 95.05 F
17:03:55.576 -> Average Cal sensor Temp.: 93.19 Ambient Temp.: 95.13 F
17:03:57.132 -> Average Cal sensor Temp.: 92.79 Ambient Temp.: 95.05 F
17:03:58.695 -> Average Cal sensor Temp.: 92.99 Ambient Temp.: 95.02 F
17:04:00.253 -> Average Cal sensor Temp.: 93.22 Ambient Temp.: 95.05 F
17:04:01.808 -> Average Cal sensor Temp.: 93.40 Ambient Temp.: 95.05 F
17:04:03.361 -> Average Cal sensor Temp.: 93.07 Ambient Temp.: 95.02 F
17:04:04.921 -> Average Cal sensor Temp.: 92.92 Ambient Temp.: 95.09 F
17:04:06.477 -> Average Cal sensor Temp.: 92.50 Ambient Temp.: 95.05 F
17:04:08.035 -> Average Cal sensor Temp.: 93.02 Ambient Temp.: 95.02 F
17:04:09.588 -> Average Cal sensor Temp.: 93.04 Ambient Temp.: 95.05 F
17:04:11.141 -> Average Cal sensor Temp.: 93.09 Ambient Temp.: 95.05 F
17:04:12.694 -> Average Cal sensor Temp.: 92.90 Ambient Temp.: 95.13 F
17:04:14.251 -> Average Cal sensor Temp.: 93.08 Ambient Temp.: 95.09 F
17:04:15.800 -> Average Cal sensor Temp.: 93.06 Ambient Temp.: 95.05 F
17:04:17.349 -> Average Cal sensor Temp.: 93.02 Ambient Temp.: 95.05 F
17:04:18.909 -> Average Cal sensor Temp.: 93.37 Ambient Temp.: 95.05 F
17:04:20.466 -> Average Cal sensor Temp.: 93.57 Ambient Temp.: 95.09 F
17:04:22.025 -> Average Cal sensor Temp.: 93.46 Ambient Temp.: 95.05 F
17:04:23.581 -> Average Cal sensor Temp.: 93.45 Ambient Temp.: 95.05 F
17:04:25.138 -> Average Cal sensor Temp.: 93.54 Ambient Temp.: 95.05 F
17:04:26.694 -> Average Cal sensor Temp.: 93.11 Ambient Temp.: 95.13 F
17:04:28.254 -> Average Cal sensor Temp.: 93.18 Ambient Temp.: 95.05 F
17:04:29.811 -> Average Cal sensor Temp.: 93.52 Ambient Temp.: 95.05 F
17:04:31.367 -> Average Cal sensor Temp.: 93.29 Ambient Temp.: 95.09 F
17:04:32.925 -> Average Cal sensor Temp.: 93.64 Ambient Temp.: 95.09 F
17:04:34.470 -> Average Cal sensor Temp.: 93.63 Ambient Temp.: 95.05 F
17:04:36.026 -> Average Cal sensor Temp.: 93.43 Ambient Temp.: 95.05 F
17:04:37.579 -> Average Cal sensor Temp.: 93.24 Ambient Temp.: 95.09 F
17:04:39.138 -> Average Cal sensor Temp.: 93.22 Ambient Temp.: 95.09 F
17:04:40.696 -> Average Cal sensor Temp.: 93.17 Ambient Temp.: 95.13 F
17:04:42.249 -> Average Cal sensor Temp.: 93.36 Ambient Temp.: 95.05 F
17:04:43.801 -> Average Cal sensor Temp.: 93.43 Ambient Temp.: 95.05 F
17:04:45.358 -> Average Cal sensor Temp.: 93.61 Ambient Temp.: 95.13 F
17:04:46.910 -> Average Cal sensor Temp.: 93.41 Ambient Temp.: 95.09 F
17:04:48.466 -> Average Cal sensor Temp.: 93.59 Ambient Temp.: 95.05 F
17:04:50.025 -> Average Cal sensor Temp.: 93.52 Ambient Temp.: 95.09 F
17:04:51.575 -> Average Cal sensor Temp.: 93.62 Ambient Temp.: 95.13 F
17:04:53.129 -> Average Cal sensor Temp.: 93.20 Ambient Temp.: 95.05 F
17:04:54.684 -> Average Cal sensor Temp.: 93.21 Ambient Temp.: 95.05 F
17:04:56.236 -> Average Cal sensor Temp.: 92.98 Ambient Temp.: 95.05 F
17:04:57.793 -> Average Cal sensor Temp.: 93.35 Ambient Temp.: 95.13 F
17:04:59.344 -> Average Cal sensor Temp.: 92.89 Ambient Temp.: 95.09 F
17:05:00.903 -> Average Cal sensor Temp.: 92.65 Ambient Temp.: 95.05 F
17:05:02.458 -> Average Cal sensor Temp.: 92.92 Ambient Temp.: 95.09 F
17:05:04.043 -> Average Cal sensor Temp.: 92.92 Ambient Temp.: 95.13 F
17:05:05.598 -> Average Cal sensor Temp.: 93.02 Ambient Temp.: 95.05 F
17:05:07.151 -> Average Cal sensor Temp.: 92.93 Ambient Temp.: 95.05 F
17:05:08.702 -> Average Cal sensor Temp.: 92.99 Ambient Temp.: 95.05 F
17:05:10.257 -> Average Cal sensor Temp.: 92.48 Ambient Temp.: 95.05 F
17:05:11.808 -> Average Cal sensor Temp.: 92.76 Ambient Temp.: 95.05 F
17:05:13.362 -> Average Cal sensor Temp.: 93.12 Ambient Temp.: 95.02 F
17:05:14.917 -> Average Cal sensor Temp.: 93.30 Ambient Temp.: 95.09 F
17:05:16.476 -> Average Cal sensor Temp.: 93.13 Ambient Temp.: 95.02 F
17:05:18.029 -> Average Cal sensor Temp.: 93.53 Ambient Temp.: 95.09 F
17:05:19.582 -> Average Cal sensor Temp.: 93.40 Ambient Temp.: 95.09 F
17:05:21.138 -> Average Cal sensor Temp.: 93.39 Ambient Temp.: 95.09 F
17:05:22.693 -> Average Cal sensor Temp.: 93.62 Ambient Temp.: 95.09 F
17:05:24.247 -> Average Cal sensor Temp.: 93.67 Ambient Temp.: 95.09 F
17:05:25.807 -> Average Cal sensor Temp.: 92.67 Ambient Temp.: 95.05 F
17:05:27.360 -> Average Cal sensor Temp.: 93.00 Ambient Temp.: 95.05 F
17:05:28.915 -> Average Cal sensor Temp.: 93.45 Ambient Temp.: 95.05 F
17:05:30.470 -> Average Cal sensor Temp.: 93.70 Ambient Temp.: 95.09 F
17:05:32.025 -> Average Cal sensor Temp.: 93.80 Ambient Temp.: 95.13 F
17:05:33.577 -> Average Cal sensor Temp.: 93.80 Ambient Temp.: 95.09 F
17:05:35.136 -> Average Cal sensor Temp.: 93.56 Ambient Temp.: 95.09 F
17:05:36.723 -> Average Cal sensor Temp.: 93.74 Ambient Temp.: 95.09 F
17:05:38.272 -> Average Cal sensor Temp.: 93.68 Ambient Temp.: 95.05 F
17:05:39.825 -> Average Cal sensor Temp.: 93.62 Ambient Temp.: 95.09 F
17:05:41.380 -> Average Cal sensor Temp.: 93.66 Ambient Temp.: 95.09 F
17:05:42.933 -> Average Cal sensor Temp.: 93.70 Ambient Temp.: 95.05 F
17:05:44.486 -> Average Cal sensor Temp.: 93.76 Ambient Temp.: 95.09 F
17:05:46.040 -> Average Cal sensor Temp.: 93.72 Ambient Temp.: 95.13 F
17:05:47.596 -> Average Cal sensor Temp.: 93.52 Ambient Temp.: 95.13 F
17:05:49.151 -> Average Cal sensor Temp.: 93.50 Ambient Temp.: 95.09 F
17:05:50.705 -> Average Cal sensor Temp.: 93.59 Ambient Temp.: 95.09 F
17:05:52.264 -> Average Cal sensor Temp.: 93.72 Ambient Temp.: 95.09 F
17:05:53.816 -> Average Cal sensor Temp.: 93.50 Ambient Temp.: 95.13 F
17:05:55.369 -> Average Cal sensor Temp.: 93.25 Ambient Temp.: 95.05 F
17:05:56.923 -> Average Cal sensor Temp.: 93.64 Ambient Temp.: 95.16 F
17:05:58.479 -> Average Cal sensor Temp.: 93.79 Ambient Temp.: 95.09 F
17:06:00.061 -> Average Cal sensor Temp.: 93.60 Ambient Temp.: 95.13 F
17:06:01.586 -> Average Cal sensor Temp.: 93.30 Ambient Temp.: 95.09 F
17:06:03.175 -> Average Cal sensor Temp.: 93.71 Ambient Temp.: 95.09 F
17:06:04.728 -> Average Cal sensor Temp.: 93.60 Ambient Temp.: 95.13 F
17:06:06.281 -> Average Cal sensor Temp.: 93.67 Ambient Temp.: 95.09 F
17:06:07.838 -> Average Cal sensor Temp.: 93.58 Ambient Temp.: 95.09 F
17:06:09.391 -> Average Cal sensor Temp.: 93.30 Ambient Temp.: 95.16 F
17:06:10.949 -> Average Cal sensor Temp.: 93.07 Ambient Temp.: 95.13 F
17:06:12.478 -> Average Cal sensor Temp.: 92.98 Ambient Temp.: 95.13 F
17:06:14.031 -> Average Cal sensor Temp.: 92.85 Ambient Temp.: 95.09 F
17:06:15.589 -> Average Cal sensor Temp.: 93.07 Ambient Temp.: 95.16 F
17:06:17.142 -> Average Cal sensor Temp.: 92.96 Ambient Temp.: 95.16 F
17:06:18.699 -> Average Cal sensor Temp.: 92.80 Ambient Temp.: 95.13 F
17:06:20.286 -> Average Cal sensor Temp.: 92.84 Ambient Temp.: 95.16 F
17:06:21.844 -> Average Cal sensor Temp.: 93.10 Ambient Temp.: 95.09 F
17:06:23.395 -> Average Cal sensor Temp.: 92.80 Ambient Temp.: 95.09 F
17:06:24.950 -> Average Cal sensor Temp.: 92.86 Ambient Temp.: 95.09 F
17:06:26.503 -> Average Cal sensor Temp.: 93.10 Ambient Temp.: 95.09 F
17:06:28.048 -> Average Cal sensor Temp.: 92.96 Ambient Temp.: 95.16 F
17:06:29.597 -> Average Cal sensor Temp.: 92.93 Ambient Temp.: 95.13 F
17:06:31.181 -> Average Cal sensor Temp.: 93.09 Ambient Temp.: 95.09 F
17:06:32.737 -> Average Cal sensor Temp.: 93.35 Ambient Temp.: 95.13 F
17:06:34.291 -> Average Cal sensor Temp.: 93.36 Ambient Temp.: 95.09 F
17:06:35.841 -> Average Cal sensor Temp.: 93.54 Ambient Temp.: 95.09 F
17:06:37.391 -> Average Cal sensor Temp.: 93.69 Ambient Temp.: 95.09 F
17:06:38.945 -> Average Cal sensor Temp.: 93.24 Ambient Temp.: 95.09 F
17:06:40.499 -> Average Cal sensor Temp.: 92.79 Ambient Temp.: 95.09 F
17:06:42.054 -> Average Cal sensor Temp.: 93.53 Ambient Temp.: 95.09 F
17:06:43.608 -> Average Cal sensor Temp.: 93.33 Ambient Temp.: 95.09 F
17:06:45.162 -> Average Cal sensor Temp.: 93.28 Ambient Temp.: 95.09 F
17:06:46.712 -> Average Cal sensor Temp.: 93.44 Ambient Temp.: 95.09 F
17:06:48.272 -> Average Cal sensor Temp.: 93.59 Ambient Temp.: 95.09 F
17:06:49.827 -> Average Cal sensor Temp.: 93.54 Ambient Temp.: 95.09 F
17:06:51.381 -> Average Cal sensor Temp.: 93.26 Ambient Temp.: 95.13 F
17:06:52.963 -> Average Cal sensor Temp.: 93.68 Ambient Temp.: 95.09 F
17:06:54.517 -> Average Cal sensor Temp.: 93.29 Ambient Temp.: 95.09 F
17:06:56.065 -> Average Cal sensor Temp.: 93.56 Ambient Temp.: 95.05 F
17:06:57.614 -> Average Cal sensor Temp.: 93.56 Ambient Temp.: 95.09 F
17:06:59.170 -> Average Cal sensor Temp.: 93.81 Ambient Temp.: 95.16 F
17:07:00.727 -> Average Cal sensor Temp.: 94.04 Ambient Temp.: 95.16 F
17:07:02.279 -> Average Cal sensor Temp.: 93.97 Ambient Temp.: 95.09 F
17:07:03.831 -> Average Cal sensor Temp.: 94.01 Ambient Temp.: 95.13 F
17:07:05.417 -> Average Cal sensor Temp.: 93.92 Ambient Temp.: 95.13 F
17:07:06.968 -> Average Cal sensor Temp.: 93.71 Ambient Temp.: 95.13 F
17:07:08.525 -> Average Cal sensor Temp.: 93.79 Ambient Temp.: 95.13 F
17:07:10.075 -> Average Cal sensor Temp.: 93.75 Ambient Temp.: 95.13 F
17:07:11.628 -> Average Cal sensor Temp.: 93.91 Ambient Temp.: 95.16 F
17:07:13.178 -> Average Cal sensor Temp.: 93.43 Ambient Temp.: 95.09 F
17:07:14.736 -> Average Cal sensor Temp.: 93.59 Ambient Temp.: 95.13 F
17:07:16.291 -> Average Cal sensor Temp.: 93.62 Ambient Temp.: 95.13 F
17:07:17.841 -> Average Cal sensor Temp.: 93.74 Ambient Temp.: 95.09 F
17:07:19.396 -> Average Cal sensor Temp.: 93.64 Ambient Temp.: 95.09 F
17:07:20.943 -> Average Cal sensor Temp.: 93.31 Ambient Temp.: 95.16 F
17:07:22.498 -> Average Cal sensor Temp.: 93.69 Ambient Temp.: 95.16 F
17:07:24.049 -> Average Cal sensor Temp.: 93.74 Ambient Temp.: 95.16 F
17:07:25.636 -> Average Cal sensor Temp.: 93.60 Ambient Temp.: 95.13 F
17:07:27.185 -> Average Cal sensor Temp.: 93.40 Ambient Temp.: 95.13 F
17:07:28.737 -> Average Cal sensor Temp.: 93.56 Ambient Temp.: 95.13 F
17:07:30.290 -> Average Cal sensor Temp.: 93.70 Ambient Temp.: 95.16 F
17:07:31.838 -> Average Cal sensor Temp.: 93.47 Ambient Temp.: 95.13 F
17:07:33.392 -> Average Cal sensor Temp.: 93.38 Ambient Temp.: 95.16 F
17:07:34.949 -> Average Cal sensor Temp.: 93.76 Ambient Temp.: 95.09 F
17:07:36.503 -> Average Cal sensor Temp.: 93.66 Ambient Temp.: 95.13 F
17:07:38.058 -> Average Cal sensor Temp.: 93.64 Ambient Temp.: 95.13 F
17:07:39.604 -> Average Cal sensor Temp.: 93.39 Ambient Temp.: 95.16 F
17:07:41.158 -> Average Cal sensor Temp.: 93.69 Ambient Temp.: 95.13 F
17:07:42.745 -> Average Cal sensor Temp.: 93.74 Ambient Temp.: 95.16 F
17:07:44.299 -> Average Cal sensor Temp.: 93.12 Ambient Temp.: 95.16 F
17:07:45.854 -> Average Cal sensor Temp.: 93.44 Ambient Temp.: 95.16 F
17:07:47.408 -> Average Cal sensor Temp.: 93.50 Ambient Temp.: 95.13 F
17:07:48.959 -> Average Cal sensor Temp.: 93.78 Ambient Temp.: 95.16 F
17:07:50.508 -> Average Cal sensor Temp.: 93.74 Ambient Temp.: 95.09 F
17:07:52.062 -> Average Cal sensor Temp.: 93.60 Ambient Temp.: 95.16 F
17:07:53.619 -> Average Cal sensor Temp.: 93.55 Ambient Temp.: 95.13 F
17:07:55.173 -> Average Cal sensor Temp.: 93.74 Ambient Temp.: 95.13 F
17:07:56.731 -> Average Cal sensor Temp.: 93.73 Ambient Temp.: 95.16 F
17:07:58.281 -> Average Cal sensor Temp.: 93.89 Ambient Temp.: 95.16 F
17:07:59.834 -> Average Cal sensor Temp.: 93.96 Ambient Temp.: 95.09 F
17:08:01.390 -> Average Cal sensor Temp.: 93.91 Ambient Temp.: 95.16 F
17:08:02.943 -> Average Cal sensor Temp.: 93.87 Ambient Temp.: 95.20 F
17:08:04.492 -> Average Cal sensor Temp.: 93.82 Ambient Temp.: 95.16 F
17:08:06.076 -> Average Cal sensor Temp.: 93.84 Ambient Temp.: 95.16 F
17:08:07.627 -> Average Cal sensor Temp.: 93.95 Ambient Temp.: 95.27 F
17:08:09.179 -> Average Cal sensor Temp.: 93.98 Ambient Temp.: 95.16 F
17:08:10.733 -> Average Cal sensor Temp.: 94.15 Ambient Temp.: 95.16 F
17:08:12.286 -> Average Cal sensor Temp.: 94.10 Ambient Temp.: 95.16 F
17:08:13.835 -> Average Cal sensor Temp.: 93.70 Ambient Temp.: 95.20 F
17:08:15.386 -> Average Cal sensor Temp.: 93.70 Ambient Temp.: 95.20 F
17:08:16.942 -> Average Cal sensor Temp.: 93.92 Ambient Temp.: 95.20 F
17:08:18.498 -> Average Cal sensor Temp.: 93.88 Ambient Temp.: 95.16 F
17:08:20.048 -> Average Cal sensor Temp.: 93.63 Ambient Temp.: 95.16 F
17:08:21.603 -> Average Cal sensor Temp.: 93.73 Ambient Temp.: 95.20 F
17:08:23.191 -> Average Cal sensor Temp.: 94.02 Ambient Temp.: 95.20 F
17:08:24.743 -> Average Cal sensor Temp.: 94.05 Ambient Temp.: 95.20 F
17:08:26.301 -> Average Cal sensor Temp.: 93.88 Ambient Temp.: 95.20 F
17:08:27.854 -> Average Cal sensor Temp.: 93.93 Ambient Temp.: 95.16 F
17:08:29.407 -> Average Cal sensor Temp.: 93.96 Ambient Temp.: 95.20 F
17:08:30.957 -> Average Cal sensor Temp.: 93.79 Ambient Temp.: 95.20 F
17:08:32.514 -> Average Cal sensor Temp.: 93.92 Ambient Temp.: 95.20 F
17:08:34.068 -> Average Cal sensor Temp.: 93.97 Ambient Temp.: 95.20 F
17:08:35.623 -> Average Cal sensor Temp.: 93.86 Ambient Temp.: 95.27 F
17:08:37.174 -> Average Cal sensor Temp.: 93.68 Ambient Temp.: 95.20 F
17:08:38.724 -> Average Cal sensor Temp.: 93.92 Ambient Temp.: 95.20 F
17:08:40.277 -> Average Cal sensor Temp.: 94.01 Ambient Temp.: 95.27 F
17:08:41.863 -> Average Cal sensor Temp.: 93.83 Ambient Temp.: 95.16 F
17:08:43.413 -> Average Cal sensor Temp.: 93.83 Ambient Temp.: 95.16 F
17:08:44.965 -> Average Cal sensor Temp.: 93.94 Ambient Temp.: 95.27 F
17:08:46.516 -> Average Cal sensor Temp.: 93.84 Ambient Temp.: 95.20 F
17:08:48.067 -> Average Cal sensor Temp.: 93.91 Ambient Temp.: 95.27 F
17:08:49.617 -> Average Cal sensor Temp.: 93.97 Ambient Temp.: 95.31 F
17:08:51.166 -> Average Cal sensor Temp.: 94.05 Ambient Temp.: 95.20 F
17:08:52.752 -> Average Cal sensor Temp.: 93.96 Ambient Temp.: 95.20 F
17:08:54.306 -> Average Cal sensor Temp.: 94.17 Ambient Temp.: 95.27 F
17:08:55.853 -> Average Cal sensor Temp.: 94.32 Ambient Temp.: 95.31 F
17:08:57.407 -> Average Cal sensor Temp.: 94.33 Ambient Temp.: 95.20 F
17:08:58.961 -> Average Cal sensor Temp.: 94.27 Ambient Temp.: 95.27 F
17:09:00.516 -> Average Cal sensor Temp.: 93.90 Ambient Temp.: 95.20 F
17:09:02.059 -> Average Cal sensor Temp.: 93.47 Ambient Temp.: 95.27 F
17:09:03.610 -> Average Cal sensor Temp.: 93.21 Ambient Temp.: 95.27 F
17:09:05.197 -> Average Cal sensor Temp.: 93.18 Ambient Temp.: 95.27 F
17:09:06.751 -> Average Cal sensor Temp.: 93.54 Ambient Temp.: 95.20 F
17:09:08.299 -> Average Cal sensor Temp.: 93.32 Ambient Temp.: 95.27 F
17:09:09.851 -> Average Cal sensor Temp.: 93.55 Ambient Temp.: 95.31 F
17:09:11.399 -> Average Cal sensor Temp.: 93.78 Ambient Temp.: 95.27 F
17:09:12.957 -> Average Cal sensor Temp.: 93.84 Ambient Temp.: 95.20 F
17:09:14.507 -> Average Cal sensor Temp.: 94.08 Ambient Temp.: 95.20 F
17:09:16.062 -> Average Cal sensor Temp.: 93.97 Ambient Temp.: 95.27 F
17:09:17.612 -> Average Cal sensor Temp.: 93.83 Ambient Temp.: 95.27 F
17:09:19.165 -> Average Cal sensor Temp.: 93.97 Ambient Temp.: 95.34 F
17:09:20.749 -> Average Cal sensor Temp.: 94.20 Ambient Temp.: 95.31 F
17:09:22.302 -> Average Cal sensor Temp.: 93.92 Ambient Temp.: 95.31 F
17:09:23.861 -> Average Cal sensor Temp.: 93.83 Ambient Temp.: 95.34 F
17:09:25.413 -> Average Cal sensor Temp.: 94.01 Ambient Temp.: 95.34 F
17:09:26.970 -> Average Cal sensor Temp.: 93.98 Ambient Temp.: 95.27 F
17:09:28.524 -> Average Cal sensor Temp.: 94.11 Ambient Temp.: 95.31 F
17:09:30.078 -> Average Cal sensor Temp.: 93.80 Ambient Temp.: 95.27 F
17:09:31.634 -> Average Cal sensor Temp.: 93.84 Ambient Temp.: 95.31 F
17:09:33.190 -> Average Cal sensor Temp.: 93.85 Ambient Temp.: 95.27 F
17:09:34.741 -> Average Cal sensor Temp.: 93.56 Ambient Temp.: 95.20 F
17:09:36.287 -> Average Cal sensor Temp.: 93.93 Ambient Temp.: 95.31 F
17:09:37.836 -> Average Cal sensor Temp.: 93.92 Ambient Temp.: 95.31 F
17:09:39.385 -> Average Cal sensor Temp.: 94.23 Ambient Temp.: 95.31 F
17:09:40.939 -> Average Cal sensor Temp.: 94.14 Ambient Temp.: 95.31 F
17:09:42.494 -> Average Cal sensor Temp.: 93.91 Ambient Temp.: 95.31 F
17:09:44.078 -> Average Cal sensor Temp.: 94.07 Ambient Temp.: 95.38 F
17:09:45.628 -> Average Cal sensor Temp.: 93.93 Ambient Temp.: 95.34 F
17:09:47.184 -> Average Cal sensor Temp.: 93.58 Ambient Temp.: 95.34 F
17:09:48.733 -> Average Cal sensor Temp.: 93.82 Ambient Temp.: 95.31 F
17:09:50.287 -> Average Cal sensor Temp.: 93.88 Ambient Temp.: 95.31 F
17:09:51.836 -> Average Cal sensor Temp.: 94.00 Ambient Temp.: 95.31 F
17:09:53.390 -> Average Cal sensor Temp.: 94.00 Ambient Temp.: 95.38 F
17:09:54.971 -> Average Cal sensor Temp.: 94.03 Ambient Temp.: 95.31 F
17:09:56.526 -> Average Cal sensor Temp.: 93.91 Ambient Temp.: 95.34 F
17:09:58.076 -> Average Cal sensor Temp.: 93.97 Ambient Temp.: 95.31 F
17:09:59.635 -> Average Cal sensor Temp.: 93.88 Ambient Temp.: 95.31 F
17:10:01.187 -> Average Cal sensor Temp.: 93.70 Ambient Temp.: 95.34 F
17:10:02.741 -> Average Cal sensor Temp.: 93.84 Ambient Temp.: 95.38 F
17:10:04.295 -> Average Cal sensor Temp.: 93.97 Ambient Temp.: 95.34 F
17:10:05.849 -> Average Cal sensor Temp.: 94.11 Ambient Temp.: 95.34 F
17:10:07.404 -> Average Cal sensor Temp.: 94.04 Ambient Temp.: 95.34 F
17:10:08.956 -> Average Cal sensor Temp.: 94.29 Ambient Temp.: 95.31 F
17:10:10.507 -> Average Cal sensor Temp.: 94.43 Ambient Temp.: 95.34 F
17:10:12.058 -> Average Cal sensor Temp.: 94.36 Ambient Temp.: 95.34 F
17:10:13.613 -> Average Cal sensor Temp.: 94.46 Ambient Temp.: 95.31 F
17:10:15.196 -> Average Cal sensor Temp.: 94.46 Ambient Temp.: 95.31 F
17:10:16.750 -> Average Cal sensor Temp.: 94.37 Ambient Temp.: 95.34 F
17:10:18.304 -> Average Cal sensor Temp.: 94.46 Ambient Temp.: 95.38 F
17:10:19.851 -> Average Cal sensor Temp.: 93.98 Ambient Temp.: 95.38 F
17:10:21.396 -> Average Cal sensor Temp.: 94.24 Ambient Temp.: 95.38 F
17:10:22.980 -> Average Cal sensor Temp.: 94.37 Ambient Temp.: 95.41 F
17:10:24.534 -> Average Cal sensor Temp.: 94.35 Ambient Temp.: 95.38 F
17:10:26.081 -> Average Cal sensor Temp.: 94.17 Ambient Temp.: 95.38 F
17:10:27.636 -> Average Cal sensor Temp.: 93.93 Ambient Temp.: 95.34 F
17:10:29.185 -> Average Cal sensor Temp.: 94.01 Ambient Temp.: 95.34 F
17:10:30.733 -> Average Cal sensor Temp.: 94.01 Ambient Temp.: 95.34 F
17:10:32.314 -> Average Cal sensor Temp.: 94.22 Ambient Temp.: 95.34 F
17:10:33.872 -> Average Cal sensor Temp.: 93.54 Ambient Temp.: 95.34 F
17:10:35.422 -> Average Cal sensor Temp.: 93.96 Ambient Temp.: 95.41 F
17:10:36.977 -> Average Cal sensor Temp.: 94.31 Ambient Temp.: 95.38 F
17:10:38.528 -> Average Cal sensor Temp.: 94.31 Ambient Temp.: 95.38 F
17:10:40.073 -> Average Cal sensor Temp.: 94.38 Ambient Temp.: 95.38 F
17:10:41.630 -> Average Cal sensor Temp.: 94.10 Ambient Temp.: 95.38 F
17:10:43.182 -> Average Cal sensor Temp.: 94.05 Ambient Temp.: 95.31 F
17:10:44.733 -> Average Cal sensor Temp.: 94.03 Ambient Temp.: 95.34 F
17:10:46.289 -> Average Cal sensor Temp.: 94.07 Ambient Temp.: 95.38 F
17:10:47.841 -> Average Cal sensor Temp.: 94.32 Ambient Temp.: 95.38 F
17:10:49.426 -> Average Cal sensor Temp.: 94.22 Ambient Temp.: 95.34 F
17:10:50.985 -> Average Cal sensor Temp.: 93.85 Ambient Temp.: 95.38 F
17:10:52.530 -> Average Cal sensor Temp.: 93.86 Ambient Temp.: 95.34 F
17:10:54.082 -> Average Cal sensor Temp.: 94.10 Ambient Temp.: 95.38 F
17:10:55.629 -> Average Cal sensor Temp.: 94.16 Ambient Temp.: 95.34 F
17:10:57.186 -> Average Cal sensor Temp.: 94.07 Ambient Temp.: 95.38 F
17:10:58.736 -> Average Cal sensor Temp.: 94.25 Ambient Temp.: 95.34 F
17:11:00.321 -> Average Cal sensor Temp.: 94.18 Ambient Temp.: 95.34 F
17:11:01.872 -> Average Cal sensor Temp.: 93.54 Ambient Temp.: 95.38 F
17:11:03.432 -> Average Cal sensor Temp.: 93.84 Ambient Temp.: 95.34 F
17:11:04.981 -> Average Cal sensor Temp.: 93.53 Ambient Temp.: 95.34 F
17:11:06.530 -> Average Cal sensor Temp.: 93.96 Ambient Temp.: 95.38 F
17:11:08.080 -> Average Cal sensor Temp.: 93.91 Ambient Temp.: 95.34 F
17:11:09.630 -> Average Cal sensor Temp.: 93.96 Ambient Temp.: 95.34 F
17:11:11.182 -> Average Cal sensor Temp.: 93.81 Ambient Temp.: 95.34 F
17:11:12.767 -> Average Cal sensor Temp.: 93.93 Ambient Temp.: 95.38 F
17:11:14.317 -> Average Cal sensor Temp.: 94.00 Ambient Temp.: 95.38 F
17:11:15.871 -> Average Cal sensor Temp.: 94.12 Ambient Temp.: 95.41 F
17:11:17.421 -> Average Cal sensor Temp.: 94.02 Ambient Temp.: 95.31 F
17:11:18.973 -> Average Cal sensor Temp.: 94.01 Ambient Temp.: 95.34 F
17:11:20.527 -> Average Cal sensor Temp.: 94.33 Ambient Temp.: 95.34 F
17:11:22.081 -> Average Cal sensor Temp.: 94.10 Ambient Temp.: 95.38 F
17:11:23.633 -> Average Cal sensor Temp.: 94.32 Ambient Temp.: 95.38 F
17:11:25.211 -> Average Cal sensor Temp.: 94.37 Ambient Temp.: 95.38 F
17:11:26.767 -> Average Cal sensor Temp.: 94.04 Ambient Temp.: 95.31 F
17:11:28.320 -> Average Cal sensor Temp.: 93.68 Ambient Temp.: 95.31 F
17:11:29.874 -> Average Cal sensor Temp.: 94.01 Ambient Temp.: 95.34 F
17:11:31.427 -> Average Cal sensor Temp.: 93.99 Ambient Temp.: 95.34 F
17:11:32.979 -> Average Cal sensor Temp.: 94.28 Ambient Temp.: 95.34 F
17:11:34.527 -> Average Cal sensor Temp.: 94.35 Ambient Temp.: 95.34 F
17:11:36.083 -> Average Cal sensor Temp.: 94.39 Ambient Temp.: 95.34 F
17:11:37.638 -> Average Cal sensor Temp.: 94.35 Ambient Temp.: 95.34 F
17:11:39.198 -> Average Cal sensor Temp.: 94.49 Ambient Temp.: 95.34 F
17:11:40.747 -> Average Cal sensor Temp.: 94.45 Ambient Temp.: 95.34 F
17:11:42.307 -> Average Cal sensor Temp.: 94.32 Ambient Temp.: 95.41 F
17:11:43.861 -> Average Cal sensor Temp.: 94.25 Ambient Temp.: 95.38 F
17:11:45.446 -> Average Cal sensor Temp.: 94.32 Ambient Temp.: 95.34 F
17:11:47.000 -> Average Cal sensor Temp.: 94.39 Ambient Temp.: 95.34 F
17:11:48.555 -> Average Cal sensor Temp.: 94.19 Ambient Temp.: 95.38 F
17:11:50.104 -> Average Cal sensor Temp.: 94.35 Ambient Temp.: 95.38 F
17:11:51.655 -> Average Cal sensor Temp.: 94.27 Ambient Temp.: 95.34 F
17:11:53.210 -> Average Cal sensor Temp.: 94.23 Ambient Temp.: 95.38 F
17:11:54.766 -> Average Cal sensor Temp.: 94.29 Ambient Temp.: 95.41 F
17:11:56.320 -> Average Cal sensor Temp.: 94.28 Ambient Temp.: 95.41 F
17:11:57.871 -> Average Cal sensor Temp.: 94.10 Ambient Temp.: 95.38 F
17:11:59.418 -> Average Cal sensor Temp.: 94.16 Ambient Temp.: 95.38 F
17:12:01.001 -> Average Cal sensor Temp.: 93.98 Ambient Temp.: 95.34 F
17:12:02.546 -> Average Cal sensor Temp.: 93.89 Ambient Temp.: 95.34 F
17:12:04.101 -> Average Cal sensor Temp.: 94.23 Ambient Temp.: 95.38 F
17:12:05.656 -> Average Cal sensor Temp.: 93.93 Ambient Temp.: 95.38 F
17:12:07.207 -> Average Cal sensor Temp.: 93.95 Ambient Temp.: 95.38 F
17:12:08.763 -> Average Cal sensor Temp.: 94.00 Ambient Temp.: 95.41 F
17:12:10.319 -> Average Cal sensor Temp.: 93.97 Ambient Temp.: 95.38 F
17:12:11.874 -> Average Cal sensor Temp.: 93.96 Ambient Temp.: 95.38 F
17:12:13.427 -> Average Cal sensor Temp.: 93.87 Ambient Temp.: 95.41 F
17:12:15.011 -> Average Cal sensor Temp.: 94.15 Ambient Temp.: 95.38 F
17:12:16.566 -> Average Cal sensor Temp.: 94.50 Ambient Temp.: 95.41 F
17:12:18.120 -> Average Cal sensor Temp.: 94.36 Ambient Temp.: 95.41 F
17:12:19.672 -> Average Cal sensor Temp.: 94.48 Ambient Temp.: 95.38 F
17:12:21.222 -> Average Cal sensor Temp.: 94.32 Ambient Temp.: 95.38 F
17:12:22.775 -> Average Cal sensor Temp.: 94.37 Ambient Temp.: 95.38 F
17:12:24.334 -> Average Cal sensor Temp.: 94.32 Ambient Temp.: 95.38 F
17:12:25.891 -> Average Cal sensor Temp.: 94.56 Ambient Temp.: 95.38 F
17:12:27.442 -> Average Cal sensor Temp.: 94.74 Ambient Temp.: 95.41 F
17:12:29.000 -> Average Cal sensor Temp.: 94.56 Ambient Temp.: 95.41 F
17:12:30.555 -> Average Cal sensor Temp.: 94.33 Ambient Temp.: 95.38 F
17:12:32.108 -> Average Cal sensor Temp.: 93.98 Ambient Temp.: 95.38 F
17:12:33.659 -> Average Cal sensor Temp.: 94.23 Ambient Temp.: 95.38 F
17:12:35.212 -> Average Cal sensor Temp.: 94.33 Ambient Temp.: 95.41 F
17:12:36.771 -> Average Cal sensor Temp.: 94.08 Ambient Temp.: 95.41 F
17:12:38.321 -> Average Cal sensor Temp.: 94.48 Ambient Temp.: 95.38 F
17:12:39.880 -> Average Cal sensor Temp.: 94.48 Ambient Temp.: 95.41 F
17:12:41.465 -> Average Cal sensor Temp.: 94.36 Ambient Temp.: 95.41 F
17:12:43.010 -> Average Cal sensor Temp.: 94.19 Ambient Temp.: 95.41 F
17:12:44.568 -> Average Cal sensor Temp.: 93.90 Ambient Temp.: 95.38 F
17:12:46.119 -> Average Cal sensor Temp.: 94.22 Ambient Temp.: 95.41 F
17:12:47.672 -> Average Cal sensor Temp.: 94.37 Ambient Temp.: 95.41 F
17:12:49.222 -> Average Cal sensor Temp.: 94.41 Ambient Temp.: 95.38 F
17:12:50.771 -> Average Cal sensor Temp.: 94.64 Ambient Temp.: 95.41 F
17:12:52.358 -> Average Cal sensor Temp.: 94.99 Ambient Temp.: 95.41 F
17:12:53.907 -> Average Cal sensor Temp.: 94.94 Ambient Temp.: 95.41 F
17:12:55.463 -> Average Cal sensor Temp.: 94.88 Ambient Temp.: 95.38 F
17:12:57.014 -> Average Cal sensor Temp.: 94.82 Ambient Temp.: 95.34 F
17:12:58.567 -> Average Cal sensor Temp.: 94.54 Ambient Temp.: 95.38 F
17:13:00.114 -> Average Cal sensor Temp.: 94.31 Ambient Temp.: 95.41 F
17:13:01.666 -> Average Cal sensor Temp.: 94.27 Ambient Temp.: 95.41 F
17:13:03.250 -> Average Cal sensor Temp.: 94.14 Ambient Temp.: 95.41 F
17:13:04.808 -> Average Cal sensor Temp.: 94.23 Ambient Temp.: 95.45 F
17:13:06.364 -> Average Cal sensor Temp.: 94.57 Ambient Temp.: 95.45 F
17:13:07.917 -> Average Cal sensor Temp.: 94.36 Ambient Temp.: 95.41 F
17:13:09.476 -> Average Cal sensor Temp.: 94.46 Ambient Temp.: 95.41 F
17:13:11.030 -> Average Cal sensor Temp.: 94.82 Ambient Temp.: 95.41 F
17:13:12.584 -> Average Cal sensor Temp.: 94.41 Ambient Temp.: 95.45 F
17:13:14.131 -> Average Cal sensor Temp.: 93.82 Ambient Temp.: 95.41 F
17:13:15.691 -> Average Cal sensor Temp.: 94.36 Ambient Temp.: 95.38 F
17:13:17.244 -> Average Cal sensor Temp.: 94.66 Ambient Temp.: 95.38 F
17:13:18.794 -> Average Cal sensor Temp.: 94.41 Ambient Temp.: 95.41 F
17:13:20.344 -> Average Cal sensor Temp.: 94.32 Ambient Temp.: 95.41 F
17:13:21.897 -> Average Cal sensor Temp.: 93.98 Ambient Temp.: 95.38 F
17:13:23.452 -> Average Cal sensor Temp.: 93.93 Ambient Temp.: 95.41 F
17:13:25.013 -> Average Cal sensor Temp.: 94.23 Ambient Temp.: 95.41 F
17:13:26.564 -> Average Cal sensor Temp.: 94.25 Ambient Temp.: 95.45 F
17:13:28.121 -> Average Cal sensor Temp.: 94.05 Ambient Temp.: 95.38 F
17:13:29.678 -> Average Cal sensor Temp.: 94.19 Ambient Temp.: 95.38 F
17:13:31.228 -> Average Cal sensor Temp.: 94.36 Ambient Temp.: 95.38 F
17:13:32.784 -> Average Cal sensor Temp.: 94.38 Ambient Temp.: 95.41 F
17:13:34.368 -> Average Cal sensor Temp.: 94.25 Ambient Temp.: 95.38 F
17:13:35.921 -> Average Cal sensor Temp.: 94.49 Ambient Temp.: 95.45 F
17:13:37.477 -> Average Cal sensor Temp.: 94.75 Ambient Temp.: 95.41 F
17:13:39.029 -> Average Cal sensor Temp.: 94.26 Ambient Temp.: 95.41 F
17:13:40.586 -> Average Cal sensor Temp.: 94.53 Ambient Temp.: 95.45 F
17:13:42.135 -> Average Cal sensor Temp.: 94.79 Ambient Temp.: 95.45 F
17:13:43.685 -> Average Cal sensor Temp.: 94.65 Ambient Temp.: 95.49 F
17:13:45.265 -> Average Cal sensor Temp.: 94.77 Ambient Temp.: 95.38 F
17:13:46.816 -> Average Cal sensor Temp.: 94.86 Ambient Temp.: 95.38 F
17:13:48.368 -> Average Cal sensor Temp.: 94.74 Ambient Temp.: 95.45 F
17:13:49.919 -> Average Cal sensor Temp.: 94.75 Ambient Temp.: 95.41 F
17:13:51.471 -> Average Cal sensor Temp.: 94.82 Ambient Temp.: 95.41 F
17:13:53.025 -> Average Cal sensor Temp.: 94.73 Ambient Temp.: 95.41 F
17:13:54.577 -> Average Cal sensor Temp.: 94.86 Ambient Temp.: 95.41 F
17:13:56.136 -> Average Cal sensor Temp.: 94.87 Ambient Temp.: 95.41 F
17:13:57.689 -> Average Cal sensor Temp.: 94.24 Ambient Temp.: 95.41 F
17:13:59.244 -> Average Cal sensor Temp.: 94.61 Ambient Temp.: 95.41 F
17:14:00.825 -> Average Cal sensor Temp.: 94.63 Ambient Temp.: 95.41 F
17:14:02.379 -> Average Cal sensor Temp.: 94.43 Ambient Temp.: 95.49 F
17:14:03.933 -> Average Cal sensor Temp.: 94.44 Ambient Temp.: 95.38 F
17:14:05.486 -> Average Cal sensor Temp.: 94.34 Ambient Temp.: 95.49 F
17:14:07.039 -> Average Cal sensor Temp.: 93.88 Ambient Temp.: 95.45 F
17:14:08.594 -> Average Cal sensor Temp.: 94.26 Ambient Temp.: 95.41 F
17:14:10.147 -> Average Cal sensor Temp.: 94.41 Ambient Temp.: 95.41 F
17:14:11.704 -> Average Cal sensor Temp.: 94.49 Ambient Temp.: 95.41 F
17:14:13.254 -> Average Cal sensor Temp.: 94.31 Ambient Temp.: 95.45 F
17:14:14.804 -> Average Cal sensor Temp.: 94.30 Ambient Temp.: 95.38 F
17:14:16.360 -> Average Cal sensor Temp.: 94.47 Ambient Temp.: 95.45 F
17:14:17.944 -> Average Cal sensor Temp.: 94.44 Ambient Temp.: 95.45 F
17:14:19.496 -> Average Cal sensor Temp.: 94.51 Ambient Temp.: 95.41 F
17:14:21.049 -> Average Cal sensor Temp.: 94.68 Ambient Temp.: 95.41 F
17:14:22.606 -> Average Cal sensor Temp.: 94.83 Ambient Temp.: 95.45 F
17:14:24.161 -> Average Cal sensor Temp.: 94.74 Ambient Temp.: 95.45 F
17:14:25.713 -> Average Cal sensor Temp.: 94.81 Ambient Temp.: 95.41 F
17:14:27.266 -> Average Cal sensor Temp.: 93.88 Ambient Temp.: 95.41 F
17:14:28.822 -> Average Cal sensor Temp.: 94.24 Ambient Temp.: 95.45 F
17:14:30.372 -> Average Cal sensor Temp.: 94.51 Ambient Temp.: 95.41 F
17:14:31.921 -> Average Cal sensor Temp.: 94.28 Ambient Temp.: 95.45 F
17:14:33.503 -> Average Cal sensor Temp.: 93.97 Ambient Temp.: 95.41 F
17:14:35.058 -> Average Cal sensor Temp.: 94.44 Ambient Temp.: 95.41 F
17:14:36.612 -> Average Cal sensor Temp.: 94.42 Ambient Temp.: 95.41 F
17:14:38.161 -> Average Cal sensor Temp.: 94.57 Ambient Temp.: 95.45 F
17:14:39.714 -> Average Cal sensor Temp.: 94.57 Ambient Temp.: 95.41 F
17:14:41.270 -> Average Cal sensor Temp.: 94.56 Ambient Temp.: 95.45 F
17:14:42.819 -> Average Cal sensor Temp.: 94.31 Ambient Temp.: 95.41 F
17:14:44.370 -> Average Cal sensor Temp.: 94.60 Ambient Temp.: 95.45 F
17:14:45.951 -> Average Cal sensor Temp.: 94.48 Ambient Temp.: 95.41 F
17:14:47.509 -> Average Cal sensor Temp.: 94.63 Ambient Temp.: 95.45 F
17:14:49.065 -> Average Cal sensor Temp.: 94.77 Ambient Temp.: 95.49 F
17:14:50.588 -> Average Cal sensor Temp.: 94.98 Ambient Temp.: 95.45 F
17:14:52.177 -> Average Cal sensor Temp.: 94.99 Ambient Temp.: 95.49 F
17:14:53.725 -> Average Cal sensor Temp.: 94.92 Ambient Temp.: 95.45 F
17:14:55.271 -> Average Cal sensor Temp.: 94.76 Ambient Temp.: 95.49 F
17:14:56.821 -> Average Cal sensor Temp.: 94.73 Ambient Temp.: 95.49 F
17:14:58.379 -> Average Cal sensor Temp.: 94.83 Ambient Temp.: 95.45 F
17:14:59.962 -> Average Cal sensor Temp.: 94.94 Ambient Temp.: 95.49 F
17:15:01.517 -> Average Cal sensor Temp.: 94.95 Ambient Temp.: 95.45 F
17:15:03.069 -> Average Cal sensor Temp.: 95.06 Ambient Temp.: 95.45 F
17:15:04.627 -> Average Cal sensor Temp.: 94.79 Ambient Temp.: 95.49 F
17:15:06.181 -> Average Cal sensor Temp.: 94.66 Ambient Temp.: 95.49 F
17:15:07.731 -> Average Cal sensor Temp.: 95.00 Ambient Temp.: 95.49 F
17:15:09.292 -> Average Cal sensor Temp.: 95.05 Ambient Temp.: 95.49 F
17:15:10.843 -> Average Cal sensor Temp.: 94.90 Ambient Temp.: 95.45 F
17:15:12.398 -> Average Cal sensor Temp.: 94.98 Ambient Temp.: 95.52 F
17:15:13.951 -> Average Cal sensor Temp.: 94.64 Ambient Temp.: 95.45 F
17:15:15.529 -> Average Cal sensor Temp.: 94.68 Ambient Temp.: 95.49 F
17:15:17.077 -> Average Cal sensor Temp.: 94.68 Ambient Temp.: 95.49 F
17:15:18.635 -> Average Cal sensor Temp.: 94.86 Ambient Temp.: 95.45 F
17:15:20.196 -> Average Cal sensor Temp.: 95.09 Ambient Temp.: 95.52 F
17:15:21.743 -> Average Cal sensor Temp.: 94.76 Ambient Temp.: 95.49 F
17:15:23.298 -> Average Cal sensor Temp.: 94.92 Ambient Temp.: 95.49 F
17:15:24.851 -> Average Cal sensor Temp.: 94.83 Ambient Temp.: 95.49 F
17:15:26.400 -> Average Cal sensor Temp.: 94.94 Ambient Temp.: 95.49 F
17:15:27.984 -> Average Cal sensor Temp.: 95.13 Ambient Temp.: 95.52 F
17:15:29.534 -> Average Cal sensor Temp.: 95.10 Ambient Temp.: 95.52 F
17:15:31.086 -> Average Cal sensor Temp.: 95.06 Ambient Temp.: 95.49 F
17:15:32.641 -> Average Cal sensor Temp.: 94.96 Ambient Temp.: 95.52 F
17:15:34.194 -> Average Cal sensor Temp.: 95.05 Ambient Temp.: 95.52 F
17:15:35.747 -> Average Cal sensor Temp.: 94.69 Ambient Temp.: 95.52 F
17:15:37.298 -> Average Cal sensor Temp.: 95.04 Ambient Temp.: 95.49 F
17:15:38.883 -> Average Cal sensor Temp.: 94.95 Ambient Temp.: 95.52 F
17:15:40.432 -> Average Cal sensor Temp.: 94.49 Ambient Temp.: 95.49 F
17:15:41.988 -> Average Cal sensor Temp.: 94.64 Ambient Temp.: 95.52 F
17:15:43.535 -> Average Cal sensor Temp.: 94.90 Ambient Temp.: 95.49 F
17:15:45.087 -> Average Cal sensor Temp.: 94.98 Ambient Temp.: 95.52 F
17:15:46.640 -> Average Cal sensor Temp.: 94.81 Ambient Temp.: 95.52 F
17:15:48.190 -> Average Cal sensor Temp.: 94.92 Ambient Temp.: 95.52 F
17:15:49.746 -> Average Cal sensor Temp.: 94.66 Ambient Temp.: 95.52 F
17:15:51.334 -> Average Cal sensor Temp.: 94.94 Ambient Temp.: 95.52 F
17:15:52.859 -> Average Cal sensor Temp.: 95.09 Ambient Temp.: 95.52 F
17:15:54.415 -> Average Cal sensor Temp.: 95.03 Ambient Temp.: 95.52 F
17:15:56.000 -> Average Cal sensor Temp.: 95.30 Ambient Temp.: 95.52 F
17:15:57.554 -> Average Cal sensor Temp.: 95.07 Ambient Temp.: 95.56 F
17:15:59.103 -> Average Cal sensor Temp.: 95.27 Ambient Temp.: 95.52 F
17:16:00.657 -> Average Cal sensor Temp.: 94.88 Ambient Temp.: 95.52 F
17:16:02.209 -> Average Cal sensor Temp.: 95.12 Ambient Temp.: 95.56 F
17:16:03.764 -> Average Cal sensor Temp.: 95.17 Ambient Temp.: 95.56 F
17:16:05.321 -> Average Cal sensor Temp.: 95.01 Ambient Temp.: 95.52 F
17:16:06.874 -> Average Cal sensor Temp.: 95.18 Ambient Temp.: 95.52 F
17:16:08.423 -> Average Cal sensor Temp.: 95.10 Ambient Temp.: 95.56 F
17:16:09.975 -> Average Cal sensor Temp.: 94.88 Ambient Temp.: 95.52 F
17:16:11.528 -> Average Cal sensor Temp.: 94.94 Ambient Temp.: 95.56 F
17:16:13.083 -> Average Cal sensor Temp.: 94.78 Ambient Temp.: 95.56 F
17:16:14.643 -> Average Cal sensor Temp.: 94.23 Ambient Temp.: 95.56 F
17:16:16.197 -> Average Cal sensor Temp.: 94.83 Ambient Temp.: 95.52 F
17:16:17.752 -> Average Cal sensor Temp.: 95.05 Ambient Temp.: 95.56 F
17:16:19.314 -> Average Cal sensor Temp.: 94.80 Ambient Temp.: 95.52 F
17:16:20.865 -> Average Cal sensor Temp.: 94.20 Ambient Temp.: 95.52 F
17:16:22.423 -> Average Cal sensor Temp.: 94.63 Ambient Temp.: 95.59 F
17:16:23.979 -> Average Cal sensor Temp.: 94.84 Ambient Temp.: 95.52 F
17:16:25.534 -> Average Cal sensor Temp.: 95.01 Ambient Temp.: 95.52 F
17:16:27.115 -> Average Cal sensor Temp.: 95.21 Ambient Temp.: 95.52 F
17:16:28.667 -> Average Cal sensor Temp.: 94.97 Ambient Temp.: 95.56 F
17:16:30.219 -> Average Cal sensor Temp.: 94.75 Ambient Temp.: 95.59 F
17:16:31.773 -> Average Cal sensor Temp.: 95.01 Ambient Temp.: 95.56 F
17:16:33.323 -> Average Cal sensor Temp.: 95.23 Ambient Temp.: 95.52 F
17:16:34.878 -> Average Cal sensor Temp.: 95.28 Ambient Temp.: 95.59 F
17:16:36.441 -> Average Cal sensor Temp.: 95.14 Ambient Temp.: 95.52 F
17:16:37.997 -> Average Cal sensor Temp.: 95.24 Ambient Temp.: 95.56 F
17:16:39.551 -> Average Cal sensor Temp.: 95.07 Ambient Temp.: 95.52 F
17:16:41.100 -> Average Cal sensor Temp.: 94.80 Ambient Temp.: 95.56 F
17:16:42.656 -> Average Cal sensor Temp.: 95.16 Ambient Temp.: 95.56 F
17:16:44.209 -> Average Cal sensor Temp.: 95.13 Ambient Temp.: 95.59 F
17:16:45.762 -> Average Cal sensor Temp.: 95.21 Ambient Temp.: 95.56 F
17:16:47.313 -> Average Cal sensor Temp.: 95.28 Ambient Temp.: 95.59 F
17:16:48.864 -> Average Cal sensor Temp.: 95.30 Ambient Temp.: 95.59 F
17:16:50.421 -> Average Cal sensor Temp.: 95.19 Ambient Temp.: 95.56 F
17:16:52.004 -> Average Cal sensor Temp.: 95.23 Ambient Temp.: 95.67 F
17:16:53.552 -> Average Cal sensor Temp.: 95.40 Ambient Temp.: 95.56 F
17:16:55.106 -> Average Cal sensor Temp.: 95.39 Ambient Temp.: 95.67 F
17:16:56.664 -> Average Cal sensor Temp.: 95.45 Ambient Temp.: 95.59 F
17:16:58.218 -> Average Cal sensor Temp.: 95.42 Ambient Temp.: 95.67 F
17:16:59.771 -> Average Cal sensor Temp.: 95.25 Ambient Temp.: 95.70 F
17:17:01.324 -> Average Cal sensor Temp.: 95.22 Ambient Temp.: 95.59 F
17:17:02.879 -> Average Cal sensor Temp.: 94.82 Ambient Temp.: 95.59 F
17:17:04.431 -> Average Cal sensor Temp.: 94.82 Ambient Temp.: 95.59 F
17:17:05.983 -> Average Cal sensor Temp.: 95.00 Ambient Temp.: 95.59 F
17:17:07.530 -> Average Cal sensor Temp.: 94.99 Ambient Temp.: 95.59 F
17:17:09.118 -> Average Cal sensor Temp.: 94.73 Ambient Temp.: 95.67 F
17:17:10.642 -> Average Cal sensor Temp.: 95.11 Ambient Temp.: 95.56 F
17:17:12.197 -> Average Cal sensor Temp.: 95.25 Ambient Temp.: 95.59 F
17:17:13.758 -> Average Cal sensor Temp.: 94.89 Ambient Temp.: 95.59 F
17:17:15.312 -> Average Cal sensor Temp.: 94.70 Ambient Temp.: 95.56 F
17:17:16.868 -> Average Cal sensor Temp.: 94.78 Ambient Temp.: 95.56 F
17:17:18.452 -> Average Cal sensor Temp.: 94.70 Ambient Temp.: 95.56 F
17:17:20.005 -> Average Cal sensor Temp.: 94.82 Ambient Temp.: 95.56 F
17:17:21.558 -> Average Cal sensor Temp.: 94.99 Ambient Temp.: 95.56 F
17:17:23.111 -> Average Cal sensor Temp.: 94.95 Ambient Temp.: 95.59 F
17:17:24.665 -> Average Cal sensor Temp.: 95.19 Ambient Temp.: 95.56 F
17:17:26.220 -> Average Cal sensor Temp.: 95.19 Ambient Temp.: 95.67 F
17:17:27.769 -> Average Cal sensor Temp.: 95.05 Ambient Temp.: 95.59 F
17:17:29.325 -> Average Cal sensor Temp.: 95.10 Ambient Temp.: 95.56 F
17:17:30.872 -> Average Cal sensor Temp.: 95.42 Ambient Temp.: 95.59 F
17:17:32.424 -> Average Cal sensor Temp.: 95.43 Ambient Temp.: 95.59 F
17:17:34.013 -> Average Cal sensor Temp.: 95.54 Ambient Temp.: 95.59 F
17:17:35.536 -> Average Cal sensor Temp.: 95.61 Ambient Temp.: 95.59 F
17:17:37.118 -> Average Cal sensor Temp.: 95.24 Ambient Temp.: 95.56 F
17:17:38.670 -> Average Cal sensor Temp.: 95.23 Ambient Temp.: 95.67 F
17:17:40.224 -> Average Cal sensor Temp.: 95.13 Ambient Temp.: 95.59 F
17:17:41.782 -> Average Cal sensor Temp.: 95.23 Ambient Temp.: 95.59 F
17:17:43.332 -> Average Cal sensor Temp.: 95.32 Ambient Temp.: 95.67 F
17:17:44.885 -> Average Cal sensor Temp.: 95.37 Ambient Temp.: 95.59 F
17:17:46.441 -> Average Cal sensor Temp.: 95.24 Ambient Temp.: 95.59 F
17:17:47.995 -> Average Cal sensor Temp.: 95.26 Ambient Temp.: 95.67 F
17:17:49.552 -> Average Cal sensor Temp.: 95.40 Ambient Temp.: 95.59 F
17:17:51.104 -> Average Cal sensor Temp.: 94.67 Ambient Temp.: 95.70 F
17:17:52.653 -> Average Cal sensor Temp.: 95.12 Ambient Temp.: 95.70 F
17:17:54.209 -> Average Cal sensor Temp.: 95.24 Ambient Temp.: 95.59 F
17:17:55.760 -> Average Cal sensor Temp.: 95.40 Ambient Temp.: 95.67 F
17:17:57.316 -> Average Cal sensor Temp.: 95.29 Ambient Temp.: 95.56 F
17:17:58.867 -> Average Cal sensor Temp.: 95.53 Ambient Temp.: 95.67 F
17:18:00.451 -> Average Cal sensor Temp.: 95.56 Ambient Temp.: 95.59 F
17:18:01.999 -> Average Cal sensor Temp.: 95.46 Ambient Temp.: 95.67 F
17:18:03.554 -> Average Cal sensor Temp.: 95.07 Ambient Temp.: 95.67 F
17:18:05.106 -> Average Cal sensor Temp.: 95.26 Ambient Temp.: 95.70 F
17:18:06.660 -> Average Cal sensor Temp.: 95.41 Ambient Temp.: 95.67 F
17:18:08.217 -> Average Cal sensor Temp.: 95.24 Ambient Temp.: 95.67 F
17:18:09.778 -> Average Cal sensor Temp.: 95.17 Ambient Temp.: 95.70 F
17:18:11.332 -> Average Cal sensor Temp.: 95.13 Ambient Temp.: 95.70 F
17:18:12.884 -> Average Cal sensor Temp.: 95.40 Ambient Temp.: 95.59 F
17:18:14.430 -> Average Cal sensor Temp.: 95.48 Ambient Temp.: 95.67 F
17:18:15.984 -> Average Cal sensor Temp.: 95.28 Ambient Temp.: 95.70 F
17:18:17.539 -> Average Cal sensor Temp.: 95.53 Ambient Temp.: 95.70 F
17:18:19.094 -> Average Cal sensor Temp.: 95.26 Ambient Temp.: 95.59 F
17:18:20.652 -> Average Cal sensor Temp.: 95.12 Ambient Temp.: 95.67 F
17:18:22.209 -> Average Cal sensor Temp.: 95.21 Ambient Temp.: 95.59 F
17:18:23.767 -> Average Cal sensor Temp.: 95.43 Ambient Temp.: 95.59 F
17:18:25.323 -> Average Cal sensor Temp.: 95.51 Ambient Temp.: 95.67 F
17:18:26.874 -> Average Cal sensor Temp.: 95.43 Ambient Temp.: 95.56 F
17:18:28.425 -> Average Cal sensor Temp.: 95.38 Ambient Temp.: 95.70 F
17:18:29.977 -> Average Cal sensor Temp.: 95.54 Ambient Temp.: 95.67 F
17:18:31.560 -> Average Cal sensor Temp.: 94.90 Ambient Temp.: 95.67 F
17:18:33.107 -> Average Cal sensor Temp.: 95.12 Ambient Temp.: 95.67 F
17:18:34.656 -> Average Cal sensor Temp.: 95.16 Ambient Temp.: 95.59 F
17:18:36.215 -> Average Cal sensor Temp.: 95.18 Ambient Temp.: 95.67 F
17:18:37.766 -> Average Cal sensor Temp.: 95.22 Ambient Temp.: 95.67 F
17:18:39.322 -> Average Cal sensor Temp.: 95.14 Ambient Temp.: 95.59 F
17:18:40.880 -> Average Cal sensor Temp.: 95.24 Ambient Temp.: 95.59 F
17:18:42.434 -> Average Cal sensor Temp.: 95.16 Ambient Temp.: 95.67 F
17:18:43.986 -> Average Cal sensor Temp.: 94.77 Ambient Temp.: 95.70 F
17:18:45.542 -> Average Cal sensor Temp.: 95.26 Ambient Temp.: 95.59 F
17:18:47.092 -> Average Cal sensor Temp.: 95.18 Ambient Temp.: 95.56 F
17:18:48.644 -> Average Cal sensor Temp.: 95.07 Ambient Temp.: 95.59 F
17:18:50.227 -> Average Cal sensor Temp.: 95.44 Ambient Temp.: 95.59 F
17:18:51.776 -> Average Cal sensor Temp.: 95.32 Ambient Temp.: 95.59 F
17:18:53.328 -> Average Cal sensor Temp.: 95.07 Ambient Temp.: 95.56 F
17:18:54.881 -> Average Cal sensor Temp.: 94.90 Ambient Temp.: 95.59 F
17:18:56.439 -> Average Cal sensor Temp.: 95.50 Ambient Temp.: 95.56 F
17:18:57.990 -> Average Cal sensor Temp.: 95.50 Ambient Temp.: 95.67 F
17:18:59.541 -> Average Cal sensor Temp.: 95.47 Ambient Temp.: 95.59 F
17:19:01.095 -> Average Cal sensor Temp.: 95.34 Ambient Temp.: 95.59 F
17:19:02.643 -> Average Cal sensor Temp.: 95.11 Ambient Temp.: 95.59 F
17:19:04.228 -> Average Cal sensor Temp.: 95.32 Ambient Temp.: 95.59 F
17:19:05.781 -> Average Cal sensor Temp.: 95.27 Ambient Temp.: 95.56 F
17:19:07.333 -> Average Cal sensor Temp.: 95.39 Ambient Temp.: 95.59 F
17:19:08.883 -> Average Cal sensor Temp.: 95.28 Ambient Temp.: 95.59 F
17:19:10.437 -> Average Cal sensor Temp.: 95.59 Ambient Temp.: 95.59 F
17:19:11.990 -> Average Cal sensor Temp.: 94.72 Ambient Temp.: 95.59 F
17:19:13.551 -> Average Cal sensor Temp.: 95.02 Ambient Temp.: 95.56 F
17:19:15.100 -> Average Cal sensor Temp.: 95.16 Ambient Temp.: 95.67 F
17:19:16.656 -> Average Cal sensor Temp.: 94.89 Ambient Temp.: 95.59 F
17:19:18.211 -> Average Cal sensor Temp.: 95.33 Ambient Temp.: 95.59 F
17:19:19.770 -> Average Cal sensor Temp.: 95.54 Ambient Temp.: 95.59 F
17:19:21.327 -> Average Cal sensor Temp.: 95.50 Ambient Temp.: 95.56 F
17:19:22.882 -> Average Cal sensor Temp.: 95.53 Ambient Temp.: 95.59 F
17:19:24.430 -> Average Cal sensor Temp.: 95.57 Ambient Temp.: 95.67 F
17:19:25.983 -> Average Cal sensor Temp.: 95.45 Ambient Temp.: 95.59 F
17:19:27.533 -> Average Cal sensor Temp.: 95.64 Ambient Temp.: 95.70 F
17:19:29.087 -> Average Cal sensor Temp.: 95.72 Ambient Temp.: 95.59 F
17:19:30.673 -> Average Cal sensor Temp.: 95.75 Ambient Temp.: 95.59 F
17:19:32.195 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 95.59 F
17:19:33.782 -> Average Cal sensor Temp.: 95.76 Ambient Temp.: 95.67 F
17:19:35.328 -> Average Cal sensor Temp.: 95.66 Ambient Temp.: 95.70 F
17:19:36.880 -> Average Cal sensor Temp.: 95.71 Ambient Temp.: 95.59 F
17:19:38.427 -> Average Cal sensor Temp.: 95.64 Ambient Temp.: 95.67 F
17:19:39.985 -> Average Cal sensor Temp.: 95.62 Ambient Temp.: 95.70 F
17:19:41.531 -> Average Cal sensor Temp.: 95.72 Ambient Temp.: 95.67 F
17:19:43.109 -> Average Cal sensor Temp.: 95.89 Ambient Temp.: 95.67 F
17:19:44.667 -> Average Cal sensor Temp.: 95.64 Ambient Temp.: 95.67 F
17:19:46.223 -> Average Cal sensor Temp.: 95.33 Ambient Temp.: 95.70 F
17:19:47.780 -> Average Cal sensor Temp.: 95.53 Ambient Temp.: 95.67 F
17:19:49.332 -> Average Cal sensor Temp.: 95.52 Ambient Temp.: 95.67 F
17:19:50.887 -> Average Cal sensor Temp.: 95.68 Ambient Temp.: 95.70 F
17:19:52.445 -> Average Cal sensor Temp.: 95.56 Ambient Temp.: 95.70 F
17:19:54.004 -> Average Cal sensor Temp.: 95.15 Ambient Temp.: 95.70 F
17:19:55.558 -> Average Cal sensor Temp.: 95.49 Ambient Temp.: 95.67 F
17:19:57.117 -> Average Cal sensor Temp.: 95.65 Ambient Temp.: 95.67 F
17:19:58.664 -> Average Cal sensor Temp.: 95.53 Ambient Temp.: 95.67 F
17:20:00.211 -> Average Cal sensor Temp.: 95.67 Ambient Temp.: 95.70 F
17:20:01.770 -> Average Cal sensor Temp.: 95.12 Ambient Temp.: 95.59 F
17:20:03.324 -> Average Cal sensor Temp.: 95.58 Ambient Temp.: 95.67 F
17:20:04.872 -> Average Cal sensor Temp.: 95.35 Ambient Temp.: 95.59 F
17:20:06.425 -> Average Cal sensor Temp.: 95.51 Ambient Temp.: 95.70 F
17:20:07.973 -> Average Cal sensor Temp.: 95.50 Ambient Temp.: 95.59 F
17:20:09.533 -> Average Cal sensor Temp.: 95.48 Ambient Temp.: 95.67 F
17:20:11.083 -> Average Cal sensor Temp.: 95.61 Ambient Temp.: 95.70 F
17:20:12.641 -> Average Cal sensor Temp.: 95.69 Ambient Temp.: 95.70 F
17:20:14.194 -> Average Cal sensor Temp.: 95.62 Ambient Temp.: 95.70 F
17:20:15.776 -> Average Cal sensor Temp.: 95.78 Ambient Temp.: 95.67 F
17:20:17.324 -> Average Cal sensor Temp.: 95.77 Ambient Temp.: 95.74 F
17:20:18.874 -> Average Cal sensor Temp.: 95.84 Ambient Temp.: 95.67 F
17:20:20.425 -> Average Cal sensor Temp.: 95.94 Ambient Temp.: 95.67 F
17:20:21.977 -> Average Cal sensor Temp.: 95.91 Ambient Temp.: 95.74 F
17:20:23.533 -> Average Cal sensor Temp.: 95.82 Ambient Temp.: 95.70 F
17:20:25.088 -> Average Cal sensor Temp.: 95.54 Ambient Temp.: 95.74 F
17:20:26.643 -> Average Cal sensor Temp.: 95.62 Ambient Temp.: 95.67 F
17:20:28.201 -> Average Cal sensor Temp.: 95.65 Ambient Temp.: 95.67 F
17:20:29.756 -> Average Cal sensor Temp.: 95.80 Ambient Temp.: 95.70 F
17:20:31.336 -> Average Cal sensor Temp.: 95.89 Ambient Temp.: 95.67 F
17:20:32.887 -> Average Cal sensor Temp.: 95.72 Ambient Temp.: 95.70 F
17:20:34.442 -> Average Cal sensor Temp.: 95.98 Ambient Temp.: 95.70 F
17:20:35.998 -> Average Cal sensor Temp.: 95.86 Ambient Temp.: 95.74 F
17:20:37.547 -> Average Cal sensor Temp.: 95.49 Ambient Temp.: 95.67 F
17:20:39.097 -> Average Cal sensor Temp.: 95.85 Ambient Temp.: 95.74 F
17:20:40.654 -> Average Cal sensor Temp.: 95.76 Ambient Temp.: 95.70 F
17:20:42.205 -> Average Cal sensor Temp.: 95.72 Ambient Temp.: 95.74 F
17:20:43.761 -> Average Cal sensor Temp.: 95.94 Ambient Temp.: 95.74 F
17:20:45.318 -> Average Cal sensor Temp.: 95.82 Ambient Temp.: 95.74 F
17:20:46.872 -> Average Cal sensor Temp.: 95.72 Ambient Temp.: 95.70 F
17:20:48.422 -> Average Cal sensor Temp.: 95.97 Ambient Temp.: 95.74 F
17:20:50.003 -> Average Cal sensor Temp.: 95.70 Ambient Temp.: 95.74 F
17:20:51.552 -> Average Cal sensor Temp.: 95.39 Ambient Temp.: 95.77 F
17:20:53.105 -> Average Cal sensor Temp.: 95.73 Ambient Temp.: 95.74 F
17:20:54.652 -> Average Cal sensor Temp.: 95.76 Ambient Temp.: 95.70 F
17:20:56.207 -> Average Cal sensor Temp.: 95.73 Ambient Temp.: 95.74 F
17:20:57.765 -> Average Cal sensor Temp.: 95.74 Ambient Temp.: 95.70 F
17:20:59.316 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 95.77 F
17:21:00.870 -> Average Cal sensor Temp.: 95.72 Ambient Temp.: 95.74 F
17:21:02.423 -> Average Cal sensor Temp.: 95.75 Ambient Temp.: 95.74 F
17:21:03.972 -> Average Cal sensor Temp.: 95.65 Ambient Temp.: 95.74 F
17:21:05.553 -> Average Cal sensor Temp.: 95.82 Ambient Temp.: 95.74 F
17:21:07.104 -> Average Cal sensor Temp.: 95.73 Ambient Temp.: 95.70 F
17:21:08.649 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 95.77 F
17:21:10.203 -> Average Cal sensor Temp.: 95.55 Ambient Temp.: 95.74 F
17:21:11.754 -> Average Cal sensor Temp.: 95.44 Ambient Temp.: 95.74 F
17:21:13.329 -> Average Cal sensor Temp.: 95.50 Ambient Temp.: 95.74 F
17:21:14.884 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 95.70 F
17:21:16.436 -> Average Cal sensor Temp.: 95.75 Ambient Temp.: 95.77 F
17:21:17.987 -> Average Cal sensor Temp.: 95.74 Ambient Temp.: 95.70 F
17:21:19.541 -> Average Cal sensor Temp.: 95.85 Ambient Temp.: 95.74 F
17:21:21.099 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 95.74 F
17:21:22.657 -> Average Cal sensor Temp.: 95.83 Ambient Temp.: 95.74 F
17:21:24.208 -> Average Cal sensor Temp.: 95.67 Ambient Temp.: 95.77 F
17:21:25.766 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 95.77 F
17:21:27.318 -> Average Cal sensor Temp.: 95.85 Ambient Temp.: 95.77 F
17:21:28.869 -> Average Cal sensor Temp.: 95.59 Ambient Temp.: 95.74 F
17:21:30.423 -> Average Cal sensor Temp.: 95.73 Ambient Temp.: 95.74 F
17:21:31.974 -> Average Cal sensor Temp.: 95.29 Ambient Temp.: 95.77 F
17:21:33.522 -> Average Cal sensor Temp.: 95.72 Ambient Temp.: 95.77 F
17:21:35.109 -> Average Cal sensor Temp.: 95.80 Ambient Temp.: 95.77 F
17:21:36.662 -> Average Cal sensor Temp.: 95.80 Ambient Temp.: 95.77 F
17:21:38.214 -> Average Cal sensor Temp.: 95.53 Ambient Temp.: 95.77 F
17:21:39.769 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 95.77 F
17:21:41.322 -> Average Cal sensor Temp.: 95.92 Ambient Temp.: 95.81 F
17:21:42.877 -> Average Cal sensor Temp.: 95.55 Ambient Temp.: 95.77 F
17:21:44.429 -> Average Cal sensor Temp.: 95.72 Ambient Temp.: 95.77 F
17:21:45.992 -> Average Cal sensor Temp.: 95.05 Ambient Temp.: 95.74 F
17:21:47.517 -> Average Cal sensor Temp.: 95.47 Ambient Temp.: 95.74 F
17:21:49.076 -> Average Cal sensor Temp.: 95.67 Ambient Temp.: 95.74 F
17:21:50.632 -> Average Cal sensor Temp.: 95.69 Ambient Temp.: 95.74 F
17:21:52.188 -> Average Cal sensor Temp.: 95.11 Ambient Temp.: 95.77 F
17:21:53.747 -> Average Cal sensor Temp.: 94.93 Ambient Temp.: 95.70 F
17:21:55.305 -> Average Cal sensor Temp.: 95.35 Ambient Temp.: 95.70 F
17:21:56.860 -> Average Cal sensor Temp.: 95.66 Ambient Temp.: 95.74 F
17:21:58.419 -> Average Cal sensor Temp.: 95.56 Ambient Temp.: 95.67 F
17:21:59.979 -> Average Cal sensor Temp.: 95.40 Ambient Temp.: 95.74 F
17:22:01.535 -> Average Cal sensor Temp.: 95.49 Ambient Temp.: 95.70 F
17:22:03.090 -> Average Cal sensor Temp.: 95.75 Ambient Temp.: 95.74 F
17:22:04.645 -> Average Cal sensor Temp.: 95.77 Ambient Temp.: 95.74 F
17:22:06.207 -> Average Cal sensor Temp.: 95.86 Ambient Temp.: 95.74 F
17:22:07.761 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 95.74 F
17:22:09.320 -> Average Cal sensor Temp.: 95.69 Ambient Temp.: 95.74 F
17:22:10.873 -> Average Cal sensor Temp.: 95.83 Ambient Temp.: 95.74 F
17:22:12.427 -> Average Cal sensor Temp.: 95.96 Ambient Temp.: 95.70 F
17:22:13.983 -> Average Cal sensor Temp.: 95.90 Ambient Temp.: 95.77 F
17:22:15.543 -> Average Cal sensor Temp.: 95.88 Ambient Temp.: 95.74 F
17:22:17.099 -> Average Cal sensor Temp.: 95.89 Ambient Temp.: 95.74 F
17:22:18.659 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 95.74 F
17:22:20.218 -> Average Cal sensor Temp.: 96.11 Ambient Temp.: 95.74 F
17:22:21.743 -> Average Cal sensor Temp.: 96.03 Ambient Temp.: 95.74 F
17:22:23.301 -> Average Cal sensor Temp.: 95.82 Ambient Temp.: 95.74 F
17:22:24.858 -> Average Cal sensor Temp.: 96.09 Ambient Temp.: 95.74 F
17:22:26.414 -> Average Cal sensor Temp.: 95.99 Ambient Temp.: 95.70 F
17:22:28.000 -> Average Cal sensor Temp.: 95.69 Ambient Temp.: 95.74 F
17:22:29.555 -> Average Cal sensor Temp.: 95.76 Ambient Temp.: 95.70 F
17:22:31.080 -> Average Cal sensor Temp.: 95.72 Ambient Temp.: 95.74 F
17:22:32.667 -> Average Cal sensor Temp.: 95.87 Ambient Temp.: 95.77 F
17:22:34.222 -> Average Cal sensor Temp.: 96.11 Ambient Temp.: 95.77 F
17:22:35.774 -> Average Cal sensor Temp.: 96.22 Ambient Temp.: 95.74 F
17:22:37.330 -> Average Cal sensor Temp.: 96.25 Ambient Temp.: 95.77 F
17:22:38.854 -> Average Cal sensor Temp.: 96.32 Ambient Temp.: 95.77 F
17:22:40.409 -> Average Cal sensor Temp.: 95.96 Ambient Temp.: 95.74 F
17:22:41.997 -> Average Cal sensor Temp.: 96.03 Ambient Temp.: 95.77 F
17:22:43.549 -> Average Cal sensor Temp.: 96.06 Ambient Temp.: 95.81 F
17:22:45.109 -> Average Cal sensor Temp.: 96.05 Ambient Temp.: 95.77 F
17:22:46.664 -> Average Cal sensor Temp.: 95.82 Ambient Temp.: 95.77 F
17:22:48.217 -> Average Cal sensor Temp.: 95.93 Ambient Temp.: 95.77 F
17:22:49.774 -> Average Cal sensor Temp.: 95.77 Ambient Temp.: 95.77 F
17:22:51.329 -> Average Cal sensor Temp.: 95.85 Ambient Temp.: 95.77 F
17:22:52.886 -> Average Cal sensor Temp.: 95.72 Ambient Temp.: 95.74 F
17:22:54.437 -> Average Cal sensor Temp.: 95.99 Ambient Temp.: 95.77 F
17:22:55.994 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 95.77 F
17:22:57.551 -> Average Cal sensor Temp.: 95.99 Ambient Temp.: 95.77 F
17:22:59.108 -> Average Cal sensor Temp.: 95.86 Ambient Temp.: 95.81 F
17:23:00.638 -> Average Cal sensor Temp.: 95.87 Ambient Temp.: 95.77 F
17:23:02.198 -> Average Cal sensor Temp.: 95.83 Ambient Temp.: 95.77 F
17:23:03.751 -> Average Cal sensor Temp.: 95.93 Ambient Temp.: 95.77 F
17:23:05.308 -> Average Cal sensor Temp.: 95.86 Ambient Temp.: 95.81 F
17:23:06.868 -> Average Cal sensor Temp.: 95.93 Ambient Temp.: 95.77 F
17:23:08.419 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 95.77 F
17:23:09.975 -> Average Cal sensor Temp.: 95.79 Ambient Temp.: 95.77 F
17:23:11.533 -> Average Cal sensor Temp.: 96.18 Ambient Temp.: 95.74 F
17:23:13.089 -> Average Cal sensor Temp.: 95.74 Ambient Temp.: 95.74 F
17:23:14.641 -> Average Cal sensor Temp.: 95.76 Ambient Temp.: 95.77 F
17:23:16.196 -> Average Cal sensor Temp.: 95.84 Ambient Temp.: 95.77 F
17:23:17.755 -> Average Cal sensor Temp.: 96.20 Ambient Temp.: 95.77 F
17:23:19.314 -> Average Cal sensor Temp.: 96.13 Ambient Temp.: 95.77 F
17:23:20.871 -> Average Cal sensor Temp.: 96.06 Ambient Temp.: 95.81 F
17:23:22.428 -> Average Cal sensor Temp.: 96.03 Ambient Temp.: 95.74 F
17:23:23.986 -> Average Cal sensor Temp.: 96.00 Ambient Temp.: 95.81 F
17:23:25.548 -> Average Cal sensor Temp.: 95.94 Ambient Temp.: 95.74 F
17:23:27.102 -> Average Cal sensor Temp.: 96.01 Ambient Temp.: 95.81 F
17:23:28.657 -> Average Cal sensor Temp.: 95.83 Ambient Temp.: 95.81 F
17:23:30.209 -> Average Cal sensor Temp.: 96.07 Ambient Temp.: 95.81 F
17:23:31.767 -> Average Cal sensor Temp.: 95.98 Ambient Temp.: 95.74 F
17:23:33.320 -> Average Cal sensor Temp.: 95.97 Ambient Temp.: 95.81 F
17:23:34.877 -> Average Cal sensor Temp.: 96.24 Ambient Temp.: 95.81 F
17:23:36.432 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 95.77 F
17:23:37.990 -> Average Cal sensor Temp.: 96.16 Ambient Temp.: 95.81 F
17:23:39.549 -> Average Cal sensor Temp.: 96.23 Ambient Temp.: 95.81 F
17:23:41.101 -> Average Cal sensor Temp.: 96.29 Ambient Temp.: 95.81 F
17:23:42.659 -> Average Cal sensor Temp.: 96.22 Ambient Temp.: 95.85 F
17:23:44.216 -> Average Cal sensor Temp.: 95.99 Ambient Temp.: 95.77 F
17:23:45.775 -> Average Cal sensor Temp.: 96.23 Ambient Temp.: 95.85 F
17:23:47.333 -> Average Cal sensor Temp.: 96.22 Ambient Temp.: 95.85 F
17:23:48.892 -> Average Cal sensor Temp.: 96.01 Ambient Temp.: 95.81 F
17:23:50.451 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 95.85 F
17:23:52.005 -> Average Cal sensor Temp.: 96.23 Ambient Temp.: 95.77 F
17:23:53.563 -> Average Cal sensor Temp.: 95.99 Ambient Temp.: 95.77 F
17:23:55.090 -> Average Cal sensor Temp.: 96.13 Ambient Temp.: 95.81 F
17:23:56.646 -> Average Cal sensor Temp.: 96.12 Ambient Temp.: 95.77 F
17:23:58.200 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 95.77 F
17:23:59.754 -> Average Cal sensor Temp.: 96.07 Ambient Temp.: 95.81 F
17:24:01.340 -> Average Cal sensor Temp.: 96.05 Ambient Temp.: 95.77 F
17:24:02.890 -> Average Cal sensor Temp.: 95.61 Ambient Temp.: 95.81 F
17:24:04.441 -> Average Cal sensor Temp.: 95.91 Ambient Temp.: 95.77 F
17:24:05.995 -> Average Cal sensor Temp.: 95.88 Ambient Temp.: 95.81 F
17:24:07.553 -> Average Cal sensor Temp.: 96.22 Ambient Temp.: 95.77 F
17:24:09.108 -> Average Cal sensor Temp.: 95.88 Ambient Temp.: 95.77 F
17:24:10.660 -> Average Cal sensor Temp.: 96.07 Ambient Temp.: 95.81 F
17:24:12.211 -> Average Cal sensor Temp.: 95.77 Ambient Temp.: 95.81 F
17:24:13.765 -> Average Cal sensor Temp.: 95.82 Ambient Temp.: 95.81 F
17:24:15.319 -> Average Cal sensor Temp.: 95.71 Ambient Temp.: 95.81 F
17:24:16.872 -> Average Cal sensor Temp.: 95.92 Ambient Temp.: 95.77 F
17:24:18.425 -> Average Cal sensor Temp.: 96.22 Ambient Temp.: 95.81 F
17:24:19.985 -> Average Cal sensor Temp.: 95.82 Ambient Temp.: 95.81 F
17:24:21.539 -> Average Cal sensor Temp.: 96.19 Ambient Temp.: 95.77 F
17:24:23.091 -> Average Cal sensor Temp.: 95.78 Ambient Temp.: 95.77 F
17:24:24.649 -> Average Cal sensor Temp.: 96.29 Ambient Temp.: 95.77 F
17:24:26.197 -> Average Cal sensor Temp.: 96.24 Ambient Temp.: 95.77 F
17:24:27.783 -> Average Cal sensor Temp.: 96.25 Ambient Temp.: 95.77 F
17:24:29.336 -> Average Cal sensor Temp.: 96.24 Ambient Temp.: 95.85 F
17:24:30.895 -> Average Cal sensor Temp.: 96.24 Ambient Temp.: 95.81 F
17:24:32.444 -> Average Cal sensor Temp.: 95.90 Ambient Temp.: 95.77 F
17:24:33.995 -> Average Cal sensor Temp.: 96.05 Ambient Temp.: 95.81 F
17:24:35.551 -> Average Cal sensor Temp.: 96.10 Ambient Temp.: 95.77 F
17:24:37.101 -> Average Cal sensor Temp.: 96.21 Ambient Temp.: 95.81 F
17:24:38.655 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 95.77 F
17:24:40.207 -> Average Cal sensor Temp.: 95.98 Ambient Temp.: 95.81 F
17:24:41.755 -> Average Cal sensor Temp.: 95.82 Ambient Temp.: 95.81 F
17:24:43.342 -> Average Cal sensor Temp.: 95.84 Ambient Temp.: 95.81 F
17:24:44.901 -> Average Cal sensor Temp.: 95.97 Ambient Temp.: 95.81 F
17:24:46.447 -> Average Cal sensor Temp.: 96.05 Ambient Temp.: 95.81 F
17:24:47.997 -> Average Cal sensor Temp.: 96.25 Ambient Temp.: 95.81 F
17:24:49.550 -> Average Cal sensor Temp.: 96.32 Ambient Temp.: 95.77 F
17:24:51.098 -> Average Cal sensor Temp.: 96.18 Ambient Temp.: 95.85 F
17:24:52.653 -> Average Cal sensor Temp.: 96.23 Ambient Temp.: 95.77 F
17:24:54.205 -> Average Cal sensor Temp.: 96.38 Ambient Temp.: 95.81 F
17:24:55.765 -> Average Cal sensor Temp.: 96.32 Ambient Temp.: 95.85 F
17:24:57.346 -> Average Cal sensor Temp.: 96.34 Ambient Temp.: 95.81 F
17:24:58.899 -> Average Cal sensor Temp.: 96.43 Ambient Temp.: 95.85 F
17:25:00.455 -> Average Cal sensor Temp.: 96.03 Ambient Temp.: 95.81 F
17:25:02.011 -> Average Cal sensor Temp.: 96.21 Ambient Temp.: 95.88 F
17:25:03.568 -> Average Cal sensor Temp.: 96.13 Ambient Temp.: 95.81 F
17:25:05.120 -> Average Cal sensor Temp.: 96.28 Ambient Temp.: 95.81 F
17:25:06.674 -> Average Cal sensor Temp.: 96.42 Ambient Temp.: 95.81 F
17:25:08.229 -> Average Cal sensor Temp.: 96.34 Ambient Temp.: 95.81 F
17:25:09.779 -> Average Cal sensor Temp.: 96.26 Ambient Temp.: 95.85 F
17:25:11.330 -> Average Cal sensor Temp.: 95.99 Ambient Temp.: 95.77 F
17:25:12.888 -> Average Cal sensor Temp.: 96.16 Ambient Temp.: 95.81 F
17:25:14.446 -> Average Cal sensor Temp.: 96.18 Ambient Temp.: 95.77 F
17:25:15.996 -> Average Cal sensor Temp.: 96.14 Ambient Temp.: 95.81 F
17:25:17.552 -> Average Cal sensor Temp.: 95.88 Ambient Temp.: 95.85 F
17:25:19.105 -> Average Cal sensor Temp.: 95.75 Ambient Temp.: 95.81 F
17:25:20.663 -> Average Cal sensor Temp.: 96.12 Ambient Temp.: 95.85 F
17:25:22.215 -> Average Cal sensor Temp.: 95.79 Ambient Temp.: 95.77 F
17:25:23.766 -> Average Cal sensor Temp.: 95.83 Ambient Temp.: 95.81 F
17:25:25.355 -> Average Cal sensor Temp.: 95.73 Ambient Temp.: 95.81 F
17:25:26.879 -> Average Cal sensor Temp.: 95.91 Ambient Temp.: 95.81 F
17:25:28.434 -> Average Cal sensor Temp.: 95.85 Ambient Temp.: 95.85 F
17:25:29.991 -> Average Cal sensor Temp.: 96.21 Ambient Temp.: 95.81 F
17:25:31.542 -> Average Cal sensor Temp.: 96.04 Ambient Temp.: 95.81 F
17:25:33.132 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 95.81 F
17:25:34.682 -> Average Cal sensor Temp.: 96.25 Ambient Temp.: 95.81 F
17:25:36.242 -> Average Cal sensor Temp.: 96.23 Ambient Temp.: 95.81 F
17:25:37.796 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 95.77 F
17:25:39.352 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 95.85 F
17:25:40.906 -> Average Cal sensor Temp.: 96.24 Ambient Temp.: 95.85 F
17:25:42.464 -> Average Cal sensor Temp.: 95.96 Ambient Temp.: 95.77 F
17:25:44.018 -> Average Cal sensor Temp.: 96.29 Ambient Temp.: 95.81 F
17:25:45.571 -> Average Cal sensor Temp.: 96.29 Ambient Temp.: 95.77 F
17:25:47.122 -> Average Cal sensor Temp.: 95.45 Ambient Temp.: 95.77 F
17:25:48.675 -> Average Cal sensor Temp.: 95.86 Ambient Temp.: 95.81 F
17:25:50.227 -> Average Cal sensor Temp.: 95.99 Ambient Temp.: 95.81 F
17:25:51.780 -> Average Cal sensor Temp.: 95.77 Ambient Temp.: 95.81 F
17:25:53.336 -> Average Cal sensor Temp.: 96.04 Ambient Temp.: 95.81 F
17:25:54.888 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 95.77 F
17:25:56.452 -> Average Cal sensor Temp.: 96.31 Ambient Temp.: 95.85 F
17:25:58.005 -> Average Cal sensor Temp.: 95.60 Ambient Temp.: 95.77 F
17:25:59.558 -> Average Cal sensor Temp.: 95.50 Ambient Temp.: 95.81 F
17:26:01.108 -> Average Cal sensor Temp.: 95.72 Ambient Temp.: 95.77 F
17:26:02.665 -> Average Cal sensor Temp.: 95.59 Ambient Temp.: 95.77 F
17:26:04.219 -> Average Cal sensor Temp.: 95.40 Ambient Temp.: 95.81 F
17:26:05.775 -> Average Cal sensor Temp.: 95.83 Ambient Temp.: 95.81 F
17:26:07.329 -> Average Cal sensor Temp.: 95.76 Ambient Temp.: 95.81 F
17:26:08.888 -> Average Cal sensor Temp.: 95.75 Ambient Temp.: 95.74 F
17:26:10.436 -> Average Cal sensor Temp.: 95.98 Ambient Temp.: 95.77 F
17:26:12.024 -> Average Cal sensor Temp.: 96.24 Ambient Temp.: 95.81 F
17:26:13.575 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 95.77 F
17:26:15.125 -> Average Cal sensor Temp.: 96.19 Ambient Temp.: 95.81 F
17:26:16.679 -> Average Cal sensor Temp.: 96.09 Ambient Temp.: 95.81 F
17:26:18.235 -> Average Cal sensor Temp.: 96.22 Ambient Temp.: 95.85 F
17:26:19.789 -> Average Cal sensor Temp.: 96.26 Ambient Temp.: 95.77 F
17:26:21.344 -> Average Cal sensor Temp.: 96.11 Ambient Temp.: 95.74 F
17:26:22.901 -> Average Cal sensor Temp.: 96.26 Ambient Temp.: 95.77 F
17:26:24.454 -> Average Cal sensor Temp.: 95.78 Ambient Temp.: 95.77 F
17:26:26.010 -> Average Cal sensor Temp.: 95.87 Ambient Temp.: 95.81 F
17:26:27.565 -> Average Cal sensor Temp.: 95.91 Ambient Temp.: 95.81 F
17:26:29.113 -> Average Cal sensor Temp.: 96.19 Ambient Temp.: 95.77 F
17:26:30.666 -> Average Cal sensor Temp.: 96.06 Ambient Temp.: 95.74 F
17:26:32.223 -> Average Cal sensor Temp.: 95.97 Ambient Temp.: 95.81 F
17:26:33.779 -> Average Cal sensor Temp.: 95.91 Ambient Temp.: 95.77 F
17:26:35.328 -> Average Cal sensor Temp.: 96.09 Ambient Temp.: 95.77 F
17:26:36.912 -> Average Cal sensor Temp.: 96.12 Ambient Temp.: 95.77 F
17:26:38.461 -> Average Cal sensor Temp.: 96.09 Ambient Temp.: 95.77 F
17:26:40.015 -> Average Cal sensor Temp.: 96.19 Ambient Temp.: 95.77 F
17:26:41.566 -> Average Cal sensor Temp.: 95.95 Ambient Temp.: 95.81 F
17:26:43.114 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 95.81 F
17:26:44.671 -> Average Cal sensor Temp.: 96.33 Ambient Temp.: 95.81 F
17:26:46.221 -> Average Cal sensor Temp.: 96.35 Ambient Temp.: 95.85 F
17:26:47.780 -> Average Cal sensor Temp.: 96.52 Ambient Temp.: 95.77 F
17:26:49.333 -> Average Cal sensor Temp.: 96.40 Ambient Temp.: 95.85 F
17:26:50.890 -> Average Cal sensor Temp.: 96.47 Ambient Temp.: 95.81 F
17:26:52.474 -> Average Cal sensor Temp.: 96.40 Ambient Temp.: 95.81 F
17:26:54.029 -> Average Cal sensor Temp.: 96.31 Ambient Temp.: 95.77 F
17:26:55.583 -> Average Cal sensor Temp.: 96.33 Ambient Temp.: 95.81 F
17:26:57.135 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 95.85 F
17:26:58.690 -> Average Cal sensor Temp.: 96.20 Ambient Temp.: 95.77 F
17:27:00.239 -> Average Cal sensor Temp.: 96.28 Ambient Temp.: 95.81 F
17:27:01.800 -> Average Cal sensor Temp.: 95.87 Ambient Temp.: 95.81 F
17:27:03.355 -> Average Cal sensor Temp.: 96.32 Ambient Temp.: 95.85 F
17:27:04.912 -> Average Cal sensor Temp.: 96.07 Ambient Temp.: 95.85 F
17:27:06.463 -> Average Cal sensor Temp.: 96.20 Ambient Temp.: 95.81 F
17:27:08.016 -> Average Cal sensor Temp.: 96.15 Ambient Temp.: 95.81 F
17:27:09.571 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 95.81 F
17:27:11.119 -> Average Cal sensor Temp.: 96.31 Ambient Temp.: 95.81 F
17:27:12.672 -> Average Cal sensor Temp.: 96.02 Ambient Temp.: 95.81 F
17:27:14.261 -> Average Cal sensor Temp.: 96.00 Ambient Temp.: 95.81 F
17:27:15.811 -> Average Cal sensor Temp.: 95.98 Ambient Temp.: 95.81 F
17:27:17.361 -> Average Cal sensor Temp.: 96.04 Ambient Temp.: 95.88 F
17:27:18.910 -> Average Cal sensor Temp.: 96.39 Ambient Temp.: 95.85 F
17:27:20.460 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 95.85 F
17:27:22.043 -> Average Cal sensor Temp.: 96.31 Ambient Temp.: 95.85 F
17:27:23.598 -> Average Cal sensor Temp.: 96.70 Ambient Temp.: 95.85 F
17:27:25.147 -> Average Cal sensor Temp.: 96.44 Ambient Temp.: 95.85 F
17:27:26.705 -> Average Cal sensor Temp.: 95.76 Ambient Temp.: 95.85 F
17:27:28.259 -> Average Cal sensor Temp.: 95.75 Ambient Temp.: 95.77 F
17:27:29.809 -> Average Cal sensor Temp.: 95.45 Ambient Temp.: 95.85 F
17:27:31.360 -> Average Cal sensor Temp.: 96.04 Ambient Temp.: 95.81 F
17:27:32.907 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 95.88 F
17:27:34.466 -> Average Cal sensor Temp.: 96.25 Ambient Temp.: 95.85 F
17:27:36.051 -> Average Cal sensor Temp.: 96.23 Ambient Temp.: 95.81 F
17:27:37.599 -> Average Cal sensor Temp.: 96.23 Ambient Temp.: 95.85 F
17:27:39.151 -> Average Cal sensor Temp.: 96.28 Ambient Temp.: 95.77 F
17:27:40.703 -> Average Cal sensor Temp.: 96.51 Ambient Temp.: 95.85 F
17:27:42.253 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 95.85 F
17:27:43.801 -> Average Cal sensor Temp.: 95.93 Ambient Temp.: 95.81 F
17:27:45.357 -> Average Cal sensor Temp.: 96.33 Ambient Temp.: 95.85 F
17:27:46.915 -> Average Cal sensor Temp.: 96.31 Ambient Temp.: 95.81 F
17:27:48.472 -> Average Cal sensor Temp.: 96.22 Ambient Temp.: 95.81 F
17:27:50.026 -> Average Cal sensor Temp.: 96.27 Ambient Temp.: 95.81 F
17:27:51.582 -> Average Cal sensor Temp.: 96.25 Ambient Temp.: 95.81 F
17:27:53.138 -> Average Cal sensor Temp.: 96.33 Ambient Temp.: 95.77 F
17:27:54.689 -> Average Cal sensor Temp.: 96.09 Ambient Temp.: 95.85 F
17:27:56.249 -> Average Cal sensor Temp.: 96.26 Ambient Temp.: 95.88 F
17:27:57.835 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 95.81 F
17:27:59.361 -> Average Cal sensor Temp.: 95.64 Ambient Temp.: 95.81 F
17:28:00.950 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 95.81 F
17:28:02.503 -> Average Cal sensor Temp.: 96.16 Ambient Temp.: 95.81 F
17:28:04.049 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 95.81 F
17:28:05.609 -> Average Cal sensor Temp.: 96.33 Ambient Temp.: 95.81 F
17:28:07.154 -> Average Cal sensor Temp.: 96.56 Ambient Temp.: 95.85 F
17:28:08.703 -> Average Cal sensor Temp.: 96.46 Ambient Temp.: 95.85 F
17:28:10.265 -> Average Cal sensor Temp.: 96.46 Ambient Temp.: 95.81 F
17:28:11.818 -> Average Cal sensor Temp.: 96.33 Ambient Temp.: 95.81 F
17:28:13.404 -> Average Cal sensor Temp.: 96.27 Ambient Temp.: 95.81 F
17:28:14.956 -> Average Cal sensor Temp.: 96.34 Ambient Temp.: 95.88 F
17:28:16.508 -> Average Cal sensor Temp.: 96.26 Ambient Temp.: 95.85 F
17:28:18.063 -> Average Cal sensor Temp.: 96.33 Ambient Temp.: 95.85 F
17:28:19.616 -> Average Cal sensor Temp.: 96.52 Ambient Temp.: 95.81 F
17:28:21.166 -> Average Cal sensor Temp.: 96.52 Ambient Temp.: 95.85 F
17:28:22.721 -> Average Cal sensor Temp.: 96.25 Ambient Temp.: 95.81 F
17:28:24.279 -> Average Cal sensor Temp.: 96.37 Ambient Temp.: 95.88 F
17:28:25.832 -> Average Cal sensor Temp.: 96.35 Ambient Temp.: 95.85 F
17:28:27.389 -> Average Cal sensor Temp.: 96.31 Ambient Temp.: 95.85 F
17:28:28.936 -> Average Cal sensor Temp.: 96.55 Ambient Temp.: 95.85 F
17:28:30.491 -> Average Cal sensor Temp.: 96.27 Ambient Temp.: 95.81 F
17:28:32.075 -> Average Cal sensor Temp.: 95.94 Ambient Temp.: 95.85 F
17:28:33.631 -> Average Cal sensor Temp.: 96.34 Ambient Temp.: 95.77 F
17:28:35.175 -> Average Cal sensor Temp.: 96.29 Ambient Temp.: 95.85 F
17:28:36.725 -> Average Cal sensor Temp.: 96.46 Ambient Temp.: 95.85 F
17:28:38.274 -> Average Cal sensor Temp.: 96.08 Ambient Temp.: 95.81 F
17:28:39.857 -> Average Cal sensor Temp.: 96.37 Ambient Temp.: 95.77 F
17:28:41.410 -> Average Cal sensor Temp.: 96.44 Ambient Temp.: 95.88 F
17:28:42.968 -> Average Cal sensor Temp.: 96.53 Ambient Temp.: 95.85 F
17:28:44.523 -> Average Cal sensor Temp.: 96.32 Ambient Temp.: 95.88 F
17:28:46.075 -> Average Cal sensor Temp.: 96.03 Ambient Temp.: 95.88 F
17:28:47.629 -> Average Cal sensor Temp.: 96.22 Ambient Temp.: 95.88 F
17:28:49.176 -> Average Cal sensor Temp.: 96.51 Ambient Temp.: 95.85 F
17:28:50.728 -> Average Cal sensor Temp.: 96.46 Ambient Temp.: 95.85 F
17:28:52.283 -> Average Cal sensor Temp.: 96.47 Ambient Temp.: 95.81 F
17:28:53.837 -> Average Cal sensor Temp.: 96.03 Ambient Temp.: 95.88 F
17:28:55.424 -> Average Cal sensor Temp.: 96.23 Ambient Temp.: 95.85 F
17:28:56.974 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 95.81 F
17:28:58.527 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 95.88 F
17:29:00.076 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 95.85 F
17:29:01.629 -> Average Cal sensor Temp.: 96.24 Ambient Temp.: 95.85 F
17:29:03.178 -> Average Cal sensor Temp.: 95.95 Ambient Temp.: 95.81 F
17:29:04.763 -> Average Cal sensor Temp.: 96.38 Ambient Temp.: 95.85 F
17:29:06.311 -> Average Cal sensor Temp.: 96.38 Ambient Temp.: 95.85 F
17:29:07.856 -> Average Cal sensor Temp.: 96.45 Ambient Temp.: 95.88 F
17:29:09.404 -> Average Cal sensor Temp.: 96.58 Ambient Temp.: 95.85 F
17:29:10.963 -> Average Cal sensor Temp.: 96.44 Ambient Temp.: 95.88 F
17:29:12.549 -> Average Cal sensor Temp.: 96.40 Ambient Temp.: 95.88 F
17:29:14.101 -> Average Cal sensor Temp.: 96.45 Ambient Temp.: 95.85 F
17:29:15.649 -> Average Cal sensor Temp.: 96.48 Ambient Temp.: 95.85 F
17:29:17.206 -> Average Cal sensor Temp.: 96.52 Ambient Temp.: 95.85 F
17:29:18.759 -> Average Cal sensor Temp.: 96.57 Ambient Temp.: 95.81 F
17:29:20.316 -> Average Cal sensor Temp.: 96.58 Ambient Temp.: 95.88 F
17:29:21.863 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 95.88 F
17:29:23.417 -> Average Cal sensor Temp.: 96.67 Ambient Temp.: 95.85 F
17:29:24.978 -> Average Cal sensor Temp.: 96.51 Ambient Temp.: 95.85 F
17:29:26.534 -> Average Cal sensor Temp.: 96.26 Ambient Temp.: 95.85 F
17:29:28.089 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 95.85 F
17:29:29.672 -> Average Cal sensor Temp.: 96.37 Ambient Temp.: 95.85 F
17:29:31.223 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 95.85 F
17:29:32.778 -> Average Cal sensor Temp.: 96.13 Ambient Temp.: 95.88 F
17:29:34.326 -> Average Cal sensor Temp.: 95.96 Ambient Temp.: 95.85 F
17:29:35.883 -> Average Cal sensor Temp.: 96.38 Ambient Temp.: 95.85 F
17:29:37.444 -> Average Cal sensor Temp.: 96.61 Ambient Temp.: 95.88 F
17:29:38.996 -> Average Cal sensor Temp.: 96.64 Ambient Temp.: 95.88 F
17:29:40.548 -> Average Cal sensor Temp.: 96.67 Ambient Temp.: 95.85 F
17:29:42.107 -> Average Cal sensor Temp.: 96.69 Ambient Temp.: 95.85 F
17:29:43.684 -> Average Cal sensor Temp.: 96.62 Ambient Temp.: 95.88 F
17:29:45.229 -> Average Cal sensor Temp.: 96.70 Ambient Temp.: 95.85 F
17:29:46.785 -> Average Cal sensor Temp.: 96.43 Ambient Temp.: 95.92 F
17:29:48.334 -> Average Cal sensor Temp.: 96.36 Ambient Temp.: 95.88 F
17:29:49.893 -> Average Cal sensor Temp.: 96.48 Ambient Temp.: 95.88 F
17:29:51.473 -> Average Cal sensor Temp.: 96.43 Ambient Temp.: 95.88 F
17:29:53.027 -> Average Cal sensor Temp.: 96.64 Ambient Temp.: 95.92 F
17:29:54.575 -> Average Cal sensor Temp.: 96.46 Ambient Temp.: 95.92 F
17:29:56.128 -> Average Cal sensor Temp.: 96.38 Ambient Temp.: 95.85 F
17:29:57.677 -> Average Cal sensor Temp.: 96.42 Ambient Temp.: 95.85 F
17:29:59.262 -> Average Cal sensor Temp.: 96.39 Ambient Temp.: 95.92 F
17:30:00.814 -> Average Cal sensor Temp.: 96.35 Ambient Temp.: 95.88 F
17:30:02.364 -> Average Cal sensor Temp.: 96.25 Ambient Temp.: 95.88 F
17:30:03.918 -> Average Cal sensor Temp.: 96.45 Ambient Temp.: 95.92 F
17:30:05.475 -> Average Cal sensor Temp.: 96.57 Ambient Temp.: 95.88 F
17:30:07.033 -> Average Cal sensor Temp.: 96.71 Ambient Temp.: 95.85 F
17:30:08.583 -> Average Cal sensor Temp.: 96.48 Ambient Temp.: 95.88 F
17:30:10.139 -> Average Cal sensor Temp.: 96.57 Ambient Temp.: 95.92 F
17:30:11.698 -> Average Cal sensor Temp.: 96.57 Ambient Temp.: 95.92 F
17:30:13.249 -> Average Cal sensor Temp.: 96.36 Ambient Temp.: 95.92 F
17:30:14.807 -> Average Cal sensor Temp.: 96.26 Ambient Temp.: 95.92 F
17:30:16.356 -> Average Cal sensor Temp.: 96.44 Ambient Temp.: 95.95 F
17:30:17.938 -> Average Cal sensor Temp.: 96.29 Ambient Temp.: 95.92 F
17:30:19.494 -> Average Cal sensor Temp.: 96.59 Ambient Temp.: 95.88 F
17:30:21.042 -> Average Cal sensor Temp.: 96.52 Ambient Temp.: 95.88 F
17:30:22.590 -> Average Cal sensor Temp.: 96.40 Ambient Temp.: 95.85 F
17:30:24.169 -> Average Cal sensor Temp.: 96.66 Ambient Temp.: 95.92 F
17:30:25.725 -> Average Cal sensor Temp.: 96.23 Ambient Temp.: 95.92 F
17:30:27.284 -> Average Cal sensor Temp.: 96.44 Ambient Temp.: 95.92 F
17:30:28.837 -> Average Cal sensor Temp.: 96.45 Ambient Temp.: 95.92 F
17:30:30.383 -> Average Cal sensor Temp.: 96.27 Ambient Temp.: 95.85 F
17:30:31.939 -> Average Cal sensor Temp.: 96.53 Ambient Temp.: 95.88 F
17:30:33.490 -> Average Cal sensor Temp.: 96.37 Ambient Temp.: 95.92 F
17:30:35.038 -> Average Cal sensor Temp.: 96.65 Ambient Temp.: 95.88 F
17:30:36.596 -> Average Cal sensor Temp.: 96.74 Ambient Temp.: 95.85 F
17:30:38.181 -> Average Cal sensor Temp.: 96.66 Ambient Temp.: 95.92 F
17:30:39.738 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 95.88 F
17:30:41.290 -> Average Cal sensor Temp.: 96.89 Ambient Temp.: 95.92 F
17:30:42.841 -> Average Cal sensor Temp.: 96.73 Ambient Temp.: 95.92 F
17:30:44.392 -> Average Cal sensor Temp.: 96.55 Ambient Temp.: 95.88 F
17:30:45.945 -> Average Cal sensor Temp.: 96.89 Ambient Temp.: 95.88 F
17:30:47.496 -> Average Cal sensor Temp.: 96.58 Ambient Temp.: 95.88 F
17:30:49.049 -> Average Cal sensor Temp.: 96.55 Ambient Temp.: 95.88 F
17:30:50.638 -> Average Cal sensor Temp.: 96.46 Ambient Temp.: 95.92 F
17:30:52.191 -> Average Cal sensor Temp.: 96.35 Ambient Temp.: 95.88 F
17:30:53.738 -> Average Cal sensor Temp.: 96.42 Ambient Temp.: 95.88 F
17:30:55.287 -> Average Cal sensor Temp.: 96.61 Ambient Temp.: 95.85 F
17:30:56.843 -> Average Cal sensor Temp.: 96.47 Ambient Temp.: 95.92 F
17:30:58.423 -> Average Cal sensor Temp.: 96.37 Ambient Temp.: 95.92 F
17:30:59.976 -> Average Cal sensor Temp.: 96.65 Ambient Temp.: 95.92 F
17:31:01.529 -> Average Cal sensor Temp.: 96.74 Ambient Temp.: 95.92 F
17:31:03.080 -> Average Cal sensor Temp.: 96.71 Ambient Temp.: 95.92 F
17:31:04.633 -> Average Cal sensor Temp.: 96.60 Ambient Temp.: 95.92 F
17:31:06.183 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 95.92 F
17:31:07.739 -> Average Cal sensor Temp.: 96.96 Ambient Temp.: 95.88 F
17:31:09.287 -> Average Cal sensor Temp.: 96.86 Ambient Temp.: 95.92 F
17:31:10.875 -> Average Cal sensor Temp.: 96.68 Ambient Temp.: 95.92 F
17:31:12.428 -> Average Cal sensor Temp.: 96.66 Ambient Temp.: 95.88 F
17:31:13.978 -> Average Cal sensor Temp.: 96.76 Ambient Temp.: 95.88 F
17:31:15.534 -> Average Cal sensor Temp.: 96.56 Ambient Temp.: 95.99 F
17:31:17.083 -> Average Cal sensor Temp.: 96.83 Ambient Temp.: 95.95 F
17:31:18.665 -> Average Cal sensor Temp.: 96.56 Ambient Temp.: 95.88 F
17:31:20.191 -> Average Cal sensor Temp.: 96.38 Ambient Temp.: 95.85 F
17:31:21.778 -> Average Cal sensor Temp.: 96.66 Ambient Temp.: 95.88 F
17:31:23.329 -> Average Cal sensor Temp.: 96.16 Ambient Temp.: 95.88 F
17:31:24.879 -> Average Cal sensor Temp.: 96.21 Ambient Temp.: 95.92 F
17:31:26.431 -> Average Cal sensor Temp.: 96.52 Ambient Temp.: 95.85 F
17:31:27.985 -> Average Cal sensor Temp.: 96.43 Ambient Temp.: 95.88 F
17:31:29.543 -> Average Cal sensor Temp.: 96.50 Ambient Temp.: 95.92 F
17:31:31.093 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 95.92 F
17:31:32.647 -> Average Cal sensor Temp.: 96.02 Ambient Temp.: 95.88 F
17:31:34.232 -> Average Cal sensor Temp.: 95.95 Ambient Temp.: 95.92 F
17:31:35.782 -> Average Cal sensor Temp.: 96.22 Ambient Temp.: 95.92 F
17:31:37.338 -> Average Cal sensor Temp.: 96.26 Ambient Temp.: 95.92 F
17:31:38.891 -> Average Cal sensor Temp.: 96.60 Ambient Temp.: 95.92 F
17:31:40.448 -> Average Cal sensor Temp.: 96.88 Ambient Temp.: 95.92 F
17:31:42.003 -> Average Cal sensor Temp.: 96.72 Ambient Temp.: 95.88 F
17:31:43.558 -> Average Cal sensor Temp.: 96.58 Ambient Temp.: 95.92 F
17:31:45.109 -> Average Cal sensor Temp.: 96.62 Ambient Temp.: 95.92 F
17:31:46.693 -> Average Cal sensor Temp.: 96.78 Ambient Temp.: 95.88 F
17:31:48.243 -> Average Cal sensor Temp.: 96.62 Ambient Temp.: 95.88 F
17:31:49.797 -> Average Cal sensor Temp.: 96.12 Ambient Temp.: 95.95 F
17:31:51.347 -> Average Cal sensor Temp.: 96.47 Ambient Temp.: 95.92 F
17:31:52.902 -> Average Cal sensor Temp.: 96.29 Ambient Temp.: 95.92 F
17:31:54.454 -> Average Cal sensor Temp.: 96.16 Ambient Temp.: 95.92 F
17:31:56.009 -> Average Cal sensor Temp.: 96.42 Ambient Temp.: 95.92 F
17:31:57.565 -> Average Cal sensor Temp.: 96.55 Ambient Temp.: 95.92 F
17:31:59.119 -> Average Cal sensor Temp.: 96.54 Ambient Temp.: 95.92 F
17:32:00.673 -> Average Cal sensor Temp.: 96.47 Ambient Temp.: 95.92 F
17:32:02.257 -> Average Cal sensor Temp.: 96.61 Ambient Temp.: 95.95 F
17:32:03.807 -> Average Cal sensor Temp.: 96.78 Ambient Temp.: 95.95 F
17:32:05.361 -> Average Cal sensor Temp.: 96.76 Ambient Temp.: 95.95 F
17:32:06.908 -> Average Cal sensor Temp.: 96.89 Ambient Temp.: 95.92 F
17:32:08.460 -> Average Cal sensor Temp.: 97.20 Ambient Temp.: 95.95 F
17:32:10.018 -> Average Cal sensor Temp.: 97.23 Ambient Temp.: 95.92 F
17:32:11.568 -> Average Cal sensor Temp.: 96.95 Ambient Temp.: 95.92 F
17:32:13.121 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 95.99 F
17:32:14.679 -> Average Cal sensor Temp.: 96.90 Ambient Temp.: 95.99 F
17:32:16.268 -> Average Cal sensor Temp.: 97.01 Ambient Temp.: 95.95 F
17:32:17.795 -> Average Cal sensor Temp.: 96.84 Ambient Temp.: 96.03 F
17:32:19.383 -> Average Cal sensor Temp.: 96.79 Ambient Temp.: 95.95 F
17:32:20.935 -> Average Cal sensor Temp.: 96.88 Ambient Temp.: 95.95 F
17:32:22.488 -> Average Cal sensor Temp.: 96.92 Ambient Temp.: 95.99 F
17:32:24.038 -> Average Cal sensor Temp.: 96.80 Ambient Temp.: 95.99 F
17:32:25.588 -> Average Cal sensor Temp.: 96.54 Ambient Temp.: 95.95 F
17:32:27.137 -> Average Cal sensor Temp.: 96.98 Ambient Temp.: 95.99 F
17:32:28.724 -> Average Cal sensor Temp.: 97.09 Ambient Temp.: 95.99 F
17:32:30.275 -> Average Cal sensor Temp.: 96.82 Ambient Temp.: 95.99 F
17:32:31.827 -> Average Cal sensor Temp.: 97.10 Ambient Temp.: 96.03 F
17:32:33.383 -> Average Cal sensor Temp.: 96.80 Ambient Temp.: 95.99 F
17:32:34.940 -> Average Cal sensor Temp.: 96.84 Ambient Temp.: 95.99 F
17:32:36.500 -> Average Cal sensor Temp.: 96.92 Ambient Temp.: 95.95 F
17:32:38.054 -> Average Cal sensor Temp.: 97.11 Ambient Temp.: 95.99 F
17:32:39.608 -> Average Cal sensor Temp.: 97.26 Ambient Temp.: 95.95 F
17:32:41.158 -> Average Cal sensor Temp.: 97.06 Ambient Temp.: 95.95 F
17:32:42.740 -> Average Cal sensor Temp.: 96.93 Ambient Temp.: 95.99 F
17:32:44.288 -> Average Cal sensor Temp.: 96.94 Ambient Temp.: 96.03 F
17:32:45.839 -> Average Cal sensor Temp.: 96.99 Ambient Temp.: 95.99 F
17:32:47.393 -> Average Cal sensor Temp.: 96.86 Ambient Temp.: 95.99 F
17:32:48.948 -> Average Cal sensor Temp.: 96.65 Ambient Temp.: 95.99 F
17:32:50.529 -> Average Cal sensor Temp.: 96.90 Ambient Temp.: 96.03 F
17:32:52.081 -> Average Cal sensor Temp.: 96.44 Ambient Temp.: 95.99 F
17:32:53.641 -> Average Cal sensor Temp.: 96.47 Ambient Temp.: 95.95 F
17:32:55.189 -> Average Cal sensor Temp.: 96.39 Ambient Temp.: 95.99 F
17:32:56.739 -> Average Cal sensor Temp.: 96.39 Ambient Temp.: 95.95 F
17:32:58.289 -> Average Cal sensor Temp.: 96.70 Ambient Temp.: 95.99 F
17:32:59.870 -> Average Cal sensor Temp.: 96.63 Ambient Temp.: 95.95 F
17:33:01.421 -> Average Cal sensor Temp.: 96.60 Ambient Temp.: 95.95 F
17:33:02.972 -> Average Cal sensor Temp.: 96.36 Ambient Temp.: 95.99 F
17:33:04.527 -> Average Cal sensor Temp.: 96.59 Ambient Temp.: 95.95 F
17:33:06.082 -> Average Cal sensor Temp.: 96.44 Ambient Temp.: 96.03 F
17:33:07.635 -> Average Cal sensor Temp.: 96.49 Ambient Temp.: 95.99 F
17:33:09.187 -> Average Cal sensor Temp.: 96.53 Ambient Temp.: 95.99 F
17:33:10.775 -> Average Cal sensor Temp.: 96.49 Ambient Temp.: 95.99 F
17:33:12.329 -> Average Cal sensor Temp.: 96.77 Ambient Temp.: 95.99 F
17:33:13.878 -> Average Cal sensor Temp.: 96.52 Ambient Temp.: 95.99 F
17:33:15.433 -> Average Cal sensor Temp.: 96.74 Ambient Temp.: 96.03 F
17:33:16.981 -> Average Cal sensor Temp.: 96.68 Ambient Temp.: 95.99 F
17:33:18.530 -> Average Cal sensor Temp.: 96.87 Ambient Temp.: 95.99 F
17:33:20.114 -> Average Cal sensor Temp.: 96.65 Ambient Temp.: 95.99 F
17:33:21.670 -> Average Cal sensor Temp.: 96.78 Ambient Temp.: 95.99 F
17:33:23.214 -> Average Cal sensor Temp.: 96.97 Ambient Temp.: 95.99 F
17:33:24.763 -> Average Cal sensor Temp.: 96.89 Ambient Temp.: 95.99 F
17:33:26.349 -> Average Cal sensor Temp.: 97.12 Ambient Temp.: 95.99 F
17:33:27.873 -> Average Cal sensor Temp.: 97.09 Ambient Temp.: 95.95 F
17:33:29.428 -> Average Cal sensor Temp.: 97.18 Ambient Temp.: 95.99 F
17:33:31.011 -> Average Cal sensor Temp.: 97.09 Ambient Temp.: 95.95 F
17:33:32.560 -> Average Cal sensor Temp.: 97.16 Ambient Temp.: 96.10 F
17:33:34.115 -> Average Cal sensor Temp.: 97.19 Ambient Temp.: 96.03 F
17:33:35.671 -> Average Cal sensor Temp.: 97.01 Ambient Temp.: 96.03 F
17:33:37.223 -> Average Cal sensor Temp.: 96.69 Ambient Temp.: 96.03 F
17:33:38.776 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 95.99 F
17:33:40.326 -> Average Cal sensor Temp.: 97.01 Ambient Temp.: 95.99 F
17:33:41.909 -> Average Cal sensor Temp.: 97.25 Ambient Temp.: 96.10 F
17:33:43.457 -> Average Cal sensor Temp.: 97.40 Ambient Temp.: 95.99 F
17:33:45.011 -> Average Cal sensor Temp.: 97.34 Ambient Temp.: 96.03 F
17:33:46.559 -> Average Cal sensor Temp.: 97.43 Ambient Temp.: 96.10 F
17:33:48.111 -> Average Cal sensor Temp.: 97.31 Ambient Temp.: 95.99 F
17:33:49.662 -> Average Cal sensor Temp.: 97.41 Ambient Temp.: 96.10 F
17:33:51.213 -> Average Cal sensor Temp.: 97.00 Ambient Temp.: 95.99 F
17:33:52.796 -> Average Cal sensor Temp.: 97.05 Ambient Temp.: 96.03 F
17:33:54.348 -> Average Cal sensor Temp.: 97.08 Ambient Temp.: 95.99 F
17:33:55.905 -> Average Cal sensor Temp.: 98.39 Ambient Temp.: 96.03 F
17:33:57.465 -> Average Cal sensor Temp.: 97.97 Ambient Temp.: 96.10 F
17:33:59.017 -> Average Cal sensor Temp.: 97.29 Ambient Temp.: 95.99 F
17:34:00.572 -> Average Cal sensor Temp.: 97.46 Ambient Temp.: 95.99 F
17:34:02.120 -> Average Cal sensor Temp.: 97.47 Ambient Temp.: 96.03 F
17:34:03.671 -> Average Cal sensor Temp.: 96.95 Ambient Temp.: 96.13 F
17:34:05.229 -> Average Cal sensor Temp.: 96.93 Ambient Temp.: 96.10 F
17:34:06.811 -> Average Cal sensor Temp.: 96.89 Ambient Temp.: 96.03 F
17:34:08.364 -> Average Cal sensor Temp.: 97.22 Ambient Temp.: 96.03 F
17:34:09.915 -> Average Cal sensor Temp.: 97.28 Ambient Temp.: 96.03 F
17:34:11.473 -> Average Cal sensor Temp.: 97.26 Ambient Temp.: 96.03 F
17:34:13.023 -> Average Cal sensor Temp.: 96.79 Ambient Temp.: 96.10 F
17:34:14.574 -> Average Cal sensor Temp.: 96.99 Ambient Temp.: 96.03 F
17:34:16.129 -> Average Cal sensor Temp.: 96.89 Ambient Temp.: 96.03 F
17:34:17.683 -> Average Cal sensor Temp.: 96.98 Ambient Temp.: 96.03 F
17:34:19.269 -> Average Cal sensor Temp.: 97.08 Ambient Temp.: 96.03 F
17:34:20.824 -> Average Cal sensor Temp.: 96.97 Ambient Temp.: 95.99 F
17:34:22.378 -> Average Cal sensor Temp.: 97.36 Ambient Temp.: 96.03 F
17:34:23.929 -> Average Cal sensor Temp.: 96.81 Ambient Temp.: 96.10 F
17:34:25.481 -> Average Cal sensor Temp.: 97.25 Ambient Temp.: 96.03 F
17:34:27.039 -> Average Cal sensor Temp.: 97.02 Ambient Temp.: 95.99 F
17:34:28.592 -> Average Cal sensor Temp.: 97.07 Ambient Temp.: 96.03 F
17:34:30.149 -> Average Cal sensor Temp.: 96.93 Ambient Temp.: 96.03 F
17:34:31.701 -> Average Cal sensor Temp.: 97.23 Ambient Temp.: 96.03 F
17:34:33.251 -> Average Cal sensor Temp.: 97.48 Ambient Temp.: 96.10 F
17:34:34.827 -> Average Cal sensor Temp.: 97.15 Ambient Temp.: 96.10 F
17:34:36.380 -> Average Cal sensor Temp.: 97.10 Ambient Temp.: 96.10 F
17:34:37.929 -> Average Cal sensor Temp.: 97.36 Ambient Temp.: 96.03 F
17:34:39.479 -> Average Cal sensor Temp.: 97.10 Ambient Temp.: 96.03 F
17:34:41.033 -> Average Cal sensor Temp.: 97.13 Ambient Temp.: 96.03 F
17:34:42.583 -> Average Cal sensor Temp.: 96.89 Ambient Temp.: 96.13 F
17:34:44.138 -> Average Cal sensor Temp.: 97.00 Ambient Temp.: 96.10 F
17:34:45.727 -> Average Cal sensor Temp.: 96.31 Ambient Temp.: 96.03 F
17:34:47.276 -> Average Cal sensor Temp.: 96.51 Ambient Temp.: 96.10 F
17:34:48.826 -> Average Cal sensor Temp.: 96.86 Ambient Temp.: 96.10 F
17:34:50.377 -> Average Cal sensor Temp.: 96.92 Ambient Temp.: 96.10 F
17:34:51.932 -> Average Cal sensor Temp.: 97.02 Ambient Temp.: 96.03 F
17:34:53.486 -> Average Cal sensor Temp.: 96.97 Ambient Temp.: 95.99 F
17:34:55.044 -> Average Cal sensor Temp.: 96.68 Ambient Temp.: 96.03 F
17:34:56.593 -> Average Cal sensor Temp.: 96.86 Ambient Temp.: 96.10 F
17:34:58.180 -> Average Cal sensor Temp.: 96.88 Ambient Temp.: 95.99 F
17:34:59.735 -> Average Cal sensor Temp.: 97.19 Ambient Temp.: 96.03 F
17:35:01.286 -> Average Cal sensor Temp.: 97.01 Ambient Temp.: 96.10 F
17:35:02.836 -> Average Cal sensor Temp.: 97.43 Ambient Temp.: 96.03 F
17:35:04.383 -> Average Cal sensor Temp.: 97.41 Ambient Temp.: 96.03 F
17:35:05.939 -> Average Cal sensor Temp.: 96.94 Ambient Temp.: 96.03 F
17:35:07.524 -> Average Cal sensor Temp.: 96.98 Ambient Temp.: 96.10 F
17:35:09.074 -> Average Cal sensor Temp.: 96.46 Ambient Temp.: 95.99 F
17:35:10.627 -> Average Cal sensor Temp.: 96.67 Ambient Temp.: 96.10 F
17:35:12.182 -> Average Cal sensor Temp.: 96.65 Ambient Temp.: 96.03 F
17:35:13.733 -> Average Cal sensor Temp.: 96.77 Ambient Temp.: 96.10 F
17:35:15.280 -> Average Cal sensor Temp.: 97.13 Ambient Temp.: 96.10 F
17:35:16.863 -> Average Cal sensor Temp.: 97.22 Ambient Temp.: 96.03 F
17:35:18.388 -> Average Cal sensor Temp.: 97.26 Ambient Temp.: 96.03 F
17:35:19.972 -> Average Cal sensor Temp.: 96.80 Ambient Temp.: 96.03 F
17:35:21.525 -> Average Cal sensor Temp.: 97.31 Ambient Temp.: 96.03 F
17:35:23.082 -> Average Cal sensor Temp.: 97.09 Ambient Temp.: 95.95 F
17:35:24.639 -> Average Cal sensor Temp.: 96.98 Ambient Temp.: 96.03 F
17:35:26.193 -> Average Cal sensor Temp.: 97.03 Ambient Temp.: 96.03 F
17:35:27.742 -> Average Cal sensor Temp.: 97.10 Ambient Temp.: 96.03 F
17:35:29.297 -> Average Cal sensor Temp.: 97.01 Ambient Temp.: 96.03 F
17:35:30.852 -> Average Cal sensor Temp.: 97.24 Ambient Temp.: 95.99 F
17:35:32.402 -> Average Cal sensor Temp.: 97.07 Ambient Temp.: 96.03 F
17:35:33.957 -> Average Cal sensor Temp.: 97.13 Ambient Temp.: 96.10 F
17:35:35.508 -> Average Cal sensor Temp.: 97.28 Ambient Temp.: 96.03 F
17:35:37.063 -> Average Cal sensor Temp.: 97.41 Ambient Temp.: 96.10 F
17:35:38.618 -> Average Cal sensor Temp.: 97.41 Ambient Temp.: 96.03 F
17:35:40.207 -> Average Cal sensor Temp.: 97.52 Ambient Temp.: 95.99 F
17:35:41.763 -> Average Cal sensor Temp.: 97.50 Ambient Temp.: 96.03 F
17:35:43.311 -> Average Cal sensor Temp.: 96.96 Ambient Temp.: 96.03 F
17:35:44.861 -> Average Cal sensor Temp.: 97.23 Ambient Temp.: 96.10 F
17:35:46.419 -> Average Cal sensor Temp.: 97.48 Ambient Temp.: 96.03 F
17:35:47.966 -> Average Cal sensor Temp.: 97.22 Ambient Temp.: 96.10 F
17:35:49.519 -> Average Cal sensor Temp.: 96.55 Ambient Temp.: 96.10 F
17:35:51.074 -> Average Cal sensor Temp.: 96.80 Ambient Temp.: 96.03 F
17:35:52.627 -> Average Cal sensor Temp.: 96.27 Ambient Temp.: 96.03 F
17:35:54.215 -> Average Cal sensor Temp.: 96.70 Ambient Temp.: 95.99 F
17:35:55.766 -> Average Cal sensor Temp.: 97.07 Ambient Temp.: 96.10 F
17:35:57.324 -> Average Cal sensor Temp.: 97.24 Ambient Temp.: 96.10 F
17:35:58.879 -> Average Cal sensor Temp.: 97.16 Ambient Temp.: 96.03 F
17:36:00.435 -> Average Cal sensor Temp.: 97.10 Ambient Temp.: 96.10 F
17:36:01.991 -> Average Cal sensor Temp.: 96.92 Ambient Temp.: 95.99 F
17:36:03.543 -> Average Cal sensor Temp.: 97.15 Ambient Temp.: 96.10 F
17:36:05.098 -> Average Cal sensor Temp.: 97.14 Ambient Temp.: 96.03 F
17:36:06.659 -> Average Cal sensor Temp.: 96.86 Ambient Temp.: 96.10 F
17:36:08.216 -> Average Cal sensor Temp.: 96.97 Ambient Temp.: 96.03 F
17:36:09.778 -> Average Cal sensor Temp.: 97.21 Ambient Temp.: 96.03 F
17:36:11.332 -> Average Cal sensor Temp.: 97.47 Ambient Temp.: 96.03 F
17:36:12.882 -> Average Cal sensor Temp.: 97.26 Ambient Temp.: 96.03 F
17:36:14.433 -> Average Cal sensor Temp.: 97.43 Ambient Temp.: 96.03 F
17:36:15.990 -> Average Cal sensor Temp.: 97.10 Ambient Temp.: 96.10 F
17:36:17.548 -> Average Cal sensor Temp.: 97.02 Ambient Temp.: 96.10 F
17:36:19.098 -> Average Cal sensor Temp.: 96.80 Ambient Temp.: 96.03 F
17:36:20.646 -> Average Cal sensor Temp.: 97.10 Ambient Temp.: 95.99 F
17:36:22.201 -> Average Cal sensor Temp.: 97.29 Ambient Temp.: 96.03 F
17:36:23.762 -> Average Cal sensor Temp.: 97.05 Ambient Temp.: 96.10 F
17:36:25.318 -> Average Cal sensor Temp.: 97.38 Ambient Temp.: 95.99 F
17:36:26.875 -> Average Cal sensor Temp.: 97.50 Ambient Temp.: 96.03 F
17:36:28.432 -> Average Cal sensor Temp.: 97.26 Ambient Temp.: 95.95 F
17:36:29.982 -> Average Cal sensor Temp.: 97.45 Ambient Temp.: 95.95 F
17:36:31.535 -> Average Cal sensor Temp.: 97.41 Ambient Temp.: 95.99 F
17:36:33.093 -> Average Cal sensor Temp.: 97.53 Ambient Temp.: 96.03 F
17:36:34.645 -> Average Cal sensor Temp.: 97.53 Ambient Temp.: 96.10 F
17:36:36.203 -> Average Cal sensor Temp.: 97.49 Ambient Temp.: 96.03 F
17:36:37.784 -> Average Cal sensor Temp.: 97.19 Ambient Temp.: 96.03 F
17:36:39.339 -> Average Cal sensor Temp.: 97.30 Ambient Temp.: 96.03 F
17:36:40.892 -> Average Cal sensor Temp.: 97.35 Ambient Temp.: 96.03 F
17:36:42.444 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 96.03 F
17:36:44.000 -> Average Cal sensor Temp.: 97.22 Ambient Temp.: 96.03 F
17:36:45.556 -> Average Cal sensor Temp.: 97.57 Ambient Temp.: 95.99 F
17:36:47.110 -> Average Cal sensor Temp.: 97.58 Ambient Temp.: 96.03 F
17:36:48.665 -> Average Cal sensor Temp.: 97.87 Ambient Temp.: 95.99 F
17:36:50.221 -> Average Cal sensor Temp.: 97.50 Ambient Temp.: 96.03 F
17:36:51.777 -> Average Cal sensor Temp.: 97.04 Ambient Temp.: 96.03 F
17:36:53.329 -> Average Cal sensor Temp.: 97.02 Ambient Temp.: 96.10 F
17:36:54.886 -> Average Cal sensor Temp.: 97.21 Ambient Temp.: 95.99 F
17:36:56.438 -> Average Cal sensor Temp.: 97.51 Ambient Temp.: 95.99 F
17:36:57.990 -> Average Cal sensor Temp.: 97.64 Ambient Temp.: 95.99 F
17:36:59.540 -> Average Cal sensor Temp.: 97.41 Ambient Temp.: 96.03 F
17:37:01.094 -> Average Cal sensor Temp.: 97.71 Ambient Temp.: 95.99 F
17:37:02.679 -> Average Cal sensor Temp.: 97.84 Ambient Temp.: 96.03 F
17:37:04.206 -> Average Cal sensor Temp.: 97.84 Ambient Temp.: 95.99 F
17:37:05.788 -> Average Cal sensor Temp.: 97.85 Ambient Temp.: 96.03 F
17:37:07.338 -> Average Cal sensor Temp.: 97.33 Ambient Temp.: 95.95 F
17:37:08.895 -> Average Cal sensor Temp.: 97.15 Ambient Temp.: 96.03 F
17:37:10.449 -> Average Cal sensor Temp.: 97.03 Ambient Temp.: 95.99 F
17:37:12.001 -> Average Cal sensor Temp.: 97.42 Ambient Temp.: 95.99 F
17:37:13.556 -> Average Cal sensor Temp.: 97.58 Ambient Temp.: 95.99 F
17:37:15.105 -> Average Cal sensor Temp.: 97.53 Ambient Temp.: 96.03 F
17:37:16.654 -> Average Cal sensor Temp.: 97.54 Ambient Temp.: 95.99 F
17:37:18.240 -> Average Cal sensor Temp.: 97.38 Ambient Temp.: 95.99 F
17:37:19.796 -> Average Cal sensor Temp.: 97.45 Ambient Temp.: 95.99 F
17:37:21.350 -> Average Cal sensor Temp.: 97.53 Ambient Temp.: 96.10 F
17:37:22.907 -> Average Cal sensor Temp.: 97.66 Ambient Temp.: 96.03 F
17:37:24.462 -> Average Cal sensor Temp.: 97.71 Ambient Temp.: 96.10 F
17:37:26.017 -> Average Cal sensor Temp.: 97.42 Ambient Temp.: 96.03 F
17:37:27.566 -> Average Cal sensor Temp.: 97.48 Ambient Temp.: 96.03 F
17:37:29.109 -> Average Cal sensor Temp.: 97.71 Ambient Temp.: 96.10 F
17:37:30.691 -> Average Cal sensor Temp.: 97.73 Ambient Temp.: 96.03 F
17:37:32.241 -> Average Cal sensor Temp.: 97.55 Ambient Temp.: 95.99 F
17:37:33.795 -> Average Cal sensor Temp.: 97.89 Ambient Temp.: 96.10 F
17:37:35.352 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 95.99 F
17:37:36.899 -> Average Cal sensor Temp.: 97.96 Ambient Temp.: 96.10 F
17:37:38.443 -> Average Cal sensor Temp.: 97.89 Ambient Temp.: 96.10 F
17:37:40.025 -> Average Cal sensor Temp.: 97.63 Ambient Temp.: 96.10 F
17:37:41.583 -> Average Cal sensor Temp.: 97.49 Ambient Temp.: 96.03 F
17:37:43.137 -> Average Cal sensor Temp.: 97.54 Ambient Temp.: 96.10 F
17:37:44.689 -> Average Cal sensor Temp.: 97.42 Ambient Temp.: 96.10 F
17:37:46.242 -> Average Cal sensor Temp.: 97.52 Ambient Temp.: 96.10 F
17:37:47.796 -> Average Cal sensor Temp.: 97.95 Ambient Temp.: 96.10 F
17:37:49.340 -> Average Cal sensor Temp.: 97.99 Ambient Temp.: 96.03 F
17:37:50.893 -> Average Cal sensor Temp.: 97.55 Ambient Temp.: 96.03 F
17:37:52.479 -> Average Cal sensor Temp.: 97.77 Ambient Temp.: 96.03 F
17:37:54.033 -> Average Cal sensor Temp.: 97.64 Ambient Temp.: 96.13 F
17:37:55.582 -> Average Cal sensor Temp.: 97.84 Ambient Temp.: 96.10 F
17:37:57.135 -> Average Cal sensor Temp.: 97.59 Ambient Temp.: 96.13 F
17:37:58.689 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 96.10 F
17:38:00.239 -> Average Cal sensor Temp.: 98.12 Ambient Temp.: 96.03 F
17:38:01.790 -> Average Cal sensor Temp.: 98.12 Ambient Temp.: 96.13 F
17:38:03.337 -> Average Cal sensor Temp.: 98.03 Ambient Temp.: 96.10 F
17:38:04.920 -> Average Cal sensor Temp.: 97.96 Ambient Temp.: 96.10 F
17:38:06.474 -> Average Cal sensor Temp.: 97.74 Ambient Temp.: 96.13 F
17:38:08.025 -> Average Cal sensor Temp.: 97.74 Ambient Temp.: 96.10 F
17:38:09.575 -> Average Cal sensor Temp.: 97.75 Ambient Temp.: 96.10 F
17:38:11.126 -> Average Cal sensor Temp.: 97.91 Ambient Temp.: 96.10 F
17:38:12.680 -> Average Cal sensor Temp.: 97.64 Ambient Temp.: 96.10 F
17:38:14.237 -> Average Cal sensor Temp.: 97.41 Ambient Temp.: 96.10 F
17:38:15.788 -> Average Cal sensor Temp.: 97.16 Ambient Temp.: 96.10 F
17:38:17.370 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 96.13 F
17:38:18.894 -> Average Cal sensor Temp.: 97.66 Ambient Temp.: 96.13 F
17:38:20.480 -> Average Cal sensor Temp.: 98.02 Ambient Temp.: 96.13 F
17:38:22.031 -> Average Cal sensor Temp.: 97.73 Ambient Temp.: 96.10 F
17:38:23.583 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 96.10 F
17:38:25.139 -> Average Cal sensor Temp.: 98.12 Ambient Temp.: 96.03 F
17:38:26.693 -> Average Cal sensor Temp.: 98.17 Ambient Temp.: 96.10 F
17:38:28.242 -> Average Cal sensor Temp.: 98.21 Ambient Temp.: 96.10 F
17:38:29.799 -> Average Cal sensor Temp.: 97.84 Ambient Temp.: 96.10 F
17:38:31.354 -> Average Cal sensor Temp.: 97.94 Ambient Temp.: 96.13 F
17:38:32.905 -> Average Cal sensor Temp.: 97.74 Ambient Temp.: 96.10 F
17:38:34.461 -> Average Cal sensor Temp.: 97.56 Ambient Temp.: 96.17 F
17:38:36.010 -> Average Cal sensor Temp.: 97.54 Ambient Temp.: 96.10 F
17:38:37.595 -> Average Cal sensor Temp.: 97.59 Ambient Temp.: 96.10 F
17:38:39.141 -> Average Cal sensor Temp.: 97.96 Ambient Temp.: 96.13 F
17:38:40.691 -> Average Cal sensor Temp.: 97.86 Ambient Temp.: 96.13 F
17:38:42.245 -> Average Cal sensor Temp.: 97.52 Ambient Temp.: 96.17 F
17:38:43.800 -> Average Cal sensor Temp.: 97.51 Ambient Temp.: 96.17 F
17:38:45.353 -> Average Cal sensor Temp.: 97.54 Ambient Temp.: 96.10 F
17:38:46.912 -> Average Cal sensor Temp.: 97.80 Ambient Temp.: 96.10 F
17:38:48.459 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 96.10 F
17:38:50.010 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 96.13 F
17:38:51.567 -> Average Cal sensor Temp.: 97.91 Ambient Temp.: 96.17 F
17:38:53.123 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 96.13 F
17:38:54.681 -> Average Cal sensor Temp.: 97.64 Ambient Temp.: 96.13 F
17:38:56.259 -> Average Cal sensor Temp.: 97.52 Ambient Temp.: 96.13 F
17:38:57.809 -> Average Cal sensor Temp.: 97.73 Ambient Temp.: 96.10 F
17:38:59.360 -> Average Cal sensor Temp.: 97.99 Ambient Temp.: 96.10 F
17:39:00.913 -> Average Cal sensor Temp.: 97.84 Ambient Temp.: 96.13 F
17:39:02.465 -> Average Cal sensor Temp.: 97.89 Ambient Temp.: 96.13 F
17:39:04.012 -> Average Cal sensor Temp.: 98.02 Ambient Temp.: 96.10 F
17:39:05.571 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 96.13 F
17:39:07.123 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 96.13 F
17:39:08.713 -> Average Cal sensor Temp.: 98.26 Ambient Temp.: 96.10 F
17:39:10.264 -> Average Cal sensor Temp.: 98.08 Ambient Temp.: 96.17 F
17:39:11.816 -> Average Cal sensor Temp.: 97.16 Ambient Temp.: 96.13 F
17:39:13.367 -> Average Cal sensor Temp.: 97.57 Ambient Temp.: 96.10 F
17:39:14.921 -> Average Cal sensor Temp.: 97.65 Ambient Temp.: 96.13 F
17:39:16.470 -> Average Cal sensor Temp.: 97.61 Ambient Temp.: 96.13 F
17:39:18.021 -> Average Cal sensor Temp.: 98.02 Ambient Temp.: 96.10 F
17:39:19.604 -> Average Cal sensor Temp.: 97.86 Ambient Temp.: 96.10 F
17:39:21.159 -> Average Cal sensor Temp.: 97.97 Ambient Temp.: 96.13 F
17:39:22.711 -> Average Cal sensor Temp.: 97.78 Ambient Temp.: 96.13 F
17:39:24.263 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 96.13 F
17:39:25.816 -> Average Cal sensor Temp.: 97.76 Ambient Temp.: 96.17 F
17:39:27.373 -> Average Cal sensor Temp.: 98.18 Ambient Temp.: 96.13 F
17:39:28.928 -> Average Cal sensor Temp.: 98.00 Ambient Temp.: 96.13 F
17:39:30.484 -> Average Cal sensor Temp.: 98.36 Ambient Temp.: 96.10 F
17:39:32.042 -> Average Cal sensor Temp.: 98.42 Ambient Temp.: 96.13 F
17:39:33.594 -> Average Cal sensor Temp.: 97.84 Ambient Temp.: 96.13 F
17:39:35.144 -> Average Cal sensor Temp.: 97.42 Ambient Temp.: 96.13 F
17:39:36.702 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 96.13 F
17:39:38.248 -> Average Cal sensor Temp.: 97.68 Ambient Temp.: 96.17 F
17:39:39.803 -> Average Cal sensor Temp.: 98.02 Ambient Temp.: 96.10 F
17:39:41.386 -> Average Cal sensor Temp.: 97.96 Ambient Temp.: 96.10 F
17:39:42.942 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 96.03 F
17:39:44.499 -> Average Cal sensor Temp.: 97.41 Ambient Temp.: 96.10 F
17:39:46.051 -> Average Cal sensor Temp.: 97.80 Ambient Temp.: 96.10 F
17:39:47.602 -> Average Cal sensor Temp.: 97.83 Ambient Temp.: 96.10 F
17:39:49.155 -> Average Cal sensor Temp.: 97.96 Ambient Temp.: 96.13 F
17:39:50.707 -> Average Cal sensor Temp.: 98.09 Ambient Temp.: 96.13 F
17:39:52.258 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 96.10 F
17:39:53.839 -> Average Cal sensor Temp.: 98.31 Ambient Temp.: 96.10 F
17:39:55.393 -> Average Cal sensor Temp.: 98.50 Ambient Temp.: 96.17 F
17:39:56.947 -> Average Cal sensor Temp.: 98.26 Ambient Temp.: 96.13 F
17:39:58.494 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 96.13 F
17:40:00.043 -> Average Cal sensor Temp.: 98.07 Ambient Temp.: 96.17 F
17:40:01.598 -> Average Cal sensor Temp.: 97.76 Ambient Temp.: 96.10 F
17:40:03.153 -> Average Cal sensor Temp.: 97.95 Ambient Temp.: 96.13 F
17:40:04.707 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 96.13 F
17:40:06.265 -> Average Cal sensor Temp.: 97.76 Ambient Temp.: 96.10 F
17:40:07.818 -> Average Cal sensor Temp.: 97.44 Ambient Temp.: 96.10 F
17:40:09.378 -> Average Cal sensor Temp.: 97.72 Ambient Temp.: 96.03 F
17:40:10.930 -> Average Cal sensor Temp.: 97.58 Ambient Temp.: 96.10 F
17:40:12.512 -> Average Cal sensor Temp.: 97.56 Ambient Temp.: 96.10 F
17:40:14.058 -> Average Cal sensor Temp.: 97.63 Ambient Temp.: 96.10 F
17:40:15.613 -> Average Cal sensor Temp.: 98.04 Ambient Temp.: 96.10 F
17:40:17.162 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 96.03 F
17:40:18.716 -> Average Cal sensor Temp.: 98.27 Ambient Temp.: 96.10 F
17:40:20.265 -> Average Cal sensor Temp.: 98.28 Ambient Temp.: 96.13 F
17:40:21.849 -> Average Cal sensor Temp.: 97.65 Ambient Temp.: 96.13 F
17:40:23.403 -> Average Cal sensor Temp.: 97.50 Ambient Temp.: 96.10 F
17:40:24.958 -> Average Cal sensor Temp.: 97.60 Ambient Temp.: 96.10 F
17:40:26.516 -> Average Cal sensor Temp.: 97.82 Ambient Temp.: 96.13 F
17:40:28.068 -> Average Cal sensor Temp.: 97.66 Ambient Temp.: 96.17 F
17:40:29.620 -> Average Cal sensor Temp.: 97.59 Ambient Temp.: 96.10 F
17:40:31.170 -> Average Cal sensor Temp.: 98.00 Ambient Temp.: 96.10 F
17:40:32.722 -> Average Cal sensor Temp.: 97.95 Ambient Temp.: 96.13 F
17:40:34.276 -> Average Cal sensor Temp.: 98.14 Ambient Temp.: 96.13 F
17:40:35.833 -> Average Cal sensor Temp.: 97.91 Ambient Temp.: 96.10 F
17:40:37.388 -> Average Cal sensor Temp.: 98.00 Ambient Temp.: 96.10 F
17:40:38.947 -> Average Cal sensor Temp.: 98.14 Ambient Temp.: 96.10 F
17:40:40.527 -> Average Cal sensor Temp.: 98.12 Ambient Temp.: 96.03 F
17:40:42.085 -> Average Cal sensor Temp.: 98.03 Ambient Temp.: 96.03 F
17:40:43.635 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 96.03 F
17:40:45.186 -> Average Cal sensor Temp.: 97.97 Ambient Temp.: 96.03 F
17:40:46.739 -> Average Cal sensor Temp.: 97.84 Ambient Temp.: 96.13 F
17:40:48.290 -> Average Cal sensor Temp.: 97.71 Ambient Temp.: 96.03 F
17:40:49.843 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 96.10 F
17:40:51.394 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 96.03 F
17:40:52.953 -> Average Cal sensor Temp.: 97.96 Ambient Temp.: 96.03 F
17:40:54.534 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 96.03 F
17:40:56.088 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 96.13 F
17:40:57.638 -> Average Cal sensor Temp.: 98.03 Ambient Temp.: 96.10 F
17:40:59.191 -> Average Cal sensor Temp.: 97.86 Ambient Temp.: 96.10 F
17:41:00.738 -> Average Cal sensor Temp.: 97.80 Ambient Temp.: 96.03 F
17:41:02.294 -> Average Cal sensor Temp.: 97.86 Ambient Temp.: 96.10 F
17:41:03.848 -> Average Cal sensor Temp.: 98.14 Ambient Temp.: 96.13 F
17:41:05.404 -> Average Cal sensor Temp.: 97.98 Ambient Temp.: 96.13 F
17:41:06.957 -> Average Cal sensor Temp.: 97.72 Ambient Temp.: 96.13 F
17:41:08.544 -> Average Cal sensor Temp.: 97.82 Ambient Temp.: 96.03 F
17:41:10.094 -> Average Cal sensor Temp.: 97.63 Ambient Temp.: 96.13 F
17:41:11.649 -> Average Cal sensor Temp.: 97.46 Ambient Temp.: 96.13 F
17:41:13.205 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 96.10 F
17:41:14.763 -> Average Cal sensor Temp.: 97.82 Ambient Temp.: 96.10 F
17:41:16.315 -> Average Cal sensor Temp.: 97.98 Ambient Temp.: 95.99 F
17:41:17.868 -> Average Cal sensor Temp.: 97.08 Ambient Temp.: 96.03 F
17:41:19.419 -> Average Cal sensor Temp.: 97.50 Ambient Temp.: 96.03 F
17:41:20.967 -> Average Cal sensor Temp.: 97.87 Ambient Temp.: 96.03 F
17:41:22.520 -> Average Cal sensor Temp.: 97.99 Ambient Temp.: 96.10 F
17:41:24.098 -> Average Cal sensor Temp.: 98.12 Ambient Temp.: 96.03 F
17:41:25.655 -> Average Cal sensor Temp.: 97.84 Ambient Temp.: 96.13 F
17:41:27.206 -> Average Cal sensor Temp.: 97.89 Ambient Temp.: 96.10 F
17:41:28.755 -> Average Cal sensor Temp.: 97.82 Ambient Temp.: 96.10 F
17:41:30.300 -> Average Cal sensor Temp.: 97.83 Ambient Temp.: 96.03 F
17:41:31.884 -> Average Cal sensor Temp.: 97.84 Ambient Temp.: 96.03 F
17:41:33.441 -> Average Cal sensor Temp.: 97.63 Ambient Temp.: 96.10 F
17:41:34.994 -> Average Cal sensor Temp.: 97.62 Ambient Temp.: 96.03 F
17:41:36.547 -> Average Cal sensor Temp.: 98.07 Ambient Temp.: 96.10 F
17:41:38.095 -> Average Cal sensor Temp.: 97.30 Ambient Temp.: 96.10 F
17:41:39.648 -> Average Cal sensor Temp.: 97.65 Ambient Temp.: 96.10 F
17:41:41.199 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 96.10 F
17:41:42.754 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 96.10 F
17:41:44.309 -> Average Cal sensor Temp.: 98.05 Ambient Temp.: 96.10 F
17:41:45.890 -> Average Cal sensor Temp.: 98.07 Ambient Temp.: 96.10 F
17:41:47.440 -> Average Cal sensor Temp.: 98.13 Ambient Temp.: 96.10 F
17:41:48.991 -> Average Cal sensor Temp.: 97.55 Ambient Temp.: 96.10 F
17:41:50.543 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 96.13 F
17:41:52.092 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 96.13 F
17:41:53.679 -> Average Cal sensor Temp.: 97.97 Ambient Temp.: 96.10 F
17:41:55.226 -> Average Cal sensor Temp.: 98.07 Ambient Temp.: 96.17 F
17:41:56.780 -> Average Cal sensor Temp.: 98.24 Ambient Temp.: 96.10 F
17:41:58.331 -> Average Cal sensor Temp.: 97.98 Ambient Temp.: 96.10 F
17:41:59.883 -> Average Cal sensor Temp.: 98.25 Ambient Temp.: 96.10 F
17:42:01.437 -> Average Cal sensor Temp.: 98.47 Ambient Temp.: 96.10 F
17:42:02.987 -> Average Cal sensor Temp.: 98.08 Ambient Temp.: 96.10 F
17:42:04.575 -> Average Cal sensor Temp.: 97.82 Ambient Temp.: 96.10 F
17:42:06.124 -> Average Cal sensor Temp.: 98.42 Ambient Temp.: 96.10 F
17:42:07.677 -> Average Cal sensor Temp.: 99.20 Ambient Temp.: 96.13 F
17:42:09.234 -> Average Cal sensor Temp.: 99.10 Ambient Temp.: 96.13 F
17:42:10.782 -> Average Cal sensor Temp.: 98.83 Ambient Temp.: 96.13 F
17:42:12.330 -> Average Cal sensor Temp.: 98.51 Ambient Temp.: 96.10 F
17:42:13.883 -> Average Cal sensor Temp.: 98.53 Ambient Temp.: 96.13 F
17:42:15.466 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 96.17 F
17:42:17.016 -> Average Cal sensor Temp.: 98.32 Ambient Temp.: 96.13 F
17:42:18.574 -> Average Cal sensor Temp.: 98.22 Ambient Temp.: 96.03 F
17:42:20.122 -> Average Cal sensor Temp.: 97.98 Ambient Temp.: 96.10 F
17:42:21.671 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 96.10 F
17:42:23.227 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 96.10 F
17:42:24.781 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 96.13 F
17:42:26.362 -> Average Cal sensor Temp.: 98.17 Ambient Temp.: 96.13 F
17:42:27.914 -> Average Cal sensor Temp.: 98.13 Ambient Temp.: 96.10 F
17:42:29.459 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 96.13 F
17:42:31.010 -> Average Cal sensor Temp.: 97.85 Ambient Temp.: 96.13 F
17:42:32.560 -> Average Cal sensor Temp.: 98.22 Ambient Temp.: 96.10 F
17:42:34.116 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 96.13 F
17:42:35.699 -> Average Cal sensor Temp.: 97.74 Ambient Temp.: 96.13 F
17:42:37.250 -> Average Cal sensor Temp.: 98.14 Ambient Temp.: 96.17 F
17:42:38.801 -> Average Cal sensor Temp.: 98.23 Ambient Temp.: 96.13 F
17:42:40.351 -> Average Cal sensor Temp.: 98.17 Ambient Temp.: 96.17 F
17:42:41.898 -> Average Cal sensor Temp.: 97.84 Ambient Temp.: 96.13 F
17:42:43.453 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 96.13 F
17:42:45.032 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 96.17 F
17:42:46.585 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 96.13 F
17:42:48.141 -> Average Cal sensor Temp.: 98.04 Ambient Temp.: 96.17 F
17:42:49.693 -> Average Cal sensor Temp.: 97.94 Ambient Temp.: 96.13 F
17:42:51.250 -> Average Cal sensor Temp.: 97.85 Ambient Temp.: 96.13 F
17:42:52.801 -> Average Cal sensor Temp.: 97.93 Ambient Temp.: 96.17 F
17:42:54.353 -> Average Cal sensor Temp.: 98.16 Ambient Temp.: 96.17 F
17:42:55.903 -> Average Cal sensor Temp.: 98.07 Ambient Temp.: 96.13 F
17:42:57.462 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 96.17 F
17:42:59.015 -> Average Cal sensor Temp.: 98.28 Ambient Temp.: 96.13 F
17:43:00.570 -> Average Cal sensor Temp.: 97.97 Ambient Temp.: 96.13 F
17:43:02.126 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 96.13 F
17:43:03.682 -> Average Cal sensor Temp.: 98.13 Ambient Temp.: 96.17 F
17:43:05.234 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 96.10 F
17:43:06.796 -> Average Cal sensor Temp.: 97.22 Ambient Temp.: 96.13 F
17:43:08.349 -> Average Cal sensor Temp.: 97.69 Ambient Temp.: 96.13 F
17:43:09.903 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 96.17 F
17:43:11.459 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 96.10 F
17:43:13.042 -> Average Cal sensor Temp.: 97.99 Ambient Temp.: 96.17 F
17:43:14.593 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 96.13 F
17:43:16.143 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 96.13 F
17:43:17.693 -> Average Cal sensor Temp.: 98.39 Ambient Temp.: 96.10 F
17:43:19.242 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 96.17 F
17:43:20.793 -> Average Cal sensor Temp.: 98.18 Ambient Temp.: 96.17 F
17:43:22.378 -> Average Cal sensor Temp.: 98.04 Ambient Temp.: 96.13 F
17:43:23.928 -> Average Cal sensor Temp.: 98.21 Ambient Temp.: 96.13 F
17:43:25.481 -> Average Cal sensor Temp.: 98.24 Ambient Temp.: 96.13 F
17:43:27.035 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 96.17 F
17:43:28.592 -> Average Cal sensor Temp.: 98.42 Ambient Temp.: 96.17 F
17:43:30.140 -> Average Cal sensor Temp.: 98.62 Ambient Temp.: 96.24 F
17:43:31.692 -> Average Cal sensor Temp.: 98.67 Ambient Temp.: 96.21 F
17:43:33.244 -> Average Cal sensor Temp.: 98.61 Ambient Temp.: 96.21 F
17:43:34.829 -> Average Cal sensor Temp.: 98.07 Ambient Temp.: 96.21 F
17:43:36.356 -> Average Cal sensor Temp.: 98.21 Ambient Temp.: 96.21 F
17:43:37.912 -> Average Cal sensor Temp.: 98.16 Ambient Temp.: 96.17 F
17:43:39.494 -> Average Cal sensor Temp.: 98.16 Ambient Temp.: 96.17 F
17:43:41.044 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 96.17 F
17:43:42.597 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 96.21 F
17:43:44.144 -> Average Cal sensor Temp.: 97.74 Ambient Temp.: 96.17 F
17:43:45.723 -> Average Cal sensor Temp.: 97.24 Ambient Temp.: 96.21 F
17:43:47.248 -> Average Cal sensor Temp.: 97.64 Ambient Temp.: 96.17 F
17:43:48.803 -> Average Cal sensor Temp.: 97.88 Ambient Temp.: 96.21 F
17:43:50.358 -> Average Cal sensor Temp.: 98.16 Ambient Temp.: 96.17 F
17:43:51.947 -> Average Cal sensor Temp.: 97.98 Ambient Temp.: 96.21 F
17:43:53.499 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 96.17 F
17:43:55.050 -> Average Cal sensor Temp.: 97.70 Ambient Temp.: 96.21 F
17:43:56.610 -> Average Cal sensor Temp.: 98.21 Ambient Temp.: 96.24 F
17:43:58.166 -> Average Cal sensor Temp.: 98.36 Ambient Temp.: 96.21 F
17:43:59.717 -> Average Cal sensor Temp.: 98.24 Ambient Temp.: 96.21 F
17:44:01.269 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 96.21 F
17:44:02.818 -> Average Cal sensor Temp.: 98.28 Ambient Temp.: 96.21 F
17:44:04.371 -> Average Cal sensor Temp.: 97.65 Ambient Temp.: 96.21 F
17:44:05.925 -> Average Cal sensor Temp.: 98.08 Ambient Temp.: 96.17 F
17:44:07.477 -> Average Cal sensor Temp.: 98.19 Ambient Temp.: 96.21 F
17:44:09.062 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 96.24 F
17:44:10.618 -> Average Cal sensor Temp.: 98.41 Ambient Temp.: 96.21 F
17:44:12.164 -> Average Cal sensor Temp.: 98.39 Ambient Temp.: 96.17 F
17:44:13.720 -> Average Cal sensor Temp.: 98.39 Ambient Temp.: 96.17 F
17:44:15.267 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 96.21 F
17:44:16.815 -> Average Cal sensor Temp.: 97.75 Ambient Temp.: 96.24 F
17:44:18.369 -> Average Cal sensor Temp.: 98.22 Ambient Temp.: 96.17 F
17:44:19.922 -> Average Cal sensor Temp.: 98.30 Ambient Temp.: 96.24 F
17:44:21.505 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 96.21 F
17:44:23.059 -> Average Cal sensor Temp.: 98.03 Ambient Temp.: 96.17 F
17:44:24.611 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 96.21 F
17:44:26.169 -> Average Cal sensor Temp.: 98.35 Ambient Temp.: 96.13 F
17:44:27.724 -> Average Cal sensor Temp.: 97.65 Ambient Temp.: 96.17 F
17:44:29.277 -> Average Cal sensor Temp.: 98.12 Ambient Temp.: 96.21 F
17:44:30.825 -> Average Cal sensor Temp.: 98.52 Ambient Temp.: 96.17 F
17:44:32.374 -> Average Cal sensor Temp.: 98.36 Ambient Temp.: 96.21 F
17:44:33.927 -> Average Cal sensor Temp.: 98.39 Ambient Temp.: 96.21 F
17:44:35.481 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 96.21 F
17:44:37.066 -> Average Cal sensor Temp.: 98.16 Ambient Temp.: 96.17 F
17:44:38.613 -> Average Cal sensor Temp.: 98.51 Ambient Temp.: 96.17 F
17:44:40.156 -> Average Cal sensor Temp.: 98.04 Ambient Temp.: 96.17 F
17:44:41.704 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 96.17 F
17:44:43.291 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 96.17 F
17:44:44.843 -> Average Cal sensor Temp.: 97.49 Ambient Temp.: 96.21 F
17:44:46.396 -> Average Cal sensor Temp.: 97.65 Ambient Temp.: 96.21 F
17:44:47.948 -> Average Cal sensor Temp.: 97.85 Ambient Temp.: 96.21 F
17:44:49.500 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 96.21 F
17:44:51.049 -> Average Cal sensor Temp.: 98.56 Ambient Temp.: 96.17 F
17:44:52.601 -> Average Cal sensor Temp.: 98.76 Ambient Temp.: 96.13 F
17:44:54.183 -> Average Cal sensor Temp.: 98.42 Ambient Temp.: 96.17 F
17:44:55.734 -> Average Cal sensor Temp.: 98.59 Ambient Temp.: 96.21 F
17:44:57.292 -> Average Cal sensor Temp.: 98.55 Ambient Temp.: 96.21 F
17:44:58.845 -> Average Cal sensor Temp.: 98.17 Ambient Temp.: 96.21 F
17:45:00.400 -> Average Cal sensor Temp.: 98.16 Ambient Temp.: 96.17 F
17:45:01.956 -> Average Cal sensor Temp.: 98.24 Ambient Temp.: 96.17 F
17:45:03.508 -> Average Cal sensor Temp.: 98.17 Ambient Temp.: 96.17 F
17:45:05.060 -> Average Cal sensor Temp.: 98.33 Ambient Temp.: 96.24 F
17:45:06.607 -> Average Cal sensor Temp.: 98.62 Ambient Temp.: 96.21 F
17:45:08.163 -> Average Cal sensor Temp.: 98.61 Ambient Temp.: 96.21 F
17:45:09.720 -> Average Cal sensor Temp.: 98.53 Ambient Temp.: 96.21 F
17:45:11.273 -> Average Cal sensor Temp.: 98.60 Ambient Temp.: 96.21 F
17:45:12.857 -> Average Cal sensor Temp.: 98.73 Ambient Temp.: 96.24 F
17:45:14.384 -> Average Cal sensor Temp.: 98.54 Ambient Temp.: 96.17 F
17:45:15.940 -> Average Cal sensor Temp.: 98.56 Ambient Temp.: 96.21 F
17:45:17.526 -> Average Cal sensor Temp.: 98.63 Ambient Temp.: 96.24 F
17:45:19.080 -> Average Cal sensor Temp.: 98.66 Ambient Temp.: 96.21 F
17:45:20.634 -> Average Cal sensor Temp.: 98.75 Ambient Temp.: 96.24 F
17:45:22.191 -> Average Cal sensor Temp.: 98.73 Ambient Temp.: 96.21 F
17:45:23.743 -> Average Cal sensor Temp.: 98.53 Ambient Temp.: 96.24 F
17:45:25.301 -> Average Cal sensor Temp.: 97.84 Ambient Temp.: 96.24 F
17:45:26.862 -> Average Cal sensor Temp.: 98.10 Ambient Temp.: 96.21 F
17:45:28.414 -> Average Cal sensor Temp.: 98.52 Ambient Temp.: 96.21 F
17:45:29.976 -> Average Cal sensor Temp.: 98.20 Ambient Temp.: 96.21 F
17:45:31.523 -> Average Cal sensor Temp.: 98.38 Ambient Temp.: 96.24 F
17:45:33.073 -> Average Cal sensor Temp.: 98.30 Ambient Temp.: 96.24 F
17:45:34.631 -> Average Cal sensor Temp.: 97.75 Ambient Temp.: 96.24 F
17:45:36.186 -> Average Cal sensor Temp.: 97.41 Ambient Temp.: 96.24 F
17:45:37.735 -> Average Cal sensor Temp.: 97.72 Ambient Temp.: 96.17 F
17:45:39.288 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 96.24 F
17:45:40.850 -> Average Cal sensor Temp.: 98.24 Ambient Temp.: 96.24 F
17:45:42.410 -> Average Cal sensor Temp.: 98.45 Ambient Temp.: 96.21 F
17:45:43.964 -> Average Cal sensor Temp.: 98.62 Ambient Temp.: 96.21 F
17:45:45.516 -> Average Cal sensor Temp.: 98.38 Ambient Temp.: 96.21 F
17:45:47.069 -> Average Cal sensor Temp.: 98.55 Ambient Temp.: 96.24 F
17:45:48.657 -> Average Cal sensor Temp.: 98.49 Ambient Temp.: 96.24 F
17:45:50.213 -> Average Cal sensor Temp.: 98.61 Ambient Temp.: 96.17 F
17:45:51.763 -> Average Cal sensor Temp.: 97.97 Ambient Temp.: 96.24 F
17:45:53.311 -> Average Cal sensor Temp.: 98.38 Ambient Temp.: 96.21 F
17:45:54.856 -> Average Cal sensor Temp.: 98.19 Ambient Temp.: 96.17 F
17:45:56.411 -> Average Cal sensor Temp.: 98.72 Ambient Temp.: 96.24 F
17:45:57.995 -> Average Cal sensor Temp.: 98.63 Ambient Temp.: 96.24 F
17:45:59.543 -> Average Cal sensor Temp.: 98.60 Ambient Temp.: 96.24 F
17:46:01.093 -> Average Cal sensor Temp.: 98.82 Ambient Temp.: 96.17 F
17:46:02.642 -> Average Cal sensor Temp.: 98.97 Ambient Temp.: 96.24 F
17:46:04.228 -> Average Cal sensor Temp.: 98.52 Ambient Temp.: 96.24 F
17:46:05.783 -> Average Cal sensor Temp.: 98.55 Ambient Temp.: 96.24 F
17:46:07.335 -> Average Cal sensor Temp.: 98.21 Ambient Temp.: 96.21 F
17:46:08.891 -> Average Cal sensor Temp.: 98.09 Ambient Temp.: 96.21 F
17:46:10.448 -> Average Cal sensor Temp.: 98.56 Ambient Temp.: 96.21 F
17:46:11.998 -> Average Cal sensor Temp.: 98.55 Ambient Temp.: 96.24 F
17:46:13.549 -> Average Cal sensor Temp.: 98.26 Ambient Temp.: 96.21 F
17:46:15.100 -> Average Cal sensor Temp.: 98.77 Ambient Temp.: 96.24 F
17:46:16.680 -> Average Cal sensor Temp.: 98.71 Ambient Temp.: 96.24 F
17:46:18.232 -> Average Cal sensor Temp.: 98.59 Ambient Temp.: 96.24 F
17:46:19.776 -> Average Cal sensor Temp.: 98.22 Ambient Temp.: 96.21 F
17:46:21.330 -> Average Cal sensor Temp.: 98.43 Ambient Temp.: 96.24 F
17:46:22.884 -> Average Cal sensor Temp.: 98.32 Ambient Temp.: 96.24 F
17:46:24.441 -> Average Cal sensor Temp.: 98.33 Ambient Temp.: 96.24 F
17:46:25.997 -> Average Cal sensor Temp.: 98.49 Ambient Temp.: 96.28 F
17:46:27.551 -> Average Cal sensor Temp.: 98.21 Ambient Temp.: 96.24 F
17:46:29.136 -> Average Cal sensor Temp.: 98.58 Ambient Temp.: 96.24 F
17:46:30.690 -> Average Cal sensor Temp.: 98.79 Ambient Temp.: 96.31 F
17:46:32.241 -> Average Cal sensor Temp.: 98.24 Ambient Temp.: 96.28 F
17:46:33.793 -> Average Cal sensor Temp.: 98.43 Ambient Temp.: 96.24 F
17:46:35.339 -> Average Cal sensor Temp.: 98.49 Ambient Temp.: 96.21 F
17:46:36.921 -> Average Cal sensor Temp.: 97.97 Ambient Temp.: 96.21 F
17:46:38.471 -> Average Cal sensor Temp.: 97.44 Ambient Temp.: 96.24 F
17:46:40.020 -> Average Cal sensor Temp.: 97.69 Ambient Temp.: 96.21 F
17:46:41.573 -> Average Cal sensor Temp.: 97.68 Ambient Temp.: 96.17 F
17:46:43.124 -> Average Cal sensor Temp.: 98.26 Ambient Temp.: 96.21 F
17:46:44.680 -> Average Cal sensor Temp.: 98.63 Ambient Temp.: 96.28 F
17:46:46.240 -> Average Cal sensor Temp.: 97.95 Ambient Temp.: 96.21 F
17:46:47.790 -> Average Cal sensor Temp.: 98.51 Ambient Temp.: 96.24 F
17:46:49.370 -> Average Cal sensor Temp.: 98.91 Ambient Temp.: 96.24 F
17:46:50.916 -> Average Cal sensor Temp.: 98.58 Ambient Temp.: 96.28 F
17:46:52.463 -> Average Cal sensor Temp.: 98.07 Ambient Temp.: 96.21 F
17:46:54.015 -> Average Cal sensor Temp.: 98.38 Ambient Temp.: 96.21 F
17:46:55.597 -> Average Cal sensor Temp.: 98.60 Ambient Temp.: 96.17 F
17:46:57.150 -> Average Cal sensor Temp.: 98.56 Ambient Temp.: 96.24 F
17:46:58.698 -> Average Cal sensor Temp.: 98.71 Ambient Temp.: 96.21 F
17:47:00.250 -> Average Cal sensor Temp.: 98.64 Ambient Temp.: 96.21 F
17:47:01.806 -> Average Cal sensor Temp.: 98.86 Ambient Temp.: 96.21 F
17:47:03.363 -> Average Cal sensor Temp.: 98.83 Ambient Temp.: 96.24 F
17:47:04.917 -> Average Cal sensor Temp.: 99.11 Ambient Temp.: 96.28 F
17:47:06.499 -> Average Cal sensor Temp.: 98.65 Ambient Temp.: 96.28 F
17:47:08.024 -> Average Cal sensor Temp.: 98.73 Ambient Temp.: 96.24 F
17:47:09.587 -> Average Cal sensor Temp.: 99.01 Ambient Temp.: 96.28 F
17:47:11.149 -> Average Cal sensor Temp.: 99.05 Ambient Temp.: 96.28 F
17:47:12.710 -> Average Cal sensor Temp.: 99.25 Ambient Temp.: 96.28 F
17:47:14.264 -> Average Cal sensor Temp.: 99.12 Ambient Temp.: 96.31 F
17:47:15.822 -> Average Cal sensor Temp.: 99.09 Ambient Temp.: 96.28 F
17:47:17.374 -> Average Cal sensor Temp.: 98.84 Ambient Temp.: 96.24 F
17:47:18.931 -> Average Cal sensor Temp.: 98.76 Ambient Temp.: 96.28 F
17:47:20.489 -> Average Cal sensor Temp.: 99.16 Ambient Temp.: 96.31 F
17:47:22.038 -> Average Cal sensor Temp.: 99.09 Ambient Temp.: 96.31 F
17:47:23.595 -> Average Cal sensor Temp.: 99.19 Ambient Temp.: 96.31 F
17:47:25.151 -> Average Cal sensor Temp.: 98.65 Ambient Temp.: 96.31 F
17:47:26.709 -> Average Cal sensor Temp.: 98.44 Ambient Temp.: 96.28 F
17:47:28.297 -> Average Cal sensor Temp.: 98.67 Ambient Temp.: 96.31 F
17:47:29.826 -> Average Cal sensor Temp.: 98.74 Ambient Temp.: 96.31 F
17:47:31.411 -> Average Cal sensor Temp.: 98.78 Ambient Temp.: 96.28 F
17:47:32.964 -> Average Cal sensor Temp.: 98.80 Ambient Temp.: 96.28 F
17:47:34.520 -> Average Cal sensor Temp.: 98.64 Ambient Temp.: 96.35 F
17:47:36.077 -> Average Cal sensor Temp.: 98.55 Ambient Temp.: 96.31 F
17:47:37.635 -> Average Cal sensor Temp.: 98.72 Ambient Temp.: 96.31 F
17:47:39.184 -> Average Cal sensor Temp.: 98.73 Ambient Temp.: 96.31 F
17:47:40.744 -> Average Cal sensor Temp.: 98.85 Ambient Temp.: 96.35 F
17:47:42.297 -> Average Cal sensor Temp.: 98.54 Ambient Temp.: 96.31 F
17:47:43.855 -> Average Cal sensor Temp.: 98.70 Ambient Temp.: 96.31 F
17:47:45.412 -> Average Cal sensor Temp.: 98.78 Ambient Temp.: 96.35 F
17:47:46.973 -> Average Cal sensor Temp.: 98.93 Ambient Temp.: 96.35 F
17:47:48.530 -> Average Cal sensor Temp.: 99.09 Ambient Temp.: 96.35 F
17:47:50.092 -> Average Cal sensor Temp.: 98.99 Ambient Temp.: 96.39 F
17:47:51.646 -> Average Cal sensor Temp.: 99.09 Ambient Temp.: 96.35 F
17:47:53.201 -> Average Cal sensor Temp.: 98.67 Ambient Temp.: 96.39 F
17:47:54.749 -> Average Cal sensor Temp.: 98.57 Ambient Temp.: 96.35 F
17:47:56.312 -> Average Cal sensor Temp.: 98.78 Ambient Temp.: 96.35 F
17:47:57.868 -> Average Cal sensor Temp.: 98.70 Ambient Temp.: 96.35 F
17:47:59.426 -> Average Cal sensor Temp.: 98.59 Ambient Temp.: 96.39 F
17:48:00.982 -> Average Cal sensor Temp.: 98.88 Ambient Temp.: 96.39 F
17:48:02.534 -> Average Cal sensor Temp.: 98.87 Ambient Temp.: 96.35 F
17:48:04.086 -> Average Cal sensor Temp.: 98.83 Ambient Temp.: 96.39 F
17:48:05.644 -> Average Cal sensor Temp.: 99.11 Ambient Temp.: 96.35 F
17:48:07.195 -> Average Cal sensor Temp.: 98.96 Ambient Temp.: 96.39 F
17:48:08.755 -> Average Cal sensor Temp.: 99.16 Ambient Temp.: 96.31 F
17:48:10.318 -> Average Cal sensor Temp.: 99.26 Ambient Temp.: 96.42 F
17:48:11.879 -> Average Cal sensor Temp.: 99.19 Ambient Temp.: 96.39 F
17:48:13.435 -> Average Cal sensor Temp.: 99.30 Ambient Temp.: 96.39 F
17:48:14.988 -> Average Cal sensor Temp.: 99.32 Ambient Temp.: 96.42 F
17:48:16.551 -> Average Cal sensor Temp.: 99.10 Ambient Temp.: 96.42 F
17:48:18.101 -> Average Cal sensor Temp.: 99.27 Ambient Temp.: 96.35 F
17:48:19.662 -> Average Cal sensor Temp.: 99.29 Ambient Temp.: 96.39 F
17:48:21.215 -> Average Cal sensor Temp.: 99.40 Ambient Temp.: 96.42 F
17:48:22.769 -> Average Cal sensor Temp.: 99.23 Ambient Temp.: 96.49 F
17:48:24.319 -> Average Cal sensor Temp.: 99.05 Ambient Temp.: 96.35 F
17:48:25.906 -> Average Cal sensor Temp.: 99.06 Ambient Temp.: 96.42 F
17:48:27.453 -> Average Cal sensor Temp.: 99.06 Ambient Temp.: 96.42 F
17:48:29.011 -> Average Cal sensor Temp.: 99.14 Ambient Temp.: 96.42 F
17:48:30.568 -> Average Cal sensor Temp.: 99.06 Ambient Temp.: 96.53 F
17:48:32.120 -> Average Cal sensor Temp.: 99.27 Ambient Temp.: 96.49 F
17:48:33.674 -> Average Cal sensor Temp.: 98.90 Ambient Temp.: 96.42 F
17:48:35.229 -> Average Cal sensor Temp.: 98.98 Ambient Temp.: 96.49 F
17:48:36.813 -> Average Cal sensor Temp.: 98.88 Ambient Temp.: 96.42 F
17:48:38.367 -> Average Cal sensor Temp.: 98.90 Ambient Temp.: 96.49 F
17:48:39.922 -> Average Cal sensor Temp.: 98.99 Ambient Temp.: 96.42 F
17:48:41.471 -> Average Cal sensor Temp.: 99.09 Ambient Temp.: 96.42 F
17:48:43.024 -> Average Cal sensor Temp.: 98.57 Ambient Temp.: 96.39 F
17:48:44.580 -> Average Cal sensor Temp.: 98.63 Ambient Temp.: 96.42 F
17:48:46.137 -> Average Cal sensor Temp.: 98.94 Ambient Temp.: 96.49 F
17:48:47.694 -> Average Cal sensor Temp.: 99.04 Ambient Temp.: 96.49 F
17:48:49.252 -> Average Cal sensor Temp.: 99.06 Ambient Temp.: 96.42 F
17:48:50.837 -> Average Cal sensor Temp.: 98.67 Ambient Temp.: 96.42 F
17:48:52.392 -> Average Cal sensor Temp.: 99.17 Ambient Temp.: 96.42 F
17:48:53.944 -> Average Cal sensor Temp.: 99.10 Ambient Temp.: 96.42 F
17:48:55.498 -> Average Cal sensor Temp.: 99.39 Ambient Temp.: 96.53 F
17:48:57.055 -> Average Cal sensor Temp.: 99.13 Ambient Temp.: 96.42 F
17:48:58.607 -> Average Cal sensor Temp.: 99.32 Ambient Temp.: 96.42 F
17:49:00.159 -> Average Cal sensor Temp.: 98.89 Ambient Temp.: 96.49 F
17:49:01.708 -> Average Cal sensor Temp.: 99.02 Ambient Temp.: 96.49 F
17:49:03.292 -> Average Cal sensor Temp.: 99.12 Ambient Temp.: 96.49 F
17:49:04.847 -> Average Cal sensor Temp.: 98.79 Ambient Temp.: 96.42 F
17:49:06.403 -> Average Cal sensor Temp.: 99.14 Ambient Temp.: 96.49 F
17:49:07.958 -> Average Cal sensor Temp.: 99.03 Ambient Temp.: 96.42 F
17:49:09.520 -> Average Cal sensor Temp.: 98.99 Ambient Temp.: 96.49 F
17:49:11.049 -> Average Cal sensor Temp.: 99.10 Ambient Temp.: 96.42 F
17:49:12.635 -> Average Cal sensor Temp.: 98.96 Ambient Temp.: 96.49 F
17:49:14.189 -> Average Cal sensor Temp.: 98.84 Ambient Temp.: 96.49 F
17:49:15.742 -> Average Cal sensor Temp.: 99.01 Ambient Temp.: 96.49 F
17:49:17.295 -> Average Cal sensor Temp.: 99.03 Ambient Temp.: 96.49 F
17:49:18.851 -> Average Cal sensor Temp.: 99.11 Ambient Temp.: 96.49 F
17:49:20.401 -> Average Cal sensor Temp.: 99.27 Ambient Temp.: 96.49 F
17:49:21.956 -> Average Cal sensor Temp.: 99.27 Ambient Temp.: 96.53 F
17:49:23.511 -> Average Cal sensor Temp.: 99.17 Ambient Temp.: 96.42 F
17:49:25.094 -> Average Cal sensor Temp.: 99.23 Ambient Temp.: 96.49 F
17:49:26.651 -> Average Cal sensor Temp.: 99.24 Ambient Temp.: 96.42 F
17:49:28.202 -> Average Cal sensor Temp.: 99.47 Ambient Temp.: 96.49 F
17:49:29.759 -> Average Cal sensor Temp.: 99.54 Ambient Temp.: 96.53 F
17:49:31.310 -> Average Cal sensor Temp.: 99.49 Ambient Temp.: 96.49 F
17:49:32.858 -> Average Cal sensor Temp.: 99.32 Ambient Temp.: 96.53 F
17:49:34.416 -> Average Cal sensor Temp.: 99.26 Ambient Temp.: 96.53 F
17:49:35.977 -> Average Cal sensor Temp.: 99.22 Ambient Temp.: 96.49 F
17:49:37.529 -> Average Cal sensor Temp.: 99.50 Ambient Temp.: 96.53 F
17:49:39.122 -> Average Cal sensor Temp.: 98.98 Ambient Temp.: 96.53 F
17:49:40.667 -> Average Cal sensor Temp.: 99.14 Ambient Temp.: 96.49 F
17:49:42.216 -> Average Cal sensor Temp.: 99.27 Ambient Temp.: 96.49 F
17:49:43.766 -> Average Cal sensor Temp.: 98.89 Ambient Temp.: 96.53 F
17:49:45.327 -> Average Cal sensor Temp.: 99.21 Ambient Temp.: 96.53 F
17:49:46.881 -> Average Cal sensor Temp.: 99.36 Ambient Temp.: 96.42 F
17:49:48.436 -> Average Cal sensor Temp.: 99.23 Ambient Temp.: 96.53 F
17:49:50.020 -> Average Cal sensor Temp.: 99.16 Ambient Temp.: 96.53 F
17:49:51.569 -> Average Cal sensor Temp.: 99.42 Ambient Temp.: 96.57 F
17:49:53.124 -> Average Cal sensor Temp.: 99.21 Ambient Temp.: 96.53 F
17:49:54.677 -> Average Cal sensor Temp.: 99.23 Ambient Temp.: 96.53 F
17:49:56.238 -> Average Cal sensor Temp.: 99.48 Ambient Temp.: 96.49 F
17:49:57.796 -> Average Cal sensor Temp.: 99.29 Ambient Temp.: 96.49 F
17:49:59.352 -> Average Cal sensor Temp.: 99.31 Ambient Temp.: 96.53 F
17:50:00.905 -> Average Cal sensor Temp.: 99.45 Ambient Temp.: 96.49 F
17:50:02.463 -> Average Cal sensor Temp.: 99.25 Ambient Temp.: 96.57 F
17:50:04.020 -> Average Cal sensor Temp.: 99.35 Ambient Temp.: 96.53 F
17:50:05.579 -> Average Cal sensor Temp.: 99.31 Ambient Temp.: 96.57 F
17:50:07.128 -> Average Cal sensor Temp.: 99.42 Ambient Temp.: 96.57 F
17:50:08.679 -> Average Cal sensor Temp.: 99.58 Ambient Temp.: 96.57 F
17:50:10.231 -> Average Cal sensor Temp.: 99.66 Ambient Temp.: 96.57 F
17:50:11.818 -> Average Cal sensor Temp.: 99.25 Ambient Temp.: 96.57 F
17:50:13.371 -> Average Cal sensor Temp.: 99.28 Ambient Temp.: 96.57 F
17:50:14.929 -> Average Cal sensor Temp.: 98.39 Ambient Temp.: 96.53 F
17:50:16.486 -> Average Cal sensor Temp.: 99.39 Ambient Temp.: 96.53 F
17:50:18.039 -> Average Cal sensor Temp.: 99.72 Ambient Temp.: 96.57 F
17:50:19.588 -> Average Cal sensor Temp.: 99.38 Ambient Temp.: 96.53 F
17:50:21.141 -> Average Cal sensor Temp.: 99.52 Ambient Temp.: 96.53 F
17:50:22.728 -> Average Cal sensor Temp.: 99.50 Ambient Temp.: 96.53 F
17:50:24.277 -> Average Cal sensor Temp.: 99.39 Ambient Temp.: 96.53 F
17:50:25.821 -> Average Cal sensor Temp.: 99.72 Ambient Temp.: 96.57 F
17:50:27.379 -> Average Cal sensor Temp.: 99.28 Ambient Temp.: 96.53 F
17:50:28.937 -> Average Cal sensor Temp.: 99.06 Ambient Temp.: 96.53 F
17:50:30.490 -> Average Cal sensor Temp.: 99.34 Ambient Temp.: 96.57 F
17:50:32.046 -> Average Cal sensor Temp.: 99.13 Ambient Temp.: 96.53 F
17:50:33.629 -> Average Cal sensor Temp.: 99.32 Ambient Temp.: 96.53 F
17:50:35.176 -> Average Cal sensor Temp.: 99.53 Ambient Temp.: 96.53 F
17:50:36.736 -> Average Cal sensor Temp.: 99.37 Ambient Temp.: 96.53 F
17:50:38.291 -> Average Cal sensor Temp.: 99.64 Ambient Temp.: 96.53 F
17:50:39.856 -> Average Cal sensor Temp.: 99.42 Ambient Temp.: 96.53 F
17:50:41.408 -> Average Cal sensor Temp.: 99.59 Ambient Temp.: 96.53 F
17:50:42.960 -> Average Cal sensor Temp.: 99.51 Ambient Temp.: 96.57 F
17:50:44.545 -> Average Cal sensor Temp.: 99.54 Ambient Temp.: 96.53 F
17:50:46.093 -> Average Cal sensor Temp.: 99.60 Ambient Temp.: 96.53 F
17:50:47.648 -> Average Cal sensor Temp.: 99.51 Ambient Temp.: 96.57 F
17:50:49.208 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 96.49 F
17:50:50.795 -> Average Cal sensor Temp.: 99.91 Ambient Temp.: 96.57 F
17:50:52.344 -> Average Cal sensor Temp.: 99.82 Ambient Temp.: 96.53 F
17:50:53.898 -> Average Cal sensor Temp.: 99.55 Ambient Temp.: 96.53 F
17:50:55.450 -> Average Cal sensor Temp.: 99.32 Ambient Temp.: 96.53 F
17:50:57.037 -> Average Cal sensor Temp.: 99.13 Ambient Temp.: 96.49 F
17:50:58.588 -> Average Cal sensor Temp.: 99.57 Ambient Temp.: 96.57 F
17:51:00.143 -> Average Cal sensor Temp.: 99.56 Ambient Temp.: 96.53 F
17:51:01.701 -> Average Cal sensor Temp.: 99.14 Ambient Temp.: 96.53 F
17:51:03.255 -> Average Cal sensor Temp.: 99.29 Ambient Temp.: 96.53 F
17:51:04.808 -> Average Cal sensor Temp.: 99.47 Ambient Temp.: 96.57 F
17:51:06.385 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 96.53 F
17:51:07.930 -> Average Cal sensor Temp.: 99.29 Ambient Temp.: 96.57 F
17:51:09.479 -> Average Cal sensor Temp.: 99.23 Ambient Temp.: 96.53 F
17:51:11.029 -> Average Cal sensor Temp.: 99.31 Ambient Temp.: 96.53 F
17:51:12.617 -> Average Cal sensor Temp.: 99.54 Ambient Temp.: 96.49 F
17:51:14.170 -> Average Cal sensor Temp.: 98.93 Ambient Temp.: 96.57 F
17:51:15.725 -> Average Cal sensor Temp.: 99.47 Ambient Temp.: 96.53 F
17:51:17.277 -> Average Cal sensor Temp.: 99.36 Ambient Temp.: 96.53 F
17:51:18.832 -> Average Cal sensor Temp.: 99.36 Ambient Temp.: 96.53 F
17:51:20.381 -> Average Cal sensor Temp.: 99.20 Ambient Temp.: 96.53 F
17:51:21.941 -> Average Cal sensor Temp.: 99.21 Ambient Temp.: 96.53 F
17:51:23.490 -> Average Cal sensor Temp.: 99.36 Ambient Temp.: 96.53 F
17:51:25.044 -> Average Cal sensor Temp.: 99.55 Ambient Temp.: 96.53 F
17:51:26.627 -> Average Cal sensor Temp.: 99.52 Ambient Temp.: 96.53 F
17:51:28.181 -> Average Cal sensor Temp.: 99.58 Ambient Temp.: 96.53 F
17:51:29.735 -> Average Cal sensor Temp.: 99.47 Ambient Temp.: 96.53 F
17:51:31.314 -> Average Cal sensor Temp.: 99.60 Ambient Temp.: 96.53 F
17:51:32.863 -> Average Cal sensor Temp.: 99.36 Ambient Temp.: 96.57 F
17:51:34.418 -> Average Cal sensor Temp.: 99.07 Ambient Temp.: 96.57 F
17:51:35.972 -> Average Cal sensor Temp.: 99.37 Ambient Temp.: 96.57 F
17:51:37.524 -> Average Cal sensor Temp.: 99.27 Ambient Temp.: 96.53 F
17:51:39.083 -> Average Cal sensor Temp.: 98.90 Ambient Temp.: 96.53 F
17:51:40.636 -> Average Cal sensor Temp.: 98.93 Ambient Temp.: 96.57 F
17:51:42.196 -> Average Cal sensor Temp.: 99.23 Ambient Temp.: 96.53 F
17:51:43.747 -> Average Cal sensor Temp.: 99.00 Ambient Temp.: 96.53 F
17:51:45.304 -> Average Cal sensor Temp.: 99.59 Ambient Temp.: 96.49 F
17:51:46.891 -> Average Cal sensor Temp.: 99.39 Ambient Temp.: 96.53 F
17:51:48.442 -> Average Cal sensor Temp.: 99.52 Ambient Temp.: 96.53 F
17:51:49.995 -> Average Cal sensor Temp.: 99.40 Ambient Temp.: 96.57 F
17:51:51.572 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 96.53 F
17:51:53.123 -> Average Cal sensor Temp.: 99.50 Ambient Temp.: 96.53 F
17:51:54.673 -> Average Cal sensor Temp.: 99.38 Ambient Temp.: 96.49 F
17:51:56.222 -> Average Cal sensor Temp.: 99.40 Ambient Temp.: 96.53 F
17:51:57.781 -> Average Cal sensor Temp.: 99.55 Ambient Temp.: 96.53 F
17:51:59.364 -> Average Cal sensor Temp.: 99.34 Ambient Temp.: 96.53 F
17:52:00.913 -> Average Cal sensor Temp.: 99.59 Ambient Temp.: 96.53 F
17:52:02.464 -> Average Cal sensor Temp.: 99.58 Ambient Temp.: 96.57 F
17:52:04.042 -> Average Cal sensor Temp.: 99.41 Ambient Temp.: 96.57 F
17:52:05.595 -> Average Cal sensor Temp.: 99.34 Ambient Temp.: 96.57 F
17:52:07.149 -> Average Cal sensor Temp.: 99.47 Ambient Temp.: 96.57 F
17:52:08.704 -> Average Cal sensor Temp.: 99.39 Ambient Temp.: 96.57 F
17:52:10.260 -> Average Cal sensor Temp.: 99.40 Ambient Temp.: 96.57 F
17:52:11.819 -> Average Cal sensor Temp.: 99.75 Ambient Temp.: 96.57 F
17:52:13.376 -> Average Cal sensor Temp.: 99.09 Ambient Temp.: 96.60 F
17:52:14.927 -> Average Cal sensor Temp.: 99.44 Ambient Temp.: 96.57 F
17:52:16.485 -> Average Cal sensor Temp.: 99.45 Ambient Temp.: 96.57 F
17:52:18.038 -> Average Cal sensor Temp.: 99.75 Ambient Temp.: 96.60 F
17:52:19.602 -> Average Cal sensor Temp.: 99.85 Ambient Temp.: 96.57 F
17:52:21.163 -> Average Cal sensor Temp.: 99.77 Ambient Temp.: 96.57 F
17:52:22.751 -> Average Cal sensor Temp.: 99.68 Ambient Temp.: 96.57 F
17:52:24.305 -> Average Cal sensor Temp.: 100.00 Ambient Temp.: 96.60 F
17:52:25.862 -> Average Cal sensor Temp.: 99.96 Ambient Temp.: 96.60 F
17:52:27.441 -> Average Cal sensor Temp.: 99.41 Ambient Temp.: 96.60 F
17:52:28.997 -> Average Cal sensor Temp.: 99.43 Ambient Temp.: 96.64 F
17:52:30.549 -> Average Cal sensor Temp.: 99.27 Ambient Temp.: 96.60 F
17:52:32.098 -> Average Cal sensor Temp.: 99.26 Ambient Temp.: 96.60 F
17:52:33.654 -> Average Cal sensor Temp.: 99.58 Ambient Temp.: 96.60 F
17:52:35.239 -> Average Cal sensor Temp.: 99.56 Ambient Temp.: 96.60 F
17:52:36.794 -> Average Cal sensor Temp.: 99.82 Ambient Temp.: 96.64 F
17:52:38.353 -> Average Cal sensor Temp.: 99.86 Ambient Temp.: 96.60 F
17:52:39.906 -> Average Cal sensor Temp.: 99.06 Ambient Temp.: 96.60 F
17:52:41.460 -> Average Cal sensor Temp.: 99.58 Ambient Temp.: 96.57 F
17:52:43.044 -> Average Cal sensor Temp.: 99.53 Ambient Temp.: 96.60 F
17:52:44.597 -> Average Cal sensor Temp.: 99.47 Ambient Temp.: 96.60 F
17:52:46.158 -> Average Cal sensor Temp.: 99.41 Ambient Temp.: 96.57 F
17:52:47.712 -> Average Cal sensor Temp.: 99.30 Ambient Temp.: 96.60 F
17:52:49.267 -> Average Cal sensor Temp.: 99.42 Ambient Temp.: 96.57 F
17:52:50.821 -> Average Cal sensor Temp.: 99.45 Ambient Temp.: 96.53 F
17:52:52.373 -> Average Cal sensor Temp.: 99.48 Ambient Temp.: 96.57 F
17:52:53.925 -> Average Cal sensor Temp.: 99.58 Ambient Temp.: 96.57 F
17:52:55.520 -> Average Cal sensor Temp.: 99.84 Ambient Temp.: 96.57 F
17:52:57.075 -> Average Cal sensor Temp.: 99.89 Ambient Temp.: 96.64 F
17:52:58.629 -> Average Cal sensor Temp.: 99.87 Ambient Temp.: 96.64 F
17:53:00.182 -> Average Cal sensor Temp.: 99.96 Ambient Temp.: 96.57 F
17:53:01.766 -> Average Cal sensor Temp.: 99.74 Ambient Temp.: 96.57 F
17:53:03.317 -> Average Cal sensor Temp.: 99.50 Ambient Temp.: 96.60 F
17:53:04.871 -> Average Cal sensor Temp.: 99.35 Ambient Temp.: 96.57 F
17:53:06.421 -> Average Cal sensor Temp.: 99.22 Ambient Temp.: 96.57 F
17:53:07.973 -> Average Cal sensor Temp.: 99.32 Ambient Temp.: 96.60 F
17:53:09.531 -> Average Cal sensor Temp.: 99.53 Ambient Temp.: 96.60 F
17:53:11.114 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 96.57 F
17:53:12.669 -> Average Cal sensor Temp.: 99.62 Ambient Temp.: 96.60 F
17:53:14.225 -> Average Cal sensor Temp.: 99.67 Ambient Temp.: 96.57 F
17:53:15.812 -> Average Cal sensor Temp.: 99.21 Ambient Temp.: 96.60 F
17:53:17.335 -> Average Cal sensor Temp.: 99.74 Ambient Temp.: 96.60 F
17:53:18.926 -> Average Cal sensor Temp.: 99.83 Ambient Temp.: 96.57 F
17:53:20.486 -> Average Cal sensor Temp.: 99.81 Ambient Temp.: 96.57 F
17:53:22.040 -> Average Cal sensor Temp.: 99.72 Ambient Temp.: 96.60 F
17:53:23.596 -> Average Cal sensor Temp.: 99.11 Ambient Temp.: 96.57 F
17:53:25.142 -> Average Cal sensor Temp.: 98.77 Ambient Temp.: 96.57 F
17:53:26.701 -> Average Cal sensor Temp.: 99.10 Ambient Temp.: 96.53 F
17:53:28.259 -> Average Cal sensor Temp.: 99.38 Ambient Temp.: 96.49 F
17:53:29.818 -> Average Cal sensor Temp.: 99.55 Ambient Temp.: 96.53 F
17:53:31.375 -> Average Cal sensor Temp.: 99.78 Ambient Temp.: 96.53 F
17:53:32.956 -> Average Cal sensor Temp.: 100.01 Ambient Temp.: 96.57 F
17:53:34.509 -> Average Cal sensor Temp.: 99.51 Ambient Temp.: 96.53 F
17:53:36.065 -> Average Cal sensor Temp.: 99.56 Ambient Temp.: 96.57 F
17:53:37.621 -> Average Cal sensor Temp.: 99.28 Ambient Temp.: 96.49 F
17:53:39.208 -> Average Cal sensor Temp.: 99.52 Ambient Temp.: 96.57 F
17:53:40.762 -> Average Cal sensor Temp.: 99.59 Ambient Temp.: 96.53 F
17:53:42.320 -> Average Cal sensor Temp.: 99.54 Ambient Temp.: 96.57 F
17:53:43.872 -> Average Cal sensor Temp.: 99.58 Ambient Temp.: 96.57 F
17:53:45.430 -> Average Cal sensor Temp.: 99.79 Ambient Temp.: 96.57 F
17:53:47.018 -> Average Cal sensor Temp.: 99.86 Ambient Temp.: 96.57 F
17:53:48.570 -> Average Cal sensor Temp.: 99.65 Ambient Temp.: 96.57 F
17:53:50.122 -> Average Cal sensor Temp.: 99.82 Ambient Temp.: 96.53 F
17:53:51.677 -> Average Cal sensor Temp.: 99.34 Ambient Temp.: 96.60 F
17:53:53.265 -> Average Cal sensor Temp.: 99.61 Ambient Temp.: 96.57 F
17:53:54.818 -> Average Cal sensor Temp.: 99.87 Ambient Temp.: 96.57 F
17:53:56.372 -> Average Cal sensor Temp.: 99.65 Ambient Temp.: 96.53 F
17:53:57.923 -> Average Cal sensor Temp.: 99.05 Ambient Temp.: 96.57 F
17:53:59.478 -> Average Cal sensor Temp.: 99.60 Ambient Temp.: 96.53 F
17:54:01.063 -> Average Cal sensor Temp.: 99.78 Ambient Temp.: 96.57 F
17:54:02.617 -> Average Cal sensor Temp.: 99.99 Ambient Temp.: 96.53 F
17:54:04.165 -> Average Cal sensor Temp.: 99.46 Ambient Temp.: 96.57 F
17:54:05.753 -> Average Cal sensor Temp.: 99.56 Ambient Temp.: 96.57 F
17:54:07.305 -> Average Cal sensor Temp.: 99.73 Ambient Temp.: 96.57 F
17:54:08.856 -> Average Cal sensor Temp.: 99.78 Ambient Temp.: 96.57 F
17:54:10.415 -> Average Cal sensor Temp.: 99.91 Ambient Temp.: 96.57 F
17:54:12.000 -> Average Cal sensor Temp.: 100.09 Ambient Temp.: 96.57 F
17:54:13.548 -> Average Cal sensor Temp.: 99.81 Ambient Temp.: 96.57 F
17:54:15.099 -> Average Cal sensor Temp.: 100.12 Ambient Temp.: 96.57 F
17:54:16.693 -> Average Cal sensor Temp.: 100.14 Ambient Temp.: 96.57 F
17:54:18.247 -> Average Cal sensor Temp.: 99.85 Ambient Temp.: 96.57 F
17:54:19.799 -> Average Cal sensor Temp.: 100.09 Ambient Temp.: 96.57 F
17:54:21.353 -> Average Cal sensor Temp.: 100.02 Ambient Temp.: 96.64 F
17:54:22.937 -> Average Cal sensor Temp.: 100.10 Ambient Temp.: 96.60 F
17:54:24.492 -> Average Cal sensor Temp.: 100.11 Ambient Temp.: 96.60 F
17:54:26.072 -> Average Cal sensor Temp.: 100.26 Ambient Temp.: 96.64 F
17:54:27.633 -> Average Cal sensor Temp.: 99.39 Ambient Temp.: 96.60 F
17:54:29.179 -> Average Cal sensor Temp.: 99.68 Ambient Temp.: 96.53 F
17:54:30.752 -> Average Cal sensor Temp.: 99.87 Ambient Temp.: 96.60 F
17:54:32.304 -> Average Cal sensor Temp.: 99.25 Ambient Temp.: 96.60 F
17:54:33.865 -> Average Cal sensor Temp.: 99.56 Ambient Temp.: 96.53 F
17:54:35.421 -> Average Cal sensor Temp.: 99.69 Ambient Temp.: 96.60 F
17:54:37.011 -> Average Cal sensor Temp.: 98.88 Ambient Temp.: 96.57 F
17:54:38.566 -> Average Cal sensor Temp.: 98.72 Ambient Temp.: 96.57 F
17:54:40.124 -> Average Cal sensor Temp.: 99.45 Ambient Temp.: 96.57 F
17:54:41.680 -> Average Cal sensor Temp.: 99.50 Ambient Temp.: 96.53 F
17:54:43.235 -> Average Cal sensor Temp.: 99.56 Ambient Temp.: 96.57 F
17:54:44.790 -> Average Cal sensor Temp.: 99.71 Ambient Temp.: 96.57 F
17:54:46.343 -> Average Cal sensor Temp.: 99.08 Ambient Temp.: 96.53 F
17:54:47.894 -> Average Cal sensor Temp.: 99.45 Ambient Temp.: 96.53 F
17:54:49.483 -> Average Cal sensor Temp.: 98.82 Ambient Temp.: 96.53 F
17:54:51.039 -> Average Cal sensor Temp.: 99.52 Ambient Temp.: 96.53 F
17:54:52.592 -> Average Cal sensor Temp.: 99.46 Ambient Temp.: 96.53 F
17:54:54.144 -> Average Cal sensor Temp.: 98.95 Ambient Temp.: 96.49 F
17:54:55.693 -> Average Cal sensor Temp.: 99.47 Ambient Temp.: 96.49 F
17:54:57.247 -> Average Cal sensor Temp.: 100.17 Ambient Temp.: 96.49 F
17:54:58.833 -> Average Cal sensor Temp.: 99.98 Ambient Temp.: 96.53 F
17:55:00.384 -> Average Cal sensor Temp.: 99.86 Ambient Temp.: 96.42 F
17:55:01.943 -> Average Cal sensor Temp.: 99.08 Ambient Temp.: 96.49 F
17:55:03.494 -> Average Cal sensor Temp.: 99.34 Ambient Temp.: 96.49 F
17:55:05.082 -> Average Cal sensor Temp.: 99.75 Ambient Temp.: 96.42 F
17:55:06.630 -> Average Cal sensor Temp.: 99.88 Ambient Temp.: 96.39 F
17:55:08.187 -> Average Cal sensor Temp.: 99.82 Ambient Temp.: 96.49 F
17:55:09.742 -> Average Cal sensor Temp.: 99.65 Ambient Temp.: 96.49 F
17:55:11.331 -> Average Cal sensor Temp.: 99.36 Ambient Temp.: 96.42 F
17:55:12.888 -> Average Cal sensor Temp.: 99.57 Ambient Temp.: 96.42 F
17:55:14.441 -> Average Cal sensor Temp.: 99.78 Ambient Temp.: 96.42 F
17:55:15.993 -> Average Cal sensor Temp.: 99.84 Ambient Temp.: 96.49 F
17:55:17.552 -> Average Cal sensor Temp.: 100.24 Ambient Temp.: 96.49 F
17:55:19.138 -> Average Cal sensor Temp.: 100.22 Ambient Temp.: 96.53 F
17:55:20.693 -> Average Cal sensor Temp.: 100.23 Ambient Temp.: 96.49 F
17:55:22.271 -> Average Cal sensor Temp.: 100.08 Ambient Temp.: 96.49 F
17:55:23.820 -> Average Cal sensor Temp.: 99.73 Ambient Temp.: 96.42 F
17:55:25.382 -> Average Cal sensor Temp.: 100.20 Ambient Temp.: 96.49 F
17:55:26.972 -> Average Cal sensor Temp.: 99.84 Ambient Temp.: 96.53 F
17:55:28.530 -> Average Cal sensor Temp.: 99.90 Ambient Temp.: 96.49 F
17:55:30.088 -> Average Cal sensor Temp.: 100.11 Ambient Temp.: 96.57 F
17:55:31.637 -> Average Cal sensor Temp.: 100.07 Ambient Temp.: 96.49 F
17:55:33.225 -> Average Cal sensor Temp.: 99.91 Ambient Temp.: 96.53 F
17:55:34.778 -> Average Cal sensor Temp.: 100.03 Ambient Temp.: 96.53 F
17:55:36.333 -> Average Cal sensor Temp.: 100.11 Ambient Temp.: 96.53 F
17:55:37.890 -> Average Cal sensor Temp.: 100.11 Ambient Temp.: 96.53 F
17:55:39.480 -> Average Cal sensor Temp.: 100.03 Ambient Temp.: 96.53 F
17:55:41.031 -> Average Cal sensor Temp.: 100.29 Ambient Temp.: 96.49 F
17:55:42.612 -> Average Cal sensor Temp.: 100.40 Ambient Temp.: 96.57 F
17:55:44.162 -> Average Cal sensor Temp.: 100.10 Ambient Temp.: 96.53 F
17:55:45.720 -> Average Cal sensor Temp.: 100.17 Ambient Temp.: 96.53 F
17:55:47.312 -> Average Cal sensor Temp.: 100.06 Ambient Temp.: 96.57 F
17:55:48.867 -> Average Cal sensor Temp.: 100.01 Ambient Temp.: 96.57 F
17:55:50.416 -> Average Cal sensor Temp.: 99.98 Ambient Temp.: 96.53 F
17:55:52.006 -> Average Cal sensor Temp.: 100.47 Ambient Temp.: 96.53 F
17:55:53.556 -> Average Cal sensor Temp.: 100.59 Ambient Temp.: 96.57 F
17:55:55.135 -> Average Cal sensor Temp.: 100.45 Ambient Temp.: 96.60 F
17:55:56.691 -> Average Cal sensor Temp.: 100.06 Ambient Temp.: 96.57 F
17:55:58.240 -> Average Cal sensor Temp.: 100.13 Ambient Temp.: 96.60 F
17:55:59.829 -> Average Cal sensor Temp.: 99.99 Ambient Temp.: 96.60 F
17:56:01.383 -> Average Cal sensor Temp.: 99.82 Ambient Temp.: 96.60 F
17:56:02.935 -> Average Cal sensor Temp.: 100.01 Ambient Temp.: 96.60 F
17:56:04.527 -> Average Cal sensor Temp.: 100.05 Ambient Temp.: 96.60 F
17:56:06.083 -> Average Cal sensor Temp.: 100.38 Ambient Temp.: 96.60 F
17:56:07.635 -> Average Cal sensor Temp.: 100.41 Ambient Temp.: 96.53 F
17:56:09.199 -> Average Cal sensor Temp.: 100.10 Ambient Temp.: 99.95 F
17:56:10.790 -> Average Cal sensor Temp.: 99.77 Ambient Temp.: 96.60 F
17:56:12.345 -> Average Cal sensor Temp.: 99.50 Ambient Temp.: 96.64 F
17:56:13.896 -> Average Cal sensor Temp.: 100.04 Ambient Temp.: 96.60 F
17:56:15.458 -> Average Cal sensor Temp.: 100.40 Ambient Temp.: 96.64 F
17:56:17.044 -> Average Cal sensor Temp.: 99.79 Ambient Temp.: 96.60 F
17:56:18.600 -> Average Cal sensor Temp.: 99.99 Ambient Temp.: 96.67 F
17:56:20.160 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 96.60 F
17:56:21.714 -> Average Cal sensor Temp.: 99.81 Ambient Temp.: 96.64 F
17:56:23.297 -> Average Cal sensor Temp.: 100.06 Ambient Temp.: 96.60 F
17:56:24.852 -> Average Cal sensor Temp.: 100.35 Ambient Temp.: 96.64 F
17:56:26.414 -> Average Cal sensor Temp.: 99.65 Ambient Temp.: 96.64 F
17:56:27.997 -> Average Cal sensor Temp.: 99.90 Ambient Temp.: 96.60 F
17:56:29.555 -> Average Cal sensor Temp.: 99.90 Ambient Temp.: 96.60 F
17:56:31.105 -> Average Cal sensor Temp.: 100.42 Ambient Temp.: 96.64 F
17:56:32.695 -> Average Cal sensor Temp.: 100.23 Ambient Temp.: 96.64 F
17:56:34.250 -> Average Cal sensor Temp.: 100.23 Ambient Temp.: 96.60 F
17:56:35.813 -> Average Cal sensor Temp.: 100.07 Ambient Temp.: 96.60 F
17:56:37.364 -> Average Cal sensor Temp.: 100.00 Ambient Temp.: 96.64 F
17:56:38.957 -> Average Cal sensor Temp.: 99.66 Ambient Temp.: 96.64 F
17:56:40.518 -> Average Cal sensor Temp.: 100.13 Ambient Temp.: 96.64 F
17:56:42.070 -> Average Cal sensor Temp.: 99.92 Ambient Temp.: 96.64 F
17:56:43.628 -> Average Cal sensor Temp.: 100.11 Ambient Temp.: 96.64 F
17:56:45.218 -> Average Cal sensor Temp.: 100.39 Ambient Temp.: 96.64 F
17:56:46.776 -> Average Cal sensor Temp.: 100.27 Ambient Temp.: 96.64 F
17:56:48.325 -> Average Cal sensor Temp.: 99.78 Ambient Temp.: 96.64 F
17:56:49.914 -> Average Cal sensor Temp.: 100.15 Ambient Temp.: 96.67 F
17:56:51.469 -> Average Cal sensor Temp.: 99.95 Ambient Temp.: 96.67 F
17:56:53.024 -> Average Cal sensor Temp.: 99.49 Ambient Temp.: 96.64 F
17:56:54.578 -> Average Cal sensor Temp.: 99.58 Ambient Temp.: 96.64 F
17:56:56.140 -> Average Cal sensor Temp.: 99.81 Ambient Temp.: 96.64 F
17:56:57.696 -> Average Cal sensor Temp.: 99.43 Ambient Temp.: 96.64 F
17:56:59.253 -> Average Cal sensor Temp.: 99.40 Ambient Temp.: 96.64 F
17:57:00.842 -> Average Cal sensor Temp.: 99.69 Ambient Temp.: 96.64 F
17:57:02.394 -> Average Cal sensor Temp.: 99.76 Ambient Temp.: 96.60 F
17:57:03.953 -> Average Cal sensor Temp.: 100.10 Ambient Temp.: 96.64 F
17:57:05.513 -> Average Cal sensor Temp.: 100.03 Ambient Temp.: 96.60 F
17:57:07.071 -> Average Cal sensor Temp.: 99.81 Ambient Temp.: 96.64 F
17:57:08.660 -> Average Cal sensor Temp.: 99.93 Ambient Temp.: 96.64 F
17:57:10.220 -> Average Cal sensor Temp.: 100.22 Ambient Temp.: 96.60 F
17:57:11.777 -> Average Cal sensor Temp.: 100.29 Ambient Temp.: 96.64 F
17:57:13.361 -> Average Cal sensor Temp.: 100.18 Ambient Temp.: 96.64 F
17:57:14.912 -> Average Cal sensor Temp.: 100.31 Ambient Temp.: 96.64 F
17:57:16.462 -> Average Cal sensor Temp.: 100.22 Ambient Temp.: 96.64 F
17:57:18.051 -> Average Cal sensor Temp.: 100.49 Ambient Temp.: 96.67 F
17:57:19.610 -> Average Cal sensor Temp.: 100.47 Ambient Temp.: 96.60 F
17:57:21.161 -> Average Cal sensor Temp.: 100.35 Ambient Temp.: 96.64 F
17:57:22.752 -> Average Cal sensor Temp.: 100.37 Ambient Temp.: 96.67 F
17:57:24.308 -> Average Cal sensor Temp.: 100.40 Ambient Temp.: 96.67 F
17:57:25.861 -> Average Cal sensor Temp.: 100.41 Ambient Temp.: 96.64 F
17:57:27.448 -> Average Cal sensor Temp.: 99.87 Ambient Temp.: 96.67 F
17:57:29.003 -> Average Cal sensor Temp.: 100.02 Ambient Temp.: 96.64 F
17:57:30.586 -> Average Cal sensor Temp.: 100.24 Ambient Temp.: 96.67 F
17:57:32.143 -> Average Cal sensor Temp.: 100.61 Ambient Temp.: 96.67 F
17:57:33.699 -> Average Cal sensor Temp.: 100.22 Ambient Temp.: 96.64 F
17:57:35.261 -> Average Cal sensor Temp.: 100.41 Ambient Temp.: 96.67 F
17:57:36.849 -> Average Cal sensor Temp.: 100.48 Ambient Temp.: 96.67 F
17:57:38.405 -> Average Cal sensor Temp.: 100.76 Ambient Temp.: 96.67 F
17:57:39.961 -> Average Cal sensor Temp.: 100.88 Ambient Temp.: 96.67 F
17:57:41.521 -> Average Cal sensor Temp.: 100.80 Ambient Temp.: 96.64 F
17:57:43.113 -> Average Cal sensor Temp.: 100.64 Ambient Temp.: 96.67 F
17:57:44.667 -> Average Cal sensor Temp.: 100.72 Ambient Temp.: 96.67 F
17:57:46.224 -> Average Cal sensor Temp.: 100.84 Ambient Temp.: 96.64 F
17:57:47.816 -> Average Cal sensor Temp.: 99.59 Ambient Temp.: 96.67 F
17:57:49.372 -> Average Cal sensor Temp.: 99.86 Ambient Temp.: 96.67 F
17:57:50.922 -> Average Cal sensor Temp.: 99.88 Ambient Temp.: 96.71 F
17:57:52.482 -> Average Cal sensor Temp.: 100.32 Ambient Temp.: 96.71 F
17:57:54.069 -> Average Cal sensor Temp.: 100.14 Ambient Temp.: 96.67 F
17:57:55.626 -> Average Cal sensor Temp.: 100.31 Ambient Temp.: 96.64 F
17:57:57.184 -> Average Cal sensor Temp.: 100.55 Ambient Temp.: 96.67 F
17:57:58.740 -> Average Cal sensor Temp.: 100.75 Ambient Temp.: 96.71 F
17:58:00.300 -> Average Cal sensor Temp.: 100.15 Ambient Temp.: 96.67 F
17:58:01.889 -> Average Cal sensor Temp.: 100.25 Ambient Temp.: 96.78 F
17:58:03.441 -> Average Cal sensor Temp.: 100.45 Ambient Temp.: 96.71 F
17:58:04.999 -> Average Cal sensor Temp.: 100.23 Ambient Temp.: 96.67 F
17:58:06.558 -> Average Cal sensor Temp.: 100.44 Ambient Temp.: 96.71 F
17:58:08.147 -> Average Cal sensor Temp.: 100.53 Ambient Temp.: 96.75 F
17:58:09.707 -> Average Cal sensor Temp.: 100.73 Ambient Temp.: 96.71 F
17:58:11.256 -> Average Cal sensor Temp.: 100.56 Ambient Temp.: 96.71 F
17:58:12.846 -> Average Cal sensor Temp.: 100.37 Ambient Temp.: 96.71 F
17:58:14.398 -> Average Cal sensor Temp.: 100.61 Ambient Temp.: 96.75 F
17:58:15.951 -> Average Cal sensor Temp.: 100.78 Ambient Temp.: 96.67 F
17:58:17.534 -> Average Cal sensor Temp.: 100.85 Ambient Temp.: 96.71 F
17:58:19.089 -> Average Cal sensor Temp.: 100.38 Ambient Temp.: 96.75 F
17:58:20.645 -> Average Cal sensor Temp.: 100.39 Ambient Temp.: 96.71 F
17:58:22.229 -> Average Cal sensor Temp.: 100.46 Ambient Temp.: 96.75 F
17:58:23.785 -> Average Cal sensor Temp.: 100.67 Ambient Temp.: 96.67 F
17:58:25.339 -> Average Cal sensor Temp.: 100.72 Ambient Temp.: 96.75 F
17:58:26.931 -> Average Cal sensor Temp.: 100.80 Ambient Temp.: 96.71 F
17:58:28.480 -> Average Cal sensor Temp.: 100.84 Ambient Temp.: 96.75 F
17:58:30.039 -> Average Cal sensor Temp.: 100.79 Ambient Temp.: 96.75 F
17:58:31.600 -> Average Cal sensor Temp.: 100.87 Ambient Temp.: 96.75 F
17:58:33.190 -> Average Cal sensor Temp.: 100.83 Ambient Temp.: 96.75 F
17:58:34.748 -> Average Cal sensor Temp.: 100.18 Ambient Temp.: 96.71 F
17:58:36.313 -> Average Cal sensor Temp.: 100.00 Ambient Temp.: 96.78 F
17:58:37.872 -> Average Cal sensor Temp.: 99.59 Ambient Temp.: 96.71 F
17:58:39.433 -> Average Cal sensor Temp.: 99.76 Ambient Temp.: 96.75 F
17:58:40.991 -> Average Cal sensor Temp.: 100.14 Ambient Temp.: 96.71 F
17:58:42.582 -> Average Cal sensor Temp.: 99.74 Ambient Temp.: 96.75 F
17:58:44.135 -> Average Cal sensor Temp.: 99.97 Ambient Temp.: 96.75 F
17:58:45.691 -> Average Cal sensor Temp.: 100.52 Ambient Temp.: 96.78 F
17:58:47.249 -> Average Cal sensor Temp.: 100.60 Ambient Temp.: 96.75 F
17:58:48.837 -> Average Cal sensor Temp.: 100.36 Ambient Temp.: 96.71 F
17:58:50.391 -> Average Cal sensor Temp.: 100.39 Ambient Temp.: 96.78 F
17:58:51.943 -> Average Cal sensor Temp.: 100.49 Ambient Temp.: 96.75 F
17:58:53.528 -> Average Cal sensor Temp.: 100.52 Ambient Temp.: 96.71 F
17:58:55.079 -> Average Cal sensor Temp.: 100.52 Ambient Temp.: 96.75 F
17:58:56.665 -> Average Cal sensor Temp.: 100.47 Ambient Temp.: 96.78 F
17:58:58.214 -> Average Cal sensor Temp.: 100.76 Ambient Temp.: 96.75 F
17:58:59.764 -> Average Cal sensor Temp.: 100.42 Ambient Temp.: 96.82 F
17:59:01.352 -> Average Cal sensor Temp.: 100.61 Ambient Temp.: 96.75 F
17:59:02.905 -> Average Cal sensor Temp.: 100.39 Ambient Temp.: 96.75 F
17:59:04.490 -> Average Cal sensor Temp.: 100.29 Ambient Temp.: 96.78 F
17:59:06.042 -> Average Cal sensor Temp.: 100.56 Ambient Temp.: 96.71 F
17:59:07.598 -> Average Cal sensor Temp.: 100.34 Ambient Temp.: 96.75 F
17:59:09.187 -> Average Cal sensor Temp.: 99.86 Ambient Temp.: 96.71 F
17:59:10.743 -> Average Cal sensor Temp.: 100.52 Ambient Temp.: 96.71 F
17:59:12.296 -> Average Cal sensor Temp.: 100.35 Ambient Temp.: 96.71 F
17:59:13.880 -> Average Cal sensor Temp.: 100.53 Ambient Temp.: 96.75 F
17:59:15.435 -> Average Cal sensor Temp.: 100.38 Ambient Temp.: 96.78 F
17:59:16.982 -> Average Cal sensor Temp.: 100.49 Ambient Temp.: 96.75 F
17:59:18.573 -> Average Cal sensor Temp.: 100.58 Ambient Temp.: 96.78 F
17:59:20.131 -> Average Cal sensor Temp.: 100.50 Ambient Temp.: 96.75 F
17:59:21.680 -> Average Cal sensor Temp.: 100.67 Ambient Temp.: 96.75 F
17:59:23.273 -> Average Cal sensor Temp.: 100.54 Ambient Temp.: 96.75 F
17:59:24.829 -> Average Cal sensor Temp.: 100.54 Ambient Temp.: 96.75 F
17:59:26.384 -> Average Cal sensor Temp.: 100.57 Ambient Temp.: 96.75 F
17:59:27.965 -> Average Cal sensor Temp.: 100.53 Ambient Temp.: 96.78 F
17:59:29.524 -> Average Cal sensor Temp.: 100.58 Ambient Temp.: 96.89 F
17:59:31.072 -> Average Cal sensor Temp.: 100.68 Ambient Temp.: 96.82 F
17:59:32.661 -> Average Cal sensor Temp.: 100.71 Ambient Temp.: 96.78 F
17:59:34.207 -> Average Cal sensor Temp.: 100.67 Ambient Temp.: 96.78 F
17:59:35.790 -> Average Cal sensor Temp.: 100.73 Ambient Temp.: 96.75 F
17:59:37.343 -> Average Cal sensor Temp.: 100.28 Ambient Temp.: 96.78 F
17:59:38.928 -> Average Cal sensor Temp.: 100.31 Ambient Temp.: 96.82 F
17:59:40.484 -> Average Cal sensor Temp.: 100.47 Ambient Temp.: 96.75 F
17:59:42.042 -> Average Cal sensor Temp.: 100.15 Ambient Temp.: 96.78 F
17:59:43.598 -> Average Cal sensor Temp.: 100.21 Ambient Temp.: 96.78 F
17:59:45.184 -> Average Cal sensor Temp.: 100.15 Ambient Temp.: 96.82 F
17:59:46.735 -> Average Cal sensor Temp.: 100.31 Ambient Temp.: 96.78 F
17:59:48.287 -> Average Cal sensor Temp.: 100.04 Ambient Temp.: 96.78 F
17:59:49.870 -> Average Cal sensor Temp.: 100.03 Ambient Temp.: 96.78 F
17:59:51.428 -> Average Cal sensor Temp.: 100.40 Ambient Temp.: 96.78 F
17:59:53.015 -> Average Cal sensor Temp.: 100.14 Ambient Temp.: 96.75 F
17:59:54.569 -> Average Cal sensor Temp.: 100.13 Ambient Temp.: 96.78 F
17:59:56.120 -> Average Cal sensor Temp.: 100.15 Ambient Temp.: 96.75 F
17:59:57.709 -> Average Cal sensor Temp.: 100.19 Ambient Temp.: 96.75 F
17:59:59.258 -> Average Cal sensor Temp.: 99.85 Ambient Temp.: 96.82 F
18:00:00.815 -> Average Cal sensor Temp.: 100.18 Ambient Temp.: 96.75 F
18:00:02.401 -> Average Cal sensor Temp.: 100.67 Ambient Temp.: 96.78 F
18:00:03.952 -> Average Cal sensor Temp.: 100.61 Ambient Temp.: 96.82 F
18:00:05.506 -> Average Cal sensor Temp.: 100.82 Ambient Temp.: 96.78 F
18:00:07.097 -> Average Cal sensor Temp.: 100.66 Ambient Temp.: 96.78 F
18:00:08.644 -> Average Cal sensor Temp.: 100.33 Ambient Temp.: 96.78 F
18:00:10.237 -> Average Cal sensor Temp.: 100.78 Ambient Temp.: 96.82 F
18:00:11.789 -> Average Cal sensor Temp.: 100.81 Ambient Temp.: 96.78 F
18:00:13.338 -> Average Cal sensor Temp.: 100.54 Ambient Temp.: 96.78 F
18:00:14.922 -> Average Cal sensor Temp.: 100.63 Ambient Temp.: 96.78 F
18:00:16.476 -> Average Cal sensor Temp.: 100.23 Ambient Temp.: 96.78 F
18:00:18.058 -> Average Cal sensor Temp.: 100.09 Ambient Temp.: 96.78 F
18:00:19.619 -> Average Cal sensor Temp.: 100.39 Ambient Temp.: 96.82 F
18:00:21.173 -> Average Cal sensor Temp.: 101.01 Ambient Temp.: 96.78 F
18:00:22.757 -> Average Cal sensor Temp.: 100.85 Ambient Temp.: 96.82 F
18:00:24.310 -> Average Cal sensor Temp.: 100.83 Ambient Temp.: 96.78 F
18:00:25.866 -> Average Cal sensor Temp.: 100.30 Ambient Temp.: 96.82 F
18:00:27.452 -> Average Cal sensor Temp.: 100.03 Ambient Temp.: 96.82 F
18:00:29.005 -> Average Cal sensor Temp.: 100.16 Ambient Temp.: 96.82 F
18:00:30.591 -> Average Cal sensor Temp.: 100.16 Ambient Temp.: 96.82 F
18:00:32.142 -> Average Cal sensor Temp.: 100.63 Ambient Temp.: 96.82 F
18:00:33.726 -> Average Cal sensor Temp.: 100.06 Ambient Temp.: 96.82 F
18:00:35.277 -> Average Cal sensor Temp.: 99.95 Ambient Temp.: 96.78 F
18:00:36.831 -> Average Cal sensor Temp.: 99.49 Ambient Temp.: 96.75 F
18:00:38.383 -> Average Cal sensor Temp.: 99.78 Ambient Temp.: 96.71 F
18:00:39.967 -> Average Cal sensor Temp.: 100.23 Ambient Temp.: 96.78 F
18:00:41.521 -> Average Cal sensor Temp.: 100.46 Ambient Temp.: 96.82 F
18:00:43.077 -> Average Cal sensor Temp.: 100.63 Ambient Temp.: 96.75 F
18:00:44.669 -> Average Cal sensor Temp.: 100.21 Ambient Temp.: 96.71 F
18:00:46.235 -> Average Cal sensor Temp.: 100.32 Ambient Temp.: 96.71 F
18:00:47.786 -> Average Cal sensor Temp.: 99.97 Ambient Temp.: 96.75 F
18:00:49.346 -> Average Cal sensor Temp.: 100.00 Ambient Temp.: 96.67 F
18:00:50.901 -> Average Cal sensor Temp.: 100.15 Ambient Temp.: 96.67 F
18:00:52.469 -> Average Cal sensor Temp.: 100.39 Ambient Temp.: 96.71 F
18:00:54.055 -> Average Cal sensor Temp.: 100.50 Ambient Temp.: 96.71 F
18:00:55.611 -> Average Cal sensor Temp.: 100.59 Ambient Temp.: 96.67 F
18:00:57.172 -> Average Cal sensor Temp.: 100.28 Ambient Temp.: 96.71 F
18:00:58.725 -> Average Cal sensor Temp.: 100.01 Ambient Temp.: 96.67 F
18:01:00.309 -> Average Cal sensor Temp.: 100.68 Ambient Temp.: 96.75 F
18:01:01.862 -> Average Cal sensor Temp.: 100.69 Ambient Temp.: 96.75 F
18:01:03.443 -> Average Cal sensor Temp.: 100.44 Ambient Temp.: 96.64 F
18:01:05.004 -> Average Cal sensor Temp.: 100.41 Ambient Temp.: 96.67 F
18:01:06.584 -> Average Cal sensor Temp.: 100.64 Ambient Temp.: 96.71 F
18:01:08.145 -> Average Cal sensor Temp.: 100.71 Ambient Temp.: 96.67 F
18:01:09.692 -> Average Cal sensor Temp.: 100.60 Ambient Temp.: 96.67 F
18:01:11.278 -> Average Cal sensor Temp.: 100.86 Ambient Temp.: 96.67 F
18:01:12.833 -> Average Cal sensor Temp.: 100.93 Ambient Temp.: 96.71 F
18:01:14.385 -> Average Cal sensor Temp.: 100.83 Ambient Temp.: 96.64 F
18:01:15.973 -> Average Cal sensor Temp.: 100.73 Ambient Temp.: 96.71 F
18:01:17.524 -> Average Cal sensor Temp.: 100.54 Ambient Temp.: 96.67 F
18:01:19.114 -> Average Cal sensor Temp.: 99.80 Ambient Temp.: 96.64 F
18:01:20.664 -> Average Cal sensor Temp.: 99.48 Ambient Temp.: 96.67 F
18:01:22.218 -> Average Cal sensor Temp.: 100.03 Ambient Temp.: 96.64 F
18:01:23.772 -> Average Cal sensor Temp.: 100.24 Ambient Temp.: 96.64 F
18:01:25.365 -> Average Cal sensor Temp.: 100.55 Ambient Temp.: 96.64 F
18:01:26.916 -> Average Cal sensor Temp.: 100.89 Ambient Temp.: 96.64 F
18:01:28.470 -> Average Cal sensor Temp.: 100.43 Ambient Temp.: 96.71 F
18:01:30.060 -> Average Cal sensor Temp.: 100.47 Ambient Temp.: 96.67 F
18:01:31.612 -> Average Cal sensor Temp.: 100.25 Ambient Temp.: 96.64 F
18:01:33.166 -> Average Cal sensor Temp.: 100.25 Ambient Temp.: 96.64 F
18:01:34.748 -> Average Cal sensor Temp.: 100.48 Ambient Temp.: 96.64 F
18:01:36.300 -> Average Cal sensor Temp.: 100.77 Ambient Temp.: 96.67 F
18:01:37.886 -> Average Cal sensor Temp.: 100.36 Ambient Temp.: 96.64 F
18:01:39.442 -> Average Cal sensor Temp.: 100.31 Ambient Temp.: 96.64 F
18:01:40.996 -> Average Cal sensor Temp.: 100.36 Ambient Temp.: 96.64 F
18:01:42.586 -> Average Cal sensor Temp.: 100.39 Ambient Temp.: 96.71 F
18:01:44.133 -> Average Cal sensor Temp.: 100.35 Ambient Temp.: 96.67 F
18:01:45.722 -> Average Cal sensor Temp.: 100.40 Ambient Temp.: 96.64 F
18:01:47.276 -> Average Cal sensor Temp.: 100.38 Ambient Temp.: 96.64 F
18:01:48.831 -> Average Cal sensor Temp.: 100.04 Ambient Temp.: 96.64 F
18:01:50.412 -> Average Cal sensor Temp.: 100.30 Ambient Temp.: 96.64 F
18:01:51.965 -> Average Cal sensor Temp.: 100.38 Ambient Temp.: 96.60 F
18:01:53.555 -> Average Cal sensor Temp.: 100.38 Ambient Temp.: 96.64 F
18:01:55.107 -> Average Cal sensor Temp.: 100.23 Ambient Temp.: 96.64 F
18:01:56.663 -> Average Cal sensor Temp.: 100.34 Ambient Temp.: 96.67 F
18:01:58.254 -> Average Cal sensor Temp.: 99.90 Ambient Temp.: 96.64 F
18:01:59.802 -> Average Cal sensor Temp.: 100.35 Ambient Temp.: 96.60 F
18:02:01.357 -> Average Cal sensor Temp.: 100.67 Ambient Temp.: 96.60 F
18:02:02.944 -> Average Cal sensor Temp.: 100.44 Ambient Temp.: 96.64 F
18:02:04.494 -> Average Cal sensor Temp.: 100.45 Ambient Temp.: 96.60 F
18:02:06.083 -> Average Cal sensor Temp.: 100.88 Ambient Temp.: 96.60 F
18:02:07.636 -> Average Cal sensor Temp.: 100.36 Ambient Temp.: 96.53 F
18:02:09.191 -> Average Cal sensor Temp.: 100.49 Ambient Temp.: 96.60 F
18:02:10.774 -> Average Cal sensor Temp.: 99.73 Ambient Temp.: 96.60 F
18:02:12.331 -> Average Cal sensor Temp.: 100.15 Ambient Temp.: 96.60 F
18:02:13.915 -> Average Cal sensor Temp.: 100.43 Ambient Temp.: 96.60 F
18:02:15.470 -> Average Cal sensor Temp.: 99.66 Ambient Temp.: 96.60 F
18:02:17.019 -> Average Cal sensor Temp.: 100.03 Ambient Temp.: 96.57 F
18:02:18.604 -> Average Cal sensor Temp.: 100.01 Ambient Temp.: 96.60 F
18:02:20.157 -> Average Cal sensor Temp.: 100.44 Ambient Temp.: 96.60 F
18:02:21.745 -> Average Cal sensor Temp.: 100.45 Ambient Temp.: 96.64 F
18:02:23.295 -> Average Cal sensor Temp.: 100.53 Ambient Temp.: 96.60 F
18:02:24.851 -> Average Cal sensor Temp.: 100.49 Ambient Temp.: 96.57 F
18:02:26.411 -> Average Cal sensor Temp.: 100.45 Ambient Temp.: 96.64 F
18:02:28.003 -> Average Cal sensor Temp.: 100.67 Ambient Temp.: 96.64 F
18:02:29.557 -> Average Cal sensor Temp.: 100.71 Ambient Temp.: 96.60 F
18:02:31.115 -> Average Cal sensor Temp.: 100.22 Ambient Temp.: 96.64 F
18:02:32.703 -> Average Cal sensor Temp.: 100.56 Ambient Temp.: 96.60 F
18:02:34.249 -> Average Cal sensor Temp.: 100.78 Ambient Temp.: 96.60 F
18:02:35.809 -> Average Cal sensor Temp.: 100.73 Ambient Temp.: 96.64 F
18:02:37.401 -> Average Cal sensor Temp.: 99.95 Ambient Temp.: 96.64 F
18:02:38.952 -> Average Cal sensor Temp.: 99.99 Ambient Temp.: 96.60 F
18:02:40.501 -> Average Cal sensor Temp.: 100.53 Ambient Temp.: 96.60 F
18:02:42.087 -> Average Cal sensor Temp.: 100.35 Ambient Temp.: 96.60 F
18:02:43.636 -> Average Cal sensor Temp.: 100.27 Ambient Temp.: 96.64 F
18:02:45.227 -> Average Cal sensor Temp.: 99.86 Ambient Temp.: 96.64 F
18:02:46.784 -> Average Cal sensor Temp.: 100.35 Ambient Temp.: 96.60 F
18:02:48.340 -> Average Cal sensor Temp.: 100.47 Ambient Temp.: 96.64 F
18:02:49.900 -> Average Cal sensor Temp.: 100.53 Ambient Temp.: 96.60 F
18:02:51.486 -> Average Cal sensor Temp.: 100.42 Ambient Temp.: 96.64 F
18:02:53.039 -> Average Cal sensor Temp.: 100.49 Ambient Temp.: 96.64 F
18:02:54.622 -> Average Cal sensor Temp.: 100.37 Ambient Temp.: 96.64 F
18:02:56.176 -> Average Cal sensor Temp.: 100.53 Ambient Temp.: 96.64 F
18:02:57.733 -> Average Cal sensor Temp.: 101.21 Ambient Temp.: 96.60 F
18:02:59.317 -> Average Cal sensor Temp.: 100.99 Ambient Temp.: 96.64 F
18:03:00.869 -> Average Cal sensor Temp.: 101.10 Ambient Temp.: 96.60 F
18:03:02.426 -> Average Cal sensor Temp.: 100.97 Ambient Temp.: 96.67 F
18:03:04.009 -> Average Cal sensor Temp.: 100.76 Ambient Temp.: 96.60 F
18:03:05.560 -> Average Cal sensor Temp.: 100.87 Ambient Temp.: 96.64 F
18:03:07.145 -> Average Cal sensor Temp.: 100.98 Ambient Temp.: 96.64 F
18:03:08.709 -> Average Cal sensor Temp.: 100.86 Ambient Temp.: 96.64 F
18:03:10.267 -> Average Cal sensor Temp.: 100.39 Ambient Temp.: 96.67 F
18:03:11.826 -> Average Cal sensor Temp.: 100.55 Ambient Temp.: 96.67 F
18:03:13.416 -> Average Cal sensor Temp.: 100.57 Ambient Temp.: 96.64 F
18:03:14.971 -> Average Cal sensor Temp.: 100.79 Ambient Temp.: 96.64 F
18:03:16.550 -> Average Cal sensor Temp.: 100.71 Ambient Temp.: 96.64 F
18:03:18.107 -> Average Cal sensor Temp.: 100.49 Ambient Temp.: 96.60 F
18:03:19.664 -> Average Cal sensor Temp.: 99.99 Ambient Temp.: 96.60 F
18:03:21.248 -> Average Cal sensor Temp.: 100.22 Ambient Temp.: 96.64 F
18:03:22.798 -> Average Cal sensor Temp.: 100.50 Ambient Temp.: 96.64 F
18:03:24.355 -> Average Cal sensor Temp.: 100.50 Ambient Temp.: 96.67 F
18:03:25.914 -> Average Cal sensor Temp.: 100.44 Ambient Temp.: 96.67 F
18:03:27.504 -> Average Cal sensor Temp.: 100.95 Ambient Temp.: 96.67 F
18:03:29.061 -> Average Cal sensor Temp.: 101.12 Ambient Temp.: 96.67 F
18:03:30.611 -> Average Cal sensor Temp.: 101.24 Ambient Temp.: 96.67 F
18:03:32.200 -> Average Cal sensor Temp.: 101.06 Ambient Temp.: 96.71 F
18:03:33.752 -> Average Cal sensor Temp.: 101.15 Ambient Temp.: 96.60 F
18:03:35.301 -> Average Cal sensor Temp.: 101.11 Ambient Temp.: 96.64 F
18:03:36.890 -> Average Cal sensor Temp.: 100.89 Ambient Temp.: 96.67 F
18:03:38.438 -> Average Cal sensor Temp.: 100.86 Ambient Temp.: 96.67 F
18:03:40.021 -> Average Cal sensor Temp.: 100.77 Ambient Temp.: 96.67 F
18:03:41.571 -> Average Cal sensor Temp.: 100.54 Ambient Temp.: 96.67 F
18:03:43.156 -> Average Cal sensor Temp.: 100.56 Ambient Temp.: 96.64 F
18:03:44.707 -> Average Cal sensor Temp.: 100.77 Ambient Temp.: 96.71 F
18:03:46.259 -> Average Cal sensor Temp.: 100.98 Ambient Temp.: 96.71 F
18:03:47.842 -> Average Cal sensor Temp.: 100.80 Ambient Temp.: 96.64 F
18:03:49.395 -> Average Cal sensor Temp.: 100.85 Ambient Temp.: 96.67 F
18:03:50.982 -> Average Cal sensor Temp.: 100.89 Ambient Temp.: 96.71 F
18:03:52.537 -> Average Cal sensor Temp.: 100.74 Ambient Temp.: 96.75 F
18:03:54.091 -> Average Cal sensor Temp.: 100.75 Ambient Temp.: 96.67 F
18:03:55.682 -> Average Cal sensor Temp.: 100.71 Ambient Temp.: 96.67 F
18:03:57.243 -> Average Cal sensor Temp.: 100.80 Ambient Temp.: 96.67 F
18:03:58.797 -> Average Cal sensor Temp.: 100.43 Ambient Temp.: 96.71 F
18:04:00.383 -> Average Cal sensor Temp.: 100.30 Ambient Temp.: 96.71 F
18:04:01.933 -> Average Cal sensor Temp.: 100.44 Ambient Temp.: 96.75 F
18:04:03.481 -> Average Cal sensor Temp.: 100.52 Ambient Temp.: 96.71 F
18:04:05.060 -> Average Cal sensor Temp.: 99.76 Ambient Temp.: 96.67 F
18:04:06.613 -> Average Cal sensor Temp.: 99.43 Ambient Temp.: 96.71 F
18:04:08.196 -> Average Cal sensor Temp.: 100.12 Ambient Temp.: 96.67 F
18:04:09.744 -> Average Cal sensor Temp.: 100.55 Ambient Temp.: 96.71 F
18:04:11.296 -> Average Cal sensor Temp.: 100.62 Ambient Temp.: 96.71 F
18:04:12.887 -> Average Cal sensor Temp.: 99.90 Ambient Temp.: 96.67 F
18:04:14.438 -> Average Cal sensor Temp.: 100.23 Ambient Temp.: 96.64 F
18:04:16.003 -> Average Cal sensor Temp.: 100.59 Ambient Temp.: 96.67 F
18:04:17.587 -> Average Cal sensor Temp.: 100.53 Ambient Temp.: 96.67 F
18:04:19.139 -> Average Cal sensor Temp.: 100.36 Ambient Temp.: 96.64 F
18:04:20.689 -> Average Cal sensor Temp.: 100.63 Ambient Temp.: 96.67 F
18:04:22.269 -> Average Cal sensor Temp.: 100.90 Ambient Temp.: 96.67 F
18:04:23.825 -> Average Cal sensor Temp.: 100.74 Ambient Temp.: 96.67 F
18:04:25.417 -> Average Cal sensor Temp.: 100.83 Ambient Temp.: 96.64 F
18:04:26.972 -> Average Cal sensor Temp.: 100.83 Ambient Temp.: 96.64 F
18:04:28.531 -> Average Cal sensor Temp.: 100.87 Ambient Temp.: 96.64 F
18:04:30.107 -> Average Cal sensor Temp.: 100.84 Ambient Temp.: 96.67 F
18:04:31.664 -> Average Cal sensor Temp.: 100.57 Ambient Temp.: 96.71 F
18:04:33.246 -> Average Cal sensor Temp.: 100.55 Ambient Temp.: 96.67 F
18:04:34.800 -> Average Cal sensor Temp.: 100.60 Ambient Temp.: 96.71 F
18:04:36.354 -> Average Cal sensor Temp.: 100.71 Ambient Temp.: 96.71 F
18:04:37.940 -> Average Cal sensor Temp.: 100.96 Ambient Temp.: 96.71 F
18:04:39.491 -> Average Cal sensor Temp.: 100.08 Ambient Temp.: 96.71 F
18:04:41.072 -> Average Cal sensor Temp.: 100.15 Ambient Temp.: 96.71 F
18:04:42.629 -> Average Cal sensor Temp.: 100.96 Ambient Temp.: 96.75 F
18:04:44.181 -> Average Cal sensor Temp.: 100.59 Ambient Temp.: 96.67 F
18:04:45.739 -> Average Cal sensor Temp.: 100.45 Ambient Temp.: 96.67 F
18:04:47.328 -> Average Cal sensor Temp.: 100.76 Ambient Temp.: 96.71 F
18:04:48.885 -> Average Cal sensor Temp.: 100.88 Ambient Temp.: 96.67 F
18:04:50.440 -> Average Cal sensor Temp.: 100.35 Ambient Temp.: 96.67 F
18:04:52.023 -> Average Cal sensor Temp.: 100.16 Ambient Temp.: 96.64 F
18:04:53.570 -> Average Cal sensor Temp.: 100.43 Ambient Temp.: 96.67 F
18:04:55.156 -> Average Cal sensor Temp.: 100.90 Ambient Temp.: 96.67 F
18:04:56.708 -> Average Cal sensor Temp.: 100.63 Ambient Temp.: 96.60 F
18:04:58.290 -> Average Cal sensor Temp.: 100.49 Ambient Temp.: 96.67 F
18:04:59.848 -> Average Cal sensor Temp.: 100.33 Ambient Temp.: 96.64 F
18:05:01.432 -> Average Cal sensor Temp.: 100.45 Ambient Temp.: 96.64 F
18:05:02.984 -> Average Cal sensor Temp.: 100.45 Ambient Temp.: 96.64 F
18:05:04.565 -> Average Cal sensor Temp.: 99.86 Ambient Temp.: 96.64 F
18:05:06.123 -> Average Cal sensor Temp.: 99.93 Ambient Temp.: 96.64 F
18:05:07.677 -> Average Cal sensor Temp.: 100.05 Ambient Temp.: 96.67 F
18:05:09.230 -> Average Cal sensor Temp.: 100.24 Ambient Temp.: 96.64 F
18:05:10.793 -> Average Cal sensor Temp.: 100.56 Ambient Temp.: 96.60 F
18:05:12.377 -> Average Cal sensor Temp.: 100.72 Ambient Temp.: 96.64 F
18:05:13.929 -> Average Cal sensor Temp.: 100.73 Ambient Temp.: 96.64 F
18:05:15.511 -> Average Cal sensor Temp.: 100.77 Ambient Temp.: 96.60 F
18:05:17.062 -> Average Cal sensor Temp.: 100.67 Ambient Temp.: 96.64 F
18:05:18.647 -> Average Cal sensor Temp.: 101.00 Ambient Temp.: 96.67 F
18:05:20.200 -> Average Cal sensor Temp.: 100.77 Ambient Temp.: 96.64 F
18:05:21.783 -> Average Cal sensor Temp.: 100.49 Ambient Temp.: 96.64 F
18:05:23.338 -> Average Cal sensor Temp.: 100.78 Ambient Temp.: 96.67 F
18:05:24.891 -> Average Cal sensor Temp.: 100.67 Ambient Temp.: 96.67 F
18:05:26.482 -> Average Cal sensor Temp.: 100.88 Ambient Temp.: 96.67 F
18:05:28.039 -> Average Cal sensor Temp.: 100.76 Ambient Temp.: 96.71 F
18:05:29.601 -> Average Cal sensor Temp.: 100.67 Ambient Temp.: 96.67 F
18:05:31.154 -> Average Cal sensor Temp.: 100.96 Ambient Temp.: 96.71 F
18:05:32.746 -> Average Cal sensor Temp.: 100.35 Ambient Temp.: 96.67 F
18:05:34.301 -> Average Cal sensor Temp.: 100.08 Ambient Temp.: 96.67 F
18:05:35.857 -> Average Cal sensor Temp.: 100.53 Ambient Temp.: 96.71 F
18:05:37.414 -> Average Cal sensor Temp.: 100.08 Ambient Temp.: 96.64 F
18:05:39.006 -> Average Cal sensor Temp.: 100.16 Ambient Temp.: 96.67 F
18:05:40.565 -> Average Cal sensor Temp.: 100.35 Ambient Temp.: 96.71 F
18:05:42.124 -> Average Cal sensor Temp.: 99.66 Ambient Temp.: 96.67 F
18:05:43.674 -> Average Cal sensor Temp.: 100.39 Ambient Temp.: 96.75 F
18:05:45.239 -> Average Cal sensor Temp.: 100.25 Ambient Temp.: 96.71 F
18:05:46.829 -> Average Cal sensor Temp.: 100.22 Ambient Temp.: 96.67 F
18:05:48.380 -> Average Cal sensor Temp.: 100.55 Ambient Temp.: 96.67 F
18:05:49.935 -> Average Cal sensor Temp.: 100.44 Ambient Temp.: 96.64 F
18:05:51.524 -> Average Cal sensor Temp.: 100.73 Ambient Temp.: 96.67 F
18:05:53.079 -> Average Cal sensor Temp.: 100.58 Ambient Temp.: 96.64 F
18:05:54.639 -> Average Cal sensor Temp.: 100.04 Ambient Temp.: 96.64 F
18:05:56.229 -> Average Cal sensor Temp.: 100.47 Ambient Temp.: 96.64 F
18:05:57.782 -> Average Cal sensor Temp.: 100.34 Ambient Temp.: 96.60 F
18:05:59.333 -> Average Cal sensor Temp.: 100.60 Ambient Temp.: 96.64 F
18:06:00.923 -> Average Cal sensor Temp.: 100.37 Ambient Temp.: 96.67 F
18:06:02.481 -> Average Cal sensor Temp.: 100.78 Ambient Temp.: 96.64 F
18:06:04.040 -> Average Cal sensor Temp.: 100.87 Ambient Temp.: 96.71 F
18:06:05.601 -> Average Cal sensor Temp.: 100.97 Ambient Temp.: 96.64 F
18:06:07.192 -> Average Cal sensor Temp.: 100.83 Ambient Temp.: 96.64 F
18:06:08.740 -> Average Cal sensor Temp.: 101.02 Ambient Temp.: 96.60 F
18:06:10.303 -> Average Cal sensor Temp.: 100.85 Ambient Temp.: 96.64 F
18:06:11.890 -> Average Cal sensor Temp.: 100.66 Ambient Temp.: 96.71 F
18:06:13.441 -> Average Cal sensor Temp.: 100.94 Ambient Temp.: 96.67 F
18:06:14.994 -> Average Cal sensor Temp.: 100.73 Ambient Temp.: 96.75 F
18:06:16.587 -> Average Cal sensor Temp.: 100.38 Ambient Temp.: 96.67 F
18:06:18.143 -> Average Cal sensor Temp.: 100.38 Ambient Temp.: 96.71 F
18:06:19.701 -> Average Cal sensor Temp.: 100.05 Ambient Temp.: 96.67 F
18:06:21.288 -> Average Cal sensor Temp.: 100.54 Ambient Temp.: 96.64 F
18:06:22.843 -> Average Cal sensor Temp.: 100.71 Ambient Temp.: 96.67 F
18:06:24.391 -> Average Cal sensor Temp.: 100.89 Ambient Temp.: 96.67 F
18:06:25.980 -> Average Cal sensor Temp.: 99.91 Ambient Temp.: 96.67 F
18:06:27.537 -> Average Cal sensor Temp.: 99.75 Ambient Temp.: 96.64 F
18:06:29.083 -> Average Cal sensor Temp.: 100.49 Ambient Temp.: 96.67 F
18:06:30.671 -> Average Cal sensor Temp.: 100.85 Ambient Temp.: 96.67 F
18:06:32.226 -> Average Cal sensor Temp.: 100.69 Ambient Temp.: 96.64 F
18:06:33.779 -> Average Cal sensor Temp.: 100.82 Ambient Temp.: 96.67 F
18:06:35.369 -> Average Cal sensor Temp.: 101.13 Ambient Temp.: 96.64 F
18:06:36.924 -> Average Cal sensor Temp.: 101.17 Ambient Temp.: 96.67 F
18:06:38.471 -> Average Cal sensor Temp.: 101.27 Ambient Temp.: 96.71 F
18:06:40.057 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 96.67 F
18:06:41.607 -> Average Cal sensor Temp.: 101.24 Ambient Temp.: 96.67 F
18:06:43.167 -> Average Cal sensor Temp.: 100.43 Ambient Temp.: 96.67 F
18:06:44.756 -> Average Cal sensor Temp.: 100.85 Ambient Temp.: 96.67 F
18:06:46.315 -> Average Cal sensor Temp.: 100.92 Ambient Temp.: 96.71 F
18:06:47.868 -> Average Cal sensor Temp.: 100.63 Ambient Temp.: 96.71 F
18:06:49.454 -> Average Cal sensor Temp.: 101.11 Ambient Temp.: 96.71 F
18:06:51.007 -> Average Cal sensor Temp.: 100.62 Ambient Temp.: 96.60 F
18:06:52.561 -> Average Cal sensor Temp.: 100.62 Ambient Temp.: 96.67 F
18:06:54.140 -> Average Cal sensor Temp.: 100.85 Ambient Temp.: 96.67 F
18:06:55.691 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 96.64 F
18:06:57.276 -> Average Cal sensor Temp.: 100.69 Ambient Temp.: 96.67 F
18:06:58.834 -> Average Cal sensor Temp.: 100.71 Ambient Temp.: 96.64 F
18:07:00.384 -> Average Cal sensor Temp.: 100.65 Ambient Temp.: 96.71 F
18:07:01.968 -> Average Cal sensor Temp.: 100.80 Ambient Temp.: 96.64 F
18:07:03.522 -> Average Cal sensor Temp.: 100.99 Ambient Temp.: 96.67 F
18:07:05.106 -> Average Cal sensor Temp.: 101.06 Ambient Temp.: 96.67 F
18:07:06.662 -> Average Cal sensor Temp.: 101.20 Ambient Temp.: 96.71 F
18:07:08.219 -> Average Cal sensor Temp.: 101.24 Ambient Temp.: 96.71 F
18:07:09.774 -> Average Cal sensor Temp.: 100.70 Ambient Temp.: 96.67 F
18:07:10.640 -> Ambient Temp.: 96.67 F
18:07:11.194 -> Ambient Temp.: 96.67 F
18:07:11.780 -> Ambient Temp.: 96.67 F
18:07:13.335 -> Average Cal sensor Temp.: 100.54 Ambient Temp.: 96.64 F
18:07:14.921 -> Average Cal sensor Temp.: 100.64 Ambient Temp.: 96.60 F
18:07:16.470 -> Average Cal sensor Temp.: 99.86 Ambient Temp.: 96.67 F
18:07:18.050 -> Average Cal sensor Temp.: 100.13 Ambient Temp.: 96.67 F
18:07:19.607 -> Average Cal sensor Temp.: 100.51 Ambient Temp.: 96.67 F
18:07:21.165 -> Average Cal sensor Temp.: 100.73 Ambient Temp.: 96.67 F
18:07:22.722 -> Average Cal sensor Temp.: 101.06 Ambient Temp.: 96.67 F
18:07:24.310 -> Average Cal sensor Temp.: 100.88 Ambient Temp.: 96.71 F
18:07:25.865 -> Average Cal sensor Temp.: 100.92 Ambient Temp.: 96.67 F
18:07:27.422 -> Average Cal sensor Temp.: 100.60 Ambient Temp.: 96.67 F
18:07:29.006 -> Average Cal sensor Temp.: 100.45 Ambient Temp.: 96.67 F
18:07:30.557 -> Average Cal sensor Temp.: 100.46 Ambient Temp.: 96.67 F
18:07:32.138 -> Average Cal sensor Temp.: 101.09 Ambient Temp.: 96.71 F
18:07:33.691 -> Average Cal sensor Temp.: 101.25 Ambient Temp.: 96.67 F
18:07:35.240 -> Average Cal sensor Temp.: 101.16 Ambient Temp.: 96.64 F
18:07:36.826 -> Average Cal sensor Temp.: 101.03 Ambient Temp.: 96.71 F
18:07:38.374 -> Average Cal sensor Temp.: 100.24 Ambient Temp.: 96.75 F
18:07:39.959 -> Average Cal sensor Temp.: 101.01 Ambient Temp.: 96.64 F
18:07:41.509 -> Average Cal sensor Temp.: 101.01 Ambient Temp.: 96.71 F
18:07:43.059 -> Average Cal sensor Temp.: 100.68 Ambient Temp.: 96.71 F
18:07:44.647 -> Average Cal sensor Temp.: 100.73 Ambient Temp.: 96.71 F
18:07:46.200 -> Average Cal sensor Temp.: 100.74 Ambient Temp.: 96.60 F
18:07:47.783 -> Average Cal sensor Temp.: 100.46 Ambient Temp.: 96.64 F
18:07:49.342 -> Average Cal sensor Temp.: 100.44 Ambient Temp.: 96.67 F
18:07:50.894 -> Average Cal sensor Temp.: 100.87 Ambient Temp.: 96.64 F
18:07:52.479 -> Average Cal sensor Temp.: 101.08 Ambient Temp.: 96.71 F
18:07:54.034 -> Average Cal sensor Temp.: 100.67 Ambient Temp.: 96.71 F
18:07:55.593 -> Average Cal sensor Temp.: 101.00 Ambient Temp.: 96.64 F
18:07:57.183 -> Average Cal sensor Temp.: 101.25 Ambient Temp.: 96.67 F
18:07:58.732 -> Average Cal sensor Temp.: 101.37 Ambient Temp.: 96.67 F
18:08:00.280 -> Average Cal sensor Temp.: 100.78 Ambient Temp.: 96.67 F
18:08:01.872 -> Average Cal sensor Temp.: 100.59 Ambient Temp.: 96.67 F
18:08:03.425 -> Average Cal sensor Temp.: 101.15 Ambient Temp.: 96.67 F
18:08:04.980 -> Average Cal sensor Temp.: 100.94 Ambient Temp.: 96.67 F
18:08:06.569 -> Average Cal sensor Temp.: 101.06 Ambient Temp.: 96.75 F
18:08:08.126 -> Average Cal sensor Temp.: 101.01 Ambient Temp.: 96.64 F
18:08:09.676 -> Average Cal sensor Temp.: 101.10 Ambient Temp.: 96.71 F
18:08:11.232 -> Average Cal sensor Temp.: 101.45 Ambient Temp.: 96.67 F
18:08:12.822 -> Average Cal sensor Temp.: 101.06 Ambient Temp.: 96.71 F
18:08:14.372 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 96.71 F
18:08:15.934 -> Average Cal sensor Temp.: 101.36 Ambient Temp.: 96.67 F
18:08:17.490 -> Average Cal sensor Temp.: 100.88 Ambient Temp.: 96.64 F
18:08:19.078 -> Average Cal sensor Temp.: 100.84 Ambient Temp.: 96.71 F
18:08:20.633 -> Average Cal sensor Temp.: 100.93 Ambient Temp.: 96.71 F
18:08:22.183 -> Average Cal sensor Temp.: 100.94 Ambient Temp.: 96.71 F
18:08:23.767 -> Average Cal sensor Temp.: 100.99 Ambient Temp.: 96.71 F
18:08:25.323 -> Average Cal sensor Temp.: 100.73 Ambient Temp.: 96.67 F
18:08:26.883 -> Average Cal sensor Temp.: 100.97 Ambient Temp.: 96.71 F
18:08:28.478 -> Average Cal sensor Temp.: 100.94 Ambient Temp.: 96.71 F
18:08:30.028 -> Average Cal sensor Temp.: 101.06 Ambient Temp.: 96.75 F
18:08:31.578 -> Average Cal sensor Temp.: 101.03 Ambient Temp.: 96.71 F
18:08:33.165 -> Average Cal sensor Temp.: 100.77 Ambient Temp.: 96.71 F
18:08:34.721 -> Average Cal sensor Temp.: 100.95 Ambient Temp.: 96.75 F
18:08:36.278 -> Average Cal sensor Temp.: 100.57 Ambient Temp.: 96.67 F
18:08:37.864 -> Average Cal sensor Temp.: 100.22 Ambient Temp.: 96.75 F
18:08:39.414 -> Average Cal sensor Temp.: 100.70 Ambient Temp.: 96.71 F
18:08:40.971 -> Average Cal sensor Temp.: 100.85 Ambient Temp.: 96.75 F
18:08:42.552 -> Average Cal sensor Temp.: 100.80 Ambient Temp.: 96.67 F
18:08:44.108 -> Average Cal sensor Temp.: 100.48 Ambient Temp.: 96.71 F
18:08:45.662 -> Average Cal sensor Temp.: 100.09 Ambient Temp.: 96.71 F
18:08:47.243 -> Average Cal sensor Temp.: 100.98 Ambient Temp.: 96.71 F
18:08:48.793 -> Average Cal sensor Temp.: 100.65 Ambient Temp.: 96.67 F
18:08:50.385 -> Average Cal sensor Temp.: 100.84 Ambient Temp.: 96.71 F
18:08:51.941 -> Average Cal sensor Temp.: 101.10 Ambient Temp.: 96.71 F
18:08:53.496 -> Average Cal sensor Temp.: 101.00 Ambient Temp.: 96.75 F
18:08:55.051 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 96.75 F
18:08:56.645 -> Average Cal sensor Temp.: 101.16 Ambient Temp.: 96.71 F
18:08:58.201 -> Average Cal sensor Temp.: 101.16 Ambient Temp.: 96.67 F
18:08:59.752 -> Average Cal sensor Temp.: 100.92 Ambient Temp.: 96.71 F
18:09:01.336 -> Average Cal sensor Temp.: 100.88 Ambient Temp.: 96.64 F
18:09:02.890 -> Average Cal sensor Temp.: 101.25 Ambient Temp.: 96.67 F
18:09:04.437 -> Average Cal sensor Temp.: 100.92 Ambient Temp.: 96.71 F
18:09:06.026 -> Average Cal sensor Temp.: 101.20 Ambient Temp.: 96.75 F
18:09:07.579 -> Average Cal sensor Temp.: 101.11 Ambient Temp.: 96.75 F
18:09:09.131 -> Average Cal sensor Temp.: 101.24 Ambient Temp.: 96.75 F
18:09:10.719 -> Average Cal sensor Temp.: 101.05 Ambient Temp.: 96.75 F
18:09:12.271 -> Average Cal sensor Temp.: 100.94 Ambient Temp.: 96.71 F
18:09:13.828 -> Average Cal sensor Temp.: 101.00 Ambient Temp.: 96.71 F
18:09:15.387 -> Average Cal sensor Temp.: 101.09 Ambient Temp.: 96.78 F
18:09:16.972 -> Average Cal sensor Temp.: 101.15 Ambient Temp.: 96.71 F
18:09:18.525 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 96.75 F
18:09:20.074 -> Average Cal sensor Temp.: 101.08 Ambient Temp.: 96.78 F
18:09:21.661 -> Average Cal sensor Temp.: 100.57 Ambient Temp.: 96.82 F
18:09:23.219 -> Average Cal sensor Temp.: 101.08 Ambient Temp.: 96.75 F
18:09:24.773 -> Average Cal sensor Temp.: 100.70 Ambient Temp.: 96.75 F
18:09:26.359 -> Average Cal sensor Temp.: 100.64 Ambient Temp.: 96.71 F
18:09:27.913 -> Average Cal sensor Temp.: 100.79 Ambient Temp.: 96.75 F
18:09:29.470 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 96.78 F
18:09:31.058 -> Average Cal sensor Temp.: 101.00 Ambient Temp.: 96.75 F
18:09:32.606 -> Average Cal sensor Temp.: 100.68 Ambient Temp.: 96.75 F
18:09:34.191 -> Average Cal sensor Temp.: 100.58 Ambient Temp.: 96.78 F
18:09:35.750 -> Average Cal sensor Temp.: 100.38 Ambient Temp.: 96.75 F
18:09:37.301 -> Average Cal sensor Temp.: 101.04 Ambient Temp.: 96.78 F
18:09:38.888 -> Average Cal sensor Temp.: 101.36 Ambient Temp.: 96.71 F
18:09:40.442 -> Average Cal sensor Temp.: 100.26 Ambient Temp.: 96.71 F
18:09:41.995 -> Average Cal sensor Temp.: 100.18 Ambient Temp.: 96.67 F
18:09:43.555 -> Average Cal sensor Temp.: 100.78 Ambient Temp.: 96.71 F
18:09:45.145 -> Average Cal sensor Temp.: 101.21 Ambient Temp.: 96.67 F
18:09:46.700 -> Average Cal sensor Temp.: 101.34 Ambient Temp.: 96.71 F
18:09:48.257 -> Average Cal sensor Temp.: 100.97 Ambient Temp.: 96.71 F
18:09:49.811 -> Average Cal sensor Temp.: 101.03 Ambient Temp.: 96.71 F
18:09:51.398 -> Average Cal sensor Temp.: 101.19 Ambient Temp.: 96.71 F
18:09:52.945 -> Average Cal sensor Temp.: 101.25 Ambient Temp.: 96.71 F
18:09:54.502 -> Average Cal sensor Temp.: 100.88 Ambient Temp.: 96.71 F
18:09:56.087 -> Average Cal sensor Temp.: 100.78 Ambient Temp.: 96.75 F
18:09:57.644 -> Average Cal sensor Temp.: 101.19 Ambient Temp.: 96.75 F
18:09:59.195 -> Average Cal sensor Temp.: 100.99 Ambient Temp.: 96.71 F
18:10:00.776 -> Average Cal sensor Temp.: 101.03 Ambient Temp.: 96.71 F
18:10:02.330 -> Average Cal sensor Temp.: 100.98 Ambient Temp.: 96.71 F
18:10:03.885 -> Average Cal sensor Temp.: 101.05 Ambient Temp.: 96.78 F
18:10:05.469 -> Average Cal sensor Temp.: 100.77 Ambient Temp.: 96.78 F
18:10:07.016 -> Average Cal sensor Temp.: 101.03 Ambient Temp.: 96.71 F
18:10:08.608 -> Average Cal sensor Temp.: 100.79 Ambient Temp.: 96.71 F
18:10:10.166 -> Average Cal sensor Temp.: 101.08 Ambient Temp.: 96.75 F
18:10:11.720 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 96.71 F
18:10:13.272 -> Average Cal sensor Temp.: 101.32 Ambient Temp.: 96.71 F
18:10:14.854 -> Average Cal sensor Temp.: 100.60 Ambient Temp.: 96.75 F
18:10:16.407 -> Average Cal sensor Temp.: 100.92 Ambient Temp.: 96.78 F
18:10:17.996 -> Average Cal sensor Temp.: 101.18 Ambient Temp.: 96.71 F
18:10:19.552 -> Average Cal sensor Temp.: 101.24 Ambient Temp.: 96.71 F
18:10:21.105 -> Average Cal sensor Temp.: 101.15 Ambient Temp.: 96.75 F
18:10:22.659 -> Average Cal sensor Temp.: 101.16 Ambient Temp.: 96.71 F
18:10:24.238 -> Average Cal sensor Temp.: 101.20 Ambient Temp.: 96.78 F
18:10:25.786 -> Average Cal sensor Temp.: 100.93 Ambient Temp.: 96.75 F
18:10:27.373 -> Average Cal sensor Temp.: 101.06 Ambient Temp.: 96.78 F
18:10:28.932 -> Average Cal sensor Temp.: 100.98 Ambient Temp.: 96.75 F
18:10:30.486 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 96.78 F
18:10:32.071 -> Average Cal sensor Temp.: 101.16 Ambient Temp.: 96.75 F
18:10:33.625 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 96.75 F
18:10:35.176 -> Average Cal sensor Temp.: 100.99 Ambient Temp.: 96.75 F
18:10:36.765 -> Average Cal sensor Temp.: 101.10 Ambient Temp.: 96.75 F
18:10:38.315 -> Average Cal sensor Temp.: 101.14 Ambient Temp.: 96.78 F
18:10:39.866 -> Average Cal sensor Temp.: 100.97 Ambient Temp.: 96.75 F
18:10:41.451 -> Average Cal sensor Temp.: 100.71 Ambient Temp.: 96.75 F
18:10:43.003 -> Average Cal sensor Temp.: 101.10 Ambient Temp.: 96.75 F
18:10:44.561 -> Average Cal sensor Temp.: 101.48 Ambient Temp.: 96.75 F
18:10:46.120 -> Average Cal sensor Temp.: 101.32 Ambient Temp.: 96.78 F
18:10:47.704 -> Average Cal sensor Temp.: 100.87 Ambient Temp.: 96.78 F
18:10:49.253 -> Average Cal sensor Temp.: 100.70 Ambient Temp.: 96.78 F
18:10:50.834 -> Average Cal sensor Temp.: 100.88 Ambient Temp.: 96.75 F
18:10:52.383 -> Average Cal sensor Temp.: 99.90 Ambient Temp.: 96.75 F
18:10:53.967 -> Average Cal sensor Temp.: 100.48 Ambient Temp.: 96.71 F
18:10:55.520 -> Average Cal sensor Temp.: 100.48 Ambient Temp.: 96.71 F
18:10:57.106 -> Average Cal sensor Temp.: 100.44 Ambient Temp.: 96.67 F
18:10:58.656 -> Average Cal sensor Temp.: 101.14 Ambient Temp.: 96.71 F
18:11:00.201 -> Average Cal sensor Temp.: 101.35 Ambient Temp.: 96.71 F
18:11:01.788 -> Average Cal sensor Temp.: 101.08 Ambient Temp.: 96.67 F
18:11:03.339 -> Average Cal sensor Temp.: 100.70 Ambient Temp.: 96.64 F
18:11:04.899 -> Average Cal sensor Temp.: 100.79 Ambient Temp.: 96.78 F
18:11:06.490 -> Average Cal sensor Temp.: 100.87 Ambient Temp.: 96.71 F
18:11:08.046 -> Average Cal sensor Temp.: 101.20 Ambient Temp.: 96.67 F
18:11:09.602 -> Average Cal sensor Temp.: 101.20 Ambient Temp.: 96.67 F
18:11:11.156 -> Average Cal sensor Temp.: 101.31 Ambient Temp.: 96.64 F
18:11:12.751 -> Average Cal sensor Temp.: 101.52 Ambient Temp.: 96.71 F
18:11:14.302 -> Average Cal sensor Temp.: 101.32 Ambient Temp.: 96.71 F
18:11:15.858 -> Average Cal sensor Temp.: 101.17 Ambient Temp.: 96.67 F
18:11:17.409 -> Average Cal sensor Temp.: 101.43 Ambient Temp.: 96.67 F
18:11:18.996 -> Average Cal sensor Temp.: 101.21 Ambient Temp.: 96.67 F
18:11:20.550 -> Average Cal sensor Temp.: 100.87 Ambient Temp.: 96.67 F
18:11:22.105 -> Average Cal sensor Temp.: 100.92 Ambient Temp.: 96.71 F
18:11:23.691 -> Average Cal sensor Temp.: 100.96 Ambient Temp.: 96.71 F
18:11:25.239 -> Average Cal sensor Temp.: 101.12 Ambient Temp.: 96.67 F
18:11:26.823 -> Average Cal sensor Temp.: 101.09 Ambient Temp.: 96.67 F
18:11:28.376 -> Average Cal sensor Temp.: 101.15 Ambient Temp.: 96.71 F
18:11:29.937 -> Average Cal sensor Temp.: 100.97 Ambient Temp.: 96.67 F
18:11:31.490 -> Average Cal sensor Temp.: 100.52 Ambient Temp.: 96.67 F
18:11:33.083 -> Average Cal sensor Temp.: 100.22 Ambient Temp.: 96.64 F
18:11:34.633 -> Average Cal sensor Temp.: 100.81 Ambient Temp.: 96.67 F
18:11:36.216 -> Average Cal sensor Temp.: 100.67 Ambient Temp.: 96.67 F
18:11:37.767 -> Average Cal sensor Temp.: 101.09 Ambient Temp.: 96.67 F
18:11:39.322 -> Average Cal sensor Temp.: 101.11 Ambient Temp.: 96.64 F
18:11:40.877 -> Average Cal sensor Temp.: 100.98 Ambient Temp.: 96.60 F
18:11:42.459 -> Average Cal sensor Temp.: 100.28 Ambient Temp.: 96.67 F
18:11:44.043 -> Average Cal sensor Temp.: 100.82 Ambient Temp.: 96.64 F
18:11:45.600 -> Average Cal sensor Temp.: 100.88 Ambient Temp.: 96.64 F
18:11:47.154 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 96.67 F
18:11:48.735 -> Average Cal sensor Temp.: 101.01 Ambient Temp.: 96.64 F
18:11:50.283 -> Average Cal sensor Temp.: 101.00 Ambient Temp.: 96.67 F
18:11:51.869 -> Average Cal sensor Temp.: 100.73 Ambient Temp.: 96.64 F
18:11:53.419 -> Average Cal sensor Temp.: 100.74 Ambient Temp.: 96.64 F
18:11:54.971 -> Average Cal sensor Temp.: 101.12 Ambient Temp.: 96.71 F
18:11:56.554 -> Average Cal sensor Temp.: 101.17 Ambient Temp.: 96.67 F
18:11:58.097 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 96.60 F
18:11:59.690 -> Average Cal sensor Temp.: 101.43 Ambient Temp.: 96.67 F
18:12:01.242 -> Average Cal sensor Temp.: 101.49 Ambient Temp.: 96.64 F
18:12:02.798 -> Average Cal sensor Temp.: 101.32 Ambient Temp.: 96.67 F
18:12:04.355 -> Average Cal sensor Temp.: 101.74 Ambient Temp.: 96.67 F
18:12:05.943 -> Average Cal sensor Temp.: 102.33 Ambient Temp.: 96.71 F
18:12:07.502 -> Average Cal sensor Temp.: 102.32 Ambient Temp.: 96.67 F
18:12:09.059 -> Average Cal sensor Temp.: 102.14 Ambient Temp.: 96.71 F
18:12:10.608 -> Average Cal sensor Temp.: 102.10 Ambient Temp.: 96.75 F
18:12:12.199 -> Average Cal sensor Temp.: 101.37 Ambient Temp.: 96.75 F
18:12:13.748 -> Average Cal sensor Temp.: 100.83 Ambient Temp.: 96.75 F
18:12:15.300 -> Average Cal sensor Temp.: 101.09 Ambient Temp.: 96.71 F
18:12:16.885 -> Average Cal sensor Temp.: 100.85 Ambient Temp.: 96.75 F
18:12:18.436 -> Average Cal sensor Temp.: 100.71 Ambient Temp.: 96.75 F
18:12:20.019 -> Average Cal sensor Temp.: 100.58 Ambient Temp.: 96.67 F
18:12:21.573 -> Average Cal sensor Temp.: 101.06 Ambient Temp.: 96.75 F
18:12:23.127 -> Average Cal sensor Temp.: 101.09 Ambient Temp.: 96.71 F
18:12:24.719 -> Average Cal sensor Temp.: 101.45 Ambient Temp.: 96.75 F
18:12:26.274 -> Average Cal sensor Temp.: 101.51 Ambient Temp.: 96.75 F
18:12:27.825 -> Average Cal sensor Temp.: 101.43 Ambient Temp.: 96.75 F
18:12:29.405 -> Average Cal sensor Temp.: 101.52 Ambient Temp.: 96.71 F
18:12:30.957 -> Average Cal sensor Temp.: 100.63 Ambient Temp.: 96.75 F
18:12:32.509 -> Average Cal sensor Temp.: 101.37 Ambient Temp.: 96.71 F
18:12:34.092 -> Average Cal sensor Temp.: 101.58 Ambient Temp.: 96.78 F
18:12:35.639 -> Average Cal sensor Temp.: 100.97 Ambient Temp.: 96.71 F
18:12:37.194 -> Average Cal sensor Temp.: 101.36 Ambient Temp.: 96.75 F
18:12:38.780 -> Average Cal sensor Temp.: 101.67 Ambient Temp.: 96.75 F
18:12:40.332 -> Average Cal sensor Temp.: 101.34 Ambient Temp.: 96.75 F
18:12:41.915 -> Average Cal sensor Temp.: 101.03 Ambient Temp.: 96.71 F
18:12:43.468 -> Average Cal sensor Temp.: 101.56 Ambient Temp.: 96.75 F
18:12:45.022 -> Average Cal sensor Temp.: 101.62 Ambient Temp.: 96.75 F
18:12:46.577 -> Average Cal sensor Temp.: 101.25 Ambient Temp.: 96.75 F
18:12:48.163 -> Average Cal sensor Temp.: 101.11 Ambient Temp.: 96.75 F
18:12:49.720 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 96.78 F
18:12:51.274 -> Average Cal sensor Temp.: 101.37 Ambient Temp.: 96.75 F
18:12:52.831 -> Average Cal sensor Temp.: 101.47 Ambient Temp.: 96.78 F
18:12:54.419 -> Average Cal sensor Temp.: 100.96 Ambient Temp.: 96.78 F
18:12:55.975 -> Average Cal sensor Temp.: 101.20 Ambient Temp.: 96.82 F
18:12:57.537 -> Average Cal sensor Temp.: 101.54 Ambient Temp.: 96.78 F
18:12:59.093 -> Average Cal sensor Temp.: 101.91 Ambient Temp.: 96.78 F
18:13:00.650 -> Average Cal sensor Temp.: 102.12 Ambient Temp.: 96.78 F
18:13:02.238 -> Average Cal sensor Temp.: 101.69 Ambient Temp.: 96.82 F
18:13:03.786 -> Average Cal sensor Temp.: 101.01 Ambient Temp.: 96.78 F
18:13:05.373 -> Average Cal sensor Temp.: 101.16 Ambient Temp.: 96.78 F
18:13:06.921 -> Average Cal sensor Temp.: 100.76 Ambient Temp.: 96.75 F
18:13:08.473 -> Average Cal sensor Temp.: 100.88 Ambient Temp.: 96.75 F
18:13:10.062 -> Average Cal sensor Temp.: 100.80 Ambient Temp.: 96.75 F
18:13:11.613 -> Average Cal sensor Temp.: 101.25 Ambient Temp.: 96.71 F
18:13:13.167 -> Average Cal sensor Temp.: 101.68 Ambient Temp.: 96.82 F
18:13:14.756 -> Average Cal sensor Temp.: 101.56 Ambient Temp.: 96.78 F
18:13:16.310 -> Average Cal sensor Temp.: 101.16 Ambient Temp.: 96.78 F
18:13:17.862 -> Average Cal sensor Temp.: 101.46 Ambient Temp.: 96.78 F
18:13:19.450 -> Average Cal sensor Temp.: 101.57 Ambient Temp.: 96.78 F
18:13:21.003 -> Average Cal sensor Temp.: 101.75 Ambient Temp.: 96.82 F
18:13:22.555 -> Average Cal sensor Temp.: 101.58 Ambient Temp.: 96.78 F
18:13:24.137 -> Average Cal sensor Temp.: 101.04 Ambient Temp.: 96.78 F
18:13:25.693 -> Average Cal sensor Temp.: 101.12 Ambient Temp.: 96.78 F
18:13:27.251 -> Average Cal sensor Temp.: 101.02 Ambient Temp.: 96.75 F
18:13:28.837 -> Average Cal sensor Temp.: 101.52 Ambient Temp.: 96.71 F
18:13:30.391 -> Average Cal sensor Temp.: 101.50 Ambient Temp.: 96.75 F
18:13:31.938 -> Average Cal sensor Temp.: 101.06 Ambient Temp.: 96.71 F
18:13:33.521 -> Average Cal sensor Temp.: 101.23 Ambient Temp.: 96.71 F
18:13:35.074 -> Average Cal sensor Temp.: 100.96 Ambient Temp.: 96.71 F
18:13:36.628 -> Average Cal sensor Temp.: 101.41 Ambient Temp.: 96.78 F
18:13:38.212 -> Average Cal sensor Temp.: 101.99 Ambient Temp.: 96.75 F
18:13:39.760 -> Average Cal sensor Temp.: 101.80 Ambient Temp.: 96.75 F
18:13:41.317 -> Average Cal sensor Temp.: 101.88 Ambient Temp.: 96.78 F
18:13:42.905 -> Average Cal sensor Temp.: 101.88 Ambient Temp.: 96.75 F
18:13:44.458 -> Average Cal sensor Temp.: 101.74 Ambient Temp.: 96.75 F
18:13:46.017 -> Average Cal sensor Temp.: 101.56 Ambient Temp.: 96.82 F
18:13:47.599 -> Average Cal sensor Temp.: 101.43 Ambient Temp.: 96.75 F
18:13:49.147 -> Average Cal sensor Temp.: 101.65 Ambient Temp.: 96.82 F
18:13:50.735 -> Average Cal sensor Temp.: 101.19 Ambient Temp.: 96.78 F
18:13:52.291 -> Average Cal sensor Temp.: 101.22 Ambient Temp.: 96.71 F
18:13:53.843 -> Average Cal sensor Temp.: 101.55 Ambient Temp.: 96.78 F
18:13:55.423 -> Average Cal sensor Temp.: 101.91 Ambient Temp.: 96.78 F
18:13:56.980 -> Average Cal sensor Temp.: 101.26 Ambient Temp.: 96.75 F
18:13:58.529 -> Average Cal sensor Temp.: 101.19 Ambient Temp.: 96.75 F
18:14:00.083 -> Average Cal sensor Temp.: 101.57 Ambient Temp.: 96.78 F
18:14:01.670 -> Average Cal sensor Temp.: 101.35 Ambient Temp.: 96.78 F
18:14:03.225 -> Average Cal sensor Temp.: 101.52 Ambient Temp.: 96.75 F
18:14:04.778 -> Average Cal sensor Temp.: 101.22 Ambient Temp.: 96.89 F
18:14:06.357 -> Average Cal sensor Temp.: 101.34 Ambient Temp.: 96.78 F
18:14:07.910 -> Average Cal sensor Temp.: 101.52 Ambient Temp.: 96.78 F
18:14:09.497 -> Average Cal sensor Temp.: 101.20 Ambient Temp.: 96.82 F
18:14:11.047 -> Average Cal sensor Temp.: 101.28 Ambient Temp.: 96.78 F
18:14:12.600 -> Average Cal sensor Temp.: 101.45 Ambient Temp.: 96.78 F
18:14:14.187 -> Average Cal sensor Temp.: 101.70 Ambient Temp.: 96.82 F
18:14:15.744 -> Average Cal sensor Temp.: 100.49 Ambient Temp.: 96.78 F
18:14:17.295 -> Average Cal sensor Temp.: 101.27 Ambient Temp.: 96.82 F
18:14:18.883 -> Average Cal sensor Temp.: 100.75 Ambient Temp.: 96.75 F
18:14:20.434 -> Average Cal sensor Temp.: 101.46 Ambient Temp.: 96.78 F
18:14:21.986 -> Average Cal sensor Temp.: 101.19 Ambient Temp.: 96.82 F
18:14:23.572 -> Average Cal sensor Temp.: 101.29 Ambient Temp.: 96.78 F
18:14:25.122 -> Average Cal sensor Temp.: 101.27 Ambient Temp.: 96.82 F
18:14:26.707 -> Average Cal sensor Temp.: 101.76 Ambient Temp.: 96.78 F
18:14:28.262 -> Average Cal sensor Temp.: 101.48 Ambient Temp.: 96.82 F
18:14:29.816 -> Average Cal sensor Temp.: 101.84 Ambient Temp.: 96.82 F
18:14:31.368 -> Average Cal sensor Temp.: 102.08 Ambient Temp.: 96.82 F
18:14:32.951 -> Average Cal sensor Temp.: 101.57 Ambient Temp.: 96.82 F
18:14:34.506 -> Average Cal sensor Temp.: 101.49 Ambient Temp.: 96.82 F
18:14:36.082 -> Average Cal sensor Temp.: 101.67 Ambient Temp.: 96.78 F
18:14:37.639 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 96.78 F
18:14:39.221 -> Average Cal sensor Temp.: 101.16 Ambient Temp.: 96.78 F
18:14:40.773 -> Average Cal sensor Temp.: 101.34 Ambient Temp.: 96.93 F
18:14:42.325 -> Average Cal sensor Temp.: 101.69 Ambient Temp.: 96.82 F
18:14:43.908 -> Average Cal sensor Temp.: 102.01 Ambient Temp.: 96.89 F
18:14:45.460 -> Average Cal sensor Temp.: 101.93 Ambient Temp.: 96.93 F
18:14:47.013 -> Average Cal sensor Temp.: 101.68 Ambient Temp.: 96.82 F
18:14:48.599 -> Average Cal sensor Temp.: 101.51 Ambient Temp.: 96.89 F
18:14:50.159 -> Average Cal sensor Temp.: 101.77 Ambient Temp.: 96.82 F
18:14:51.712 -> Average Cal sensor Temp.: 101.80 Ambient Temp.: 96.93 F
18:14:53.300 -> Average Cal sensor Temp.: 101.94 Ambient Temp.: 96.89 F
18:14:54.852 -> Average Cal sensor Temp.: 101.79 Ambient Temp.: 96.89 F
18:14:56.413 -> Average Cal sensor Temp.: 102.15 Ambient Temp.: 96.93 F
18:14:57.998 -> Average Cal sensor Temp.: 101.84 Ambient Temp.: 96.89 F
18:14:59.547 -> Average Cal sensor Temp.: 101.26 Ambient Temp.: 96.93 F
18:15:01.106 -> Average Cal sensor Temp.: 101.27 Ambient Temp.: 96.93 F
18:15:02.688 -> Average Cal sensor Temp.: 101.33 Ambient Temp.: 96.89 F
18:15:04.242 -> Average Cal sensor Temp.: 101.48 Ambient Temp.: 96.93 F
18:15:05.792 -> Average Cal sensor Temp.: 101.71 Ambient Temp.: 96.93 F
18:15:07.379 -> Average Cal sensor Temp.: 101.71 Ambient Temp.: 96.96 F
18:15:08.933 -> Average Cal sensor Temp.: 101.57 Ambient Temp.: 96.89 F
18:15:10.490 -> Average Cal sensor Temp.: 101.73 Ambient Temp.: 96.93 F
18:15:12.042 -> Average Cal sensor Temp.: 101.86 Ambient Temp.: 96.96 F
18:15:13.635 -> Average Cal sensor Temp.: 101.45 Ambient Temp.: 96.93 F
18:15:15.191 -> Average Cal sensor Temp.: 101.74 Ambient Temp.: 96.96 F
18:15:16.738 -> Average Cal sensor Temp.: 101.82 Ambient Temp.: 96.93 F
18:15:18.329 -> Average Cal sensor Temp.: 102.06 Ambient Temp.: 96.96 F
18:15:19.880 -> Average Cal sensor Temp.: 102.00 Ambient Temp.: 96.96 F
18:15:21.434 -> Average Cal sensor Temp.: 101.41 Ambient Temp.: 96.96 F
18:15:23.026 -> Average Cal sensor Temp.: 101.63 Ambient Temp.: 96.96 F
18:15:24.573 -> Average Cal sensor Temp.: 101.50 Ambient Temp.: 96.93 F
18:15:26.157 -> Average Cal sensor Temp.: 101.74 Ambient Temp.: 96.96 F
18:15:27.716 -> Average Cal sensor Temp.: 101.99 Ambient Temp.: 96.96 F
18:15:29.271 -> Average Cal sensor Temp.: 101.71 Ambient Temp.: 97.00 F
18:15:30.822 -> Average Cal sensor Temp.: 101.66 Ambient Temp.: 96.93 F
18:15:32.411 -> Average Cal sensor Temp.: 101.43 Ambient Temp.: 97.00 F
18:15:33.962 -> Average Cal sensor Temp.: 101.70 Ambient Temp.: 97.00 F
18:15:35.515 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 97.03 F
18:15:37.102 -> Average Cal sensor Temp.: 100.97 Ambient Temp.: 96.96 F
18:15:38.656 -> Average Cal sensor Temp.: 101.28 Ambient Temp.: 96.96 F
18:15:40.213 -> Average Cal sensor Temp.: 101.49 Ambient Temp.: 96.96 F
18:15:41.798 -> Average Cal sensor Temp.: 101.59 Ambient Temp.: 97.03 F
18:15:43.354 -> Average Cal sensor Temp.: 101.22 Ambient Temp.: 97.03 F
18:15:44.903 -> Average Cal sensor Temp.: 101.62 Ambient Temp.: 97.03 F
18:15:46.492 -> Average Cal sensor Temp.: 101.88 Ambient Temp.: 97.00 F
18:15:48.045 -> Average Cal sensor Temp.: 101.80 Ambient Temp.: 97.03 F
18:15:49.602 -> Average Cal sensor Temp.: 101.59 Ambient Temp.: 97.00 F
18:15:51.175 -> Average Cal sensor Temp.: 101.74 Ambient Temp.: 96.96 F
18:15:52.728 -> Average Cal sensor Temp.: 102.04 Ambient Temp.: 96.96 F
18:15:54.310 -> Average Cal sensor Temp.: 101.92 Ambient Temp.: 96.96 F
18:15:55.869 -> Average Cal sensor Temp.: 102.05 Ambient Temp.: 97.00 F
18:15:57.425 -> Average Cal sensor Temp.: 101.79 Ambient Temp.: 97.00 F
18:15:59.008 -> Average Cal sensor Temp.: 101.75 Ambient Temp.: 97.00 F
18:16:00.558 -> Average Cal sensor Temp.: 101.68 Ambient Temp.: 97.00 F
18:16:02.110 -> Average Cal sensor Temp.: 101.32 Ambient Temp.: 96.96 F
18:16:03.698 -> Average Cal sensor Temp.: 101.38 Ambient Temp.: 97.00 F
18:16:05.252 -> Average Cal sensor Temp.: 101.41 Ambient Temp.: 97.00 F
18:16:06.807 -> Average Cal sensor Temp.: 101.38 Ambient Temp.: 97.00 F
18:16:08.394 -> Average Cal sensor Temp.: 101.34 Ambient Temp.: 97.00 F
18:16:09.944 -> Average Cal sensor Temp.: 101.74 Ambient Temp.: 96.93 F
18:16:11.494 -> Average Cal sensor Temp.: 101.86 Ambient Temp.: 97.00 F
18:16:13.084 -> Average Cal sensor Temp.: 101.43 Ambient Temp.: 97.00 F
18:16:14.631 -> Average Cal sensor Temp.: 101.70 Ambient Temp.: 97.00 F
18:16:16.215 -> Average Cal sensor Temp.: 101.79 Ambient Temp.: 97.00 F
18:16:17.763 -> Average Cal sensor Temp.: 101.78 Ambient Temp.: 97.03 F
18:16:19.315 -> Average Cal sensor Temp.: 101.77 Ambient Temp.: 97.00 F
18:16:20.905 -> Average Cal sensor Temp.: 102.09 Ambient Temp.: 96.96 F
18:16:22.455 -> Average Cal sensor Temp.: 102.14 Ambient Temp.: 96.96 F
18:16:24.039 -> Average Cal sensor Temp.: 101.91 Ambient Temp.: 97.00 F
18:16:25.581 -> Average Cal sensor Temp.: 101.75 Ambient Temp.: 97.00 F
18:16:27.136 -> Average Cal sensor Temp.: 101.46 Ambient Temp.: 97.03 F
18:16:28.729 -> Average Cal sensor Temp.: 101.75 Ambient Temp.: 97.03 F
18:16:30.286 -> Average Cal sensor Temp.: 101.85 Ambient Temp.: 96.96 F
18:16:31.837 -> Average Cal sensor Temp.: 101.90 Ambient Temp.: 96.96 F
18:16:33.394 -> Average Cal sensor Temp.: 101.41 Ambient Temp.: 97.03 F
18:16:34.979 -> Average Cal sensor Temp.: 101.28 Ambient Temp.: 97.00 F
18:16:36.537 -> Average Cal sensor Temp.: 101.89 Ambient Temp.: 97.00 F
18:16:38.085 -> Average Cal sensor Temp.: 101.97 Ambient Temp.: 97.00 F
18:16:39.678 -> Average Cal sensor Temp.: 101.72 Ambient Temp.: 97.00 F
18:16:41.230 -> Average Cal sensor Temp.: 102.09 Ambient Temp.: 97.00 F
18:16:42.779 -> Average Cal sensor Temp.: 101.81 Ambient Temp.: 97.00 F
18:16:44.366 -> Average Cal sensor Temp.: 102.06 Ambient Temp.: 97.03 F
18:16:45.907 -> Average Cal sensor Temp.: 102.17 Ambient Temp.: 97.03 F
18:16:47.497 -> Average Cal sensor Temp.: 101.36 Ambient Temp.: 96.96 F
18:16:49.050 -> Average Cal sensor Temp.: 101.72 Ambient Temp.: 97.00 F
18:16:50.601 -> Average Cal sensor Temp.: 101.93 Ambient Temp.: 97.00 F
18:16:52.187 -> Average Cal sensor Temp.: 101.41 Ambient Temp.: 97.00 F
18:16:53.738 -> Average Cal sensor Temp.: 101.36 Ambient Temp.: 97.00 F
18:16:55.294 -> Average Cal sensor Temp.: 101.65 Ambient Temp.: 97.00 F
18:16:56.878 -> Average Cal sensor Temp.: 101.90 Ambient Temp.: 96.96 F
18:16:58.433 -> Average Cal sensor Temp.: 101.28 Ambient Temp.: 97.00 F
18:17:00.013 -> Average Cal sensor Temp.: 101.72 Ambient Temp.: 97.00 F
18:17:01.563 -> Average Cal sensor Temp.: 101.24 Ambient Temp.: 97.00 F
18:17:03.119 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 97.00 F
18:17:04.703 -> Average Cal sensor Temp.: 101.65 Ambient Temp.: 97.00 F
18:17:06.256 -> Average Cal sensor Temp.: 101.80 Ambient Temp.: 97.03 F
18:17:07.804 -> Average Cal sensor Temp.: 101.84 Ambient Temp.: 97.00 F
18:17:09.395 -> Average Cal sensor Temp.: 102.11 Ambient Temp.: 97.00 F
18:17:10.954 -> Average Cal sensor Temp.: 101.25 Ambient Temp.: 97.00 F
18:17:12.509 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 97.03 F
18:17:14.067 -> Average Cal sensor Temp.: 101.45 Ambient Temp.: 97.00 F
18:17:15.617 -> Average Cal sensor Temp.: 102.07 Ambient Temp.: 97.00 F
18:17:17.204 -> Average Cal sensor Temp.: 101.68 Ambient Temp.: 97.00 F
18:17:18.756 -> Average Cal sensor Temp.: 102.05 Ambient Temp.: 97.00 F
18:17:20.341 -> Average Cal sensor Temp.: 101.99 Ambient Temp.: 97.00 F
18:17:21.893 -> Average Cal sensor Temp.: 102.25 Ambient Temp.: 97.00 F
18:17:23.451 -> Average Cal sensor Temp.: 101.97 Ambient Temp.: 97.00 F
18:17:25.032 -> Average Cal sensor Temp.: 102.01 Ambient Temp.: 97.00 F
18:17:26.587 -> Average Cal sensor Temp.: 101.61 Ambient Temp.: 97.00 F
18:17:28.138 -> Average Cal sensor Temp.: 101.75 Ambient Temp.: 97.00 F
18:17:29.728 -> Average Cal sensor Temp.: 101.59 Ambient Temp.: 97.03 F
18:17:31.286 -> Average Cal sensor Temp.: 101.80 Ambient Temp.: 97.00 F
18:17:32.840 -> Average Cal sensor Temp.: 101.66 Ambient Temp.: 97.00 F
18:17:34.390 -> Average Cal sensor Temp.: 101.49 Ambient Temp.: 97.03 F
18:17:35.975 -> Average Cal sensor Temp.: 101.38 Ambient Temp.: 97.03 F
18:17:37.528 -> Average Cal sensor Temp.: 101.49 Ambient Temp.: 97.00 F
18:17:39.113 -> Average Cal sensor Temp.: 101.78 Ambient Temp.: 96.96 F
18:17:40.666 -> Average Cal sensor Temp.: 101.51 Ambient Temp.: 97.00 F
18:17:42.215 -> Average Cal sensor Temp.: 101.39 Ambient Temp.: 97.00 F
18:17:43.798 -> Average Cal sensor Temp.: 101.74 Ambient Temp.: 97.03 F
18:17:45.348 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 96.96 F
18:17:46.903 -> Average Cal sensor Temp.: 101.26 Ambient Temp.: 97.00 F
18:17:48.496 -> Average Cal sensor Temp.: 101.43 Ambient Temp.: 97.00 F
18:17:50.054 -> Average Cal sensor Temp.: 101.49 Ambient Temp.: 97.03 F
18:17:51.608 -> Average Cal sensor Temp.: 101.92 Ambient Temp.: 96.96 F
18:17:53.162 -> Average Cal sensor Temp.: 102.02 Ambient Temp.: 97.03 F
18:17:54.749 -> Average Cal sensor Temp.: 102.18 Ambient Temp.: 97.03 F
18:17:56.314 -> Average Cal sensor Temp.: 102.25 Ambient Temp.: 97.07 F
18:17:57.869 -> Average Cal sensor Temp.: 102.16 Ambient Temp.: 97.03 F
18:17:59.429 -> Average Cal sensor Temp.: 101.99 Ambient Temp.: 97.03 F
18:18:00.981 -> Average Cal sensor Temp.: 102.09 Ambient Temp.: 97.00 F
18:18:02.570 -> Average Cal sensor Temp.: 101.61 Ambient Temp.: 97.00 F
18:18:04.122 -> Average Cal sensor Temp.: 101.96 Ambient Temp.: 97.07 F
18:18:05.680 -> Average Cal sensor Temp.: 101.96 Ambient Temp.: 97.03 F
18:18:07.234 -> Average Cal sensor Temp.: 101.97 Ambient Temp.: 97.03 F
18:18:08.828 -> Average Cal sensor Temp.: 102.20 Ambient Temp.: 97.03 F
18:18:10.386 -> Average Cal sensor Temp.: 102.30 Ambient Temp.: 97.03 F
18:18:11.941 -> Average Cal sensor Temp.: 102.14 Ambient Temp.: 97.07 F
18:18:13.500 -> Average Cal sensor Temp.: 101.73 Ambient Temp.: 97.07 F
18:18:15.052 -> Average Cal sensor Temp.: 101.66 Ambient Temp.: 97.00 F
18:18:16.638 -> Average Cal sensor Temp.: 102.19 Ambient Temp.: 97.07 F
18:18:18.197 -> Average Cal sensor Temp.: 101.96 Ambient Temp.: 97.03 F
18:18:19.753 -> Average Cal sensor Temp.: 101.93 Ambient Temp.: 97.07 F
18:18:21.311 -> Average Cal sensor Temp.: 101.86 Ambient Temp.: 97.11 F
18:18:22.902 -> Average Cal sensor Temp.: 101.82 Ambient Temp.: 97.03 F
18:18:24.448 -> Average Cal sensor Temp.: 102.02 Ambient Temp.: 97.03 F
18:18:26.003 -> Average Cal sensor Temp.: 102.19 Ambient Temp.: 97.11 F
18:18:27.566 -> Average Cal sensor Temp.: 101.93 Ambient Temp.: 97.00 F
18:18:29.154 -> Average Cal sensor Temp.: 101.12 Ambient Temp.: 97.03 F
18:18:30.707 -> Average Cal sensor Temp.: 101.73 Ambient Temp.: 97.03 F
18:18:32.260 -> Average Cal sensor Temp.: 101.84 Ambient Temp.: 97.00 F
18:18:33.815 -> Average Cal sensor Temp.: 102.17 Ambient Temp.: 97.03 F
18:18:35.407 -> Average Cal sensor Temp.: 102.29 Ambient Temp.: 97.07 F
18:18:36.965 -> Average Cal sensor Temp.: 101.79 Ambient Temp.: 97.00 F
18:18:38.520 -> Average Cal sensor Temp.: 101.54 Ambient Temp.: 97.00 F
18:18:40.074 -> Average Cal sensor Temp.: 101.52 Ambient Temp.: 97.03 F
18:18:41.663 -> Average Cal sensor Temp.: 101.36 Ambient Temp.: 97.03 F
18:18:43.219 -> Average Cal sensor Temp.: 101.08 Ambient Temp.: 97.00 F
18:18:44.770 -> Average Cal sensor Temp.: 101.48 Ambient Temp.: 97.00 F
18:18:46.331 -> Average Cal sensor Temp.: 102.04 Ambient Temp.: 97.00 F
18:18:47.919 -> Average Cal sensor Temp.: 102.06 Ambient Temp.: 97.03 F
18:18:49.469 -> Average Cal sensor Temp.: 102.29 Ambient Temp.: 97.03 F
18:18:51.024 -> Average Cal sensor Temp.: 101.92 Ambient Temp.: 97.03 F
18:18:52.594 -> Average Cal sensor Temp.: 102.14 Ambient Temp.: 97.07 F
18:18:54.153 -> Average Cal sensor Temp.: 101.72 Ambient Temp.: 97.00 F
18:18:55.711 -> Average Cal sensor Temp.: 101.86 Ambient Temp.: 97.03 F
18:18:57.296 -> Average Cal sensor Temp.: 102.12 Ambient Temp.: 97.03 F
18:18:58.852 -> Average Cal sensor Temp.: 102.31 Ambient Temp.: 97.03 F
18:19:00.404 -> Average Cal sensor Temp.: 102.29 Ambient Temp.: 97.03 F
18:19:01.991 -> Average Cal sensor Temp.: 102.33 Ambient Temp.: 97.00 F
18:19:03.542 -> Average Cal sensor Temp.: 102.16 Ambient Temp.: 97.00 F
18:19:05.098 -> Average Cal sensor Temp.: 102.41 Ambient Temp.: 97.00 F
18:19:06.685 -> Average Cal sensor Temp.: 102.03 Ambient Temp.: 97.03 F
18:19:08.235 -> Average Cal sensor Temp.: 102.31 Ambient Temp.: 97.03 F
18:19:09.785 -> Average Cal sensor Temp.: 102.35 Ambient Temp.: 97.03 F
18:19:11.375 -> Average Cal sensor Temp.: 102.62 Ambient Temp.: 97.03 F
18:19:12.931 -> Average Cal sensor Temp.: 102.33 Ambient Temp.: 97.03 F
18:19:14.480 -> Average Cal sensor Temp.: 101.77 Ambient Temp.: 97.03 F
18:19:16.037 -> Average Cal sensor Temp.: 101.75 Ambient Temp.: 97.03 F
18:19:17.624 -> Average Cal sensor Temp.: 102.08 Ambient Temp.: 97.03 F
18:19:19.175 -> Average Cal sensor Temp.: 102.47 Ambient Temp.: 97.03 F
18:19:20.759 -> Average Cal sensor Temp.: 102.32 Ambient Temp.: 97.03 F
18:19:22.304 -> Average Cal sensor Temp.: 102.14 Ambient Temp.: 97.03 F
18:19:23.853 -> Average Cal sensor Temp.: 102.00 Ambient Temp.: 97.00 F
18:19:25.438 -> Average Cal sensor Temp.: 101.71 Ambient Temp.: 97.03 F
18:19:26.995 -> Average Cal sensor Temp.: 102.14 Ambient Temp.: 97.03 F
18:19:28.573 -> Average Cal sensor Temp.: 102.00 Ambient Temp.: 97.07 F
18:19:30.127 -> Average Cal sensor Temp.: 102.05 Ambient Temp.: 97.11 F
18:19:31.684 -> Average Cal sensor Temp.: 101.98 Ambient Temp.: 97.07 F
18:19:33.239 -> Average Cal sensor Temp.: 101.92 Ambient Temp.: 97.07 F
18:19:34.826 -> Average Cal sensor Temp.: 102.01 Ambient Temp.: 97.07 F
18:19:36.382 -> Average Cal sensor Temp.: 102.34 Ambient Temp.: 97.07 F
18:19:37.935 -> Average Cal sensor Temp.: 102.20 Ambient Temp.: 97.03 F
18:19:39.491 -> Average Cal sensor Temp.: 101.70 Ambient Temp.: 97.11 F
18:19:41.081 -> Average Cal sensor Temp.: 102.05 Ambient Temp.: 97.11 F
18:19:42.639 -> Average Cal sensor Temp.: 102.03 Ambient Temp.: 97.07 F
18:19:44.189 -> Average Cal sensor Temp.: 101.37 Ambient Temp.: 97.07 F
18:19:45.775 -> Average Cal sensor Temp.: 102.22 Ambient Temp.: 97.07 F
18:19:47.325 -> Average Cal sensor Temp.: 102.37 Ambient Temp.: 97.07 F
18:19:48.881 -> Average Cal sensor Temp.: 102.33 Ambient Temp.: 97.03 F
18:19:50.463 -> Average Cal sensor Temp.: 102.65 Ambient Temp.: 97.07 F
18:19:52.002 -> Average Cal sensor Temp.: 102.49 Ambient Temp.: 97.11 F
18:19:53.575 -> Average Cal sensor Temp.: 101.76 Ambient Temp.: 97.00 F
18:19:55.143 -> Average Cal sensor Temp.: 102.20 Ambient Temp.: 97.00 F
18:19:56.711 -> Average Cal sensor Temp.: 101.08 Ambient Temp.: 97.03 F
18:19:58.283 -> Average Cal sensor Temp.: 101.88 Ambient Temp.: 97.03 F
18:19:59.822 -> Average Cal sensor Temp.: 102.13 Ambient Temp.: 97.03 F
18:20:01.394 -> Average Cal sensor Temp.: 102.60 Ambient Temp.: 97.03 F
18:20:02.959 -> Average Cal sensor Temp.: 102.47 Ambient Temp.: 97.03 F
18:20:04.526 -> Average Cal sensor Temp.: 101.78 Ambient Temp.: 97.03 F
18:20:06.091 -> Average Cal sensor Temp.: 102.11 Ambient Temp.: 97.00 F
18:20:07.666 -> Average Cal sensor Temp.: 102.56 Ambient Temp.: 97.00 F
18:20:09.231 -> Average Cal sensor Temp.: 102.06 Ambient Temp.: 97.03 F
18:20:10.791 -> Average Cal sensor Temp.: 101.80 Ambient Temp.: 97.03 F
18:20:12.362 -> Average Cal sensor Temp.: 102.04 Ambient Temp.: 97.03 F
18:20:13.900 -> Average Cal sensor Temp.: 102.49 Ambient Temp.: 97.03 F
18:20:15.472 -> Average Cal sensor Temp.: 102.47 Ambient Temp.: 97.03 F
18:20:17.043 -> Average Cal sensor Temp.: 102.74 Ambient Temp.: 97.07 F
18:20:18.613 -> Average Cal sensor Temp.: 102.66 Ambient Temp.: 97.07 F
18:20:20.152 -> Average Cal sensor Temp.: 102.51 Ambient Temp.: 97.03 F
18:20:21.717 -> Average Cal sensor Temp.: 102.78 Ambient Temp.: 97.03 F
18:20:23.286 -> Average Cal sensor Temp.: 102.48 Ambient Temp.: 97.07 F
18:20:24.853 -> Average Cal sensor Temp.: 102.49 Ambient Temp.: 97.03 F
18:20:26.414 -> Average Cal sensor Temp.: 102.56 Ambient Temp.: 97.07 F
18:20:27.983 -> Average Cal sensor Temp.: 102.31 Ambient Temp.: 97.07 F
18:20:29.553 -> Average Cal sensor Temp.: 102.12 Ambient Temp.: 97.07 F
18:20:31.123 -> Average Cal sensor Temp.: 101.84 Ambient Temp.: 97.11 F
18:20:32.662 -> Average Cal sensor Temp.: 101.89 Ambient Temp.: 97.11 F
18:20:34.236 -> Average Cal sensor Temp.: 102.07 Ambient Temp.: 97.11 F
18:20:35.805 -> Average Cal sensor Temp.: 102.55 Ambient Temp.: 97.07 F
18:20:37.374 -> Average Cal sensor Temp.: 102.11 Ambient Temp.: 97.11 F
18:20:38.913 -> Average Cal sensor Temp.: 102.40 Ambient Temp.: 97.07 F
18:20:40.479 -> Average Cal sensor Temp.: 102.60 Ambient Temp.: 97.11 F
18:20:42.067 -> Average Cal sensor Temp.: 102.28 Ambient Temp.: 97.07 F
18:20:43.619 -> Average Cal sensor Temp.: 102.23 Ambient Temp.: 97.11 F
18:20:45.171 -> Average Cal sensor Temp.: 102.56 Ambient Temp.: 97.11 F
18:20:46.752 -> Average Cal sensor Temp.: 102.47 Ambient Temp.: 97.11 F
18:20:48.298 -> Average Cal sensor Temp.: 102.75 Ambient Temp.: 97.14 F
18:20:49.885 -> Average Cal sensor Temp.: 102.28 Ambient Temp.: 97.07 F
18:20:51.437 -> Average Cal sensor Temp.: 102.64 Ambient Temp.: 97.07 F
18:20:52.982 -> Average Cal sensor Temp.: 102.70 Ambient Temp.: 97.07 F
18:20:54.573 -> Average Cal sensor Temp.: 102.84 Ambient Temp.: 97.14 F
18:20:56.124 -> Average Cal sensor Temp.: 102.79 Ambient Temp.: 97.07 F
18:20:57.681 -> Average Cal sensor Temp.: 102.90 Ambient Temp.: 97.11 F
18:20:59.270 -> Average Cal sensor Temp.: 102.70 Ambient Temp.: 97.07 F
18:21:00.822 -> Average Cal sensor Temp.: 102.06 Ambient Temp.: 97.11 F
18:21:02.375 -> Average Cal sensor Temp.: 102.30 Ambient Temp.: 97.03 F
18:21:03.963 -> Average Cal sensor Temp.: 101.75 Ambient Temp.: 97.11 F
18:21:05.519 -> Average Cal sensor Temp.: 101.97 Ambient Temp.: 97.07 F
18:21:07.074 -> Average Cal sensor Temp.: 102.37 Ambient Temp.: 97.07 F
18:21:08.623 -> Average Cal sensor Temp.: 102.65 Ambient Temp.: 97.07 F
18:21:10.181 -> Average Cal sensor Temp.: 102.58 Ambient Temp.: 97.03 F
18:21:11.770 -> Average Cal sensor Temp.: 102.56 Ambient Temp.: 97.11 F
18:21:13.326 -> Average Cal sensor Temp.: 102.92 Ambient Temp.: 97.07 F
18:21:14.880 -> Average Cal sensor Temp.: 102.57 Ambient Temp.: 97.03 F
18:21:16.436 -> Average Cal sensor Temp.: 102.55 Ambient Temp.: 97.07 F
18:21:18.024 -> Average Cal sensor Temp.: 102.45 Ambient Temp.: 97.07 F
18:21:19.579 -> Average Cal sensor Temp.: 102.26 Ambient Temp.: 97.03 F
18:21:21.130 -> Average Cal sensor Temp.: 102.18 Ambient Temp.: 97.03 F
18:21:22.720 -> Average Cal sensor Temp.: 102.28 Ambient Temp.: 97.07 F
18:21:24.278 -> Average Cal sensor Temp.: 102.37 Ambient Temp.: 97.07 F
18:21:25.833 -> Average Cal sensor Temp.: 102.76 Ambient Temp.: 97.07 F
18:21:27.382 -> Average Cal sensor Temp.: 102.46 Ambient Temp.: 97.07 F
18:21:28.968 -> Average Cal sensor Temp.: 102.57 Ambient Temp.: 97.03 F
18:21:30.521 -> Average Cal sensor Temp.: 102.64 Ambient Temp.: 97.07 F
18:21:32.106 -> Average Cal sensor Temp.: 102.42 Ambient Temp.: 97.03 F
18:21:33.660 -> Average Cal sensor Temp.: 102.77 Ambient Temp.: 97.07 F
18:21:35.213 -> Average Cal sensor Temp.: 102.55 Ambient Temp.: 97.07 F
18:21:36.795 -> Average Cal sensor Temp.: 102.70 Ambient Temp.: 97.11 F
18:21:38.342 -> Average Cal sensor Temp.: 102.80 Ambient Temp.: 97.07 F
18:21:39.892 -> Average Cal sensor Temp.: 102.43 Ambient Temp.: 97.14 F
18:21:41.475 -> Average Cal sensor Temp.: 102.50 Ambient Temp.: 97.03 F
18:21:43.023 -> Average Cal sensor Temp.: 102.60 Ambient Temp.: 97.07 F
18:21:44.615 -> Average Cal sensor Temp.: 102.29 Ambient Temp.: 97.11 F
18:21:46.171 -> Average Cal sensor Temp.: 102.26 Ambient Temp.: 97.11 F
18:21:47.729 -> Average Cal sensor Temp.: 102.47 Ambient Temp.: 97.07 F
18:21:49.289 -> Average Cal sensor Temp.: 102.35 Ambient Temp.: 97.11 F
18:21:50.844 -> Average Cal sensor Temp.: 101.91 Ambient Temp.: 97.11 F
18:21:52.424 -> Average Cal sensor Temp.: 101.52 Ambient Temp.: 97.11 F
18:21:53.972 -> Average Cal sensor Temp.: 102.02 Ambient Temp.: 97.03 F
18:21:55.525 -> Average Cal sensor Temp.: 101.40 Ambient Temp.: 97.03 F
18:21:57.111 -> Average Cal sensor Temp.: 101.84 Ambient Temp.: 97.03 F
18:21:58.660 -> Average Cal sensor Temp.: 101.69 Ambient Temp.: 97.03 F
18:22:00.249 -> Average Cal sensor Temp.: 102.34 Ambient Temp.: 97.03 F
18:22:01.801 -> Average Cal sensor Temp.: 102.17 Ambient Temp.: 97.00 F
18:22:03.355 -> Average Cal sensor Temp.: 102.54 Ambient Temp.: 97.00 F
18:22:04.913 -> Average Cal sensor Temp.: 102.86 Ambient Temp.: 97.03 F
18:22:06.506 -> Average Cal sensor Temp.: 102.30 Ambient Temp.: 97.03 F
18:22:08.057 -> Average Cal sensor Temp.: 102.79 Ambient Temp.: 97.03 F
18:22:09.611 -> Average Cal sensor Temp.: 102.60 Ambient Temp.: 97.03 F
18:22:11.166 -> Average Cal sensor Temp.: 102.74 Ambient Temp.: 97.03 F
18:22:12.749 -> Average Cal sensor Temp.: 102.78 Ambient Temp.: 97.03 F
18:22:14.302 -> Average Cal sensor Temp.: 102.81 Ambient Temp.: 97.03 F
18:22:15.857 -> Average Cal sensor Temp.: 102.81 Ambient Temp.: 97.03 F
18:22:17.451 -> Average Cal sensor Temp.: 102.92 Ambient Temp.: 97.07 F
18:22:19.012 -> Average Cal sensor Temp.: 102.78 Ambient Temp.: 97.07 F
18:22:20.574 -> Average Cal sensor Temp.: 101.50 Ambient Temp.: 97.03 F
18:22:22.131 -> Average Cal sensor Temp.: 102.35 Ambient Temp.: 97.03 F
18:22:23.689 -> Average Cal sensor Temp.: 102.49 Ambient Temp.: 97.03 F
18:22:25.248 -> Average Cal sensor Temp.: 102.52 Ambient Temp.: 97.03 F
18:22:26.836 -> Average Cal sensor Temp.: 102.74 Ambient Temp.: 97.07 F
18:22:28.389 -> Average Cal sensor Temp.: 102.56 Ambient Temp.: 97.11 F
18:22:29.949 -> Average Cal sensor Temp.: 102.42 Ambient Temp.: 97.07 F
18:22:31.497 -> Average Cal sensor Temp.: 102.41 Ambient Temp.: 97.03 F
18:22:33.084 -> Average Cal sensor Temp.: 102.50 Ambient Temp.: 97.03 F
18:22:34.638 -> Average Cal sensor Temp.: 102.50 Ambient Temp.: 97.03 F
18:22:36.196 -> Average Cal sensor Temp.: 102.55 Ambient Temp.: 97.07 F
18:22:37.748 -> Average Cal sensor Temp.: 102.75 Ambient Temp.: 97.03 F
18:22:39.331 -> Average Cal sensor Temp.: 102.55 Ambient Temp.: 97.07 F
18:22:40.882 -> Average Cal sensor Temp.: 102.75 Ambient Temp.: 97.03 F
18:22:42.440 -> Average Cal sensor Temp.: 102.31 Ambient Temp.: 97.07 F
18:22:44.027 -> Average Cal sensor Temp.: 102.19 Ambient Temp.: 97.03 F
18:22:45.584 -> Average Cal sensor Temp.: 102.42 Ambient Temp.: 97.07 F
18:22:47.139 -> Average Cal sensor Temp.: 102.02 Ambient Temp.: 97.03 F
18:22:48.692 -> Average Cal sensor Temp.: 102.19 Ambient Temp.: 97.00 F
18:22:50.286 -> Average Cal sensor Temp.: 102.26 Ambient Temp.: 97.07 F
18:22:51.839 -> Average Cal sensor Temp.: 102.09 Ambient Temp.: 97.03 F
18:22:53.397 -> Average Cal sensor Temp.: 102.33 Ambient Temp.: 97.07 F
18:22:54.949 -> Average Cal sensor Temp.: 102.06 Ambient Temp.: 97.07 F
18:22:56.543 -> Average Cal sensor Temp.: 102.40 Ambient Temp.: 97.07 F
18:22:58.104 -> Average Cal sensor Temp.: 102.36 Ambient Temp.: 97.07 F
18:22:59.669 -> Average Cal sensor Temp.: 102.45 Ambient Temp.: 97.07 F
18:23:01.221 -> Average Cal sensor Temp.: 102.43 Ambient Temp.: 97.07 F
18:23:02.773 -> Average Cal sensor Temp.: 102.12 Ambient Temp.: 97.07 F
18:23:04.364 -> Average Cal sensor Temp.: 101.79 Ambient Temp.: 97.07 F
18:23:05.921 -> Average Cal sensor Temp.: 102.31 Ambient Temp.: 97.07 F
18:23:07.477 -> Average Cal sensor Temp.: 102.16 Ambient Temp.: 97.07 F
18:23:09.032 -> Average Cal sensor Temp.: 102.19 Ambient Temp.: 97.07 F
18:23:10.586 -> Average Cal sensor Temp.: 102.01 Ambient Temp.: 97.07 F
18:23:12.182 -> Average Cal sensor Temp.: 101.57 Ambient Temp.: 97.07 F
18:23:13.742 -> Average Cal sensor Temp.: 101.78 Ambient Temp.: 97.03 F
18:23:15.295 -> Average Cal sensor Temp.: 102.11 Ambient Temp.: 97.03 F
18:23:16.856 -> Average Cal sensor Temp.: 102.18 Ambient Temp.: 97.03 F
18:23:18.410 -> Average Cal sensor Temp.: 102.03 Ambient Temp.: 97.07 F
18:23:19.975 -> Average Cal sensor Temp.: 101.93 Ambient Temp.: 97.07 F
18:23:21.538 -> Average Cal sensor Temp.: 101.86 Ambient Temp.: 97.07 F
18:23:23.095 -> Average Cal sensor Temp.: 101.82 Ambient Temp.: 97.07 F
18:23:24.684 -> Average Cal sensor Temp.: 101.88 Ambient Temp.: 97.07 F
18:23:26.244 -> Average Cal sensor Temp.: 101.74 Ambient Temp.: 97.03 F
18:23:27.799 -> Average Cal sensor Temp.: 101.84 Ambient Temp.: 97.03 F
18:23:29.357 -> Average Cal sensor Temp.: 101.71 Ambient Temp.: 97.11 F
18:23:30.942 -> Average Cal sensor Temp.: 101.63 Ambient Temp.: 97.03 F
18:23:32.498 -> Average Cal sensor Temp.: 101.77 Ambient Temp.: 97.03 F
18:23:34.056 -> Average Cal sensor Temp.: 102.01 Ambient Temp.: 97.03 F
18:23:35.613 -> Average Cal sensor Temp.: 101.99 Ambient Temp.: 97.03 F
18:23:37.176 -> Average Cal sensor Temp.: 101.98 Ambient Temp.: 97.07 F
18:23:38.728 -> Average Cal sensor Temp.: 101.66 Ambient Temp.: 97.03 F
18:23:40.322 -> Average Cal sensor Temp.: 101.40 Ambient Temp.: 97.07 F
18:23:41.873 -> Average Cal sensor Temp.: 100.62 Ambient Temp.: 97.00 F
18:23:43.426 -> Average Cal sensor Temp.: 101.04 Ambient Temp.: 97.07 F
18:23:45.016 -> Average Cal sensor Temp.: 101.44 Ambient Temp.: 97.03 F
18:23:46.570 -> Average Cal sensor Temp.: 101.34 Ambient Temp.: 97.03 F
18:23:48.132 -> Average Cal sensor Temp.: 100.98 Ambient Temp.: 97.00 F
18:23:49.694 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 97.11 F
18:23:51.246 -> Average Cal sensor Temp.: 101.17 Ambient Temp.: 97.03 F
18:23:52.828 -> Average Cal sensor Temp.: 101.19 Ambient Temp.: 97.00 F
18:23:54.382 -> Average Cal sensor Temp.: 101.09 Ambient Temp.: 97.03 F
18:23:55.942 -> Average Cal sensor Temp.: 100.98 Ambient Temp.: 97.03 F
18:23:57.527 -> Average Cal sensor Temp.: 101.10 Ambient Temp.: 97.07 F
18:23:59.088 -> Average Cal sensor Temp.: 101.24 Ambient Temp.: 97.07 F
18:24:00.650 -> Average Cal sensor Temp.: 101.19 Ambient Temp.: 97.07 F
18:24:02.210 -> Average Cal sensor Temp.: 101.03 Ambient Temp.: 97.03 F
18:24:03.771 -> Average Cal sensor Temp.: 100.94 Ambient Temp.: 97.07 F
18:24:05.335 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 97.03 F
18:24:06.892 -> Average Cal sensor Temp.: 101.14 Ambient Temp.: 97.00 F
18:24:08.448 -> Average Cal sensor Temp.: 101.25 Ambient Temp.: 97.03 F
18:24:10.039 -> Average Cal sensor Temp.: 101.20 Ambient Temp.: 97.03 F
18:24:11.595 -> Average Cal sensor Temp.: 100.80 Ambient Temp.: 97.00 F
18:24:13.146 -> Average Cal sensor Temp.: 100.77 Ambient Temp.: 97.07 F
18:24:14.732 -> Average Cal sensor Temp.: 101.10 Ambient Temp.: 97.07 F
18:24:16.292 -> Average Cal sensor Temp.: 100.76 Ambient Temp.: 97.03 F
18:24:17.855 -> Average Cal sensor Temp.: 100.74 Ambient Temp.: 97.07 F
18:24:19.409 -> Average Cal sensor Temp.: 100.70 Ambient Temp.: 97.03 F
18:24:20.990 -> Average Cal sensor Temp.: 100.59 Ambient Temp.: 97.03 F
18:24:22.536 -> Average Cal sensor Temp.: 100.77 Ambient Temp.: 96.96 F
18:24:24.093 -> Average Cal sensor Temp.: 100.76 Ambient Temp.: 97.03 F
18:24:25.688 -> Average Cal sensor Temp.: 100.69 Ambient Temp.: 96.96 F
18:24:27.244 -> Average Cal sensor Temp.: 101.21 Ambient Temp.: 97.03 F
18:24:28.795 -> Average Cal sensor Temp.: 101.26 Ambient Temp.: 97.00 F
18:24:30.355 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 97.00 F
18:24:31.913 -> Average Cal sensor Temp.: 100.87 Ambient Temp.: 97.03 F
18:24:33.503 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 97.00 F
18:24:35.056 -> Average Cal sensor Temp.: 100.95 Ambient Temp.: 97.03 F
18:24:36.610 -> Average Cal sensor Temp.: 101.08 Ambient Temp.: 97.00 F
18:24:38.203 -> Average Cal sensor Temp.: 101.33 Ambient Temp.: 97.00 F
18:24:39.758 -> Average Cal sensor Temp.: 101.18 Ambient Temp.: 97.07 F
18:24:41.311 -> Average Cal sensor Temp.: 101.33 Ambient Temp.: 97.03 F
18:24:42.869 -> Average Cal sensor Temp.: 101.34 Ambient Temp.: 97.03 F
18:24:44.456 -> Average Cal sensor Temp.: 101.34 Ambient Temp.: 97.03 F
18:24:46.019 -> Average Cal sensor Temp.: 101.38 Ambient Temp.: 97.03 F
18:24:47.574 -> Average Cal sensor Temp.: 101.28 Ambient Temp.: 97.03 F
18:24:49.130 -> Average Cal sensor Temp.: 101.25 Ambient Temp.: 97.03 F
18:24:50.688 -> Average Cal sensor Temp.: 100.85 Ambient Temp.: 97.00 F
18:24:52.277 -> Average Cal sensor Temp.: 100.78 Ambient Temp.: 97.03 F
18:24:53.827 -> Average Cal sensor Temp.: 101.08 Ambient Temp.: 97.03 F
18:24:55.376 -> Average Cal sensor Temp.: 101.23 Ambient Temp.: 97.00 F
18:24:56.970 -> Average Cal sensor Temp.: 101.36 Ambient Temp.: 97.00 F
18:24:58.520 -> Average Cal sensor Temp.: 101.41 Ambient Temp.: 97.00 F
18:25:00.076 -> Average Cal sensor Temp.: 101.32 Ambient Temp.: 97.03 F
18:25:01.629 -> Average Cal sensor Temp.: 101.24 Ambient Temp.: 97.00 F
18:25:03.222 -> Average Cal sensor Temp.: 101.36 Ambient Temp.: 97.00 F
18:25:04.780 -> Average Cal sensor Temp.: 101.34 Ambient Temp.: 97.03 F
18:25:06.339 -> Average Cal sensor Temp.: 101.33 Ambient Temp.: 97.03 F
18:25:07.897 -> Average Cal sensor Temp.: 101.35 Ambient Temp.: 97.07 F
18:25:09.455 -> Average Cal sensor Temp.: 101.39 Ambient Temp.: 97.07 F
18:25:11.039 -> Average Cal sensor Temp.: 101.33 Ambient Temp.: 97.07 F
18:25:12.592 -> Average Cal sensor Temp.: 101.17 Ambient Temp.: 97.07 F
18:25:14.146 -> Average Cal sensor Temp.: 101.36 Ambient Temp.: 97.07 F
18:25:15.709 -> Average Cal sensor Temp.: 101.32 Ambient Temp.: 97.03 F
18:25:17.260 -> Average Cal sensor Temp.: 101.22 Ambient Temp.: 97.07 F
18:25:18.849 -> Average Cal sensor Temp.: 101.26 Ambient Temp.: 97.07 F
18:25:20.410 -> Average Cal sensor Temp.: 101.32 Ambient Temp.: 97.03 F
18:25:21.967 -> Average Cal sensor Temp.: 101.22 Ambient Temp.: 97.07 F
18:25:23.517 -> Average Cal sensor Temp.: 101.23 Ambient Temp.: 97.11 F
18:25:25.075 -> Average Cal sensor Temp.: 101.23 Ambient Temp.: 97.07 F
18:25:26.634 -> Average Cal sensor Temp.: 101.36 Ambient Temp.: 97.11 F
18:25:28.223 -> Average Cal sensor Temp.: 101.41 Ambient Temp.: 97.07 F
18:25:29.780 -> Average Cal sensor Temp.: 101.34 Ambient Temp.: 97.07 F
18:25:31.336 -> Average Cal sensor Temp.: 101.13 Ambient Temp.: 97.11 F
18:25:32.890 -> Average Cal sensor Temp.: 100.74 Ambient Temp.: 97.07 F
18:25:34.479 -> Average Cal sensor Temp.: 100.53 Ambient Temp.: 97.11 F
18:25:36.037 -> Average Cal sensor Temp.: 100.82 Ambient Temp.: 97.07 F
18:25:37.594 -> Average Cal sensor Temp.: 101.02 Ambient Temp.: 97.14 F
18:25:39.151 -> Average Cal sensor Temp.: 101.11 Ambient Temp.: 97.14 F
18:25:40.704 -> Average Cal sensor Temp.: 101.03 Ambient Temp.: 97.11 F
18:25:42.292 -> Average Cal sensor Temp.: 101.09 Ambient Temp.: 97.11 F
18:25:43.849 -> Average Cal sensor Temp.: 100.88 Ambient Temp.: 97.11 F
18:25:45.406 -> Average Cal sensor Temp.: 100.00 Ambient Temp.: 97.11 F
18:25:46.962 -> Average Cal sensor Temp.: 99.90 Ambient Temp.: 97.07 F
18:25:48.518 -> Average Cal sensor Temp.: 99.90 Ambient Temp.: 97.11 F
18:25:50.081 -> Average Cal sensor Temp.: 100.24 Ambient Temp.: 97.11 F
18:25:51.663 -> Average Cal sensor Temp.: 100.48 Ambient Temp.: 97.11 F
18:25:53.220 -> Average Cal sensor Temp.: 100.58 Ambient Temp.: 97.18 F
18:25:54.776 -> Average Cal sensor Temp.: 100.82 Ambient Temp.: 97.14 F
18:25:56.370 -> Average Cal sensor Temp.: 100.84 Ambient Temp.: 97.14 F
18:25:57.926 -> Average Cal sensor Temp.: 100.85 Ambient Temp.: 97.14 F
18:25:59.483 -> Average Cal sensor Temp.: 100.80 Ambient Temp.: 97.14 F
18:26:01.036 -> Average Cal sensor Temp.: 100.61 Ambient Temp.: 97.11 F
18:26:02.623 -> Average Cal sensor Temp.: 100.58 Ambient Temp.: 97.21 F
18:26:04.178 -> Average Cal sensor Temp.: 100.62 Ambient Temp.: 97.14 F
18:26:05.732 -> Average Cal sensor Temp.: 100.63 Ambient Temp.: 97.18 F
18:26:07.320 -> Average Cal sensor Temp.: 100.62 Ambient Temp.: 97.14 F
18:26:08.874 -> Average Cal sensor Temp.: 100.87 Ambient Temp.: 97.18 F
18:26:10.430 -> Average Cal sensor Temp.: 100.79 Ambient Temp.: 97.21 F
18:26:11.991 -> Average Cal sensor Temp.: 100.31 Ambient Temp.: 97.18 F
18:26:13.577 -> Average Cal sensor Temp.: 100.55 Ambient Temp.: 97.18 F
18:26:15.136 -> Average Cal sensor Temp.: 100.59 Ambient Temp.: 97.18 F
18:26:16.695 -> Average Cal sensor Temp.: 100.74 Ambient Temp.: 97.18 F
18:26:18.245 -> Average Cal sensor Temp.: 100.92 Ambient Temp.: 97.21 F
18:26:19.834 -> Average Cal sensor Temp.: 100.86 Ambient Temp.: 97.21 F
18:26:21.389 -> Average Cal sensor Temp.: 100.51 Ambient Temp.: 97.21 F
18:26:22.944 -> Average Cal sensor Temp.: 100.60 Ambient Temp.: 97.21 F
18:26:24.529 -> Average Cal sensor Temp.: 100.54 Ambient Temp.: 97.21 F
18:26:26.085 -> Average Cal sensor Temp.: 100.59 Ambient Temp.: 97.21 F
18:26:27.638 -> Average Cal sensor Temp.: 100.59 Ambient Temp.: 97.21 F
18:26:29.227 -> Average Cal sensor Temp.: 100.52 Ambient Temp.: 97.21 F
18:26:30.781 -> Average Cal sensor Temp.: 100.45 Ambient Temp.: 97.21 F
18:26:32.330 -> Average Cal sensor Temp.: 100.51 Ambient Temp.: 97.32 F
18:26:33.889 -> Average Cal sensor Temp.: 100.58 Ambient Temp.: 97.29 F
18:26:35.479 -> Average Cal sensor Temp.: 100.57 Ambient Temp.: 97.29 F
18:26:37.036 -> Average Cal sensor Temp.: 100.27 Ambient Temp.: 97.29 F
18:26:38.591 -> Average Cal sensor Temp.: 99.81 Ambient Temp.: 97.21 F
18:26:40.152 -> Average Cal sensor Temp.: 99.76 Ambient Temp.: 97.18 F
18:26:41.740 -> Average Cal sensor Temp.: 99.80 Ambient Temp.: 97.21 F
18:26:43.289 -> Average Cal sensor Temp.: 100.05 Ambient Temp.: 97.21 F
18:26:44.846 -> Average Cal sensor Temp.: 100.17 Ambient Temp.: 97.18 F
18:26:46.406 -> Average Cal sensor Temp.: 100.15 Ambient Temp.: 97.21 F
18:26:47.963 -> Average Cal sensor Temp.: 99.80 Ambient Temp.: 97.18 F
18:26:49.552 -> Average Cal sensor Temp.: 99.64 Ambient Temp.: 97.21 F
18:26:51.105 -> Average Cal sensor Temp.: 99.80 Ambient Temp.: 97.21 F
18:26:52.657 -> Average Cal sensor Temp.: 99.76 Ambient Temp.: 97.18 F
18:26:54.239 -> Average Cal sensor Temp.: 99.92 Ambient Temp.: 97.18 F
18:26:55.799 -> Average Cal sensor Temp.: 100.22 Ambient Temp.: 97.18 F
18:26:57.358 -> Average Cal sensor Temp.: 100.26 Ambient Temp.: 97.21 F
18:26:58.911 -> Average Cal sensor Temp.: 100.25 Ambient Temp.: 97.18 F
18:27:00.469 -> Average Cal sensor Temp.: 100.18 Ambient Temp.: 97.18 F
18:27:02.052 -> Average Cal sensor Temp.: 100.14 Ambient Temp.: 97.18 F
18:27:03.610 -> Average Cal sensor Temp.: 100.29 Ambient Temp.: 97.18 F
18:27:05.163 -> Average Cal sensor Temp.: 100.26 Ambient Temp.: 97.18 F
18:27:06.754 -> Average Cal sensor Temp.: 99.50 Ambient Temp.: 97.18 F
18:27:08.312 -> Average Cal sensor Temp.: 99.99 Ambient Temp.: 97.18 F
18:27:09.871 -> Average Cal sensor Temp.: 100.03 Ambient Temp.: 97.11 F
18:27:11.424 -> Average Cal sensor Temp.: 100.05 Ambient Temp.: 97.14 F
18:27:13.008 -> Average Cal sensor Temp.: 100.11 Ambient Temp.: 97.14 F
18:27:14.565 -> Average Cal sensor Temp.: 99.90 Ambient Temp.: 97.14 F
18:27:16.125 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 97.14 F
18:27:17.682 -> Average Cal sensor Temp.: 99.78 Ambient Temp.: 97.14 F
18:27:19.229 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 97.11 F
18:27:20.818 -> Average Cal sensor Temp.: 99.82 Ambient Temp.: 97.18 F
18:27:22.374 -> Average Cal sensor Temp.: 99.85 Ambient Temp.: 97.14 F
18:27:23.929 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 97.14 F
18:27:25.481 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 97.14 F
18:27:27.042 -> Average Cal sensor Temp.: 99.40 Ambient Temp.: 97.14 F
18:27:28.596 -> Average Cal sensor Temp.: 99.60 Ambient Temp.: 97.14 F
18:27:30.183 -> Average Cal sensor Temp.: 99.85 Ambient Temp.: 97.07 F
18:27:31.744 -> Average Cal sensor Temp.: 99.62 Ambient Temp.: 97.14 F
18:27:33.293 -> Average Cal sensor Temp.: 99.72 Ambient Temp.: 97.14 F
18:27:34.847 -> Average Cal sensor Temp.: 99.64 Ambient Temp.: 97.14 F
18:27:36.408 -> Average Cal sensor Temp.: 99.62 Ambient Temp.: 97.14 F
18:27:37.962 -> Average Cal sensor Temp.: 99.74 Ambient Temp.: 97.14 F
18:27:39.555 -> Average Cal sensor Temp.: 99.12 Ambient Temp.: 97.14 F
18:27:41.108 -> Average Cal sensor Temp.: 99.02 Ambient Temp.: 97.18 F
18:27:42.656 -> Average Cal sensor Temp.: 99.09 Ambient Temp.: 97.14 F
18:27:44.215 -> Average Cal sensor Temp.: 99.06 Ambient Temp.: 97.07 F
18:27:45.775 -> Average Cal sensor Temp.: 99.35 Ambient Temp.: 97.18 F
18:27:47.331 -> Average Cal sensor Temp.: 99.29 Ambient Temp.: 97.11 F
18:27:48.882 -> Average Cal sensor Temp.: 99.46 Ambient Temp.: 97.11 F
18:27:50.442 -> Average Cal sensor Temp.: 99.61 Ambient Temp.: 97.14 F
18:27:51.996 -> Average Cal sensor Temp.: 99.77 Ambient Temp.: 97.14 F
18:27:53.549 -> Average Cal sensor Temp.: 99.67 Ambient Temp.: 97.14 F
18:27:55.098 -> Average Cal sensor Temp.: 99.34 Ambient Temp.: 97.11 F
18:27:56.662 -> Average Cal sensor Temp.: 99.51 Ambient Temp.: 97.11 F
18:27:58.213 -> Average Cal sensor Temp.: 99.50 Ambient Temp.: 97.11 F
18:27:59.770 -> Average Cal sensor Temp.: 99.68 Ambient Temp.: 97.11 F
18:28:01.330 -> Average Cal sensor Temp.: 99.55 Ambient Temp.: 97.11 F
18:28:02.918 -> Average Cal sensor Temp.: 99.35 Ambient Temp.: 97.14 F
18:28:04.473 -> Average Cal sensor Temp.: 99.22 Ambient Temp.: 97.07 F
18:28:06.002 -> Average Cal sensor Temp.: 99.31 Ambient Temp.: 97.11 F
18:28:07.584 -> Average Cal sensor Temp.: 99.53 Ambient Temp.: 97.11 F
18:28:09.137 -> Average Cal sensor Temp.: 99.55 Ambient Temp.: 97.07 F
18:28:10.697 -> Average Cal sensor Temp.: 99.53 Ambient Temp.: 97.11 F
18:28:12.256 -> Average Cal sensor Temp.: 99.45 Ambient Temp.: 97.11 F
18:28:13.807 -> Average Cal sensor Temp.: 99.55 Ambient Temp.: 97.07 F
18:28:15.365 -> Average Cal sensor Temp.: 99.45 Ambient Temp.: 97.11 F
18:28:16.918 -> Average Cal sensor Temp.: 99.28 Ambient Temp.: 97.11 F
18:28:18.465 -> Average Cal sensor Temp.: 99.09 Ambient Temp.: 97.07 F
18:28:20.019 -> Average Cal sensor Temp.: 98.80 Ambient Temp.: 97.07 F
18:28:21.579 -> Average Cal sensor Temp.: 99.00 Ambient Temp.: 97.11 F
18:28:23.134 -> Average Cal sensor Temp.: 99.09 Ambient Temp.: 97.07 F
18:28:24.690 -> Average Cal sensor Temp.: 98.70 Ambient Temp.: 97.07 F
18:28:26.249 -> Average Cal sensor Temp.: 98.90 Ambient Temp.: 97.03 F
18:28:27.803 -> Average Cal sensor Temp.: 99.06 Ambient Temp.: 97.03 F
18:28:29.390 -> Average Cal sensor Temp.: 99.21 Ambient Temp.: 97.03 F
18:28:30.936 -> Average Cal sensor Temp.: 99.38 Ambient Temp.: 97.03 F
18:28:32.494 -> Average Cal sensor Temp.: 99.31 Ambient Temp.: 97.03 F
18:28:34.044 -> Average Cal sensor Temp.: 99.06 Ambient Temp.: 97.11 F
18:28:35.606 -> Average Cal sensor Temp.: 99.06 Ambient Temp.: 97.03 F
18:28:37.162 -> Average Cal sensor Temp.: 99.09 Ambient Temp.: 97.07 F
18:28:38.712 -> Average Cal sensor Temp.: 99.18 Ambient Temp.: 97.03 F
18:28:40.269 -> Average Cal sensor Temp.: 99.35 Ambient Temp.: 97.07 F
18:28:41.828 -> Average Cal sensor Temp.: 99.30 Ambient Temp.: 97.07 F
18:28:43.388 -> Average Cal sensor Temp.: 98.51 Ambient Temp.: 97.03 F
18:28:44.944 -> Average Cal sensor Temp.: 98.91 Ambient Temp.: 97.03 F
18:28:46.503 -> Average Cal sensor Temp.: 99.13 Ambient Temp.: 97.07 F
18:28:48.058 -> Average Cal sensor Temp.: 98.97 Ambient Temp.: 97.07 F
18:28:49.611 -> Average Cal sensor Temp.: 98.95 Ambient Temp.: 97.00 F
18:28:51.164 -> Average Cal sensor Temp.: 99.00 Ambient Temp.: 97.00 F
18:28:52.717 -> Average Cal sensor Temp.: 99.10 Ambient Temp.: 97.03 F
18:28:54.274 -> Average Cal sensor Temp.: 99.18 Ambient Temp.: 97.00 F
18:28:55.831 -> Average Cal sensor Temp.: 99.16 Ambient Temp.: 97.00 F
18:28:57.381 -> Average Cal sensor Temp.: 98.97 Ambient Temp.: 97.03 F
18:28:58.938 -> Average Cal sensor Temp.: 98.95 Ambient Temp.: 97.03 F
18:29:00.496 -> Average Cal sensor Temp.: 99.00 Ambient Temp.: 97.03 F
18:29:02.050 -> Average Cal sensor Temp.: 98.82 Ambient Temp.: 97.03 F
18:29:03.611 -> Average Cal sensor Temp.: 98.69 Ambient Temp.: 97.03 F
18:29:05.170 -> Average Cal sensor Temp.: 98.82 Ambient Temp.: 97.07 F
18:29:06.724 -> Average Cal sensor Temp.: 98.77 Ambient Temp.: 97.03 F
18:29:08.278 -> Average Cal sensor Temp.: 98.60 Ambient Temp.: 97.03 F
18:29:09.834 -> Average Cal sensor Temp.: 98.58 Ambient Temp.: 97.03 F
18:29:11.417 -> Average Cal sensor Temp.: 98.82 Ambient Temp.: 97.03 F
18:29:12.977 -> Average Cal sensor Temp.: 98.63 Ambient Temp.: 97.00 F
18:29:14.532 -> Average Cal sensor Temp.: 98.57 Ambient Temp.: 97.03 F
18:29:16.091 -> Average Cal sensor Temp.: 98.62 Ambient Temp.: 97.03 F
18:29:17.647 -> Average Cal sensor Temp.: 98.69 Ambient Temp.: 97.03 F
18:29:19.202 -> Average Cal sensor Temp.: 98.49 Ambient Temp.: 97.11 F
18:29:20.760 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 97.03 F
18:29:22.309 -> Average Cal sensor Temp.: 98.33 Ambient Temp.: 97.03 F
18:29:23.864 -> Average Cal sensor Temp.: 98.47 Ambient Temp.: 96.96 F
18:29:25.424 -> Average Cal sensor Temp.: 98.66 Ambient Temp.: 96.96 F
18:29:26.986 -> Average Cal sensor Temp.: 98.83 Ambient Temp.: 97.00 F
18:29:28.539 -> Average Cal sensor Temp.: 98.36 Ambient Temp.: 97.00 F
18:29:30.095 -> Average Cal sensor Temp.: 98.00 Ambient Temp.: 97.00 F
18:29:31.648 -> Average Cal sensor Temp.: 98.37 Ambient Temp.: 96.96 F
18:29:33.206 -> Average Cal sensor Temp.: 98.45 Ambient Temp.: 97.00 F
18:29:34.759 -> Average Cal sensor Temp.: 98.65 Ambient Temp.: 96.96 F
18:29:36.324 -> Average Cal sensor Temp.: 98.65 Ambient Temp.: 97.03 F
18:29:37.875 -> Average Cal sensor Temp.: 98.53 Ambient Temp.: 97.00 F
18:29:39.433 -> Average Cal sensor Temp.: 98.70 Ambient Temp.: 97.00 F
18:29:40.990 -> Average Cal sensor Temp.: 98.51 Ambient Temp.: 97.00 F
18:29:42.540 -> Average Cal sensor Temp.: 98.65 Ambient Temp.: 97.00 F
18:29:44.098 -> Average Cal sensor Temp.: 98.68 Ambient Temp.: 96.96 F
18:29:45.652 -> Average Cal sensor Temp.: 98.61 Ambient Temp.: 96.96 F
18:29:47.211 -> Average Cal sensor Temp.: 98.54 Ambient Temp.: 97.00 F
18:29:48.797 -> Average Cal sensor Temp.: 98.29 Ambient Temp.: 97.00 F
18:29:50.325 -> Average Cal sensor Temp.: 98.35 Ambient Temp.: 97.00 F
18:29:51.881 -> Average Cal sensor Temp.: 98.29 Ambient Temp.: 96.96 F
18:29:53.470 -> Average Cal sensor Temp.: 98.40 Ambient Temp.: 96.96 F
18:29:55.020 -> Average Cal sensor Temp.: 98.43 Ambient Temp.: 96.93 F
18:29:56.576 -> Average Cal sensor Temp.: 98.38 Ambient Temp.: 96.96 F
18:29:58.128 -> Average Cal sensor Temp.: 98.27 Ambient Temp.: 96.96 F
18:29:59.688 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 96.93 F
18:30:01.241 -> Average Cal sensor Temp.: 98.37 Ambient Temp.: 96.89 F
18:30:02.792 -> Average Cal sensor Temp.: 98.41 Ambient Temp.: 96.89 F
18:30:04.344 -> Average Cal sensor Temp.: 98.41 Ambient Temp.: 96.93 F
18:30:05.904 -> Average Cal sensor Temp.: 98.41 Ambient Temp.: 96.96 F
18:30:07.461 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 96.89 F
18:30:09.016 -> Average Cal sensor Temp.: 98.25 Ambient Temp.: 96.96 F
18:30:10.580 -> Average Cal sensor Temp.: 98.30 Ambient Temp.: 96.89 F
18:30:12.132 -> Average Cal sensor Temp.: 98.38 Ambient Temp.: 96.93 F
18:30:13.717 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 96.93 F
18:30:15.270 -> Average Cal sensor Temp.: 98.30 Ambient Temp.: 96.96 F
18:30:16.827 -> Average Cal sensor Temp.: 97.81 Ambient Temp.: 96.93 F
18:30:18.382 -> Average Cal sensor Temp.: 98.11 Ambient Temp.: 96.93 F
18:30:19.942 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 96.82 F
18:30:21.495 -> Average Cal sensor Temp.: 98.07 Ambient Temp.: 96.89 F
18:30:23.043 -> Average Cal sensor Temp.: 98.27 Ambient Temp.: 96.82 F
18:30:24.596 -> Average Cal sensor Temp.: 98.27 Ambient Temp.: 96.89 F
18:30:26.160 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 96.82 F
18:30:27.716 -> Average Cal sensor Temp.: 98.15 Ambient Temp.: 96.93 F
18:30:29.267 -> Average Cal sensor Temp.: 98.27 Ambient Temp.: 96.82 F
18:30:30.824 -> Average Cal sensor Temp.: 97.96 Ambient Temp.: 96.89 F
18:30:32.381 -> Average Cal sensor Temp.: 97.47 Ambient Temp.: 96.82 F
18:30:33.940 -> Average Cal sensor Temp.: 97.63 Ambient Temp.: 96.78 F
18:30:35.492 -> Average Cal sensor Temp.: 97.64 Ambient Temp.: 96.75 F
18:30:37.046 -> Average Cal sensor Temp.: 97.78 Ambient Temp.: 96.89 F
18:30:38.601 -> Average Cal sensor Temp.: 97.66 Ambient Temp.: 96.75 F
18:30:40.160 -> Average Cal sensor Temp.: 97.52 Ambient Temp.: 96.75 F
18:30:41.712 -> Average Cal sensor Temp.: 97.68 Ambient Temp.: 96.78 F
18:30:43.301 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 96.78 F
18:30:44.824 -> Average Cal sensor Temp.: 97.86 Ambient Temp.: 96.78 F
18:30:46.385 -> Average Cal sensor Temp.: 97.99 Ambient Temp.: 96.75 F
18:30:47.943 -> Average Cal sensor Temp.: 98.02 Ambient Temp.: 96.75 F
18:30:49.501 -> Average Cal sensor Temp.: 97.95 Ambient Temp.: 96.78 F
18:30:51.062 -> Average Cal sensor Temp.: 97.71 Ambient Temp.: 96.75 F
18:30:52.619 -> Average Cal sensor Temp.: 97.82 Ambient Temp.: 96.75 F
18:30:54.170 -> Average Cal sensor Temp.: 97.63 Ambient Temp.: 96.78 F
18:30:55.726 -> Average Cal sensor Temp.: 97.74 Ambient Temp.: 96.75 F
18:30:57.277 -> Average Cal sensor Temp.: 97.84 Ambient Temp.: 96.78 F
18:30:58.832 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 96.71 F
18:31:00.418 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 96.71 F
18:31:01.974 -> Average Cal sensor Temp.: 97.83 Ambient Temp.: 96.75 F
18:31:03.526 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 96.75 F
18:31:05.085 -> Average Cal sensor Temp.: 97.95 Ambient Temp.: 96.75 F
18:31:06.641 -> Average Cal sensor Temp.: 97.91 Ambient Temp.: 96.75 F
18:31:08.189 -> Average Cal sensor Temp.: 97.84 Ambient Temp.: 96.71 F
18:31:09.745 -> Average Cal sensor Temp.: 97.80 Ambient Temp.: 96.75 F
18:31:11.293 -> Average Cal sensor Temp.: 98.02 Ambient Temp.: 96.71 F
18:31:12.847 -> Average Cal sensor Temp.: 97.86 Ambient Temp.: 96.75 F
18:31:14.403 -> Average Cal sensor Temp.: 97.81 Ambient Temp.: 96.75 F
18:31:15.964 -> Average Cal sensor Temp.: 97.51 Ambient Temp.: 96.71 F
18:31:17.519 -> Average Cal sensor Temp.: 97.84 Ambient Temp.: 96.71 F
18:31:19.069 -> Average Cal sensor Temp.: 97.92 Ambient Temp.: 96.75 F
18:31:20.622 -> Average Cal sensor Temp.: 97.79 Ambient Temp.: 96.75 F
18:31:22.172 -> Average Cal sensor Temp.: 97.90 Ambient Temp.: 96.71 F
18:31:23.755 -> Average Cal sensor Temp.: 97.57 Ambient Temp.: 96.75 F
18:31:25.308 -> Average Cal sensor Temp.: 97.55 Ambient Temp.: 96.75 F
18:31:26.858 -> Average Cal sensor Temp.: 97.75 Ambient Temp.: 96.71 F
18:31:28.413 -> Average Cal sensor Temp.: 97.62 Ambient Temp.: 96.67 F
18:31:29.973 -> Average Cal sensor Temp.: 97.76 Ambient Temp.: 96.67 F
18:31:31.530 -> Average Cal sensor Temp.: 97.76 Ambient Temp.: 96.71 F
18:31:33.057 -> Average Cal sensor Temp.: 97.73 Ambient Temp.: 96.71 F
18:31:34.638 -> Average Cal sensor Temp.: 97.78 Ambient Temp.: 96.67 F
18:31:36.192 -> Average Cal sensor Temp.: 97.56 Ambient Temp.: 96.75 F
18:31:37.741 -> Average Cal sensor Temp.: 97.56 Ambient Temp.: 96.64 F
18:31:39.298 -> Average Cal sensor Temp.: 97.59 Ambient Temp.: 96.67 F
18:31:40.848 -> Average Cal sensor Temp.: 97.69 Ambient Temp.: 96.71 F
18:31:42.406 -> Average Cal sensor Temp.: 97.50 Ambient Temp.: 96.67 F
18:31:43.957 -> Average Cal sensor Temp.: 97.51 Ambient Temp.: 96.71 F
18:31:45.512 -> Average Cal sensor Temp.: 97.43 Ambient Temp.: 96.71 F
18:31:47.071 -> Average Cal sensor Temp.: 97.44 Ambient Temp.: 96.71 F
18:31:48.623 -> Average Cal sensor Temp.: 97.37 Ambient Temp.: 96.67 F
18:31:50.179 -> Average Cal sensor Temp.: 97.28 Ambient Temp.: 96.71 F
18:31:51.725 -> Average Cal sensor Temp.: 97.44 Ambient Temp.: 96.71 F
18:31:53.305 -> Average Cal sensor Temp.: 97.32 Ambient Temp.: 96.67 F
18:31:54.858 -> Average Cal sensor Temp.: 97.25 Ambient Temp.: 96.71 F
18:31:56.417 -> Average Cal sensor Temp.: 97.30 Ambient Temp.: 96.67 F
18:31:57.970 -> Average Cal sensor Temp.: 97.29 Ambient Temp.: 96.64 F
18:31:59.525 -> Average Cal sensor Temp.: 97.51 Ambient Temp.: 96.67 F
18:32:01.078 -> Average Cal sensor Temp.: 97.45 Ambient Temp.: 96.64 F
18:32:02.639 -> Average Cal sensor Temp.: 97.44 Ambient Temp.: 96.64 F
18:32:04.193 -> Average Cal sensor Temp.: 97.30 Ambient Temp.: 96.67 F
18:32:05.745 -> Average Cal sensor Temp.: 97.36 Ambient Temp.: 96.71 F
18:32:07.300 -> Average Cal sensor Temp.: 97.32 Ambient Temp.: 96.67 F
18:32:08.848 -> Average Cal sensor Temp.: 97.43 Ambient Temp.: 96.67 F
18:32:10.404 -> Average Cal sensor Temp.: 97.43 Ambient Temp.: 96.67 F
18:32:11.955 -> Average Cal sensor Temp.: 97.35 Ambient Temp.: 96.64 F
18:32:13.506 -> Average Cal sensor Temp.: 97.30 Ambient Temp.: 96.67 F
18:32:15.095 -> Average Cal sensor Temp.: 97.40 Ambient Temp.: 96.64 F
18:32:16.623 -> Average Cal sensor Temp.: 97.25 Ambient Temp.: 96.67 F
18:32:18.177 -> Average Cal sensor Temp.: 97.36 Ambient Temp.: 96.67 F
18:32:19.732 -> Average Cal sensor Temp.: 97.42 Ambient Temp.: 96.60 F
18:32:21.293 -> Average Cal sensor Temp.: 97.30 Ambient Temp.: 96.64 F
18:32:22.849 -> Average Cal sensor Temp.: 97.19 Ambient Temp.: 96.67 F
18:32:24.405 -> Average Cal sensor Temp.: 97.32 Ambient Temp.: 96.60 F
18:32:25.962 -> Average Cal sensor Temp.: 97.25 Ambient Temp.: 96.67 F
18:32:27.521 -> Average Cal sensor Temp.: 97.52 Ambient Temp.: 96.64 F
18:32:29.077 -> Average Cal sensor Temp.: 97.30 Ambient Temp.: 96.64 F
18:32:30.634 -> Average Cal sensor Temp.: 97.15 Ambient Temp.: 96.64 F
18:32:32.187 -> Average Cal sensor Temp.: 97.21 Ambient Temp.: 96.60 F
18:32:33.739 -> Average Cal sensor Temp.: 97.35 Ambient Temp.: 96.60 F
18:32:35.296 -> Average Cal sensor Temp.: 97.51 Ambient Temp.: 96.64 F
18:32:36.860 -> Average Cal sensor Temp.: 97.52 Ambient Temp.: 96.64 F
18:32:38.412 -> Average Cal sensor Temp.: 97.42 Ambient Temp.: 96.67 F
18:32:39.972 -> Average Cal sensor Temp.: 96.96 Ambient Temp.: 96.64 F
18:32:41.523 -> Average Cal sensor Temp.: 97.03 Ambient Temp.: 96.53 F
18:32:43.077 -> Average Cal sensor Temp.: 96.96 Ambient Temp.: 96.60 F
18:32:44.635 -> Average Cal sensor Temp.: 96.55 Ambient Temp.: 96.60 F
18:32:46.226 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 96.57 F
18:32:47.779 -> Average Cal sensor Temp.: 96.85 Ambient Temp.: 96.57 F
18:32:49.338 -> Average Cal sensor Temp.: 97.00 Ambient Temp.: 96.57 F
18:32:50.860 -> Average Cal sensor Temp.: 96.88 Ambient Temp.: 96.57 F
18:32:52.452 -> Average Cal sensor Temp.: 97.10 Ambient Temp.: 96.53 F
18:32:54.007 -> Average Cal sensor Temp.: 96.99 Ambient Temp.: 96.60 F
18:32:55.533 -> Average Cal sensor Temp.: 97.06 Ambient Temp.: 96.53 F
18:32:57.090 -> Average Cal sensor Temp.: 97.10 Ambient Temp.: 96.57 F
18:32:58.646 -> Average Cal sensor Temp.: 97.05 Ambient Temp.: 96.57 F
18:33:00.210 -> Average Cal sensor Temp.: 96.72 Ambient Temp.: 96.53 F
18:33:01.763 -> Average Cal sensor Temp.: 96.94 Ambient Temp.: 96.64 F
18:33:03.347 -> Average Cal sensor Temp.: 96.94 Ambient Temp.: 96.57 F
18:33:04.903 -> Average Cal sensor Temp.: 96.75 Ambient Temp.: 96.64 F
18:33:06.462 -> Average Cal sensor Temp.: 96.39 Ambient Temp.: 96.53 F
18:33:08.016 -> Average Cal sensor Temp.: 96.56 Ambient Temp.: 96.53 F
18:33:09.576 -> Average Cal sensor Temp.: 96.73 Ambient Temp.: 96.53 F
18:33:11.103 -> Average Cal sensor Temp.: 96.91 Ambient Temp.: 96.53 F
18:33:12.655 -> Average Cal sensor Temp.: 96.88 Ambient Temp.: 96.53 F
18:33:14.239 -> Average Cal sensor Temp.: 96.50 Ambient Temp.: 96.53 F
18:33:15.802 -> Average Cal sensor Temp.: 96.72 Ambient Temp.: 96.53 F
18:33:17.325 -> Average Cal sensor Temp.: 96.47 Ambient Temp.: 96.57 F
18:33:18.886 -> Average Cal sensor Temp.: 96.65 Ambient Temp.: 96.42 F
18:33:20.444 -> Average Cal sensor Temp.: 96.89 Ambient Temp.: 96.57 F
18:33:22.029 -> Average Cal sensor Temp.: 96.66 Ambient Temp.: 96.53 F
18:33:23.577 -> Average Cal sensor Temp.: 96.57 Ambient Temp.: 96.42 F
18:33:25.136 -> Average Cal sensor Temp.: 96.57 Ambient Temp.: 96.53 F
18:33:26.693 -> Average Cal sensor Temp.: 96.53 Ambient Temp.: 96.42 F
18:33:28.250 -> Average Cal sensor Temp.: 96.63 Ambient Temp.: 96.42 F
18:33:29.805 -> Average Cal sensor Temp.: 96.52 Ambient Temp.: 96.49 F
18:33:31.362 -> Average Cal sensor Temp.: 96.74 Ambient Temp.: 96.49 F
18:33:32.914 -> Average Cal sensor Temp.: 97.57 Ambient Temp.: 96.49 F
18:33:34.468 -> Average Cal sensor Temp.: 97.46 Ambient Temp.: 96.39 F
18:33:36.026 -> Average Cal sensor Temp.: 97.24 Ambient Temp.: 96.49 F
18:33:37.575 -> Average Cal sensor Temp.: 97.15 Ambient Temp.: 96.53 F
18:33:39.131 -> Average Cal sensor Temp.: 97.21 Ambient Temp.: 96.42 F
18:33:40.686 -> Average Cal sensor Temp.: 97.20 Ambient Temp.: 96.42 F
18:33:42.243 -> Average Cal sensor Temp.: 96.87 Ambient Temp.: 96.42 F
18:33:43.793 -> Average Cal sensor Temp.: 96.49 Ambient Temp.: 96.42 F
18:33:45.349 -> Average Cal sensor Temp.: 96.71 Ambient Temp.: 96.39 F
18:33:46.904 -> Average Cal sensor Temp.: 96.81 Ambient Temp.: 96.39 F
18:33:48.459 -> Average Cal sensor Temp.: 96.86 Ambient Temp.: 96.42 F
18:33:50.020 -> Average Cal sensor Temp.: 96.77 Ambient Temp.: 96.42 F
18:33:51.578 -> Average Cal sensor Temp.: 96.50 Ambient Temp.: 96.42 F
18:33:53.137 -> Average Cal sensor Temp.: 96.37 Ambient Temp.: 96.42 F
18:33:54.691 -> Average Cal sensor Temp.: 96.51 Ambient Temp.: 96.42 F
18:33:56.247 -> Average Cal sensor Temp.: 96.33 Ambient Temp.: 96.42 F
18:33:57.799 -> Average Cal sensor Temp.: 96.37 Ambient Temp.: 96.42 F
18:33:59.354 -> Average Cal sensor Temp.: 96.39 Ambient Temp.: 96.39 F
18:34:00.944 -> Average Cal sensor Temp.: 96.44 Ambient Temp.: 96.35 F
18:34:02.500 -> Average Cal sensor Temp.: 96.44 Ambient Temp.: 96.49 F
18:34:04.060 -> Average Cal sensor Temp.: 95.98 Ambient Temp.: 96.35 F
18:34:05.587 -> Average Cal sensor Temp.: 96.27 Ambient Temp.: 96.35 F
18:34:07.144 -> Average Cal sensor Temp.: 96.34 Ambient Temp.: 96.39 F
18:34:08.705 -> Average Cal sensor Temp.: 96.41 Ambient Temp.: 96.42 F
18:34:10.262 -> Average Cal sensor Temp.: 96.51 Ambient Temp.: 96.42 F
18:34:11.819 -> Average Cal sensor Temp.: 96.47 Ambient Temp.: 96.39 F
18:34:13.369 -> Average Cal sensor Temp.: 96.45 Ambient Temp.: 96.42 F
18:34:14.954 -> Average Cal sensor Temp.: 96.22 Ambient Temp.: 96.42 F
18:34:16.510 -> Average Cal sensor Temp.: 96.15 Ambient Temp.: 96.39 F
18:34:18.067 -> Average Cal sensor Temp.: 96.37 Ambient Temp.: 96.35 F
18:34:19.613 -> Average Cal sensor Temp.: 96.48 Ambient Temp.: 96.39 F
18:34:21.167 -> Average Cal sensor Temp.: 96.44 Ambient Temp.: 96.35 F
18:34:22.717 -> Average Cal sensor Temp.: 96.38 Ambient Temp.: 96.35 F
18:34:24.269 -> Average Cal sensor Temp.: 96.37 Ambient Temp.: 96.42 F
18:34:25.854 -> Average Cal sensor Temp.: 96.32 Ambient Temp.: 96.35 F
18:34:27.379 -> Average Cal sensor Temp.: 96.32 Ambient Temp.: 96.39 F
18:34:28.934 -> Average Cal sensor Temp.: 96.32 Ambient Temp.: 96.39 F
18:34:30.492 -> Average Cal sensor Temp.: 96.33 Ambient Temp.: 96.39 F
18:34:32.082 -> Average Cal sensor Temp.: 95.93 Ambient Temp.: 96.28 F
18:34:33.633 -> Average Cal sensor Temp.: 95.73 Ambient Temp.: 96.31 F
18:34:35.190 -> Average Cal sensor Temp.: 95.96 Ambient Temp.: 96.35 F
18:34:36.747 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 96.35 F
18:34:38.300 -> Average Cal sensor Temp.: 96.13 Ambient Temp.: 96.35 F
18:34:39.858 -> Average Cal sensor Temp.: 96.26 Ambient Temp.: 96.31 F
18:34:41.415 -> Average Cal sensor Temp.: 96.22 Ambient Temp.: 96.31 F
18:34:42.968 -> Average Cal sensor Temp.: 96.18 Ambient Temp.: 96.35 F
18:34:44.522 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 96.31 F
18:34:46.082 -> Average Cal sensor Temp.: 96.26 Ambient Temp.: 96.35 F
18:34:47.630 -> Average Cal sensor Temp.: 96.29 Ambient Temp.: 96.31 F
18:34:49.186 -> Average Cal sensor Temp.: 96.28 Ambient Temp.: 96.35 F
18:34:50.737 -> Average Cal sensor Temp.: 96.30 Ambient Temp.: 96.35 F
18:34:52.289 -> Average Cal sensor Temp.: 96.21 Ambient Temp.: 96.31 F
18:34:53.840 -> Average Cal sensor Temp.: 96.31 Ambient Temp.: 96.31 F
18:34:55.396 -> Average Cal sensor Temp.: 96.28 Ambient Temp.: 96.28 F
18:34:56.953 -> Average Cal sensor Temp.: 96.00 Ambient Temp.: 96.39 F
18:34:58.504 -> Average Cal sensor Temp.: 95.84 Ambient Temp.: 96.31 F
18:35:00.094 -> Average Cal sensor Temp.: 95.80 Ambient Temp.: 96.28 F
18:35:01.619 -> Average Cal sensor Temp.: 95.74 Ambient Temp.: 96.28 F
18:35:03.174 -> Average Cal sensor Temp.: 95.97 Ambient Temp.: 96.31 F
18:35:04.759 -> Average Cal sensor Temp.: 95.98 Ambient Temp.: 96.28 F
18:35:06.315 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 96.28 F
18:35:07.869 -> Average Cal sensor Temp.: 95.98 Ambient Temp.: 96.28 F
18:35:09.426 -> Average Cal sensor Temp.: 95.93 Ambient Temp.: 96.31 F
18:35:10.984 -> Average Cal sensor Temp.: 95.92 Ambient Temp.: 96.31 F
18:35:12.535 -> Average Cal sensor Temp.: 96.15 Ambient Temp.: 96.28 F
18:35:14.088 -> Average Cal sensor Temp.: 96.14 Ambient Temp.: 96.28 F
18:35:15.643 -> Average Cal sensor Temp.: 95.66 Ambient Temp.: 96.28 F
18:35:17.195 -> Average Cal sensor Temp.: 95.71 Ambient Temp.: 96.31 F
18:35:18.752 -> Average Cal sensor Temp.: 95.76 Ambient Temp.: 96.28 F
18:35:20.305 -> Average Cal sensor Temp.: 95.55 Ambient Temp.: 96.24 F
18:35:21.860 -> Average Cal sensor Temp.: 95.82 Ambient Temp.: 96.24 F
18:35:23.421 -> Average Cal sensor Temp.: 95.83 Ambient Temp.: 96.28 F
18:35:24.978 -> Average Cal sensor Temp.: 95.87 Ambient Temp.: 96.24 F
18:35:26.537 -> Average Cal sensor Temp.: 96.02 Ambient Temp.: 96.28 F
18:35:28.092 -> Average Cal sensor Temp.: 95.90 Ambient Temp.: 96.31 F
18:35:29.651 -> Average Cal sensor Temp.: 96.03 Ambient Temp.: 96.21 F
18:35:31.209 -> Average Cal sensor Temp.: 95.83 Ambient Temp.: 96.28 F
18:35:32.755 -> Average Cal sensor Temp.: 95.30 Ambient Temp.: 96.21 F
18:35:34.313 -> Average Cal sensor Temp.: 95.23 Ambient Temp.: 96.21 F
18:35:35.872 -> Average Cal sensor Temp.: 94.96 Ambient Temp.: 96.21 F
18:35:37.426 -> Average Cal sensor Temp.: 95.12 Ambient Temp.: 96.24 F
18:35:38.981 -> Average Cal sensor Temp.: 95.17 Ambient Temp.: 96.24 F
18:35:40.537 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 96.21 F
18:35:42.091 -> Average Cal sensor Temp.: 95.76 Ambient Temp.: 96.21 F
18:35:43.645 -> Average Cal sensor Temp.: 95.77 Ambient Temp.: 96.24 F
18:35:45.199 -> Average Cal sensor Temp.: 95.49 Ambient Temp.: 96.17 F
18:35:46.752 -> Average Cal sensor Temp.: 95.35 Ambient Temp.: 96.21 F
18:35:48.306 -> Average Cal sensor Temp.: 95.38 Ambient Temp.: 96.21 F
18:35:49.859 -> Average Cal sensor Temp.: 95.19 Ambient Temp.: 96.17 F
18:35:51.410 -> Average Cal sensor Temp.: 95.64 Ambient Temp.: 96.17 F
18:35:52.998 -> Average Cal sensor Temp.: 95.78 Ambient Temp.: 96.17 F
18:35:54.522 -> Average Cal sensor Temp.: 95.50 Ambient Temp.: 96.17 F
18:35:56.078 -> Average Cal sensor Temp.: 95.55 Ambient Temp.: 96.24 F
18:35:57.666 -> Average Cal sensor Temp.: 95.73 Ambient Temp.: 96.21 F
18:35:59.218 -> Average Cal sensor Temp.: 95.70 Ambient Temp.: 96.17 F
18:36:00.775 -> Average Cal sensor Temp.: 95.64 Ambient Temp.: 96.21 F
18:36:02.328 -> Average Cal sensor Temp.: 95.57 Ambient Temp.: 96.21 F
18:36:03.886 -> Average Cal sensor Temp.: 95.49 Ambient Temp.: 96.13 F
18:36:05.409 -> Average Cal sensor Temp.: 95.67 Ambient Temp.: 96.17 F
18:36:06.995 -> Average Cal sensor Temp.: 95.58 Ambient Temp.: 96.17 F
18:36:08.520 -> Average Cal sensor Temp.: 95.60 Ambient Temp.: 96.10 F
18:36:10.078 -> Average Cal sensor Temp.: 95.74 Ambient Temp.: 96.13 F
18:36:11.632 -> Average Cal sensor Temp.: 95.68 Ambient Temp.: 96.13 F
18:36:13.192 -> Average Cal sensor Temp.: 95.69 Ambient Temp.: 96.13 F
18:36:14.749 -> Average Cal sensor Temp.: 95.78 Ambient Temp.: 96.17 F
18:36:16.333 -> Average Cal sensor Temp.: 95.80 Ambient Temp.: 96.13 F
18:36:17.886 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 96.13 F
18:36:19.443 -> Average Cal sensor Temp.: 95.74 Ambient Temp.: 96.13 F
18:36:21.001 -> Average Cal sensor Temp.: 95.73 Ambient Temp.: 96.21 F
18:36:22.552 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 96.17 F
18:36:24.102 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 96.17 F
18:36:25.666 -> Average Cal sensor Temp.: 95.59 Ambient Temp.: 96.17 F
18:36:27.219 -> Average Cal sensor Temp.: 95.54 Ambient Temp.: 96.13 F
18:36:28.772 -> Average Cal sensor Temp.: 95.62 Ambient Temp.: 96.13 F
18:36:30.300 -> Average Cal sensor Temp.: 95.63 Ambient Temp.: 96.13 F
18:36:31.887 -> Average Cal sensor Temp.: 95.59 Ambient Temp.: 96.13 F
18:36:33.442 -> Average Cal sensor Temp.: 95.54 Ambient Temp.: 96.13 F
18:36:34.996 -> Average Cal sensor Temp.: 95.50 Ambient Temp.: 96.17 F
18:36:36.554 -> Average Cal sensor Temp.: 95.57 Ambient Temp.: 96.17 F
18:36:38.079 -> Average Cal sensor Temp.: 95.55 Ambient Temp.: 96.13 F
18:36:39.631 -> Average Cal sensor Temp.: 95.62 Ambient Temp.: 96.17 F
18:36:41.184 -> Average Cal sensor Temp.: 95.69 Ambient Temp.: 96.17 F
18:36:42.741 -> Average Cal sensor Temp.: 95.41 Ambient Temp.: 96.13 F
18:36:44.325 -> Average Cal sensor Temp.: 95.37 Ambient Temp.: 96.21 F
18:36:45.881 -> Average Cal sensor Temp.: 95.34 Ambient Temp.: 96.17 F
18:36:47.438 -> Average Cal sensor Temp.: 95.00 Ambient Temp.: 96.17 F
18:36:48.989 -> Average Cal sensor Temp.: 94.85 Ambient Temp.: 96.13 F
18:36:50.547 -> Average Cal sensor Temp.: 95.22 Ambient Temp.: 96.13 F
18:36:52.103 -> Average Cal sensor Temp.: 95.36 Ambient Temp.: 96.17 F
18:36:53.662 -> Average Cal sensor Temp.: 95.29 Ambient Temp.: 96.10 F
18:36:55.188 -> Average Cal sensor Temp.: 95.43 Ambient Temp.: 96.13 F
18:36:56.749 -> Average Cal sensor Temp.: 95.44 Ambient Temp.: 96.13 F
18:36:58.311 -> Average Cal sensor Temp.: 95.45 Ambient Temp.: 96.13 F
18:36:59.875 -> Average Cal sensor Temp.: 95.44 Ambient Temp.: 96.13 F
18:37:01.437 -> Average Cal sensor Temp.: 95.56 Ambient Temp.: 96.17 F
18:37:02.986 -> Average Cal sensor Temp.: 95.49 Ambient Temp.: 96.10 F
18:37:04.538 -> Average Cal sensor Temp.: 95.57 Ambient Temp.: 96.13 F
18:37:06.096 -> Average Cal sensor Temp.: 95.62 Ambient Temp.: 96.17 F
18:37:07.652 -> Average Cal sensor Temp.: 95.68 Ambient Temp.: 96.13 F
18:37:09.205 -> Average Cal sensor Temp.: 95.58 Ambient Temp.: 96.10 F
18:37:10.763 -> Average Cal sensor Temp.: 95.54 Ambient Temp.: 96.10 F
18:37:12.318 -> Average Cal sensor Temp.: 95.61 Ambient Temp.: 96.13 F
18:37:13.876 -> Average Cal sensor Temp.: 95.71 Ambient Temp.: 96.17 F
18:37:15.434 -> Average Cal sensor Temp.: 95.53 Ambient Temp.: 96.13 F
18:37:16.986 -> Average Cal sensor Temp.: 95.51 Ambient Temp.: 96.13 F
18:37:18.541 -> Average Cal sensor Temp.: 95.51 Ambient Temp.: 96.13 F
18:37:20.100 -> Average Cal sensor Temp.: 95.43 Ambient Temp.: 96.13 F
18:37:21.656 -> Average Cal sensor Temp.: 95.35 Ambient Temp.: 96.13 F
18:37:23.207 -> Average Cal sensor Temp.: 95.30 Ambient Temp.: 96.10 F
18:37:24.761 -> Average Cal sensor Temp.: 95.10 Ambient Temp.: 96.03 F
18:37:26.322 -> Average Cal sensor Temp.: 95.39 Ambient Temp.: 96.10 F
18:37:27.877 -> Average Cal sensor Temp.: 95.35 Ambient Temp.: 96.13 F
18:37:29.430 -> Average Cal sensor Temp.: 96.02 Ambient Temp.: 96.13 F
18:37:30.976 -> Average Cal sensor Temp.: 96.07 Ambient Temp.: 96.10 F
18:37:32.531 -> Average Cal sensor Temp.: 96.04 Ambient Temp.: 96.13 F
18:37:34.080 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 96.10 F
18:37:35.635 -> Average Cal sensor Temp.: 96.00 Ambient Temp.: 96.17 F
18:37:37.190 -> Average Cal sensor Temp.: 95.99 Ambient Temp.: 96.13 F
18:37:38.745 -> Average Cal sensor Temp.: 95.81 Ambient Temp.: 96.13 F
18:37:40.300 -> Average Cal sensor Temp.: 95.90 Ambient Temp.: 96.13 F
18:37:41.882 -> Average Cal sensor Temp.: 95.71 Ambient Temp.: 96.10 F
18:37:43.440 -> Average Cal sensor Temp.: 95.62 Ambient Temp.: 96.13 F
18:37:44.966 -> Average Cal sensor Temp.: 95.40 Ambient Temp.: 96.21 F
18:37:46.548 -> Average Cal sensor Temp.: 95.40 Ambient Temp.: 96.10 F
18:37:48.103 -> Average Cal sensor Temp.: 95.25 Ambient Temp.: 96.17 F
18:37:49.655 -> Average Cal sensor Temp.: 95.29 Ambient Temp.: 96.13 F
18:37:51.213 -> Average Cal sensor Temp.: 95.15 Ambient Temp.: 96.17 F
18:37:52.764 -> Average Cal sensor Temp.: 95.16 Ambient Temp.: 96.17 F
18:37:54.317 -> Average Cal sensor Temp.: 95.02 Ambient Temp.: 96.10 F
18:37:55.870 -> Average Cal sensor Temp.: 95.26 Ambient Temp.: 96.21 F
18:37:57.424 -> Average Cal sensor Temp.: 95.38 Ambient Temp.: 96.13 F
18:37:58.981 -> Average Cal sensor Temp.: 95.37 Ambient Temp.: 96.17 F
18:38:00.538 -> Average Cal sensor Temp.: 95.23 Ambient Temp.: 96.17 F
18:38:02.100 -> Average Cal sensor Temp.: 95.17 Ambient Temp.: 96.13 F
18:38:03.651 -> Average Cal sensor Temp.: 95.22 Ambient Temp.: 96.03 F
18:38:05.204 -> Average Cal sensor Temp.: 95.29 Ambient Temp.: 96.13 F
18:38:06.758 -> Average Cal sensor Temp.: 95.03 Ambient Temp.: 96.17 F
18:38:08.307 -> Average Cal sensor Temp.: 94.85 Ambient Temp.: 96.10 F
18:38:09.854 -> Average Cal sensor Temp.: 94.83 Ambient Temp.: 96.13 F
18:38:11.408 -> Average Cal sensor Temp.: 95.14 Ambient Temp.: 96.13 F
18:38:12.959 -> Average Cal sensor Temp.: 95.16 Ambient Temp.: 96.10 F
18:38:14.514 -> Average Cal sensor Temp.: 94.80 Ambient Temp.: 96.13 F
18:38:16.097 -> Average Cal sensor Temp.: 95.08 Ambient Temp.: 96.10 F
18:38:17.650 -> Average Cal sensor Temp.: 95.28 Ambient Temp.: 96.10 F
18:38:19.200 -> Average Cal sensor Temp.: 95.22 Ambient Temp.: 96.10 F
18:38:20.755 -> Average Cal sensor Temp.: 95.16 Ambient Temp.: 96.13 F
18:38:22.307 -> Average Cal sensor Temp.: 95.51 Ambient Temp.: 96.13 F
18:38:23.861 -> Average Cal sensor Temp.: 95.60 Ambient Temp.: 96.10 F
18:38:25.410 -> Average Cal sensor Temp.: 95.13 Ambient Temp.: 96.13 F
18:38:26.966 -> Average Cal sensor Temp.: 94.87 Ambient Temp.: 96.13 F
18:38:28.519 -> Average Cal sensor Temp.: 95.12 Ambient Temp.: 96.13 F
18:38:30.073 -> Average Cal sensor Temp.: 95.29 Ambient Temp.: 96.10 F
18:38:31.659 -> Average Cal sensor Temp.: 95.84 Ambient Temp.: 96.13 F
18:38:33.213 -> Average Cal sensor Temp.: 96.17 Ambient Temp.: 96.13 F
18:38:34.764 -> Average Cal sensor Temp.: 96.25 Ambient Temp.: 96.13 F
18:38:36.313 -> Average Cal sensor Temp.: 95.98 Ambient Temp.: 96.10 F
18:38:37.868 -> Average Cal sensor Temp.: 96.20 Ambient Temp.: 96.13 F
18:38:39.430 -> Average Cal sensor Temp.: 96.46 Ambient Temp.: 96.13 F
18:38:40.984 -> Average Cal sensor Temp.: 96.77 Ambient Temp.: 96.03 F
18:38:42.532 -> Average Cal sensor Temp.: 96.79 Ambient Temp.: 96.13 F
18:38:44.084 -> Average Cal sensor Temp.: 96.64 Ambient Temp.: 96.10 F
18:38:45.639 -> Average Cal sensor Temp.: 96.90 Ambient Temp.: 96.10 F
18:38:47.192 -> Average Cal sensor Temp.: 97.00 Ambient Temp.: 96.13 F
18:38:48.775 -> Average Cal sensor Temp.: 96.73 Ambient Temp.: 96.13 F
18:38:50.326 -> Average Cal sensor Temp.: 96.90 Ambient Temp.: 96.13 F
18:38:51.879 -> Average Cal sensor Temp.: 97.03 Ambient Temp.: 96.13 F
18:38:53.432 -> Average Cal sensor Temp.: 96.96 Ambient Temp.: 96.13 F
18:38:54.989 -> Average Cal sensor Temp.: 97.06 Ambient Temp.: 96.10 F
18:38:56.544 -> Average Cal sensor Temp.: 97.12 Ambient Temp.: 96.17 F
18:38:58.094 -> Average Cal sensor Temp.: 97.05 Ambient Temp.: 96.10 F
18:38:59.647 -> Average Cal sensor Temp.: 97.34 Ambient Temp.: 96.17 F
18:39:01.197 -> Average Cal sensor Temp.: 96.99 Ambient Temp.: 96.13 F
18:39:02.780 -> Average Cal sensor Temp.: 97.16 Ambient Temp.: 96.13 F
18:39:04.331 -> Average Cal sensor Temp.: 97.46 Ambient Temp.: 96.13 F
18:39:05.888 -> Average Cal sensor Temp.: 97.57 Ambient Temp.: 96.13 F
18:39:07.437 -> Average Cal sensor Temp.: 98.01 Ambient Temp.: 96.13 F
18:39:08.989 -> Average Cal sensor Temp.: 98.10 Ambient Temp.: 96.17 F
18:39:10.538 -> Average Cal sensor Temp.: 98.06 Ambient Temp.: 96.10 F
18:39:12.090 -> Average Cal sensor Temp.: 98.14 Ambient Temp.: 96.13 F
18:39:13.644 -> Average Cal sensor Temp.: 98.23 Ambient Temp.: 96.17 F
18:39:15.203 -> Average Cal sensor Temp.: 97.99 Ambient Temp.: 96.21 F
18:39:16.758 -> Average Cal sensor Temp.: 97.98 Ambient Temp.: 96.21 F
18:39:18.313 -> Average Cal sensor Temp.: 98.33 Ambient Temp.: 96.17 F
18:39:19.875 -> Average Cal sensor Temp.: 98.68 Ambient Temp.: 96.17 F
18:39:21.424 -> Average Cal sensor Temp.: 99.05 Ambient Temp.: 96.17 F
18:39:23.011 -> Average Cal sensor Temp.: 99.31 Ambient Temp.: 96.17 F
18:39:24.560 -> Average Cal sensor Temp.: 99.23 Ambient Temp.: 96.21 F
18:39:26.110 -> Average Cal sensor Temp.: 99.20 Ambient Temp.: 96.21 F
18:39:27.663 -> Average Cal sensor Temp.: 99.39 Ambient Temp.: 96.21 F
18:39:29.214 -> Average Cal sensor Temp.: 99.26 Ambient Temp.: 96.21 F
18:39:30.770 -> Average Cal sensor Temp.: 99.56 Ambient Temp.: 96.28 F
18:39:32.356 -> Average Cal sensor Temp.: 99.55 Ambient Temp.: 96.24 F
18:39:33.905 -> Average Cal sensor Temp.: 99.60 Ambient Temp.: 96.24 F
18:39:35.464 -> Average Cal sensor Temp.: 99.74 Ambient Temp.: 96.24 F
18:39:37.017 -> Average Cal sensor Temp.: 99.98 Ambient Temp.: 96.31 F
18:39:38.576 -> Average Cal sensor Temp.: 99.77 Ambient Temp.: 96.28 F
18:39:40.136 -> Average Cal sensor Temp.: 99.85 Ambient Temp.: 96.28 F
18:39:41.726 -> Average Cal sensor Temp.: 100.00 Ambient Temp.: 96.31 F
18:39:43.276 -> Average Cal sensor Temp.: 99.56 Ambient Temp.: 96.28 F
18:39:44.831 -> Average Cal sensor Temp.: 99.62 Ambient Temp.: 96.31 F
18:39:46.412 -> Average Cal sensor Temp.: 99.75 Ambient Temp.: 96.35 F
18:39:47.966 -> Average Cal sensor Temp.: 99.96 Ambient Temp.: 96.35 F
18:39:49.513 -> Average Cal sensor Temp.: 100.25 Ambient Temp.: 96.28 F
18:39:51.103 -> Average Cal sensor Temp.: 100.74 Ambient Temp.: 96.35 F
18:39:52.657 -> Average Cal sensor Temp.: 100.73 Ambient Temp.: 96.31 F
18:39:54.209 -> Average Cal sensor Temp.: 100.94 Ambient Temp.: 96.39 F
18:39:55.798 -> Average Cal sensor Temp.: 100.76 Ambient Temp.: 96.39 F
18:39:57.348 -> Average Cal sensor Temp.: 100.68 Ambient Temp.: 96.39 F
18:39:58.938 -> Average Cal sensor Temp.: 101.15 Ambient Temp.: 96.39 F
18:40:00.490 -> Average Cal sensor Temp.: 101.33 Ambient Temp.: 96.31 F
18:40:02.045 -> Average Cal sensor Temp.: 101.29 Ambient Temp.: 96.35 F
18:40:03.600 -> Average Cal sensor Temp.: 101.04 Ambient Temp.: 96.39 F
18:40:05.184 -> Average Cal sensor Temp.: 101.18 Ambient Temp.: 96.42 F
18:40:06.741 -> Average Cal sensor Temp.: 101.35 Ambient Temp.: 96.49 F
18:40:08.294 -> Average Cal sensor Temp.: 101.39 Ambient Temp.: 96.42 F
18:40:09.878 -> Average Cal sensor Temp.: 101.53 Ambient Temp.: 96.39 F
18:40:11.429 -> Average Cal sensor Temp.: 101.05 Ambient Temp.: 96.49 F
18:40:13.010 -> Average Cal sensor Temp.: 101.48 Ambient Temp.: 96.49 F
18:40:14.561 -> Average Cal sensor Temp.: 102.02 Ambient Temp.: 96.49 F
18:40:16.141 -> Average Cal sensor Temp.: 102.02 Ambient Temp.: 96.57 F
18:40:17.690 -> Average Cal sensor Temp.: 101.90 Ambient Temp.: 96.53 F
18:40:19.246 -> Average Cal sensor Temp.: 102.06 Ambient Temp.: 96.53 F
18:40:20.832 -> Average Cal sensor Temp.: 101.88 Ambient Temp.: 96.57 F
18:40:22.386 -> Average Cal sensor Temp.: 101.82 Ambient Temp.: 96.53 F
18:40:23.938 -> Average Cal sensor Temp.: 101.90 Ambient Temp.: 96.60 F
18:40:25.528 -> Average Cal sensor Temp.: 102.05 Ambient Temp.: 96.60 F
18:40:27.080 -> Average Cal sensor Temp.: 102.10 Ambient Temp.: 96.60 F
18:40:28.638 -> Average Cal sensor Temp.: 102.21 Ambient Temp.: 96.64 F
18:40:30.222 -> Average Cal sensor Temp.: 102.38 Ambient Temp.: 96.67 F
18:40:31.775 -> Average Cal sensor Temp.: 102.46 Ambient Temp.: 96.67 F
18:40:33.360 -> Average Cal sensor Temp.: 102.42 Ambient Temp.: 96.71 F
18:40:34.916 -> Average Cal sensor Temp.: 102.59 Ambient Temp.: 96.67 F
18:40:36.468 -> Average Cal sensor Temp.: 102.53 Ambient Temp.: 96.71 F
18:40:38.024 -> Average Cal sensor Temp.: 102.60 Ambient Temp.: 96.75 F
18:40:39.585 -> Average Cal sensor Temp.: 102.79 Ambient Temp.: 96.71 F
18:40:41.168 -> Average Cal sensor Temp.: 102.17 Ambient Temp.: 96.75 F
18:40:42.719 -> Average Cal sensor Temp.: 102.69 Ambient Temp.: 96.71 F
18:40:44.304 -> Average Cal sensor Temp.: 102.72 Ambient Temp.: 96.78 F
18:40:45.859 -> Average Cal sensor Temp.: 102.79 Ambient Temp.: 96.78 F
18:40:47.414 -> Average Cal sensor Temp.: 103.03 Ambient Temp.: 96.78 F
18:40:48.971 -> Average Cal sensor Temp.: 102.97 Ambient Temp.: 96.78 F
18:40:50.552 -> Average Cal sensor Temp.: 103.08 Ambient Temp.: 96.89 F
18:40:52.109 -> Average Cal sensor Temp.: 103.13 Ambient Temp.: 96.89 F
18:40:53.663 -> Average Cal sensor Temp.: 103.29 Ambient Temp.: 96.89 F
18:40:55.250 -> Average Cal sensor Temp.: 103.18 Ambient Temp.: 96.89 F
18:40:56.806 -> Average Cal sensor Temp.: 103.08 Ambient Temp.: 96.93 F
18:40:58.363 -> Average Cal sensor Temp.: 102.83 Ambient Temp.: 96.93 F
18:40:59.956 -> Average Cal sensor Temp.: 102.82 Ambient Temp.: 96.93 F
18:41:01.507 -> Average Cal sensor Temp.: 102.78 Ambient Temp.: 96.93 F
18:41:03.061 -> Average Cal sensor Temp.: 103.10 Ambient Temp.: 96.93 F
18:41:04.651 -> Average Cal sensor Temp.: 102.96 Ambient Temp.: 96.93 F
18:41:06.204 -> Average Cal sensor Temp.: 102.97 Ambient Temp.: 96.96 F
18:41:07.756 -> Average Cal sensor Temp.: 103.06 Ambient Temp.: 96.96 F
18:41:09.341 -> Average Cal sensor Temp.: 102.85 Ambient Temp.: 97.03 F
18:41:10.897 -> Average Cal sensor Temp.: 103.06 Ambient Temp.: 97.00 F
18:41:12.451 -> Average Cal sensor Temp.: 103.18 Ambient Temp.: 97.07 F
18:41:14.041 -> Average Cal sensor Temp.: 103.01 Ambient Temp.: 97.07 F
18:41:15.594 -> Average Cal sensor Temp.: 103.21 Ambient Temp.: 97.11 F
18:41:17.150 -> Average Cal sensor Temp.: 103.15 Ambient Temp.: 97.03 F
18:41:18.706 -> Average Cal sensor Temp.: 102.95 Ambient Temp.: 97.11 F
18:41:20.289 -> Average Cal sensor Temp.: 103.09 Ambient Temp.: 97.11 F
18:41:21.842 -> Average Cal sensor Temp.: 102.49 Ambient Temp.: 97.07 F
18:41:23.397 -> Average Cal sensor Temp.: 102.32 Ambient Temp.: 97.11 F
18:41:24.985 -> Average Cal sensor Temp.: 102.16 Ambient Temp.: 97.11 F
18:41:26.536 -> Average Cal sensor Temp.: 102.47 Ambient Temp.: 97.11 F
18:41:28.091 -> Average Cal sensor Temp.: 102.77 Ambient Temp.: 97.14 F
18:41:29.671 -> Average Cal sensor Temp.: 102.90 Ambient Temp.: 97.18 F
18:41:31.219 -> Average Cal sensor Temp.: 102.88 Ambient Temp.: 97.18 F
18:41:32.802 -> Average Cal sensor Temp.: 102.91 Ambient Temp.: 97.18 F
18:41:34.362 -> Average Cal sensor Temp.: 102.57 Ambient Temp.: 97.18 F
18:41:35.917 -> Average Cal sensor Temp.: 102.19 Ambient Temp.: 97.21 F
18:41:37.476 -> Average Cal sensor Temp.: 101.99 Ambient Temp.: 97.29 F
18:41:39.034 -> Average Cal sensor Temp.: 102.23 Ambient Temp.: 97.29 F
18:41:40.589 -> Average Cal sensor Temp.: 102.38 Ambient Temp.: 97.29 F
18:41:42.181 -> Average Cal sensor Temp.: 102.60 Ambient Temp.: 97.32 F
18:41:43.734 -> Average Cal sensor Temp.: 102.84 Ambient Temp.: 97.32 F
18:41:45.286 -> Average Cal sensor Temp.: 102.89 Ambient Temp.: 97.36 F
18:41:46.871 -> Average Cal sensor Temp.: 102.99 Ambient Temp.: 97.32 F
18:41:48.419 -> Average Cal sensor Temp.: 102.71 Ambient Temp.: 97.32 F
18:41:49.974 -> Average Cal sensor Temp.: 102.46 Ambient Temp.: 97.32 F
18:41:51.567 -> Average Cal sensor Temp.: 102.62 Ambient Temp.: 97.32 F
18:41:53.120 -> Average Cal sensor Temp.: 102.65 Ambient Temp.: 97.36 F
18:41:54.670 -> Average Cal sensor Temp.: 102.56 Ambient Temp.: 97.39 F
18:41:56.261 -> Average Cal sensor Temp.: 102.73 Ambient Temp.: 97.36 F
18:41:57.814 -> Average Cal sensor Temp.: 102.72 Ambient Temp.: 97.36 F
18:41:59.370 -> Average Cal sensor Temp.: 102.67 Ambient Temp.: 97.39 F
18:42:00.926 -> Average Cal sensor Temp.: 102.60 Ambient Temp.: 97.43 F
18:42:02.516 -> Average Cal sensor Temp.: 102.65 Ambient Temp.: 97.43 F
18:42:04.067 -> Average Cal sensor Temp.: 102.61 Ambient Temp.: 97.43 F
18:42:05.626 -> Average Cal sensor Temp.: 101.93 Ambient Temp.: 97.43 F
18:42:07.179 -> Average Cal sensor Temp.: 102.07 Ambient Temp.: 97.39 F
18:42:08.772 -> Average Cal sensor Temp.: 102.07 Ambient Temp.: 97.43 F
18:42:10.322 -> Average Cal sensor Temp.: 102.38 Ambient Temp.: 97.43 F
18:42:11.877 -> Average Cal sensor Temp.: 102.39 Ambient Temp.: 97.43 F
18:42:13.432 -> Average Cal sensor Temp.: 102.40 Ambient Temp.: 97.43 F
18:42:15.021 -> Average Cal sensor Temp.: 102.35 Ambient Temp.: 97.43 F
18:42:16.575 -> Average Cal sensor Temp.: 102.26 Ambient Temp.: 97.43 F
18:42:18.128 -> Average Cal sensor Temp.: 102.57 Ambient Temp.: 97.43 F
18:42:19.693 -> Average Cal sensor Temp.: 102.58 Ambient Temp.: 97.43 F
18:42:21.282 -> Average Cal sensor Temp.: 102.31 Ambient Temp.: 97.47 F
18:42:22.830 -> Average Cal sensor Temp.: 102.62 Ambient Temp.: 97.47 F
18:42:24.386 -> Average Cal sensor Temp.: 102.73 Ambient Temp.: 97.50 F
18:42:25.942 -> Average Cal sensor Temp.: 102.31 Ambient Temp.: 97.43 F
18:42:27.534 -> Average Cal sensor Temp.: 102.42 Ambient Temp.: 97.43 F
18:42:29.088 -> Average Cal sensor Temp.: 102.54 Ambient Temp.: 97.47 F
18:42:30.645 -> Average Cal sensor Temp.: 102.39 Ambient Temp.: 97.43 F
18:42:32.198 -> Average Cal sensor Temp.: 102.49 Ambient Temp.: 97.47 F
18:42:33.777 -> Average Cal sensor Temp.: 102.19 Ambient Temp.: 97.50 F
18:42:35.328 -> Average Cal sensor Temp.: 102.53 Ambient Temp.: 97.50 F
18:42:36.918 -> Average Cal sensor Temp.: 102.88 Ambient Temp.: 97.54 F
18:42:38.475 -> Average Cal sensor Temp.: 102.75 Ambient Temp.: 97.57 F
18:42:40.032 -> Average Cal sensor Temp.: 102.65 Ambient Temp.: 97.57 F
18:42:41.588 -> Average Cal sensor Temp.: 102.79 Ambient Temp.: 97.54 F
18:42:43.176 -> Average Cal sensor Temp.: 102.72 Ambient Temp.: 97.54 F
18:42:44.730 -> Average Cal sensor Temp.: 102.47 Ambient Temp.: 97.54 F
18:42:46.282 -> Average Cal sensor Temp.: 102.67 Ambient Temp.: 97.54 F
18:42:47.867 -> Average Cal sensor Temp.: 102.59 Ambient Temp.: 97.61 F
18:42:49.419 -> Average Cal sensor Temp.: 102.75 Ambient Temp.: 97.57 F
18:42:50.972 -> Average Cal sensor Temp.: 102.87 Ambient Temp.: 97.50 F
18:42:52.530 -> Average Cal sensor Temp.: 102.87 Ambient Temp.: 97.54 F
18:42:54.119 -> Average Cal sensor Temp.: 103.04 Ambient Temp.: 97.61 F
18:42:55.672 -> Average Cal sensor Temp.: 102.85 Ambient Temp.: 97.57 F
18:42:57.225 -> Average Cal sensor Temp.: 103.04 Ambient Temp.: 97.61 F
18:42:58.812 -> Average Cal sensor Temp.: 103.05 Ambient Temp.: 97.68 F
18:43:00.372 -> Average Cal sensor Temp.: 102.57 Ambient Temp.: 97.61 F
18:43:01.924 -> Average Cal sensor Temp.: 102.69 Ambient Temp.: 97.68 F
18:43:03.479 -> Average Cal sensor Temp.: 102.42 Ambient Temp.: 97.68 F
18:43:05.068 -> Average Cal sensor Temp.: 102.65 Ambient Temp.: 97.61 F
18:43:06.624 -> Average Cal sensor Temp.: 102.75 Ambient Temp.: 97.68 F
18:43:08.175 -> Average Cal sensor Temp.: 102.77 Ambient Temp.: 97.68 F
18:43:09.763 -> Average Cal sensor Temp.: 103.00 Ambient Temp.: 97.68 F
18:43:11.321 -> Average Cal sensor Temp.: 102.99 Ambient Temp.: 97.72 F
18:43:12.874 -> Average Cal sensor Temp.: 103.16 Ambient Temp.: 97.72 F
18:43:14.429 -> Average Cal sensor Temp.: 102.98 Ambient Temp.: 97.75 F
18:43:16.015 -> Average Cal sensor Temp.: 103.16 Ambient Temp.: 97.68 F
18:43:17.569 -> Average Cal sensor Temp.: 103.20 Ambient Temp.: 97.75 F
18:43:19.123 -> Average Cal sensor Temp.: 103.09 Ambient Temp.: 97.72 F
18:43:20.703 -> Average Cal sensor Temp.: 103.21 Ambient Temp.: 97.75 F
18:43:22.257 -> Average Cal sensor Temp.: 102.78 Ambient Temp.: 97.79 F
18:43:23.838 -> Average Cal sensor Temp.: 102.56 Ambient Temp.: 97.75 F
18:43:25.394 -> Average Cal sensor Temp.: 102.86 Ambient Temp.: 97.72 F
18:43:26.951 -> Average Cal sensor Temp.: 103.26 Ambient Temp.: 97.75 F
18:43:28.506 -> Average Cal sensor Temp.: 103.29 Ambient Temp.: 97.75 F
18:43:30.062 -> Average Cal sensor Temp.: 103.25 Ambient Temp.: 97.83 F
18:43:31.652 -> Average Cal sensor Temp.: 103.27 Ambient Temp.: 97.83 F
18:43:33.203 -> Average Cal sensor Temp.: 103.39 Ambient Temp.: 97.79 F
18:43:34.757 -> Average Cal sensor Temp.: 103.52 Ambient Temp.: 97.79 F
18:43:36.351 -> Average Cal sensor Temp.: 103.45 Ambient Temp.: 97.83 F
18:43:37.902 -> Average Cal sensor Temp.: 103.39 Ambient Temp.: 97.83 F
18:43:39.455 -> Average Cal sensor Temp.: 103.31 Ambient Temp.: 97.86 F
18:43:41.009 -> Average Cal sensor Temp.: 102.69 Ambient Temp.: 97.86 F
18:43:42.597 -> Average Cal sensor Temp.: 103.31 Ambient Temp.: 97.83 F
18:43:44.151 -> Average Cal sensor Temp.: 102.75 Ambient Temp.: 97.83 F
18:43:45.706 -> Average Cal sensor Temp.: 103.26 Ambient Temp.: 97.86 F
18:43:47.295 -> Average Cal sensor Temp.: 103.36 Ambient Temp.: 97.86 F
18:43:48.851 -> Average Cal sensor Temp.: 103.41 Ambient Temp.: 97.83 F
18:43:50.401 -> Average Cal sensor Temp.: 103.44 Ambient Temp.: 97.86 F
18:43:51.985 -> Average Cal sensor Temp.: 103.41 Ambient Temp.: 97.90 F
18:43:53.542 -> Average Cal sensor Temp.: 103.41 Ambient Temp.: 97.86 F
18:43:55.089 -> Average Cal sensor Temp.: 103.57 Ambient Temp.: 97.90 F
18:43:56.679 -> Average Cal sensor Temp.: 103.64 Ambient Temp.: 97.90 F
18:43:58.236 -> Average Cal sensor Temp.: 103.01 Ambient Temp.: 97.90 F
18:43:59.781 -> Average Cal sensor Temp.: 102.52 Ambient Temp.: 97.86 F
18:44:01.367 -> Average Cal sensor Temp.: 103.22 Ambient Temp.: 97.90 F
18:44:02.916 -> Average Cal sensor Temp.: 103.49 Ambient Temp.: 97.86 F
18:44:04.501 -> Average Cal sensor Temp.: 102.59 Ambient Temp.: 97.90 F
18:44:06.056 -> Average Cal sensor Temp.: 102.82 Ambient Temp.: 97.83 F
18:44:07.607 -> Average Cal sensor Temp.: 103.10 Ambient Temp.: 97.86 F
18:44:09.198 -> Average Cal sensor Temp.: 103.38 Ambient Temp.: 97.83 F
18:44:10.749 -> Average Cal sensor Temp.: 103.34 Ambient Temp.: 97.90 F
18:44:11.337 -> Ambient Temp.: 97.90 F
18:44:11.920 -> Ambient Temp.: 97.90 F
18:44:12.509 -> Ambient Temp.: 97.90 F
18:44:13.097 -> Ambient Temp.: 97.90 F
18:44:13.684 -> Ambient Temp.: 97.90 F
18:44:14.233 -> Ambient Temp.: 97.90 F
18:44:15.811 -> Average Cal sensor Temp.: 103.51 Ambient Temp.: 97.86 F
18:44:17.366 -> Average Cal sensor Temp.: 103.68 Ambient Temp.: 97.83 F
18:44:18.090 -> Ambient Temp.: 97.83 F
18:44:18.677 -> Ambient Temp.: 97.83 F
18:44:20.195 -> Average Cal sensor Temp.: 103.70 Ambient Temp.: 97.86 F
18:44:21.784 -> Average Cal sensor Temp.: 103.51 Ambient Temp.: 97.86 F
18:44:23.335 -> Average Cal sensor Temp.: 104.08 Ambient Temp.: 97.83 F
18:44:24.917 -> Average Cal sensor Temp.: 104.04 Ambient Temp.: 97.86 F
18:44:26.469 -> Average Cal sensor Temp.: 103.65 Ambient Temp.: 97.90 F
18:44:28.028 -> Average Cal sensor Temp.: 104.01 Ambient Temp.: 97.90 F
18:44:29.611 -> Average Cal sensor Temp.: 104.01 Ambient Temp.: 97.93 F
18:44:31.161 -> Average Cal sensor Temp.: 103.94 Ambient Temp.: 97.90 F
18:44:32.747 -> Average Cal sensor Temp.: 104.07 Ambient Temp.: 97.90 F
18:44:34.298 -> Average Cal sensor Temp.: 103.96 Ambient Temp.: 97.93 F
18:44:35.855 -> Average Cal sensor Temp.: 103.85 Ambient Temp.: 97.93 F
18:44:37.440 -> Average Cal sensor Temp.: 104.08 Ambient Temp.: 97.90 F
18:44:38.996 -> Average Cal sensor Temp.: 104.03 Ambient Temp.: 97.93 F
18:44:40.553 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 97.90 F
18:44:42.110 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 97.90 F
18:44:43.697 -> Average Cal sensor Temp.: 104.05 Ambient Temp.: 97.90 F
18:44:45.244 -> Average Cal sensor Temp.: 104.10 Ambient Temp.: 97.90 F
18:44:46.831 -> Average Cal sensor Temp.: 104.13 Ambient Temp.: 97.97 F
18:44:48.385 -> Average Cal sensor Temp.: 104.20 Ambient Temp.: 97.97 F
18:44:49.937 -> Average Cal sensor Temp.: 103.93 Ambient Temp.: 98.01 F
18:44:51.517 -> Average Cal sensor Temp.: 103.55 Ambient Temp.: 97.93 F
18:44:53.066 -> Average Cal sensor Temp.: 103.98 Ambient Temp.: 97.86 F
18:44:54.648 -> Average Cal sensor Temp.: 104.25 Ambient Temp.: 97.90 F
18:44:56.195 -> Average Cal sensor Temp.: 104.09 Ambient Temp.: 97.93 F
18:44:57.783 -> Average Cal sensor Temp.: 104.32 Ambient Temp.: 97.90 F
18:44:59.339 -> Average Cal sensor Temp.: 104.11 Ambient Temp.: 98.01 F
18:45:00.890 -> Average Cal sensor Temp.: 104.08 Ambient Temp.: 97.93 F
18:45:02.478 -> Average Cal sensor Temp.: 104.05 Ambient Temp.: 97.93 F
18:45:04.030 -> Average Cal sensor Temp.: 104.20 Ambient Temp.: 97.97 F
18:45:05.589 -> Average Cal sensor Temp.: 104.48 Ambient Temp.: 97.90 F
18:45:07.174 -> Average Cal sensor Temp.: 104.25 Ambient Temp.: 97.97 F
18:45:08.719 -> Average Cal sensor Temp.: 104.23 Ambient Temp.: 97.93 F
18:45:10.275 -> Average Cal sensor Temp.: 104.17 Ambient Temp.: 97.93 F
18:45:11.862 -> Average Cal sensor Temp.: 104.17 Ambient Temp.: 97.97 F
18:45:13.414 -> Average Cal sensor Temp.: 104.17 Ambient Temp.: 97.93 F
18:45:14.965 -> Average Cal sensor Temp.: 104.18 Ambient Temp.: 97.97 F
18:45:16.548 -> Average Cal sensor Temp.: 104.33 Ambient Temp.: 97.97 F
18:45:18.104 -> Average Cal sensor Temp.: 104.62 Ambient Temp.: 98.01 F
18:45:19.655 -> Average Cal sensor Temp.: 104.72 Ambient Temp.: 98.01 F
18:45:21.243 -> Average Cal sensor Temp.: 104.71 Ambient Temp.: 97.97 F
18:45:22.794 -> Average Cal sensor Temp.: 104.59 Ambient Temp.: 98.01 F
18:45:24.352 -> Average Cal sensor Temp.: 104.67 Ambient Temp.: 98.01 F
18:45:25.941 -> Average Cal sensor Temp.: 104.77 Ambient Temp.: 98.08 F
18:45:27.493 -> Average Cal sensor Temp.: 104.71 Ambient Temp.: 98.01 F
18:45:29.049 -> Average Cal sensor Temp.: 104.66 Ambient Temp.: 98.01 F
18:45:30.631 -> Average Cal sensor Temp.: 104.66 Ambient Temp.: 98.11 F
18:45:32.188 -> Average Cal sensor Temp.: 104.33 Ambient Temp.: 98.08 F
18:45:33.745 -> Average Cal sensor Temp.: 104.29 Ambient Temp.: 98.11 F
18:45:35.327 -> Average Cal sensor Temp.: 104.08 Ambient Temp.: 98.08 F
18:45:36.880 -> Average Cal sensor Temp.: 103.91 Ambient Temp.: 98.08 F
18:45:38.435 -> Average Cal sensor Temp.: 104.08 Ambient Temp.: 98.01 F
18:45:40.016 -> Average Cal sensor Temp.: 104.23 Ambient Temp.: 98.11 F
18:45:41.572 -> Average Cal sensor Temp.: 104.20 Ambient Temp.: 98.11 F
18:45:43.157 -> Average Cal sensor Temp.: 104.23 Ambient Temp.: 98.08 F
18:45:44.709 -> Average Cal sensor Temp.: 104.36 Ambient Temp.: 98.11 F
18:45:46.261 -> Average Cal sensor Temp.: 104.57 Ambient Temp.: 98.11 F
18:45:47.845 -> Average Cal sensor Temp.: 104.35 Ambient Temp.: 98.11 F
18:45:49.398 -> Average Cal sensor Temp.: 104.54 Ambient Temp.: 98.15 F
18:45:50.950 -> Average Cal sensor Temp.: 104.55 Ambient Temp.: 98.19 F
18:45:52.544 -> Average Cal sensor Temp.: 103.77 Ambient Temp.: 98.15 F
18:45:54.093 -> Average Cal sensor Temp.: 103.33 Ambient Temp.: 98.08 F
18:45:55.641 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 98.11 F
18:45:57.232 -> Average Cal sensor Temp.: 104.14 Ambient Temp.: 98.11 F
18:45:58.787 -> Average Cal sensor Temp.: 104.39 Ambient Temp.: 98.11 F
18:46:00.345 -> Average Cal sensor Temp.: 103.94 Ambient Temp.: 98.08 F
18:46:01.899 -> Average Cal sensor Temp.: 104.07 Ambient Temp.: 98.01 F
18:46:03.490 -> Average Cal sensor Temp.: 104.21 Ambient Temp.: 98.08 F
18:46:05.042 -> Average Cal sensor Temp.: 104.62 Ambient Temp.: 98.11 F
18:46:06.598 -> Average Cal sensor Temp.: 104.64 Ambient Temp.: 98.08 F
18:46:08.186 -> Average Cal sensor Temp.: 104.83 Ambient Temp.: 98.15 F
18:46:09.736 -> Average Cal sensor Temp.: 105.09 Ambient Temp.: 98.11 F
18:46:11.291 -> Average Cal sensor Temp.: 104.58 Ambient Temp.: 98.08 F
18:46:12.881 -> Average Cal sensor Temp.: 104.85 Ambient Temp.: 98.11 F
18:46:14.431 -> Average Cal sensor Temp.: 104.73 Ambient Temp.: 98.15 F
18:46:16.017 -> Average Cal sensor Temp.: 104.95 Ambient Temp.: 98.11 F
18:46:17.566 -> Average Cal sensor Temp.: 104.72 Ambient Temp.: 98.11 F
18:46:19.118 -> Average Cal sensor Temp.: 104.75 Ambient Temp.: 98.15 F
18:46:20.696 -> Average Cal sensor Temp.: 104.81 Ambient Temp.: 98.19 F
18:46:22.257 -> Average Cal sensor Temp.: 104.63 Ambient Temp.: 98.19 F
18:46:23.816 -> Average Cal sensor Temp.: 104.45 Ambient Temp.: 98.19 F
18:46:25.403 -> Average Cal sensor Temp.: 104.50 Ambient Temp.: 98.19 F
18:46:26.960 -> Average Cal sensor Temp.: 104.28 Ambient Temp.: 98.19 F
18:46:28.511 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 98.19 F
18:46:30.067 -> Average Cal sensor Temp.: 104.35 Ambient Temp.: 98.15 F
18:46:31.657 -> Average Cal sensor Temp.: 104.61 Ambient Temp.: 98.22 F
18:46:33.208 -> Average Cal sensor Temp.: 104.72 Ambient Temp.: 98.22 F
18:46:34.792 -> Average Cal sensor Temp.: 104.92 Ambient Temp.: 98.19 F
18:46:36.352 -> Average Cal sensor Temp.: 105.19 Ambient Temp.: 98.26 F
18:46:37.905 -> Average Cal sensor Temp.: 105.20 Ambient Temp.: 98.22 F
18:46:39.466 -> Average Cal sensor Temp.: 104.81 Ambient Temp.: 98.22 F
18:46:41.051 -> Average Cal sensor Temp.: 104.74 Ambient Temp.: 98.22 F
18:46:42.606 -> Average Cal sensor Temp.: 104.79 Ambient Temp.: 98.19 F
18:46:44.159 -> Average Cal sensor Temp.: 104.54 Ambient Temp.: 98.26 F
18:46:45.746 -> Average Cal sensor Temp.: 104.53 Ambient Temp.: 98.26 F
18:46:47.301 -> Average Cal sensor Temp.: 104.35 Ambient Temp.: 98.26 F
18:46:48.855 -> Average Cal sensor Temp.: 103.74 Ambient Temp.: 98.22 F
18:46:50.437 -> Average Cal sensor Temp.: 104.03 Ambient Temp.: 98.26 F
18:46:51.990 -> Average Cal sensor Temp.: 104.10 Ambient Temp.: 98.22 F
18:46:53.544 -> Average Cal sensor Temp.: 104.30 Ambient Temp.: 98.22 F
18:46:55.131 -> Average Cal sensor Temp.: 104.52 Ambient Temp.: 98.22 F
18:46:56.683 -> Average Cal sensor Temp.: 104.53 Ambient Temp.: 98.19 F
18:46:58.239 -> Average Cal sensor Temp.: 104.45 Ambient Temp.: 98.26 F
18:46:59.828 -> Average Cal sensor Temp.: 104.55 Ambient Temp.: 98.26 F
18:47:01.382 -> Average Cal sensor Temp.: 104.48 Ambient Temp.: 98.26 F
18:47:02.937 -> Average Cal sensor Temp.: 104.09 Ambient Temp.: 98.19 F
18:47:04.525 -> Average Cal sensor Temp.: 103.79 Ambient Temp.: 98.26 F
18:47:06.078 -> Average Cal sensor Temp.: 104.18 Ambient Temp.: 98.22 F
18:47:07.633 -> Average Cal sensor Temp.: 104.14 Ambient Temp.: 98.22 F
18:47:09.188 -> Average Cal sensor Temp.: 103.99 Ambient Temp.: 98.19 F
18:47:10.772 -> Average Cal sensor Temp.: 104.36 Ambient Temp.: 98.22 F
18:47:12.332 -> Average Cal sensor Temp.: 104.67 Ambient Temp.: 98.22 F
18:47:13.887 -> Average Cal sensor Temp.: 104.38 Ambient Temp.: 98.19 F
18:47:15.472 -> Average Cal sensor Temp.: 104.45 Ambient Temp.: 98.19 F
18:47:17.023 -> Average Cal sensor Temp.: 101.82 Ambient Temp.: 98.15 F
18:47:18.578 -> Average Cal sensor Temp.: 103.77 Ambient Temp.: 98.15 F
18:47:20.158 -> Average Cal sensor Temp.: 103.80 Ambient Temp.: 98.19 F
18:47:21.713 -> Average Cal sensor Temp.: 103.95 Ambient Temp.: 98.11 F
18:47:23.303 -> Average Cal sensor Temp.: 103.89 Ambient Temp.: 98.19 F
18:47:24.856 -> Average Cal sensor Temp.: 104.19 Ambient Temp.: 98.19 F
18:47:26.413 -> Average Cal sensor Temp.: 104.36 Ambient Temp.: 98.19 F
18:47:27.993 -> Average Cal sensor Temp.: 104.24 Ambient Temp.: 98.22 F
18:47:29.551 -> Average Cal sensor Temp.: 103.67 Ambient Temp.: 98.19 F
18:47:31.107 -> Average Cal sensor Temp.: 103.76 Ambient Temp.: 98.19 F
18:47:32.668 -> Average Cal sensor Temp.: 104.05 Ambient Temp.: 98.19 F
18:47:34.254 -> Average Cal sensor Temp.: 104.10 Ambient Temp.: 98.22 F
18:47:35.802 -> Average Cal sensor Temp.: 103.72 Ambient Temp.: 98.22 F
18:47:37.354 -> Average Cal sensor Temp.: 103.37 Ambient Temp.: 98.22 F
18:47:38.939 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 98.19 F
18:47:40.515 -> Average Cal sensor Temp.: 103.41 Ambient Temp.: 98.19 F
18:47:42.067 -> Average Cal sensor Temp.: 102.81 Ambient Temp.: 98.19 F
18:47:43.621 -> Average Cal sensor Temp.: 102.71 Ambient Temp.: 98.19 F
18:47:45.174 -> Average Cal sensor Temp.: 102.66 Ambient Temp.: 98.19 F
18:47:46.756 -> Average Cal sensor Temp.: 102.79 Ambient Temp.: 98.15 F
18:47:48.306 -> Average Cal sensor Temp.: 102.85 Ambient Temp.: 98.19 F
18:47:49.894 -> Average Cal sensor Temp.: 102.83 Ambient Temp.: 98.22 F
18:47:51.453 -> Average Cal sensor Temp.: 102.60 Ambient Temp.: 98.15 F
18:47:53.008 -> Average Cal sensor Temp.: 102.54 Ambient Temp.: 98.11 F
18:47:54.564 -> Average Cal sensor Temp.: 102.82 Ambient Temp.: 98.15 F
18:47:56.158 -> Average Cal sensor Temp.: 102.66 Ambient Temp.: 98.15 F
18:47:57.717 -> Average Cal sensor Temp.: 102.38 Ambient Temp.: 98.15 F
18:47:59.269 -> Average Cal sensor Temp.: 102.54 Ambient Temp.: 98.19 F
18:48:00.831 -> Average Cal sensor Temp.: 102.65 Ambient Temp.: 98.11 F
18:48:02.385 -> Average Cal sensor Temp.: 102.73 Ambient Temp.: 98.08 F
18:48:03.973 -> Average Cal sensor Temp.: 102.37 Ambient Temp.: 98.15 F
18:48:05.533 -> Average Cal sensor Temp.: 102.33 Ambient Temp.: 98.15 F
18:48:07.089 -> Average Cal sensor Temp.: 102.39 Ambient Temp.: 98.19 F
18:48:08.647 -> Average Cal sensor Temp.: 102.37 Ambient Temp.: 98.15 F
18:48:10.202 -> Average Cal sensor Temp.: 102.32 Ambient Temp.: 98.19 F
18:48:11.795 -> Average Cal sensor Temp.: 102.22 Ambient Temp.: 98.15 F
18:48:13.354 -> Average Cal sensor Temp.: 101.99 Ambient Temp.: 98.19 F
18:48:14.911 -> Average Cal sensor Temp.: 101.93 Ambient Temp.: 98.19 F
18:48:16.472 -> Average Cal sensor Temp.: 101.76 Ambient Temp.: 98.19 F
18:48:18.055 -> Average Cal sensor Temp.: 101.94 Ambient Temp.: 98.19 F
18:48:19.609 -> Average Cal sensor Temp.: 102.29 Ambient Temp.: 98.11 F
18:48:21.163 -> Average Cal sensor Temp.: 102.33 Ambient Temp.: 98.15 F
18:48:22.715 -> Average Cal sensor Temp.: 102.09 Ambient Temp.: 98.11 F
18:48:24.307 -> Average Cal sensor Temp.: 102.11 Ambient Temp.: 98.08 F
18:48:25.864 -> Average Cal sensor Temp.: 101.88 Ambient Temp.: 98.11 F
18:48:27.417 -> Average Cal sensor Temp.: 101.59 Ambient Temp.: 98.08 F
18:48:28.974 -> Average Cal sensor Temp.: 101.48 Ambient Temp.: 98.11 F
18:48:30.562 -> Average Cal sensor Temp.: 101.41 Ambient Temp.: 98.08 F
18:48:32.121 -> Average Cal sensor Temp.: 101.25 Ambient Temp.: 98.11 F
18:48:33.678 -> Average Cal sensor Temp.: 101.01 Ambient Temp.: 98.01 F
18:48:35.240 -> Average Cal sensor Temp.: 101.17 Ambient Temp.: 98.08 F
18:48:36.797 -> Average Cal sensor Temp.: 101.22 Ambient Temp.: 98.08 F
18:48:38.355 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 98.11 F
18:48:39.946 -> Average Cal sensor Temp.: 101.39 Ambient Temp.: 98.08 F
18:48:41.495 -> Average Cal sensor Temp.: 101.44 Ambient Temp.: 98.01 F
18:48:43.053 -> Average Cal sensor Temp.: 101.44 Ambient Temp.: 98.08 F
18:48:44.642 -> Average Cal sensor Temp.: 101.45 Ambient Temp.: 97.97 F
18:48:46.198 -> Average Cal sensor Temp.: 101.29 Ambient Temp.: 97.97 F
18:48:47.755 -> Average Cal sensor Temp.: 101.50 Ambient Temp.: 97.93 F
18:48:49.310 -> Average Cal sensor Temp.: 101.87 Ambient Temp.: 98.01 F
18:48:50.867 -> Average Cal sensor Temp.: 101.61 Ambient Temp.: 97.93 F
18:48:52.454 -> Average Cal sensor Temp.: 101.84 Ambient Temp.: 97.93 F
18:48:54.012 -> Average Cal sensor Temp.: 101.54 Ambient Temp.: 97.93 F
18:48:55.561 -> Average Cal sensor Temp.: 101.53 Ambient Temp.: 97.97 F
18:48:57.145 -> Average Cal sensor Temp.: 101.14 Ambient Temp.: 97.93 F
18:48:58.701 -> Average Cal sensor Temp.: 101.09 Ambient Temp.: 97.90 F
18:49:00.263 -> Average Cal sensor Temp.: 101.10 Ambient Temp.: 97.90 F
18:49:01.822 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 97.90 F
18:49:03.379 -> Average Cal sensor Temp.: 101.13 Ambient Temp.: 97.90 F
18:49:04.965 -> Average Cal sensor Temp.: 100.94 Ambient Temp.: 97.83 F
18:49:06.525 -> Average Cal sensor Temp.: 100.89 Ambient Temp.: 97.83 F
18:49:08.078 -> Average Cal sensor Temp.: 101.04 Ambient Temp.: 97.83 F
18:49:09.636 -> Average Cal sensor Temp.: 100.93 Ambient Temp.: 97.83 F
18:49:11.206 -> Average Cal sensor Temp.: 101.05 Ambient Temp.: 97.75 F
18:49:12.771 -> Average Cal sensor Temp.: 101.06 Ambient Temp.: 97.75 F
18:49:14.357 -> Average Cal sensor Temp.: 100.76 Ambient Temp.: 97.75 F
18:49:15.912 -> Average Cal sensor Temp.: 100.98 Ambient Temp.: 97.72 F
18:49:17.467 -> Average Cal sensor Temp.: 101.01 Ambient Temp.: 97.79 F
18:49:19.025 -> Average Cal sensor Temp.: 101.02 Ambient Temp.: 97.72 F
18:49:20.582 -> Average Cal sensor Temp.: 101.41 Ambient Temp.: 97.68 F
18:49:22.164 -> Average Cal sensor Temp.: 101.61 Ambient Temp.: 97.72 F
18:49:23.719 -> Average Cal sensor Temp.: 101.81 Ambient Temp.: 97.75 F
18:49:25.276 -> Average Cal sensor Temp.: 101.88 Ambient Temp.: 97.72 F
18:49:26.864 -> Average Cal sensor Temp.: 102.29 Ambient Temp.: 97.68 F
18:49:28.421 -> Average Cal sensor Temp.: 102.57 Ambient Temp.: 97.68 F
18:49:29.974 -> Average Cal sensor Temp.: 102.34 Ambient Temp.: 97.72 F
18:49:31.528 -> Average Cal sensor Temp.: 102.11 Ambient Temp.: 97.68 F
18:49:33.114 -> Average Cal sensor Temp.: 101.41 Ambient Temp.: 97.61 F
18:49:34.667 -> Average Cal sensor Temp.: 101.68 Ambient Temp.: 97.72 F
18:49:36.220 -> Average Cal sensor Temp.: 101.41 Ambient Temp.: 97.72 F
18:49:37.807 -> Average Cal sensor Temp.: 101.41 Ambient Temp.: 97.61 F
18:49:39.361 -> Average Cal sensor Temp.: 101.71 Ambient Temp.: 97.68 F
18:49:40.916 -> Average Cal sensor Temp.: 101.54 Ambient Temp.: 97.75 F
18:49:42.471 -> Average Cal sensor Temp.: 101.40 Ambient Temp.: 97.68 F
18:49:44.052 -> Average Cal sensor Temp.: 101.68 Ambient Temp.: 97.61 F
18:49:45.607 -> Average Cal sensor Temp.: 101.36 Ambient Temp.: 97.61 F
18:49:47.158 -> Average Cal sensor Temp.: 101.53 Ambient Temp.: 97.57 F
18:49:48.746 -> Average Cal sensor Temp.: 101.28 Ambient Temp.: 97.61 F
18:49:50.302 -> Average Cal sensor Temp.: 101.30 Ambient Temp.: 97.61 F
18:49:51.854 -> Average Cal sensor Temp.: 101.15 Ambient Temp.: 97.61 F
18:49:53.441 -> Average Cal sensor Temp.: 100.91 Ambient Temp.: 97.61 F
18:49:54.993 -> Average Cal sensor Temp.: 100.80 Ambient Temp.: 97.57 F
18:49:56.556 -> Average Cal sensor Temp.: 100.55 Ambient Temp.: 97.57 F
18:49:58.108 -> Average Cal sensor Temp.: 100.38 Ambient Temp.: 97.54 F
18:49:59.700 -> Average Cal sensor Temp.: 100.38 Ambient Temp.: 97.54 F
18:50:01.260 -> Average Cal sensor Temp.: 100.04 Ambient Temp.: 97.54 F
18:50:02.813 -> Average Cal sensor Temp.: 99.90 Ambient Temp.: 97.47 F
18:50:04.364 -> Average Cal sensor Temp.: 99.42 Ambient Temp.: 97.47 F
18:50:05.928 -> Average Cal sensor Temp.: 99.36 Ambient Temp.: 97.43 F
18:50:07.487 -> Average Cal sensor Temp.: 99.51 Ambient Temp.: 97.39 F
18:50:09.039 -> Average Cal sensor Temp.: 99.53 Ambient Temp.: 97.43 F
18:50:10.596 -> Average Cal sensor Temp.: 99.77 Ambient Temp.: 97.39 F
18:50:12.184 -> Average Cal sensor Temp.: 99.33 Ambient Temp.: 97.36 F
18:50:13.739 -> Average Cal sensor Temp.: 99.38 Ambient Temp.: 97.39 F
18:50:15.295 -> Average Cal sensor Temp.: 99.33 Ambient Temp.: 97.36 F
18:50:16.848 -> Average Cal sensor Temp.: 99.43 Ambient Temp.: 97.32 F
18:50:18.399 -> Average Cal sensor Temp.: 99.65 Ambient Temp.: 97.29 F
18:50:19.963 -> Average Cal sensor Temp.: 99.70 Ambient Temp.: 97.29 F
18:50:21.521 -> Average Cal sensor Temp.: 99.74 Ambient Temp.: 97.21 F
18:50:23.073 -> Average Cal sensor Temp.: 99.17 Ambient Temp.: 97.29 F
18:50:24.629 -> Average Cal sensor Temp.: 99.59 Ambient Temp.: 97.21 F
18:50:26.186 -> Average Cal sensor Temp.: 99.46 Ambient Temp.: 97.18 F
18:50:27.774 -> Average Cal sensor Temp.: 99.59 Ambient Temp.: 97.18 F
18:50:29.331 -> Average Cal sensor Temp.: 99.26 Ambient Temp.: 97.14 F
18:50:30.881 -> Average Cal sensor Temp.: 99.26 Ambient Temp.: 97.14 F
18:50:32.435 -> Average Cal sensor Temp.: 99.26 Ambient Temp.: 97.11 F
18:50:33.991 -> Average Cal sensor Temp.: 99.40 Ambient Temp.: 97.14 F
18:50:35.550 -> Average Cal sensor Temp.: 99.71 Ambient Temp.: 97.11 F
18:50:37.104 -> Average Cal sensor Temp.: 99.65 Ambient Temp.: 97.03 F
18:50:38.660 -> Average Cal sensor Temp.: 99.41 Ambient Temp.: 97.11 F
18:50:40.212 -> Average Cal sensor Temp.: 98.88 Ambient Temp.: 97.07 F
18:50:41.771 -> Average Cal sensor Temp.: 99.02 Ambient Temp.: 97.07 F
18:50:43.326 -> Average Cal sensor Temp.: 98.88 Ambient Temp.: 97.03 F
18:50:44.882 -> Average Cal sensor Temp.: 99.37 Ambient Temp.: 97.00 F
18:50:46.441 -> Average Cal sensor Temp.: 99.51 Ambient Temp.: 97.00 F
18:50:47.995 -> Average Cal sensor Temp.: 99.18 Ambient Temp.: 97.03 F
18:50:49.554 -> Average Cal sensor Temp.: 99.53 Ambient Temp.: 97.00 F
18:50:51.142 -> Average Cal sensor Temp.: 99.62 Ambient Temp.: 97.00 F
18:50:52.701 -> Average Cal sensor Temp.: 99.45 Ambient Temp.: 96.93 F
18:50:54.255 -> Average Cal sensor Temp.: 99.49 Ambient Temp.: 96.96 F
18:50:55.813 -> Average Cal sensor Temp.: 99.95 Ambient Temp.: 97.00 F
18:50:57.369 -> Average Cal sensor Temp.: 99.62 Ambient Temp.: 97.00 F
18:50:58.925 -> Average Cal sensor Temp.: 99.59 Ambient Temp.: 96.93 F
18:51:00.486 -> Average Cal sensor Temp.: 99.82 Ambient Temp.: 96.96 F
18:51:02.073 -> Average Cal sensor Temp.: 99.85 Ambient Temp.: 96.93 F
18:51:03.627 -> Average Cal sensor Temp.: 99.75 Ambient Temp.: 96.93 F
18:51:05.185 -> Average Cal sensor Temp.: 99.68 Ambient Temp.: 96.93 F
18:51:06.746 -> Average Cal sensor Temp.: 99.56 Ambient Temp.: 96.93 F
18:51:08.299 -> Average Cal sensor Temp.: 99.27 Ambient Temp.: 96.89 F
18:51:09.855 -> Average Cal sensor Temp.: 99.06 Ambient Temp.: 96.93 F
18:51:11.412 -> Average Cal sensor Temp.: 99.49 Ambient Temp.: 96.89 F
18:51:12.998 -> Average Cal sensor Temp.: 99.70 Ambient Temp.: 96.78 F
18:51:14.553 -> Average Cal sensor Temp.: 99.72 Ambient Temp.: 96.89 F
18:51:16.109 -> Average Cal sensor Temp.: 99.50 Ambient Temp.: 96.89 F
18:51:17.663 -> Average Cal sensor Temp.: 98.83 Ambient Temp.: 96.82 F
18:51:19.218 -> Average Cal sensor Temp.: 99.41 Ambient Temp.: 96.82 F
18:51:20.800 -> Average Cal sensor Temp.: 99.58 Ambient Temp.: 96.78 F
18:51:22.352 -> Average Cal sensor Temp.: 99.67 Ambient Temp.: 96.82 F
18:51:23.900 -> Average Cal sensor Temp.: 99.49 Ambient Temp.: 96.82 F
18:51:25.457 -> Average Cal sensor Temp.: 99.41 Ambient Temp.: 96.78 F
18:51:27.041 -> Average Cal sensor Temp.: 99.44 Ambient Temp.: 96.75 F
18:51:28.564 -> Average Cal sensor Temp.: 99.70 Ambient Temp.: 96.78 F
18:51:30.157 -> Average Cal sensor Temp.: 99.74 Ambient Temp.: 96.78 F
18:51:31.711 -> Average Cal sensor Temp.: 99.78 Ambient Temp.: 96.78 F
18:51:33.261 -> Average Cal sensor Temp.: 99.40 Ambient Temp.: 96.75 F
18:51:34.814 -> Average Cal sensor Temp.: 99.54 Ambient Temp.: 96.71 F
18:51:36.401 -> Average Cal sensor Temp.: 99.46 Ambient Temp.: 96.75 F
18:51:37.953 -> Average Cal sensor Temp.: 99.57 Ambient Temp.: 96.71 F
18:51:39.511 -> Average Cal sensor Temp.: 99.52 Ambient Temp.: 96.75 F
18:51:41.066 -> Average Cal sensor Temp.: 99.66 Ambient Temp.: 96.78 F
18:51:42.626 -> Average Cal sensor Temp.: 99.14 Ambient Temp.: 96.75 F
18:51:44.178 -> Average Cal sensor Temp.: 99.07 Ambient Temp.: 96.71 F
18:51:45.731 -> Average Cal sensor Temp.: 99.28 Ambient Temp.: 96.67 F
18:51:47.288 -> Average Cal sensor Temp.: 98.95 Ambient Temp.: 96.71 F
18:51:48.874 -> Average Cal sensor Temp.: 98.65 Ambient Temp.: 96.67 F
18:51:50.429 -> Average Cal sensor Temp.: 98.78 Ambient Temp.: 96.71 F
18:51:51.983 -> Average Cal sensor Temp.: 99.18 Ambient Temp.: 96.64 F
18:51:53.533 -> Average Cal sensor Temp.: 99.48 Ambient Temp.: 96.64 F
18:51:55.088 -> Average Cal sensor Temp.: 99.27 Ambient Temp.: 96.64 F
18:51:56.643 -> Average Cal sensor Temp.: 99.23 Ambient Temp.: 96.64 F
18:51:58.195 -> Average Cal sensor Temp.: 98.95 Ambient Temp.: 96.67 F
18:51:59.753 -> Average Cal sensor Temp.: 99.05 Ambient Temp.: 96.64 F
18:52:01.308 -> Average Cal sensor Temp.: 99.17 Ambient Temp.: 96.60 F
18:52:02.864 -> Average Cal sensor Temp.: 99.28 Ambient Temp.: 96.64 F
18:52:04.423 -> Average Cal sensor Temp.: 99.51 Ambient Temp.: 96.60 F
18:52:05.980 -> Average Cal sensor Temp.: 99.54 Ambient Temp.: 96.67 F
18:52:07.567 -> Average Cal sensor Temp.: 99.64 Ambient Temp.: 96.60 F
18:52:09.125 -> Average Cal sensor Temp.: 99.33 Ambient Temp.: 96.60 F
18:52:10.675 -> Average Cal sensor Temp.: 99.80 Ambient Temp.: 96.64 F
18:52:12.233 -> Average Cal sensor Temp.: 100.04 Ambient Temp.: 96.64 F
18:52:13.825 -> Average Cal sensor Temp.: 100.13 Ambient Temp.: 96.64 F
18:52:15.380 -> Average Cal sensor Temp.: 99.95 Ambient Temp.: 96.60 F
18:52:16.938 -> Average Cal sensor Temp.: 99.95 Ambient Temp.: 96.67 F
18:52:18.496 -> Average Cal sensor Temp.: 99.54 Ambient Temp.: 96.64 F
18:52:20.051 -> Average Cal sensor Temp.: 99.71 Ambient Temp.: 96.64 F
18:52:21.609 -> Average Cal sensor Temp.: 99.69 Ambient Temp.: 96.67 F
18:52:23.204 -> Average Cal sensor Temp.: 99.35 Ambient Temp.: 96.57 F
18:52:24.754 -> Average Cal sensor Temp.: 99.38 Ambient Temp.: 96.60 F
18:52:26.312 -> Average Cal sensor Temp.: 99.82 Ambient Temp.: 96.57 F
18:52:27.868 -> Average Cal sensor Temp.: 99.99 Ambient Temp.: 96.60 F
18:52:29.422 -> Average Cal sensor Temp.: 100.00 Ambient Temp.: 96.57 F
18:52:31.010 -> Average Cal sensor Temp.: 99.15 Ambient Temp.: 96.60 F
18:52:32.564 -> Average Cal sensor Temp.: 99.14 Ambient Temp.: 96.57 F
18:52:34.117 -> Average Cal sensor Temp.: 99.44 Ambient Temp.: 96.60 F
18:52:35.676 -> Average Cal sensor Temp.: 99.62 Ambient Temp.: 96.60 F
18:52:37.229 -> Average Cal sensor Temp.: 99.72 Ambient Temp.: 96.57 F
18:52:38.780 -> Average Cal sensor Temp.: 99.73 Ambient Temp.: 96.60 F
18:52:40.369 -> Average Cal sensor Temp.: 99.25 Ambient Temp.: 96.64 F
18:52:41.915 -> Average Cal sensor Temp.: 99.23 Ambient Temp.: 96.60 F
18:52:43.475 -> Average Cal sensor Temp.: 99.62 Ambient Temp.: 96.60 F
18:52:45.030 -> Average Cal sensor Temp.: 99.90 Ambient Temp.: 96.57 F
18:52:46.592 -> Average Cal sensor Temp.: 99.81 Ambient Temp.: 96.60 F
18:52:48.149 -> Average Cal sensor Temp.: 99.66 Ambient Temp.: 96.60 F
18:52:49.707 -> Average Cal sensor Temp.: 100.08 Ambient Temp.: 96.60 F
18:52:51.296 -> Average Cal sensor Temp.: 100.04 Ambient Temp.: 96.60 F
18:52:52.852 -> Average Cal sensor Temp.: 99.24 Ambient Temp.: 96.64 F
18:52:54.410 -> Average Cal sensor Temp.: 99.65 Ambient Temp.: 96.60 F
18:52:55.966 -> Average Cal sensor Temp.: 99.64 Ambient Temp.: 96.57 F
18:52:57.524 -> Average Cal sensor Temp.: 99.93 Ambient Temp.: 96.60 F
18:52:59.111 -> Average Cal sensor Temp.: 100.13 Ambient Temp.: 96.57 F
18:53:00.666 -> Average Cal sensor Temp.: 99.90 Ambient Temp.: 96.57 F
18:53:02.220 -> Average Cal sensor Temp.: 100.01 Ambient Temp.: 96.60 F
18:53:03.808 -> Average Cal sensor Temp.: 99.47 Ambient Temp.: 96.64 F
18:53:05.366 -> Average Cal sensor Temp.: 99.43 Ambient Temp.: 96.57 F
18:53:06.923 -> Average Cal sensor Temp.: 99.59 Ambient Temp.: 96.64 F
18:53:08.485 -> Average Cal sensor Temp.: 99.53 Ambient Temp.: 96.57 F
18:53:10.039 -> Average Cal sensor Temp.: 99.79 Ambient Temp.: 96.57 F
18:53:11.581 -> Average Cal sensor Temp.: 99.55 Ambient Temp.: 96.60 F
18:53:13.174 -> Average Cal sensor Temp.: 99.08 Ambient Temp.: 96.57 F
18:53:14.728 -> Average Cal sensor Temp.: 98.52 Ambient Temp.: 96.53 F
18:53:16.257 -> Average Cal sensor Temp.: 98.53 Ambient Temp.: 96.49 F
18:53:17.813 -> Average Cal sensor Temp.: 98.74 Ambient Temp.: 96.53 F
18:53:19.372 -> Average Cal sensor Temp.: 98.98 Ambient Temp.: 96.57 F
18:53:20.956 -> Average Cal sensor Temp.: 99.11 Ambient Temp.: 96.57 F
18:53:22.510 -> Average Cal sensor Temp.: 99.56 Ambient Temp.: 96.57 F
18:53:24.070 -> Average Cal sensor Temp.: 99.22 Ambient Temp.: 96.60 F
18:53:25.623 -> Average Cal sensor Temp.: 99.33 Ambient Temp.: 96.57 F
18:53:27.181 -> Average Cal sensor Temp.: 99.68 Ambient Temp.: 96.57 F
18:53:28.736 -> Average Cal sensor Temp.: 99.62 Ambient Temp.: 96.57 F
18:53:30.295 -> Average Cal sensor Temp.: 99.87 Ambient Temp.: 96.57 F
18:53:31.848 -> Average Cal sensor Temp.: 99.51 Ambient Temp.: 96.49 F
18:53:33.408 -> Average Cal sensor Temp.: 99.77 Ambient Temp.: 96.57 F
18:53:34.998 -> Average Cal sensor Temp.: 99.92 Ambient Temp.: 96.49 F
18:53:36.553 -> Average Cal sensor Temp.: 99.82 Ambient Temp.: 96.49 F
18:53:38.112 -> Average Cal sensor Temp.: 99.55 Ambient Temp.: 96.57 F
18:53:39.668 -> Average Cal sensor Temp.: 99.69 Ambient Temp.: 96.53 F
18:53:41.224 -> Average Cal sensor Temp.: 99.69 Ambient Temp.: 96.53 F
18:53:42.812 -> Average Cal sensor Temp.: 99.74 Ambient Temp.: 96.53 F
18:53:44.361 -> Average Cal sensor Temp.: 99.49 Ambient Temp.: 96.53 F
18:53:45.918 -> Average Cal sensor Temp.: 99.42 Ambient Temp.: 96.60 F
18:53:47.472 -> Average Cal sensor Temp.: 99.10 Ambient Temp.: 96.53 F
18:53:49.024 -> Average Cal sensor Temp.: 99.86 Ambient Temp.: 96.53 F
18:53:50.612 -> Average Cal sensor Temp.: 99.09 Ambient Temp.: 96.53 F
18:53:52.163 -> Average Cal sensor Temp.: 99.47 Ambient Temp.: 96.57 F
18:53:53.715 -> Average Cal sensor Temp.: 99.83 Ambient Temp.: 96.57 F
18:53:55.270 -> Average Cal sensor Temp.: 100.04 Ambient Temp.: 96.57 F
18:53:56.824 -> Average Cal sensor Temp.: 99.79 Ambient Temp.: 96.57 F
18:53:58.416 -> Average Cal sensor Temp.: 99.67 Ambient Temp.: 96.57 F
18:53:59.976 -> Average Cal sensor Temp.: 99.84 Ambient Temp.: 96.57 F
18:54:01.528 -> Average Cal sensor Temp.: 99.45 Ambient Temp.: 96.60 F
18:54:03.083 -> Average Cal sensor Temp.: 99.84 Ambient Temp.: 96.57 F
18:54:04.638 -> Average Cal sensor Temp.: 99.80 Ambient Temp.: 96.53 F
18:54:06.194 -> Average Cal sensor Temp.: 99.85 Ambient Temp.: 96.53 F
18:54:07.754 -> Average Cal sensor Temp.: 99.85 Ambient Temp.: 96.57 F
18:54:09.347 -> Average Cal sensor Temp.: 99.67 Ambient Temp.: 96.57 F
18:54:10.905 -> Average Cal sensor Temp.: 99.66 Ambient Temp.: 96.57 F
18:54:12.457 -> Average Cal sensor Temp.: 99.62 Ambient Temp.: 96.53 F
18:54:14.010 -> Average Cal sensor Temp.: 99.44 Ambient Temp.: 96.53 F
18:54:15.568 -> Average Cal sensor Temp.: 99.45 Ambient Temp.: 96.57 F
18:54:17.126 -> Average Cal sensor Temp.: 99.45 Ambient Temp.: 96.53 F
18:54:18.681 -> Average Cal sensor Temp.: 99.52 Ambient Temp.: 96.49 F
18:54:20.270 -> Average Cal sensor Temp.: 99.84 Ambient Temp.: 96.57 F
18:54:21.821 -> Average Cal sensor Temp.: 100.26 Ambient Temp.: 96.57 F
18:54:23.383 -> Average Cal sensor Temp.: 99.91 Ambient Temp.: 96.53 F
18:54:24.938 -> Average Cal sensor Temp.: 99.54 Ambient Temp.: 96.53 F
18:54:26.496 -> Average Cal sensor Temp.: 99.85 Ambient Temp.: 96.57 F
18:54:28.056 -> Average Cal sensor Temp.: 99.61 Ambient Temp.: 96.64 F
18:54:29.614 -> Average Cal sensor Temp.: 99.48 Ambient Temp.: 96.53 F
18:54:31.173 -> Average Cal sensor Temp.: 99.56 Ambient Temp.: 96.60 F
18:54:32.733 -> Average Cal sensor Temp.: 99.40 Ambient Temp.: 96.57 F
18:54:34.323 -> Average Cal sensor Temp.: 98.84 Ambient Temp.: 96.53 F
18:54:35.881 -> Average Cal sensor Temp.: 99.41 Ambient Temp.: 96.57 F
18:54:37.432 -> Average Cal sensor Temp.: 99.06 Ambient Temp.: 96.57 F
18:54:38.985 -> Average Cal sensor Temp.: 99.92 Ambient Temp.: 96.53 F
18:54:40.537 -> Average Cal sensor Temp.: 99.76 Ambient Temp.: 96.60 F
18:54:42.089 -> Average Cal sensor Temp.: 99.63 Ambient Temp.: 96.57 F
18:54:43.677 -> Average Cal sensor Temp.: 99.89 Ambient Temp.: 96.57 F
18:54:45.233 -> Average Cal sensor Temp.: 100.03 Ambient Temp.: 96.57 F
18:54:46.792 -> Average Cal sensor Temp.: 99.67 Ambient Temp.: 96.57 F
18:54:48.342 -> Average Cal sensor Temp.: 99.64 Ambient Temp.: 96.53 F
18:54:49.935 -> Average Cal sensor Temp.: 99.64 Ambient Temp.: 96.53 F
18:54:51.489 -> Average Cal sensor Temp.: 99.80 Ambient Temp.: 96.57 F
18:54:53.043 -> Average Cal sensor Temp.: 100.05 Ambient Temp.: 96.60 F
18:54:54.596 -> Average Cal sensor Temp.: 100.56 Ambient Temp.: 96.57 F
18:54:56.189 -> Average Cal sensor Temp.: 100.31 Ambient Temp.: 96.60 F
18:54:57.743 -> Average Cal sensor Temp.: 100.26 Ambient Temp.: 96.60 F
18:54:59.296 -> Average Cal sensor Temp.: 100.44 Ambient Temp.: 96.57 F
18:55:00.883 -> Average Cal sensor Temp.: 100.64 Ambient Temp.: 96.64 F
18:55:02.442 -> Average Cal sensor Temp.: 100.81 Ambient Temp.: 96.57 F
18:55:03.998 -> Average Cal sensor Temp.: 100.42 Ambient Temp.: 96.57 F
18:55:05.589 -> Average Cal sensor Temp.: 99.98 Ambient Temp.: 96.60 F
18:55:07.150 -> Average Cal sensor Temp.: 99.95 Ambient Temp.: 96.57 F
18:55:08.703 -> Average Cal sensor Temp.: 99.86 Ambient Temp.: 96.57 F
18:55:10.262 -> Average Cal sensor Temp.: 99.73 Ambient Temp.: 96.57 F
18:55:11.817 -> Average Cal sensor Temp.: 100.06 Ambient Temp.: 96.57 F
18:55:13.402 -> Average Cal sensor Temp.: 100.34 Ambient Temp.: 96.57 F
18:55:14.961 -> Average Cal sensor Temp.: 100.29 Ambient Temp.: 96.57 F
18:55:16.519 -> Average Cal sensor Temp.: 100.29 Ambient Temp.: 96.60 F
18:55:18.077 -> Average Cal sensor Temp.: 100.27 Ambient Temp.: 96.60 F
18:55:19.641 -> Average Cal sensor Temp.: 99.91 Ambient Temp.: 96.57 F
18:55:21.229 -> Average Cal sensor Temp.: 100.15 Ambient Temp.: 96.60 F
18:55:22.779 -> Average Cal sensor Temp.: 100.33 Ambient Temp.: 96.60 F
18:55:24.332 -> Average Cal sensor Temp.: 100.31 Ambient Temp.: 96.60 F
18:55:25.896 -> Average Cal sensor Temp.: 99.60 Ambient Temp.: 96.64 F
18:55:27.487 -> Average Cal sensor Temp.: 100.04 Ambient Temp.: 96.57 F
18:55:29.044 -> Average Cal sensor Temp.: 100.51 Ambient Temp.: 96.64 F
18:55:30.594 -> Average Cal sensor Temp.: 100.63 Ambient Temp.: 96.60 F
18:55:32.177 -> Average Cal sensor Temp.: 100.44 Ambient Temp.: 96.60 F
18:55:33.739 -> Average Cal sensor Temp.: 100.07 Ambient Temp.: 96.60 F
18:55:35.297 -> Average Cal sensor Temp.: 100.27 Ambient Temp.: 96.60 F
18:55:36.886 -> Average Cal sensor Temp.: 100.42 Ambient Temp.: 96.64 F
18:55:38.437 -> Average Cal sensor Temp.: 100.54 Ambient Temp.: 96.60 F
18:55:39.994 -> Average Cal sensor Temp.: 100.44 Ambient Temp.: 96.57 F
18:55:41.582 -> Average Cal sensor Temp.: 100.56 Ambient Temp.: 96.64 F
18:55:43.136 -> Average Cal sensor Temp.: 100.37 Ambient Temp.: 96.60 F
18:55:44.693 -> Average Cal sensor Temp.: 100.67 Ambient Temp.: 96.64 F
18:55:46.281 -> Average Cal sensor Temp.: 100.62 Ambient Temp.: 96.60 F
18:55:47.831 -> Average Cal sensor Temp.: 100.40 Ambient Temp.: 96.64 F
18:55:49.392 -> Average Cal sensor Temp.: 100.44 Ambient Temp.: 96.64 F
18:55:50.981 -> Average Cal sensor Temp.: 100.45 Ambient Temp.: 96.60 F
18:55:52.534 -> Average Cal sensor Temp.: 100.91 Ambient Temp.: 96.64 F
18:55:54.121 -> Average Cal sensor Temp.: 100.74 Ambient Temp.: 96.60 F
18:55:55.682 -> Average Cal sensor Temp.: 100.31 Ambient Temp.: 96.64 F
18:55:57.240 -> Average Cal sensor Temp.: 100.63 Ambient Temp.: 96.64 F
18:55:58.796 -> Average Cal sensor Temp.: 100.76 Ambient Temp.: 96.60 F
18:56:00.351 -> Average Cal sensor Temp.: 100.79 Ambient Temp.: 96.60 F
18:56:01.941 -> Average Cal sensor Temp.: 100.50 Ambient Temp.: 96.60 F
18:56:03.498 -> Average Cal sensor Temp.: 100.44 Ambient Temp.: 96.64 F
18:56:05.052 -> Average Cal sensor Temp.: 99.98 Ambient Temp.: 96.60 F
18:56:06.644 -> Average Cal sensor Temp.: 100.63 Ambient Temp.: 96.71 F
18:56:08.189 -> Average Cal sensor Temp.: 100.86 Ambient Temp.: 96.67 F
18:56:09.773 -> Average Cal sensor Temp.: 100.80 Ambient Temp.: 96.64 F
18:56:11.327 -> Average Cal sensor Temp.: 100.49 Ambient Temp.: 96.67 F
18:56:12.915 -> Average Cal sensor Temp.: 100.79 Ambient Temp.: 96.67 F
18:56:14.468 -> Average Cal sensor Temp.: 101.00 Ambient Temp.: 96.64 F
18:56:16.022 -> Average Cal sensor Temp.: 100.62 Ambient Temp.: 96.67 F
18:56:17.580 -> Average Cal sensor Temp.: 100.62 Ambient Temp.: 96.67 F
18:56:19.173 -> Average Cal sensor Temp.: 100.96 Ambient Temp.: 96.67 F
18:56:20.728 -> Average Cal sensor Temp.: 101.00 Ambient Temp.: 96.67 F
18:56:22.284 -> Average Cal sensor Temp.: 100.76 Ambient Temp.: 96.67 F
18:56:23.878 -> Average Cal sensor Temp.: 100.56 Ambient Temp.: 96.67 F
18:56:25.429 -> Average Cal sensor Temp.: 100.91 Ambient Temp.: 96.71 F
18:56:26.988 -> Average Cal sensor Temp.: 100.58 Ambient Temp.: 96.64 F
18:56:28.575 -> Average Cal sensor Temp.: 100.85 Ambient Temp.: 96.67 F
18:56:30.136 -> Average Cal sensor Temp.: 100.64 Ambient Temp.: 96.71 F
18:56:31.692 -> Average Cal sensor Temp.: 100.90 Ambient Temp.: 96.75 F
18:56:33.247 -> Average Cal sensor Temp.: 100.75 Ambient Temp.: 96.71 F
18:56:34.832 -> Average Cal sensor Temp.: 100.70 Ambient Temp.: 96.71 F
18:56:36.385 -> Average Cal sensor Temp.: 100.87 Ambient Temp.: 96.75 F
18:56:37.938 -> Average Cal sensor Temp.: 100.46 Ambient Temp.: 96.71 F
18:56:39.523 -> Average Cal sensor Temp.: 100.84 Ambient Temp.: 96.75 F
18:56:41.077 -> Average Cal sensor Temp.: 100.42 Ambient Temp.: 96.71 F
18:56:42.634 -> Average Cal sensor Temp.: 100.57 Ambient Temp.: 96.78 F
18:56:44.221 -> Average Cal sensor Temp.: 100.32 Ambient Temp.: 96.71 F
18:56:45.777 -> Average Cal sensor Temp.: 100.41 Ambient Temp.: 96.71 F
18:56:47.332 -> Average Cal sensor Temp.: 100.00 Ambient Temp.: 96.71 F
18:56:48.918 -> Average Cal sensor Temp.: 100.34 Ambient Temp.: 96.67 F
18:56:50.477 -> Average Cal sensor Temp.: 100.53 Ambient Temp.: 96.75 F
18:56:52.033 -> Average Cal sensor Temp.: 100.89 Ambient Temp.: 96.75 F
18:56:53.617 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 96.71 F
18:56:55.170 -> Average Cal sensor Temp.: 101.10 Ambient Temp.: 96.75 F
18:56:56.726 -> Average Cal sensor Temp.: 100.76 Ambient Temp.: 96.67 F
18:56:58.279 -> Average Cal sensor Temp.: 100.40 Ambient Temp.: 96.71 F
18:56:59.872 -> Average Cal sensor Temp.: 100.73 Ambient Temp.: 96.71 F
18:57:01.427 -> Average Cal sensor Temp.: 100.76 Ambient Temp.: 96.71 F
18:57:02.985 -> Average Cal sensor Temp.: 100.60 Ambient Temp.: 96.71 F
18:57:04.572 -> Average Cal sensor Temp.: 100.73 Ambient Temp.: 96.71 F
18:57:06.126 -> Average Cal sensor Temp.: 100.94 Ambient Temp.: 96.75 F
18:57:07.682 -> Average Cal sensor Temp.: 101.12 Ambient Temp.: 96.78 F
18:57:09.233 -> Average Cal sensor Temp.: 100.86 Ambient Temp.: 96.75 F
18:57:10.823 -> Average Cal sensor Temp.: 101.04 Ambient Temp.: 96.75 F
18:57:12.377 -> Average Cal sensor Temp.: 101.07 Ambient Temp.: 96.75 F
18:57:13.936 -> Average Cal sensor Temp.: 100.99 Ambient Temp.: 96.82 F
18:57:15.526 -> Average Cal sensor Temp.: 100.71 Ambient Temp.: 96.89 F
18:57:17.083 -> Average Cal sensor Temp.: 101.04 Ambient Temp.: 96.78 F
18:57:18.633 -> Average Cal sensor Temp.: 101.05 Ambient Temp.: 96.75 F
18:57:20.192 -> Average Cal sensor Temp.: 101.04 Ambient Temp.: 96.78 F
18:57:21.777 -> Average Cal sensor Temp.: 101.22 Ambient Temp.: 96.75 F
18:57:23.331 -> Average Cal sensor Temp.: 100.91 Ambient Temp.: 96.78 F
18:57:24.884 -> Average Cal sensor Temp.: 101.08 Ambient Temp.: 96.78 F
18:57:26.475 -> Average Cal sensor Temp.: 101.15 Ambient Temp.: 96.82 F
18:57:28.030 -> Average Cal sensor Temp.: 100.91 Ambient Temp.: 96.75 F
18:57:29.584 -> Average Cal sensor Temp.: 100.90 Ambient Temp.: 96.78 F
18:57:31.174 -> Average Cal sensor Temp.: 100.66 Ambient Temp.: 96.78 F
18:57:32.733 -> Average Cal sensor Temp.: 100.31 Ambient Temp.: 96.82 F
18:57:34.287 -> Average Cal sensor Temp.: 100.67 Ambient Temp.: 96.75 F
18:57:35.842 -> Average Cal sensor Temp.: 100.89 Ambient Temp.: 96.78 F
18:57:37.430 -> Average Cal sensor Temp.: 101.10 Ambient Temp.: 96.78 F
18:57:38.988 -> Average Cal sensor Temp.: 101.13 Ambient Temp.: 96.89 F
18:57:40.543 -> Average Cal sensor Temp.: 101.40 Ambient Temp.: 96.82 F
18:57:42.129 -> Average Cal sensor Temp.: 101.37 Ambient Temp.: 96.82 F
18:57:43.687 -> Average Cal sensor Temp.: 101.19 Ambient Temp.: 96.78 F
18:57:45.239 -> Average Cal sensor Temp.: 101.34 Ambient Temp.: 96.82 F
18:57:46.796 -> Average Cal sensor Temp.: 101.55 Ambient Temp.: 96.82 F
18:57:48.381 -> Average Cal sensor Temp.: 101.36 Ambient Temp.: 96.89 F
18:57:49.934 -> Average Cal sensor Temp.: 101.00 Ambient Temp.: 96.78 F
18:57:51.492 -> Average Cal sensor Temp.: 101.02 Ambient Temp.: 96.82 F
18:57:53.079 -> Average Cal sensor Temp.: 101.17 Ambient Temp.: 96.82 F
18:57:54.624 -> Average Cal sensor Temp.: 101.24 Ambient Temp.: 96.82 F
18:57:56.176 -> Average Cal sensor Temp.: 101.08 Ambient Temp.: 96.82 F
18:57:57.769 -> Average Cal sensor Temp.: 101.21 Ambient Temp.: 96.82 F
18:57:59.325 -> Average Cal sensor Temp.: 101.11 Ambient Temp.: 96.89 F
18:58:00.879 -> Average Cal sensor Temp.: 101.11 Ambient Temp.: 96.82 F
18:58:02.430 -> Average Cal sensor Temp.: 101.04 Ambient Temp.: 96.89 F
18:58:04.021 -> Average Cal sensor Temp.: 101.03 Ambient Temp.: 96.93 F
18:58:05.581 -> Average Cal sensor Temp.: 101.10 Ambient Temp.: 96.75 F
18:58:07.136 -> Average Cal sensor Temp.: 101.18 Ambient Temp.: 96.82 F
18:58:08.689 -> Average Cal sensor Temp.: 101.21 Ambient Temp.: 96.82 F
18:58:10.281 -> Average Cal sensor Temp.: 101.02 Ambient Temp.: 96.78 F
18:58:11.832 -> Average Cal sensor Temp.: 101.08 Ambient Temp.: 96.82 F
18:58:13.386 -> Average Cal sensor Temp.: 101.21 Ambient Temp.: 96.78 F
18:58:14.976 -> Average Cal sensor Temp.: 100.94 Ambient Temp.: 96.89 F
18:58:16.531 -> Average Cal sensor Temp.: 101.25 Ambient Temp.: 96.82 F
18:58:18.086 -> Average Cal sensor Temp.: 101.12 Ambient Temp.: 96.89 F
18:58:19.641 -> Average Cal sensor Temp.: 100.66 Ambient Temp.: 96.89 F
18:58:21.224 -> Average Cal sensor Temp.: 100.76 Ambient Temp.: 96.93 F
18:58:22.778 -> Average Cal sensor Temp.: 101.03 Ambient Temp.: 96.82 F
18:58:24.363 -> Average Cal sensor Temp.: 101.08 Ambient Temp.: 96.82 F
18:58:25.919 -> Average Cal sensor Temp.: 100.58 Ambient Temp.: 96.82 F
18:58:27.475 -> Average Cal sensor Temp.: 101.17 Ambient Temp.: 96.82 F
18:58:29.033 -> Average Cal sensor Temp.: 101.13 Ambient Temp.: 96.82 F
18:58:30.619 -> Average Cal sensor Temp.: 101.12 Ambient Temp.: 96.78 F
18:58:32.173 -> Average Cal sensor Temp.: 100.32 Ambient Temp.: 96.78 F
18:58:33.723 -> Average Cal sensor Temp.: 100.76 Ambient Temp.: 96.82 F
18:58:35.318 -> Average Cal sensor Temp.: 100.84 Ambient Temp.: 96.82 F
18:58:36.869 -> Average Cal sensor Temp.: 101.19 Ambient Temp.: 96.78 F
18:58:38.418 -> Average Cal sensor Temp.: 100.95 Ambient Temp.: 96.82 F
18:58:40.004 -> Average Cal sensor Temp.: 101.10 Ambient Temp.: 96.82 F
18:58:41.557 -> Average Cal sensor Temp.: 101.16 Ambient Temp.: 96.78 F
18:58:43.143 -> Average Cal sensor Temp.: 101.04 Ambient Temp.: 96.75 F
18:58:44.693 -> Average Cal sensor Temp.: 101.31 Ambient Temp.: 96.82 F
18:58:46.251 -> Average Cal sensor Temp.: 101.69 Ambient Temp.: 96.89 F
18:58:47.833 -> Average Cal sensor Temp.: 101.55 Ambient Temp.: 96.78 F
18:58:49.396 -> Average Cal sensor Temp.: 101.21 Ambient Temp.: 96.82 F
18:58:50.952 -> Average Cal sensor Temp.: 101.10 Ambient Temp.: 96.82 F
18:58:52.504 -> Average Cal sensor Temp.: 101.43 Ambient Temp.: 96.78 F
18:58:54.059 -> Average Cal sensor Temp.: 100.83 Ambient Temp.: 96.82 F
18:58:55.649 -> Average Cal sensor Temp.: 101.15 Ambient Temp.: 96.82 F
18:58:57.203 -> Average Cal sensor Temp.: 101.26 Ambient Temp.: 96.82 F
18:58:58.756 -> Average Cal sensor Temp.: 101.54 Ambient Temp.: 96.82 F
18:59:00.317 -> Average Cal sensor Temp.: 101.37 Ambient Temp.: 96.78 F
18:59:01.908 -> Average Cal sensor Temp.: 101.49 Ambient Temp.: 96.89 F
18:59:03.466 -> Average Cal sensor Temp.: 101.51 Ambient Temp.: 96.82 F
18:59:05.018 -> Average Cal sensor Temp.: 101.68 Ambient Temp.: 96.89 F
18:59:06.578 -> Average Cal sensor Temp.: 101.29 Ambient Temp.: 96.89 F
18:59:08.164 -> Average Cal sensor Temp.: 101.19 Ambient Temp.: 96.82 F
18:59:09.727 -> Average Cal sensor Temp.: 101.22 Ambient Temp.: 96.89 F
18:59:11.274 -> Average Cal sensor Temp.: 100.95 Ambient Temp.: 96.82 F
18:59:12.864 -> Average Cal sensor Temp.: 101.08 Ambient Temp.: 96.89 F
18:59:14.422 -> Average Cal sensor Temp.: 101.17 Ambient Temp.: 96.96 F
18:59:15.985 -> Average Cal sensor Temp.: 101.13 Ambient Temp.: 96.89 F
18:59:17.539 -> Average Cal sensor Temp.: 101.34 Ambient Temp.: 96.89 F
18:59:19.100 -> Average Cal sensor Temp.: 101.56 Ambient Temp.: 96.89 F
18:59:20.692 -> Average Cal sensor Temp.: 100.82 Ambient Temp.: 96.82 F
18:59:22.248 -> Average Cal sensor Temp.: 101.33 Ambient Temp.: 96.89 F
18:59:23.800 -> Average Cal sensor Temp.: 101.21 Ambient Temp.: 96.89 F
18:59:25.358 -> Average Cal sensor Temp.: 101.56 Ambient Temp.: 96.93 F
18:59:26.952 -> Average Cal sensor Temp.: 101.61 Ambient Temp.: 96.93 F
18:59:28.502 -> Average Cal sensor Temp.: 101.84 Ambient Temp.: 96.89 F
18:59:30.063 -> Average Cal sensor Temp.: 101.74 Ambient Temp.: 96.93 F
18:59:31.617 -> Average Cal sensor Temp.: 101.64 Ambient Temp.: 96.93 F
18:59:33.205 -> Average Cal sensor Temp.: 101.34 Ambient Temp.: 96.89 F
18:59:34.756 -> Average Cal sensor Temp.: 101.87 Ambient Temp.: 96.93 F
18:59:36.320 -> Average Cal sensor Temp.: 101.47 Ambient Temp.: 96.93 F
18:59:37.876 -> Average Cal sensor Temp.: 101.48 Ambient Temp.: 96.82 F
18:59:39.466 -> Average Cal sensor Temp.: 101.52 Ambient Temp.: 96.93 F
18:59:41.020 -> Average Cal sensor Temp.: 101.72 Ambient Temp.: 96.96 F
18:59:42.581 -> Average Cal sensor Temp.: 101.79 Ambient Temp.: 97.00 F
18:59:44.131 -> Average Cal sensor Temp.: 101.64 Ambient Temp.: 96.93 F
18:59:45.691 -> Average Cal sensor Temp.: 101.70 Ambient Temp.: 96.93 F
18:59:47.273 -> Average Cal sensor Temp.: 101.51 Ambient Temp.: 96.93 F
18:59:48.827 -> Average Cal sensor Temp.: 100.99 Ambient Temp.: 96.89 F
18:59:50.386 -> Average Cal sensor Temp.: 101.35 Ambient Temp.: 96.93 F
18:59:51.978 -> Average Cal sensor Temp.: 101.16 Ambient Temp.: 96.96 F
18:59:53.531 -> Average Cal sensor Temp.: 101.36 Ambient Temp.: 96.96 F
18:59:55.081 -> Average Cal sensor Temp.: 101.33 Ambient Temp.: 96.93 F
18:59:56.642 -> Average Cal sensor Temp.: 101.19 Ambient Temp.: 96.93 F
18:59:58.234 -> Average Cal sensor Temp.: 101.48 Ambient Temp.: 96.96 F
18:59:59.796 -> Average Cal sensor Temp.: 101.18 Ambient Temp.: 96.89 F
19:00:01.350 -> Average Cal sensor Temp.: 101.44 Ambient Temp.: 96.96 F
19:00:02.903 -> Average Cal sensor Temp.: 100.98 Ambient Temp.: 96.93 F
19:00:04.489 -> Average Cal sensor Temp.: 101.14 Ambient Temp.: 96.96 F
19:00:06.046 -> Average Cal sensor Temp.: 101.31 Ambient Temp.: 96.89 F
19:00:07.603 -> Average Cal sensor Temp.: 101.46 Ambient Temp.: 96.89 F
19:00:09.191 -> Average Cal sensor Temp.: 101.54 Ambient Temp.: 96.93 F
19:00:10.747 -> Average Cal sensor Temp.: 101.89 Ambient Temp.: 96.89 F
19:00:12.305 -> Average Cal sensor Temp.: 101.97 Ambient Temp.: 96.93 F
19:00:13.858 -> Average Cal sensor Temp.: 101.22 Ambient Temp.: 96.82 F
19:00:15.449 -> Average Cal sensor Temp.: 101.33 Ambient Temp.: 96.89 F
19:00:17.005 -> Average Cal sensor Temp.: 101.57 Ambient Temp.: 96.82 F
19:00:18.555 -> Average Cal sensor Temp.: 101.85 Ambient Temp.: 96.89 F
19:00:20.141 -> Average Cal sensor Temp.: 101.67 Ambient Temp.: 96.89 F
19:00:21.692 -> Average Cal sensor Temp.: 101.66 Ambient Temp.: 96.89 F
19:00:23.246 -> Average Cal sensor Temp.: 101.22 Ambient Temp.: 96.93 F
19:00:24.831 -> Average Cal sensor Temp.: 101.43 Ambient Temp.: 96.78 F
19:00:26.388 -> Average Cal sensor Temp.: 102.08 Ambient Temp.: 96.89 F
19:00:27.946 -> Average Cal sensor Temp.: 101.81 Ambient Temp.: 96.89 F
19:00:29.538 -> Average Cal sensor Temp.: 102.16 Ambient Temp.: 96.93 F
19:00:31.092 -> Average Cal sensor Temp.: 102.03 Ambient Temp.: 96.89 F
19:00:32.646 -> Average Cal sensor Temp.: 101.88 Ambient Temp.: 96.89 F
19:00:34.201 -> Average Cal sensor Temp.: 101.68 Ambient Temp.: 96.89 F
19:00:35.791 -> Average Cal sensor Temp.: 101.31 Ambient Temp.: 96.93 F
19:00:37.348 -> Average Cal sensor Temp.: 101.76 Ambient Temp.: 96.96 F
19:00:38.906 -> Average Cal sensor Temp.: 101.42 Ambient Temp.: 97.00 F
19:00:40.461 -> Average Cal sensor Temp.: 101.66 Ambient Temp.: 96.93 F
19:00:42.048 -> Average Cal sensor Temp.: 101.34 Ambient Temp.: 96.96 F
19:00:43.602 -> Average Cal sensor Temp.: 101.92 Ambient Temp.: 96.93 F
19:00:45.160 -> Average Cal sensor Temp.: 102.10 Ambient Temp.: 96.89 F
19:00:46.750 -> Average Cal sensor Temp.: 101.97 Ambient Temp.: 96.96 F
19:00:48.305 -> Average Cal sensor Temp.: 101.81 Ambient Temp.: 96.93 F
19:00:49.864 -> Average Cal sensor Temp.: 101.88 Ambient Temp.: 96.93 F
19:00:51.419 -> Average Cal sensor Temp.: 101.76 Ambient Temp.: 96.96 F
19:00:53.001 -> Average Cal sensor Temp.: 101.89 Ambient Temp.: 96.93 F
19:00:54.555 -> Average Cal sensor Temp.: 101.97 Ambient Temp.: 96.96 F
19:00:56.112 -> Average Cal sensor Temp.: 102.10 Ambient Temp.: 96.96 F
19:00:57.703 -> Average Cal sensor Temp.: 101.95 Ambient Temp.: 96.93 F
19:00:59.255 -> Average Cal sensor Temp.: 101.81 Ambient Temp.: 96.96 F
19:01:00.813 -> Average Cal sensor Temp.: 101.90 Ambient Temp.: 96.96 F
19:01:02.367 -> Average Cal sensor Temp.: 101.75 Ambient Temp.: 96.96 F
19:01:03.950 -> Average Cal sensor Temp.: 101.49 Ambient Temp.: 97.03 F
19:01:05.502 -> Average Cal sensor Temp.: 101.93 Ambient Temp.: 96.96 F
19:01:07.060 -> Average Cal sensor Temp.: 101.99 Ambient Temp.: 96.96 F
19:01:08.645 -> Average Cal sensor Temp.: 101.97 Ambient Temp.: 96.93 F
19:01:10.196 -> Average Cal sensor Temp.: 102.10 Ambient Temp.: 97.03 F
19:01:11.754 -> Average Cal sensor Temp.: 102.22 Ambient Temp.: 97.00 F
19:01:13.346 -> Average Cal sensor Temp.: 102.19 Ambient Temp.: 96.96 F
19:01:14.895 -> Average Cal sensor Temp.: 101.99 Ambient Temp.: 97.00 F
19:01:16.447 -> Average Cal sensor Temp.: 102.30 Ambient Temp.: 96.96 F
19:01:18.040 -> Average Cal sensor Temp.: 101.92 Ambient Temp.: 97.03 F
19:01:19.601 -> Average Cal sensor Temp.: 101.88 Ambient Temp.: 97.03 F
19:01:21.156 -> Average Cal sensor Temp.: 102.17 Ambient Temp.: 97.00 F
19:01:22.708 -> Average Cal sensor Temp.: 102.49 Ambient Temp.: 97.00 F
19:01:24.297 -> Average Cal sensor Temp.: 102.26 Ambient Temp.: 97.03 F
19:01:25.847 -> Average Cal sensor Temp.: 102.50 Ambient Temp.: 97.03 F
19:01:27.426 -> Average Cal sensor Temp.: 102.51 Ambient Temp.: 97.00 F
19:01:28.979 -> Average Cal sensor Temp.: 102.15 Ambient Temp.: 97.03 F
19:01:30.535 -> Average Cal sensor Temp.: 101.55 Ambient Temp.: 97.03 F
19:01:32.088 -> Average Cal sensor Temp.: 102.07 Ambient Temp.: 97.00 F
19:01:33.677 -> Average Cal sensor Temp.: 101.47 Ambient Temp.: 97.07 F
19:01:35.230 -> Average Cal sensor Temp.: 101.85 Ambient Temp.: 97.03 F
19:01:36.792 -> Average Cal sensor Temp.: 101.23 Ambient Temp.: 97.00 F
19:01:38.342 -> Average Cal sensor Temp.: 101.66 Ambient Temp.: 97.00 F
19:01:39.930 -> Average Cal sensor Temp.: 101.66 Ambient Temp.: 97.00 F
19:01:41.485 -> Average Cal sensor Temp.: 101.46 Ambient Temp.: 97.00 F
19:01:43.035 -> Average Cal sensor Temp.: 101.88 Ambient Temp.: 97.00 F
19:01:44.627 -> Average Cal sensor Temp.: 101.72 Ambient Temp.: 97.00 F
19:01:46.186 -> Average Cal sensor Temp.: 102.09 Ambient Temp.: 97.00 F
19:01:47.743 -> Average Cal sensor Temp.: 101.83 Ambient Temp.: 96.96 F
19:01:49.296 -> Average Cal sensor Temp.: 102.06 Ambient Temp.: 96.96 F
19:01:50.879 -> Average Cal sensor Temp.: 101.34 Ambient Temp.: 97.00 F
19:01:52.429 -> Average Cal sensor Temp.: 101.68 Ambient Temp.: 97.03 F
19:01:54.016 -> Average Cal sensor Temp.: 101.79 Ambient Temp.: 96.96 F
19:01:55.569 -> Average Cal sensor Temp.: 101.41 Ambient Temp.: 96.96 F
19:01:57.124 -> Average Cal sensor Temp.: 101.88 Ambient Temp.: 97.00 F
19:01:58.703 -> Average Cal sensor Temp.: 102.04 Ambient Temp.: 96.96 F
19:02:00.261 -> Average Cal sensor Temp.: 102.06 Ambient Temp.: 97.00 F
19:02:01.817 -> Average Cal sensor Temp.: 102.20 Ambient Temp.: 97.00 F
19:02:03.403 -> Average Cal sensor Temp.: 102.05 Ambient Temp.: 97.00 F
19:02:04.961 -> Average Cal sensor Temp.: 101.99 Ambient Temp.: 97.07 F
19:02:05.929 -> Ambient Temp.: 97.00 F
19:02:06.484 -> Ambient Temp.: 97.00 F
19:02:07.071 -> Ambient Temp.: 97.00 F
19:02:08.625 -> Average Cal sensor Temp.: 102.33 Ambient Temp.: 97.00 F
19:02:10.186 -> Average Cal sensor Temp.: 102.61 Ambient Temp.: 96.96 F
19:02:11.772 -> Average Cal sensor Temp.: 101.68 Ambient Temp.: 97.03 F
19:02:13.326 -> Average Cal sensor Temp.: 102.04 Ambient Temp.: 97.00 F
19:02:14.876 -> Average Cal sensor Temp.: 101.53 Ambient Temp.: 96.96 F
19:02:16.466 -> Average Cal sensor Temp.: 102.14 Ambient Temp.: 97.00 F
19:02:18.024 -> Average Cal sensor Temp.: 101.87 Ambient Temp.: 96.96 F
19:02:19.580 -> Average Cal sensor Temp.: 101.97 Ambient Temp.: 96.96 F
19:02:21.139 -> Average Cal sensor Temp.: 102.06 Ambient Temp.: 97.00 F
19:02:22.727 -> Average Cal sensor Temp.: 102.15 Ambient Temp.: 97.03 F
19:02:24.286 -> Average Cal sensor Temp.: 102.51 Ambient Temp.: 97.00 F
19:02:25.846 -> Average Cal sensor Temp.: 101.71 Ambient Temp.: 97.03 F
19:02:27.402 -> Average Cal sensor Temp.: 101.96 Ambient Temp.: 97.00 F
19:02:28.951 -> Average Cal sensor Temp.: 101.90 Ambient Temp.: 97.03 F
19:02:30.543 -> Average Cal sensor Temp.: 102.02 Ambient Temp.: 97.00 F
19:02:32.094 -> Average Cal sensor Temp.: 102.50 Ambient Temp.: 96.96 F
19:02:33.648 -> Average Cal sensor Temp.: 103.07 Ambient Temp.: 97.03 F
19:02:35.237 -> Average Cal sensor Temp.: 102.76 Ambient Temp.: 97.03 F
19:02:36.789 -> Average Cal sensor Temp.: 102.72 Ambient Temp.: 97.00 F
19:02:38.339 -> Average Cal sensor Temp.: 102.47 Ambient Temp.: 97.03 F
19:02:39.933 -> Average Cal sensor Temp.: 102.43 Ambient Temp.: 97.00 F
19:02:41.482 -> Average Cal sensor Temp.: 101.86 Ambient Temp.: 97.03 F
19:02:43.038 -> Average Cal sensor Temp.: 101.66 Ambient Temp.: 97.03 F
19:02:44.597 -> Average Cal sensor Temp.: 101.44 Ambient Temp.: 97.07 F
19:02:46.186 -> Average Cal sensor Temp.: 101.77 Ambient Temp.: 97.03 F
19:02:47.741 -> Average Cal sensor Temp.: 102.11 Ambient Temp.: 97.03 F
19:02:49.297 -> Average Cal sensor Temp.: 102.39 Ambient Temp.: 97.00 F
19:02:50.853 -> Average Cal sensor Temp.: 102.15 Ambient Temp.: 97.00 F
19:02:52.409 -> Average Cal sensor Temp.: 102.32 Ambient Temp.: 97.00 F
19:02:53.989 -> Average Cal sensor Temp.: 101.77 Ambient Temp.: 97.00 F
19:02:55.548 -> Average Cal sensor Temp.: 102.38 Ambient Temp.: 97.03 F
19:02:57.103 -> Average Cal sensor Temp.: 102.34 Ambient Temp.: 97.03 F
19:02:58.691 -> Average Cal sensor Temp.: 102.59 Ambient Temp.: 97.00 F
19:03:00.243 -> Average Cal sensor Temp.: 102.42 Ambient Temp.: 97.03 F
19:03:01.793 -> Average Cal sensor Temp.: 101.82 Ambient Temp.: 97.07 F
19:03:03.378 -> Average Cal sensor Temp.: 101.63 Ambient Temp.: 97.03 F
19:03:04.930 -> Average Cal sensor Temp.: 102.11 Ambient Temp.: 97.00 F
19:03:06.485 -> Average Cal sensor Temp.: 102.32 Ambient Temp.: 97.07 F
19:03:08.072 -> Average Cal sensor Temp.: 102.71 Ambient Temp.: 97.07 F
19:03:09.633 -> Average Cal sensor Temp.: 102.86 Ambient Temp.: 97.03 F
19:03:11.185 -> Average Cal sensor Temp.: 103.04 Ambient Temp.: 97.07 F
19:03:12.744 -> Average Cal sensor Temp.: 102.61 Ambient Temp.: 97.03 F
19:03:14.329 -> Average Cal sensor Temp.: 102.74 Ambient Temp.: 97.03 F
19:03:15.885 -> Average Cal sensor Temp.: 102.56 Ambient Temp.: 97.03 F
19:03:17.436 -> Average Cal sensor Temp.: 102.49 Ambient Temp.: 97.03 F
19:03:19.025 -> Average Cal sensor Temp.: 102.71 Ambient Temp.: 97.00 F
19:03:20.575 -> Average Cal sensor Temp.: 102.47 Ambient Temp.: 97.07 F
19:03:22.131 -> Average Cal sensor Temp.: 102.67 Ambient Temp.: 97.03 F
19:03:23.715 -> Average Cal sensor Temp.: 102.54 Ambient Temp.: 97.03 F
19:03:25.271 -> Average Cal sensor Temp.: 102.61 Ambient Temp.: 97.07 F
19:03:26.822 -> Average Cal sensor Temp.: 102.51 Ambient Temp.: 97.07 F
19:03:28.406 -> Average Cal sensor Temp.: 102.04 Ambient Temp.: 97.07 F
19:03:29.961 -> Average Cal sensor Temp.: 102.18 Ambient Temp.: 97.00 F
19:03:31.513 -> Average Cal sensor Temp.: 102.17 Ambient Temp.: 97.11 F
19:03:33.068 -> Average Cal sensor Temp.: 102.60 Ambient Temp.: 97.03 F
19:03:34.659 -> Average Cal sensor Temp.: 102.63 Ambient Temp.: 97.07 F
19:03:36.220 -> Average Cal sensor Temp.: 102.60 Ambient Temp.: 97.07 F
19:03:37.777 -> Average Cal sensor Temp.: 102.87 Ambient Temp.: 97.03 F
19:03:39.325 -> Average Cal sensor Temp.: 102.79 Ambient Temp.: 97.07 F
19:03:40.911 -> Average Cal sensor Temp.: 102.60 Ambient Temp.: 97.07 F
19:03:42.468 -> Average Cal sensor Temp.: 102.28 Ambient Temp.: 97.07 F
19:03:44.023 -> Average Cal sensor Temp.: 102.50 Ambient Temp.: 97.07 F
19:03:45.612 -> Average Cal sensor Temp.: 102.90 Ambient Temp.: 97.11 F
19:03:47.170 -> Average Cal sensor Temp.: 102.85 Ambient Temp.: 97.11 F
19:03:48.721 -> Average Cal sensor Temp.: 102.54 Ambient Temp.: 97.11 F
19:03:50.276 -> Average Cal sensor Temp.: 102.25 Ambient Temp.: 97.07 F
19:03:51.863 -> Average Cal sensor Temp.: 102.45 Ambient Temp.: 97.07 F
19:03:53.420 -> Average Cal sensor Temp.: 102.74 Ambient Temp.: 97.07 F
19:03:54.969 -> Average Cal sensor Temp.: 102.39 Ambient Temp.: 97.11 F
19:03:56.526 -> Average Cal sensor Temp.: 102.50 Ambient Temp.: 97.07 F
19:03:58.117 -> Average Cal sensor Temp.: 102.21 Ambient Temp.: 97.11 F
19:03:59.675 -> Average Cal sensor Temp.: 102.77 Ambient Temp.: 97.07 F
19:04:01.229 -> Average Cal sensor Temp.: 103.12 Ambient Temp.: 97.03 F
19:04:02.778 -> Average Cal sensor Temp.: 102.99 Ambient Temp.: 97.07 F
19:04:04.365 -> Average Cal sensor Temp.: 102.76 Ambient Temp.: 97.07 F
19:04:05.925 -> Average Cal sensor Temp.: 102.72 Ambient Temp.: 97.07 F
19:04:07.476 -> Average Cal sensor Temp.: 102.76 Ambient Temp.: 97.07 F
19:04:09.060 -> Average Cal sensor Temp.: 102.41 Ambient Temp.: 97.03 F
19:04:10.619 -> Average Cal sensor Temp.: 102.33 Ambient Temp.: 97.11 F
19:04:12.174 -> Average Cal sensor Temp.: 102.56 Ambient Temp.: 97.03 F
19:04:13.726 -> Average Cal sensor Temp.: 102.38 Ambient Temp.: 97.03 F
19:04:15.315 -> Average Cal sensor Temp.: 102.35 Ambient Temp.: 97.07 F
19:04:16.878 -> Average Cal sensor Temp.: 102.32 Ambient Temp.: 97.11 F
19:04:18.433 -> Average Cal sensor Temp.: 102.40 Ambient Temp.: 97.03 F
19:04:19.992 -> Average Cal sensor Temp.: 102.69 Ambient Temp.: 97.11 F
19:04:21.547 -> Average Cal sensor Temp.: 102.50 Ambient Temp.: 97.03 F
19:04:23.102 -> Average Cal sensor Temp.: 102.32 Ambient Temp.: 97.03 F
19:04:24.690 -> Average Cal sensor Temp.: 102.43 Ambient Temp.: 97.00 F
19:04:26.250 -> Average Cal sensor Temp.: 102.84 Ambient Temp.: 97.03 F
19:04:27.806 -> Average Cal sensor Temp.: 102.74 Ambient Temp.: 97.07 F
19:04:29.361 -> Average Cal sensor Temp.: 102.72 Ambient Temp.: 97.11 F
19:04:30.955 -> Average Cal sensor Temp.: 102.53 Ambient Temp.: 97.03 F
19:04:32.510 -> Average Cal sensor Temp.: 102.61 Ambient Temp.: 97.03 F
19:04:34.063 -> Average Cal sensor Temp.: 102.66 Ambient Temp.: 97.03 F
19:04:35.622 -> Average Cal sensor Temp.: 102.78 Ambient Temp.: 97.03 F
19:04:37.176 -> Average Cal sensor Temp.: 102.76 Ambient Temp.: 97.07 F
19:04:38.766 -> Average Cal sensor Temp.: 102.46 Ambient Temp.: 97.03 F
19:04:40.318 -> Average Cal sensor Temp.: 101.95 Ambient Temp.: 97.07 F
19:04:41.874 -> Average Cal sensor Temp.: 102.34 Ambient Temp.: 97.07 F
19:04:43.462 -> Average Cal sensor Temp.: 101.59 Ambient Temp.: 97.00 F
19:04:45.017 -> Average Cal sensor Temp.: 102.11 Ambient Temp.: 97.03 F
19:04:46.575 -> Average Cal sensor Temp.: 102.57 Ambient Temp.: 96.96 F
19:04:48.134 -> Average Cal sensor Temp.: 102.79 Ambient Temp.: 97.03 F
19:04:49.690 -> Average Cal sensor Temp.: 102.87 Ambient Temp.: 97.03 F
19:04:51.278 -> Average Cal sensor Temp.: 102.60 Ambient Temp.: 97.00 F
19:04:52.829 -> Average Cal sensor Temp.: 102.64 Ambient Temp.: 97.00 F
19:04:54.384 -> Average Cal sensor Temp.: 102.50 Ambient Temp.: 97.03 F
19:04:55.979 -> Average Cal sensor Temp.: 102.65 Ambient Temp.: 97.07 F
19:04:57.535 -> Average Cal sensor Temp.: 102.52 Ambient Temp.: 97.00 F
19:04:59.089 -> Average Cal sensor Temp.: 102.68 Ambient Temp.: 97.03 F
19:05:00.672 -> Average Cal sensor Temp.: 102.80 Ambient Temp.: 97.03 F
19:05:02.227 -> Average Cal sensor Temp.: 102.69 Ambient Temp.: 97.03 F
19:05:03.777 -> Average Cal sensor Temp.: 101.90 Ambient Temp.: 97.00 F
19:05:05.366 -> Average Cal sensor Temp.: 102.32 Ambient Temp.: 97.07 F
19:05:06.920 -> Average Cal sensor Temp.: 102.56 Ambient Temp.: 97.03 F
19:05:08.475 -> Average Cal sensor Temp.: 102.87 Ambient Temp.: 97.03 F
19:05:10.032 -> Average Cal sensor Temp.: 102.97 Ambient Temp.: 97.03 F
19:05:11.619 -> Average Cal sensor Temp.: 102.72 Ambient Temp.: 97.00 F
19:05:13.175 -> Average Cal sensor Temp.: 102.78 Ambient Temp.: 97.07 F
19:05:14.729 -> Average Cal sensor Temp.: 102.90 Ambient Temp.: 97.07 F
19:05:16.314 -> Average Cal sensor Temp.: 102.76 Ambient Temp.: 97.11 F
19:05:17.870 -> Average Cal sensor Temp.: 102.85 Ambient Temp.: 97.03 F
19:05:19.426 -> Average Cal sensor Temp.: 103.20 Ambient Temp.: 97.07 F
19:05:20.983 -> Average Cal sensor Temp.: 103.18 Ambient Temp.: 97.07 F
19:05:22.542 -> Average Cal sensor Temp.: 103.24 Ambient Temp.: 97.03 F
19:05:24.102 -> Average Cal sensor Temp.: 102.83 Ambient Temp.: 97.07 F
19:05:25.693 -> Average Cal sensor Temp.: 102.54 Ambient Temp.: 97.11 F
19:05:27.252 -> Average Cal sensor Temp.: 102.54 Ambient Temp.: 97.07 F
19:05:28.809 -> Average Cal sensor Temp.: 102.64 Ambient Temp.: 97.07 F
19:05:30.366 -> Average Cal sensor Temp.: 102.45 Ambient Temp.: 97.07 F
19:05:31.925 -> Average Cal sensor Temp.: 102.67 Ambient Temp.: 97.07 F
19:05:33.513 -> Average Cal sensor Temp.: 102.44 Ambient Temp.: 97.03 F
19:05:35.072 -> Average Cal sensor Temp.: 103.01 Ambient Temp.: 97.11 F
19:05:36.632 -> Average Cal sensor Temp.: 103.32 Ambient Temp.: 97.11 F
19:05:38.192 -> Average Cal sensor Temp.: 102.89 Ambient Temp.: 97.11 F
19:05:39.749 -> Average Cal sensor Temp.: 102.91 Ambient Temp.: 97.14 F
19:05:41.304 -> Average Cal sensor Temp.: 102.97 Ambient Temp.: 97.07 F
19:05:42.895 -> Average Cal sensor Temp.: 102.28 Ambient Temp.: 97.18 F
19:05:44.452 -> Average Cal sensor Temp.: 102.51 Ambient Temp.: 97.11 F
19:05:46.008 -> Average Cal sensor Temp.: 102.51 Ambient Temp.: 97.03 F
19:05:47.561 -> Average Cal sensor Temp.: 102.68 Ambient Temp.: 97.14 F
19:05:49.150 -> Average Cal sensor Temp.: 102.82 Ambient Temp.: 97.07 F
19:05:50.681 -> Average Cal sensor Temp.: 102.69 Ambient Temp.: 97.07 F
19:05:52.270 -> Average Cal sensor Temp.: 102.12 Ambient Temp.: 97.07 F
19:05:53.826 -> Average Cal sensor Temp.: 102.44 Ambient Temp.: 97.07 F
19:05:55.383 -> Average Cal sensor Temp.: 102.87 Ambient Temp.: 97.07 F
19:05:56.938 -> Average Cal sensor Temp.: 103.09 Ambient Temp.: 97.14 F
19:05:58.530 -> Average Cal sensor Temp.: 103.05 Ambient Temp.: 97.11 F
19:06:00.084 -> Average Cal sensor Temp.: 102.99 Ambient Temp.: 97.11 F
19:06:01.638 -> Average Cal sensor Temp.: 102.38 Ambient Temp.: 97.03 F
19:06:03.195 -> Average Cal sensor Temp.: 102.78 Ambient Temp.: 97.11 F
19:06:04.786 -> Average Cal sensor Temp.: 103.11 Ambient Temp.: 97.11 F
19:06:06.346 -> Average Cal sensor Temp.: 102.89 Ambient Temp.: 97.11 F
19:06:07.901 -> Average Cal sensor Temp.: 103.21 Ambient Temp.: 97.11 F
19:06:09.457 -> Average Cal sensor Temp.: 102.67 Ambient Temp.: 97.14 F
19:06:11.011 -> Average Cal sensor Temp.: 102.57 Ambient Temp.: 97.11 F
19:06:12.601 -> Average Cal sensor Temp.: 102.87 Ambient Temp.: 97.18 F
19:06:14.152 -> Average Cal sensor Temp.: 102.72 Ambient Temp.: 97.11 F
19:06:15.711 -> Average Cal sensor Temp.: 102.84 Ambient Temp.: 97.14 F
19:06:17.257 -> Average Cal sensor Temp.: 102.81 Ambient Temp.: 97.03 F
19:06:18.847 -> Average Cal sensor Temp.: 102.94 Ambient Temp.: 97.07 F
19:06:20.404 -> Average Cal sensor Temp.: 102.93 Ambient Temp.: 97.14 F
19:06:21.960 -> Average Cal sensor Temp.: 103.07 Ambient Temp.: 97.14 F
19:06:23.515 -> Average Cal sensor Temp.: 102.76 Ambient Temp.: 97.14 F
19:06:25.105 -> Average Cal sensor Temp.: 103.04 Ambient Temp.: 97.14 F
19:06:26.662 -> Average Cal sensor Temp.: 103.10 Ambient Temp.: 97.11 F
19:06:28.216 -> Average Cal sensor Temp.: 102.91 Ambient Temp.: 97.18 F
19:06:29.771 -> Average Cal sensor Temp.: 102.83 Ambient Temp.: 97.11 F
19:06:31.360 -> Average Cal sensor Temp.: 102.84 Ambient Temp.: 97.14 F
19:06:32.921 -> Average Cal sensor Temp.: 103.01 Ambient Temp.: 97.14 F
19:06:34.473 -> Average Cal sensor Temp.: 103.19 Ambient Temp.: 97.11 F
19:06:36.028 -> Average Cal sensor Temp.: 103.05 Ambient Temp.: 97.14 F
19:06:37.616 -> Average Cal sensor Temp.: 103.21 Ambient Temp.: 97.11 F
19:06:39.172 -> Average Cal sensor Temp.: 103.01 Ambient Temp.: 97.18 F
19:06:40.729 -> Average Cal sensor Temp.: 102.85 Ambient Temp.: 97.14 F
19:06:42.279 -> Average Cal sensor Temp.: 102.89 Ambient Temp.: 97.14 F
19:06:43.835 -> Average Cal sensor Temp.: 102.93 Ambient Temp.: 97.14 F
19:06:45.418 -> Average Cal sensor Temp.: 102.67 Ambient Temp.: 97.14 F
19:06:46.971 -> Average Cal sensor Temp.: 102.90 Ambient Temp.: 97.18 F
19:06:48.559 -> Average Cal sensor Temp.: 102.74 Ambient Temp.: 97.14 F
19:06:50.111 -> Average Cal sensor Temp.: 102.98 Ambient Temp.: 97.14 F
19:06:51.666 -> Average Cal sensor Temp.: 102.54 Ambient Temp.: 97.14 F
19:06:53.223 -> Average Cal sensor Temp.: 102.90 Ambient Temp.: 97.14 F
19:06:54.814 -> Average Cal sensor Temp.: 103.21 Ambient Temp.: 97.14 F
19:06:56.372 -> Average Cal sensor Temp.: 102.90 Ambient Temp.: 97.18 F
19:06:57.929 -> Average Cal sensor Temp.: 102.98 Ambient Temp.: 97.11 F
19:06:59.484 -> Average Cal sensor Temp.: 102.34 Ambient Temp.: 97.11 F
19:07:01.043 -> Average Cal sensor Temp.: 102.10 Ambient Temp.: 97.14 F
19:07:02.632 -> Average Cal sensor Temp.: 102.66 Ambient Temp.: 97.11 F
19:07:04.187 -> Average Cal sensor Temp.: 102.45 Ambient Temp.: 97.11 F
19:07:05.738 -> Average Cal sensor Temp.: 102.88 Ambient Temp.: 97.14 F
19:07:07.292 -> Average Cal sensor Temp.: 103.19 Ambient Temp.: 97.14 F
19:07:08.881 -> Average Cal sensor Temp.: 102.75 Ambient Temp.: 97.07 F
19:07:10.437 -> Average Cal sensor Temp.: 103.30 Ambient Temp.: 97.18 F
19:07:11.989 -> Average Cal sensor Temp.: 103.31 Ambient Temp.: 97.18 F
19:07:13.550 -> Average Cal sensor Temp.: 102.59 Ambient Temp.: 97.14 F
19:07:15.140 -> Average Cal sensor Temp.: 102.90 Ambient Temp.: 97.14 F
19:07:16.695 -> Average Cal sensor Temp.: 102.68 Ambient Temp.: 97.14 F
19:07:18.252 -> Average Cal sensor Temp.: 102.82 Ambient Temp.: 97.18 F
19:07:19.810 -> Average Cal sensor Temp.: 103.20 Ambient Temp.: 97.11 F
19:07:21.395 -> Average Cal sensor Temp.: 103.16 Ambient Temp.: 97.14 F
19:07:22.948 -> Average Cal sensor Temp.: 103.19 Ambient Temp.: 97.14 F
19:07:24.501 -> Average Cal sensor Temp.: 103.40 Ambient Temp.: 97.11 F
19:07:26.054 -> Average Cal sensor Temp.: 102.92 Ambient Temp.: 97.11 F
19:07:27.642 -> Average Cal sensor Temp.: 102.89 Ambient Temp.: 97.14 F
19:07:29.195 -> Average Cal sensor Temp.: 103.00 Ambient Temp.: 97.11 F
19:07:30.747 -> Average Cal sensor Temp.: 102.96 Ambient Temp.: 97.11 F
19:07:32.336 -> Average Cal sensor Temp.: 102.25 Ambient Temp.: 97.11 F
19:07:33.890 -> Average Cal sensor Temp.: 103.19 Ambient Temp.: 97.11 F
19:07:35.443 -> Average Cal sensor Temp.: 103.37 Ambient Temp.: 97.07 F
19:07:37.001 -> Average Cal sensor Temp.: 103.11 Ambient Temp.: 97.07 F
19:07:38.557 -> Average Cal sensor Temp.: 103.35 Ambient Temp.: 97.07 F
19:07:40.148 -> Average Cal sensor Temp.: 103.19 Ambient Temp.: 97.11 F
19:07:41.695 -> Average Cal sensor Temp.: 103.16 Ambient Temp.: 97.11 F
19:07:43.251 -> Average Cal sensor Temp.: 103.11 Ambient Temp.: 97.11 F
19:07:44.810 -> Average Cal sensor Temp.: 103.01 Ambient Temp.: 97.11 F
19:07:46.399 -> Average Cal sensor Temp.: 103.07 Ambient Temp.: 97.11 F
19:07:47.952 -> Average Cal sensor Temp.: 102.75 Ambient Temp.: 97.07 F
19:07:49.510 -> Average Cal sensor Temp.: 102.93 Ambient Temp.: 97.11 F
19:07:51.064 -> Average Cal sensor Temp.: 103.07 Ambient Temp.: 97.14 F
19:07:52.651 -> Average Cal sensor Temp.: 103.17 Ambient Temp.: 97.07 F
19:07:54.209 -> Average Cal sensor Temp.: 103.01 Ambient Temp.: 97.14 F
19:07:55.768 -> Average Cal sensor Temp.: 103.09 Ambient Temp.: 97.14 F
19:07:57.324 -> Average Cal sensor Temp.: 103.30 Ambient Temp.: 97.14 F
19:07:58.875 -> Average Cal sensor Temp.: 103.21 Ambient Temp.: 97.11 F
19:08:00.461 -> Average Cal sensor Temp.: 103.33 Ambient Temp.: 97.11 F
19:08:02.017 -> Average Cal sensor Temp.: 103.56 Ambient Temp.: 97.18 F
19:08:03.572 -> Average Cal sensor Temp.: 102.97 Ambient Temp.: 97.14 F
19:08:05.131 -> Average Cal sensor Temp.: 103.37 Ambient Temp.: 97.14 F
19:08:06.688 -> Average Cal sensor Temp.: 103.33 Ambient Temp.: 97.14 F
19:08:08.248 -> Average Cal sensor Temp.: 103.06 Ambient Temp.: 97.14 F
19:08:09.839 -> Average Cal sensor Temp.: 103.16 Ambient Temp.: 97.14 F
19:08:11.395 -> Average Cal sensor Temp.: 103.23 Ambient Temp.: 97.14 F
19:08:12.956 -> Average Cal sensor Temp.: 103.36 Ambient Temp.: 97.18 F
19:08:14.508 -> Average Cal sensor Temp.: 103.30 Ambient Temp.: 97.14 F
19:08:16.098 -> Average Cal sensor Temp.: 102.73 Ambient Temp.: 97.14 F
19:08:17.653 -> Average Cal sensor Temp.: 103.37 Ambient Temp.: 97.18 F
19:08:19.212 -> Average Cal sensor Temp.: 103.38 Ambient Temp.: 97.14 F
19:08:20.769 -> Average Cal sensor Temp.: 102.90 Ambient Temp.: 97.11 F
19:08:22.323 -> Average Cal sensor Temp.: 102.87 Ambient Temp.: 97.11 F
19:08:23.908 -> Average Cal sensor Temp.: 102.89 Ambient Temp.: 97.18 F
19:08:25.465 -> Average Cal sensor Temp.: 103.20 Ambient Temp.: 97.11 F
19:08:27.022 -> Average Cal sensor Temp.: 103.37 Ambient Temp.: 97.14 F
19:08:28.580 -> Average Cal sensor Temp.: 102.56 Ambient Temp.: 97.14 F
19:08:30.136 -> Average Cal sensor Temp.: 103.29 Ambient Temp.: 97.11 F
19:08:31.722 -> Average Cal sensor Temp.: 102.98 Ambient Temp.: 97.14 F
19:08:33.280 -> Average Cal sensor Temp.: 102.90 Ambient Temp.: 97.14 F
19:08:34.830 -> Average Cal sensor Temp.: 103.08 Ambient Temp.: 97.18 F
19:08:36.393 -> Average Cal sensor Temp.: 103.48 Ambient Temp.: 97.18 F
19:08:37.982 -> Average Cal sensor Temp.: 103.51 Ambient Temp.: 97.11 F
19:08:39.540 -> Average Cal sensor Temp.: 103.60 Ambient Temp.: 97.18 F
19:08:41.093 -> Average Cal sensor Temp.: 103.79 Ambient Temp.: 97.18 F
19:08:42.649 -> Average Cal sensor Temp.: 103.31 Ambient Temp.: 97.11 F
19:08:44.240 -> Average Cal sensor Temp.: 103.15 Ambient Temp.: 97.14 F
19:08:45.794 -> Average Cal sensor Temp.: 103.46 Ambient Temp.: 97.18 F
19:08:47.344 -> Average Cal sensor Temp.: 102.96 Ambient Temp.: 97.18 F
19:08:48.903 -> Average Cal sensor Temp.: 103.24 Ambient Temp.: 97.14 F
19:08:50.490 -> Average Cal sensor Temp.: 103.38 Ambient Temp.: 97.18 F
19:08:52.041 -> Average Cal sensor Temp.: 103.34 Ambient Temp.: 97.14 F
19:08:53.598 -> Average Cal sensor Temp.: 102.83 Ambient Temp.: 97.14 F
19:08:55.158 -> Average Cal sensor Temp.: 103.22 Ambient Temp.: 97.18 F
19:08:56.747 -> Average Cal sensor Temp.: 103.34 Ambient Temp.: 97.18 F
19:08:58.294 -> Average Cal sensor Temp.: 103.28 Ambient Temp.: 97.14 F
19:08:59.849 -> Average Cal sensor Temp.: 102.66 Ambient Temp.: 97.14 F
19:09:01.436 -> Average Cal sensor Temp.: 103.32 Ambient Temp.: 97.18 F
19:09:02.993 -> Average Cal sensor Temp.: 103.17 Ambient Temp.: 97.18 F
19:09:04.544 -> Average Cal sensor Temp.: 103.51 Ambient Temp.: 97.14 F
19:09:06.104 -> Average Cal sensor Temp.: 103.22 Ambient Temp.: 97.14 F
19:09:07.687 -> Average Cal sensor Temp.: 103.17 Ambient Temp.: 97.14 F
19:09:09.241 -> Average Cal sensor Temp.: 103.08 Ambient Temp.: 97.14 F
19:09:10.799 -> Average Cal sensor Temp.: 102.63 Ambient Temp.: 97.14 F
19:09:12.353 -> Average Cal sensor Temp.: 102.78 Ambient Temp.: 97.14 F
19:09:13.939 -> Average Cal sensor Temp.: 102.90 Ambient Temp.: 97.21 F
19:09:15.499 -> Average Cal sensor Temp.: 103.10 Ambient Temp.: 97.14 F
19:09:17.053 -> Average Cal sensor Temp.: 102.93 Ambient Temp.: 97.14 F
19:09:18.608 -> Average Cal sensor Temp.: 103.09 Ambient Temp.: 97.14 F
19:09:20.200 -> Average Cal sensor Temp.: 103.32 Ambient Temp.: 97.14 F
19:09:21.757 -> Average Cal sensor Temp.: 103.26 Ambient Temp.: 97.18 F
19:09:23.300 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 97.14 F
19:09:24.891 -> Average Cal sensor Temp.: 103.42 Ambient Temp.: 97.14 F
19:09:26.447 -> Average Cal sensor Temp.: 103.27 Ambient Temp.: 97.11 F
19:09:27.991 -> Average Cal sensor Temp.: 103.14 Ambient Temp.: 97.11 F
19:09:29.551 -> Average Cal sensor Temp.: 103.29 Ambient Temp.: 97.14 F
19:09:31.134 -> Average Cal sensor Temp.: 103.33 Ambient Temp.: 97.11 F
19:09:32.694 -> Average Cal sensor Temp.: 102.93 Ambient Temp.: 97.18 F
19:09:34.246 -> Average Cal sensor Temp.: 103.11 Ambient Temp.: 97.18 F
19:09:35.805 -> Average Cal sensor Temp.: 103.43 Ambient Temp.: 97.18 F
19:09:37.363 -> Average Cal sensor Temp.: 103.58 Ambient Temp.: 97.18 F
19:09:38.955 -> Average Cal sensor Temp.: 103.47 Ambient Temp.: 97.18 F
19:09:40.514 -> Average Cal sensor Temp.: 103.62 Ambient Temp.: 97.18 F
19:09:42.067 -> Average Cal sensor Temp.: 103.49 Ambient Temp.: 97.11 F
19:09:43.623 -> Average Cal sensor Temp.: 103.45 Ambient Temp.: 97.18 F
19:09:45.208 -> Average Cal sensor Temp.: 103.41 Ambient Temp.: 97.14 F
19:09:46.771 -> Average Cal sensor Temp.: 103.39 Ambient Temp.: 97.14 F
19:09:48.328 -> Average Cal sensor Temp.: 103.09 Ambient Temp.: 97.14 F
19:09:49.895 -> Average Cal sensor Temp.: 103.32 Ambient Temp.: 97.14 F
19:09:51.452 -> Average Cal sensor Temp.: 103.55 Ambient Temp.: 97.14 F
19:09:53.004 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 97.18 F
19:09:54.593 -> Average Cal sensor Temp.: 102.99 Ambient Temp.: 97.14 F
19:09:56.152 -> Average Cal sensor Temp.: 103.46 Ambient Temp.: 97.11 F
19:09:57.708 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 97.18 F
19:09:59.264 -> Average Cal sensor Temp.: 103.81 Ambient Temp.: 97.18 F
19:10:00.851 -> Average Cal sensor Temp.: 103.32 Ambient Temp.: 97.21 F
19:10:02.409 -> Average Cal sensor Temp.: 103.46 Ambient Temp.: 97.14 F
19:10:03.959 -> Average Cal sensor Temp.: 103.55 Ambient Temp.: 97.14 F
19:10:05.513 -> Average Cal sensor Temp.: 103.70 Ambient Temp.: 97.21 F
19:10:07.075 -> Average Cal sensor Temp.: 103.47 Ambient Temp.: 97.18 F
19:10:08.663 -> Average Cal sensor Temp.: 103.32 Ambient Temp.: 97.14 F
19:10:10.221 -> Average Cal sensor Temp.: 102.90 Ambient Temp.: 97.21 F
19:10:11.775 -> Average Cal sensor Temp.: 103.10 Ambient Temp.: 97.14 F
19:10:13.333 -> Average Cal sensor Temp.: 103.37 Ambient Temp.: 97.14 F
19:10:14.889 -> Average Cal sensor Temp.: 103.23 Ambient Temp.: 97.14 F
19:10:16.477 -> Average Cal sensor Temp.: 102.95 Ambient Temp.: 97.14 F
19:10:18.033 -> Average Cal sensor Temp.: 103.23 Ambient Temp.: 97.14 F
19:10:19.585 -> Average Cal sensor Temp.: 103.80 Ambient Temp.: 97.21 F
19:10:21.176 -> Average Cal sensor Temp.: 103.91 Ambient Temp.: 97.14 F
19:10:22.729 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 97.14 F
19:10:24.285 -> Average Cal sensor Temp.: 103.75 Ambient Temp.: 97.21 F
19:10:25.847 -> Average Cal sensor Temp.: 103.76 Ambient Temp.: 97.18 F
19:10:27.398 -> Average Cal sensor Temp.: 104.07 Ambient Temp.: 97.18 F
19:10:28.989 -> Average Cal sensor Temp.: 103.62 Ambient Temp.: 97.14 F
19:10:30.542 -> Average Cal sensor Temp.: 103.77 Ambient Temp.: 97.14 F
19:10:32.098 -> Average Cal sensor Temp.: 103.87 Ambient Temp.: 97.21 F
19:10:33.655 -> Average Cal sensor Temp.: 103.82 Ambient Temp.: 97.18 F
19:10:35.213 -> Average Cal sensor Temp.: 104.11 Ambient Temp.: 97.18 F
19:10:36.806 -> Average Cal sensor Temp.: 103.79 Ambient Temp.: 97.18 F
19:10:38.364 -> Average Cal sensor Temp.: 103.84 Ambient Temp.: 97.29 F
19:10:39.919 -> Average Cal sensor Temp.: 103.93 Ambient Temp.: 97.21 F
19:10:41.468 -> Average Cal sensor Temp.: 103.21 Ambient Temp.: 97.18 F
19:10:43.030 -> Average Cal sensor Temp.: 103.44 Ambient Temp.: 97.18 F
19:10:44.621 -> Average Cal sensor Temp.: 103.51 Ambient Temp.: 97.21 F
19:10:46.174 -> Average Cal sensor Temp.: 103.46 Ambient Temp.: 97.18 F
19:10:47.729 -> Average Cal sensor Temp.: 102.92 Ambient Temp.: 97.18 F
19:10:49.283 -> Average Cal sensor Temp.: 103.01 Ambient Temp.: 97.29 F
19:10:50.875 -> Average Cal sensor Temp.: 103.10 Ambient Temp.: 97.21 F
19:10:52.427 -> Average Cal sensor Temp.: 103.41 Ambient Temp.: 97.21 F
19:10:53.983 -> Average Cal sensor Temp.: 103.09 Ambient Temp.: 97.21 F
19:10:55.542 -> Average Cal sensor Temp.: 102.94 Ambient Temp.: 97.21 F
19:10:57.102 -> Average Cal sensor Temp.: 103.26 Ambient Temp.: 97.29 F
19:10:58.691 -> Average Cal sensor Temp.: 103.41 Ambient Temp.: 97.29 F
19:11:00.243 -> Average Cal sensor Temp.: 103.37 Ambient Temp.: 97.21 F
19:11:01.791 -> Average Cal sensor Temp.: 103.51 Ambient Temp.: 97.18 F
19:11:03.353 -> Average Cal sensor Temp.: 103.88 Ambient Temp.: 97.21 F
19:11:04.945 -> Average Cal sensor Temp.: 103.89 Ambient Temp.: 97.21 F
19:11:06.502 -> Average Cal sensor Temp.: 103.82 Ambient Temp.: 97.21 F
19:11:08.054 -> Average Cal sensor Temp.: 103.82 Ambient Temp.: 97.21 F
19:11:09.613 -> Average Cal sensor Temp.: 103.44 Ambient Temp.: 97.29 F
19:11:11.203 -> Average Cal sensor Temp.: 103.87 Ambient Temp.: 97.29 F
19:11:12.758 -> Average Cal sensor Temp.: 103.76 Ambient Temp.: 97.29 F
19:11:14.308 -> Average Cal sensor Temp.: 103.80 Ambient Temp.: 97.21 F
19:11:15.868 -> Average Cal sensor Temp.: 103.70 Ambient Temp.: 97.29 F
19:11:17.460 -> Average Cal sensor Temp.: 103.47 Ambient Temp.: 97.29 F
19:11:19.017 -> Average Cal sensor Temp.: 103.25 Ambient Temp.: 97.29 F
19:11:20.577 -> Average Cal sensor Temp.: 103.14 Ambient Temp.: 97.29 F
19:11:22.128 -> Average Cal sensor Temp.: 103.45 Ambient Temp.: 97.21 F
19:11:23.713 -> Average Cal sensor Temp.: 103.54 Ambient Temp.: 97.29 F
19:11:25.266 -> Average Cal sensor Temp.: 103.70 Ambient Temp.: 97.32 F
19:11:26.820 -> Average Cal sensor Temp.: 103.55 Ambient Temp.: 97.29 F
19:11:28.374 -> Average Cal sensor Temp.: 103.55 Ambient Temp.: 97.29 F
19:11:29.969 -> Average Cal sensor Temp.: 103.82 Ambient Temp.: 97.39 F
19:11:31.521 -> Average Cal sensor Temp.: 103.74 Ambient Temp.: 97.29 F
19:11:33.072 -> Average Cal sensor Temp.: 103.91 Ambient Temp.: 97.32 F
19:11:34.659 -> Average Cal sensor Temp.: 104.25 Ambient Temp.: 97.32 F
19:11:36.221 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 97.29 F
19:11:37.779 -> Average Cal sensor Temp.: 103.02 Ambient Temp.: 97.29 F
19:11:39.337 -> Average Cal sensor Temp.: 103.08 Ambient Temp.: 97.21 F
19:11:40.895 -> Average Cal sensor Temp.: 103.41 Ambient Temp.: 97.21 F
19:11:42.481 -> Average Cal sensor Temp.: 103.68 Ambient Temp.: 97.29 F
19:11:44.039 -> Average Cal sensor Temp.: 104.26 Ambient Temp.: 97.32 F
19:11:45.596 -> Average Cal sensor Temp.: 103.98 Ambient Temp.: 97.21 F
19:11:47.144 -> Average Cal sensor Temp.: 103.79 Ambient Temp.: 97.29 F
19:11:48.728 -> Average Cal sensor Temp.: 103.78 Ambient Temp.: 97.32 F
19:11:50.286 -> Average Cal sensor Temp.: 103.90 Ambient Temp.: 97.36 F
19:11:51.839 -> Average Cal sensor Temp.: 104.05 Ambient Temp.: 97.32 F
19:11:53.427 -> Average Cal sensor Temp.: 103.65 Ambient Temp.: 97.29 F
19:11:54.984 -> Average Cal sensor Temp.: 103.40 Ambient Temp.: 97.21 F
19:11:56.544 -> Average Cal sensor Temp.: 103.46 Ambient Temp.: 97.21 F
19:11:58.103 -> Average Cal sensor Temp.: 103.25 Ambient Temp.: 97.29 F
19:11:59.685 -> Average Cal sensor Temp.: 102.92 Ambient Temp.: 97.29 F
19:12:01.243 -> Average Cal sensor Temp.: 103.37 Ambient Temp.: 97.29 F
19:12:02.792 -> Average Cal sensor Temp.: 103.06 Ambient Temp.: 97.18 F
19:12:04.348 -> Average Cal sensor Temp.: 102.95 Ambient Temp.: 97.21 F
19:12:05.911 -> Average Cal sensor Temp.: 103.47 Ambient Temp.: 97.18 F
19:12:07.502 -> Average Cal sensor Temp.: 103.41 Ambient Temp.: 97.18 F
19:12:09.053 -> Average Cal sensor Temp.: 103.35 Ambient Temp.: 97.18 F
19:12:10.608 -> Average Cal sensor Temp.: 103.74 Ambient Temp.: 97.21 F
19:12:12.162 -> Average Cal sensor Temp.: 103.70 Ambient Temp.: 97.11 F
19:12:13.752 -> Average Cal sensor Temp.: 103.75 Ambient Temp.: 97.21 F
19:12:15.311 -> Average Cal sensor Temp.: 103.88 Ambient Temp.: 97.18 F
19:12:16.869 -> Average Cal sensor Temp.: 103.71 Ambient Temp.: 97.21 F
19:12:18.425 -> Average Cal sensor Temp.: 103.49 Ambient Temp.: 97.18 F
19:12:19.983 -> Average Cal sensor Temp.: 103.43 Ambient Temp.: 97.29 F
19:12:21.569 -> Average Cal sensor Temp.: 103.16 Ambient Temp.: 97.21 F
19:12:23.127 -> Average Cal sensor Temp.: 103.26 Ambient Temp.: 97.21 F
19:12:24.684 -> Average Cal sensor Temp.: 103.45 Ambient Temp.: 97.21 F
19:12:26.241 -> Average Cal sensor Temp.: 103.53 Ambient Temp.: 97.21 F
19:12:27.798 -> Average Cal sensor Temp.: 103.68 Ambient Temp.: 97.18 F
19:12:29.382 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 97.21 F
19:12:30.935 -> Average Cal sensor Temp.: 103.98 Ambient Temp.: 97.21 F
19:12:32.499 -> Average Cal sensor Temp.: 103.51 Ambient Temp.: 97.21 F
19:12:34.082 -> Average Cal sensor Temp.: 103.45 Ambient Temp.: 97.29 F
19:12:35.638 -> Average Cal sensor Temp.: 103.28 Ambient Temp.: 97.21 F
19:12:37.189 -> Average Cal sensor Temp.: 103.43 Ambient Temp.: 97.21 F
19:12:38.746 -> Average Cal sensor Temp.: 103.88 Ambient Temp.: 97.21 F
19:12:40.329 -> Average Cal sensor Temp.: 103.68 Ambient Temp.: 97.18 F
19:12:41.888 -> Average Cal sensor Temp.: 102.83 Ambient Temp.: 97.21 F
19:12:43.442 -> Average Cal sensor Temp.: 103.42 Ambient Temp.: 97.18 F
19:12:45.029 -> Average Cal sensor Temp.: 103.41 Ambient Temp.: 97.14 F
19:12:46.584 -> Average Cal sensor Temp.: 103.72 Ambient Temp.: 97.18 F
19:12:48.141 -> Average Cal sensor Temp.: 103.84 Ambient Temp.: 97.18 F
19:12:49.701 -> Average Cal sensor Temp.: 103.47 Ambient Temp.: 97.21 F
19:12:51.251 -> Average Cal sensor Temp.: 103.56 Ambient Temp.: 97.29 F
19:12:52.844 -> Average Cal sensor Temp.: 103.57 Ambient Temp.: 97.21 F
19:12:54.397 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 97.21 F
19:12:55.954 -> Average Cal sensor Temp.: 103.17 Ambient Temp.: 97.21 F
19:12:57.501 -> Average Cal sensor Temp.: 103.49 Ambient Temp.: 97.21 F
19:12:59.089 -> Average Cal sensor Temp.: 103.82 Ambient Temp.: 97.18 F
19:13:00.643 -> Average Cal sensor Temp.: 103.28 Ambient Temp.: 97.18 F
19:13:02.223 -> Average Cal sensor Temp.: 103.00 Ambient Temp.: 97.14 F
19:13:03.777 -> Average Cal sensor Temp.: 102.85 Ambient Temp.: 97.18 F
19:13:05.334 -> Average Cal sensor Temp.: 103.38 Ambient Temp.: 97.14 F
19:13:06.891 -> Average Cal sensor Temp.: 102.69 Ambient Temp.: 97.14 F
19:13:08.478 -> Average Cal sensor Temp.: 102.81 Ambient Temp.: 97.18 F
19:13:10.025 -> Average Cal sensor Temp.: 103.34 Ambient Temp.: 97.14 F
19:13:11.585 -> Average Cal sensor Temp.: 103.58 Ambient Temp.: 97.18 F
19:13:13.139 -> Average Cal sensor Temp.: 103.71 Ambient Temp.: 97.18 F
19:13:14.723 -> Average Cal sensor Temp.: 103.49 Ambient Temp.: 97.14 F
19:13:16.276 -> Average Cal sensor Temp.: 103.94 Ambient Temp.: 97.14 F
19:13:17.862 -> Average Cal sensor Temp.: 104.31 Ambient Temp.: 97.18 F
19:13:19.414 -> Average Cal sensor Temp.: 103.77 Ambient Temp.: 97.18 F
19:13:20.970 -> Average Cal sensor Temp.: 103.28 Ambient Temp.: 97.18 F
19:13:22.520 -> Average Cal sensor Temp.: 102.94 Ambient Temp.: 97.14 F
19:13:24.110 -> Average Cal sensor Temp.: 103.42 Ambient Temp.: 97.14 F
19:13:25.666 -> Average Cal sensor Temp.: 102.75 Ambient Temp.: 97.14 F
19:13:27.225 -> Average Cal sensor Temp.: 103.54 Ambient Temp.: 97.21 F
19:13:28.778 -> Average Cal sensor Temp.: 103.23 Ambient Temp.: 97.14 F
19:13:30.369 -> Average Cal sensor Temp.: 103.76 Ambient Temp.: 97.11 F
19:13:31.924 -> Average Cal sensor Temp.: 103.66 Ambient Temp.: 97.18 F
19:13:33.479 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 97.14 F
19:13:35.030 -> Average Cal sensor Temp.: 103.45 Ambient Temp.: 97.18 F
19:13:36.615 -> Average Cal sensor Temp.: 103.41 Ambient Temp.: 97.14 F
19:13:38.164 -> Average Cal sensor Temp.: 103.03 Ambient Temp.: 97.11 F
19:13:39.721 -> Average Cal sensor Temp.: 103.36 Ambient Temp.: 97.11 F
19:13:41.309 -> Average Cal sensor Temp.: 103.69 Ambient Temp.: 97.14 F
19:13:42.854 -> Average Cal sensor Temp.: 103.80 Ambient Temp.: 97.14 F
19:13:44.441 -> Average Cal sensor Temp.: 103.54 Ambient Temp.: 97.14 F
19:13:45.993 -> Average Cal sensor Temp.: 103.89 Ambient Temp.: 97.14 F
19:13:47.541 -> Average Cal sensor Temp.: 103.71 Ambient Temp.: 97.14 F
19:13:49.128 -> Average Cal sensor Temp.: 103.96 Ambient Temp.: 97.07 F
19:13:50.681 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 97.11 F
19:13:52.259 -> Average Cal sensor Temp.: 104.03 Ambient Temp.: 97.18 F
19:13:53.802 -> Average Cal sensor Temp.: 103.67 Ambient Temp.: 97.11 F
19:13:55.356 -> Average Cal sensor Temp.: 103.55 Ambient Temp.: 97.14 F
19:13:56.937 -> Average Cal sensor Temp.: 103.31 Ambient Temp.: 97.14 F
19:13:58.485 -> Average Cal sensor Temp.: 103.32 Ambient Temp.: 97.11 F
19:14:00.070 -> Average Cal sensor Temp.: 103.51 Ambient Temp.: 97.07 F
19:14:01.619 -> Average Cal sensor Temp.: 103.66 Ambient Temp.: 97.14 F
19:14:03.174 -> Average Cal sensor Temp.: 103.66 Ambient Temp.: 97.18 F
19:14:04.755 -> Average Cal sensor Temp.: 103.45 Ambient Temp.: 97.14 F
19:14:06.309 -> Average Cal sensor Temp.: 103.99 Ambient Temp.: 97.21 F
19:14:07.890 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 97.14 F
19:14:09.440 -> Average Cal sensor Temp.: 104.22 Ambient Temp.: 97.14 F
19:14:10.990 -> Average Cal sensor Temp.: 104.21 Ambient Temp.: 97.14 F
19:14:12.568 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 97.18 F
19:14:14.127 -> Average Cal sensor Temp.: 104.08 Ambient Temp.: 97.18 F
19:14:15.683 -> Average Cal sensor Temp.: 104.02 Ambient Temp.: 97.21 F
19:14:17.267 -> Average Cal sensor Temp.: 103.49 Ambient Temp.: 97.18 F
19:14:18.821 -> Average Cal sensor Temp.: 103.41 Ambient Temp.: 97.18 F
19:14:20.399 -> Average Cal sensor Temp.: 103.77 Ambient Temp.: 97.21 F
19:14:21.949 -> Average Cal sensor Temp.: 104.05 Ambient Temp.: 97.18 F
19:14:23.499 -> Average Cal sensor Temp.: 103.96 Ambient Temp.: 97.21 F
19:14:25.081 -> Average Cal sensor Temp.: 103.94 Ambient Temp.: 97.21 F
19:14:26.637 -> Average Cal sensor Temp.: 103.92 Ambient Temp.: 97.21 F
19:14:28.221 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 97.21 F
19:14:29.770 -> Average Cal sensor Temp.: 103.76 Ambient Temp.: 97.21 F
19:14:31.323 -> Average Cal sensor Temp.: 103.87 Ambient Temp.: 97.29 F
19:14:32.881 -> Average Cal sensor Temp.: 103.46 Ambient Temp.: 97.29 F
19:14:34.466 -> Average Cal sensor Temp.: 103.59 Ambient Temp.: 97.29 F
19:14:36.021 -> Average Cal sensor Temp.: 103.68 Ambient Temp.: 97.21 F
19:14:37.605 -> Average Cal sensor Temp.: 103.78 Ambient Temp.: 97.21 F
19:14:39.159 -> Average Cal sensor Temp.: 103.76 Ambient Temp.: 97.29 F
19:14:40.714 -> Average Cal sensor Temp.: 103.77 Ambient Temp.: 97.32 F
19:14:42.297 -> Average Cal sensor Temp.: 103.70 Ambient Temp.: 97.32 F
19:14:43.852 -> Average Cal sensor Temp.: 103.66 Ambient Temp.: 97.29 F
19:14:45.404 -> Average Cal sensor Temp.: 103.82 Ambient Temp.: 97.29 F
19:14:46.963 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 97.36 F
19:14:48.523 -> Average Cal sensor Temp.: 103.87 Ambient Temp.: 97.32 F
19:14:50.115 -> Average Cal sensor Temp.: 104.19 Ambient Temp.: 97.32 F
19:14:51.670 -> Average Cal sensor Temp.: 103.83 Ambient Temp.: 97.32 F
19:14:53.218 -> Average Cal sensor Temp.: 103.49 Ambient Temp.: 97.36 F
19:14:54.808 -> Average Cal sensor Temp.: 103.41 Ambient Temp.: 97.36 F
19:14:56.366 -> Average Cal sensor Temp.: 103.50 Ambient Temp.: 97.39 F
19:14:57.913 -> Average Cal sensor Temp.: 103.56 Ambient Temp.: 97.36 F
19:14:59.500 -> Average Cal sensor Temp.: 103.34 Ambient Temp.: 97.36 F
19:15:01.058 -> Average Cal sensor Temp.: 103.54 Ambient Temp.: 97.36 F
19:15:02.613 -> Average Cal sensor Temp.: 103.61 Ambient Temp.: 97.36 F
19:15:04.165 -> Average Cal sensor Temp.: 103.63 Ambient Temp.: 97.36 F
19:15:05.753 -> Average Cal sensor Temp.: 103.59 Ambient Temp.: 97.39 F
19:15:07.305 -> Average Cal sensor Temp.: 103.62 Ambient Temp.: 97.39 F
19:15:08.857 -> Average Cal sensor Temp.: 102.96 Ambient Temp.: 97.39 F
19:15:10.453 -> Average Cal sensor Temp.: 103.48 Ambient Temp.: 97.39 F
19:15:12.006 -> Average Cal sensor Temp.: 103.92 Ambient Temp.: 97.36 F
19:15:13.561 -> Average Cal sensor Temp.: 103.79 Ambient Temp.: 97.39 F
19:15:15.142 -> Average Cal sensor Temp.: 103.90 Ambient Temp.: 97.36 F
19:15:16.700 -> Average Cal sensor Temp.: 103.88 Ambient Temp.: 97.39 F
19:15:18.253 -> Average Cal sensor Temp.: 103.73 Ambient Temp.: 97.39 F
19:15:19.833 -> Average Cal sensor Temp.: 103.66 Ambient Temp.: 97.39 F
19:15:21.386 -> Average Cal sensor Temp.: 104.05 Ambient Temp.: 97.39 F
19:15:22.968 -> Average Cal sensor Temp.: 104.01 Ambient Temp.: 97.39 F
19:15:24.520 -> Average Cal sensor Temp.: 104.13 Ambient Temp.: 97.36 F
19:15:26.069 -> Average Cal sensor Temp.: 104.40 Ambient Temp.: 97.43 F
19:15:27.659 -> Average Cal sensor Temp.: 104.24 Ambient Temp.: 97.47 F
19:15:29.215 -> Average Cal sensor Temp.: 104.13 Ambient Temp.: 97.47 F
19:15:30.770 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 97.43 F
19:15:32.350 -> Average Cal sensor Temp.: 104.15 Ambient Temp.: 97.47 F
19:15:33.903 -> Average Cal sensor Temp.: 104.06 Ambient Temp.: 97.47 F
19:15:35.461 -> Average Cal sensor Temp.: 104.29 Ambient Temp.: 97.47 F
19:15:37.019 -> Average Cal sensor Temp.: 103.80 Ambient Temp.: 97.47 F
19:15:38.606 -> Average Cal sensor Temp.: 104.26 Ambient Temp.: 97.50 F
19:15:40.159 -> Average Cal sensor Temp.: 103.81 Ambient Temp.: 97.47 F
19:15:41.711 -> Average Cal sensor Temp.: 104.08 Ambient Temp.: 97.50 F
19:15:43.300 -> Average Cal sensor Temp.: 103.70 Ambient Temp.: 97.50 F
19:15:44.854 -> Average Cal sensor Temp.: 103.44 Ambient Temp.: 97.50 F
19:15:46.407 -> Average Cal sensor Temp.: 103.75 Ambient Temp.: 97.47 F
19:15:47.991 -> Average Cal sensor Temp.: 103.78 Ambient Temp.: 97.50 F
19:15:49.545 -> Average Cal sensor Temp.: 103.60 Ambient Temp.: 97.47 F
19:15:51.099 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 97.50 F
19:15:52.685 -> Average Cal sensor Temp.: 103.75 Ambient Temp.: 97.47 F
19:15:54.235 -> Average Cal sensor Temp.: 105.30 Ambient Temp.: 97.47 F
19:15:55.787 -> Average Cal sensor Temp.: 104.86 Ambient Temp.: 97.47 F
19:15:57.377 -> Average Cal sensor Temp.: 104.47 Ambient Temp.: 97.47 F
19:15:58.930 -> Average Cal sensor Temp.: 104.12 Ambient Temp.: 97.50 F
19:16:00.513 -> Average Cal sensor Temp.: 103.61 Ambient Temp.: 97.47 F
19:16:02.067 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 97.43 F
19:16:03.619 -> Average Cal sensor Temp.: 103.98 Ambient Temp.: 97.50 F
19:16:05.173 -> Average Cal sensor Temp.: 103.70 Ambient Temp.: 97.43 F
19:16:06.771 -> Average Cal sensor Temp.: 103.12 Ambient Temp.: 97.43 F
19:16:08.321 -> Average Cal sensor Temp.: 103.44 Ambient Temp.: 97.39 F
19:16:09.876 -> Average Cal sensor Temp.: 103.33 Ambient Temp.: 97.39 F
19:16:11.427 -> Average Cal sensor Temp.: 103.59 Ambient Temp.: 97.36 F
19:16:13.013 -> Average Cal sensor Temp.: 103.28 Ambient Temp.: 97.36 F
19:16:14.567 -> Average Cal sensor Temp.: 103.22 Ambient Temp.: 97.39 F
19:16:16.126 -> Average Cal sensor Temp.: 103.33 Ambient Temp.: 97.36 F
19:16:17.709 -> Average Cal sensor Temp.: 103.39 Ambient Temp.: 97.32 F
19:16:19.262 -> Average Cal sensor Temp.: 103.47 Ambient Temp.: 97.36 F
19:16:20.819 -> Average Cal sensor Temp.: 103.38 Ambient Temp.: 97.29 F
19:16:22.403 -> Average Cal sensor Temp.: 103.36 Ambient Temp.: 97.32 F
19:16:23.960 -> Average Cal sensor Temp.: 103.29 Ambient Temp.: 97.32 F
19:16:25.511 -> Average Cal sensor Temp.: 103.58 Ambient Temp.: 97.32 F
19:16:27.066 -> Average Cal sensor Temp.: 103.82 Ambient Temp.: 97.32 F
19:16:28.655 -> Average Cal sensor Temp.: 103.70 Ambient Temp.: 97.32 F
19:16:30.206 -> Average Cal sensor Temp.: 103.57 Ambient Temp.: 97.32 F
19:16:31.758 -> Average Cal sensor Temp.: 103.64 Ambient Temp.: 97.32 F
19:16:33.350 -> Average Cal sensor Temp.: 103.46 Ambient Temp.: 97.29 F
19:16:34.907 -> Average Cal sensor Temp.: 103.90 Ambient Temp.: 97.32 F
19:16:36.464 -> Average Cal sensor Temp.: 104.43 Ambient Temp.: 97.32 F
19:16:38.046 -> Average Cal sensor Temp.: 104.42 Ambient Temp.: 97.36 F
19:16:39.592 -> Average Cal sensor Temp.: 103.57 Ambient Temp.: 97.29 F
19:16:41.150 -> Average Cal sensor Temp.: 103.93 Ambient Temp.: 97.36 F
19:16:42.708 -> Average Cal sensor Temp.: 103.95 Ambient Temp.: 97.32 F
19:16:44.298 -> Average Cal sensor Temp.: 104.19 Ambient Temp.: 97.32 F
19:16:45.854 -> Average Cal sensor Temp.: 103.51 Ambient Temp.: 97.21 F
19:16:47.403 -> Average Cal sensor Temp.: 104.06 Ambient Temp.: 97.32 F
19:16:48.960 -> Average Cal sensor Temp.: 104.24 Ambient Temp.: 97.29 F
19:16:50.551 -> Average Cal sensor Temp.: 104.18 Ambient Temp.: 97.32 F
19:16:52.107 -> Average Cal sensor Temp.: 103.52 Ambient Temp.: 97.32 F
19:16:53.659 -> Average Cal sensor Temp.: 103.23 Ambient Temp.: 97.36 F
19:16:55.243 -> Average Cal sensor Temp.: 103.67 Ambient Temp.: 97.29 F
19:16:56.804 -> Average Cal sensor Temp.: 103.45 Ambient Temp.: 97.32 F
19:16:58.357 -> Average Cal sensor Temp.: 103.31 Ambient Temp.: 97.29 F
19:16:59.907 -> Average Cal sensor Temp.: 103.35 Ambient Temp.: 97.29 F
19:17:01.498 -> Average Cal sensor Temp.: 103.46 Ambient Temp.: 97.32 F
19:17:03.049 -> Average Cal sensor Temp.: 104.04 Ambient Temp.: 97.32 F
19:17:04.631 -> Average Cal sensor Temp.: 104.07 Ambient Temp.: 97.29 F
19:17:06.189 -> Average Cal sensor Temp.: 103.68 Ambient Temp.: 97.29 F
19:17:07.745 -> Average Cal sensor Temp.: 103.84 Ambient Temp.: 97.21 F
19:17:09.301 -> Average Cal sensor Temp.: 104.33 Ambient Temp.: 97.21 F
19:17:10.886 -> Average Cal sensor Temp.: 104.25 Ambient Temp.: 97.29 F
19:17:12.443 -> Average Cal sensor Temp.: 103.85 Ambient Temp.: 97.29 F
19:17:13.999 -> Average Cal sensor Temp.: 104.14 Ambient Temp.: 97.21 F
19:17:15.551 -> Average Cal sensor Temp.: 103.62 Ambient Temp.: 97.21 F
19:17:17.139 -> Average Cal sensor Temp.: 103.80 Ambient Temp.: 97.21 F
19:17:18.694 -> Average Cal sensor Temp.: 103.55 Ambient Temp.: 97.29 F
19:17:20.245 -> Average Cal sensor Temp.: 103.68 Ambient Temp.: 97.18 F
19:17:21.829 -> Average Cal sensor Temp.: 103.92 Ambient Temp.: 97.21 F
19:17:23.384 -> Average Cal sensor Temp.: 103.97 Ambient Temp.: 97.29 F
19:17:24.934 -> Average Cal sensor Temp.: 103.90 Ambient Temp.: 97.29 F
19:17:26.527 -> Average Cal sensor Temp.: 104.31 Ambient Temp.: 97.32 F
19:17:28.072 -> Average Cal sensor Temp.: 104.36 Ambient Temp.: 97.29 F
19:17:29.652 -> Average Cal sensor Temp.: 104.07 Ambient Temp.: 97.21 F
19:17:31.199 -> Average Cal sensor Temp.: 104.22 Ambient Temp.: 97.32 F
19:17:32.753 -> Average Cal sensor Temp.: 104.25 Ambient Temp.: 97.29 F
19:17:34.336 -> Average Cal sensor Temp.: 104.14 Ambient Temp.: 97.32 F
19:17:35.893 -> Average Cal sensor Temp.: 103.84 Ambient Temp.: 97.29 F
19:17:37.444 -> Average Cal sensor Temp.: 103.99 Ambient Temp.: 97.29 F
19:17:39.026 -> Average Cal sensor Temp.: 103.83 Ambient Temp.: 97.29 F
19:17:40.584 -> Average Cal sensor Temp.: 104.08 Ambient Temp.: 97.32 F
19:17:42.137 -> Average Cal sensor Temp.: 104.42 Ambient Temp.: 97.32 F
19:17:43.728 -> Average Cal sensor Temp.: 104.04 Ambient Temp.: 97.32 F
19:17:45.280 -> Average Cal sensor Temp.: 103.78 Ambient Temp.: 97.29 F
19:17:46.832 -> Average Cal sensor Temp.: 103.72 Ambient Temp.: 97.29 F
19:17:48.425 -> Average Cal sensor Temp.: 104.02 Ambient Temp.: 97.29 F
19:17:49.975 -> Average Cal sensor Temp.: 104.26 Ambient Temp.: 97.21 F
19:17:51.524 -> Average Cal sensor Temp.: 103.79 Ambient Temp.: 97.29 F
19:17:53.113 -> Average Cal sensor Temp.: 104.04 Ambient Temp.: 97.21 F
19:17:54.667 -> Average Cal sensor Temp.: 104.31 Ambient Temp.: 97.29 F
19:17:56.222 -> Average Cal sensor Temp.: 104.27 Ambient Temp.: 97.32 F
19:17:57.804 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 97.29 F
19:17:59.355 -> Average Cal sensor Temp.: 104.21 Ambient Temp.: 97.36 F
19:18:00.936 -> Average Cal sensor Temp.: 104.12 Ambient Temp.: 97.32 F
19:18:02.484 -> Average Cal sensor Temp.: 104.32 Ambient Temp.: 97.32 F
19:18:04.036 -> Average Cal sensor Temp.: 104.48 Ambient Temp.: 97.29 F
19:18:05.624 -> Average Cal sensor Temp.: 104.26 Ambient Temp.: 97.29 F
19:18:07.174 -> Average Cal sensor Temp.: 104.40 Ambient Temp.: 97.39 F
19:18:08.731 -> Average Cal sensor Temp.: 104.38 Ambient Temp.: 97.36 F
19:18:10.323 -> Average Cal sensor Temp.: 104.14 Ambient Temp.: 97.39 F
19:18:11.874 -> Average Cal sensor Temp.: 104.09 Ambient Temp.: 97.32 F
19:18:13.426 -> Average Cal sensor Temp.: 104.47 Ambient Temp.: 97.39 F
19:18:15.017 -> Average Cal sensor Temp.: 104.03 Ambient Temp.: 97.29 F
19:18:16.575 -> Average Cal sensor Temp.: 103.61 Ambient Temp.: 97.36 F
19:18:18.130 -> Average Cal sensor Temp.: 103.87 Ambient Temp.: 97.43 F
19:18:19.681 -> Average Cal sensor Temp.: 103.93 Ambient Temp.: 97.32 F
19:18:21.272 -> Average Cal sensor Temp.: 103.94 Ambient Temp.: 97.39 F
19:18:22.822 -> Average Cal sensor Temp.: 104.18 Ambient Temp.: 97.39 F
19:18:24.376 -> Average Cal sensor Temp.: 104.50 Ambient Temp.: 97.39 F
19:18:25.964 -> Average Cal sensor Temp.: 104.36 Ambient Temp.: 97.36 F
19:18:27.516 -> Average Cal sensor Temp.: 104.45 Ambient Temp.: 97.39 F
19:18:29.071 -> Average Cal sensor Temp.: 104.59 Ambient Temp.: 97.39 F
19:18:30.660 -> Average Cal sensor Temp.: 104.51 Ambient Temp.: 97.43 F
19:18:32.208 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 97.39 F
19:18:33.791 -> Average Cal sensor Temp.: 103.88 Ambient Temp.: 97.39 F
19:18:35.345 -> Average Cal sensor Temp.: 104.53 Ambient Temp.: 97.43 F
19:18:36.902 -> Average Cal sensor Temp.: 104.41 Ambient Temp.: 97.36 F
19:18:38.487 -> Average Cal sensor Temp.: 103.92 Ambient Temp.: 97.39 F
19:18:40.035 -> Average Cal sensor Temp.: 103.71 Ambient Temp.: 97.36 F
19:18:41.586 -> Average Cal sensor Temp.: 103.62 Ambient Temp.: 97.39 F
19:18:43.176 -> Average Cal sensor Temp.: 103.23 Ambient Temp.: 97.39 F
19:18:44.728 -> Average Cal sensor Temp.: 104.32 Ambient Temp.: 97.36 F
19:18:46.291 -> Average Cal sensor Temp.: 104.43 Ambient Temp.: 97.32 F
19:18:47.845 -> Average Cal sensor Temp.: 104.40 Ambient Temp.: 97.36 F
19:18:49.432 -> Average Cal sensor Temp.: 104.45 Ambient Temp.: 97.36 F
19:18:50.984 -> Average Cal sensor Temp.: 104.43 Ambient Temp.: 97.36 F
19:18:52.543 -> Average Cal sensor Temp.: 104.26 Ambient Temp.: 97.39 F
19:18:54.101 -> Average Cal sensor Temp.: 104.40 Ambient Temp.: 97.39 F
19:18:55.690 -> Average Cal sensor Temp.: 104.26 Ambient Temp.: 97.39 F
19:18:57.251 -> Average Cal sensor Temp.: 104.44 Ambient Temp.: 97.36 F
19:18:58.799 -> Average Cal sensor Temp.: 103.77 Ambient Temp.: 97.36 F
19:19:00.387 -> Average Cal sensor Temp.: 103.95 Ambient Temp.: 97.36 F
19:19:01.940 -> Average Cal sensor Temp.: 104.27 Ambient Temp.: 97.32 F
19:19:03.495 -> Average Cal sensor Temp.: 103.53 Ambient Temp.: 97.32 F
19:19:05.049 -> Average Cal sensor Temp.: 103.86 Ambient Temp.: 97.36 F
19:19:06.636 -> Average Cal sensor Temp.: 104.07 Ambient Temp.: 97.36 F
19:19:08.190 -> Average Cal sensor Temp.: 104.33 Ambient Temp.: 97.36 F
19:19:09.742 -> Average Cal sensor Temp.: 104.37 Ambient Temp.: 97.39 F
19:19:11.327 -> Average Cal sensor Temp.: 104.56 Ambient Temp.: 97.36 F
19:19:12.886 -> Average Cal sensor Temp.: 104.36 Ambient Temp.: 97.36 F
19:19:14.438 -> Average Cal sensor Temp.: 104.36 Ambient Temp.: 97.36 F
19:19:16.021 -> Average Cal sensor Temp.: 104.35 Ambient Temp.: 97.39 F
19:19:17.570 -> Average Cal sensor Temp.: 104.27 Ambient Temp.: 97.39 F
19:19:19.159 -> Average Cal sensor Temp.: 104.55 Ambient Temp.: 97.39 F
19:19:20.714 -> Average Cal sensor Temp.: 104.89 Ambient Temp.: 97.36 F
19:19:22.267 -> Average Cal sensor Temp.: 104.31 Ambient Temp.: 97.36 F
19:19:23.853 -> Average Cal sensor Temp.: 104.00 Ambient Temp.: 97.43 F
19:19:25.410 -> Average Cal sensor Temp.: 103.70 Ambient Temp.: 97.36 F
19:19:26.963 -> Average Cal sensor Temp.: 103.96 Ambient Temp.: 97.32 F
19:19:28.515 -> Average Cal sensor Temp.: 104.16 Ambient Temp.: 97.39 F
19:19:30.101 -> Average Cal sensor Temp.: 104.24 Ambient Temp.: 97.36 F
19:19:31.655 -> Average Cal sensor Temp.: 104.54 Ambient Temp.: 97.39 F
19:19:33.209 -> Average Cal sensor Temp.: 104.53 Ambient Temp.: 97.36 F
19:19:34.798 -> Average Cal sensor Temp.: 104.26 Ambient Temp.: 97.39 F
19:19:36.351 -> Average Cal sensor Temp.: 104.32 Ambient Temp.: 97.39 F
19:19:37.901 -> Average Cal sensor Temp.: 104.03 Ambient Temp.: 97.39 F
19:19:39.482 -> Average Cal sensor Temp.: 104.01 Ambient Temp.: 97.39 F
19:19:41.040 -> Average Cal sensor Temp.: 103.82 Ambient Temp.: 97.39 F
19:19:42.625 -> Average Cal sensor Temp.: 104.13 Ambient Temp.: 97.36 F
19:19:44.183 -> Average Cal sensor Temp.: 104.23 Ambient Temp.: 97.39 F
19:19:45.737 -> Average Cal sensor Temp.: 104.08 Ambient Temp.: 97.32 F
19:19:47.286 -> Average Cal sensor Temp.: 103.91 Ambient Temp.: 97.36 F
19:19:48.875 -> Average Cal sensor Temp.: 104.36 Ambient Temp.: 97.36 F